## Sheet1
Event Date | Event Type | Manufacturer | Date Received | Brand Name | Device Problem | IMDRF Code | Patient Problem | Event Text
16-02-2018 | Malfunction | ethicon endo-surgery | 02-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A BILATERAL LAPAROSCOPIC OOPHORECTOMY THE FIRST STAPLER, WITH A GRAY RELOAD, WOULD STAPLE BUT WOULDN'T CUT. A SECOND LIKE RELOAD WAS USED WITH THE SAME STAPLER AND WOULD STAPLE BUT NOT CUT. A SECOND STAPLER WAS TRIED WITH A WHITE RELOAD AND THE SECOND STAPLER STAPLED BUT DIDN'T CUT. A SECOND LIKE WHITE RELOAD WAS TRIED WITH THE SECOND STAPLER AND THE SAME ISSUE OF STAPLING BUT NOT CUTTING OCCURRED. THE PROCEDURE WAS COMPLETED USING ENDO SHEAR SCISSORS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-02-2018 | Malfunction | ethicon endo-surgery | 02-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A BILATERAL LAPAROSCOPIC OOPHERECTOMY THE FIRST STAPLER, WITH A GRAY RELOAD, WOULD STAPLE BUT WOULDN'T CUT. A SECOND LIKE RELOAD WAS USED WITH THE SAME STAPLER AND WOULD STAPLE BUT NOT CUT. A SECOND STAPLER WAS TRIED WITH A WHITE RELOAD AND THE SECOND STAPLER STAPLED BUT DIDN'T CUT. A SECOND LIKE WHITE RELOAD WAS TRIED WITH THE SECOND STAPLER AND THE SAME ISSUE OF STAPLING BUT NOT CUTTING OCCURRED. THE PROCEDURE WAS COMPLETED USING ENDO SHEAR SCISSORS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
05-03-2018 | Injury | ethicon endo-surgery | 02-04-2018 | ECHELON ENDOPATH STAPLER | Device Operates Differently Than Expected |  | Fistula | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE: TITLE: LAPAROSCOPIC CENTRAL PANCREATECTOMY FOR SOLID PSEUDOPAPILLARY TUMORS OF THE PANCREAS: OUR EXPERIENCE WITH TEN CASES AUTHOR(S): XUE-MIN CHEN, YUE ZHANG AND DONG-LIN SUN CITATION: WORLD JOURNAL OF SURGICAL ONCOLOGY 2014, 12:312; DOI: 10.1186/1477-7819-12-312 THIS RETROSPECTIVE COHORT STUDY AIMED TO EVALUATE THE FEASIBILITY, SAFETY AND LONG-TERM OUTCOME OF LAPAROSCOPIC CENTRAL PANCREATECTOMIES (LCPS) BASED ON A SERIES OF SOLID PSEUDOPAPILLARY TUMORS (SPTS). BETWEEN FEB2009 AND DEC2013, 10 PATIENTS (N=2 MALE, N=8 FEMALE) WITH SPT OF THE PANCREAS WAS TREATED WITH LCP. AFTER COMPLETING THE PANCREATIC RESECTION, AN END-TO-SIDE PANCREATOJEJUNOSTOMY WAS PERFORMED. 3-0 VICRYL SUTURES WERE PLACED BETWEEN THE JEJUNAL SEROSA AND THE POSTERIOR SIDE OF THE PANCREATIC CAPSULE FOR APPOSITION. THE JEJUNUM WAS OPENED WITH A HARMONIC SCALPEL, SUITABLE FOR A DISTAL PANCREATIC STUMP. THE POSTERIOR AND ANTERIOR LAYERS BETWEEN THE PANCREAS (PARENCHYMA AND CAPSULE) AND THE FULL THICKNESS OF THE JEJUNUM WAS PERFORMED WITH CONTINUOUS OR INTERRUPTED 3-0 VICRYL SUTURES. A SIDE TO-SIDE JEJUNOJEJUNOSTOMY WAS PERFORMED WITH AN ENDOCUTTER 60 STAPLER. TWO PATIENTS ((B)(6) FEMALE AND (B)(6) MALE) EXPERIENCED A POSTOPERATIVE PANCREATIC FISTULA (GRADE A), WHICH WAS MANAGED CONSERVATIVELY AND ULTIMATELY CURED. THE END-TO-SIDE PANCREATICOJEJUNAL WAS USED TO ACCOMPLISH THE PANCREATICOJEJUNAL RECONSTRUCTION AND WE HAD A PANCREATIC FISTULA RATE OF 20%. LCP IS A SAFE AND EFFECTIVE TECHNIQUE FOR RESECTING SPT OF THE NECK AND PROXIMAL BODY OF THE PANCREAS WHILE PRESERVING PANCREATIC ENDOCRINE AND EXOCRINE FUNCTION, AND THE SPLEEN.
05-03-2018 | Injury | ethicon endo-surgery | 02-04-2018 | ECHELON ENDOPATH STAPLER | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE: TITLE: LAPAROSCOPIC CENTRAL PANCREATECTOMY FOR SOLID PSEUDOPAPILLARY TUMORS OF THE PANCREAS: OUR EXPERIENCE WITH TEN CASES AUTHOR(S): XUE-MIN CHEN, YUE ZHANG AND DONG-LIN SUN CITATION: WORLD JOURNAL OF SURGICAL ONCOLOGY 2014, 12:312; DOI: 10.1186/1477-7819-12-312 THIS RETROSPECTIVE COHORT STUDY AIMED TO EVALUATE THE FEASIBILITY, SAFETY AND LONG-TERM OUTCOME OF LAPAROSCOPIC CENTRAL PANCREATECTOMIES (LCPS) BASED ON A SERIES OF SOLID PSEUDOPAPILLARY TUMORS (SPTS). BETWEEN FEB2009 AND DEC2013, 10 PATIENTS (N=2 MALE, N=8 FEMALE) WITH SPT OF THE PANCREAS WAS TREATED WITH LCP. AFTER COMPLETING THE PANCREATIC RESECTION, AN END-TO-SIDE PANCREATOJEJUNOSTOMY WAS PERFORMED. 3-0 VICRYL SUTURES WERE PLACED BETWEEN THE JEJUNAL SEROSA AND THE POSTERIOR SIDE OF THE PANCREATIC CAPSULE FOR APPOSITION. THE JEJUNUM WAS OPENED WITH A HARMONIC SCALPEL, SUITABLE FOR A DISTAL PANCREATIC STUMP. THE POSTERIOR AND ANTERIOR LAYERS BETWEEN THE PANCREAS (PARENCHYMA AND CAPSULE) AND THE FULL THICKNESS OF THE JEJUNUM WAS PERFORMED WITH CONTINUOUS OR INTERRUPTED 3-0 VICRYL SUTURES. A SIDE TO-SIDE JEJUNOJEJUNOSTOMY WAS PERFORMED WITH AN ENDOCUTTER 60 STAPLER. TWO PATIENTS ((B)(6) FEMALE AND (B)(6) MALE) EXPERIENCED A POSTOPERATIVE PANCREATIC FISTULA (GRADE A), WHICH WAS MANAGED CONSERVATIVELY AND ULTIMATELY CURED. THE END-TO-SIDE PANCREATICOJEJUNAL WAS USED TO ACCOMPLISH THE PANCREATICOJEJUNAL RECONSTRUCTION AND WE HAD A PANCREATIC FISTULA RATE OF 20%. LCP IS A SAFE AND EFFECTIVE TECHNIQUE FOR RESECTING SPT OF THE NECK AND PROXIMAL BODY OF THE PANCREAS WHILE PRESERVING PANCREATIC ENDOCRINE AND EXOCRINE FUNCTION, AND THE SPLEEN.
01-01-2018 | Malfunction | ethicon endo-surgery | 02-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AFTER FIRED COMPLETED, HCP FOUND SOME STAPLE'S B-SHAPE WAS NOT MADE WELL AND SOME STAPLES WERE OPENED IT SHAPE.
05-03-2018 | Injury | ethicon endo-surgery | 02-04-2018 | ECHELON ENDOPATH STAPLER | Insufficient Information | E2401 | Fistula | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE: TITLE: LAPAROSCOPIC CENTRAL PANCREATECTOMY FOR SOLID PSEUDOPAPILLARY TUMORS OF THE PANCREAS: OUR EXPERIENCE WITH TEN CASES AUTHOR(S): XUE-MIN CHEN, YUE ZHANG AND DONG-LIN SUN CITATION: WORLD JOURNAL OF SURGICAL ONCOLOGY 2014, 12:312; DOI: 10.1186/1477-7819-12-312. THIS RETROSPECTIVE COHORT STUDY AIMED TO EVALUATE THE FEASIBILITY, SAFETY AND LONG-TERM OUTCOME OF LAPAROSCOPIC CENTRAL PANCREATECTOMIES (LCPS) BASED ON A SERIES OF SOLID PSEUDOPAPILLARY TUMORS (SPTS). BETWEEN (B)(6) 2009 AND (B)(6) 2013, 10 PATIENTS (N=2 MALE, N=8 FEMALE) WITH SPT OF THE PANCREAS WAS TREATED WITH LCP. AFTER COMPLETING THE PANCREATIC RESECTION, AN END-TO-SIDE PANCREATOJEJUNOSTOMY WAS PERFORMED. 3-0 VICRYL SUTURES WERE PLACED BETWEEN THE JEJUNAL SEROSA AND THE POSTERIOR SIDE OF THE PANCREATIC CAPSULE FOR APPOSITION. THE JEJUNUM WAS OPENED WITH A HARMONIC SCALPEL, SUITABLE FOR A DISTAL PANCREATIC STUMP. THE POSTERIOR AND ANTERIOR LAYERS BETWEEN THE PANCREAS (PARENCHYMA AND CAPSULE) AND THE FULL THICKNESS OF THE JEJUNUM WAS PERFORMED WITH CONTINUOUS OR INTERRUPTED 3-0 VICRYL SUTURES. A SIDETO-SIDE JEJUNOJEJUNOSTOMY WAS PERFORMED WITH AN ENDOCUTTER 60 STAPLER. TWO PATIENTS ((B)(6) FEMALE AND (B)(6) MALE) EXPERIENCED A POSTOPERATIVE PANCREATIC FISTULA (GRADE A), WHICH WAS MANAGED CONSERVATIVELY AND ULTIMATELY CURED. THE END-TO-SIDE PANCREATICOJEJUNAL WAS USED TO ACCOMPLISH THE PANCREATICOJEJUNAL RECONSTRUCTION AND WE HAD A PANCREATIC FISTULA RATE OF 20%. LCP IS A SAFE AND EFFECTIVE TECHNIQUE FOR RESECTING SPT OF THE NECK AND PROXIMAL BODY OF THE PANCREAS WHILE PRESERVING PANCREATIC ENDOCRINE AND EXOCRINE FUNCTION, AND THE SPLEEN.
06-01-2018 | Malfunction | ethicon endo-surgery | 02-04-2018 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THERE WAS AN ACCIDENTAL LOCK-OUT. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
08-01-2018 | Injury | ethicon endo-surgery | 02-04-2018 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Code Available | nan
03-07-2018 | Malfunction | covidien lp llc north haven | 02-04-2018 | EEA ORVIL | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY WHILE BEING APPLIED AT THE ANASTOMOSIS, THE SUTURE BROKE AND WAS DIFFICULT/UNABLE TO OPEN. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 02-04-2018 | ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, THE RELOAD MISFIRED.
21-03-2018 | Malfunction | covidien mfg dc boulder | 02-04-2018 | LIGASURE | Therapy Delivered to Incorrect Body Area | A090407 | Burn, Thermal | Event Description: ACCORDING TO THE REPORTER, DURING GLOSSECTOMY, THE SURGEON HAS ACTIVATED THE DEVICE INSIDE THE ORAL CAVITY AND HE BASED THE DEVICE ON THE PATIENT¿S LIPS. THE METAL RECTANGLE LOCATED ON THE JAWS HEATED UP AND CAUSED BURNS. THE OPERATOR REALIZED WHAT HAPPENED AFTER RESECTING THE TONGUE. THE SCRUB NURSE CHECKED THE DEVICE ON THE WET GAUZE PAD AND SHE FELT THE HEAT COMING FROM THE METAL RECTANGLE OF THE DEVICE. THE DEVICE WAS ACTIVATED SEVERAL DOZEN TIMES; THE SURFACE THEY WERE WORKING ON WAS WET. MONOPOLAR ENERGY WAS ALSO USED BUT INTERCHANGEABLY. AFTER THIS INCIDENT, THEY CONTINUED USING THE DEVICE FOR LYMPHADENECTOMY WITHOUT ANY PROBLEMS BUT NOT TOUCHING ANYTHING WITH THIS METAL RECTANGLE.
02-04-2018 | Malfunction | covidien lp llc north haven | 02-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, THE RELOAD MISFIRED.
07-03-2018 | Malfunction | covidien lp llc north haven | 02-04-2018 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER,DURING THE FIRST FIRING OF THE DEVICE WHILE PERFORMING GASTRIC TUBE CONSTRUCTION IN A THORACO/ESOPHAGUS PROCEDURE, THE STAPLING STOPPED HALFWAY AND THE DEVICE COULD NOT BE FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | us surgical puerto rico | 02-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ROBOTIC LAPAROSCOPIC NEPHRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED OFF KIDNEY GATE, THE STAPLER WAS ABLE TO FIRE BUT THE STAPLES DID NOT FORM "B" SHAPE. INSTEAD, THE STAPLE LEGS WERE STRAIGHT UPRIGHT ON THE RELOAD. THE BROKEN BLOOD VESSEL WAS CLIPPED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-02-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ENDOPATH | Insufficient Information | E2401 | Tissue Damage | Event Description: DURING USE OF THE ARTICULATING ENDO STAPLER A VASCULAR TEAR OCCURRED. STAPLER REMOVED FROM FIELD. STAPLER AND STAPLER LOAD TAGGED AND LABELED.
23-02-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PNEUMONECTOMY, THE JAWS DID NOT OPEN WHEN THE DEVICE WAS INSERTED INTO THE CHEST CAVITY BEFORE THE 4TH FIRING. THE DEVICE WAS USED ON THE LUNG. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT SINCE THE EVENT OCCURRED BEFORE THE TARGET TISSUE WAS CLAMPED.
06-02-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LADG, THE CARTRIDGE COULD NOT BE LOADED INTO THE JAW AT THE 4TH FIRING. THE DEVICE WAS USED ON THE STOMACH. WHEN THE SALES REP CHECKED THE DEVICE, THE KNIFE WAS NOT AT THE HOME POSITION BUT THE JAWS WERE OPEN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
06-02-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Malposition of Device | A150202 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LADG, THE CARTRIDGE COULD NOT BE LOADED INTO THE JAW AT THE 4TH FIRING. THE DEVICE WAS USED ON THE STOMACH. WHEN THE SALES REP CHECKED THE DEVICE, THE KNIFE WAS NOT AT THE HOME POSITION BUT THE JAWS WERE OPEN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-01-2018 | Injury | ethicon endo-surgery | 03-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Device Operates Differently Than Expected |  | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN OPEN PRIMARY GASTRIC BYPASS. THE STOMACH WAS EXTREMELY FIBROTIC AND THICK SO BLACK RELOADS WITH SEAMGUARD WERE USED FOR ALL FIRINGS. SIXTEEN (16) HOURS LATER THERE WAS GASTRIC FLUID NOTED IN THE DRAINS. PATIENT TAKEN BACK TO THEATRE AND THE SURGEON SAW AN OPENING IN THE STAPLE LINE AROUND THE TISSUE WHERE THE 3RD FIRING WOULD HAVE TAKEN PLACE. HE SUTURED THIS TISSUE BACK TOGETHER AND PATIENT WAS SENT TO ICU FOR CLOSE MONITORING FOR 24 HRS.
01-01-2018 | Injury | ethicon endo-surgery | 03-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN OPEN PRIMARY GASTRIC BYPASS. THE STOMACH WAS EXTREMELY FIBROTIC AND THICK SO BLACK RELOADS WITH SEAMGUARD WERE USED FOR ALL FIRINGS. SIXTEEN (16) HOURS LATER THERE WAS GASTRIC FLUID NOTED IN THE DRAINS. PATIENT TAKEN BACK TO THEATRE AND THE SURGEON SAW AN OPENING IN THE STAPLE LINE AROUND THE TISSUE WHERE THE 3RD FIRING WOULD HAVE TAKEN PLACE. HE SUTURED THIS TISSUE BACK TOGETHER AND PATIENT WAS SENT TO ICU FOR CLOSE MONITORING FOR 24 HRS.
01-03-2018 | Injury | ethicon endo-surgery | 03-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Device Inoperable |  | No Code Available | Event Description: IT WAS REPORTED THAT THE BATTERY OF THE DEVICE DID NOT WORK THEREFORE A MANUAL DEVICE WAS USED. A SLEEVE PROCEDURE WAS REPLACED BY A BYPASS PROCEDURE DUE TO THE CHANGES OF STAPLERS AND DUE TO THE CARTRIDGE WHICH DID NOT WORK ON THE MANUAL DEVICE ( THE CARTRIDGE USED WAS ASSOCIATED TO THE AUTOMATIC STAPLER). THE PROBE WAS RAISED AND THE STOMACH WAS CUT TOO MUCH NEAR. TO AVOID THE STENOSIS, A BY PASS PROCEDURE WAS PERFORMED.
05-03-2018 | Malfunction | justright surgical | 03-04-2018 | JUSTRIGHT 5MM STAPLER | Improper or Incorrect Procedure or Method | A2303 | No Consequences Or Impact To Patient | Event Description: DURING A THORACOSCOPIC LEFT LOWER LOBECTOMY, THE STAPLER MISFIRED AND STAPLES RELEASED INTO THE BODY CAVITY. A SECOND STAPLER WAS USED, AND IT WORKED PERFECTLY. PATIENT IS FINE AND WAS DISCHARGED. BOTH STAPLERS AND RELOADS WERE RETURNED FOR EVALUATION.
21-02-2018 | Malfunction | covidien lp llc north haven | 03-04-2018 | EGIA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE STAPLING, A DEVICE COMPONENT BROKE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINS. PATIENT IS HOME.
05-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOID COLECTOMY WHEN THE DEVICE WAS FIRST FIRED THE STAPLES ON THE POSTERIOR SIDE OF THE ANASTOMOSES DID NOT FORM. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCES REPORTED. THE LEAK TEST WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
06-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | PROXIMATE*RL LINEAR STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A RECTAL ANASTOMOSIS PROCEDURE, THE CLIP DID NOT HOLD. THE SUTURING WAS PERFORMED MANUALLY. NO PATIENT CONSEQUENCES REPORTED.
06-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER HAS CUT BUT NOT STAPLED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
08-03-2018 | Malfunction | davol inc sub c.r bard | 03-04-2018 | OPTIFIX 30 CT STRAIGHT 39CM | Bent |  | No Consequences Or Impact To Patient | Event Description: INFORMATION AS REPORTED BY THE USER FACILITY: DURING A RIGHT LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE SURGEON WAS USING THE OPTIFIX FIXATION DEVICE AND THE "WHOLE TIP BROKE OFF." THE SUBJECT PRODUCT WAS RETURNED WITH A DETACHED BARREL INSERT COMPONENT. PHOTOS PROVIDED SHOW THE BARREL INSERT DETACHED INSIDE OF THE PATIENT. THE INSERT WAS REMOVED AND RETURNED WITH THE SUBJECT PRODUCT. NO PATIENT INJURY WAS REPORTED.
08-03-2018 | Malfunction | davol inc sub c.r bard | 03-04-2018 | OPTIFIX 30 CT STRAIGHT 39CM | Break | A0401 | No Consequences Or Impact To Patient | Event Description: INFORMATION AS REPORTED BY THE USER FACILITY: DURING A RIGHT LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE SURGEON WAS USING THE OPTIFIX FIXATION DEVICE AND THE "WHOLE TIP BROKE OFF." THE SUBJECT PRODUCT WAS RETURNED WITH A DETACHED BARREL INSERT COMPONENT. PHOTOS PROVIDED SHOW THE BARREL INSERT DETACHED INSIDE OF THE PATIENT. THE INSERT WAS REMOVED AND RETURNED WITH THE SUBJECT PRODUCT. NO PATIENT INJURY WAS REPORTED.
08-03-2018 | Malfunction | davol inc sub c.r bard | 03-04-2018 | OPTIFIX 30 CT STRAIGHT 39CM | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: INFORMATION AS REPORTED BY THE USER FACILITY: DURING A RIGHT LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE SURGEON WAS USING THE OPTIFIX FIXATION DEVICE AND THE "WHOLE TIP BROKE OFF." THE SUBJECT PRODUCT WAS RETURNED WITH A DETACHED BARREL INSERT COMPONENT. PHOTOS PROVIDED SHOW THE BARREL INSERT DETACHED INSIDE OF THE PATIENT. THE INSERT WAS REMOVED AND RETURNED WITH THE SUBJECT PRODUCT. NO PATIENT INJURY WAS REPORTED.
06-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Device Inoperable |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP TOTAL NEPHRECTOMY THE SURGEON TOOK THE DEVICE AND USED IT FOR CUTTING THE VESSELS. HE START TO SHOOT AND THE DEVICE STOPPED HALF WAY. HE COULD SEE THAT HALF OF THE STAPLERS CAME OUT BUT THERE WAS NO OPTION TO CONTINUE SHOOTING. THE SURGEON TOOK A NEW AND SAME DEVICE AND IT HAPPENED AGAIN.AFTER ALL THAT THE SURGEON PUT CLIPS AND CONTINUE TO CUT WITH SCISSORS.
06-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP TOTAL NEPHRECTOMY THE SURGEON TOOK THE DEVICE AND USED IT FOR CUTTING THE VESSELS. HE START TO SHOOT AND THE DEVICE STOPPED HALF WAY. HE COULD SEE THAT HALF OF THE STAPLERS CAME OUT BUT THERE WAS NO OPTION TO CONTINUE SHOOTING. THE SURGEON TOOK A NEW AND SAME DEVICE AND IT HAPPENED AGAIN.AFTER ALL THAT THE SURGEON PUT CLIPS AND CONTINUE TO CUT WITH SCISSORS.
06-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Advance | A150204 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP TOTAL NEPHRECTOMY THE SURGEON TOOK THE DEVICE AND USED IT FOR CUTTING THE VESSELS. HE START TO SHOOT AND THE DEVICE STOPPED HALF WAY. HE COULD SEE THAT HALF OF THE STAPLERS CAME OUT BUT THERE WAS NO OPTION TO CONTINUE SHOOTING. THE SURGEON TOOK A NEW AND SAME DEVICE AND IT HAPPENED AGAIN. AFTER ALL THAT THE SURGEON PUT CLIPS AND CONTINUE TO CUT WITH SCISSORS.
06-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP TOTAL NEPHRECTOMY THE SURGEON TOOK THE DEVICE AND USED IT FOR CUTTING THE VESSELS. HE START TO SHOOT AND THE DEVICE STOPPED HALF WAY. HE COULD SEE THAT HALF OF THE STAPLERS CAME OUT BUT THERE WAS NO OPTION TO CONTINUE SHOOTING. THE SURGEON TOOK A NEW AND SAME DEVICE AND IT HAPPENED AGAIN. AFTER ALL THAT THE SURGEON PUT CLIPS AND CONTINUE TO CUT WITH SCISSORS.
06-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP TOTAL NEPHRECTOMY THE SURGEON TOOK THE DEVICE AND USED IT FOR CUTTING THE VESSELS. HE START TO SHOOT AND THE DEVICE STOPPED HALF WAY. HE COULD SEE THAT HALF OF THE STAPLERS CAME OUT BUT THERE WAS NO OPTION TO CONTINUE SHOOTING. THE SURGEON TOOK A NEW AND SAME DEVICE AND IT HAPPENED AGAIN. AFTER ALL THAT THE SURGEON PUT CLIPS AND CONTINUE TO CUT WITH SCISSORS.
06-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Migration or Expulsion of Device | A0104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE FIRST FIRING, WITH A BLACK RELOAD, SEAM GUARD BUTTRESSING MATERIAL, THE DOCTOR CLOSED ON TISSUE, WAITED FIFTEEN SECONDS, ATTEMPTED TO USE A PULSING TECHNIQUE ON THE FIRING TRIGGER AND THE KNIFE BLADE INSTANTLY MOVED TO THE DISTAL END OF THE STAPLER. AFTER FIVE SECONDS, THE KNIFE BLADE RETURNED. THE DOCTOR REMOVED THE STAPLER AND THE DEVICE SUCCESSFULLY CUT AND STAPLED THE TISSUE. A SECOND LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLER WAS USED WITH A CARTRIDGE, ONCE COMPLETED THE SCRUB NURSE ATTEMPTED TO RELEASE THE STAPLER TO RE-LOAD AND IT WOUNLD NOT RELEASE.
06-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A RECONSTITUTION OF POST HARTMANN TRANSIT, THE PART OF THE RECTUM THAT WILL BE THE POSTERIOR PART OF THE ANASTOMOSIS IS CUT WITH A LINEAR STAPLER. THE CIRCULAR DEVICE WAS USED FOR THE TERMINAL TERM ANASTOMOSIS. WHEN PERFORMING THE PNEUMATIC TEST, IT IS OBSERVED HOW THE ANASTOMOSIS FILTERS BY THE PRESENCE OF BUBBLES IN AN ABUNDANT FORM, MAINLY IN THE UNION OF THE LINE OF SQUARE BRACKETS WITH THE CIRCULAR ONE. SUBSEQUENTLY A REINFORCEMENT IS MADE WITH MANUAL SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
05-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THE DURING AN OPEN RECTUM SURGERY, AFTER FIRED, FOUND THE STAPLES MALFORMED. SUTURE WAS USED TO SEW THE WOUND. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
07-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AT THE FIRST TIME, KNIFE IN RELOAD WAS STUCK. AFTER TRYING TO PUSH, IT WORKED BUT PUSHING KNOB WAS BROKEN AFTER KNIFE RAN A HALF OF RELOAD. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
07-03-2018 | Malfunction | ethicon endo-surgery | 03-04-2018 | 75MM SELECTABLE NEW TLC | Sticking |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AT THE FIRST TIME, KNIFE IN RELOAD WAS STUCK. AFTER TRYING TO PUSH, IT WORKED BUT PUSHING KNOB WAS BROKEN AFTER KNIFE RAN A HALF OF RELOAD. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
07-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN RIGHT UPPER LOBECTOMY, THE STAPLES AT THE MIDDLE PART OF THE CARTRIDGE WERE UNFORMED AT THE 2ND FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
07-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED PLUS LONG 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: AFTER FIRING THE DEVICE THE BRACKET EXPULSION WAS NOT SUCCEED COMPLETELY. THE BRACKETS WERE NOT FORM COMPLETELY. NO HARM TO THE PATIENT.
08-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | 75MM SELECTABLE NEW TLC | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE DID NOT OPEN AFTER FIRING. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
07-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PNEUMONECTOMY, THE STAPLES AT THE DISTAL END OF THE CARTRIDGE WERE UNFORMED AT THE 5TH FIRING, AND BLEEDING OCCURRED. THE DEVICE WAS USED ON THE PULMONARY VEIN, AND ITS THICKNESS WAS 1CM. THE SIZE OF THE WOUND WHERE THE BLEEDING OCCURRED WAS ABOUT 1MM. ASTRICTION WAS PERFORMED FOR 5 MINUTES AND TACO-SEAL WAS PUT ON THE TISSUE TO STOP THE BLEEDING. THE AMOUNT OF BLEEDING IS UNKNOWN. NO BLOOD TRANSFUSION WAS REQUIRED. NO OTHER DEVICE WAS REQUIRED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
08-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | ILS 29MM, CURVED | Material Separation | A0413 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ANVIL WAS SEPARATED FROM THE DEVICE WHEN THE TARGET TISSUE WAS ANASTOMOSED. THE SAME EVENT CONTINUED THOUGH THE ANVIL WAS ATTACHED TO THE DEVICE AGAIN. THEN, THE SAME ANVIL AND ANOTHER DEVICE WERE USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-02-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Positioning Failure | A150201 | No Consequences Or Impact To Patient | Event Description: STAPLER BLOCKED. THE DEVICE WAS MANUALLY OPENED. THE STAPLER DID NOT WORK IN AUTOMATICALLY. NO PATIENT CONSEQUENCE.
01-02-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Device Inoperable |  | No Consequences Or Impact To Patient | Event Description: STAPLER BLOCKED. THE DEVICE WAS MANUALLY OPENED. THE STAPLER DID NOT WORK IN AUTOMATICALLY. NO PATIENT CONSEQUENCE.
07-03-2018 | Injury | ethicon endo-surgery | 04-04-2018 | ENDO ILS 25MM, CURVED | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT DURING A GASTRIC ROUX PROCEDURE, THE DOCTOR CLOSED DOWN, FIRED THE DEVICE BUT NO OUTFLOW OF LUMEN ON THE BOWEL, POSSIBLE PART OF LUMEN WAS CAUGHT ON THE BACK WALL. ANASTOMOSES WAS NOT COMPLETE. THE DOCTOR MADE THE DECISION TO CONVERT TO OPEN AND COMPLETED THE ANASTOMOSES BY HAND SEW TECHNIQUE. DOCTOR FEELS EVERYTHING LOOKED OPTIMAL BUT LUMEN WAS PUCKERED. PARTIAL CIRCLE CLOSED BUT AIR COULD GET THROUGH. DOCTOR FELT THE DEVICE FIRED ACCURATELY.
02-02-2018 | Malfunction | davol | 04-04-2018 | OPTIFIX ABSORBABLE FIXATION SYSTEM | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: THE ABSORBABLE FASTENERS STARTED TO BREAK APART AND THEN JAMMED IN THE DEVICE. ONLY A FEW FASTENERS WORKED TO AFFIX THE MESH IN PLACE, AND THEN NO MORE WOULD COME OUT. THERE ARE SUPPOSED TO BE 15 FASTENERS IN EACH DEVICE.
02-02-2018 | Malfunction | davol | 04-04-2018 | OPTIFIX ABSORBABLE FIXATION SYSTEM | Mechanical Jam | A0506 | No Known Impact Or Consequence To Patient | Event Description: THE ABSORBABLE FASTENERS STARTED TO BREAK APART AND THEN JAMMED IN THE DEVICE. ONLY A FEW FASTENERS WORKED TO AFFIX THE MESH IN PLACE, AND THEN NO MORE WOULD COME OUT. THERE ARE SUPPOSED TO BE 15 FASTENERS IN EACH DEVICE.
07-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE THE DEVICE WAS STUCK ON STOMACH TISSUE ON THE FIRST FIRING WITH A GREEN GST RELOAD AND BUTTRESSING MATERIAL. THE POWERED REVERSE BUTTON DID NOT RETRACT THE KNIFE BLADE. THE MANUAL OVERRIDE MECHANISM BROKE WHILE BEING USED AND IT DID NOT RETRACT THE KNIFE BLADE. THE SURGEON RESECTED THE TISSUE OUT OF THE JAWS TO REMOVE THE DEVICE. PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THE POST-OP LEAK TEST SHOWED NO ISSUES WITH THE STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
07-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE THE DEVICE WAS STUCK ON STOMACH TISSUE ON THE FIRST FIRING WITH A GREEN GST RELOAD AND BUTTRESSING MATERIAL. THE POWERED REVERSE BUTTON DID NOT RETRACT THE KNIFE BLADE. THE MANUAL OVERRIDE MECHANISM BROKE WHILE BEING USED AND IT DID NOT RETRACT THE KNIFE BLADE. THE SURGEON RESECTED THE TISSUE OUT OF THE JAWS TO REMOVE THE DEVICE. PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THE POST-OP LEAK TEST SHOWED NO ISSUES WITH THE STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-03-2018 | Malfunction | ethicon | 04-04-2018 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Tear, Rip or Hole in Device Packaging | A020504 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT VENTRAL HERNIA PROCEDURE ON (B)(6) 2018 AND ABSORBABLE STRAPS WERE USED. THE HEALTHCARE PROFESSIONAL FOUND BOTH THE INTERNAL AND EXTERNAL PACKAGING HAD A PUNCTURE THROUGH THE MIDDLE. THE DEVICE WAS NOT USED ON THE PATIENT. PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED.
05-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Detachment Of Device Component |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY, OOZING OCCURRED AFTER THE FIRST FIRING. THE SITE OF OOZING WAS CHECKED, AND UNFORMED STAPLES FELL OFF THE TISSUE. THE DEVICE WAS USED ON THE HEPATIC VEIN. THE AMOUNT OF THE OOZING WAS UNKNOWN. NO BLOOD TRANSFUSION WAS REQUIRED. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE STAPLE LINE ON THE SPECIMEN SEEMED PROPERLY CLOSED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
05-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY, OOZING OCCURRED AFTER THE FIRST FIRING. THE SITE OF OOZING WAS CHECKED, AND UNFORMED STAPLES FELL OFF THE TISSUE. THE DEVICE WAS USED ON THE HEPATIC VEIN. THE AMOUNT OF THE OOZING WAS UNKNOWN. NO BLOOD TRANSFUSION WAS REQUIRED. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE STAPLE LINE ON THE SPECIMEN SEEMED PROPERLY CLOSED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
19-03-2018 | Malfunction | covidien north haven - mfg | 04-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE AND DEVICE DID NOT LOCK ON TISSUE. ACCORDINGLY, THEY HAD FIRED THE STAPLER 3 TIMES DURING THE LOBECTOMY AND WHEN ATTEMPTED TO FIRE FOR THE 4TH TIME THE STAPLER WOULD NOT FIRE. TO FINISH THE CASE THEY REPLACED BOTH INSTRUMENTS AND THE PROCEDURE CONTINUED WITHOUT INCIDENT. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | covidien north haven - mfg | 04-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ATYPICAL PULMONARY RESECTION, WHEN ON PARENCHYMA RESECTION, THE FIRST STAPLER WAS ABLE TO FIRE, THERE WAS POOR FORMATION OF DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE AND WAS FIXED BY NEW FIRINGS. THE SECOND STAPLER, DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, THERE WAS POOR STAPLE FORMATION OF THE DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE. AS STATED, THEY REINFORCED AND PROTECTED THE STAPLE LINE TO PREVENT AIR LEAKS WITH SEALANT AND USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTED, DURING SLEEVE GASTRECTOMY PROCEDURE, THE RELOADS DID NOT FIRE AFTER PRESSING THE GREEN BUTTON OF THE STAPLER. ALSO, THE STAPLER MISFIRED THREE GREEN 60MM LOADING UNITS AFTER PUSHING GREEN BUTTON AND IT DIDN'T FIRE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | covidien lp llc north haven | 04-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RECTOSIGMOIDECTOMY, WHEN ON THE TRANSECTION OF THE RECTUM, THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. THERE WAS NO PATIENT INJURY.
24-12-2017 | Malfunction | covidien lp llc north haven | 04-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE DEVICE MISFIRED. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE GREEN SAFETY BUTTON COULD NOT BE PRESSED.
24-12-2017 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE DEVICE MISFIRED. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RECTOSIGMOIDECTOMY, WHEN ON THE TRANSECTION OF THE RECTUM, THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. REPLACED BY A NEW DEVICE TO COMPLETE THE CASE TO COMPLETE THE CASE. NO PATIENT INJURY.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON THE FOURTH FIRING DURING A SUB TOTAL GASTRECTOMY PROCEDURE, THE DEVICE WAS SET THE SAME WAY AS THE THREE PREVIOUS FIRINGS, HOWEVER, THE SURGEON COULD NOT FIRE THE DEVICE AT ALL. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | covidien lp llc north haven | 04-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON THE FOURTH FIRING DURING A SUB TOTAL GASTRECTOMY PROCEDURE, THE DEVICE WAS SET THE SAME WAY AS THE THREE PREVIOUS FIRINGS, HOWEVER, THE SURGEON COULD NOT FIRE THE DEVICE AT ALL. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | covidien lp llc north haven | 04-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE GREEN SAFETY BUTTON COULD NOT BE PRESSED.
07-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY, AN UNEXPECTED SOUND WAS HEARD AND THE STAPLES WERE UNFORMED AT THE FIRING ON AN EXISTING STAPLE LINE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT HEMIHEPATECTOMY, WHEN CUTTING THE LEFT HEMI LIVER, THE DEVICE DID NOT FIRE AND THE BLACK PUSHER THUMB PIECE CANNOT BE MOVED AT ALL. THERE WAS NO PATIENT INJURY.
01-02-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | 75MM SELECTABLE NEW TLC | Fitting Problem | A1208 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE WHEN THE DOCTOR NEEDED THE DEVICE FOR TRANSECTION AND ANASTOMOSIS, THE THEATRE STAFF LOADED THE DEVICE WITH RELOAD. THE DOCTOR SAID THAT THE RELOAD DID NOT WANT TO CORRECTLY FIT INTO DEVICE. THE DOCTOR THEN DID A HAND ANASTOMOSIS, THERE WERE NO PATIENT CONSEQUENCES REPORTED.
09-03-2018 | Injury | ethicon endo-surgery | 04-04-2018 | ILS 25MM, CURVED | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN OPEN LOW ANTERIOR RESECTION, LEAKAGE OCCURRED. THE DEVICE WAS USED ON THE RECTUM. THERE WAS NO FEEDBACK THAT THE WASHER WAS CUT, AND NO BREAKING SOUND OF THE WASHER WAS HEARD AT THE FIRING. THE SURGEON FIRED THE DEVICE BY HIS BOTH HANDS. THEN, IT WAS FOUND THAT A HALF OF THE STAPLE LINE WAS INCOMPLETE, AND THE STAPLES WERE MALFORMED. THE DONUTS WERE CREATED PROPERLY. THE AMOUNT OF THE BLEEDING WAS UNKNOWN, AND NO BLOOD TRANSFUSION WAS REQUIRED. PRIOR TO THE EVENT, THE OPERATION WAS QUITE DIFFICULT AND CONVERTED FROM LAPAROSCOPIC TO OPEN. THE PROCEDURE WAS CONVERTED TO PERMANENT COLOSTOMY TO COMPLETE THE CASE.
09-03-2018 | Injury | ethicon endo-surgery | 04-04-2018 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN OPEN LOW ANTERIOR RESECTION, LEAKAGE OCCURRED. THE DEVICE WAS USED ON THE RECTUM. THERE WAS NO FEEDBACK THAT THE WASHER WAS CUT, AND NO BREAKING SOUND OF THE WASHER WAS HEARD AT THE FIRING. THE SURGEON FIRED THE DEVICE BY HIS BOTH HANDS. THEN, IT WAS FOUND THAT A HALF OF THE STAPLE LINE WAS INCOMPLETE, AND THE STAPLES WERE MALFORMED. THE DONUTS WERE CREATED PROPERLY. THE AMOUNT OF THE BLEEDING WAS UNKNOWN, AND NO BLOOD TRANSFUSION WAS REQUIRED. PRIOR TO THE EVENT, THE OPERATION WAS QUITE DIFFICULT AND CONVERTED FROM LAPAROSCOPIC TO OPEN. THE PROCEDURE WAS CONVERTED TO PERMANENT COLOSTOMY TO COMPLETE THE CASE.
09-03-2018 | Injury | ethicon endo-surgery | 04-04-2018 | ILS 25MM, CURVED | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN OPEN LOW ANTERIOR RESECTION, LEAKAGE OCCURRED. THE DEVICE WAS USED ON THE RECTUM. THERE WAS NO FEEDBACK THAT THE WASHER WAS CUT, AND NO BREAKING SOUND OF THE WASHER WAS HEARD AT THE FIRING. THE SURGEON FIRED THE DEVICE BY HIS BOTH HANDS. THEN, IT WAS FOUND THAT A HALF OF THE STAPLE LINE WAS INCOMPLETE, AND THE STAPLES WERE MALFORMED. THE DONUTS WERE CREATED PROPERLY. THE AMOUNT OF THE BLEEDING WAS UNKNOWN, AND NO BLOOD TRANSFUSION WAS REQUIRED. PRIOR TO THE EVENT, THE OPERATION WAS QUITE DIFFICULT AND CONVERTED FROM LAPAROSCOPIC TO OPEN. THE PROCEDURE WAS CONVERTED TO PERMANENT COLOSTOMY TO COMPLETE THE CASE.
23-02-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER MISFIRED AND INCOMPLETE STAPLES WERE DEPLOYED. IT WASN'T KNOWN WHICH FIRING. THREE WHITE RELOADS AND ONE BLUE RELOAD WERE USED WITH THE SAME ISSUE, NO BUTTRESSING MATERIAL USED. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
23-02-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER MISFIRED AND INCOMPLETE STAPLES WERE DEPLOYED. IT WASN'T KNOWN WHICH FIRING. THREE WHITE RELOADS AND ONE BLUE RELOAD WERE USED WITH THE SAME ISSUE, NO BUTTRESSING MATERIAL USED. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
10-02-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED 60 ECHELON +, 280MM SHAFT | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SMALL BOWEL RESECTION, THE CARTRIDGE COULD NOT BE LOADED INTO THE JAW PROPERLY AT THE 3RD FIRING. THE DEVICE WAS USED BETWEEN THE JEJUNUM AND THE JEJUNUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
14-02-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED 60 ECHELON +, 280MM SHAFT | Detachment Of Device Component |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN SIGMOIDECTOMY, THE CARTRIDGE PAN WAS LEFT IN THE JAW WHEN THE CARTRIDGE WAS REMOVED AFTER THE FIRST FIRING. THE DEVICE WAS USED ON THE COLON. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ATYPICAL PULMONARY RESECTION, WHEN ON PARENCHYMA RESECTION, THE FIRST STAPLER WAS ABLE TO FIRE, THERE WAS POOR FORMATION OF DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE AND WAS FIXED BY NEW FIRINGS. THE SECOND STAPLER, DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, THERE WAS POOR STAPLE FORMATION OF THE DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE. AS STATED, THEY REINFORCED AND PROTECTED THE STAPLE LINE TO PREVENT AIR LEAKS WITH SEALANT AND USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ATYPICAL PULMONARY RESECTION, WHEN ON PARENCHYMA RESECTION, THE FIRST STAPLER WAS ABLE TO FIRE, THERE WAS POOR FORMATION OF DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE AND WAS FIXED BY NEW FIRINGS. THE SECOND STAPLER, DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, THERE WAS POOR STAPLE FORMATION OF THE DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE. AS STATED, THEY REINFORCED AND PROTECTED THE STAPLE LINE TO PREVENT AIR LEAKS WITH SEALANT AND USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: THE DEVICE WAS LOCKED AND DID NOT OPEN ITS JAWS KEEPING THE CHARGE BETWEEN THEM. ANOTHER CHARGE WAS USED AND THE DEVICE WAS LOCKED AGAIN. THERE WERE NO PATIENT CONSEQUENCES.
08-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: PROF. FIRED VASECR35 TWICE USING PVE35A FOR LOBECTOMY. WHEN HE TRIED TO FIRE THIRD VASECR35, THE KNIFE DID NOT MOVE FORWARD. KNIFE REVERSE BUTTON DID NOT WORK AND THE MANNUAL LEVER DID NOT WORK. HE REMOVED THE DEVICE SAFTELY FROM THE PATIENT'S BODY.AND, HE FINISHED THE SURGERY USING PCE60A USING GST60W SUCCESSFULLY. PATIENT CONSEQUENCES IS UNKNOWN.
28-02-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | FLEX 60 ARTICULATING | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON WENT TO FIRE THE STAPLER AND THE STAPLES WERE FORMED BUT THE DEVICE DID NOT CUT AND BECAME STUCK ON THE TISSUE. THEY REMOVED THE DEVICE AND THEN OPENED ANOTHER DEVICE WHICH WORKED OK. THE OVER SEWED THE STAPLE LINE OF THE ORIGINAL FIRING AS THE SURGEON WASN¿T HAPPY TO LEAVE IT AS IT WAS.
28-02-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | FLEX 60 ARTICULATING | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON WENT TO FIRE THE STAPLER AND THE STAPLES WERE FORMED BUT THE DEVICE DID NOT CUT AND BECAME STUCK ON THE TISSUE. THEY REMOVED THE DEVICE AND THEN OPENED ANOTHER DEVICE WHICH WORKED OK. THE OVER SEWED THE STAPLE LINE OF THE ORIGINAL FIRING AS THE SURGEON WASN¿T HAPPY TO LEAVE IT AS IT WAS.
08-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | GST BLUE RELOAD, 45MM, 6 ROW | Tear, Rip or Hole in Device Packaging | A020504 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT PRIOR TO THE START OF AN UNKNOWN PROCEDURE, THE CUSTOMER NOTICED A HOLE IN THE PACKAGING FOR THE RELOAD. A SECOND LIKE RELOAD WAS USED TO CONTINUE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING STAPLING OF SIGMOID COLON, THE DEVICE STAYED JAMMED CLOSED ON TISSUE. IT WAS IMPOSSIBLE TO STAPLE NOR REMOVE THE DEVICE WHILE THE TWO JAWS WERE CLOSED. THE SURGEON HAS SUCCEEDED TO REMOVE THE DEVICE BY REPEATING SEVERAL MOVES (PULLING ON THE DEVICE). USE OF ANOTHER DEVICE TO FINISH THE PROCEDURE. NO PATIENT CONSEQUENCES.
08-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Component Missing | A020602 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LDG, ONE THIRD OF THE STAPLES AT THE DISTAL END OF THE CARTRIDGE WERE UNFORMED ON THE CENTRAL SIDE AT THE 2ND FIRING. THE DEVICE WAS USED ON THE DUODENUM. THE CUT LINE WAS REINFORCED BY ADDITIONAL HAND SUTURE WITH MONOCRYL. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. PRIOR TO THE EVENT, THE MARKING CLIP ON THE TUMOR COULD NOT BE LOCATED, AND THUS THE RECONSTRUCTION METHOD WAS CHANGED FROM DELTA-SHAPED ANASTOMOSIS TO ROUX-EN-Y ANASTOMOSIS. THE SURGEON IS CONCERNED THAT THE MISSING CLIP MIGHT HAVE AFFECTED THE EVENT.
08-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LDG, ONE THIRD OF THE STAPLES AT THE DISTAL END OF THE CARTRIDGE WERE UNFORMED ON THE CENTRAL SIDE AT THE 2ND FIRING. THE DEVICE WAS USED ON THE DUODENUM. THE CUT LINE WAS REINFORCED BY ADDITIONAL HAND SUTURE WITH MONOCRYL. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. PRIOR TO THE EVENT, THE MARKING CLIP ON THE TUMOR COULD NOT BE LOCATED, AND THUS THE RECONSTRUCTION METHOD WAS CHANGED FROM DELTA-SHAPED ANASTOMOSIS TO ROUX-EN-Y ANASTOMOSIS. THE SURGEON IS CONCERNED THAT THE MISSING CLIP MIGHT HAVE AFFECTED THE EVENT.
08-03-2018 | Malfunction | ethicon endo-surgery | 04-04-2018 | ECHELON FLEX 60 LONG | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE ECHELON DID NOT OPEN AFTER TRIGGERING. REVERSE BUTTON DID NOT WORK. NO HARM TO THE PATIENT.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | MULTIFIRE ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE AND DEVICE DID NOT LOCK ON TISSUE. ACCORDINGLY, THEY HAD FIRED THE STAPLER 3 TIMES DURING THE LOBECTOMY AND WHEN ATTEMPTED TO FIRE FOR THE 4TH TIME THE STAPLER WOULD NOT FIRE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ATYPICAL PULMONARY RESECTION, WHEN ON PARENCHYMA RESECTION, THE FIRST STAPLER WAS ABLE TO FIRE, THERE WAS POOR FORMATION OF DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE AND WAS FIXED BY NEW FIRINGS. THE SECOND STAPLER, DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, THERE WAS POOR STAPLE FORMATION OF THE DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | us surgical puerto rico | 04-04-2018 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ATYPICAL PULMONARY RESECTION, WHEN ON PARENCHYMA RESECTION, THE FIRST STAPLER WAS ABLE TO FIRE, THERE WAS POOR FORMATION OF DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE AND WAS FIXED BY NEW FIRINGS. THE SECOND STAPLER, DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, THERE WAS POOR STAPLE FORMATION OF THE DISTAL STAPLES AND THE STAPLE LINE WAS INCOMPLETE. THERE WAS NO PATIENT INJURY.
09-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | ILS 33MM, CURVED | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SIGMOID RESECTION PROCEDURE THE DEVICE APPEARED TO CUT BUT DID NOT RELEASE ANY STAPLES. THE PROCEDURE WAS COMPLETED BY SUTURING THE CLOSURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A PROXIMAL PANCEATICUDUODECTOMY PROCEDURE, THE TISSUE AT THE SIDE OF THE STAPLER WAS TORN AND OOZING BLEEDING HAD OCCURRED. THE JAWS OF THE DEVICE WAS ARTICULATED AND FIRED AND THE FIRING MODE WAS CHANGED TO MEDIUM SPEED DURING THE FIRING. THE TISSUE WAS SUTURED TO SOLVE THE OOZING.
19-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AFTER LOADING THE REINFORCED RELOAD AND AFTER FOUR SUCCESSFUL FIRINGS, THE DOCTOR PRESSED THE GREEN READY TO FIRE BUTTON BUT THE FIRING LEVER WAS LOOSE AND IT DIDN'T FIRE. THE DOCTOR REPEATED THE FIRING CYCLE MORE TWO TIMES BUT STILL IT HAD THE SAME RESULT. THERE WAS NO PATIENT INJURY.
07-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC UNKNOWN PROCEDURE, THE FIRING FORCE WAS HIGHER THAN EXPECTED AND THE FIRING KNOB WAS BROKEN OFF WHEN THE DEVICE WAS USED FOR FEEA. NO PIECES WERE LEFT INSIDE THE PATIENT. THE FORCEPS WERE USED TO MOVE THE REST OF THE KNOB AND TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
07-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | 75MM SELECTABLE NEW TLC | Detachment Of Device Component |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC UNKNOWN PROCEDURE, THE FIRING FORCE WAS HIGHER THAN EXPECTED AND THE FIRING KNOB WAS BROKEN OFF WHEN THE DEVICE WAS USED FOR FEEA. NO PIECES WERE LEFT INSIDE THE PATIENT. THE FORCEPS WERE USED TO MOVE THE REST OF THE KNOB AND TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
20-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ENDOSCOPIC LOBECTOMY THORACOSCOPIC SURGERY, WHILE THE PULMONARY ARTERY WAS DISCONNECTED, THE SURGEON WAS ABLE TO THE PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE SO THE STAPLER WAS UNABLE TO FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-01-2018 | Injury | ethicon endo-surgery | 05-04-2018 | ECHELON FLEX 60 LONG | Difficult to Open or Close | A051104 | No Code Available | Event Description: IT WAS REPORTED THAT DURING MINI GASTRIC BYPASS OPERATION, STAPLER JAWS WERE LOCKED ON THE STOMACH TISSUE AND THEY DIDN'T OPEN. THE CASE WAS DRAGGED ON FOR 50 MINUTES. AFTER THE EFFORTS MADE TO OPEN THE JAWS, BLEEDING OCCURRED ON STOMACH SMALL CURVATURA. THE STAPLER'S JAWS WERE OPENED LATER ON BY ITSELF. THE BLEEDING WAS STOPPED WITH HEMOSTATS.
27-01-2018 | Injury | ethicon endo-surgery | 05-04-2018 | ECHELON FLEX 60 LONG | Mechanical Jam | A0506 | No Code Available | Event Description: IT WAS REPORTED THAT DURING MINI GASTRIC BYPASS OPERATION, STAPLER JAWS WERE LOCKED ON THE STOMACH TISSUE AND THEY DIDN'T OPEN. THE CASE WAS DRAGGED ON FOR 50 MINUTES. AFTER THE EFFORTS MADE TO OPEN THE JAWS, BLEEDING OCCURRED ON STOMACH SMALL CURVATURA. THE STAPLER'S JAWS WERE OPENED LATER ON BY ITSELF. THE BLEEDING WAS STOPPED WITH HEMOSTATS.
05-04-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS FOUND IN DECONTAMINATION WITH NO INFORMATION AND HAD AN UNKNOWN ISSUE. PATIENT STATUS WAS UNKNOWN.
19-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE EXCISION OF STOMACH ON A SLEEVE GASTRECTOMY LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO THE PRESS GREEN BUTTON AND SQUEEZES THE HANDLE BUT THE STAPLE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
16-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT FIRST FIRING OF LOBECTOMY, THE SURGEON TRIED TO FIRE THE DEVICE; HOWEVER, COULD NOT. REPLACED BY NEW HANDLE AND CONNECTED THE SAME RELOAD; HOWEVER COULD NOT FIRE THE DEVICE. REPLACED BY NEW RELOAD AND CONNECTED TO THE SECOND HANDLE AND THE FIRING WAS COMPLETED. THERE WAS NO PATIENT INJURY.
28-02-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | CONTOUR TRANSTAR STAPLER SET | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN ABLATION OF ENDOMETRIOSIS, AFTER THE FIRST RELOADING, THE DEVICE JAMMED AFTER HAVING POSITIONED ON RECTUM AND CLOSED IT. "THE PEAK" STAYED JAMMED IN CLOSED POSITION. A PIECE OF THE RECTAL WALL HAS BEEN TORN OFF. USE OF ANOTHER DEVICE WITHOUT DIFFICULTY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
28-02-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | CONTOUR TRANSTAR STAPLER SET | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN ABLATION OF ENDOMETRIOSIS, AFTER THE FIRST RELOADING, THE DEVICE JAMMED AFTER HAVING POSITIONED ON RECTUM AND CLOSED IT. "THE PEAK" STAYED JAMMED IN CLOSED POSITION. A PIECE OF THE RECTAL WALL HAS BEEN TORN OFF. USE OF ANOTHER DEVICE WITHOUT DIFFICULTY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | GIA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT OPEN HEMI-COLECTOMY, WHILE STAPLING OF THE RIGHT COLON, THE STAPLER AND THE RELOAD JAMMED AND WOULD NOT FIRE. ANOTHER STAPLER AND RELOAD WAS OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
19-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE EXCISION OF STOMACH ON A SLEEVE GASTRECTOMY LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO THE PRESS GREEN BUTTON AND SQUEEZES THE HANDLE BUT THE STAPLE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Abscess | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Bacterial Infection | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Cardiac Arrest | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Cellulitis | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Stroke/CVA | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Coagulation Disorder | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Death | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Dehydration | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Dysphagia/ Odynophagia | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Encephalopathy | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Low Blood Pressure/ Hypotension | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Hypothermia | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Hypoxia | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Incontinence | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Memory Loss/Impairment | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Muscle Weakness/Atrophy | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Internal Organ Perforation | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Pain | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Pleural Effusion | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Renal Failure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Sepsis | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Shock | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Tachycardia | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Urinary Tract Infection | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Electrolyte Imbalance | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Hypernatremia | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Ulcer | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Anxiety | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Depression | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Respiratory Alkalosis | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Respiratory Failure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Lethargy | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Ascites | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Fluid Discharge | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Fibrosis | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Convulsion/Seizure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Thrombosis/Thrombus | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Stomach Ulceration | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Hyponatremia | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Death | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
19-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RECTOSIGMOIDECTOMY, WHEN ON THE TRANSECTION OF THE RECTUM, THE FIRST AND SECOND STAPLERS DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. AS STATED, THEY OPENED THE DEVICE WITH A NEW RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-02-2018 | Injury | ethicon endo-surgery | 05-04-2018 | PROXIMATE*75MM LINEAR CUTTER | Device Operates Differently Than Expected |  | No Code Available | nan
19-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESECTION OF SIGMA, THE STAPLER DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THERE WAS NO PATIENT INJURY.
15-06-2017 | Malfunction | us surgical puerto rico | 05-04-2018 | GIA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN EXPLORATORY LAPAROTOMY/RIGHT HEMICOLECTOMY, THE DEVICE FELL APART ON THE FIELD. ALL PARTS WERE RETRIEVED AND SEQUESTERED. IT WAS NOT KNOWN HOW IT FELL APART. NEW STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | IDRIVE ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: HOLD FOR LL 4/5/18ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN TRYING TO CUT THE STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE AND ITS HANDLE WAS UNABLE TO SQUEEZE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, WHEN THE SURGEON FIRED THE CONTOUR THE STAPLES CLOSED PROPERLY ON THE SPECIMEN END HOWEVER ON THE BOTTOM END NO STAPLES WERE FOUND AND THE LUMEN REMAINED OPEN. THE SURGEON APPEARED TO FIRE THE CONTOUR CORRECTLY AND ALL THE STAPLE DRIVERS APPEARED TO HAVE BEEN DRIVEN UPWARDS AND NO STAPLES WERE SEEN ON THE DEVICE. THE SURGEON HAD TO HAND SEW THE ANASTOMOSIS RATHER THAN USE A CIRCULAR DEVICE AS PLANNED. THE PROCEDURE WAS EXTENDED BY JUST OVER 30 MINUTES. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
19-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE STAPLER WOULD NOT FIRE AFTER THE THIRD LOAD SO A NEW STAPLER WAS OPENED. THERE WAS NO PATIENT OUTCOME.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ENDOSCOPIC LOBECTOMY THORACOSCOPIC SURGERY, WHILE THE PULMONARY ARTERY WAS DISCONNECTED, THE SURGEON WAS ABLE TO THE PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE SO THE STAPLER WAS UNABLE TO FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | nan | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WOULD NOT FIT INTO THE ADAPTER, DESCRIBED AS SOMETHING WAS BROKE INSIDE THE SHAFT. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | nan | Human-Device Interface Problem | A22 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WOULD NOT FIT INTO THE ADAPTER, DESCRIBED AS SOMETHING WAS BROKE INSIDE THE SHAFT. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
16-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT FIRST FIRING OF LOBECTOMY, THE SURGEON TRIED TO FIRE THE DEVICE; HOWEVER, COULD NOT. REPLACED BY NEW HANDLE AND CONNECTED THE SAME RELOAD; HOWEVER COULD NOT FIRE THE DEVICE. THERE WAS NO PATIENT INJURY.
14-07-2015 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING THE PROCEDURE, THE FIRST SHOT WITH PURPLE LOAD, MADE A NOISE. THEY PLACED A SECOND RELOAD WHICH DID NOT WORK AND EXCHANGED FOR A THIRD RELOAD THAT DID NOT WORK EITHER. THEY USED ANOTHER STAPLER WHICH FUNCTIONED NORMALLY IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ENDOSCOPIC PULMONARY LOBECTOMY THORACOSCOPIC SURGERY, WHILE CLOSING OF THE LOBAR BRONCHI, THE STAPLER WAS ABLE TO FIRE BUT JAWS COULD NOT BE CLOSED. THERE WAS NO PATIENT INJURY.
13-03-2018 | Death | ethicon endo-surgery | 05-04-2018 | 75MM SELECTABLE NEW TLC | Leak/Splash | A0504 | Death | Event Description: IT WAS REPORTED THAT DURING AN OPEN COLORECTAL SURGERY, A VERY PROFESSIONAL SURGEON USED THE DEVICE. AFTER ONE DAY, THERE WAS A MAJOR LEAK IN THE STAPLE LINE AND THE PATIENT HAS EXPIRED.
19-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | TA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING OPEN HEMICOLECTOMY PROCEDURE, WHILE SHOOTING THE TA AT THE PROXIMAL RECTUS, THE DEVICE HAD POOR STAPLE FORMATION AS PART OF THE STAPLE LINE WAS OK, BUT THERE WAS ABOUT 1 CENTIMETER WITH OPEN STAPLES AFTER SHOOTING. THE RETAINING PIN SEATED PROPERLY UPON SQUEEZING THE HANDLE AND AN OVERSEW WAS DONE TO ON THE OPEN STAPLE LINE. NO PATIENT INJURY.
15-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | EEA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE MISFIRED.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, WHILE RESECTING THE SIGMOID COLON, ONCE THE DEVICE WAS LOADED AND THE GREEN BUTTON WAS PUSHED, THE RELOAD FAILED TO FIRE ONCE THE TRIGGER WAS PULLED, IT WOULD NOT ADVANCE THE STAPLES NOR THE KNIFE. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE. THEY OPENED AN ADDITIONAL RELOAD, WHICH WORKED CORRECTLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | EGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE STAPLER DID NOT FIRE. NO PATIENT INJURY.
17-02-2018 | Injury | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
17-02-2018 | Injury | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HAND ASSISTED RIGHT HEPATECTOMY WITH ABSCESS PROCEDURE, THE SCRUB CYCLED THE RELOAD PRIOR TO HANDING IT TO THE SURGEON. DURING THE LAST FIRING OVER THE HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, ONLY A FEW OF THE STAPLES WERE FIRED, MOST DID NOT DEPLOY, BUT THE KNIFE BLADE DEPLOYED. THIS RESULTED TO POOR STAPLE FORMATION OR INCOMPLETE STAPLE LINE. THERE WAS UNANTICIPATED TISSUE LOSS AND DAMAGE. THE OPERATIVE TIME WAS DEFINITELY INCREASED BY MORE THAN 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED MORE THAN 1 INCH. THERE WAS ALSO BLOOD LOSS OF MORE THAN 500CC AND WHICH WOULD MEASURE IN LITERS. THE PROCEDURE WAS THEN CONVERTED FROM LAP PROCEDURE TO OPEN TO STOP BLEEDING WITH STITCHES IN THE HEPATIC VEIN BUT DURING THE CONVERSION TO OPEN, THE PATIENT CODED FOR ABOUT 22 MINUTES. THE PATIENT'S COGNITIVE IMPACT WAS A SCERTAIN. THE PATIENT WAS STILL IN THE HOSPITAL UNTIL RECENTLY AND SENT OVER TO A LONG TERM CARE FACILITY, STILL INTUBATED.
08-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: THERE WAS A BAD FORMATION OF THE STAPLES IN THE THIRD SHOT. PROCEDURE: SLEEVE GASTRECTOMY, PATIENT CONSEQUENCE: UNKNOWN.
19-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AFTER LOADING THE REINFORCED RELOAD AND AFTER FOUR SUCCESSFUL FIRINGS, THE DOCTOR PRESSED THE GREEN READY TO FIRE BUTTON BUT THE FIRING LEVER WAS LOOSE AND IT DIDN'T FIRE. THE DOCTOR REPEATED THE FIRING CYCLE MORE TWO TIMES BUT STILL IT HAD THE SAME RESULT. THERE WAS NO PATIENT INJURY.
01-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: THE STAPLER LOCKED DURING THE PROCEDURE AND THEY FAILED TO UNLOCK IT. THERE WAS NO HARM TO THE PATIENT. THEY TRADED THE DEFECTED STAPLER FOR ANOTHER OF THE SAME TYPE. PROCEDURE: GASTROPLASTY
01-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: THE STAPLER LOCKED DURING THE PROCEDURE AND THEY FAILED TO UNLOCK IT. THERE WAS NO HARM TO THE PATIENT. THEY TRADED THE DEFECTED STAPLER FOR ANOTHER OF THE SAME TYPE. PROCEDURE: GASTROPLASTY
19-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACIC PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO PARENCHYMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THEY USED ANOTHER RELOAD TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
19-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACIC PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO PARENCHYMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THEY USED ANOTHER RELOAD TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
06-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | 75MM SELECTABLE NEW TLC | Component Falling |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC PARTIAL GASTRIC RESECTION, FOUND ONE METAL PIECE WAS IN THE GASTRIC SPECIMEN. AFTER THEY CHECKED, NO OTHER PIECES WERE IN THE PATIENT OR SPECIMEN, SO THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
19-03-2018 | Malfunction | covidien north haven-mfg | 05-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE EXCISION OF STOMACH ON A SLEEVE GASTRECTOMY LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO THE PRESS GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLE DID NOT FIRE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
01-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | SIGNIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAP DISTAL GASTRECTOMY, AN ADAPTER WAS USED IN QUESTION. THE DEVICE'S JAWS OPENING/CLOSING WERE CHECKED, THEN THE SURGEON INSERTED THE DEVICE INTO THE CAVITY AND TRIED TO OPEN THE JAWS, HOWEVER COULD NOT. REPLACED BY ANOTHER ADAPTER AND THE PROCEDURE WAS COMPLETED. THE EVENT DID NOT REPEAT AFTER THE SURGERY. NO PATIENT CONDITION REPORTED. THEY CONTINUED TO USE THE ORIGINAL ADAPTER IN A PROCEDURE.
01-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: THE STAPLER LOCKED AND THE DOCTOR TRIED TO REVERT MANUALLY BUT FAILED. REQUESTED EXCHANGE. THERE WAS NO HARM TO THE PATIENT. ACTION USER TOOK TO MITIGATE/ RESOLVE PROBLEM DURING PROCEDURE - WITH SAME LIKE PRODUCT
19-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RECTOSIGMOIDECTOMY, WHEN ON THE TRANSECTION OF THE RECTUM, THE FIRST AND SECOND STAPLERS DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN TRYING TO CUT THE STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE AND ITS HANDLE WAS UNABLE TO SQUEEZE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
08-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: DURING THE ENDOSCOPIC LUNG CANCER SURGERY, AFTER FIRED WITH THE BLUE RELOAD, FOUND THE STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES TO THE PATIENT.
13-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: AT THE 3RD OR THE 4TH FIRING OF LOBECTOMY, THE SURGEON COULD NOT FIRE THE DEVICE. THE SALES REP NOTED THE TRACE OF FIRING HALFWAY. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT INJURY.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ENDOSCOPIC PULMONARY LOBECTOMY THORACOSCOPIC SURGERY, WHILE CLOSING OF THE LOBAR BRONCHI, THE STAPLER WAS ABLE TO FIRE BUT JAWS COULD NOT BE CLOSED. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN THORACOSCOPIC LOBECTOMY PROCEDURE, WHILE THE LOBAR BRONCHUS WAS SEVERED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE STAPLER WAS UNABLE TO FIRE. THE RELOAD WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: COULD NOT FIRE. DURING THE ENDOSCOPIC LOBECTOMY, COULD NOT FIRE WITH THE BLUE RELOAD. AFTER SURGERY, THE SURGEON WANT TO TRY TEST THE DEVICE, FOUND COULD NOT OPEN THE JAW. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
08-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: COULD NOT FIRE. DURING THE ENDOSCOPIC LOBECTOMY, COULD NOT FIRE WITH THE BLUE RELOAD. AFTER SURGERY, THE SURGEON WANT TO TRY TEST THE DEVICE, FOUND COULD NOT OPEN THE JAW. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
19-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN TRYING TO CUT THE STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE AND ITS HANDLE WAS UNABLE TO SQUEEZE. THERE WAS NO PATIENT INJURY.
09-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION WITH OSTOMY, THE DEVICE DID NOT FIRE OR MALFUNCTIONED. AS STATED, ANOTHER DEVICE WAS OPENED AND THE CASE PROCEEDED WITHOUT INCIDENT. THERE WAS NO PATIENT HARM.
09-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION WITH OSTOMY, THE DEVICE DID NOT FIRE OR MALFUNCTIONED. AS STATED, ANOTHER DEVICE WAS OPENED AND THE CASE PROCEEDED WITHOUT INCIDENT. THERE WAS NO PATIENT HARM.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER. DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT THE RECTUM RESECTION, THE OPENING/CLOSING OF JAWS WERE CHECKED THEN THE SURGEON TRIED TO INSERT THE DEVICE TO THE TROCAR, HOWEVER COULD NOT CLOSE THE JAWS. NO COMMUNICATION ERROR OF POWER HANDLE, ADAPTER AND RELOAD OCCURRED. THE EVENT DID NOT REPEAT AFTER SURGERY. THERE WAS NO PATIENT INJURY.
19-03-2018 | Injury | ethicon endo-surgery | 05-04-2018 | FLEX 60 ARTICULATING | Leak/Splash | A0504 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE TITLED: TITLE: ISOPERISTALTIC VERSUS ANTIPERISTALTIC STAPLED SIDE-TO-SIDE ANASTOMOSIS FOR COLON CANCER SURGERY: A RANDOMIZED CONTROLLED TRIAL. AUTHOR(S): AKIHISA MATSUDA, MD, MASAO MIYASHITA, MD, SATOSHI MATSUMOTO, MD,A NOBUYUKI SAKURAZAWA, MD, GORO TAKAHASHI, MD MARINA YAMADA, PHD AND EIJI UCHIDA, MD CITATION: JOURNAL OF SURGICAL RESEARCH 196 (2015) 107-112; HTTP://DX.DOI.ORG/10.1016/J.JSS.2015.02.059. THE AIM OF THIS RANDOMIZED CONTROLLED TRIAL WAS TO EVALUATE WHETHER ISOPERISTALTIC STAPLED SIDE-TO-SIDE ANASTOMOSIS IS COMPARABLE WITH ANTIPERISTALTIC SSSA FOR COLON CANCER SURGERY IN TERMS OF SHORT TERM OUTCOMES. BETWEEN JUL2012 AND JAN2014, 40 PATIENTS WITH COLON CANCER WHO UNDERWENT ELECTIVE CURATIVE RESECTION AND HAD ENOUGH MOBILITY AT ANASTOMOSIS WERE RANDOMIZED INTO TWO GROUPS: ISOPERISTALTIC SSSA GROUP (N=20; N=11 MALE AND N=9 FEMALE; AGE OF 66±12 YEARS) AND ANTIPERISTALTIC SSSA GROUP (N=20; N=11 MALE AND N=9 FEMALE; AGE OF 68±10 YEARS). A 60-MM ECHELON ENDOPATH WITH BLUE CARTRIDGE WAS UTILIZED IN RESECTION AND ANASTOMOSIS FOR BOTH GROUPS. THE ENTRY HOLE OF THE LINEAR STAPLER WAS CLOSED WITH A 1 LAYER OF RUNNING 4-0 PDS II SUTURES. THE COLONIC MESENTERIC DEFECT WAS CLOSED WITH INTERRUPTED SUTURES WITH 4-0 PDS II IN OPEN SURGERY. SURGICAL COMPLICATIONS INCLUDED WOUND INFECTION (N=1 ANTIPERISTALTIC SSSA GROUP AND N=1 ISOPERISTALTIC SSSA GROUP), PROLONGED ILEUS (N=1 ANTIPERISTALTIC SSSA GROUP AND N=2 ISOPERISTALTIC SSSA GROUP), INTRA-ABDOMINAL ABSCESS (N=1 ANTIPERISTALTIC SSSA GROUP AND N=2 ISOPERISTALTIC SSSA GROUP), ANASTOMOTIC LEAKAGE (N=2 ISOPERISTALTIC SSSA GROUP) AND ANASTOMOTIC STENOSIS (N=1 ANTIPERISTALTIC SSSA GROUP). ONE PATIENT IN THE ANTIPERISTALTIC SSSA GROUP HAD ANASTOMOTIC STENOSIS, WHICH IMPROVED CONSERVATIVELY. IN ISOPERISTALTIC SSSA GROUP, 2 PATIENTS WHO HAD INTRA-ABDOMINAL ABSCESS DUE TO ANASTOMOTIC LEAK, AND ONE PATIENT WHO HAD PROLONGED ILEUS HAD REOPERATIONS. THE PRESENT STUDY DEMONSTRATED THAT ALTHOUGH THE DIFFERENCE WAS NOT STATISTICALLY SIGNIFICANT, ANASTOMOTIC LEAKAGE OCCURRED ONLY IN THE ISOPERISTALTIC SSSAGROUP.
19-03-2018 | Injury | ethicon endo-surgery | 05-04-2018 | FLEX 60 ARTICULATING | Leak/Splash | A0504 | Abscess | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE TITLED: TITLE: ISOPERISTALTIC VERSUS ANTIPERISTALTIC STAPLED SIDE-TO-SIDE ANASTOMOSIS FOR COLON CANCER SURGERY: A RANDOMIZED CONTROLLED TRIAL. AUTHOR(S): AKIHISA MATSUDA, MD, MASAO MIYASHITA, MD, SATOSHI MATSUMOTO, MD,A NOBUYUKI SAKURAZAWA, MD, GORO TAKAHASHI, MD MARINA YAMADA, PHD AND EIJI UCHIDA, MD CITATION: JOURNAL OF SURGICAL RESEARCH 196 (2015) 107-112; HTTP://DX.DOI.ORG/10.1016/J.JSS.2015.02.059. THE AIM OF THIS RANDOMIZED CONTROLLED TRIAL WAS TO EVALUATE WHETHER ISOPERISTALTIC STAPLED SIDE-TO-SIDE ANASTOMOSIS IS COMPARABLE WITH ANTIPERISTALTIC SSSA FOR COLON CANCER SURGERY IN TERMS OF SHORT TERM OUTCOMES. BETWEEN JUL2012 AND JAN2014, 40 PATIENTS WITH COLON CANCER WHO UNDERWENT ELECTIVE CURATIVE RESECTION AND HAD ENOUGH MOBILITY AT ANASTOMOSIS WERE RANDOMIZED INTO TWO GROUPS: ISOPERISTALTIC SSSA GROUP (N=20; N=11 MALE AND N=9 FEMALE; AGE OF 66±12 YEARS) AND ANTIPERISTALTIC SSSA GROUP (N=20; N=11 MALE AND N=9 FEMALE; AGE OF 68±10 YEARS). A 60-MM ECHELON ENDOPATH WITH BLUE CARTRIDGE WAS UTILIZED IN RESECTION AND ANASTOMOSIS FOR BOTH GROUPS. THE ENTRY HOLE OF THE LINEAR STAPLER WAS CLOSED WITH A 1 LAYER OF RUNNING 4-0 PDS II SUTURES. THE COLONIC MESENTERIC DEFECT WAS CLOSED WITH INTERRUPTED SUTURES WITH 4-0 PDS II IN OPEN SURGERY. SURGICAL COMPLICATIONS INCLUDED WOUND INFECTION (N=1 ANTIPERISTALTIC SSSA GROUP AND N=1 ISOPERISTALTIC SSSA GROUP), PROLONGED ILEUS (N=1 ANTIPERISTALTIC SSSA GROUP AND N=2 ISOPERISTALTIC SSSA GROUP), INTRA-ABDOMINAL ABSCESS (N=1 ANTIPERISTALTIC SSSA GROUP AND N=2 ISOPERISTALTIC SSSA GROUP), ANASTOMOTIC LEAKAGE (N=2 ISOPERISTALTIC SSSA GROUP) AND ANASTOMOTIC STENOSIS (N=1 ANTIPERISTALTIC SSSA GROUP). ONE PATIENT IN THE ANTIPERISTALTIC SSSA GROUP HAD ANASTOMOTIC STENOSIS, WHICH IMPROVED CONSERVATIVELY. IN ISOPERISTALTIC SSSA GROUP, 2 PATIENTS WHO HAD INTRA-ABDOMINAL ABSCESS DUE TO ANASTOMOTIC LEAK, AND ONE PATIENT WHO HAD PROLONGED ILEUS HAD REOPERATIONS. THE PRESENT STUDY DEMONSTRATED THAT ALTHOUGH THE DIFFERENCE WAS NOT STATISTICALLY SIGNIFICANT, ANASTOMOTIC LEAKAGE OCCURRED ONLY IN THE ISOPERISTALTIC SSSAGROUP.
19-03-2018 | Injury | ethicon endo-surgery | 05-04-2018 | FLEX 60 ARTICULATING | Leak/Splash | A0504 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE TITLED: TITLE: ISOPERISTALTIC VERSUS ANTIPERISTALTIC STAPLED SIDE-TO-SIDE ANASTOMOSIS FOR COLON CANCER SURGERY: A RANDOMIZED CONTROLLED TRIAL. AUTHOR(S): AKIHISA MATSUDA, MD, MASAO MIYASHITA, MD, SATOSHI MATSUMOTO, MD,A NOBUYUKI SAKURAZAWA, MD, GORO TAKAHASHI, MD MARINA YAMADA, PHD AND EIJI UCHIDA, MD CITATION: JOURNAL OF SURGICAL RESEARCH 196 (2015) 107-112; HTTP://DX.DOI.ORG/10.1016/J.JSS.2015.02.059. THE AIM OF THIS RANDOMIZED CONTROLLED TRIAL WAS TO EVALUATE WHETHER ISOPERISTALTIC STAPLED SIDE-TO-SIDE ANASTOMOSIS IS COMPARABLE WITH ANTIPERISTALTIC SSSA FOR COLON CANCER SURGERY IN TERMS OF SHORT TERM OUTCOMES. BETWEEN JUL2012 AND JAN2014, 40 PATIENTS WITH COLON CANCER WHO UNDERWENT ELECTIVE CURATIVE RESECTION AND HAD ENOUGH MOBILITY AT ANASTOMOSIS WERE RANDOMIZED INTO TWO GROUPS: ISOPERISTALTIC SSSA GROUP (N=20; N=11 MALE AND N=9 FEMALE; AGE OF 66±12 YEARS) AND ANTIPERISTALTIC SSSA GROUP (N=20; N=11 MALE AND N=9 FEMALE; AGE OF 68±10 YEARS). A 60-MM ECHELON ENDOPATH WITH BLUE CARTRIDGE WAS UTILIZED IN RESECTION AND ANASTOMOSIS FOR BOTH GROUPS. THE ENTRY HOLE OF THE LINEAR STAPLER WAS CLOSED WITH A 1 LAYER OF RUNNING 4-0 PDS II SUTURES. THE COLONIC MESENTERIC DEFECT WAS CLOSED WITH INTERRUPTED SUTURES WITH 4-0 PDS II IN OPEN SURGERY. SURGICAL COMPLICATIONS INCLUDED WOUND INFECTION (N=1 ANTIPERISTALTIC SSSA GROUP AND N=1 ISOPERISTALTIC SSSA GROUP), PROLONGED ILEUS (N=1 ANTIPERISTALTIC SSSA GROUP AND N=2 ISOPERISTALTIC SSSA GROUP), INTRA-ABDOMINAL ABSCESS (N=1 ANTIPERISTALTIC SSSA GROUP AND N=2 ISOPERISTALTIC SSSA GROUP), ANASTOMOTIC LEAKAGE (N=2 ISOPERISTALTIC SSSA GROUP) AND ANASTOMOTIC STENOSIS (N=1 ANTIPERISTALTIC SSSA GROUP). ONE PATIENT IN THE ANTIPERISTALTIC SSSA GROUP HAD ANASTOMOTIC STENOSIS, WHICH IMPROVED CONSERVATIVELY. IN ISOPERISTALTIC SSSA GROUP, 2 PATIENTS WHO HAD INTRA-ABDOMINAL ABSCESS DUE TO ANASTOMOTIC LEAK, AND ONE PATIENT WHO HAD PROLONGED ILEUS HAD REOPERATIONS. THE PRESENT STUDY DEMONSTRATED THAT ALTHOUGH THE DIFFERENCE WAS NOT STATISTICALLY SIGNIFICANT, ANASTOMOTIC LEAKAGE OCCURRED ONLY IN THE ISOPERISTALTIC SSSAGROUP.
19-03-2018 | Injury | ethicon endo-surgery | 05-04-2018 | FLEX 60 ARTICULATING | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE TITLED: TITLE: ISOPERISTALTIC VERSUS ANTIPERISTALTIC STAPLED SIDE-TO-SIDE ANASTOMOSIS FOR COLON CANCER SURGERY: A RANDOMIZED CONTROLLED TRIAL. AUTHOR(S): AKIHISA MATSUDA, MD, MASAO MIYASHITA, MD, SATOSHI MATSUMOTO, MD,A NOBUYUKI SAKURAZAWA, MD, GORO TAKAHASHI, MD MARINA YAMADA, PHD AND EIJI UCHIDA, MD CITATION: JOURNAL OF SURGICAL RESEARCH 196 (2015) 107-112; HTTP://DX.DOI.ORG/10.1016/J.JSS.2015.02.059. THE AIM OF THIS RANDOMIZED CONTROLLED TRIAL WAS TO EVALUATE WHETHER ISOPERISTALTIC STAPLED SIDE-TO-SIDE ANASTOMOSIS IS COMPARABLE WITH ANTIPERISTALTIC SSSA FOR COLON CANCER SURGERY IN TERMS OF SHORT TERM OUTCOMES. BETWEEN JUL2012 AND JAN2014, 40 PATIENTS WITH COLON CANCER WHO UNDERWENT ELECTIVE CURATIVE RESECTION AND HAD ENOUGH MOBILITY AT ANASTOMOSIS WERE RANDOMIZED INTO TWO GROUPS: ISOPERISTALTIC SSSA GROUP (N=20; N=11 MALE AND N=9 FEMALE; AGE OF 66±12 YEARS) AND ANTIPERISTALTIC SSSA GROUP (N=20; N=11 MALE AND N=9 FEMALE; AGE OF 68±10 YEARS). A 60-MM ECHELON ENDOPATH WITH BLUE CARTRIDGE WAS UTILIZED IN RESECTION AND ANASTOMOSIS FOR BOTH GROUPS. THE ENTRY HOLE OF THE LINEAR STAPLER WAS CLOSED WITH A 1 LAYER OF RUNNING 4-0 PDS II SUTURES. THE COLONIC MESENTERIC DEFECT WAS CLOSED WITH INTERRUPTED SUTURES WITH 4-0 PDS II IN OPEN SURGERY. SURGICAL COMPLICATIONS INCLUDED WOUND INFECTION (N=1 ANTIPERISTALTIC SSSA GROUP AND N=1 ISOPERISTALTIC SSSA GROUP), PROLONGED ILEUS (N=1 ANTIPERISTALTIC SSSA GROUP AND N=2 ISOPERISTALTIC SSSA GROUP), INTRA-ABDOMINAL ABSCESS (N=1 ANTIPERISTALTIC SSSA GROUP AND N=2 ISOPERISTALTIC SSSA GROUP), ANASTOMOTIC LEAKAGE (N=2 ISOPERISTALTIC SSSA GROUP) AND ANASTOMOTIC STENOSIS (N=1 ANTIPERISTALTIC SSSA GROUP). ONE PATIENT IN THE ANTIPERISTALTIC SSSA GROUP HAD ANASTOMOTIC STENOSIS, WHICH IMPROVED CONSERVATIVELY. IN ISOPERISTALTIC SSSA GROUP, 2 PATIENTS WHO HAD INTRA-ABDOMINAL ABSCESS DUE TO ANASTOMOTIC LEAK, AND ONE PATIENT WHO HAD PROLONGED ILEUS HAD REOPERATIONS. THE PRESENT STUDY DEMONSTRATED THAT ALTHOUGH THE DIFFERENCE WAS NOT STATISTICALLY SIGNIFICANT, ANASTOMOTIC LEAKAGE OCCURRED ONLY IN THE ISOPERISTALTIC SSSAGROUP.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | nan | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS ABLE TO FIRE BUT HAD AN INCOMPLETE STAPLE LINE ON THE DISTAL PART AND IT ONLY FIRED HALFWAY THOUGH THE ENTIRETY.
07-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN THE THIRD SHOT THE STAPLER LOCKED AND THERE WAS A DISPLACEMENT OF THE TISSUE, THERE WAS A MALFORMATION OF THE STAPLES, AND THE LINE OF STAPLES WAS VERY IRREGULAR. IN THIS LINE OF STAPLES, AFTER CHECKING THE STOMACH, YOU CAN SEE A GAP. THE TISSUE WAS LACERATED.
07-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | BLUE RELOAD FOR ECHELON 60 | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN THE THIRD SHOT THE STAPLER LOCKED AND THERE WAS A DISPLACEMENT OF THE TISSUE, THERE WAS A MALFORMATION OF THE STAPLES, AND THE LINE OF STAPLES WAS VERY IRREGULAR. IN THIS LINE OF STAPLES, AFTER CHECKING THE STOMACH, YOU CAN SEE A GAP. THE TISSUE WAS LACERATED.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | SIGNIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT THE RECTUM RESECTION, THE OPENING/CLOSING OF JAWS WERE CHECKED THEN THE SURGEON TRIED TO INSERT THE DEVICE TO THE TROCAR, HOWEVER COULD NOT CLOSE THE JAWS. NO COMMUNICATION ERROR OF POWER HANDLE, ADAPTER AND RELOAD OCCURRED. THE EVENT DID NOT REPEAT AFTER SURGERY. THERE WAS NO PATIENT INJURY.
12-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: COULD NOT FIRE&COULD NOT RETURN BLADE. DURING THE ENDOSCOPIC LUNG CANCER SURGERY, COULD NOT FIRE WITH ECR45G. AFTER FIRED WITH ECR45G, COULD NOT RETURN THE BLADE. USED ANOTHER NEW DEVICE TO CUT THE TISSUE AND TAKE OUT OUT THE DEVICE. AS THE CUTTING-OFF TISSUE IS NOT KEY PART, THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
13-03-2018 | Malfunction | covidien north haven-mfg | 05-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: AT THE 3RD OR THE 4TH FIRING OF LOBECTOMY, THE SURGEON COULD NOT FIRE THE DEVICE. THE SALES REP NOTED THE TRACE OF FIRING HALFWAY. THERE WAS NO PATIENT INJURY.
09-01-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE DID NOT FIRE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-03-2018 | Malfunction | covidien lp llc north haven | 05-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS ABLE TO FIRE BUT HAD AN INCOMPLETE STAPLE LINE ON THE DISTAL PART AND IT ONLY FIRED HALFWAY THOUGH THE ENTIRETY.
12-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS SURGERY, FROM THE SPA BY SURGEONS . IT WAS USED THE ATS 45 MM GUN INPUT AND A 6R45B BLUE REFILL, DURING THE USE OF THE DEVICE AT THE TIME OF MAKING THE CUT AND CLOSURE OF THE STAPLE THE TISSUE WAS DISPLACED AND NOT PROPERLY STAPLED. THE SURGEON HAD TO OPEN OTHER DEVICE TO CONTINUE THE SURGERY.
12-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE HANDLE WOULD NOT OPEN, STUCK CLOSED. THIS HAPPENED TODAY IN A GASTRIC BYPASS. PROCEDURE FINISHED WITH ANOTHER DEVICE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-03-2018 | Malfunction | ethicon endo-surgery | 05-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE ONLY THREE-FOURTHS OF THE STAPLES FORMED PROPERLY DURING AN UNSPECIFIED FIRING WITH AN UNSPECIFIED GREEN RELOAD. IT WAS UNKNOWN IF THERE WAS ANY BUTTRESSING MATERIAL USED. THE REMAINING FOURTH OF THE STAPLES WERE "GOAL POSTS"/UNFORMED. SURGEON OVERSEWED ENTIRE STAPLE LINE TO PREVENT LEAK. THERE WERE NO PATIENT CONSEQUENCES.
20-03-2018 | Malfunction | us surgical puerto rico | 05-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, WHILE RESECTING THE SIGMOID COLON, ONCE THE DEVICE WAS LOADED AND THE GREEN BUTTON WAS PUSHED, THE RELOAD FAILED TO FIRE ONCE THE TRIGGER WAS PULLED, IT WOULD NOT ADVANCE THE STAPLES NOR THE KNIFE. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
09-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | ECHELON ENDOPATH STAPLER | Device Operates Differently Than Expected |  | Abscess | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE: TITLE: SLEEVE GASTRECTOMY POSTOPERATIVE HEMORRHAGE IS LINKED TO TYPE-2 DIABETES AND NOT TO SURGICAL TECHNIQUE AUTHOR(S): HADAR SPIVAK, CARMIL AZRAN, GALIA SPECTRE, GALINA LIDERMANN, ORIT BLUMENFELD CITATION: OBES SURG (2017) 27:2927¿2932; DOI 10.1007/S11695-017-2731-5. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO REPORT THE PREVALENCE, SEVERITY, AND RISK FACTORS OF LAPAROSCOPIC SLEEVE GASTRECTOMY POSTOPERATIVE HAEMORRHAGE (ACUTE AND SUBCLINICAL) IN A COHORT WITH STANDARD SETTINGS. FROM 01JAN2014 TO 31DEC2015, 394 PATIENTS UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY AND WERE ASSIGNED INTO THREE SURGICAL TECHNIQUES: GROUP A (N=137) UNDERWENT CONVENTIONAL ORDER OF STAPLING STARTING TO WHERE ECHELON FLEX BLACK CARTRIDGE OR IDRIVE ULTRA PURPLE CARTRIDGE WAS USED AND ENDING WITH BLUE OR BEIGE, GROUP B (N=145) UNDERWENT TIGHT STAPLING WITH SMALLER CLOSING STAPLE HEIGHT RELOAD (WHITE ECHELON FLEX RELOAD OR BEIGE IDRIVE ULTRA), AND GROUP C (N=115) UNDERWENT OVERSUTURING THE STAPLE LINE WITH 2¿0 VICRYL IN A CONTINUOUS BASEBALL FASHION OVER A CONVENTIONAL ORDER OF RELOADS. POSTOPERATIVE COMPLICATIONS INCLUDED 2 PATIENTS HAD LEAKS (GROUP A AND GROUP C) WHICH WERE TREATED BY SURGICAL RE-EXPLORATION AND DRAINAGE, FOLLOWED BY TPN; 2 PATIENTS HAD SLEEVE KINK AND PARTIAL OBSTRUCTION (GROUP A AND GROUP B) TREATED WITH ROUX-EN-Y GASTRIC BYPASS (N=1) AND MONTHS OF LIQUID-ONLY INTAKE (N=1); ONE PATIENT HAD ABDOMINAL WALL ADHESION (GROUP B) CAUSING UNSOLVED SEVERE PAIN WHICH WAS TREATED WITH LAPAROSCOPIC EXPLORATION; ONE PATIENT HAD PELVIC ABSCESS (GROUP A) TREATED WITH DRAINAGE; TWO PATIENTS HAD WOUND INFECTIONS (GROUP A AND GROUP B) REQUIRING BEDSIDE DRAINAGE; AND ONE PATIENT WAS DIAGNOSED WITH CLOSTRIDIUM DIFFICILE INFECTION (GROUP C). SINCE PATIENTS WITH DIABETES HAVE IMPAIRED WOUND HEALING, A DEFECTIVE STAPLE LINE HEALING, TISSUE MATURING, AND THUS FRAIL SEALING OF EMBEDDED VESSELS COULD BE THE REASON FOR EXCESSIVE POSTOPERATIVE HAEMORRHAGE IN TYPE 2 DIABETES.
09-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | ECHELON ENDOPATH STAPLER | Device Operates Differently Than Expected |  | Adhesion(s) | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE: TITLE: SLEEVE GASTRECTOMY POSTOPERATIVE HEMORRHAGE IS LINKED TO TYPE-2 DIABETES AND NOT TO SURGICAL TECHNIQUE AUTHOR(S): HADAR SPIVAK, CARMIL AZRAN, GALIA SPECTRE, GALINA LIDERMANN, ORIT BLUMENFELD CITATION: OBES SURG (2017) 27:2927¿2932; DOI 10.1007/S11695-017-2731-5. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO REPORT THE PREVALENCE, SEVERITY, AND RISK FACTORS OF LAPAROSCOPIC SLEEVE GASTRECTOMY POSTOPERATIVE HAEMORRHAGE (ACUTE AND SUBCLINICAL) IN A COHORT WITH STANDARD SETTINGS. FROM 01JAN2014 TO 31DEC2015, 394 PATIENTS UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY AND WERE ASSIGNED INTO THREE SURGICAL TECHNIQUES: GROUP A (N=137) UNDERWENT CONVENTIONAL ORDER OF STAPLING STARTING TO WHERE ECHELON FLEX BLACK CARTRIDGE OR IDRIVE ULTRA PURPLE CARTRIDGE WAS USED AND ENDING WITH BLUE OR BEIGE, GROUP B (N=145) UNDERWENT TIGHT STAPLING WITH SMALLER CLOSING STAPLE HEIGHT RELOAD (WHITE ECHELON FLEX RELOAD OR BEIGE IDRIVE ULTRA), AND GROUP C (N=115) UNDERWENT OVERSUTURING THE STAPLE LINE WITH 2¿0 VICRYL IN A CONTINUOUS BASEBALL FASHION OVER A CONVENTIONAL ORDER OF RELOADS. POSTOPERATIVE COMPLICATIONS INCLUDED 2 PATIENTS HAD LEAKS (GROUP A AND GROUP C) WHICH WERE TREATED BY SURGICAL RE-EXPLORATION AND DRAINAGE, FOLLOWED BY TPN; 2 PATIENTS HAD SLEEVE KINK AND PARTIAL OBSTRUCTION (GROUP A AND GROUP B) TREATED WITH ROUX-EN-Y GASTRIC BYPASS (N=1) AND MONTHS OF LIQUID-ONLY INTAKE (N=1); ONE PATIENT HAD ABDOMINAL WALL ADHESION (GROUP B) CAUSING UNSOLVED SEVERE PAIN WHICH WAS TREATED WITH LAPAROSCOPIC EXPLORATION; ONE PATIENT HAD PELVIC ABSCESS (GROUP A) TREATED WITH DRAINAGE; TWO PATIENTS HAD WOUND INFECTIONS (GROUP A AND GROUP B) REQUIRING BEDSIDE DRAINAGE; AND ONE PATIENT WAS DIAGNOSED WITH CLOSTRIDIUM DIFFICILE INFECTION (GROUP C). SINCE PATIENTS WITH DIABETES HAVE IMPAIRED WOUND HEALING, A DEFECTIVE STAPLE LINE HEALING, TISSUE MATURING, AND THUS FRAIL SEALING OF EMBEDDED VESSELS COULD BE THE REASON FOR EXCESSIVE POSTOPERATIVE HAEMORRHAGE IN TYPE 2 DIABETES.
09-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | ECHELON ENDOPATH STAPLER | Device Operates Differently Than Expected |  | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE: TITLE: SLEEVE GASTRECTOMY POSTOPERATIVE HEMORRHAGE IS LINKED TO TYPE-2 DIABETES AND NOT TO SURGICAL TECHNIQUE AUTHOR(S): HADAR SPIVAK, CARMIL AZRAN, GALIA SPECTRE, GALINA LIDERMANN, ORIT BLUMENFELD CITATION: OBES SURG (2017) 27:2927¿2932; DOI 10.1007/S11695-017-2731-5. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO REPORT THE PREVALENCE, SEVERITY, AND RISK FACTORS OF LAPAROSCOPIC SLEEVE GASTRECTOMY POSTOPERATIVE HAEMORRHAGE (ACUTE AND SUBCLINICAL) IN A COHORT WITH STANDARD SETTINGS. FROM 01JAN2014 TO 31DEC2015, 394 PATIENTS UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY AND WERE ASSIGNED INTO THREE SURGICAL TECHNIQUES: GROUP A (N=137) UNDERWENT CONVENTIONAL ORDER OF STAPLING STARTING TO WHERE ECHELON FLEX BLACK CARTRIDGE OR IDRIVE ULTRA PURPLE CARTRIDGE WAS USED AND ENDING WITH BLUE OR BEIGE, GROUP B (N=145) UNDERWENT TIGHT STAPLING WITH SMALLER CLOSING STAPLE HEIGHT RELOAD (WHITE ECHELON FLEX RELOAD OR BEIGE IDRIVE ULTRA), AND GROUP C (N=115) UNDERWENT OVERSUTURING THE STAPLE LINE WITH 2¿0 VICRYL IN A CONTINUOUS BASEBALL FASHION OVER A CONVENTIONAL ORDER OF RELOADS. POSTOPERATIVE COMPLICATIONS INCLUDED 2 PATIENTS HAD LEAKS (GROUP A AND GROUP C) WHICH WERE TREATED BY SURGICAL RE-EXPLORATION AND DRAINAGE, FOLLOWED BY TPN; 2 PATIENTS HAD SLEEVE KINK AND PARTIAL OBSTRUCTION (GROUP A AND GROUP B) TREATED WITH ROUX-EN-Y GASTRIC BYPASS (N=1) AND MONTHS OF LIQUID-ONLY INTAKE (N=1); ONE PATIENT HAD ABDOMINAL WALL ADHESION (GROUP B) CAUSING UNSOLVED SEVERE PAIN WHICH WAS TREATED WITH LAPAROSCOPIC EXPLORATION; ONE PATIENT HAD PELVIC ABSCESS (GROUP A) TREATED WITH DRAINAGE; TWO PATIENTS HAD WOUND INFECTIONS (GROUP A AND GROUP B) REQUIRING BEDSIDE DRAINAGE; AND ONE PATIENT WAS DIAGNOSED WITH CLOSTRIDIUM DIFFICILE INFECTION (GROUP C). SINCE PATIENTS WITH DIABETES HAVE IMPAIRED WOUND HEALING, A DEFECTIVE STAPLE LINE HEALING, TISSUE MATURING, AND THUS FRAIL SEALING OF EMBEDDED VESSELS COULD BE THE REASON FOR EXCESSIVE POSTOPERATIVE HAEMORRHAGE IN TYPE 2 DIABETES.
09-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | ECHELON ENDOPATH STAPLER | Device Operates Differently Than Expected |  | Obstruction/Occlusion | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE: TITLE: SLEEVE GASTRECTOMY POSTOPERATIVE HEMORRHAGE IS LINKED TO TYPE-2 DIABETES AND NOT TO SURGICAL TECHNIQUE AUTHOR(S): HADAR SPIVAK, CARMIL AZRAN, GALIA SPECTRE, GALINA LIDERMANN, ORIT BLUMENFELD CITATION: OBES SURG (2017) 27:2927¿2932; DOI 10.1007/S11695-017-2731-5. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO REPORT THE PREVALENCE, SEVERITY, AND RISK FACTORS OF LAPAROSCOPIC SLEEVE GASTRECTOMY POSTOPERATIVE HAEMORRHAGE (ACUTE AND SUBCLINICAL) IN A COHORT WITH STANDARD SETTINGS. FROM 01JAN2014 TO 31DEC2015, 394 PATIENTS UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY AND WERE ASSIGNED INTO THREE SURGICAL TECHNIQUES: GROUP A (N=137) UNDERWENT CONVENTIONAL ORDER OF STAPLING STARTING TO WHERE ECHELON FLEX BLACK CARTRIDGE OR IDRIVE ULTRA PURPLE CARTRIDGE WAS USED AND ENDING WITH BLUE OR BEIGE, GROUP B (N=145) UNDERWENT TIGHT STAPLING WITH SMALLER CLOSING STAPLE HEIGHT RELOAD (WHITE ECHELON FLEX RELOAD OR BEIGE IDRIVE ULTRA), AND GROUP C (N=115) UNDERWENT OVERSUTURING THE STAPLE LINE WITH 2¿0 VICRYL IN A CONTINUOUS BASEBALL FASHION OVER A CONVENTIONAL ORDER OF RELOADS. POSTOPERATIVE COMPLICATIONS INCLUDED 2 PATIENTS HAD LEAKS (GROUP A AND GROUP C) WHICH WERE TREATED BY SURGICAL RE-EXPLORATION AND DRAINAGE, FOLLOWED BY TPN; 2 PATIENTS HAD SLEEVE KINK AND PARTIAL OBSTRUCTION (GROUP A AND GROUP B) TREATED WITH ROUX-EN-Y GASTRIC BYPASS (N=1) AND MONTHS OF LIQUID-ONLY INTAKE (N=1); ONE PATIENT HAD ABDOMINAL WALL ADHESION (GROUP B) CAUSING UNSOLVED SEVERE PAIN WHICH WAS TREATED WITH LAPAROSCOPIC EXPLORATION; ONE PATIENT HAD PELVIC ABSCESS (GROUP A) TREATED WITH DRAINAGE; TWO PATIENTS HAD WOUND INFECTIONS (GROUP A AND GROUP B) REQUIRING BEDSIDE DRAINAGE; AND ONE PATIENT WAS DIAGNOSED WITH CLOSTRIDIUM DIFFICILE INFECTION (GROUP C). SINCE PATIENTS WITH DIABETES HAVE IMPAIRED WOUND HEALING, A DEFECTIVE STAPLE LINE HEALING, TISSUE MATURING, AND THUS FRAIL SEALING OF EMBEDDED VESSELS COULD BE THE REASON FOR EXCESSIVE POSTOPERATIVE HAEMORRHAGE IN TYPE 2 DIABETES.
09-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | ECHELON ENDOPATH STAPLER | Device Operates Differently Than Expected |  | Post Operative Wound Infection | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE: TITLE: SLEEVE GASTRECTOMY POSTOPERATIVE HEMORRHAGE IS LINKED TO TYPE-2 DIABETES AND NOT TO SURGICAL TECHNIQUE AUTHOR(S): HADAR SPIVAK, CARMIL AZRAN, GALIA SPECTRE, GALINA LIDERMANN, ORIT BLUMENFELD CITATION: OBES SURG (2017) 27:2927¿2932; DOI 10.1007/S11695-017-2731-5. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO REPORT THE PREVALENCE, SEVERITY, AND RISK FACTORS OF LAPAROSCOPIC SLEEVE GASTRECTOMY POSTOPERATIVE HAEMORRHAGE (ACUTE AND SUBCLINICAL) IN A COHORT WITH STANDARD SETTINGS. FROM 01JAN2014 TO 31DEC2015, 394 PATIENTS UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY AND WERE ASSIGNED INTO THREE SURGICAL TECHNIQUES: GROUP A (N=137) UNDERWENT CONVENTIONAL ORDER OF STAPLING STARTING TO WHERE ECHELON FLEX BLACK CARTRIDGE OR IDRIVE ULTRA PURPLE CARTRIDGE WAS USED AND ENDING WITH BLUE OR BEIGE, GROUP B (N=145) UNDERWENT TIGHT STAPLING WITH SMALLER CLOSING STAPLE HEIGHT RELOAD (WHITE ECHELON FLEX RELOAD OR BEIGE IDRIVE ULTRA), AND GROUP C (N=115) UNDERWENT OVERSUTURING THE STAPLE LINE WITH 2¿0 VICRYL IN A CONTINUOUS BASEBALL FASHION OVER A CONVENTIONAL ORDER OF RELOADS. POSTOPERATIVE COMPLICATIONS INCLUDED 2 PATIENTS HAD LEAKS (GROUP A AND GROUP C) WHICH WERE TREATED BY SURGICAL RE-EXPLORATION AND DRAINAGE, FOLLOWED BY TPN; 2 PATIENTS HAD SLEEVE KINK AND PARTIAL OBSTRUCTION (GROUP A AND GROUP B) TREATED WITH ROUX-EN-Y GASTRIC BYPASS (N=1) AND MONTHS OF LIQUID-ONLY INTAKE (N=1); ONE PATIENT HAD ABDOMINAL WALL ADHESION (GROUP B) CAUSING UNSOLVED SEVERE PAIN WHICH WAS TREATED WITH LAPAROSCOPIC EXPLORATION; ONE PATIENT HAD PELVIC ABSCESS (GROUP A) TREATED WITH DRAINAGE; TWO PATIENTS HAD WOUND INFECTIONS (GROUP A AND GROUP B) REQUIRING BEDSIDE DRAINAGE; AND ONE PATIENT WAS DIAGNOSED WITH CLOSTRIDIUM DIFFICILE INFECTION (GROUP C). SINCE PATIENTS WITH DIABETES HAVE IMPAIRED WOUND HEALING, A DEFECTIVE STAPLE LINE HEALING, TISSUE MATURING, AND THUS FRAIL SEALING OF EMBEDDED VESSELS COULD BE THE REASON FOR EXCESSIVE POSTOPERATIVE HAEMORRHAGE IN TYPE 2 DIABETES.
09-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | ECHELON ENDOPATH STAPLER | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT DURING REVIEW OF JOURNAL ARTICLE: TITLE: SLEEVE GASTRECTOMY POSTOPERATIVE HEMORRHAGE IS LINKED TO TYPE-2 DIABETES AND NOT TO SURGICAL TECHNIQUE AUTHOR(S): HADAR SPIVAK, CARMIL AZRAN, GALIA SPECTRE, GALINA LIDERMANN, ORIT BLUMENFELD CITATION: OBES SURG (2017) 27:2927¿2932; DOI 10.1007/S11695-017-2731-5. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO REPORT THE PREVALENCE, SEVERITY, AND RISK FACTORS OF LAPAROSCOPIC SLEEVE GASTRECTOMY POSTOPERATIVE HAEMORRHAGE (ACUTE AND SUBCLINICAL) IN A COHORT WITH STANDARD SETTINGS. FROM 01JAN2014 TO 31DEC2015, 394 PATIENTS UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY AND WERE ASSIGNED INTO THREE SURGICAL TECHNIQUES: GROUP A (N=137) UNDERWENT CONVENTIONAL ORDER OF STAPLING STARTING TO WHERE ECHELON FLEX BLACK CARTRIDGE OR IDRIVE ULTRA PURPLE CARTRIDGE WAS USED AND ENDING WITH BLUE OR BEIGE, GROUP B (N=145) UNDERWENT TIGHT STAPLING WITH SMALLER CLOSING STAPLE HEIGHT RELOAD (WHITE ECHELON FLEX RELOAD OR BEIGE IDRIVE ULTRA), AND GROUP C (N=115) UNDERWENT OVERSUTURING THE STAPLE LINE WITH 2¿0 VICRYL IN A CONTINUOUS BASEBALL FASHION OVER A CONVENTIONAL ORDER OF RELOADS. POSTOPERATIVE COMPLICATIONS INCLUDED 2 PATIENTS HAD LEAKS (GROUP A AND GROUP C) WHICH WERE TREATED BY SURGICAL RE-EXPLORATION AND DRAINAGE, FOLLOWED BY TPN; 2 PATIENTS HAD SLEEVE KINK AND PARTIAL OBSTRUCTION (GROUP A AND GROUP B) TREATED WITH ROUX-EN-Y GASTRIC BYPASS (N=1) AND MONTHS OF LIQUID-ONLY INTAKE (N=1); ONE PATIENT HAD ABDOMINAL WALL ADHESION (GROUP B) CAUSING UNSOLVED SEVERE PAIN WHICH WAS TREATED WITH LAPAROSCOPIC EXPLORATION; ONE PATIENT HAD PELVIC ABSCESS (GROUP A) TREATED WITH DRAINAGE; TWO PATIENTS HAD WOUND INFECTIONS (GROUP A AND GROUP B) REQUIRING BEDSIDE DRAINAGE; AND ONE PATIENT WAS DIAGNOSED WITH CLOSTRIDIUM DIFFICILE INFECTION (GROUP C). SINCE PATIENTS WITH DIABETES HAVE IMPAIRED WOUND HEALING, A DEFECTIVE STAPLE LINE HEALING, TISSUE MATURING, AND THUS FRAIL SEALING OF EMBEDDED VESSELS COULD BE THE REASON FOR EXCESSIVE POSTOPERATIVE HAEMORRHAGE IN TYPE 2 DIABETES.
19-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Leak/Splash | A0504 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DOCTOR REPORTED THE PATIENT RETURNED TO THE HOSPITAL WITH AN ANASTOMOTIC LEAK. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR TWENTY-FOUR DAYS. IT WAS NOT REPORTED IF THE PATIENT HAD A REOPERATION. NO OTHER INFORMATION IS KNOWN.
19-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DOCTOR REPORTED THE PATIENT RETURNED TO THE HOSPITAL WITH AN ANASTOMOTIC LEAK. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR TWENTY-FOUR DAYS. IT WAS NOT REPORTED IF THE PATIENT HAD A REOPERATION. NO OTHER INFORMATION IS KNOWN.
21-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | PROXIMATE*ILS CURVED CIR STAPL | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: STAPLED GASTRO/DUODENOJEJUNOSTOMY SHORTENS RECONSTRUCTION TIME DURING PYLORUS-PRESERVING PANCREATICODUODENECTOMY AUTHORS: NORIHIRO SATO, KEI YABUKI, SHIRO KOHI, YASUHISA MORI, NORITAKA MINAGAWA, TOSHIHISA TAMURA, AIICHIRO HIGURE, AND KOJI YAMAGUCHI CITATION: WORLD J GASTROENTEROL 2013 DECEMBER 28; 19(48): 9399-9404; DOI:10.3748/WJG.V19.I48.9399; PUBLISHED ONLINE: DECEMBER 28, 2013 THE AIM OF THIS STUDY WAS TO INVESTIGATE WHETHER A STAPLED TECHNIQUE WAS SUPERIOR TO THE CONVENTIONAL HAND-SEWN TECHNIQUE FOR GASTRO/DUODENOJEJUNOSTOMY DURING PYLORUS-PRESERVING PANCREATICODUODENECTOMY (PPPD). THE STUDY INCLUDED 38 PATIENTS (25 MEN AND 13 WOMEN), 19 (11 MALES AND 8 FEMALES) PATIENTS AGED 56 TO 79 YEARS WHO UNDERWENT PPPD WITH A STAPLED GASTRO/DUODENOJEJUNOSTOMY (STAPLED ANASTOMOSIS GROUP) AND 19 (14 MALES AND 5 FEMALES) PATIENTS AGED 54 TO 76 YEARS WHO UNDERWENT PPPD WITH A CONVENTIONAL HAND-SEWN DUODENOJEJUNOSTOMY (HAND-SEWN ANASTOMOSIS GROUP) BETWEEN JANUARY 2009 AND MARCH 2012. IN OCTOBER OF 2010, THE AUTHORS INTRODUCED A MECHANICAL ANASTOMOTIC TECHNIQUE OF GASTRO OR DUODENOJEJUNOSTOMY USING STAPLERS DURING PPPD. A MUCOSA-TO-MUCOSA ANASTOMOSIS OF THE PANCREATICOJEJUNOSTOMY WAS PERFORMED WITH INTERRUPTED SUTURES USING 5-0 MONOFILAMENT ABSORBABLE SUTURES (PDS, ETHICON INC., TOKYO, JAPAN). THE HEPATICOJEJUNOSTOMY WAS PERFORMED BY INTERRUPTED SUTURES USING 4-0 MONOFILAMENT ABSORBABLE SUTURES (PDS II, ETHICON INC.), AND THE BILIARY TUBE WAS PLACED FROM THE JEJUNAL LUMEN TO THE HEPATIC DUCT OF THE LIVER. THE CONVENTIONAL HAND-SEWN DUODENOJEJUNOSTOMY (END-TO-SIDE ANASTOMOSIS) WAS PERFORMED BY A TWO-LAYER ALBERT-LEMBERT METHOD (WHOLE LAYER, RUNNING SUTURES OF 4-0 PDS, AND SEROMUSCULAR LAYER, INTERRUPTED SUTURES OF 4-0 SILK). THE STAPLED GASTRO/DUODENOJEJUNOSTOMY WAS PERFORMED USING A CIRCULAR STAPLER (CDH25, ETHICON INC.). OVERALL, POSTOPERATIVE COMPLICATIONS OCCURRED IN 18 PATIENTS, INCLUDING 6 PATIENTS IN THE STAPLED ANASTOMOSIS GROUP AND 12 PATIENTS IN THE HAND-SEWN GROUP. AMONG THE COMPLICATIONS, PANCREATIC FISTULA/ANASTOMOTIC LEAKAGE OCCURRED IN 1 PATIENT IN THE STAPLED ANASTOMOSIS GROUP AND IN 3 PATIENTS IN THE HAND-SEWN ANASTOMOSIS GROUP. INTRA-ABDOMINAL ABSCESS WAS OBSERVED IN 1 PATIENT IN THE STAPLED ANASTOMOSIS GROUP AND IN 4 PATIENTS IN THE HAND-SEWN ANASTOMOSIS GROUP. NO PATIENT IN EITHER GROUP DEVELOPED LEAKAGE OF THE GASTRO/DUODENOJEJUNOSTOMY. DELAYED GASTRIC EMPTYING (DGE) WAS OBSERVED IN 1 PATIENT IN THE STAPLED ANASTOMOSIS GROUP AND IN 3 PATIENTS IN THE HAND-SEWN ANASTOMOSIS GROUP. IN CONCLUSION, THE AUTHORS¿ PRELIMINARY RESULTS SUGGEST THAT STAPLED GASTRO/DUODENOJEJUNOSTOMY WAS A FEASIBLE TECHNIQUE THAT COULD SHORTEN THE RECONSTRUCTION AND OPERATIVE TIME DURING PPPD WITHOUT INCREASING THE INCIDENCE OF POSTOPERATIVE COMPLICATIONS INCLUDING DELAYED GASTRIC EMPTYING.
21-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | PROXIMATE*ILS CURVED CIR STAPL | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: STAPLED GASTRO/DUODENOJEJUNOSTOMY SHORTENS RECONSTRUCTION TIME DURING PYLORUS-PRESERVING PANCREATICODUODENECTOMY AUTHORS: NORIHIRO SATO, KEI YABUKI, SHIRO KOHI, YASUHISA MORI, NORITAKA MINAGAWA, TOSHIHISA TAMURA, AIICHIRO HIGURE, AND KOJI YAMAGUCHI CITATION: WORLD J GASTROENTEROL 2013 DECEMBER 28; 19(48): 9399-9404; DOI:10.3748/WJG.V19.I48.9399; PUBLISHED ONLINE: DECEMBER 28, 2013 THE AIM OF THIS STUDY WAS TO INVESTIGATE WHETHER A STAPLED TECHNIQUE WAS SUPERIOR TO THE CONVENTIONAL HAND-SEWN TECHNIQUE FOR GASTRO/DUODENOJEJUNOSTOMY DURING PYLORUS-PRESERVING PANCREATICODUODENECTOMY (PPPD). THE STUDY INCLUDED 38 PATIENTS (25 MEN AND 13 WOMEN), 19 (11 MALES AND 8 FEMALES) PATIENTS AGED 56 TO 79 YEARS WHO UNDERWENT PPPD WITH A STAPLED GASTRO/DUODENOJEJUNOSTOMY (STAPLED ANASTOMOSIS GROUP) AND 19 (14 MALES AND 5 FEMALES) PATIENTS AGED 54 TO 76 YEARS WHO UNDERWENT PPPD WITH A CONVENTIONAL HAND-SEWN DUODENOJEJUNOSTOMY (HAND-SEWN ANASTOMOSIS GROUP) BETWEEN JANUARY 2009 AND MARCH 2012. IN OCTOBER OF 2010, THE AUTHORS INTRODUCED A MECHANICAL ANASTOMOTIC TECHNIQUE OF GASTRO OR DUODENOJEJUNOSTOMY USING STAPLERS DURING PPPD. A MUCOSA-TO-MUCOSA ANASTOMOSIS OF THE PANCREATICOJEJUNOSTOMY WAS PERFORMED WITH INTERRUPTED SUTURES USING 5-0 MONOFILAMENT ABSORBABLE SUTURES (PDS, ETHICON INC., TOKYO, JAPAN). THE HEPATICOJEJUNOSTOMY WAS PERFORMED BY INTERRUPTED SUTURES USING 4-0 MONOFILAMENT ABSORBABLE SUTURES (PDS II, ETHICON INC.), AND THE BILIARY TUBE WAS PLACED FROM THE JEJUNAL LUMEN TO THE HEPATIC DUCT OF THE LIVER. THE CONVENTIONAL HAND-SEWN DUODENOJEJUNOSTOMY (END-TO-SIDE ANASTOMOSIS) WAS PERFORMED BY A TWO-LAYER ALBERT-LEMBERT METHOD (WHOLE LAYER, RUNNING SUTURES OF 4-0 PDS, AND SEROMUSCULAR LAYER, INTERRUPTED SUTURES OF 4-0 SILK). THE STAPLED GASTRO/DUODENOJEJUNOSTOMY WAS PERFORMED USING A CIRCULAR STAPLER (CDH25, ETHICON INC.). OVERALL, POSTOPERATIVE COMPLICATIONS OCCURRED IN 18 PATIENTS, INCLUDING 6 PATIENTS IN THE STAPLED ANASTOMOSIS GROUP AND 12 PATIENTS IN THE HAND-SEWN GROUP. AMONG THE COMPLICATIONS, PANCREATIC FISTULA/ANASTOMOTIC LEAKAGE OCCURRED IN 1 PATIENT IN THE STAPLED ANASTOMOSIS GROUP AND IN 3 PATIENTS IN THE HAND-SEWN ANASTOMOSIS GROUP. INTRA-ABDOMINAL ABSCESS WAS OBSERVED IN 1 PATIENT IN THE STAPLED ANASTOMOSIS GROUP AND IN 4 PATIENTS IN THE HAND-SEWN ANASTOMOSIS GROUP. NO PATIENT IN EITHER GROUP DEVELOPED LEAKAGE OF THE GASTRO/DUODENOJEJUNOSTOMY. DELAYED GASTRIC EMPTYING (DGE) WAS OBSERVED IN 1 PATIENT IN THE STAPLED ANASTOMOSIS GROUP AND IN 3 PATIENTS IN THE HAND-SEWN ANASTOMOSIS GROUP. IN CONCLUSION, THE AUTHORS¿ PRELIMINARY RESULTS SUGGEST THAT STAPLED GASTRO/DUODENOJEJUNOSTOMY WAS A FEASIBLE TECHNIQUE THAT COULD SHORTEN THE RECONSTRUCTION AND OPERATIVE TIME DURING PPPD WITHOUT INCREASING THE INCIDENCE OF POSTOPERATIVE COMPLICATIONS INCLUDING DELAYED GASTRIC EMPTYING.
01-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | Fistula | Event Description: IT WAS REPORTED THAT DURING COLON SURGERY, RIGHT COLECTOMY, THE DEVICE WAS USED FOR THE ANASTOMOSIS. ON THE THIRD DAY IN THE POSTOPERATIVE PERIOD, THE PATIENT PRESENTED FISTULA. DEHISCENCE OF STAPLING IN THE MIDDLE THIRD.
12-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | FLEX 60 ARTICULATING | Positioning Failure | A150201 | No Code Available | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE FARTHER. DURING THE LAPAROSCOPIC LOW RECTAL CANCER RADICAL OPERATION, WHEN SEVERING THE RECTUM, COULD NOT FIRE FARTHER, FELT THE FIRING TRIGGER WAS LOOSE. CONSIDERING SAFETY RISK, THE SURGERY CHANGED TO SMALL INCISION OPERATION, USED CONTOUR TO COMPLETE THE SURGERY. THE SURGERY DELAYED ABOUT 40 MINUTES. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS STANDARD FOR THIS PROCEDURE AND IS NOT EXTENDED.
12-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | FLEX 60 ARTICULATING | Loose or Intermittent Connection | A1205 | No Code Available | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE FARTHER. DURING THE LAPAROSCOPIC LOW RECTAL CANCER RADICAL OPERATION, WHEN SEVERING THE RECTUM, COULD NOT FIRE FARTHER, FELT THE FIRING TRIGGER WAS LOOSE. CONSIDERING SAFETY RISK, THE SURGERY CHANGED TO SMALL INCISION OPERATION, USED CONTOUR TO COMPLETE THE SURGERY. THE SURGERY DELAYED ABOUT 40 MINUTES. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS STANDARD FOR THIS PROCEDURE AND IS NOT EXTENDED.
12-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | FLEX 60 ARTICULATING | Structural Problem | A0511 | No Code Available | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE FARTHER. DURING THE LAPAROSCOPIC LOW RECTAL CANCER RADICAL OPERATION, WHEN SEVERING THE RECTUM, COULD NOT FIRE FARTHER, FELT THE FIRING TRIGGER WAS LOOSE. CONSIDERING SAFETY RISK, THE SURGERY CHANGED TO SMALL INCISION OPERATION, USED CONTOUR TO COMPLETE THE SURGERY. THE SURGERY DELAYED ABOUT 40 MINUTES. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS STANDARD FOR THIS PROCEDURE AND IS NOT EXTENDED.
12-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | FLEX 60 ARTICULATING | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE FARTHER. DURING THE LAPAROSCOPIC LOW RECTAL CANCER RADICAL OPERATION, WHEN SEVERING THE RECTUM, COULD NOT FIRE FARTHER, FELT THE FIRING TRIGGER WAS LOOSE. CONSIDERING SAFETY RISK, THE SURGERY CHANGED TO SMALL INCISION OPERATION, USED CONTOUR TO COMPLETE THE SURGERY. THE SURGERY DELAYED ABOUT 40 MINUTES. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS STANDARD FOR THIS PROCEDURE AND IS NOT EXTENDED.
04-12-2017 | Malfunction | ethicon | 06-04-2018 | ETHICON SECURESTRAP | Failure to Fire | A050501 | No Information | Event Description: DEVICE WOULD NOT FIRE AND IT IS NOT USABLE FOR PROCEDURE. A SECOND SECURE STRAP HAD TO BE OPENED TO COMPLETE PROCEDURE.
07-02-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION, THERE WAS AN INCOMPLETE STAPLE "DONUT" INSIDE THE DEVICE. HE IDENTIFIED "INCOMPLETE DONUTS", WHICH THE FAILURE INDICATOR IN THE STAPLING INSTRUMENT, BECAUSE THE SEAL WAS NOT MADE AT 100 OF THE ANASTOMOSIS. THE HCP IS A HIGH USER OF THE PRODUCT AND HAS A LOT OF EXPERIENCE IN ITS USE. HE IDENTIFIED THE PROBLEM AT THE END OF THE ANASTOMOSIS AND COULD REPAIR THE DEFECT OF THE STAPLING USING MANUAL SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
07-02-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION, THERE WAS AN INCOMPLETE STAPLE "DONUT" INSIDE THE DEVICE. HE IDENTIFIED "INCOMPLETE DONUTS", WHICH THE FAILURE INDICATOR IN THE STAPLING INSTRUMENT, BECAUSE THE SEAL WAS NOT MADE AT 100 OF THE ANASTOMOSIS. THE HCP IS A HIGH USER OF THE PRODUCT AND AND HAS A LOT OF EXPERIENCE IN ITS USE. HE IDENTIFIED THE PROBLEM AT THE END OF THE ANASTOMOSIS AND COULD REPAIR THE DEFECT OF THE STAPLING USING MANUAL SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
09-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | GST BLUE RELOAD, 60MM, 6 ROW | Incorrect Device Or Component Shipped |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN OPENING A NEW BOX OF GST60B, IT WAS DISCOVERED THE BOX CONTAINED ALL GST60G. BOX WAS REMOVED FROM STOCK AND GIVEN TO REP. NO PATIENT CONSEQUENCE REPORTED.
07-02-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION, THERE WAS AN INCOMPLETE STAPLE "DONUT" INSIDE THE DEVICE. HE IDENTIFIED "INCOMPLETE DONUTS", WHICH THE FAILURE INDICATOR IN THE STAPLING INSTRUMENT, BECAUSE THE SEAL WAS NOT MADE AT 100 OF THE ANASTOMOSIS. THE HCP IS A HIGH USER OF THE PRODUCT AND HAS A LOT OF EXPERIENCE IN ITS USE. HE IDENTIFIED THE PROBLEM AT THE END OF THE ANASTOMOSIS AND COULD REPAIR THE DEFECT OF THE STAPLING USING MANUAL SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
06-04-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD UNKNOWN ISSUE. PATIENT OUTCOME WAS UNKNOWN.
21-03-2018 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA ULTRA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESECTION PROCEDURE, THE DEVICE'S BUTTON THAT PROVIDED THE PRODUCTION ARTICULATION BROKE OFF OR RUPTURED. IT WAS THEN REPLACED WITH A NEW ONE SO AS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER RADICAL SURGERY, THE STAPLE WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION AND AN INCOMPLETE DISTAL STAPLE LINE WHICH WAS FIXED BY HAND SEWING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER RADICAL SURGERY, THE STAPLE WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION AND AN INCOMPLETE DISTAL STAPLE LINE WHICH WAS FIXED BY HAND SEWING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER RADICAL SURGERY, THE STAPLE WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION AND AN INCOMPLETE DISTAL STAPLE LINE WHICH WAS FIXED BY HAND SEWING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER RADICAL SURGERY, THE STAPLE WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION AND AN INCOMPLETE DISTAL STAPLE LINE WHICH WAS FIXED BY HAND SEWING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2017 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING SURGICAL STAPLING DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER AND THE RELOAD DID NOT PERFORM STAPLING FROM THE THIRD SHOT. THE LOAD THAT WAS ON THE THIRD TRIP HAS STOPPED AND CANNOT BE USED AND BOTH PRODUCTS COULD NO LONGER BE USED. THE STAPLER DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE COULDN'T BE SQUEEZED. THEY OPENED ANOTHER STAPLER AND ANOTHER LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-02-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE SURGEON INFORMED THE NURSE THAT THE FIRST RELOAD ON THIS HANDLE FIRED BUT THE SECOND ONE DID NOT. THERE WAS A STRUGGLE IN PRESSING THE GREEN BUTTON BUT MANAGED TO PUSHED IT EVENTUALLY WITH A LITTLE BIT OF FORCE AND THE HANDLE WENT INTO FIRE MODE. HANDLE WAS ABLE TO SQUEEZED AND MADE A CRACKING NOISE BUT THE RELOAD DID FIRE. THE SURGEON PUT THE SECOND RELOAD ON A NEW HANDLE AND IT FIRED FINE. THERE WAS NO PATIENT INJURY.
21-02-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE SURGEON INFORMED THE NURSE THAT THE FIRST RELOAD ON THIS HANDLE FIRED BUT THE SECOND ONE DID NOT. THERE WAS A STRUGGLE IN PRESSING THE GREEN BUTTON BUT MANAGED TO PUSHED IT EVENTUALLY WITH A LITTLE BIT OF FORCE AND THE HANDLE WENT INTO FIRE MODE. HANDLE WAS ABLE TO SQUEEZED AND MADE A CRACKING NOISE BUT THE RELOAD DID FIRE. THE SURGEON PUT THE SECOND RELOAD ON A NEW HANDLE AND IT FIRED FINE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE THE DEVICE WAS BEING APPLIED ON THE STOMACH, ANTRUM AREA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS NOT ABLE TO FIRE AND THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE THE DEVICE WAS BEING APPLIED ON THE STOMACH, ANTRUM AREA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS NOT ABLE TO FIRE AND THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | covidien lp llc north haven | 06-04-2018 | SIGNIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER. DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT THE RECTUM RESECTION, THE OPENING/CLOSING OF JAWS WERE CHECKED THEN THE SURGEON TRIED TO INSERT THE DEVICE TO THE TROCAR, HOWEVER COULD NOT CLOSE THE JAWS. NO COMMUNICATION ERROR OF POWER HANDLE, ADAPTER AND RELOAD OCCURRED. THE EVENT DID NOT REPEAT AFTER SURGERY. THERE WAS NO PATIENT INJURY.
14-12-2017 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING SURGICAL STAPLING DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER AND THE RELOAD DID NOT PERFORM STAPLING FROM THE THIRD SHOT. THE LOAD THAT WAS ON THE THIRD TRIP HAS STOPPED AND CANNOT BE USED AND BOTH PRODUCTS COULD NO LONGER BE USED. THE STAPLER DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE COULDN'T BE SQUEEZED. THEY OPENED ANOTHER STAPLER AND ANOTHER LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA SULU | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, WHILE RELOADING THE DEVICE, THEY TRIED TO UNLOAD IT BUT PIECE OF THE RELOAD WAS STUCK WITH THE ADAPTER SHAFT. THE ADAPTER INDICATOR WAS FLASHING AND THE RELOAD INDICATOR WAS SOLID. REPLACED BY A NEW ADAPTED TO COMPLETE THE CASE. NO PATIENT INVOLVEMENT DURING SETUP.
20-03-2018 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA SULU | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, WHILE RELOADING THE DEVICE, THEY TRIED TO UNLOAD IT BUT PIECE OF THE RELOAD WAS STUCK WITH THE ADAPTER SHAFT. THE ADAPTER INDICATOR WAS FLASHING AND THE RELOAD INDICATOR WAS SOLID. REPLACED BY A NEW ADAPTED TO COMPLETE THE CASE. NO PATIENT INVOLVEMENT DURING SETUP.
15-03-2018 | Malfunction | covidien lp llc north haven | 06-04-2018 | EEA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLE WAS ABLE TO FIRE ONLY ABOUT 50% RESULTING TO A POOR STAPLE FORMATION. A MANUAL SUTURE WAS CONDUCTED TO COMPLETE THE CASE. NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | TA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN LOBECTOMY PROCEDURE, WHILE THE STAPLER BEING APPLIED ON THE LOBE, THE DEVICE DID NOT FIRE AND DID NOT PERFORM FULL AND COMPLETE SQUEEZES OF THE HANDLE. THE DOCTOR USED ANOTHER STAPLER TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
08-03-2017 | Malfunction | ethicon endo-surgery | 06-04-2018 | CONTOUR CURVED CUTTER STAPLER | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AFTER A PROCTOCOLECTOMY WITH LOYAL RESERVOIR PROCEDURE, THE DEVICE DID NOT STAPLE NOR CUT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-11-2017 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPY, WHILE PERFORMING SURGICAL STAPLING, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS UNABLE TO FIRE AND THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THEY OPENED ANOTHER DEVICE WITH THE SAME LOT NUMBER TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
14-11-2017 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPY, WHILE PERFORMING SURGICAL STAPLING, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS UNABLE TO FIRE AND THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THEY OPENED ANOTHER DEVICE WITH THE SAME LOT NUMBER TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
20-03-2018 | Injury | us surgical puerto rico | 06-04-2018 | TA PREMIUM | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN ANTERIOR RESECTION OF THE RECTUM PROCEDURE, WHILE CLOSING OF THE STUMP, THERE WERE 3 LOADS DID NOT CREATE A CORRECT LINE OF STAPLES ON THE INTESTINE DUE TO A LEAK IN THE INTESTINAL STUMP. THERE WERE 4 ATTEMPTS TO CLOSE THE STUMP SO ANOTHER LOAD FROM ANOTHER PACKAGING WAS USED AND SEWN MANUALLY TO RESOLVE THE ISSUE. A CREATION OF A PERMANENT COLOSTOMY WAS NEEDED TO PREVENT PERMANENT IMPAIRMENT OF A FUNCTION.
20-03-2018 | Injury | us surgical puerto rico | 06-04-2018 | TA PREMIUM | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN ANTERIOR RESECTION OF THE RECTUM PROCEDURE, WHILE CLOSING OF THE STUMP, THERE WERE 3 LOADS DID NOT CREATE A CORRECT LINE OF STAPLES ON THE INTESTINE DUE TO A LEAK IN THE INTESTINAL STUMP. THERE WERE 4 ATTEMPTS TO CLOSE THE STUMP SO ANOTHER LOAD FROM ANOTHER PACKAGING WAS USED AND SEWN MANUALLY TO RESOLVE THE ISSUE. A CREATION OF A PERMANENT COLOSTOMY WAS NEEDED TO PREVENT PERMANENT IMPAIRMENT OF A FUNCTION.
02-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE STAPLER DID NOT FIRE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | MULTIFIRE ENDO TA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LEFT DONOR NEPHRECTOMY, THE DEVICE WAS OPENED AND APPLIED TO THE TISSUE THAT HAD STUCK TO THE CARTRIDGE. THE DEVICE WORKED FINE ON THE RENAL ARTERY BUT WHEN RELOADED AND APPLIED TO THE VEIN, THE DEVICE FIRED STAPLES AND THE DEVICE WAS ABLE TO BE OPENED. BUT FOR SOME REASON, THE TISSUE STAPLED STUCK TO THE PLASTIC RELOAD AND HAD TO BE PRIED OUT BY ANOTHER ENDOSCOPIC KNITTER. UPON CLOSE INSPECTION, IT WAS DECIDED TO APPLY ANOTHER DEVICE RELOAD TO THE RENAL VEIN AND THE PROCEDURE WAS COMPLETED. ANOTHER DEVICE WAS USED TO PRY THE TISSUE OUT OF THE PLASTIC CARTRIDGE. ANOTHER DEVICE WAS OPENED AND APPLIED. NO PATIENT INJURY.
19-10-2017 | Malfunction | covidien lp llc north haven | 06-04-2018 | IDRIVE ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON SURGICAL STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
01-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: (B)(4)
20-03-2018 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS ABLE TO FIRE BUT HAD AN INCOMPLETE STAPLE LINE ON THE DISTAL PART AND IT ONLY FIRED HALFWAY THOUGH THE ENTIRETY.
14-02-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, THE FIRST FIRING WAS SUCCESSFUL, FOR THE SECOND FIRING, THE NURSE CONNECTED NEW RELOAD TO STAPLER WITH NO PROBLEM. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THE SALES REP NOTICED A FEW STAPLES OF THE PROXIMAL END WERE ALREADY FIRED EVEN THOUGH THE GREEN BUTTON WAS NOT PUSHED YET. THERE WAS NO PATIENT INJURY.
21-03-2018 | Injury | us surgical puerto rico | 06-04-2018 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LOWER ANTERIOR RESECTION, AFTER FIRING THE STAPLER, THE INSTRUMENT HANDLE COULD BE FULLY SQUEEZED AND THE DONUTS WERE COMPLETE BUT THERE WAS A POOR AND INCOMPLETE STAPLE FORMATION AND THE STAPLES WERE NOT DEPLOYED ON THE ANTERIOR SURFACE OF THE ANASTOMOSIS. THE ONE THAT DID DEPLOY WAS NOT CLOSED. THE SURGICAL TIME HAD BEEN EXTENDED FOR 90 MINUTES AND THE INCISION WAS EXTENDED. THERE WAS A TISSUE DAMAGE WHICH RESULTED FROM INCOMPLETE ANASTOMOSIS. THEY HAD TO OVER-SEW TO FIX THE STAPLE LINE. THE ANASTOMOSIS LEAKED THREE DAYS LATER AND THE PATIENT NEEDED REOPERATION AND A PERMANENT END COLOSTOMY.
21-03-2018 | Injury | us surgical puerto rico | 06-04-2018 | EEA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LOWER ANTERIOR RESECTION, AFTER FIRING THE STAPLER, THE INSTRUMENT HANDLE COULD BE FULLY SQUEEZED AND THE DONUTS WERE COMPLETE BUT THERE WAS A POOR AND INCOMPLETE STAPLE FORMATION AND THE STAPLES WERE NOT DEPLOYED ON THE ANTERIOR SURFACE OF THE ANASTOMOSIS. THE ONE THAT DID DEPLOY WAS NOT CLOSED. THE SURGICAL TIME HAD BEEN EXTENDED FOR 90 MINUTES AND THE INCISION WAS EXTENDED. THERE WAS A TISSUE DAMAGE WHICH RESULTED FROM INCOMPLETE ANASTOMOSIS. THEY HAD TO OVER-SEW TO FIX THE STAPLE LINE. THE ANASTOMOSIS LEAKED THREE DAYS LATER AND THE PATIENT NEEDED REOPERATION AND A PERMANENT END COLOSTOMY.
21-03-2018 | Injury | us surgical puerto rico | 06-04-2018 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LOWER ANTERIOR RESECTION, AFTER FIRING THE STAPLER, THE INSTRUMENT HANDLE COULD BE FULLY SQUEEZED AND THE DONUTS WERE COMPLETE BUT THERE WAS A POOR AND INCOMPLETE STAPLE FORMATION AND THE STAPLES WERE NOT DEPLOYED ON THE ANTERIOR SURFACE OF THE ANASTOMOSIS. THE ONE THAT DID DEPLOY WAS NOT CLOSED. THE SURGICAL TIME HAD BEEN EXTENDED FOR 90 MINUTES AND THE INCISION WAS EXTENDED. THERE WAS A TISSUE DAMAGE WHICH RESULTED FROM INCOMPLETE ANASTOMOSIS. THEY HAD TO OVER-SEW TO FIX THE STAPLE LINE. THE ANASTOMOSIS LEAKED THREE DAYS LATER AND THE PATIENT NEEDED REOPERATION AND A PERMANENT END COLOSTOMY.
14-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE RELOAD MISFIRED AND THE SURGEON TRIED ANOTHER RELOAD BUT STILL DID NOT WORK. THE SURGEON BROUGHT THE CUTTER OUT OF THE PATIENT AND THE CUTTER WAS FROZEN. THE SURGEON THEN GOT ANOTHER LOT NUMBER AND EVERYTHING WORKED FINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
14-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Operating System Becomes Nonfunctional | A110401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE RELOAD MISFIRED AND THE SURGEON TRIED ANOTHER RELOAD BUT STILL DID NOT WORK. THE SURGEON BROUGHT THE CUTTER OUT OF THE PATIENT AND THE CUTTER WAS FROZEN. THE SURGEON THEN GOT ANOTHER LOT NUMBER AND EVERYTHING WORKED FINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-12-2017 | Malfunction | ethicon endo-surgery | 06-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IT WAS NOT POSSIBLE TO PLACE A SECOND RECHARGE, BECAUSE WHEN REMOVING THE FIRST ONE, IT WAS LEFT IN THE RECHARGE CHANNEL ITS LOWER METAL PART. PREVENTING THE PLACEMENT OF ANOTHER RECHARGE. THAT IS, THE RECHARGE BROKE WHEN IT WAS REMOVED, LEAVING THE DEVICE UNABLE TO FIRE AGAIN. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-12-2017 | Malfunction | ethicon endo-surgery | 06-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IT WAS NOT POSSIBLE TO PLACE A SECOND RECHARGE, BECAUSE WHEN REMOVING THE FIRST ONE, IT WAS LEFT IN THE RECHARGE CHANNEL ITS LOWER METAL PART. PREVENTING THE PLACEMENT OF ANOTHER RECHARGE. THAT IS, THE RECHARGE BROKE WHEN IT WAS REMOVED, LEAVING THE DEVICE UNABLE TO FIRE AGAIN. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
13-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN OPEN COLORECTAL SURGERY, THE SURGEON USED THE DEVICE WITH RELOAD. HE STATED SURELY THAT THE STAPLE LINE HAD GOT OPENED. THE SURGEON COULD SAVE THE PATIENT'S LIFE AFTER 2 ADDITIONAL SURGERIES.
12-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING THE ANASTOMOSIS LEFT COLON RECTUM, THE STAPLER FORMATION WASN¿T COMPLETE. 1/3 OUT OF STAPLES. THEY HAD TO DO MANUAL SUTURE AND MADE THE BLUE METHYLENE CHECK. THEY ALSO HAD TO DO A ILEOSTOMY. 24 HOURS LATER THE PATIENT IS GOOD. DISTAL DONUT WAS OK BUT PROXIMAL DONUT WAS NOT COMPLETE.
12-02-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | ENDO ILS 29MM, CURVED | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE; THE STAPLER MISFIRED DURING THE ANASTOMOSIS REQUIRING THE SURGEON TO RESECT MORE BOWEL AND RE-ANASTOMOSIS THE BOWEL WITH A SECOND STAPLER. HALF OF THE ANASTOMOSIS HAD THE STAPLES AND HALF OF THE ANASTOMOSIS WAS COMPLETELY WITHOUT THE STAPLES. RE-RESECTION OF THE ANASTOMOSIS USING A COMPETITOR DEVICE AND ANOTHER ANASTOMOSIS USING EES STAPLING DEVICE. SURGEON WAS ABLE TO COMPLETE THE SURGERY BUT METHOD WAS NOT REPORTED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-02-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | ENDO ILS 29MM, CURVED | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, THE STAPLER MISFIRED. SURGEON WAS ABLE TO COMPLETE THE SURGERY BUT METHOD WAS NOT REPORTED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
13-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN RIGHT HEMICOLECTOMY, THE FIRING FORCE WAS HIGHER THAN EXPECTED AT THE 2ND FIRING, AND THE FIRING KNOB WAS BROKEN OFF AFTER FIRING FORCIBLY. THE DEVICE WAS USED ON THE COLON TO CLOSE INSERTION SITE FOR FEEA. THE TARGET TISSUE WAS NOT TOO THICK OR HARD. NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN PROCEDURE, THE FIRING KNOB WAS BROKEN OFF WHEN THE SURGEON PUSHED THE KNOB FORCIBLY SINCE THE FIRING KNOB GOT STUCK ON THE WAY OF THE FIRST FIRING. NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | 75MM SELECTABLE NEW TLC | Detachment Of Device Component |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN PROCEDURE, THE FIRING KNOB WAS BROKEN OFF WHEN THE SURGEON PUSHED THE KNOB FORCIBLY SINCE THE FIRING KNOB GOT STUCK ON THE WAY OF THE FIRST FIRING. NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
04-03-2018 | Injury | ethicon endo-surgery | 06-04-2018 | PROXIMATE*75MM LINEAR CUTTER | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A COLECTOMY PROCEDURE ON THE RIGHT SIDE ((B)(6), 2018). A STAPLER AND 2 RELOADS WERE USED IN ADDITION TO THE ONE THAT WAS ALREADY IN THE STAPLER (3 RELOADS TOTAL). THE PATIENT PRESENTED SWELLING IN THE BELLY AND A CLINICAL WORSENING. A REOPERATION WAS PERFORMED (REOPENING ON (B)(6) 2018). IT WAS FOUND THAT THE STAPLE LINE WAS BROKEN. THERE WAS A TOTAL OPENING OF ALL STAPLES, WITH EXTRAVASATION OF FECES BY THE ABDOMEN OF THE PATIENT. PRESENTED PERITONITIS. STAYED 2 DAYS AT CTI ON (B)(6) AND PATIENT IS IN THE ROOM IN RECOVERY.
12-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | 75MM SELECTABLE NEW TLC | Difficult To Position |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE DISTAL GASTRIC CANCER SURGERY, WHEN SEVERING THE DUODENUM, COULD NOT FIRE AND THE FIRING KNOB COULD NOT BE RETURNED TO THE HOME POSITION. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE DISTAL GASTRIC CANCER SURGERY, WHEN SEVERING THE DUODENUM, COULD NOT FIRE AND THE FIRING KNOB COULD NOT BE RETURNED TO THE HOME POSITION. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC ASSISTED LEFT HEMICOLON SURGERY, WHEN 4TH FIRING, THE FIRING KNOB WAS BROKEN. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ALL THE PIECES WERE BE DISCARDED BY HOSPITAL.
14-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTED, DURING SLEEVE GASTRECTOMY PROCEDURE, THE RELOADS DID NOT FIRE AFTER PRESSING THE GREEN BUTTON OF THE STAPLER. IMMEDIATELY THE PRODUCTS WERE REPLACED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INJURY.
12-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | ILS 29MM, CURVED | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS HARDER THAN USUAL TO REMOVE THE ANVIL FROM THE DEVICE. MOREOVER, THE ANVIL AND THE HOUSING COULD NOT BE RE-ATTACHED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
14-03-2018 | Malfunction | ethicon endo-surgery | 06-04-2018 | PROXIMATE*ILS CURVED CIR STAPL | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP SIGMA PROCEDURE, MALFORMED STAPLES ON ONE TENTH OF THE STAPLE LINE. SURGEON CLOSED THIS PART BY HAND SUTURING. SURGERY SUCCESSFULLY FINISHED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY, AFTER LOADING THE RELOAD INTO THE STAPLER AND AFTER PRESSING THE GREEN BUTTON TO FIRE, THE RELOAD DID NOT FIRE. THERE WAS NO PATIENT INJURY.
21-02-2018 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON SURGICAL STAPLING, THE SURGEON WAS ABLE TO PRESS TO GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE LOAD DID NOT FIRE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY, AFTER LOADING THE RELOAD INTO THE STAPLER AND AFTER PRESSING THE GREEN BUTTON TO FIRE, THE RELOAD DID NOT FIRE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY, AFTER LOADING THE RELOAD INTO THE STAPLER AND AFTER PRESSING THE GREEN BUTTON TO FIRE, THE RELOAD DID NOT FIRE. THERE WAS NO PATIENT INJURY.
19-10-2017 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON SURGICAL STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER LOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-02-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, WHEN PERFORMING SURGICAL STAPLING, THE STAPLER JAMMED AND THE LOAD DID NOT FIRE. AS STATED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY OPENED ANOTHER NEW STAPLER AND NEW LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-02-2018 | Malfunction | us surgical puerto rico | 06-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY PROCEDURE, WHILE DOING SURGICAL STAPLING, THE STAPLER WAS READY TO FIRE, BUT IT WOULD OPEN BY ITSELF AND CAN NOT BE USED. AS STATED, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. NEW PRODUCT WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-02-2018 | Malfunction | covidien lp llc north haven | 06-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY PROCEDURE, WHILE DOING SURGICAL STAPLING, THE STAPLER WAS READY TO FIRE, BUT IT WOULD OPEN BY ITSELF AND CAN NOT BE USED. AS STATED, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. NEW PRODUCT WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THE DEVICE FIRED AND CUT MOST OF THE WAY, BUT NOT ALL. ABOUT 1/4 OF THE STAPLES DEPLOYED, BUT WERE NOT FORMED CORRECTLY. IT WAS NOT KNOWN HOW THEY APPEARED. THEY WERE LEFT IN TISSUE. MOST STAPLES WERE STILL IN THE UNKNOWN RELOAD. THE PROCEDURE WAS COMPLETED USING A LIKE DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
07-02-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND THE LOAD DID NOT FIRE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE FIRST FIRING OF EMERGENT APPENDECTOMY, THE SURGEON PUSHED THE GREEN BUTTON, HOWEVER COULD NOT SQUEEZE THE HANDLE AND COULD NOT FIRE THE DEVICE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AFTER FIRING, SURGEON FOUND STAPLES B-SHAPE WAS NOT COMPLETED
20-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/PNEUMOTHORAX PROCEDURE, AT THE FOURTH FIRING OF BULLA RESECTION OF THE LUNGS, THE SURGEON CLAMPED THE PULMONARY PARENCHYMA, HOWEVER, THE JAWS OPENED ON THEIR OWN. THE SURGEON RE-CLAMPED THE TISSUE BUT THE SAME EVENT OCCURRED. THE SURGEON SAID THE TISSUE MIGHT HAVE BEEN SO THICK THAT THE JAWS COULD NOT CLOSE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION SURGERY, THE SURGEON CLAMPED THE RECTUM AND CLOSED THE JAWS, TRIED TO FIRE AND PRESS THE GREEN BUTTON BUT THE BUTTON DID NOT LIGHT UP, AN ERROR SOUND WAS HEARD AND THE DEVICE DID NOT FIRE. THE TISSUE WAS NOT THICK, EXCESSIVE FORCE INDICATOR WAS NOT SHOWN AND ALL OF THE OTHER INDICATORS WERE GREEN. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | SIGNIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION SURGERY, THE SURGEON CLAMPED THE RECTUM AND CLOSED THE JAWS, TRIED TO FIRE AND PRESS THE GREEN BUTTON BUT THE BUTTON DID NOT LIGHT UP, AN ERROR SOUND WAS HEARD AND THE DEVICE DID NOT FIRE. THE TISSUE WAS NOT THICK, EXCESSIVE FORCE INDICATOR WAS NOT SHOWN AND ALL OF THE OTHER INDICATORS WERE GREEN. THERE WAS NO PATIENT INJURY.
14-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: THE DEVICE WAS RECEIVED WITH NO COMPLAINT INFORMATION. ATTEMPTS WERE MADE TO RECONCILE THE DEVICE WITH NO SUCCESS RESULTING IN THE CREATION OF A NEW COMPLAINT.
21-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS ABLE TO FIRE BUT THE CLAMP HAD JAMMED AT THE TIME OF STAPLING. THE GREEN BUTTON WAS ABLE TO BE PRESSED, BUT THE HANDLE CANNOT BE SQUEEZED. THE PATIENT STATUS WAS REPORTED AS OK.
02-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Sticking |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE WOULD NOT RELEASE. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | SIGNIA | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, MULTIPLE ISSUES HAPPENED WITH DIFFERENT DEVICES. TWO ADAPTERS CAN'T BE LOADED ONTO THE HANDLE. ONE OF WHICH WAS BROKEN. RESOLUTION ON SOLVING THE ISSUE WITH THE ADAPTERS WAS NOT STATED. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. THE HANDLES THAT CAN'T RECOGNIZE RELOAD AND DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | SIGNIA | Device Displays Incorrect Message | A090201 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, MULTIPLE ISSUES HAPPENED WITH DIFFERENT DEVICES. TWO ADAPTERS CAN'T BE LOADED ONTO THE HANDLE. ONE OF WHICH WAS BROKEN. RESOLUTION ON SOLVING THE ISSUE WITH THE ADAPTERS WAS NOT STATED. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. THE HANDLES THAT CAN'T RECOGNIZE RELOAD AND DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
14-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE. THEY OPENED A NEW RELOAD TO COMPLETE THE CASE. NO PATIENT INJURY.
06-07-2017 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE LOADING THE DEVICE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED BY ANOTHER DIFFERENT BATCH OF STAPLER. NO PATIENT INJURY.
07-02-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE LOAD DID NOT FIRE. A NEW LOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS GREEN BUTTON BUT THE LOADER DID NOT FIRE.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | SIGNIA | Failure to Sense | A070908 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, MULTIPLE ISSUES HAPPENED WITH DIFFERENT DEVICES. FIRSTLY, 2 ADAPTERS THAT CAN'T BE LOADED ONTO THE HANDLE. ONE OF WHICH WAS BROKEN. RESOLUTION ON SOLVING THE ISSUE WITH THE ADAPTERS WAS NOT STATED. AND LASTLY, THE HANDLE THAT CAN'T RECOGNIZE RELOAD AND DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INVOLVEMENT.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | SIGNIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, MULTIPLE ISSUES HAPPENED WITH DIFFERENT DEVICES. FIRSTLY, 2 ADAPTERS THAT CAN'T BE LOADED ONTO THE HANDLE. ONE OF WHICH WAS BROKEN. RESOLUTION ON SOLVING THE ISSUE WITH THE ADAPTERS WAS NOT STATED. AND LASTLY, THE HANDLE THAT CAN'T RECOGNIZE RELOAD AND DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INVOLVEMENT.
21-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS GREEN BUTTON BUT THE LOADER DID NOT FIRE.
01-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER , DURING SURGICAL STAPLING ON A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE STAPLER STOPPED AFTER THE FIRST SHOT. IT DID NOT WORK ANYMORE AND HAD TO BE CHANGED WITH A NEW STAPLER. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT COLECTOMY PROCEDURE, WHILE ON RESECTION OF THE BOWEL, THE DEVICE WAS PARTIALLY FIRE AND THIS RESULTED TO AN INCOMPLETE STAPLE LINE. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
20-02-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, WHEN PERFORMING SURGICAL STAPLING, THE STAPLER JAMMED AND THE LOAD DID NOT FIRE. AS STATED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY OPENED ANOTHER NEW STAPLER AND NEW LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ECHELON FLEX 45 STANDARD | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE FARTHER. DURING THE ESOPHAGUS CANCER SURGERY, COULD NOT FIRE AFTER FIRED 3/4 ON THE 2ND FIRING, OPENED THE JAW, FOUND THE STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE FARTHER. DURING THE ESOPHAGUS CANCER SURGERY, COULD NOT FIRE AFTER FIRED 3/4 ON THE 2ND FIRING, OPENED THE JAW, FOUND THE STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-01-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Sticking |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE STOPPED WORKING DURING THE FIRING. THE STAPLER STUCK. IT WAS A LIVER RESECTION AND THEY USED THE DEVICE FOR LIVER VENA. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Device Inoperable |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE STOPPED WORKING DURING THE FIRING. THE STAPLER STUCK. IT WAS A LIVER RESECTION AND THEY USED THE DEVICE FOR LIVER VENA. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
08-07-2017 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY PROCEDURE, WHILE STAPLING, THE STAPLER WAS UNABLE TO FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. AFTER PRESSING THE GREEN BUTTON THE SURGEON ATTEMPTED TO FIRE THE CHARGE, BUT THE HANDLE MADE A LOUD NOISE AND BECAME ANNOYING. HE OPENED ANOTHER NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-07-2017 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY PROCEDURE, WHILE STAPLING, THE STAPLER WAS UNABLE TO FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. AFTER PRESSING THE GREEN BUTTON THE SURGEON ATTEMPTED TO FIRE THE CHARGE, BUT THE HANDLE MADE A LOUD NOISE AND BECAME ANNOYING. HE OPENED ANOTHER NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-06-2016 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE STAPLING, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON. THE LOAD CAUGHT AND DID NOT WORK. THEY OPENED ANOTHER LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Activation, Positioning or Separation Problem | A15 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY ON THE SECOND FIRING OF THE DEVICE IT MADE AN UNUSUAL NOISE AND ONLY TWO OF THE THREE ROWS OF STAPLES DEPLOYED. THE PROCEDURE WAS COMPLETED WITH THIS DEVICE. THE SURGEON OVER SEWED THE AREA WHERE THE STAPLES WERE MISSING. THERE WERE NO PATIENT CONSEQUENCES REPORTED
14-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY ON THE SECOND FIRING OF THE DEVICE IT MADE AN UNUSUAL NOISE AND ONLY TWO OF THE THREE ROWS OF STAPLES DEPLOYED. THE PROCEDURE WAS COMPLETED WITH THIS DEVICE. THE SURGEON OVER SEWED THE AREA WHERE THE STAPLES WERE MISSING. THERE WERE NO PATIENT CONSEQUENCES REPORTED
22-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SURGERY LOBECTOMY PROCEDURE, PART OF THE RELOAD HAD NOT BEEN FULLY FIRED AND SOME PART OF THE STAPLES REMAINED IN THE RELOAD. PATIENT OUTCOME WAS UNKNOWN.
22-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE FIRST STAPLER WAS ABLE TO FIRE HOWEVER, THERE WERE POOR STAPLE FORMATION. AND THE SECOND STAPLER WAS UNABLE TO FIRE. AS STATED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY REPLACED WITH NEW STAPLER IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/PNEUMOTHORAX PROCEDURE, WHILE AT THE BULLA RESECTION OF THE LUNG, THE GREEN BUTTON WAS PUSHED AND THE HANDLE WAS STARTED TO BE SQUEEZED, HOWEVER THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SURGERY LOBECTOMY PROCEDURE, PART OF THE RELOAD HAD NOT BEEN FULLY FIRED AND SOME PART OF THE STAPLES REMAINED IN THE RELOAD. PATIENT OUTCOME WAS UNKNOWN.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SURGERY LOBECTOMY PROCEDURE, PART OF THE RELOAD HAD NOT BEEN FULLY FIRED AND SOME PART OF THE STAPLES REMAINED IN THE RELOAD. PATIENT OUTCOME WAS UNKNOWN.
20-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE FIRST RELOAD HAD AN UNLOADING PROBLEM WHEREIN THE JOINT WAS BROKEN. WHILE THE SECOND AND THIRD RELOAD WERE ABLE TO FIRE BUT THERE WERE POOR STAPLE FORMATION AND WERE NOT SMOOTH DURING THE STAPLER FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE OPERATION OF RIGHT UPPER LOBECTOMY, WHEN THE DOCTOR DISCONNECTED PULMONARY LOBECTOMY, THE DOCTOR WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE STAPLER WAS UNABLE TO FIRE. THE JAWS WAS OPENED AND WAS REPLACED WITH A NEW LOAD BUT IT WAS STILL UNABLE TO FIRE. AFTER THE REPLACING OF A NEW STAPLER AND LOAD, IT WAS FIRED NORMALLY. THE OPERATION WAS COMPLETED SMOOTHLY, AND THERE WAS NO EFFECT TO THE PATIENT.
20-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/PNEUMOTHORAX PROCEDURE, WHILE AT THE BULLA RESECTION OF THE LUNG, THE GREEN BUTTON WAS PUSHED AND THE HANDLE WAS STARTED TO BE SQUEEZED, HOWEVER THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
08-07-2017 | Malfunction | covidien north haven - mfg | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESCTOSIGMOIDECTOMY PROCEDURE, WHILE STAPLING, THE STAPLER WAS UNABLE TO FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. AFTER PRESSING THE GREEN BUTTON THE SURGEON ATTEMPTED TO FIRE THE CHARGE, BUT THE HANDLE MADE A LOUD NOISE AND BECAME ANNOYING. HE OPENED ANOTHER NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-07-2017 | Malfunction | covidien north haven - mfg | 09-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESCTOSIGMOIDECTOMY PROCEDURE, WHILE STAPLING, THE STAPLER WAS UNABLE TO FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. AFTER PRESSING THE GREEN BUTTON THE SURGEON ATTEMPTED TO FIRE THE CHARGE, BUT THE HANDLE MADE A LOUD NOISE AND BECAME ANNOYING. HE OPENED ANOTHER NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-06-2016 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE STAPLING, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON. THE LOAD CAUGHT AND DID NOT WORK. THEY OPENED ANOTHER LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE BEING APPLIED THE STAPLING OF THE STOMACH, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON. THE HANDLE HAD ALSO BROKEN OFF. THERE WAS NO PATIENT INJURY.
13-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE BEING APPLIED THE STAPLING OF THE STOMACH, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON. THE HANDLE HAD ALSO BROKEN OFF. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT COLECTOMY PROCEDURE, WHILE ON RESECTION OF THE BOWEL, THE DEVICE WAS PARTIALLY FIRE AND THIS RESULTED TO AN INCOMPLETE STAPLE LINE. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC LIVER RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE RELOAD WAS LOADED ON TO THE HANDLE, AFTER REMOVING THE SHIPPING WEDGE, THE STAPLER WOULD NOT FIRE. ANOTHER DEVICE WAS USED WITH NO ISSUES. THERE WAS NO PATIENT INJURY.
23-01-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY PROCEDURE, ON THE COLON SECTION, THE DEVICE¿S KNIFE JAMMED MID-FIRE AND DID NOT CUT THE TISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Injury | us surgical puerto rico | 09-04-2018 | GIA | Failure to Fire | A050501 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE STAPLER FIRED ONCE THEN JAMMED AND WOULD NOT WORK AFTERWARD. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THE PATIENT LOSS ABOUT 500 CC OR MORE DUE TO THE PRODUCT PROBLEM.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY, FOR THE FIRST FIRING OF GASTRIC TUBE CONSTRUCTION, THE DEVICE WAS SET AND OPENING/ CLOSING WERE CHECKED WITH NO PROBLEM. THEN, THE SURGEON TRIED TO USE THE DEVICE; HOWEVER, THE JAWS COULD NOT CLOSE PROPERLY. THE CARTRIDGE WAS CHECKED AND NOTED THAT IT WAS PRE-FIRED. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY, FOR THE FIRST FIRING OF GASTRIC TUBE CONSTRUCTION, THE DEVICE WAS SET AND OPENING/ CLOSING WERE CHECKED WITH NO PROBLEM. THEN, THE SURGEON TRIED TO USE THE DEVICE; HOWEVER, THE JAWS COULD NOT CLOSE PROPERLY. THE CARTRIDGE WAS CHECKED AND NOTED THAT IT WAS PRE-FIRED. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
13-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ILA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING OF GASTROENTEROSTOMY (OPEN), THE NURSE FELT STRANGE WHEN HE/SHE CONNECTED THE CARTRIDGE TO THE INSTRUMENT. THE SURGEON SAID PLASTIC PART WAS SO SMALL THAT COULD NOT CONNECT THE CARTRIDGE TO THE INSTRUMENT PROPERLY. THE SURGEON TRIED TO FIRE THE DEVICE IN QUESTION; HOWEVER, IT DID NOT WORK SMOOTHLY. THE CARTRIDGE WAS NOT USABLE. REPLACED BY NEW CARTRIDGE; HOWEVER, THE SAME EVENT OCCURRED. REPLACED BY ANOTHER NEW CARTRIDGE, THE FIRING WAS COMPLETED. THE STATUS OF THE PATIENT WAS REPORTED AS NO PROBLEM.
13-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ILA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING OF GASTROENTEROSTOMY (OPEN), THE NURSE FELT STRANGE WHEN HE/SHE CONNECTED THE CARTRIDGE TO THE INSTRUMENT. THE SURGEON SAID PLASTIC PART WAS SO SMALL THAT COULD NOT CONNECT THE CARTRIDGE TO THE INSTRUMENT PROPERLY. THE SURGEON TRIED TO FIRE THE DEVICE IN QUESTION; HOWEVER, IT DID NOT WORK SMOOTHLY. THE CARTRIDGE WAS NOT USABLE. REPLACED BY NEW CARTRIDGE; HOWEVER, THE SAME EVENT OCCURRED. REPLACED BY ANOTHER NEW CARTRIDGE, THE FIRING WAS COMPLETED. THE STATUS OF THE PATIENT WAS REPORTED AS NO PROBLEM.
14-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | PROXIMATE*ILS CURVED CIR STAPL | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP SIGMA PROCEDURE, THERE WERE MALFORMED STAPLES ON ONE TENTH OF THE STAPLE LINE. SURGEON CLOSED THIS PART BY HAND SUTURING. SURGERY SUCCESSFULLY FINISHED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GENERAL GASTROINTESTINAL SURGERY RADICAL OPERATION OF GASTRIC CANCER, AFTER NORMALLY INSTALLING A NEW LOAD TWICE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE SQUEEZE THE HANDLE SO THE STAPLER WAS UNABLE TO FIRE. A NEW DEVICE WAS USED TO OPERATE NORMALLY. THERE WAS NO PATIENT INJURY.
14-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ILS 29MM, CURVED | Component Missing | A020602 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE COLONIC OPERATION, AFTER FIRED, FOUND NO STAPLE BUT WITH CUT LINE. SUTURE WAS USED TO SEW THE WOUND. THE SURGERY DELAYED ABOUT 30 MINUTES. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS STANDARD FOR THIS PROCEDURE AND IS NOT EXTENDED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | EGIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SIGMOIDECTOMY, THE SURGEON STARTED TO SQUEEZE THE HANDLE HOWEVER COULD NOT FIRE THE DEVICE. REPLACED THE DEVICE WITH A NEW CARTRIDGE AND CONNECTED TO THE SAME HANDLE, AND THE FIRING WAS COMPLETED.
22-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE OPERATION OF RIGHT UPPER LOBECTOMY , WHEN THE DOCTOR DISCONNECTED PULMONARY LOBECTOMY, THE DOCTOR WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE STAPLER WAS UNABLE TO FIRE. THE JAWS WAS OPENED AND WAS REPLACED WITH A NEW LOAD BUT IT WAS STILL UNABLE TO FIRE. AFTER THE REPLACING OF A NEW STAPLER AND LOAD, IT WAS FIRED NORMALLY . THE OPERATION WAS COMPLETED SMOOTHLY , AND THERE WAS NO EFFECT TO THE PATIENT.
21-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER NORMALLY INSTALLING A NEW LOAD TWICE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE SQUEEZE THE HANDLE SO THE STAPLER WAS UNABLE TO FIRE. A NEW DEVICE WAS USED TO OPERATE NORMALLY. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE STAPLING OF STOMACH THE DEVICE DID NOT FIRE, THE GREEN BUTTON WAS DEPRESSED BUT THE HANDLE WOULD NOT ADVANCE THE I-BEAM. THEY ATTEMPTED TO FIRE A RELOAD WITH THE STAPLER HOWEVER, IT HAD FAILED TO FIRE AGAIN. ANOTHER STAPLER WAS OPENED AND THE RELOADS FIRED AS INTENDED AND THE PROCEDURE CONTINUED. THERE WAS NO PATIENT INJURY.
12-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE DID NOT RETURN AFTER THE KNIFE MOVED FORWARD ALL THE WAY OR ON THE WAY AT THE 2ND FIRING. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE SINCE THE KNIFE REVERSE SWITCH DID NOT FUNCTION. THE STAPLES WERE FORMED PROPERLY. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
15-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN HIGH ANTERIOR RESECTION, THE FIRING HANDLE COULD NOT BE FULLY GRASPED AND NO BREAKING SOUND WAS HEARD AT THE FIRST FIRING. THE DEVICE WAS USED FOR TRANSANAL DST ANASTOMOSIS BETWEEN COLON AND RECTUM. THE GAP SETTING SCALE MARKED 2.0 MM. THE TARGET TISSUE WAS NOT CUT FULLY, AND SOME STAPLES WERE UNFORMED. THE WASHER WAS REMAINED UNBROKEN IN THE ANVIL SIDE OF THE DEVICE. THE SURGEON INSERTED THE DEVICE UPSIDE DOWN. THE INCISED WOUND WAS EXPANDED, AND THE ANASTOMOSED SITE WAS RECUT, AND ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
15-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ILS 25MM, CURVED | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN HIGH ANTERIOR RESECTION, THE FIRING HANDLE COULD NOT BE FULLY GRASPED AND NO BREAKING SOUND WAS HEARD AT THE FIRST FIRING. THE DEVICE WAS USED FOR TRANSANAL DST ANASTOMOSIS BETWEEN COLON AND RECTUM. THE GAP SETTING SCALE MARKED 2.0 MM. THE TARGET TISSUE WAS NOT CUT FULLY, AND SOME STAPLES WERE UNFORMED. THE WASHER WAS REMAINED UNBROKEN IN THE ANVIL SIDE OF THE DEVICE. THE SURGEON INSERTED THE DEVICE UPSIDE DOWN. THE INCISED WOUND WAS EXPANDED, AND THE ANASTOMOSED SITE WAS RECUT, AND ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE FIRST STAPLER WAS ABLE TO FIRE HOWEVER, THERE WERE POOR STAPLE FORMATION. AND THE SECOND STAPLER WAS UNABLE TO FIRE. AS STATED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY REPLACED WITH NEW STAPLER IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC LIVER RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE RELOAD WAS LOADED ON TO THE HANDLE, AFTER REMOVING THE SHIPPING WEDGE, THE STAPLER WOULD NOT FIRE. ANOTHER DEVICE WAS USED WITH NO ISSUES. THERE WAS NO PATIENT INJURY.
05-02-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | GOLD RELOAD FOR ECHELON 45 | Material Protrusion/Extrusion | A0411 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT STAPLES WERE PROTRUDING. DURING THE RESECTION OF PULMONARY NODULES, AFTER INSERT IN THE PATIENT, BEFORE FIRING, FOUND THE STAPLES PROTRUDING, BUT DID NOT FALL OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
22-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, THE DEVICE WAS UNABLE TO FIRE. WHEN THE SURGEON USED THE DEVICE AND CLAMPED IT, THEY TRIED TO SQUEEZE THE HANDLE OF THE DEVICE FOR FIRING ITS STAPLES AND TRANSECTING THE SPOT. AFTER THE GREEN BUTTON WAS PRESSED, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE¿S CARTRIDGE STAPLE AND KNIFE DID NOT FIRE. WHEN THE TRIGGER WAS PULLED, THE SURGEON HEARD SOMETHING BROKEN. NO PATIENT INJURY.
22-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, THE DEVICE WAS UNABLE TO FIRE. WHEN THE SURGEON USED THE DEVICE AND CLAMPED IT, THEY TRIED TO SQUEEZE THE HANDLE OF THE DEVICE FOR FIRING ITS STAPLES AND TRANSECTING THE SPOT. AFTER THE GREEN BUTTON WAS PRESSED, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE¿S CARTRIDGE STAPLE AND KNIFE DID NOT FIRE. WHEN THE TRIGGER WAS PULLED, THE SURGEON HEARD SOMETHING BROKEN. NO PATIENT INJURY.
21-02-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON SURGICAL STAPLING, THE SURGEON WAS ABLE TO PRESS TO GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE LOAD DID NOT FIRE. THERE WAS NO PATIENT INJURY.
02-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, BEFORE FIRING, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON. THE SURGEON TOOK ANOTHER LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION SURGERY, THE SURGEON CLAMPED THE RECTUM AND CLOSED THE JAWS, TRIED TO FIRE AND PRESS THE GREEN BUTTON, BUT THE BUTTON DID NOT LIGHT UP, AN ERROR SOUND WAS HEARD AND THE DEVICE DID NOT FIRE. THE TISSUE WAS NOT THICK, EXCESSIVE FORCE INDICATOR WAS NOT SHOWN AND ALL OF THE OTHER INDICATORS WERE GREEN. ANOTHER HANDLE, ADAPTER AND A NEW SHELL WERE OPENED AND WORKED FINE WITH THE SAME RELOAD. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS ABLE TO FIRE BUT THE CLAMP HAD JAMMED AT THE TIME OF STAPLING. THE GREEN BUTTON WAS ABLE TO BE PRESSED, BUT THE HANDLE CANNOT BE SQUEEZED. THE PATIENT STATUS WAS REPORTED AS FINE.
14-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ECHELON FLEX 60 COMPACT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE JAW COULD NOT OPEN. DURING THE ENDOSCOPIC GASTRIC SURGERY, AFTER FIRED WITH ECR60D, COULD NOT OPEN THE JAW. AS THE TISSUE WAS CUT OFF, REMOVED THE DEVICE FROM THE PATIENT WITH CLOSED JAW. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
14-03-2018 | Injury | ethicon endo-surgery | 09-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE THE SURGEON FOUND A 1 CM HOLE BELOW THE ANASTAMOSIS AFTER COMPLETING THE FIRING. THE CUSTOMER WAS NOT SURE IF THE DEVICE STAPLED, BUT THE CUSTOMER DID CONFIRM THAT THERE WERE NO STAPLES LEFT IN THE DEVICE AND THAT IT HAD CUT PROPERLY AS THERE WERE TWO PROPERLY SIZED AND SHAPED TISSUE DONUTS IN THE DEVICE. IT WAS NOTED THAT THE PATIENT HAS EXTENSIVE ADHESIONS AND THAT THE SIGMOID WAS ADHERED TO THE SIDE WALL, SO TO CORRECT THE HOLE AND CONTINUE THE PROCEDURE THE CUSTOMER CONVERTED TO AN OPEN PROCEDURE. PROCEDURE WAS THEN COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE.
13-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | BLUE RELOAD FOR ECHELON 60 | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THERE WAS CARTRIDGE BREAKAGE. DURING THE LOBECTOMY, WHEN SEVERING THE LOBE OF THE LUNG, COULD NOT FIRE FARTHER AFTER FIRED A HALF ON THE SECOND FIRING. THE METAL PIECE OF THE CARTRIDGE WAS LOOSE, AND THE BLADE PULLED THE STAPLE DRIVER ABOUT OF THE SLOT. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
13-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THERE WAS CARTRIDGE BREAKAGE. DURING THE LOBECTOMY, WHEN SEVERING THE LOBE OF THE LUNG, COULD NOT FIRE FARTHER AFTER FIRED A HALF ON THE SECOND FIRING. THE METAL PIECE OF THE CARTRIDGE WAS LOOSE, AND THE BLADE PULLED THE STAPLE DRIVER ABOUT OF THE SLOT. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
14-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Device Inoperable |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLER STARTED ALONE. AFTER CARTRIDGE IS CHANGE IT DOES NOT WORK. ANOTHER DEVICE WAS USED. NO PATIENT CONSEQUENCE.
09-01-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | 75MM SELECTABLE NEW TLC | Device Displays Incorrect Message | A090201 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A WHIPPLE PROCEDURE, THE DEVICE WAS ERROR AT THE JOINT POINT. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
13-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ILA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING OF GASTROENTEROSTOMY (OPEN), THE NURSE FELT STRANGE WHEN HE/SHE CONNECTED THE CARTRIDGE TO THE INSTRUMENT. THE SURGEON SAID PLASTIC PART WAS SO SMALL THAT COULD NOT CONNECT THE CARTRIDGE TO THE INSTRUMENT PROPERLY. THE SURGEON TRIED TO FIRE THE DEVICE IN QUESTION; HOWEVER, IT DID NOT WORK SMOOTHLY. THE CARTRIDGE WAS NOT USABLE. REPLACED BY NEW CARTRIDGE; HOWEVER, THE SAME EVENT OCCURRED. THE STATUS OF THE PATIENT WAS REPORTED AS NO PROBLEM.
13-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ILA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING OF GASTROENTEROSTOMY (OPEN), THE NURSE FELT STRANGE WHEN HE/SHE CONNECTED THE CARTRIDGE TO THE INSTRUMENT. THE SURGEON SAID PLASTIC PART WAS SO SMALL THAT COULD NOT CONNECT THE CARTRIDGE TO THE INSTRUMENT PROPERLY. THE SURGEON TRIED TO FIRE THE DEVICE IN QUESTION; HOWEVER, IT DID NOT WORK SMOOTHLY. THE CARTRIDGE WAS NOT USABLE. REPLACED BY NEW CARTRIDGE; HOWEVER, THE SAME EVENT OCCURRED. THE STATUS OF THE PATIENT WAS REPORTED AS NO PROBLEM.
22-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ESOPHAGECTOMY, THE DEVICE DID NOT FIRE. THEY USED ANOTHER DEVICE AND FOUND IT SMOOTH AND THE CASE HAS BEEN COMPLETED. THERE WAS NO PATIENT INJURY.
16-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, THE SURGEON FIRED THE STAPLER ON THE LIVER PART TO CUT OUT. HE PUSHED THE LEVER AND IT WAS STUCK AT THE APPROXIMATELY 4/5 POINT OF THE PATH. HE PUT MORE POWER TO FINISH IT AND THE LEVER WAS BROKEN AND DETACHED FROM THE BODY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | 75MM SELECTABLE NEW TLC | Detachment Of Device Component |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, THE SURGEON FIRED THE STAPLER ON THE LIVER PART TO CUT OUT. HE PUSHED THE LEVER AND IT WAS STUCK AT THE APPROXIMATELY 4/5 POINT OF THE PATH. HE PUT MORE POWER TO FINISH IT AND THE LEVER WAS BROKEN AND DETACHED FROM THE BODY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | ILS 25MM, CURVED | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE TARGET TISSUE COULD NOT BE ANASTOMOSED THOUGH A BREAKING SOUND WAS HEARD. THE ANVIL WAS ATTACHED TO THE HOUSING PROPERLY. IT WAS FOUND THAT THE WASHER COULD NOT BE PUNCHED OUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
22-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SIGMOIDECTOMY, THE SURGEON STARTED TO SQUEEZE THE HANDLE HOWEVER COULD NOT FIRE THE DEVICE. THERE WAS NO PATIENT INVOLVEMENT .
01-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE FIRING HANDLE BROKE AND SNAPPED OFF DURING FIRING. IT OCCURRED ACROSS THE STOMACH STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-03-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | 75MM SELECTABLE NEW TLC | Detachment Of Device Component |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE FIRING HANDLE BROKE AND SNAPPED OFF DURING FIRING. IT OCCURRED ACROSS THE STOMACH STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE FIRST RELOAD HAD AN UNLOADING PROBLEM WHEREIN THE JOINT WAS BROKEN. WHILE THE SECOND AND THIRD RELOAD WERE ABLE TO FIRE BUT THERE WERE POOR STAPLE FORMATION AND WERE NOT SMOOTH DURING THE STAPLER FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE STAPLING OF STOMACH THE DEVICE DID NOT FIRE, THE GREEN BUTTON WAS DEPRESSED BUT THE HANDLE WOULD NOT ADVANCE THE I-BEAM. THEY ATTEMPTED TO FIRE A RELOAD WITH THE STAPLER HOWEVER, IT HAD FAILED TO FIRE AGAIN. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, WHILE THE DEVICE WAS BEING APPLIED ON THE STOMACH, THE DEVICE LOCKED MAKING IT IMPOSSIBLE TO CLOSE AND SEIZE THE FABRIC. THE RELOAD STOPPED AND IT WAS IMPOSSIBLE TO OPEN AND CLOSE THE MOVEMENT. THEY REPLACED THE DEVICE WITH ANOTHER LOAD TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien north haven-mfg | 09-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE FIRST STAPLER WAS ABLE TO FIRE HOWEVER, THERE WERE POOR STAPLE FORMATION. AND THE SECOND STAPLER WAS UNABLE TO FIRE. AS STATED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: FOR THE FIRST THREE FIRINGS FOR THE DEVICE WORKED FINE. ON THE FOURTH FIRING, THE DEVICE STOPPED FUNCTIONING EVEN THOUGH THE SURGEON OPERATED THE CONTROL TOGGLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
22-03-2018 | Malfunction | covidien north haven-mfg | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, THE DEVICE WAS UNABLE TO FIRE. WHEN THE SURGEON USED THE DEVICE AND CLAMPED IT, THEY TRIED TO SQUEEZE THE HANDLE OF THE DEVICE FOR FIRING ITS STAPLES AND TRANSECTING THE SPOT. AFTER THE GREEN BUTTON WAS PRESSED, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE¿S CARTRIDGE STAPLE AND KNIFE DID NOT FIRE. WHEN THE TRIGGER WAS PULLED, THE SURGEON HEARD SOMETHING BROKEN. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien north haven-mfg | 09-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, THE DEVICE WAS UNABLE TO FIRE. WHEN THE SURGEON USED THE DEVICE AND CLAMPED IT, THEY TRIED TO SQUEEZE THE HANDLE OF THE DEVICE FOR FIRING ITS STAPLES AND TRANSECTING THE SPOT. AFTER THE GREEN BUTTON WAS PRESSED, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE¿S CARTRIDGE STAPLE AND KNIFE DID NOT FIRE. WHEN THE TRIGGER WAS PULLED, THE SURGEON HEARD SOMETHING BROKEN. THERE WAS NO PATIENT INJURY.
11-02-2018 | Malfunction | ethicon endo-surgery | 09-04-2018 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RADICAL GASTRECTOMY FOR GASTRIC CANCER, COULD NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: FOR THE FIRST THREE FIRINGS FOR THE DEVICE WORKED FINE. ON THE FOURTH FIRING, THE DEVICE STOPPED FUNCTIONING EVEN THOUGH SURGEON OPERATED THE CONTROL TOGGLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
22-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | EEA | Failure To Adhere Or Bond |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OSOPHAGECTOMY PROCEDURE, WHILE ON THE FINAL ANASTOMOSIS, THE TROCAR AND ANVIL COULD NOT BE ATTACHED TOGETHER. THE ANVIL OF THIS DEVICE REMAINED IN SITU AND ANOTHER DEVICE WAS OPENED AND IT ATTACHED SUCCESSFULLY. THERE WAS NO PATIENT INJURY IN THIS EVENT.
20-02-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | EEA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SURGICAL STAPLING ON AN OPEN PROCEDURE, THE HANDLE WAS FULLY SQUEEZED AND CONTACTED THE STAPLER BODY TO THE FULLEST EXTENT, THE DONUTS WERE COMPLETE AND THE ANVIL COULD BE TILTED AFTER FIRING BUT THE CLIPS BECAME CROOKED AND MALFORMED. IT WAS NECESSARY TO REDO THE PROCEDURE WITH ANOTHER STAPLER. THERE WAS NO PATIENT INJURY IN THIS EVENT.
20-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/ PNEUMOTHORAX PROCEDURE, AT THE FOURTH FIRING OF BULLA RESECTION OF THE LUNGS, THE SURGEON CLAMPED THE PULMONARY PARENCHYMA, HOWEVER, THE JAWS OPENED ON THEIR OWN. THE SURGEON RE-CLAMPED THE TISSUE BUT THE SAME EVENT OCCURRED. THE SURGEON SAID THE TISSUE MIGHT HAVE BEEN SO THICK THAT THE JAWS COULD NOT CLOSE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 09-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE FIRST FIRING OF EMERGENT APPENDECTOMY, THE SURGEON PUSHED THE GREEN BUTTON, HOWEVER COULD NOT SQUEEZE THE HANDLE AND COULD NOT FIRE THE DEVICE. THERE WAS NO PATIENT INJURY.
13-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | EGIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE BEING APPLIED AT THE FIRST STAPLING OF THE STOMACH, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON. THE HANDLE HAD ALSO BROKEN OFF. THEY CHANGED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | EGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE BEING APPLIED AT THE FIRST STAPLING OF THE STOMACH, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON. THE HANDLE HAD ALSO BROKEN OFF. THEY CHANGED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY, FOR THE FIRST FIRING OF GASTRIC TUBE CONSTRUCTION, THE DEVICE WAS SET AND OPENING/ CLOSING WERE CHECKED WITH NO PROBLEM. THEN, THE SURGEON TRIED TO USE THE DEVICE, HOWEVER, THE JAWS COULD NOT CLOSE PROPERLY. THE CARTRIDGE WAS CHECKED AND NOTED THAT IT WAS PRE-FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
01-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | EGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SURGICAL STAPLING ON A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE STAPLER STOPPED AFTER THE FIRST SHOT. IT DID NOT WORK ANYMORE AND HAD TO BE CHANGED WITH A NEW STAPLER. THERE WAS NO PATIENT INJURY IN THIS EVENT.
20-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, WHILE THE DEVICE WAS BEING APPLIED ON THE STOMACH, THE DEVICE LOCKED MAKING IT IMPOSSIBLE TO CLOSE AND SEIZE THE FABRIC. THE RELOAD STOPPED AND IT WAS IMPOSSIBLE TO OPEN AND CLOSE THE MOVEMENT. THEY REPLACED THE DEVICE WITH ANOTHER LOAD TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | covidien lp llc north haven | 09-04-2018 | IDRIVE ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY, FOR THE FIRST FIRING OF GASTRIC TUBE CONSTRUCTION, THE DEVICE WAS SET AND OPENING AND CLOSING WERE CHECKED WITH NO PROBLEM. THEN, THE SURGEON TRIED TO USE THE DEVICE, HOWEVER, THE JAWS COULD NOT CLOSE PROPERLY. THE CARTRIDGE WAS CHECKED AND NOTED THAT IT WAS PRE-FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
15-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY, ONLY A HALF OF THE TARGET TISSUE WAS CUT AND STAPLED AT THE FIRST FIRING WITH GST60W, AND THE DEPLOYED STAPLES AT THE MIDDLE OF THE CARTRIDGE WERE UNFORMED. THE KNIFE MOVED FORWARD HALF THE WAY AND RETURNED TO HOME POSITION. THE DEVICE WAS USED ON THE DUODENUM. THEN, TWO GST60B WERE USED WITH THE SAME DEVICE, BUT THE SAME EVENT CONTINUED. ANOTHER 2ND DEVICE LOADING ANOTHER 4TH CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
22-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | ECHELON FLEX | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: DEVICE MISFIRED AND THE SURGEON HAD TO MANUALLY RELEASE THE DEVICE FROM THE TISSUE. THE TISSUE WAS NOT CUT, BUT THE DEVICE WAS LOCKED/"FROZEN" AND WOULD NOT OPEN. ANOTHER HANDLE WAS RETRIEVED. NO INJURY OCCURRED TO THE PATIENT.
22-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | ECHELON FLEX | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: DEVICE MISFIRED AND THE SURGEON HAD TO MANUALLY RELEASE THE DEVICE FROM THE TISSUE. THE TISSUE WAS NOT CUT, BUT THE DEVICE WAS LOCKED/"FROZEN" AND WOULD NOT OPEN. ANOTHER HANDLE WAS RETRIEVED. NO INJURY OCCURRED TO THE PATIENT.
21-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY WITH APPY PROCEDURE, WHILE TRANSECTING THE BOWEL, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT WHEN THE SURGEON ATTEMPTED TO STAPLE ACROSS THE BOWEL WITH THE BLUE RELOAD, A CRACK WAS HEARD IN THE HANDLE AND THE STAPLES MALFORMED. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY WITH APPY PROCEDURE, WHILE TRANSECTING THE BOWEL, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT WHEN THE SURGEON ATTEMPTED TO STAPLE ACROSS THE BOWEL WITH THE BLUE RELOAD, A CRACK WAS HEARD IN THE HANDLE AND THE STAPLES MALFORMED. THERE WAS NO PATIENT INJURY IN THIS EVENT.
23-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PUSH THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE.
02-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ROTICULATOR | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RECTOSIGMOIDECTOMY PROCEDURE, WHILE CLOSING THE RECTUM, THERE WAS AN INCOMPLETE STAPLE LINE. THE STAPLER WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
05-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | ECHELON FLEX | Misfire | A050502 | No Information | Event Description: WHILE USING ETHICON ECHELON FLEX POWERED PLUS STAPLER 60MM, STAPLER MISFIRED.
28-02-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING STAPLING, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPY SURGERY, WHILE THE DEVICE WAS BEING USED TO CUT LUNG, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE AND THE HANDLE WAS UNABLE TO SQUEEZE BY THE SURGEON. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER FIRING, THE STAPLER WAS NOT ABLE TO OPEN AND HAD A LOADING PROBLEM. THE JAWS HAD ALSO LOCKED ON TISSUE AND COULD NOT BE OPENED. THE BLADE WAS ALSO UNABLE TO RETRACT FULLY. PATIENT INVOLVEMENT NOT REPORTED.
23-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER FIRING, THE STAPLER WAS NOT ABLE TO OPEN AND HAD A LOADING PROBLEM. THE JAWS HAD ALSO LOCKED ON TISSUE AND COULD NOT BE OPENED. THE BLADE WAS ALSO UNABLE TO RETRACT FULLY. PATIENT INVOLVEMENT NOT REPORTED.
16-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Break | A0401 | Injury | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE CARTRIDGE, THE PUSHER WHERE FALLEN OUT OF THE CARTRIDGE. A PART OF THE DEVICE MAY HAVE FALLEN INTO THE PATIENT. THERE WAS PATIENT INJURY.
16-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE CARTRIDGE, THE PUSHER WHERE FALLEN OUT OF THE CARTRIDGE. A PART OF THE DEVICE MAY HAVE FALLEN INTO THE PATIENT. THERE WAS PATIENT INJURY.
16-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Detachment Of Device Component |  | Injury | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE CARTRIDGE, THE PUSHER WHERE FALLEN OUT OF THE CARTRIDGE. A PART OF THE DEVICE MAY HAVE FALLEN INTO THE PATIENT. THERE WAS PATIENT INJURY.
16-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE CARTRIDGE, THE PUSHER WHERE FALLEN OUT OF THE CARTRIDGE. A PART OF THE DEVICE MAY HAVE FALLEN INTO THE PATIENT. THERE WAS PATIENT INJURY.
16-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Detachment Of Device Component |  | Injury | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE CARTRIDGE, THE PUSHER WHERE FALLEN OUT OF THE CARTRIDGE. A PART OF THE DEVICE MAY HAVE FALLEN INTO THE PATIENT. THERE WAS PATIENT INJURY.
16-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE CARTRIDGE, THE PUSHER WHERE FALLEN OUT OF THE CARTRIDGE. A PART OF THE DEVICE MAY HAVE FALLEN INTO THE PATIENT. THERE WAS PATIENT INJURY.
16-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Injury | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE CARTRIDGE, THE PUSHER WHERE FALLEN OUT OF THE CARTRIDGE. A PART OF THE DEVICE MAY HAVE FALLEN INTO THE PATIENT. THERE WAS PATIENT INJURY.
16-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE CARTRIDGE, THE PUSHER WHERE FALLEN OUT OF THE CARTRIDGE. A PART OF THE DEVICE MAY HAVE FALLEN INTO THE PATIENT. THERE WAS PATIENT INJURY.
23-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER FIRING, THE STAPLER WAS NOT ABLE TO OPEN AND HAD A LOADING PROBLEM. THE JAWS HAD ALSO LOCKED ON TISSUE AND COULD NOT BE OPENED. THE BLADE WAS ALSO UNABLE TO RETRACT FULLY. PATIENT INVOLVEMENT NOT REPORTED.
23-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER FIRING, THE STAPLER WAS NOT ABLE TO OPEN AND HAD A LOADING PROBLEM. THE JAWS HAD ALSO LOCKED ON TISSUE AND COULD NOT BE OPENED. THE BLADE WAS ALSO UNABLE TO RETRACT FULLY. PATIENT INVOLVEMENT NOT REPORTED.
28-02-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING STAPLING, THE DEVICE DID NOT FIRE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
16-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | GIA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SUB-TOTAL COLECTOMY, WHILE TRANSECTING THE COLON, THE DEVICE FIRED BUT NO STAPLES CAME OUT SO IT DID NOT STAPLE, SO THE DOCTOR HAD TO HAND SEW TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDOGIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN STOMA RETURNING PROCEDURE, WHILE THE DEVICE WAS USED TO CLOSE THE JOINT, THE DEVICE WAS ABLE TO FIRE BUT HAD A POOR STAPLE FORMATION. THE JAW COULD NOT BE CLOSED, IT ONLY HAD THE KNIFE AND THE NAIL AND IT WAS NOT SHAPED. THEY TRIED TO HIT IT AGAIN TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
23-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDOGIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN STOMA RETURNING PROCEDURE, WHILE THE DEVICE WAS USED TO CLOSE THE JOINT, THE DEVICE WAS ABLE TO FIRE BUT HAD A POOR STAPLE FORMATION. THE JAW COULD NOT BE CLOSED, IT ONLY HAD THE KNIFE AND THE NAIL AND IT WAS NOT SHAPED. THEY TRIED TO HIT IT AGAIN TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
23-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PUSH THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE.
23-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THERE WAS A DIFFICULTY WITH LOADING THE RELOAD ONTO THE HANDLE AND RELOAD WAS NOT FASTENED SECURELY ONTO THE SHAFT. THE SURGEON WAS NOT ABLE TO PUT SECOND RELOAD IN PLACE. THERE WAS NO REPORTED PATIENT OUTCOME.
23-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN STOMA RETURNING PROCEDURE, WHILE THE DEVICE WAS USED TO CLOSE THE JOINT, THE DEVICE WAS ABLE TO FIRE BUT HAD A POOR STAPLE FORMATION. THE JAW COULD NOT BE CLOSED, IT ONLY HAD THE KNIFE AND THE NAIL AND IT WAS NOT SHAPED. THEY TRIED TO HIT IT AGAIN TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
23-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN STOMA RETURNING PROCEDURE, WHILE THE DEVICE WAS USED TO CLOSE THE JOINT, THE DEVICE WAS ABLE TO FIRE BUT HAD A POOR STAPLE FORMATION. THE JAW COULD NOT BE CLOSED, IT ONLY HAD THE KNIFE AND THE NAIL AND IT WAS NOT SHAPED. THEY TRIED TO HIT IT AGAIN TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
10-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | GIA | Component Missing | A020602 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRECTOMY, WHEN CUTTING ON THE DUODENUM, THE RELOAD HAD NO DLU INSIDE. AS STATED, THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | SIGNIA | Failure to Charge | A070603 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LADG (LAP. ASSISTED DISTAL GASTRECTOMY), THE SURGEON FIRED 3 SHOTS WITH THE STAPLER SUCCESSFULLY. AFTER A WHILE, THEY CONNECTED A NEW CARTRIDGE TO THE ADAPTER, HOWEVER THE DISPLAY SCREEN BLACKOUT AND IT DID NOT REACT. IT WAS REPLACED BY ANOTHER HANDLE AND THE PROCEDURE WAS COMPLETED. THE HANDLE IN QUESTION WAS INSERTED TO A CHARGER, HOWEVER IT DID NOT REACT AT ALL AND THE DEVICE GOT HOT. THERE WAS NO PATIENT HARM.
23-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | SIGNIA | Overheating of Device | A1005 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LADG (LAP. ASSISTED DISTAL GASTRECTOMY), THE SURGEON FIRED 3 SHOTS WITH THE STAPLER SUCCESSFULLY. AFTER A WHILE, THEY CONNECTED A NEW CARTRIDGE TO THE ADAPTER, HOWEVER THE DISPLAY SCREEN BLACKOUT AND IT DID NOT REACT. IT WAS REPLACED BY ANOTHER HANDLE AND THE PROCEDURE WAS COMPLETED. THE HANDLE IN QUESTION WAS INSERTED TO A CHARGER, HOWEVER IT DID NOT REACT AT ALL AND THE DEVICE GOT HOT. THERE WAS NO PATIENT HARM.
23-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | SIGNIA | Device Inoperable |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LADG (LAP. ASSISTED DISTAL GASTRECTOMY), THE SURGEON FIRED 3 SHOTS WITH THE STAPLER SUCCESSFULLY. AFTER A WHILE, THEY CONNECTED A NEW CARTRIDGE TO THE ADAPTER, HOWEVER THE DISPLAY SCREEN BLACKOUT AND IT DID NOT REACT. IT WAS REPLACED BY ANOTHER HANDLE AND THE PROCEDURE WAS COMPLETED. THE HANDLE IN QUESTION WAS INSERTED TO A CHARGER, HOWEVER IT DID NOT REACT AT ALL AND THE DEVICE GOT HOT. THERE WAS NO PATIENT HARM.
12-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT HEMI COLECTOMY, WHEN CUTTING ON THE COLON, THE DEVICE BROKE. AS STATED, THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | CONTOUR CURVED CUTTER STAPLER | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER WAS NOT WORKING. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: DURING LOBECTOMY, 4 FIRINGS WERE SUCCESSFUL. FOR THE 5TH FIRING IN BRONCHUS RESECTION, THE SURGEON CLAMPED THE TISSUE, PUSHED THE GREEN BUTTON AND STARTED SQUEEZING THE HANDLE, HOWEVER COULD NOT FIRE THE DEVICE. THE SURGEON RE-CLAMPED THE TISSUE HOWEVER THE STATUS WAS SAME. REPLACED BY NEW DEVICE, AND THE RESECTION WAS COMPLETED. NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPY SURGERY, WHILE THE DEVICE WAS BEING USED TO CUT LUNG, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE HOSPITAL COLLECTED PRE-FIRED RELOAD THROUGHOUT CERTAIN PERIOD AND REPORTED IT TO THE COMPANY AT ONCE. IT WAS NOT POSSIBLE TO IDENTIFY WHICH RELOAD WAS USED TO WHOM AND WHEN WITH WHAT TYPE OF HANDLE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
21-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY WITH APPY PROCEDURE, WHILE TRANSECTING THE BOWEL, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT WHEN THE SURGEON ATTEMPTED TO STAPLE ACROSS THE BOWEL WITH THE BLUE RELOAD, A CRACK WAS HEARD IN THE HANDLE AND THE STAPLES MALFORMED. A NEW RELOAD WAS USED WITH A NEW HANDLE AND IT FIRED PROPERLY. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY WITH APPY PROCEDURE, WHILE TRANSECTING THE BOWEL, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT WHEN THE SURGEON ATTEMPTED TO STAPLE ACROSS THE BOWEL WITH THE BLUE RELOAD, A CRACK WAS HEARD IN THE HANDLE AND THE STAPLES MALFORMED. A NEW RELOAD WAS USED WITH A NEW HANDLE AND IT FIRED PROPERLY. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | TA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS DIFFICULT/UNABLE TO CLOSE. THE SURGEON WAS NOT ABLE TO EITHER TO MOVE THE HANDLE, NOR CLOSE THE JAWS (ANVIL). HE TRIED ANOTHER ONE WITH THE SAME RESULT. THE DEVICE WAS USED DURING THE PROCEDURE BUT WAS NOT USED TO THE PATIENT. THERE WAS NO PATIENT INJURY.
01-02-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER TRANSPLANT PROCEDURE, WHILE STAPLING THE RIGHT HEPATIC VEIN, THE HANDLE BROKE UPON FIRING. THERE WAS NO PATIENT INJURY.
26-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE BEING APPLIED AT THE LIGATION OF PULMONARY UPPER VEIN AND LIGATION OF FIRST BRANCH OF PULMONARY ARTERY, THE STAPLER WAS ABLE TO FIRE BUT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. THE EVENT RESULTED TO PERMANENT PULMONARY ARTERY RUPTURE DUE TO THE STAPLER JAWS NOT EXPANDING ON THE VEIN, BLOOD LOSS OF MORE THAN 500CC, AN EXTENDED INCISION OF MORE THAN 1 INCH, AND AN EXTEND IN SURGICAL TIME BY MORE THAN 30 MINUTES -- DUE TO PERFORMING AN OPEN PROCEDURE. THE PATIENT REQUIRED AN EMERGENCY THORACOTOMY; AND THE WOUND INCISION IN THE PERICARDIUM REQUIRED MANUAL SUTURING OF ARTERY AND VEIN. THE PATIENT IS IN THE ICU.
26-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE BEING APPLIED AT THE LIGATION OF PULMONARY UPPER VEIN AND LIGATION OF FIRST BRANCH OF PULMONARY ARTERY, THE STAPLER WAS ABLE TO FIRE BUT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. THE EVENT RESULTED TO PERMANENT PULMONARY ARTERY RUPTURE DUE TO THE STAPLER JAWS NOT EXPANDING ON THE VEIN, BLOOD LOSS OF MORE THAN 500CC, AN EXTENDED INCISION OF MORE THAN 1 INCH, AND AN EXTEND IN SURGICAL TIME BY MORE THAN 30 MINUTES -- DUE TO PERFORMING AN OPEN PROCEDURE. THE PATIENT REQUIRED AN EMERGENCY THORACOTOMY; AND THE WOUND INCISION IN THE PERICARDIUM REQUIRED MANUAL SUTURING OF ARTERY AND VEIN. THE PATIENT IS IN THE ICU.
26-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE BEING APPLIED AT THE LIGATION OF PULMONARY UPPER VEIN AND LIGATION OF FIRST BRANCH OF PULMONARY ARTERY, THE STAPLER WAS ABLE TO FIRE BUT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. THE EVENT RESULTED TO PERMANENT PULMONARY ARTERY RUPTURE DUE TO THE STAPLER JAWS NOT EXPANDING ON THE VEIN, BLOOD LOSS OF MORE THAN 500CC, AN EXTENDED INCISION OF MORE THAN 1 INCH, AND AN EXTEND IN SURGICAL TIME BY MORE THAN 30 MINUTES -- DUE TO PERFORMING AN OPEN PROCEDURE. THE PATIENT REQUIRED AN EMERGENCY THORACOTOMY; AND THE WOUND INCISION IN THE PERICARDIUM REQUIRED MANUAL SUTURING OF ARTERY AND VEIN. THE PATIENT IS IN THE ICU.
26-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPY SURGERY, WHILE THE DEVICE WAS BEING USED TO CUT LUNG, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE AND THE HANDLE WAS UNABLE TO SQUEEZE BY THE SURGEON. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien north haven-mfg | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPY SURGERY, WHILE THE DEVICE WAS BEING USED TO CUT LUNG, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE HOSPITAL COLLECTED PRE-FIRED RELOAD THROUGHOUT CERTAIN PERIOD AND REPORTED IT TO THE COMPANY AT ONCE. IT WAS NOT POSSIBLE TO IDENTIFY WHICH RELOAD WAS USED TO WHOM AND WHEN WITH WHAT TYPE OF HANDLE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A AN OPEN "RHC" PROCEDURE, THE DEVICE DID NOT FIRE, AFTER THE FIRST FIRING WHILE THE SINGLE LOADING UNIT WAS TAKEN OUT FOR ADDITIONAL FIRING, A WHITE SAFETY LEVER POPPED OUT AT THE STAPLER SIDE AND AS THE ANVIL SIDE WAS TRYING TO BE CLAMPED TOGETHER. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | GIA | Fail-Safe Problem | A1602 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A AN OPEN "RHC" PROCEDURE, THE DEVICE DID NOT FIRE, AFTER THE FIRST FIRING WHILE THE SINGLE LOADING UNIT WAS TAKEN OUT FOR ADDITIONAL FIRING, A WHITE SAFETY LEVER POPPED OUT AT THE STAPLER SIDE AND AS THE ANVIL SIDE WAS TRYING TO BE CLAMPED TOGETHER. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC ESOPHAGECTOMY FOR CARCINOMA OF ESOPHAGUS FIRED THE DEVICE, FOUND THE TISSUE WAS CUT OFF, BUT STAPLES MALFORMED, AND DID NOT CUT THE BREAKAWAY WASHER. SUTURE WAS USED TO SEW THE WOUND. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
09-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | ILS 25MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC ESOPHAGECTOMY FOR CARCINOMA OF ESOPHAGUS FIRED THE DEVICE, FOUND THE TISSUE WAS CUT OFF, BUT STAPLES MALFORMED, AND DID NOT CUT THE BREAKAWAY WASHER. SUTURE WAS USED TO SEW THE WOUND. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
24-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | EEA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR), WHEN ON ANASTOMOSIS, THE STAPLER HAD POOR STAPLE FORMATION, THE HANDLE WAS FULLY SQUEEZED SO THAT THE METAL UNDERSIDE OF THE HANDLE CONTACTED THE STAPLER BODY TO THE FULLEST EXTENT, THE DONUTS WERE COMPLETE, THE ANVIL CAN BE TILTED AFTER FIRING AND WAS NOT BENT AND SIZERS WERE USED. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE RELOADING THE STAPLE AND TESTING THE DEVICE DURING A LAPAROSCOPIC LUNG WEDGE RESECTION PROCEDURE, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY CHANGED IT TO A NEW STAPLE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien north haven-mfg | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPY SURGERY, WHILE THE DEVICE WAS BEING USED TO CUT LUNG, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE HOSPITAL COLLECTED PRE-FIRED RELOAD THROUGHOUT CERTAIN PERIOD AND REPORTED IT TO THE COMPANY AT ONCE. IT WAS NOT POSSIBLE TO IDENTIFY WHICH RELOAD WAS USED TO WHOM AND WHEN WITH WHAT TYPE OF HANDLE. THERE WAS NO PATIENT INJURY.
16-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE OPEN LOW RECTAL CANCER RESECTION, AFTER FIRED, THERE WERE NO STAPLES ISSUED AFTER FIRING. SUTURE WAS USED TO SEW THE WOUND. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
26-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ASSISTED APPENDECTOMY PROCEDURE, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES AND ANOTHER STAPLER HANDLE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO INJURY.
08-02-2018 | Injury | ethicon endo-surgery | 10-04-2018 | PROXIMATE*RL TITANIUM STAPLER | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON EXPERIENCED AN ANASTOMOTIC LEAK WHERE THE DEVICE WAS USED. ON RE-EXPLORATION IT WAS NOTED THAT THE STAPLE LINE HAD AN OBVIOUS DEFECT.
08-02-2018 | Injury | ethicon endo-surgery | 10-04-2018 | PROXIMATE*RL TITANIUM STAPLER | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON EXPERIENCED AN ANASTOMOTIC LEAK WHERE THE DEVICE WAS USED. ON RE-EXPLORATION IT WAS NOTED THAT THE STAPLE LINE HAD AN OBVIOUS DEFECT.
23-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE RELOADING THE STAPLE AND TESTING THE DEVICE DURING A LAPAROSCOPIC LUNG WEDGE RESECTION PROCEDURE, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY CHANGED IT TO A NEW STAPLE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ASSISTED APPENDECTOMY PROCEDURE, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES AND ANOTHER STAPLER HANDLE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO INJURY.
26-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ASSISTED APPENDECTOMY PROCEDURE, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES AND ANOTHER STAPLER HANDLE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO INJURY.
31-01-2018 | Injury | ethicon endo-surgery | 10-04-2018 | PROXIMATE*RL TITANIUM STAPLER | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN ILEOSTOMY PROCEDURE, THE SURGEON EXPERIENCED AN ANASTOMOTIC LEAK. THE PATIENT WAS GETTING SEPTIC; RETURNED TO OPERATING ROOM FOR RE-EXPLORATION. UPON RE-EXPLORATION, LEAKAGE WAS NOTED FROM THE DEVICE'S STAPLE LINE. ILEOSTOMY CLOSURE.
22-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BILOBECTOMY VIDEO ASSISTED THORACOSCOPIC SURGERY, WHILE THE DEVICE WAS BEING APPLIED ON THE LUNG TISSUE/VESSEL, ALL THE STAPLERS WERE ABLE TO FIRE AND DID NOT HAVE ANY ISSUE WITH LOADING BUT THE PATIENT HAD SIGNIFICANT BLEEDING AND THE INCISION EXTENDED MORE THAN 1 INCH. TO COMPLETE THE CASE, PERFORATIONS WERE SUTURED AND CLIPPED AND HEMOSTATIC AGENTS WERE APPLIED. BLOOD TRANSFUSION IS POSSIBLE. THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE SURGEON DID NOT COMPLAIN ABOUT THE STAPLER DURING THE CASE.
22-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BILOBECTOMY VIDEO ASSISTED THORACOSCOPIC SURGERY, WHILE THE DEVICE WAS BEING APPLIED ON THE LUNG TISSUE/VESSEL, ALL THE STAPLERS WERE ABLE TO FIRE AND DID NOT HAVE ANY ISSUE WITH LOADING BUT THE PATIENT HAD SIGNIFICANT BLEEDING AND THE INCISION EXTENDED MORE THAN 1 INCH. TO COMPLETE THE CASE, PERFORATIONS WERE SUTURED AND CLIPPED AND HEMOSTATIC AGENTS WERE APPLIED. BLOOD TRANSFUSION IS POSSIBLE. THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE SURGEON DID NOT COMPLAIN ABOUT THE STAPLER DURING THE CASE.
26-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ASSISTED APPENDECTOMY PROCEDURE, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES AND ANOTHER STAPLER HANDLE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, THE STAPLES WERE PRE-FIRED BEFORE USE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | ENDO ILS 29MM, CURVED | Leak/Splash | A0504 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE THE STAPLES ON THE POSTERIOR SIDE OF THE ANASTAMOSIS DID NOT FORM CORRECTLY LEAVING A LEAK IN THAT AREA. THE SURGEON EXCISED THE FIRST ANASTAMOSIS AND USED A SECOND DEVICE TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THE CUSTOMER CONFIRMED THAT THE SECOND ANASTAMOSIS PASSED THE LEAK TEST.
17-03-2018 | Malfunction | ethicon endo-surgery | 10-04-2018 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER STOPPED DURING SURGERY AND DID NOT FIRE, HAVING TO BE REPLACED BY ANOTHER STAPLER. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
26-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, THE STAPLES WERE PRE-FIRED BEFORE USE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BILOBECTOMY VIDEO ASSISTED THORACOSCOPIC SURGERY, WHILE THE DEVICE WAS BEING APPLIED ON THE LUNG TISSUE/VESSEL, ALL THE STAPLERS WERE ABLE TO FIRE AND DID NOT HAVE ANY ISSUE WITH LOADING BUT THE PATIENT HAD SIGNIFICANT BLEEDING AND THE INCISION EXTENDED MORE THAN 1 INCH. TO COMPLETE THE CASE, PERFORATIONS WERE SUTURED AND CLIPPED AND HEMOSTATIC AGENTS WERE APPLIED. BLOOD TRANSFUSION IS POSSIBLE. THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE SURGEON DID NOT COMPLAIN ABOUT THE STAPLER DURING THE CASE.
22-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BILOBECTOMY VIDEO ASSISTED THORACOSCOPIC SURGERY, WHILE THE DEVICE WAS BEING APPLIED ON THE LUNG TISSUE/VESSEL, ALL THE STAPLERS WERE ABLE TO FIRE AND DID NOT HAVE ANY ISSUE WITH LOADING BUT THE PATIENT HAD SIGNIFICANT BLEEDING AND THE INCISION EXTENDED MORE THAN 1 INCH. TO COMPLETE THE CASE, PERFORATIONS WERE SUTURED AND CLIPPED AND HEMOSTATIC AGENTS WERE APPLIED. BLOOD TRANSFUSION IS POSSIBLE. THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE SURGEON DID NOT COMPLAIN ABOUT THE STAPLER DURING THE CASE.
08-03-2018 | Injury | ethicon endo-surgery | 10-04-2018 | PROXIMATE*PPH PROCEDURE SET | Insufficient Information | E2401 | No Code Available | Event Description: IT WAS REPORTED THAT DURING TREATMENT OF HEMORRHOID PROLAPSE, POSTOPERATIVE BLEEDING OCCURRED IN THE NAPPA OF THE ENTIRE SUTURE LINE. RE-ENTERED THE OPERATING ROOM, PROCEEDED TO SUTURE BY QUADRANT OF THE CIRCLE. PATIENT CONDITION IS GOOD. THE DEVICE WAS DISCARDED.
22-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Exsanguination | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BILOBECTOMY VIDEO ASSISTED THORACOSCOPIC SURGERY, WHILE THE DEVICE WAS BEING APPLIED ON THE LUNG TISSUE/VESSEL, ALL THE STAPLERS WERE ABLE TO FIRE AND DID NOT HAVE ANY ISSUE WITH LOADING BUT THE PATIENT HAD SIGNIFICANT BLEEDING AND THE INCISION EXTENDED MORE THAN 1 INCH. TO COMPLETE THE CASE, PERFORATIONS WERE SUTURED AND CLIPPED AND HEMOSTATIC AGENTS WERE APPLIED. BLOOD TRANSFUSION IS POSSIBLE. THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE SURGEON DID NOT COMPLAIN ABOUT THE STAPLER DURING THE CASE.
22-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BILOBECTOMY VIDEO ASSISTED THORACOSCOPIC SURGERY, WHILE THE DEVICE WAS BEING APPLIED ON THE LUNG TISSUE/VESSEL, ALL THE STAPLERS WERE ABLE TO FIRE AND DID NOT HAVE ANY ISSUE WITH LOADING BUT THE PATIENT HAD SIGNIFICANT BLEEDING AND THE INCISION EXTENDED MORE THAN 1 INCH. TO COMPLETE THE CASE, PERFORATIONS WERE SUTURED AND CLIPPED AND HEMOSTATIC AGENTS WERE APPLIED. BLOOD TRANSFUSION IS POSSIBLE. THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE SURGEON DID NOT COMPLAIN ABOUT THE STAPLER DURING THE CASE.
22-03-2018 | Injury | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BILOBECTOMY VIDEO ASSISTED THORACOSCOPIC SURGERY, WHILE THE DEVICE WAS BEING APPLIED ON THE LUNG TISSUE/VESSEL, ALL THE STAPLERS WERE ABLE TO FIRE AND DID NOT HAVE ANY ISSUE WITH LOADING BUT THE PATIENT HAD SIGNIFICANT BLEEDING AND THE INCISION EXTENDED MORE THAN 1 INCH. TO COMPLETE THE CASE, PERFORATIONS WERE SUTURED AND CLIPPED AND HEMOSTATIC AGENTS WERE APPLIED. BLOOD TRANSFUSION IS POSSIBLE. THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE SURGEON DID NOT COMPLAIN ABOUT THE STAPLER DURING THE CASE.
01-02-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LEFT VATS WEDGE RESECTION ON LEFT UPPER LOBE AND PLEURECTOMY, TO CLIP A PORTION OF THE LEFT LUNG FOR PATHOLOGY, THE DOCTOR USED A STAPLER WHICH MISFIRED. THERE WAS NO INJURY.
01-02-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LEFT VATS WEDGE RESECTION ON LEFT UPPER LOBE AND PLEURECTOMY, TO CLIP A PORTION OF THE LEFT LUNG FOR PATHOLOGY, THE DOCTOR USED A STAPLER WHICH MISFIRED. THERE WAS NO INJURY.
22-03-2018 | Malfunction | covidien north haven-mfg | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: DURING LOBECTOMY, 4 FIRINGS WERE SUCCESSFUL. FOR THE 5TH FIRING IN BRONCHUS RESECTION, THE SURGEON CLAMPED THE TISSUE, PUSHED THE GREEN BUTTON AND STARTED SQUEEZING THE HANDLE, HOWEVER COULD NOT FIRE THE DEVICE. THE SURGEON RE-CLAMPED THE TISSUE HOWEVER THE STATUS WAS SAME. THERE WAS NO PATIENT INJURY.
25-01-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPE RADICAL RESECTION OF PULMONARY CARCINOMA, UPON CUTTING LOBI PULMONIS, THE DEVICE COULD NOT LOAD OFF. THERE WAS NO REPORTED PATIENT OUTCOME.
02-02-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY THORAX PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE. THE STAPLER WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
25-01-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPE RADICAL RESECTION OF PULMONARY CARCINOMA, UPON CUTTING LOBI PULMONIS, THE DEVICE COULD NOT LOAD OFF. THERE WAS NO REPORTED PATIENT OUTCOME.
26-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | EGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPY SURGERY, WHILE THE DEVICE WAS BEING USED TO CUT LUNG, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE HOSPITAL COLLECTED PRE-FIRED RELOAD THROUGHOUT CERTAIN PERIOD AND REPORTED IT TO THE COMPANY AT ONCE. IT WAS NOT POSSIBLE TO IDENTIFY WHICH RELOAD WAS USED TO WHOM AND WHEN WITH WHAT TYPE OF HANDLE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
02-02-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY THORAX PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE. THE STAPLER WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
23-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ROTICULATOR | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOWER ANTERIOR RESECTION, WHILE CLOSING THE RECTUM, THE JAWS WERE UNABLE TO BE OPENED TO BE REMOVE FROM TISSUE. THE DEVICE WAS PRIED OFF USING ADDITIONAL TOOL AND USED ANOTHER INSTRUMENT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE BEING APPLIED AT THE LIGATION OF PULMONARY UPPER VEIN AND LIGATION OF FIRST BRANCH OF PULMONARY ARTERY, THE STAPLER WAS ABLE TO FIRE BUT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. THE EVENT RESULTED TO PERMANENT PULMONARY ARTERY RUPTURE DUE TO THE STAPLER JAWS NOT EXPANDING ON THE VEIN, BLOOD LOSS OF MORE THAN 500CC, AN EXTENDED INCISION OF MORE THAN 1 INCH, AND AN EXTEND IN SURGICAL TIME BY MORE THAN 30 MINUTES -- DUE TO PERFORMING AN OPEN PROCEDURE. THE PATIENT REQUIRED AN EMERGENCY THORACOTOMY; AND THE WOUND INCISION IN THE PERICARDIUM REQUIRED MANUAL SUTURING OF ARTERY AND VEIN. THE PATIENT IS IN THE ICU.
26-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Fluid/Blood Leak | A050401 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE BEING APPLIED AT THE LIGATION OF PULMONARY UPPER VEIN AND LIGATION OF FIRST BRANCH OF PULMONARY ARTERY, THE STAPLER WAS ABLE TO FIRE BUT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. THE EVENT RESULTED TO PERMANENT PULMONARY ARTERY RUPTURE DUE TO THE STAPLER JAWS NOT EXPANDING ON THE VEIN, BLOOD LOSS OF MORE THAN 500CC, AN EXTENDED INCISION OF MORE THAN 1 INCH, AND AN EXTEND IN SURGICAL TIME BY MORE THAN 30 MINUTES -- DUE TO PERFORMING AN OPEN PROCEDURE. THE PATIENT REQUIRED AN EMERGENCY THORACOTOMY; AND THE WOUND INCISION IN THE PERICARDIUM REQUIRED MANUAL SUTURING OF ARTERY AND VEIN. THE PATIENT IS IN THE ICU.
26-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE BEING APPLIED AT THE LIGATION OF PULMONARY UPPER VEIN AND LIGATION OF FIRST BRANCH OF PULMONARY ARTERY, THE STAPLER WAS ABLE TO FIRE BUT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. THE EVENT RESULTED TO PERMANENT PULMONARY ARTERY RUPTURE DUE TO THE STAPLER JAWS NOT EXPANDING ON THE VEIN, BLOOD LOSS OF MORE THAN 500CC, AN EXTENDED INCISION OF MORE THAN 1 INCH, AND AN EXTEND IN SURGICAL TIME BY MORE THAN 30 MINUTES -- DUE TO PERFORMING AN OPEN PROCEDURE. THE PATIENT REQUIRED AN EMERGENCY THORACOTOMY; AND THE WOUND INCISION IN THE PERICARDIUM REQUIRED MANUAL SUTURING OF ARTERY AND VEIN. THE PATIENT IS IN THE ICU.
26-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE BEING APPLIED AT THE LIGATION OF PULMONARY UPPER VEIN AND LIGATION OF FIRST BRANCH OF PULMONARY ARTERY, THE STAPLER WAS ABLE TO FIRE BUT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. THE EVENT RESULTED TO PERMANENT PULMONARY ARTERY RUPTURE DUE TO THE STAPLER JAWS NOT EXPANDING ON THE VEIN, BLOOD LOSS OF MORE THAN 500CC, AN EXTENDED INCISION OF MORE THAN 1 INCH, AND AN EXTEND IN SURGICAL TIME BY MORE THAN 30 MINUTES -- DUE TO PERFORMING AN OPEN PROCEDURE. THE PATIENT REQUIRED AN EMERGENCY THORACOTOMY; AND THE WOUND INCISION IN THE PERICARDIUM REQUIRED MANUAL SUTURING OF ARTERY AND VEIN. THE PATIENT IS IN THE ICU.
26-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE BEING APPLIED AT THE LIGATION OF PULMONARY UPPER VEIN AND LIGATION OF FIRST BRANCH OF PULMONARY ARTERY, THE STAPLER WAS ABLE TO FIRE BUT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. THE EVENT RESULTED TO PERMANENT PULMONARY ARTERY RUPTURE DUE TO THE STAPLER JAWS NOT EXPANDING ON THE VEIN, BLOOD LOSS OF MORE THAN 500CC, AN EXTENDED INCISION OF MORE THAN 1 INCH, AND AN EXTEND IN SURGICAL TIME BY MORE THAN 30 MINUTES -- DUE TO PERFORMING AN OPEN PROCEDURE. THE PATIENT REQUIRED AN EMERGENCY THORACOTOMY; AND THE WOUND INCISION IN THE PERICARDIUM REQUIRED MANUAL SUTURING OF ARTERY AND VEIN. THE PATIENT IS IN THE ICU.
26-03-2018 | Injury | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE BEING APPLIED AT THE LIGATION OF PULMONARY UPPER VEIN AND LIGATION OF FIRST BRANCH OF PULMONARY ARTERY, THE STAPLER WAS ABLE TO FIRE BUT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. THE EVENT RESULTED TO PERMANENT PULMONARY ARTERY RUPTURE DUE TO THE STAPLER JAWS NOT EXPANDING ON THE VEIN, BLOOD LOSS OF MORE THAN 500CC, AN EXTENDED INCISION OF MORE THAN 1 INCH, AND AN EXTEND IN SURGICAL TIME BY MORE THAN 30 MINUTES -- DUE TO PERFORMING AN OPEN PROCEDURE. THE PATIENT REQUIRED AN EMERGENCY THORACOTOMY; AND THE WOUND INCISION IN THE PERICARDIUM REQUIRED MANUAL SUTURING OF ARTERY AND VEIN. THE PATIENT IS IN THE ICU.
06-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE.
06-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE.
22-03-2018 | Malfunction | us surgical puerto rico | 10-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WEDGE RESECTION, WHILE RESECTING THE PART OF A LUNG, THE STAPLER DID NOT FIRE AT ALL, THE PUSHER THUMB DIDN'T MAKE ANY MOVE. THE OPERATOR DECIDED TO TAKE ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-11-2017 | Malfunction | covidien lp llc north haven | 10-04-2018 | nan | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE STAPLE LINE WAS INCOMPLETE. THERE WAS NO PATIENT INJURY.
14-03-2018 | Malfunction | covidien lp llc north haven | 10-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHEN CUTTING THE STOMACH, THE DEVICE DID NOT OPEN WHEN TRIED TO APPLY THE JAW INTO THE TISSUE. THERE WAS NO PATIENT INJURY.
06-04-2018 | Malfunction | 3m | 10-04-2018 | PRECISE | Mechanical Jam | A0506 | No Known Impact Or Consequence To Patient | Event Description: THE 3M PRECISE DISPOSABLE SKIN STAPLER PGX-35W JAMMED DURING USE.
22-02-2018 | Malfunction | ethicon endo-surgery | 11-04-2018 | 75MM SELECTABLE NEW TLC | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THE DURING AN UNKNOWN PROCEDURE, WHEN THE SURGEON TRIED TO FIRE THE DEVICE, HE HAD A SAFETY LOCK OUT, AND DIDN¿T REALIZE IT WAS THE RELOAD THAT HAD NOT BEEN LOADED CORRECTLY. A LIKE STAPLE GUN WAS OPENED TO COMPLETE THE PROCEDURE ONCE THEY HAD REALIZED THEIR MISTAKE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-02-2018 | Malfunction | ethicon endo-surgery | 11-04-2018 | 75MM SELECTABLE NEW TLC | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE WHEN THE SURGEON TRIED TO FIRE THE DEVICE, HE HAD A SAFETY LOCK OUT, AND DIDN¿T REALIZE IT WAS THE RELOAD THAT HAD NOT BEEN LOADED CORRECTLY. A LIKE DEVICE WAS OPENED TO COMPLETE THE PROCEDURE ONCE THEY HAD REALIZED THEIR MISTAKE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-02-2018 | Malfunction | ethicon endo-surgery | 11-04-2018 | 75MM SELECTABLE NEW TLC | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE WOULD NOT LOAD. THE PROCEDURE WAS COMPLETED USING A LIKE DEVICE OF THE SAME PRODUCT CODE. NO ADDITIONAL INFORMATION IS AVAILABLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-02-2018 | Malfunction | ethicon endo-surgery | 11-04-2018 | CONTOUR CURVED CUTTER STAPLER | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN LAR PROCEDURE, THE STAPLES FIRED DURING THE INTERMEDIATE LOCKING POSITION. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
23-03-2018 | Injury | ethicon endo-surgery | 11-04-2018 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: ENDOLOOP VERSUS ENDOSTAPLER CLOSURE OF THE APPENDICEAL STUMP IN PEDIATRIC LAPAROSCOPIC APPENDECTOMY, CITATION: CAN J SURG, VOL. 55, NO. 1, FEBRUARY 2012; DOI: 10.1503/CJS.023810. THE AIM OF THIS STUDY WAS TO COMPARE INFECTIOUS OUTCOMES BETWEEN 2 LAPAROSCOPIC TECHNIQUES OF APPENDICEAL STUMP CLOSURE: ENDOLOOP (EL) AND ENDOSTAPLER (ES). A TOTAL OF 242 PATIENTS UNDERWENT LA DURING THE STUDY PERIOD AND HAD COMPLETE DATA SETS ABSTRACTED RETROSPECTIVELY FROM THEIR MEDICAL RECORDS. THE PATIENT WERE DIVIDED INTO 2 GROUPS: NON-PERFORATED APPENDICITIS (NPA) N=185, PERFORATED APPENDICITIS (PA) N=57. ONE SURGEON USED THE ES (ETHICON ENDOSURGERY) IN ALL CASES (PA AND NONPERFORATED APPENDICITIS [NPA]) WHEREAS 2 SURGEONS ROUTINELY USED EL UNLESS THE CONDITION OF THE APPENDICEAL BASE FAVOURED USE OF AN ES TO ENSURE SEALING OF HEALTHY TISSUE. ENDOLOOP STUMP CLOSURE CONSISTED OF 2 PROXIMAL AND 1 DISTAL 2¿0 POLYDIOXANONE (PDS; ETHICON) ELS, WITH DIVISION OF THE APPENDICEAL BASE BETWEEN THE MIDDLE AND DISTAL EL. SURGICAL SITE INFECTIONS (NPA (N=2); PA (N=2)) AND INTRA-ABDOMINAL ABSCESS FORMATION (NPA (N=5); PA (N=6) WITH THE USE OF EL WERE REPORTED. INTRA-ABDOMINAL ABSCESS FORMATION (N=5) WITH THE USE OF ENDOSTAPLER WAS ALSO REPORTED. IN CONCLUSION, THIS STUDY SUGGESTS THAT PREFERENTIAL USE OF EL FOR APPENDICEAL STUMP CLOSURE IN LA IS ASSOCIATED WITH AT LEAST EQUIVALENT INFECTIOUS OUTCOMES, AS WELL AS COST SAVINGS, IN COMPARISON TO ROUTINE USE OF ES.
23-03-2018 | Injury | ethicon endo-surgery | 11-04-2018 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: ENDOLOOP VERSUS ENDOSTAPLER CLOSURE OF THE APPENDICEAL STUMP IN PEDIATRIC LAPAROSCOPIC APPENDECTOMY, CITATION: CAN J SURG, VOL. 55, NO. 1, FEBRUARY 2012; DOI: 10.1503/CJS.023810. THE AIM OF THIS STUDY WAS TO COMPARE INFECTIOUS OUTCOMES BETWEEN 2 LAPAROSCOPIC TECHNIQUES OF APPENDICEAL STUMP CLOSURE: ENDOLOOP (EL) AND ENDOSTAPLER (ES). A TOTAL OF 242 PATIENTS UNDERWENT LA DURING THE STUDY PERIOD AND HAD COMPLETE DATA SETS ABSTRACTED RETROSPECTIVELY FROM THEIR MEDICAL RECORDS. THE PATIENT WERE DIVIDED INTO 2 GROUPS: NON-PERFORATED APPENDICITIS (NPA) N=185, PERFORATED APPENDICITIS (PA) N=57. ONE SURGEON USED THE ES (ETHICON ENDOSURGERY) IN ALL CASES (PA AND NONPERFORATED APPENDICITIS [NPA]) WHEREAS 2 SURGEONS ROUTINELY USED EL UNLESS THE CONDITION OF THE APPENDICEAL BASE FAVOURED USE OF AN ES TO ENSURE SEALING OF HEALTHY TISSUE. ENDOLOOP STUMP CLOSURE CONSISTED OF 2 PROXIMAL AND 1 DISTAL 2¿0 POLYDIOXANONE (PDS; ETHICON) ELS, WITH DIVISION OF THE APPENDICEAL BASE BETWEEN THE MIDDLE AND DISTAL EL. SURGICAL SITE INFECTIONS (NPA (N=2); PA (N=2)) AND INTRA-ABDOMINAL ABSCESS FORMATION (NPA (N=5); PA (N=6) WITH THE USE OF EL WERE REPORTED. INTRA-ABDOMINAL ABSCESS FORMATION (N=5) WITH THE USE OF ENDOSTAPLER WAS ALSO REPORTED. IN CONCLUSION, THIS STUDY SUGGESTS THAT PREFERENTIAL USE OF EL FOR APPENDICEAL STUMP CLOSURE IN LA IS ASSOCIATED WITH AT LEAST EQUIVALENT INFECTIOUS OUTCOMES, AS WELL AS COST SAVINGS, IN COMPARISON TO ROUTINE USE OF ES.
18-11-2017 | Injury | ethicon endo-surgery | 11-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Leak/Splash | A0504 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER AN LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THE PATIENT PRESENTED WITH A LEAK. THE SURGEON USED STENTS AND TPN TO TREAT THE PATIENT. THE PATIENT IS NOW STABLE.
28-12-2017 | Injury | ethicon endo-surgery | 11-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER AN LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THE PATIENT PRESENTED WITH A LEAK. THE SURGEON USED STENTS AND TPN TO TREAT THE PATIENT. THE PATIENT IS NOW STABLE.
26-11-2017 | Injury | ethicon endo-surgery | 11-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THE PATIENT PRESENTED WITH A LEAK. THE SURGEON USED STENTS AND TPN TO TREAT THE PATIENT. THE PATIENT IS NOW STABLE.
12-11-2017 | Injury | ethicon endo-surgery | 11-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER AN LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THE PATIENT PRESENTED WITH A LEAK. THE SURGEON USED STENTS AND TPN TO TREAT THE PATIENT. THE PATIENT IS NOW STABLE. NO DEVICE WILL BE RETURNED.
10-11-2017 | Injury | ethicon endo-surgery | 11-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER AN LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THE PATIENT PRESENTED WITH A LEAK. THE SURGEON USED STENTS AND TPN TO TREAT THE PATIENT. THE PATIENT IS NOW STABLE.
01-01-2018 | Malfunction | ethicon | 11-04-2018 | ETHICON SECURESTRAP | Failure To Adhere Or Bond |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND ABSORBABLE STRAPS WERE USED. DURING THE PROCEDURE, THE DEVICE DID NOT WORK, DID NOT HOLD THE MESH. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. NO ADDITIONAL INFORMATION WAS PROVIDED.
25-11-2017 | Malfunction | us surgical puerto rico | 11-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE STAPLE LINE WAS INCOMPLETE. THERE WAS NO PATIENT INJURY.
16-03-2018 | Malfunction | ethicon endo-surgery | 12-04-2018 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURNG A ROBOTIC LOW ANTERIOR RESECTION ON THE FIRST FIRING OF THE DEVICE THE STAPLES DEPLOYED BUT ONLY HALF OF THE STAPLES FORMED PROPERLY. THE SURGEON TURNED THE KNOB THREE QUARTERS OF A TURN BUT THE TISSUES WOULD NOT DISENGAGE. THE TISSUE WAS CUT AWAY FROM AROUND THE STAPLER HEAD SO THAT STAPLER COULD BE REMOVED. BECAUSE OF THIS THE ANASTOMOSIS WAS NOT FORMED. THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE WITH NO PATIENT CONSEQUENCES, HOWEVER THE PROCEDURE WAS PROLONGED BY THREE HOURS.
16-01-2018 | Malfunction | davol | 12-04-2018 | OPTIFIX ABSORBABLE FIXATION SYSTEM | Break | A0401 | No Information | Event Description: THE BARD OPTIFIX ABSORBABLE FIXATION SYSTEM DEVICE WAS IN USE LAPAROSCOPIC TO AFFIX MESH IN PLACE. IT IS SUPPOSED TO HAVE 15 LOADS OF FASTENERS, BUT OVER HALF OF THEM CAME OUT BROKEN. A NEW OPTIFIX WAS USED, BUT THE SCRUB TECH ATTEMPTED TO TRY AND RELEASE THE REST OF THE FASTENERS ONTO THE BACK TABLE AND MOST OF THEM ALSO BROKE WHEN COMING OUT OF THE DEVICE.
15-03-2018 | Malfunction | ethicon endo-surgery | 12-04-2018 | ILS 33MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: DURING USE OF THE DEVICE FOR COLORECTAL ANASTOMOSIS, WE DID NOT HEAR THE FIRING OF THE CLIPPING MOMENT. IT WAS OBSERVED THAT THE DEVICE HAD STAPLED THE TISSUE BUT DID NOT CUT IT. SEGMENTAL RESECTION AND RE-STAPLING WERE REQUIRED (A NEW STAPLER WAS USED) TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-03-2018 | Injury | ethicon endo-surgery | 12-04-2018 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMICOLECTOMY PROCEDURE, THEY USED THE DEVICE TO DO THE ANASTOMOSIS, THEY MADE THE SHOT, THEN TURNED THE KNOB TO OPEN THE DEVICE AND REMOVE IT. THEY THEN MADE A LEAK AND REVIEW OF THE TWO DONUTS. THEY WERE OK. IN THE POSTOPERATIVE PERIOD THERE IS EVIDENCE THAT THE PATIENT PRESENTED FISTULA.
21-03-2018 | Injury | ethicon endo-surgery | 12-04-2018 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMICOLECTOMY PROCEDURE, THEY USED THE DEVICE TO DO THE ANASTOMOSIS, THEY MADE THE SHOT, THEN TURNED THE KNOB TO OPEN THE DEVICE AND REMOVE IT. THEY THEN MADE A LEAK AND REVIEW OF THE TWO DONUTS. THEY WERE OK. IN THE POSTOPERATIVE PERIOD THERE IS EVIDENCE THAT THE PATIENT PRESENTED FISTULA.
06-02-2018 | Injury | ethicon endo-surgery | 12-04-2018 | POWERED PLUS COMPACT 45 | Insufficient Information | E2401 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCEE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS). NO FURTHER INFORMATION AVAILABLE FOR NOW.
13-02-2018 | Injury | ethicon endo-surgery | 12-04-2018 | POWERED PLUS COMPACT 45 | Leak/Splash | A0504 | Not Applicable | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS). THE DEVICE HAS BEEN USED WITH: 1 CARTRIDGE ECR45D, LOT: P4T629, 1 CARTRIDGE ECR45D, LOT: P4T20D.
13-02-2018 | Injury | ethicon endo-surgery | 12-04-2018 | POWERED PLUS COMPACT 45 | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS). THE DEVICE HAS BEEN USED WITH: 1 CARTRIDGE ECR45D, LOT: P4T629, 1 CARTRIDGE ECR45D, LOT: P4T20D.
13-02-2018 | Injury | ethicon endo-surgery | 12-04-2018 | POWERED PLUS COMPACT 45 | Device Operates Differently Than Expected |  | Not Applicable | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS). THE DEVICE HAS BEEN USED WITH: 1 CARTRIDGE ECR45D, LOT: P4T629, 1 CARTRIDGE ECR45D, LOT: P4T20D.
13-02-2018 | Injury | ethicon endo-surgery | 12-04-2018 | POWERED PLUS COMPACT 45 | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS). THE DEVICE HAS BEEN USED WITH: 1 CARTRIDGE ECR45D, LOT: P4T629, 1 CARTRIDGE ECR45D, LOT: P4T20D.
27-02-2018 | Injury | ethicon endo-surgery | 12-04-2018 | POWERED PLUS COMPACT 45 | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS).
15-03-2018 | Malfunction | ethicon endo-surgery | 12-04-2018 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY; LAST FIRING ON THE FISSURE. APPROX 1CM OF TISSUE LEFT TO DIVIDE. A BLUE RELOAD WAS IN THE DEVICE. THE DEVICE FIRED HOWEVER, THE SURGEON COULD NOT GET THE DEVICE OPEN. UPON PRESSING THE REVERSE BUTTON NOTHING HAPPENED. NO ADVERSE EVENT TO PATIENT. NO DELAY IN SURGERY.
15-03-2018 | Malfunction | ethicon endo-surgery | 12-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE THE CUSTOMER OBSERVED THAT AFTER COMPLETING AN UNSPECIFIED FIRING THAT NOT ALL OF THE STAPLES HAD DEPLOYED AND THE CUT LINE APPEARED JAGGED. THE SURGEON OVERSEWED THE STAPLELINE TO SECURE IT. PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-03-2018 | Malfunction | ethicon endo-surgery | 12-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY ON THE FOUR FIRING OF THE DEVICE THE BLADE CUT THE ENTIRE LENGTH BUT THE STAPLES ONLY DEPLOYED ON THE SPECIMEN SIDE. THE DOCTOR COMPLETED THE SLEEVE AND OVER SEWED A ONE INCH OPENING THAT REMAINED. A DRAIN HAD TO BE PLACE IN THE PATIENT¿S ABDOMEN. THE PATIENTS CURRENT CONDITION IS UNKNOWN.
21-12-2017 | Malfunction | ethicon endo-surgery | 12-04-2018 | CONTOUR CURVED CUTTER STAPLER | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW RETINOSIGMOIDECTOMY DISTAL, THE DEVICE HAD A DEFECT, THE SIDE LOCK DID NOT OPEN, IT WAS NOT POSSIBLE TO USE IT AND IT WAS REQUESTED TO REPLACE IT WITH ANOTHER LIKE DEVICE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-12-2017 | Malfunction | ethicon endo-surgery | 12-04-2018 | CONTOUR CURVED CUTTER STAPLER | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW RETINOSIGMOIDECTOMY DISTAL, THE DEVICE HAD A DEFECT, THE SIDE LOCK DID NOT OPEN, IT WAS NOT POSSIBLE TO USE IT AND IT WAS REQUESTED TO REPLACE IT WITH ANOTHER LIKE DEVICE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-03-2018 | Malfunction | ethicon endo-surgery | 12-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY ONLY THE STAPLES ON THE PATIENT SIDE DEPLOYED. THE SURGEON HAND SEWED THE OPENING ON THE OTHER SIDE CLOSED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
09-03-2018 | Malfunction | ethicon endo-surgery | 12-04-2018 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THAT DURING AN OPEN LOBECTOMY PROCEDURE THE DEVICE WOULD NOT OPEN. NO ADDITIONAL INFORMATION WAS AVAILABLE AT THE TIME OF THE CALL. IT WAS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
15-03-2018 | Malfunction | ethicon endo-surgery | 12-04-2018 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT STAPLES MALFORMED AFTER FIRING. DURING THE ENDOSCOPIC RIGHT UPPER LOBECTOMY, WHEN SEVERING THE PULMONARY FISSURE WITH ECR60D, AFTER FIRED, FOUND THE STAPLES MALFORMED WITH IRREGULAR SHAPE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
25-01-2018 | Injury | intuitive surgical | 12-04-2018 | ENDOWRIST | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER UNDERGOING A DA VINCI-ASSISTED GASTRECTOMY PROCEDURE, THE PATIENT EXPERIENCED A STAPLE LINE LEAK AND BECAME SEPTIC. AS A RESULT, THE PATIENT UNDERWENT A TRADITIONAL OPEN PROCEDURE TO REPAIR THE ANASTOMOSIS. ON 28-MAR-2018, INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: THE CSR WAS NOT PRESENT DURING THE PRIMARY GASTRECTOMY PROCEDURE WHICH WAS PERFORMED ON (B)(6) 2018. DURING THE DA VINCI SURGICAL PROCEDURE, IT WAS REPORTED THAT 3 GREEN STAPLER RELOADS WERE USED AND THERE WERE NO REPORTED INTRA-OPERATIVE COMPLICATIONS OR A MALFUNCTION OF THE DA VINCI SURGICAL SYSTEM. THERE WERE ALSO NO REPORTED STAPLER ERRORS OR CLAMPING, FIRING, UNCLAMPING ISSUES. A COUPLE OF MONTHS POST-OPERATIVELY, THE PATIENT BECAME SEPTIC AND A PROXIMAL STAPLE LINE LEAK WAS FOUND. ACCORDING TO THE CSR, THE SURGEON CLAIMED THAT DURING THE OPEN SURGERY TO REPAIR THE LEAK, SOME OF THE STAPLES DID NOT HAVE THE B-SHAPED FORMATION. THEREFORE, THE SURGEON ALLEGED THAT THE MALFORMED STAPLES CAUSED THE STAPLE LINE LEAK. THE CSR DID NOT KNOW WHAT DATE THE OPEN SURGERY WAS PERFORMED. IT IS REPORTED THAT THE PATIENT IS IN A STABLE CONDITION. ISI HAD MADE MULTIPLE ATTEMPTS TO CONTACT THE SURGEON TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, THERE WAS NO RESPONSE.
26-01-2018 | Malfunction | us surgical puerto rico | 12-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING THE FIRST FIRING OF PULMONARY PARENCHYMA RESECTION, THE SURGEON CONNECTED THE RELOAD TO THE HANDLE AND UPON FIRING, A CRACK SOUND WAS HEARD. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE PATIENT IS ALIVE AND HAS NO INJURY.
12-04-2018 | Malfunction | us surgical puerto rico | 12-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE PURPLE STAPLER MISFIRED. THERE WAS NO REPORTED PATIENT INVOLVEMENT AND OUTCOME.
09-09-2016 | Malfunction | covidien lp llc north haven | 12-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY, THE STAPLER DID NOT FIRE. NO PATIENT INJURY.
14-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | ILA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ROUX-EN-Y PROCEDURE, WHILE BEING APPLIED ON JEJUNAL TRANSECTION, THE DEVICE SHOOTS VERY ROUGHLY AND A BULGING OF THE STAPLER WAS NOTED. DURING FIRING, THE DEVICE COMPLETELY FALLEN APART AND THE SURGEON USED FORCE, SO BOTH STAPLERS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
14-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | ILA | Difficult or Delayed Positioning | A150203 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ROUX-EN-Y PROCEDURE, WHILE BEING APPLIED ON JEJUNAL TRANSECTION, THE DEVICE SHOOTS VERY ROUGHLY AND A BULGING OF THE STAPLER WAS NOTED. DURING FIRING, THE DEVICE COMPLETELY FALLEN APART AND THE SURGEON USED FORCE, SO BOTH STAPLERS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
06-11-2017 | Injury | us surgical puerto rico | 12-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE LOAD HANGED AND THE STAPLER ALSO DID NOT WORK. IT HAS BEEN WITH THE MOLE HANDLE AND COULD NOT BE USED. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE SO THE STAPLER DID NOT FIRE. ANOTHER STAPLER WAS USED AND THE SURGERY WAS CONVERTED FROM THE LAPAROSCOPIC SURGERY TO AN OPEN SURGERY.
15-03-2018 | Malfunction | us surgical puerto rico | 12-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE BEING APPLIED AT THE LAST STEP, THE STAPLER WAS ABLE TO FIRE BUT THE CARTRIDGE OPENED ONLY IN A PARTIAL MANNER AND THE PART OF THE CARTRIDGE THAT CLOSES WAS HARD TO CLOSE. THEY CHANGED THE CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-03-2018 | Malfunction | us surgical puerto rico | 12-04-2018 | ENDO GIA | Positioning Problem | A1502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE BEING APPLIED AT THE LAST STEP, THE STAPLER WAS ABLE TO FIRE BUT THE CARTRIDGE OPENED ONLY IN A PARTIAL MANNER AND THE PART OF THE CARTRIDGE THAT CLOSES WAS HARD TO CLOSE. THEY CHANGED THE CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION WHICH WAS SUSPECTED AS DUE TO A DISCRETE BLEEDING THAT OCCURRED IN THE MECHANICAL SUTURE LINE INSIDE THE POUCH OF THE STOMACH FOR THE BYPASS. A REINFORCEMENT WITH THE SUTURE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION WHICH WAS SUSPECTED AS DUE TO A DISCRETE BLEEDING THAT OCCURRED IN THE MECHANICAL SUTURE LINE INSIDE THE POUCH OF THE STOMACH FOR THE BYPASS. A REINFORCEMENT WITH THE SUTURE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION WHICH WAS SUSPECTED AS DUE TO A DISCRETE BLEEDING THAT OCCURRED IN THE MECHANICAL SUTURE LINE INSIDE THE POUCH OF THE STOMACH FOR THE BYPASS. A REINFORCEMENT WITH THE SUTURE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION WHICH WAS SUSPECTED AS DUE TO A DISCRETE BLEEDING THAT OCCURRED IN THE MECHANICAL SUTURE LINE INSIDE THE POUCH OF THE STOMACH FOR THE BYPASS. A REINFORCEMENT WITH THE SUTURE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION WHICH WAS SUSPECTED AS DUE TO A DISCRETE BLEEDING THAT OCCURRED IN THE MECHANICAL SUTURE LINE INSIDE THE POUCH OF THE STOMACH FOR THE BYPASS. A REINFORCEMENT WITH THE SUTURE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION WHICH WAS SUSPECTED AS DUE TO A DISCRETE BLEEDING THAT OCCURRED IN THE MECHANICAL SUTURE LINE INSIDE THE POUCH OF THE STOMACH FOR THE BYPASS. A REINFORCEMENT WITH THE SUTURE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
06-11-2017 | Injury | covidien lp llc north haven | 12-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE LOAD HANGED AND THE STAPLER ALSO DID NOT WORK. IT HAS BEEN WITH THE MOLE HANDLE AND COULD NOT BE USED. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE SO THE STAPLER DID NOT FIRE. ANOTHER STAPLER WAS USED AND THE SURGERY WAS CONVERTED FROM THE LAPAROSCOPIC SURGERY TO AN OPEN SURGERY.
20-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | SIGNIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE, THE JAWS OF THE DEVICE COULD NOT BE CLOSED ALTHOUGH THE CLAMP TEST WAS SUCCESSFUL. THE CENTER CONTROL BUTTON WAS PUSHED UP AND THE JAWS WERE RETURNED TO THE STRAIGHT POSITION. THE RELOAD WAS THEN REINSERTED INTO THE ADAPTER, HOWEVER, THE DEVICE DID NOT REACT AT ALL. THE ADAPTER AND THE SHELL WAS REPLACED, AND THE FIRING WAS COMPLETED. THERE WAS NO PATIENT INJURY. AFTER THE PROCEDURE, THE DEVICE WAS CHECKED AGAIN AND THE DISPLAY SCREEN SHOWED THE RELOAD CYCLE TEST INDICATOR. THE JAWS COULD BE ARTICULATED, HOWEVER, THE JAWS COULD NOT BE CLOSED STILL.
12-04-2018 | Malfunction | us surgical puerto rico | 12-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE PURPLE STAPLER MISFIRED. THERE WAS NO REPORTED PATIENT INVOLVEMENT.
19-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | RELIATACK | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR, WHILE FIXING THE MESH TO THE PERITONEUM, THERE WAS DIFFICULTY FIRING THE TACK AND THERE WAS A CRUNCHING SOUND WHEN SQUEEZING. A PORTION OF THE RELOAD DISENGAGED INTO THE CAVITY OF THE PATIENT BUT WAS RETRIEVED RIGHT AWAY. SOME TACKS WERE ABLE TO BE FIRED NORMALLY, PENETRATE AND HOLD ONTO TISSUE CORRECTLY. THEY OPENED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR, WHILE FIXING THE MESH TO THE PERITONEUM, THERE WAS DIFFICULTY FIRING THE TACK AND THERE WAS A CRUNCHING SOUND WHEN SQUEEZING. A PORTION OF THE RELOAD DISENGAGED INTO THE CAVITY OF THE PATIENT BUT WAS RETRIEVED RIGHT AWAY. SOME TACKS WERE ABLE TO BE FIRED NORMALLY, PENETRATE AND HOLD ONTO TISSUE CORRECTLY. THEY OPENED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | covidien lp llc north haven | 12-04-2018 | RELIATACK | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR, WHILE FIXING THE MESH TO THE PERITONEUM, THERE WAS DIFFICULTY FIRING THE TACK AND THERE WAS A CRUNCHING SOUND WHEN SQUEEZING. A PORTION OF THE RELOAD DISENGAGED INTO THE CAVITY OF THE PATIENT BUT WAS RETRIEVED RIGHT AWAY. SOME TACKS WERE ABLE TO BE FIRED NORMALLY, PENETRATE AND HOLD ONTO TISSUE CORRECTLY. THEY OPENED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-02-2018 | Malfunction | ethicon endo-surgery | 13-04-2018 | CONTOUR CURVED CUTTER STAPLER | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLON PROCEDURE THE DEVICE WAS TAKEN OUT OF THE BOX AND WHEN IT WAS TESTED THE JAWS WOULD NOT CLOSE. THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED.
20-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE, THE JAWS OF THE DEVICE COULD NOT BE CLOSED ALTHOUGH THE CLAMP TEST WAS SUCCESSFUL. THE CENTER CONTROL BUTTON WAS PUSHED UP AND THE JAWS WERE RETURNED TO THE STRAIGHT POSITION. THE RELOAD WAS THEN REINSERTED INTO THE ADAPTER HOWEVER; THE DEVICE DID NOT REACT AT ALL. THE ADAPTER AND THE SHELL WERE REPLACED, AND THE FIRING WAS COMPLETED. AFTER THE PROCEDURE, THE DEVICE WAS CHECKED AGAIN AND THE DISPLAY SCREEN SHOWED THE RELOAD CYCLE TEST INDICATOR. THE JAWS COULD BE ARTICULATED HOWEVER, THE JAWS COULD NOT BE CLOSED STILL. THERE WAS NO PATIENT INJURY.
01-12-2017 | Malfunction | covidien lp llc north haven | 13-04-2018 | SIGNIA | Failure to Charge | A070603 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING SERVICING. THE HANDLE HAS NO POWER AND WILL NOT TAKE A CHARGE FROM THE CHARGER. NO PATIENT INVOLVED.
26-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE, THE STAPLER WAS ABLE TO FIRE BUT THE STAPLE LINE WAS INCOMPLETE ON THE PROXIMAL PART. THE SURGEON PLACED THE CARTRIDGE AND AT THE MOMENT OF MAKING THE TRIGGER, THE CARTRIDGE ADVANCED A PART AND GOT STUCK. THE SURGEON PLACED THE SECOND CARTRIDGE AND AT THE TIME OF SHOOTING, THE POCKET KNIFE DID NOT SHORT THE TISSUE. THE SURGEON HAD TO ¿INVAGINATE¿ THE LINE OF STAPLES TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
20-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: DURING SEGMENTAL RESECTION OF LUNG, THE CARTRIDGE WAS CONNECTED TO THE HANDLE WITH NO PROBLEM . WHEN THE SURGEON CLAMPED THE TISSUE, HE/SHE FELT THE HANDLE WAS STIFF. THE SURGEON TRIED TO FIRE THE DEVICE, HOWEVER THE KNIFE DID NOT ADVANCE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | SIGNIA | Device Inoperable |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THE HANDLE WAS REMOVED FROM THE CHARGING STATION AND A 2/3 BATTERY CHARGE STATUS WAS DISPLAYED. THE HANDLE RECOGNIZED THE DISPOSABLE POWER SHELL AND ALSO THE STANDARD ADAPTER. THEN, THE HANDLE WAS CONNECTED WITH THE RELOAD AND THE HANDLE RECOGNIZED IT. THE NURSE SET ASIDE THE HANDLE. AFTER A SHORT WHILE, THE DISPLAY TURNED OFF AND THE HANDLE WAS NO LONGER FUNCTIONAL. NEITHER THE RELOAD COULD NOT BE OPENED OR CLOSED. THERE WAS NO PATIENT INJURY.
12-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Cut | A050702 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SEPTATION PROCEDURE, WHILE SURGICAL STAPLING, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD AND THE STAPLER WAS ABLE TO FIRE BUT THE FIRING BLADE DID NOT MOVE. ANOTHER PRODUCT WAS USE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THE EVENT.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE RESECTING THE APPENDIX AND MESO-APPENDIX, THE TWO STAPLERS WITH THEIR LOADING UNITS WERE UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | GIA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING UROLOGICAL PROCEDURE, THE DOCTOR PLACED THE STAPLER ON THE INTESTINE, FIRED IT BUT THERE WERE NO STAPLES IN THE LOADING UNIT AND THE INTESTINE WAS CUT. THE INCIDENT HAPPENED IN 2 LOADING UNITS WITH THE SAME ISSUE. ON THE THIRD TIME THAT THE DOCTOR USED SEPARATE LOADING UNIT, THERE WERE CLIPS ON IT AND ALL WORKED WELL. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE, THE STAPLER WAS ABLE TO FIRE BUT THE STAPLE LINE WAS INCOMPLETE ON THE PROXIMAL PART. THE SURGEON PLACED THE CARTRIDGE AND AT THE MOMENT OF MAKING THE TRIGGER, THE CARTRIDGE ADVANCED A PART AND GOT STUCK. THE SURGEON PLACED THE SECOND CARTRIDGE AND AT THE TIME OF SHOOTING, THE POCKET KNIFE DID NOT SHORT THE TISSUE. THE SURGEON HAD TO ¿INVAGINATE¿ THE LINE OF STAPLES TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Cut | A050702 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE, THE STAPLER WAS ABLE TO FIRE BUT THE STAPLE LINE WAS INCOMPLETE ON THE PROXIMAL PART. THE SURGEON PLACED THE CARTRIDGE AND AT THE MOMENT OF MAKING THE TRIGGER, THE CARTRIDGE ADVANCED A PART AND GOT STUCK. THE SURGEON PLACED THE SECOND CARTRIDGE AND AT THE TIME OF SHOOTING, THE POCKET KNIFE DID NOT SHORT THE TISSUE. THE SURGEON HAD TO ¿INVAGINATE¿ THE LINE OF STAPLES TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
13-04-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE PURPLE STAPLER MISFIRED. THERE WAS NO REPORTED PATIENT INVOLVEMENT AND OUTCOME.
26-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, A PLASTIC FOREIGN MATERIAL WAS FOUND ON THE COVER CLOTH. THE SAME HANDLE WAS USED THROUGH THE PROCEDURE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/WEDGE/SEGMENTAL LUNG RESECTION PROCEDURE, FOR THE SECOND FIRING, INSTRUMENT DID NOT FIRE. THE GREEN BUTTON WAS PRESSED HOWEVER; THE HANDLE WAS STIFF AND COULD NOT BE SQUEEZED. THERE WAS NO PATIENT INJURY.
12-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDOGIA | Failure to Cut | A050702 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SEPTATION PROCEDURE, WHILE SURGICAL STAPLING, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD AND THE STAPLER WAS ABLE TO FIRE BUT THE FIRING BLADE DID NOT MOVE. ANOTHER PRODUCT WAS USE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THE EVENT.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE RESECTING THE APPENDIX AND MESO-APPENDIX, THE TWO STAPLERS WITH THEIR LOADING UNITS WERE UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. THEY OPENED A NEW INSTRUMENT OF SAME TYPE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | GIA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING UROLOGICAL PROCEDURE, THE DOCTOR PLACED THE STAPLER ON THE INTESTINE, FIRED IT BUT THERE WERE NO STAPLES IN THE LOADING UNIT AND THE INTESTINE WAS CUT. THE INCIDENT HAPPENED IN 2 LOADING UNITS WITH THE SAME ISSUE. ON THE THIRD TIME THAT THE DOCTOR USED SEPARATE LOADING UNIT, THERE WERE CLIPS ON IT AND ALL WORKED WELL. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SUBTOTAL GASTRECTOMY, ON STOMACH RESECTION, AFTER THE SURGEON HELD THE TISSUE WITH THE DEVICE, IT DID NOT FIRE AND BLACK PUSHER THUMB PIECE CANNOT BE MOVED AT ALL. AS STATED, THE OPERATING ROOM NURSE PREPARED A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Device Slipped | A051204 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING PULMONARY PARENCHYMA RESECTION IN A LOBECTOMY PROCEDURE, THE HANDLE RAN IDLE AND THE KNIFE STOPPED ON THE HALFWAY. THEY RECOGNIZED THE TISSUE WAS THICK BUT THEY THOUGHT THE CARTRIDGE WAS APPROPRIATE FOR THE THICKNESS. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
01-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING PULMONARY PARENCHYMA RESECTION IN A LOBECTOMY PROCEDURE, THE HANDLE RAN IDLE AND THE KNIFE STOPPED ON THE HALFWAY. THEY RECOGNIZED THE TISSUE WAS THICK BUT THEY THOUGHT THE CARTRIDGE WAS APPROPRIATE FOR THE THICKNESS. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE RESECTING THE APPENDIX AND MESO-APPENDIX, THE TWO STAPLERS WITH THEIR LOADING UNITS WERE UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. THEY OPENED A NEW INSTRUMENT OF SAME TYPE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | covidien north haven-mfg | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: DURING SEGMENTAL RESECTION OF LUNG, THE CARTRIDGE WAS CONNECTED TO THE HANDLE WITH NO PROBLEM . WHEN THE SURGEON CLAMPED THE TISSUE, HE/SHE FELT THE HANDLE WAS STIFF. THE SURGEON TRIED TO FIRE THE DEVICE, HOWEVER THE KNIFE DID NOT ADVANCE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON WAS UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
01-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING PULMONARY PARENCHYMA RESECTION IN A LOBECTOMY PROCEDURE, THE HANDLE RAN IDLE AND THE KNIFE STOPPED ON THE HALFWAY. THEY RECOGNIZED THE TISSUE WAS THICK BUT THEY THOUGHT THE CARTRIDGE WAS APPROPRIATE FOR THE THICKNESS. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE RESECTING THE APPENDIX AND MESO-APPENDIX, THE TWO STAPLERS WITH THEIR LOADING UNITS WERE UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. THEY OPENED A NEW INSTRUMENT OF SAME TYPE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Hematoma | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHILE STAPLING THE ARTERY, THE STAPLER, WITH ITS LOADING UNIT, WAS UNABLE TO FIRE, ALTHOUGH THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE. THE JAWS HAD ALSO LOCKED ON TISSUE. THE DEVICE WAS CHANGED TO COMPLETE THE PROCEDURE. THE PATIENT HAD INCURRED A HEMATOMA ON THE WALL OF THE ARTERY.
21-03-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHILE STAPLING THE ARTERY, THE STAPLER, WITH ITS LOADING UNIT, WAS UNABLE TO FIRE, ALTHOUGH THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE. THE JAWS HAD ALSO LOCKED ON TISSUE. THE DEVICE WAS CHANGED TO COMPLETE THE PROCEDURE. THE PATIENT HAD INCURRED A HEMATOMA ON THE WALL OF THE ARTERY.
21-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Entrapment of Device | A150208 | Hematoma | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHILE STAPLING THE ARTERY, THE STAPLER, WITH ITS LOADING UNIT, WAS UNABLE TO FIRE, ALTHOUGH THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE. THE JAWS HAD ALSO LOCKED ON TISSUE. THE DEVICE WAS CHANGED TO COMPLETE THE PROCEDURE. THE PATIENT HAD INCURRED A HEMATOMA ON THE WALL OF THE ARTERY.
21-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHILE STAPLING THE ARTERY, THE STAPLER, WITH ITS LOADING UNIT, WAS UNABLE TO FIRE, ALTHOUGH THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE. THE JAWS HAD ALSO LOCKED ON TISSUE. THE DEVICE WAS CHANGED TO COMPLETE THE PROCEDURE. THE PATIENT HAD INCURRED A HEMATOMA ON THE WALL OF THE ARTERY.
13-04-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | EGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WOULD NOT FIRE. THE SALES REP RECEIVED THE DAMAGED DEVICE, UNKNOWN ON WHAT HAPPENED AND WHY.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE RESECTING THE APPENDIX AND MESO-APPENDIX, THE TWO STAPLERS WITH THEIR LOADING UNITS WERE UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. THEY OPENED A NEW INSTRUMENT OF SAME TYPE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE- OPERATIVELY, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON WAS UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
13-04-2018 | Malfunction | us surgical puerto rico | 13-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WOULD NOT FIRE. THE SALES REP RECEIVED THE DAMAGED DEVICE, UNKNOWN ON WHAT HAPPENED.
27-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | SIGNIA | Device Inoperable |  | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER HAD ITS STATUS, HANDLE, ADAPTER AND RELOAD INDICATOR TURNED OFF BUT WAS STILL ABLE TO RECOGNIZE THE RELOAD. HOWEVER, IT WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THE HANDLE WAS ALSO PUT INTO A SHELL AND THE LED SCREEN WENT BLACK AND IT WOULD NOT COME BACK UP AFTER TRYING TO RESET IT. THERE WAS NO PATIENT INVOLVED.
27-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | SIGNIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER HAD ITS STATUS, HANDLE, ADAPTER AND RELOAD INDICATOR TURNED OFF BUT WAS STILL ABLE TO RECOGNIZE THE RELOAD. HOWEVER, IT WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THE HANDLE WAS ALSO PUT INTO A SHELL AND THE LED SCREEN WENT BLACK AND IT WOULD NOT COME BACK UP AFTER TRYING TO RESET IT. THERE WAS NO PATIENT INVOLVED.
27-03-2018 | Malfunction | covidien north haven-mfg | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/WEDGE/SEGMENTAL LUNG RESECTION PROCEDURE, FOR THE SECOND FIRING, INSTRUMENT DID NOT FIRE. THE GREEN BUTTON WAS PRESSED HOWEVER, THE HANDLE WAS STIFF AND COULD NOT BE SQUEEZED. THE RELOAD WAS REPLACED AND THE FIRING WAS COMPLETED. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | SIGNIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER HAD ITS STATUS, HANDLE, ADAPTER AND RELOAD INDICATOR TURNED OFF BUT WAS STILL ABLE TO RECOGNIZE THE RELOAD. HOWEVER, IT WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THE HANDLE WAS ALSO PUT INTO A SHELL AND THE LED SCREEN WENT BLACK AND IT WOULD NOT COME BACK UP AFTER TRYING TO RESET IT. THERE WAS NO PATIENT INVOLVED.
27-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER HAD ITS STATUS, HANDLE, ADAPTER AND RELOAD INDICATOR TURNED OFF BUT WAS STILL ABLE TO RECOGNIZE THE RELOAD. HOWEVER, IT WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THE HANDLE WAS ALSO PUT INTO A SHELL AND THE LED SCREEN WENT BLACK AND IT WOULD NOT COME BACK UP AFTER TRYING TO RESET IT. THERE WAS NO PATIENT INVOLVED.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | SIGNIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE DID NOT PRODUCE A SOUND INDICATING THE DEVICE'S STATUS. THE SURGEON PUSHED THE 4 WAY TOGGLE UP BUT DID NOT RETURN TO THE NEUTRAL POSITION. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | SIGNIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE DID NOT PRODUCE A SOUND INDICATING THE DEVICE'S STATUS. THE SURGEON PUSHED THE 4 WAY TOGGLE UP BUT DID NOT RETURN TO THE NEUTRAL POSITION. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Injury | us surgical puerto rico | 13-04-2018 | GIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC PULMONARY BIOPSY PROCEDURE, WHILE DOING THE WEDGE RESECTION, THE DEVICE HAD LOCKED ON THE TISSUE. THE SURROUNDING TISSUE WAS RESECTED TO REMOVE THE DEVICE WHICH EXTENDED THE SURGICAL TIME FOR MORE THAN 30 MINUTES. THE RELOAD WAS ALSO STUCK ON THE STAPLER AND COULD NOT BE UNLOADED.
27-03-2018 | Injury | us surgical puerto rico | 13-04-2018 | GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC PULMONARY BIOPSY PROCEDURE, WHILE DOING THE WEDGE RESECTION, THE DEVICE HAD LOCKED ON THE TISSUE. THE SURROUNDING TISSUE WAS RESECTED TO REMOVE THE DEVICE WHICH EXTENDED THE SURGICAL TIME FOR MORE THAN 30 MINUTES. THE RELOAD WAS ALSO STUCK ON THE STAPLER AND COULD NOT BE UNLOADED.
27-03-2018 | Injury | us surgical puerto rico | 13-04-2018 | GIA | Human-Device Interface Problem | A22 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC PULMONARY BIOPSY PROCEDURE, WHILE DOING THE WEDGE RESECTION, THE DEVICE HAD LOCKED ON THE TISSUE. THE SURROUNDING TISSUE WAS RESECTED TO REMOVE THE DEVICE WHICH EXTENDED THE SURGICAL TIME FOR MORE THAN 30 MINUTES. THE RELOAD WAS ALSO STUCK ON THE STAPLER AND COULD NOT BE UNLOADED.
27-03-2018 | Injury | us surgical puerto rico | 13-04-2018 | GIA | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC PULMONARY BIOPSY PROCEDURE, WHILE DOING THE WEDGE RESECTION, THE DEVICE HAD LOCKED ON THE TISSUE. THE SURROUNDING TISSUE WAS RESECTED TO REMOVE THE DEVICE WHICH EXTENDED THE SURGICAL TIME FOR MORE THAN 30 MINUTES. THE RELOAD WAS ALSO STUCK ON THE STAPLER AND COULD NOT BE UNLOADED.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RE-CONSTRUCTION (FUNCTIONAL END-TO-END ANASTOMOSIS) OF MINI LAPAROSCOPIC COLECTOMY, THE FIRST AND THE SECOND FIRINGS WERE SUCCESSFUL. AT THIRD FIRING (SIDE-TO-SIDE ANASTOMOSIS),THE SURGEON STARTED FIRING, HOWEVER A STRANGE SOUND WAS HEARD FROM THE HANDLE AND COULD NOT FIRE THE DEVICE. HE NOTED THE STAPLES APPEARED FROM THE STAPLE POCKET OF PROXIMAL END BY 1CM. HE OPENED THE JAWS. THE DEVICE WAS REPLACED WITH A NEW CARTRIDGE, AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RE-CONSTRUCTION (FUNCTIONAL END-TO-END ANASTOMOSIS) OF MINI LAPAROSCOPIC COLECTOMY, THE FIRST AND THE SECOND FIRINGS WERE SUCCESSFUL. AT THIRD FIRING (SIDE-TO-SIDE ANASTOMOSIS),THE SURGEON STARTED FIRING, HOWEVER A STRANGE SOUND WAS HEARD FROM THE HANDLE AND COULD NOT FIRE THE DEVICE. HE NOTED THE STAPLES APPEARED FROM THE STAPLE POCKET OF PROXIMAL END BY 1CM. HE OPENED THE JAWS. THE DEVICE WAS REPLACED WITH A NEW CARTRIDGE, AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | IDRIVE ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RE-CONSTRUCTION (FUNCTIONAL END-TO-END ANASTOMOSIS) OF MINI LAPAROSCOPIC COLECTOMY, THE FIRST AND THE SECOND FIRINGS WERE SUCCESSFUL. AT THIRD FIRING (SIDE-TO-SIDE ANASTOMOSIS),THE SURGEON STARTED FIRING, HOWEVER A STRANGE SOUND WAS HEARD FROM THE HANDLE AND COULD NOT FIRE THE DEVICE. HE NOTED THE STAPLES APPEARED FROM THE STAPLE POCKET OF PROXIMAL END BY 1CM. HE OPENED THE JAWS. THE DEVICE WAS REPLACED WITH A NEW CARTRIDGE, AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | IDRIVE ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RE-CONSTRUCTION (FUNCTIONAL END-TO-END ANASTOMOSIS) OF MINI LAPAROSCOPIC COLECTOMY, THE FIRST AND THE SECOND FIRINGS WERE SUCCESSFUL. AT THIRD FIRING (SIDE-TO-SIDE ANASTOMOSIS),THE SURGEON STARTED FIRING, HOWEVER A STRANGE SOUND WAS HEARD FROM THE HANDLE AND COULD NOT FIRE THE DEVICE. HE NOTED THE STAPLES APPEARED FROM THE STAPLE POCKET OF PROXIMAL END BY 1CM. HE OPENED THE JAWS. THE DEVICE WAS REPLACED WITH A NEW CARTRIDGE, AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDOGIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, THE DEVICE DID NOT PRODUCE A SOUND INDICATING THE DEVICE'S STATUS. THE SURGEON PUSHED THE FOUR WAY TOGGLE UP BUT DID NOT RETURN TO THE NEUTRAL POSITION. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | IDRIVE ULTRA | Unintended Head Motion |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT ASSISTED LAPAROSCOPIC LOBECTOMY, BEFORE ENTERING THE TROCAR WHICH WAS NOT ON THE PATIENT AND WHEN THE RELOAD WAS PLACED, THERE WERE 2 DEVICES THAT STARTED TO CALIBRATE, STARTED SPINNING AND WOULD NOT STOP AFTER THE ADAPTER WAS TAKEN OFF. A NEW POWERED STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
14-03-2018 | Malfunction | ethicon | 13-04-2018 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Failure To Adhere Or Bond |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT INGUINAL HERNIA REPAIR ON (B)(6) 2018 AND ABSORBABLE STRAPS WERE USED. DURING THE PROCEDURE, STAPLES OF DEVICE DID NOT FIX ONTO THE MESH. WHEN SURGEON FIRED THE STAPLES PULLED OUT. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
23-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY PROCEDURE, AFTER CLAMPING THE TISSUE AND CLOSING THE DEVICE JAWS, THE GREEN BUTTON WAS PUSHED TO FIRE THE DEVICE. THE LOWER SIDE OF THE ROTATION BUTTON WAS CLICKED ON A SLOW FIRING MODE, HOWEVER THE DEVICE DID NOT REACT. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LIVER RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT IT DID NOT FIRE AND WAS UNABLE TO SQUEEZE THE HANDLE. THE SCRUB NURSE CHANGED THE DEVICE AND IT DID WORK. THERE WAS NO PATIENT CONDITION REPORTED.
23-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY PROCEDURE, AFTER CLAMPING THE TISSUE AND CLOSING THE DEVICE JAWS, THE GREEN BUTTON WAS PUSHED TO FIRE THE DEVICE, THE LOWER SIDE OF THE ROTATION BUTTON WAS CLICKED FOR A SLOW FIRING MODE, HOWEVER THE DEVICE DID NOT REACT. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE FIRING AT THE STOMACH SLEEVE, THE STAPLER WAS ABLE TO FIRE BUT THE RELOAD TORE THE SEROSA AND THE SURGEON HAD TO OVER SEW THE DEFECT.
27-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE FIRING AT THE STOMACH SLEEVE, THE STAPLER WAS ABLE TO FIRE BUT THE RELOAD TORE THE SEROSA AND THE SURGEON HAD TO OVER SEW THE DEFECT.
16-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER WAS UNABLE TO FIRE WHEN IT WAS CONNECTED TO AN APPLICATOR. THE PRODUCT WAS CHANGED TO ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVED.
14-11-2017 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, IT APPEARED THAT A LITTLE PIECE OF PLASTIC FELL FROM THE HANDLE WHEN THE TRIGGER WAS PULLED. THE COMPONENT FELL INTO THE PATIENT CAVITY BUT WAS RETRIEVED. THEY REPLACED THE RELOAD AND USED THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Unintended Head Motion |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT ASSISTED LAPAROSCOPIC LOBECTOMY, BEFORE ENTERING THE TROCAR WHICH WAS NOT ON THE PATIENT AND WHEN THE RELOAD WAS PLACED, THERE WERE TWO DEVICES THAT STARTED TO CALIBRATE, STARTED SPINNING AND WOULD NOT STOP AFTER THE ADAPTER WAS TAKEN OFF. A NEW POWERED STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
23-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Unintended Head Motion |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT ASSISTED LAPAROSCOPIC LOBECTOMY, BEFORE ENTERING THE TROCAR WHICH WAS NOT ON THE PATIENT AND WHEN THE RELOAD WAS PLACED, THERE WERE TWO DEVICES THAT STARTED TO CALIBRATE, STARTED SPINNING AND WOULD NOT STOP AFTER THE ADAPTER WAS TAKEN OFF. A NEW POWERED STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
23-03-2018 | Malfunction | covidien lp llc north haven | 13-04-2018 | ENDO GIA | Unintended Head Motion |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT ASSISTED LAPAROSCOPIC LOBECTOMY, BEFORE ENTERING THE TROCAR WHICH WAS NOT ON THE PATIENT AND WHEN THE RELOAD WAS PLACED, THERE WERE 2 DEVICES THAT STARTED TO CALIBRATE, STARTED SPINNING AND WOULD NOT STOP AFTER THE ADAPTER WAS TAKEN OFF. A NEW POWERED STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
21-02-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | GST BLUE RELOAD, 60MM, 6 ROW | Material Protrusion/Extrusion | A0411 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC GASTRECTOMY, IT WAS FOUND THAT THE SLEDS WERE PROTRUDED BEFORE THE CARTRIDGE WAS LOADED INTO THE JAW. THE CARTRIDGE COULD NOT BE LOADED INTO THE JAW. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
21-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER RADICAL SURGERY, THE STAPLE WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION AND AN INCOMPLETE DISTAL STAPLE LINE WHICH WAS FIXED BY HAND SEWING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER RADICAL SURGERY, THE STAPLE WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION AND AN INCOMPLETE DISTAL STAPLE LINE WHICH WAS FIXED BY HAND SEWING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-07-2017 | Malfunction | covidien lp llc north haven | 16-04-2018 | nan | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE LOADING THE DEVICE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED BY ANOTHER DIFFERENT BATCH OF STAPLER. NO PATIENT INJURY.
15-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THAT THERE WAS AN UNKNOWN ISSUE WITH THE STAPLER. NO OTHER INFORMATION IS KNOWN AT THIS TIME.
23-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY PROCEDURE, AFTER CLAMPING THE TISSUE AND CLOSING THE DEVICE JAWS, THE GREEN BUTTON WAS PUSHED TO FIRE THE DEVICE. THE LOWER SIDE OF THE ROTATION BUTTON WAS AS CLICKED FOR A SLOW FIRING MODE, HOWEVER THE DEVICE DID NOT REACT. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WERE TWO STAPLES MISSING FROM THE SPECIMEN SIDE OF THE FIRST FIRING USING A BLACK GST RELOAD AND GORE SEAM GUARD BUTTRESSING MATERIAL. PROCEDURE WAS COMPLETED WITH THE SAME DEVICE. THERE WERE NO ISSUES ON THE PATIENT SIDE OF THE STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
19-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THERE WERE TWO STAPLES MISSING FROM THE SPECIMEN SIDE OF THE FIRST FIRING USING A BLACK GST RELOAD AND GORE SEAM GUARD BUTTRESSING MATERIAL. PROCEDURE WAS COMPLETED WITH THE SAME DEVICE. THERE WERE NO ISSUES ON THE PATIENT SIDE OF THE STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
27-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LIVER RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT IT DID NOT FIRE AND WAS UNABLE TO SQUEEZE THE HANDLE. THE SCRUB NURSE CHANGED THE DEVICE AND IT DID WORK. NO PATIENT CONDITION REPORTED.
28-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACIC PROCEDURE, THE STAPLE RELOAD WAS NOT CLOSED WHEN PRE-TEST. THEY USED THE OTHER NEW ONES TO COMPLETE THE CASE. NO PATIENT INJURY.
19-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | ILS 25MM, CURVED | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF ESOPHAGEAL CARCINOMA, COULD NOT SQUEEZE DOWN THE FIRING TRIGGER. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
13-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER BEING FIRED, THE DEVICE'S BLADE WAS RETRACTED, BUT THE JAW WAS NOT ABLE TO OPEN. THE JAW OF THE DEVICE LOCKED ON THE TISSUE.
21-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN OPEN COLECTOMY PROCEDURE THE FIRST DEVICE CUT AND DEPLOYED ABOUT TWO THIRDS OF THE STAPLES UNFORMED. THE SECOND DEVICE CUT AND DEPLOYED APPROXIMATELY 1/3 OF THE STAPLES; IT WAS UNKNOWN HOW THEY FORMED. PROCEDURE WAS COMPLETED BY OVERSEWING AND HAVING TO CREATE AN OSTOMY FOR THE PATIENT.
19-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THE DEVICE DEPLOYED MALFORMED STAPLES THAT DID NOT FORM THE COMPLETE "B" SHAPE. THE SURGEON CLIPPED OVER THE MALFORMED STAPLE LINE TO SECURE IT. PROCEDURE WAS COMPLETED WITH THE SAME STAPLER AND A NEW RELOAD. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN OPEN COLECTOMY PROCEDURE THE FIRST DEVICE CUT AND DEPLOYED ABOUT TWO THIRDS OF THE STAPLES UNFORMED. THE SECOND DEVICE CUT AND DEPLOYED APPROXIMATELY 1/3 OF THE STAPLES; IT WAS UNKNOWN HOW THEY FORMED. PROCEDURE WAS COMPLETED BY OVERSEWING AND HAVING TO CREATE AN OSTOMY FOR THE PATIENT.
20-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | ILS 29MM, CURVED | Device Alarm System | A1601 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE DIDN'T MAKE ANY SOUND FEEDBACK AND THE WASHER DIDN'T BREAK. THE OPERATOR REPORTED THAT THE WASHER PUT RESISTANCE. THE INTERVENTION WAS CONCLUDED USING THE COMPLAINT DEVICE BECAUSE IT WAS ABLE TO CUT AND SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | ILS 29MM, CURVED | Physical Resistance |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE DIDN'T MAKE ANY SOUND FEEDBACK AND THE WASHER DIDN'T BREAK. THE OPERATOR REPORTED THAT THE WASHER PUT RESISTANCE. THE INTERVENTION WAS CONCLUDED USING THE COMPLAINT DEVICE BECAUSE IT WAS ABLE TO CUT AND SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
26-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA | Power Problem | A0708 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION OF RECTUM, WHILE PERFORMING DOUBLE STAPLING, THE DEVICE STOPPED WORKING WHILE CLAMPING TISSUE ON THE FIRST FIRING. THE BATTERY WAS RESET TO REMOVE THE DEVICE FROM THE TISSUE, THEN REASSEMBLED THE DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INVOLVEMENT AND OUTCOME.
28-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR), TO CUT THE INFERIOR MESENTERIC VEIN, THE STAPLER AND RELOAD DID NOT FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. NEW DEVICES WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY, DURING 1ST FIRING, THE FIRING KNOB WAS BROKEN. ALL THE PIECES WOULD BE RETURNED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
28-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE/SEGMENTAL RESECTION, AT THE SECOND FIRING SEGMENTAL RESECTION OF THE LUNG, THE HANDLE AND RELOAD WERE USED. THE SURGEON TRIED TO CLAMP THE TISSUE, HOWEVER, A CRACK SOUND WAS HEARD AT PIVOT POINT AND THE CARTRIDGE SEEM TO BREAK. CARTRIDGE WAS REPLACED BY A NEW ONE AND THE FIRING WAS COMPLETED. PATIENT STATUS IS UNKNOWN.
13-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | CONTOUR CURVED CUTTER STAPLER | Break | A0401 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A RECTUM RESECTION SURGERY, THE DEVICE WAS USED. WHEN THE DOCTOR CUT THE RECTUM THE TRIGGER OF THE DEVICE WAS BROKEN. THE SURGERY HAD TO BE CONVERTED INTO RECTUM EXTIRPATION.
13-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | CONTOUR CURVED CUTTER STAPLER | Break | A0401 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A RECTUM RESECTION SURGERY, THE DEVICE WAS USED. WHEN THE DOCTOR CUT THE RECTUM THE TRIGGER OF THE DEVICE WAS BROKEN. THE SURGERY HAD TO BE CONVERTED INTO RECTUM EXTIRPATION.
26-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | IDRIVE | Power Problem | A0708 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION OF RECTUM, WHILE PERFORMING DOUBLE STAPLING, THE DEVICE STOPPED WORKING WHILE CLAMPING TISSUE ON THE FIRST FIRING. THE BATTERY WAS RESET TO REMOVE THE DEVICE FROM THE TISSUE, THEN REASSEMBLED THE DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INVOLVEMENT AND OUTCOME.
27-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, AFTER THE SURGEON FIRED THE DEVICE, THE FIRING STOPPED ON THE HALFWAY. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO INJURY TO THE PATIENT.
27-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | IDRIVE ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, AFTER THE SURGEON FIRED THE DEVICE, THE FIRING STOPPED ON THE HALFWAY. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO INJURY TO THE PATIENT.
27-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, AFTER THE SURGEON FIRED THE DEVICE, THE FIRING STOPPED ON THE HALFWAY. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO INJURY TO THE PATIENT.
01-01-2018 | Injury | ethicon endo-surgery | 16-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT THE SURGEON USED A POWERED VASCULAR STAPLER IN A VATS INFERIOR RIGHT LOBECTOMY. HE FIRST FIRED ON THE VEIN AND THEN ON THE ARTERY WITHOUT PROBLEM. NO BLEEDING WAS NOTICED. AFTER ABOUT 15 MINUTES THE STAPLE LINE ON THE VEIN OPENED AT A POINT AND BLEEDING BEGAN. THE SURGEON IMMEDIATELY CONVERTED TO THORACOTOMY. THE BLEEDING WAS CONTROLLED BUT THE PATIENT RECEIVED A BAG OF BLOOD. THE SURGEON EXPLAINED THAT HE HESITATED TO USE THE STAPLER ON THE VEIN BECAUSE IT WAS BIG AND HE ALREADY EXPERIENCED ONCE THIS PROBLEM. IN THE FUTURE, HE WILL NO MORE USE THE POWERED VASCULAR STAPLER ON BIG VEINS.
01-01-2018 | Injury | ethicon endo-surgery | 16-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USED A POWERED VASCULAR STAPLER IN A VATS INFERIOR RIGHT LOBECTOMY. HE FIRST FIRED ON THE VEIN AND THEN ON THE ARTERY WITHOUT PROBLEM. NO BLEEDING WAS NOTICED. AFTER ABOUT 15 MINUTES THE STAPLE LINE ON THE VEIN OPENED AT A POINT AND BLEEDING BEGAN. THE SURGEON IMMEDIATELY CONVERTED TO THORACOTOMY. THE BLEEDING WAS CONTROLLED BUT THE PATIENT RECEIVED A BAG OF BLOOD. THE SURGEON EXPLAINED THAT HE HESITATED TO USE THE STAPLER ON THE VEIN BECAUSE IT WAS BIG AND HE ALREADY EXPERIENCED ONCE THIS PROBLEM. IN THE FUTURE, HE WILL NO MORE USE THE POWERED VASCULAR STAPLER ON BIG VEINS.
27-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL PANCREATECTOMY, THE RELOAD WAS CONNECTED TO THE DEVICE FOR THE FIRST FIRING. THE SURGEON TRIED TO FIRE THE DEVICE HOWEVER IT COULD NOT AND IT WAS NOTED THAT THE RELOAD WAS PRE-FIRED. THE RELOAD WAS REPLACED AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL PANCREATECTOMY, THE RELOAD WAS CONNECTED TO THE DEVICE FOR THE FIRST FIRING. THE SURGEON TRIED TO FIRE THE DEVICE HOWEVER IT COULD NOT AND IT WAS NOTED THAT THE RELOAD WAS PRE-FIRED. THE RELOAD WAS REPLACED AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL PANCREATECTOMY, THE RELOAD WAS CONNECTED TO THE DEVICE FOR THE FIRST FIRING. THE SURGEON TRIED TO FIRE THE DEVICE HOWEVER IT COULD NOT AND IT WAS NOTED THAT THE RELOAD WAS PRE-FIRED. THE RELOAD WAS REPLACED AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
19-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLES MALFORMED AFTER FIRING. DURING THE WEDGE RESECTION OF THE LUNG, AFTER FIRED, FOUND THE STAPLES MALFORMED, WHEN SEVERING THE PULMONARY FISSURE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
20-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOWER ANTERIOR RESECTION, THE DOCTOR FIRED THE INSTRUMENT BUT THE STAPLES DIDN'T DEPLOY OR FORM. THE DOCTOR HAD TO CUT OUT THE ANVIL AND HAND SEWED THE ANASTOMOSIS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
19-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE STAPLES WERE UNFORMED. THE SURGEON JUDGED THE DEVICE WAS UNCUT THE TARGET TISSUE SINCE THE INTESTINE WAS DISLODGED FROM THE DEVICE AFTER THE FIRING. THEN, THE FIRING HANDLE WAS GRASPED AGAIN, AND THE DEVICE WAS REMOVED FROM THE PATIENT. HOWEVER, THE WASHER REMAINED UNBROKEN IN THE ANVIL SIDE OF THE DEVICE AND THE TARGET TISSUE WAS NOT ANASTOMOSED. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE DID NOT CONVERT TO OPEN. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
27-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL PANCREATECTOMY, THE RELOAD WAS CONNECTED TO THE DEVICE FOR THE FIRST FIRING. THE SURGEON TRIED TO FIRE THE DEVICE HOWEVER IT COULD NOT AND IT WAS NOTED THAT THE RELOAD WAS PRE-FIRED. THE RELOAD WAS REPLACED AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST FIRING OF STOMACH RESECTION IN AN OPEN DISTAL GASTRECTOMY PROCEDURE, THE SURGEON CLAMPED THE TISSUE WITH NO PROBLEM AND TRIED TO PUSH THE FIRING KNOB ,HOWEVER IT WAS LOCKED AND COULD NOT ADVANCE AT ALL. THEY REPLACED IT WITH A NEW STAPLER, AND THE SURGEON CLAMPED THE TISSUE WITH NO PROBLEM AND PUSHED THE FIRING KNOB BY 1CM, HOWEVER, IT WAS LOCKED AND COULD NOT ADVANCE ON THE HALFWAY. THEY REPLACED IT WITH ANOTHER STAPLER, AND THE SURGEON COULD FIRE THE DEVICE WITH NO PROBLEM AS USUAL. THE SURGEON DID NOT RECOGNIZE THE TISSUE WAS THICK AND HARD. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST FIRING OF STOMACH RESECTION IN AN OPEN DISTAL GASTRECTOMY PROCEDURE, THE SURGEON CLAMPED THE TISSUE WITH NO PROBLEM AND TRIED TO PUSH THE FIRING KNOB ,HOWEVER IT WAS LOCKED AND COULD NOT ADVANCE AT ALL. THEY REPLACED IT WITH A NEW STAPLER, AND THE SURGEON CLAMPED THE TISSUE WITH NO PROBLEM AND PUSHED THE FIRING KNOB BY 1CM, HOWEVER, IT WAS LOCKED AND COULD NOT ADVANCE ON THE HALFWAY. THEY REPLACED IT WITH ANOTHER STAPLER, AND THE SURGEON COULD FIRE THE DEVICE WITH NO PROBLEM AS USUAL. THE SURGEON DID NOT RECOGNIZE THE TISSUE WAS THICK AND HARD. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY PROCEDURE, AFTER CLAMPING THE TISSUE AND CLOSING THE DEVICE JAWS, THE GREEN BUTTON WAS PUSHED TO FIRE THE DEVICE. THE LOWER SIDE OF THE ROTATION BUTTON WAS AS CLICKED FOR A SLOW FIRING MODE; HOWEVER THE DEVICE DID NOT REACT. THERE WAS NO PATIENT INJURY.
22-02-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | ILS 29MM, CURVED | Leak/Splash | A0504 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMICOLECTOMY, THE DEVICE WAS USED TO MAKE THE ANASTOMOSIS AND THEY STAPLED THE COLON AND NOTICED THERE WAS A LEAK. THE DEVICE DID CUT APPROPRIATELY BUT THE STAPLE WAS NOT FORMED CORRECT. THEY MADE A NEW ANASTOMOSES AND NO LONG TERM CONSEQUENCES FOR THE PATIENT OCCURRED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-02-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMICOLECTOMY, THE DEVICE WAS USED TO MAKE THE ANASTOMOSIS AND THEY STAPLED THE COLON AND NOTICED THERE WAS A LEAK. THE DEVICE DID CUT APPROPRIATELY BUT THE STAPLE WAS NOT FORMED CORRECT. THEY MADE A NEW ANASTOMOSES AND NO LONG TERM CONSEQUENCES FOR THE PATIENT OCCURRED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
27-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | TA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRECTOMY, WHEN TRANSCENDING THE STOMACH, THE STAPLER DID NOT FIRE, WERE ABLE TO PERFORM A FULL, COMPLETE SQUEEZE OF THE HANDLE, THE HANDLE RETURNED TO THE OPEN POSITION FOLLOWING THE FIRST FULL SQUEEZE AND THE HANDLE DID NOT REMAIN IN THE CLOSED POSITION FOLLOWING THE SECOND COMPLETE SQUEEZE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL PANCREATECTOMY, THE RELOAD WAS CONNECTED TO THE DEVICE FOR THE FIRST FIRING. THE SURGEON TRIED TO FIRE THE DEVICE HOWEVER IT COULD NOT AND IT WAS NOTED THAT THE RELOAD WAS PRE-FIRED. THE RELOAD WAS REPLACED AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE FIRING AT THE STOMACH SLEEVE, THE STAPLER WAS ABLE TO FIRE BUT THE RELOAD TORE THE SEROSA AND THE SURGEON HAD TO OVER SEW THE DEFECT.
27-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE FIRING AT THE STOMACH SLEEVE, THE STAPLER WAS ABLE TO FIRE BUT THE RELOAD TORE THE SEROSA AND THE SURGEON HAD TO OVER SEW THE DEFECT.
16-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER WAS UNABLE TO FIRE WHEN IT WAS CONNECTED TO AN APPLICATOR. THERE WAS NO PATIENT INVOLVED.
27-02-2018 | Injury | ethicon endo-surgery | 16-04-2018 | POWERED PLUS COMPACT 45 | Patient-Device Incompatibility | A0101 | Not Applicable | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS).
27-02-2018 | Injury | ethicon endo-surgery | 16-04-2018 | POWERED PLUS COMPACT 45 | Patient-Device Incompatibility | A0101 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS).
28-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE THE DEVICE WAS BEING USED TO LOCK THE JAW, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY CHANGED THE HANDLE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
21-03-2018 | Malfunction | ethicon endo-surgery | 16-04-2018 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC LOW RECTAL CANCER RESECTION, WHEN SEVERING THE RECTUM, COULD NOT FIRE, WITHOUT AUDIBLE FEEDBACK AND BREAKAWAY WASHER WAS NOT CUT THROUGH. THE STAPLES MALFORMED. THE PATIENT LOST BLOOD, SUTURE WAS USED TO SEW THE WOUND. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS STANDARD FOR THIS PROCEDURE AND IS NOT EXTENDED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
06-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | POWERED PLUS COMPACT 45 | Device Displays Incorrect Message | A090201 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS).
06-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | POWERED PLUS COMPACT 45 | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS).
13-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VATS LOBECTOMY, AFTER BEING FIRED, THE DEVICE'S BLADE WAS RETRACTED, BUT THE JAW WAS NOT ABLE TO OPEN. THE JAW OF THE DEVICE LOCKED ON THE TISSUE, THE LUNG WAS RESECTED AND THE DEVICE WAS PULLED BACK THROUGH THE PORT. A NEW DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO INJURY.
13-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | POWERED PLUS COMPACT 45 | Fluid/Blood Leak | A050401 | Not Applicable | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS).
13-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | POWERED PLUS COMPACT 45 | Fluid/Blood Leak | A050401 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS).
26-03-2018 | Malfunction | covidien lp llc north haven | 16-04-2018 | IDRIVE ULTRA | Power Problem | A0708 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION OF RECTUM, WHILE PERFORMING DOUBLE STAPLING, THE DEVICE STOPPED WORKING WHILE CLAMPING TISSUE ON THE FIRST FIRING. THE BATTERY WAS RESET TO REMOVE THE DEVICE FROM THE TISSUE, THEN REASSEMBLED THE DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INVOLVEMENT AND OUTCOME.
28-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | EEA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SIGMOIDECTOMY, WHEN FIRING THE DEVICE TO PLACE THE ANASTOMOSIS, THE STAPLER DID NOT FIRE, THE SPACE BETWEEN THE CARTRIDGE AND THE ANVIL CLOSED AND THE GREEN BAR WAS VISIBLE IN THE INDICATOR WINDOW, THE KNIFE BLADE DID NOT CUT, THE ANVIL CANNOT BE TILTED AFTER FIRING AND WAS NOT BENT. AS STATED, THEY OPENED ANOTHER STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR), TO CUT THE INFERIOR MESENTERIC VEIN, THE STAPLER AND RELOAD DID NOT FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRECTOMY, WHILE APPLYING AN ANASTOMOSE STOMACH TO THE ESOPHAGUS, THERE WAS BLEEDING FROM THE STAPLE LINE. THEY USED ANOTHER STAPLER TO COMPLETE THE CASE.
13-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | POWERED PLUS COMPACT 45 | Fluid/Blood Leak | A050401 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS).
28-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE THE DEVICE WAS BEING USED TO LOCK THE JAW, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY CHANGED THE HANDLE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-03-2018 | Injury | ethicon endo-surgery | 16-04-2018 | POWERED PLUS COMPACT 45 | Appropriate Term/Code Not Available |  | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON USES SINCE 3 WEEKS THE DEVICES PCE45A. HE REPORTS THAT ALL HIS PATIENTS GOT EMPHYSEMA (LEAK) THAT PREVENT TO REMOVE THE DRAINS AND INCREASE THE AVERAGE LENGTH OF STAY (4 TO 8 WEEKS).
28-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE FIRING THE 3RD CARTRIDGE, THE HANDLE WAS SQUEEZED AND GOT FREE BUT UNABLE TO FIRE THE RELOAD. THE GREEN BUTTON WAS ABLE TO PRESS. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | 195 mcdermott rd | 16-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACIC PROCEDURE, THE STAPLE RELOAD WAS NOT CLOSED WHEN PRE-TEST. THEY USED THE OTHER NEW ONES TO COMPLETE THE CASE. NO PATIENT INJURY.
28-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACIC PROCEDURE, THE STAPLE RELOAD WAS NOT CLOSED WHEN PRE-TEST. THEY USED THE OTHER NEW ONES TO COMPLETE THE CASE. NO PATIENT INJURY.
22-01-2018 | Injury | ethicon endo-surgery | 16-04-2018 | FLEX 60 ARTICULATING | Partial Blockage | A140902 | No Code Available | Event Description: IT WAS REPORTED THAT THERE WAS A FAILURE IN ECHELON FLEX 60 STAPLER THAT WAS USED IN SURGICAL PROCEDURE, STUCK IN THE MECHANISM OF ECHELON 60, DUE TO THIS THEY HAD TO CONVERT THE SURGERY. FOR WHICH ALSO THE DOCTORS DECIDED, DISCARD THE STAPLER. MEDICAL REPORT: WISHING YOU SUCCESS IN YOUR DAILY TASKS, WE INFORM YOU THE FOLLOWING. THAT DURING THE SURGICAL PROCEDURE PERFORMED ON (B)(6) OF THE CURRENT, THAT CONSISTED OF A LAPAROSCOPIC SUBTOTAL GASTRECTOMY AND THAT DURING THE SECTION OF THE DUODENUM WITH AN ECHELON STAPLING CLAMP WAS BLOCKED AND IT WAS CAUGHT THE DUODENUM, WHICH WAS IMPOSSIBLE TO UNBLOCK, THIS LED US TO CONVERT THE SURGERY TO A CONVENTIONAL (LAPAROTOMY);
22-01-2018 | Injury | ethicon endo-surgery | 16-04-2018 | FLEX 60 ARTICULATING | Entrapment of Device | A150208 | No Code Available | Event Description: IT WAS REPORTED THAT THERE WAS A FAILURE IN ECHELON FLEX 60 STAPLER THAT WAS USED IN SURGICAL PROCEDURE, STUCK IN THE MECHANISM OF ECHELON 60, DUE TO THIS THEY HAD TO CONVERT THE SURGERY. FOR WHICH ALSO THE DOCTORS DECIDED, DISCARD THE STAPLER. MEDICAL REPORT: WISHING YOU SUCCESS IN YOUR DAILY TASKS, WE INFORM YOU THE FOLLOWING. THAT DURING THE SURGICAL PROCEDURE PERFORMED ON (B)(6) OF THE CURRENT, THAT CONSISTED OF A LAPAROSCOPIC SUBTOTAL GASTRECTOMY AND THAT DURING THE SECTION OF THE DUODENUM WITH AN ECHELON STAPLING CLAMP WAS BLOCKED AND IT WAS CAUGHT THE DUODENUM, WHICH WAS IMPOSSIBLE TO UNBLOCK, THIS LED US TO CONVERT THE SURGERY TO A CONVENTIONAL (LAPAROTOMY);
22-01-2018 | Injury | ethicon endo-surgery | 16-04-2018 | FLEX 60 ARTICULATING | Difficult to Remove | A150207 | No Code Available | Event Description: IT WAS REPORTED THAT THERE WAS A FAILURE IN ECHELON FLEX 60 STAPLER THAT WAS USED IN SURGICAL PROCEDURE, STUCK IN THE MECHANISM OF ECHELON 60, DUE TO THIS THEY HAD TO CONVERT THE SURGERY. FOR WHICH ALSO THE DOCTORS DECIDED, DISCARD THE STAPLER. MEDICAL REPORT: WISHING YOU SUCCESS IN YOUR DAILY TASKS, WE INFORM YOU THE FOLLOWING. THAT DURING THE SURGICAL PROCEDURE PERFORMED ON (B)(6) OF THE CURRENT, THAT CONSISTED OF A LAPAROSCOPIC SUBTOTAL GASTRECTOMY AND THAT DURING THE SECTION OF THE DUODENUM WITH AN ECHELON STAPLING CLAMP WAS BLOCKED AND IT WAS CAUGHT THE DUODENUM, WHICH WAS IMPOSSIBLE TO UNBLOCK, THIS LED US TO CONVERT THE SURGERY TO A CONVENTIONAL (LAPAROTOMY);
20-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT OPEN HEMI-COLECTOMY, WHILE STAPLING OF THE RIGHT COLON, THE STAPLER AND THE RELOAD JAMMED AND WOULD NOT FIRE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
20-03-2018 | Malfunction | us surgical puerto rico | 16-04-2018 | GIA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT OPEN HEMI-COLECTOMY, WHILE STAPLING OF THE RIGHT COLON, THE STAPLER AND THE RELOAD JAMMED AND WOULD NOT FIRE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
27-03-2018 | Injury | ethicon endo-surgery | 17-04-2018 | FLEX 60 ARTICULATING | Device Operates Differently Than Expected |  | Blood Loss | Event Description: IT WAS REPORTED THAT THERE WAS BLEEDING AFTER SURGERY. DURING THE ENDOSCOPIC RIGHT SUPERIOR PNEUMONECTOMY, USED EC60A AND ECR60W(1ST FIRING) TO SEVERING THE RIGHT SUPERIOR PULMONARY VEIN, AND NO ABNORMAL SITUATION DURING THE SURGERY. ONE DAY AFTER SURGERY, THE PATIENT'S HEMOGLOBIN WAS ABOUT 5. DID THE 2ND SURGERY, FOUND THE BLEEDING WAS FROM THE ANASTOMOTIC STOMA OF THE RIGHT SUPERIOR PULMONARY VEIN. SUTURE WAS USED TO SEW THE WOUND AND STOPPED THE BLEEDING. THE PATIENT WAS WITH BLOOD TRANSFUSION. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
27-03-2018 | Injury | ethicon endo-surgery | 17-04-2018 | FLEX 60 ARTICULATING | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT THERE WAS BLEEDING AFTER SURGERY. DURING THE ENDOSCOPIC RIGHT SUPERIOR PNEUMONECTOMY, USED EC60A AND ECR60W(1ST FIRING) TO SEVERING THE RIGHT SUPERIOR PULMONARY VEIN, AND NO ABNORMAL SITUATION DURING THE SURGERY. ONE DAY AFTER SURGERY, THE PATIENT'S HEMOGLOBIN WAS ABOUT 5. DID THE 2ND SURGERY, FOUND THE BLEEDING WAS FROM THE ANASTOMOTIC STOMA OF THE RIGHT SUPERIOR PULMONARY VEIN. SUTURE WAS USED TO SEW THE WOUND AND STOPPED THE BLEEDING. THE PATIENT WAS WITH BLOOD TRANSFUSION. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
26-02-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: COULD NOT FIRE. DURING THE RADICAL GASTRECTOMY FOR GASTRIC CANCER, COULD NOT FIRE ON THE 2ND FIRING WITH ONE ALOUD SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
23-02-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE THE DEVICE LOCKED ONTO TISSUE AND WAS REMOVED BY USING THE MANUAL OVER-RIDE. IT IS UNKNOWN IF THE REVERSE BUTTON WAS USED. IT IS UNKNOWN WHETHER THE DEVICE ADVANCED, CUT OR STAPLED, ONLY THAT IT JAMMED. NO ADDITIONAL INFORMATION IS AVAILABLE. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
27-02-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DOCTOR USED AN ETS45 IN A LAP APPY. STAPLER LOCKED OUT. RELOAD WAS REPLACED AND WORKED AS INTENDED. THERE WERE NO PATIENT CONSEQUENCES.
27-02-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DOCTOR USED AN ETS45 IN A LAP APPY. STAPLER LOCKED OUT. RELOAD WAS REPLACED AND WORKED AS INTENDED. THERE WERE NO PATIENT CONSEQUENCES.
23-01-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | RELOADS FOR LINEAR CUTTERS | Defective Component | A0202 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY "THE LOAD IS DEFECTIVE, WOULD NOT FIRE". THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED. THERE IS NO ADDITIONAL INFORMATION.
23-01-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | RELOADS FOR LINEAR CUTTERS | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY "THE LOAD IS DEFECTIVE, WOULD NOT FIRE". THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED. THERE IS NO ADDITIONAL INFORMATION.
13-03-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | POWERED 60 ECHELON +, 280MM SHAFT | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER MISFIRED, UNKNOWN FIRING, UNKNOWN RELOAD, UNKNOWN BUTTRESSING MATERIAL. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
29-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE POSITIONING THE STAPLER FOR LUNG LOBE ACCESS IN A LUNG LOBECTOMY PROCEDURE, THE STAPLER BROKE AT THE RELOAD AND SHAFT CONNECTION POINT, AND IT REMAINED ATTACHED. THIS HAPPENED WHILE THE SURGEON WAS MANIPULATING THE STAPLER THROUGH THE PATIENT¿S THORACIC CAVITY, WITH NO TROCAR, AND HIS OTHER HAND WAS ATTEMPTING TO BEND THE SHAFT UPWARD TO ACCESS A PORTION OF THE LUNG THAT WAS TOWARD THE APEX OF THE THORACIC WALL. THERE WAS NO PATIENT INJURY.
28-02-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: THE DEVICE HAS BEEN USED FOR THE SUTURE OF THE RIGHT SUPRAHEPATIC VEIN: THE SUTURE WAS DONE WITH THE STAPLER ATS45, BUT IT HAS TURNED OUT TO BE NOT EFFICIENT SO THE SURGEON DID A LINEAR SUTURE ON THE VEIN CAVE. THE SURGEON USED THE SATINSKY CLAMPS AND PROLENE 5/0 TO CONTROL THE BLEEDING AND CLOSE THE HOLE. ONCE STOPPED THE BLEEDING, THE SURGEON ENDED THE INTERVENTION.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | EEA ORVIL | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, THE DEVICE WAS INSERTED DOWN THE ESOPHAGUS. THE SURGEON GRASPED THE DELIVERY TUBE THROUGH THE GASTROTOMY; HE FELT NO TENSION AND CONTINUED TO PULL AS NORMAL. THE SURGEON REALIZED THE ANVIL WAS NOT ATTACHED TO THE DELIVERY TUBE AS THEY REMOVED THE TUBE. THERE WERE NO SUTURES OR WHITE CONNECTION PIECE ON THE TUBE. THE SURGEON DID NOT CUT THE SUTURE TO REMOVE THE ANVIL AND WAS STILL IN THE PATIENT¿S ESOPHAGUS. WHEN LOOKING AT THE ANVIL, THE WHITE CONNECTION PIECE CONNECTING THE ANVIL, THE SUTURE AND THE DELIVERY TUBE WAS FULLY INTACT. IT LOOKED LIKE THE MATERIAL THAT WAS USED TO SECURE THE WHITE CONNECTION PIECE TO THE DELIVERY TUBE WAS MISSING, ALLOWING THE ANVIL TO DISENGAGE TO THE DELIVERY TUBE UNDER NORMAL PROCEDURAL STEPS. THE ANVIL WAS SUCCESSFULLY REMOVED WITH THE RETRIEVAL STRINGS. THE SURGEON COMPLETED THE PROCEDURE WITH A NEW ANVIL. THERE WAS NO REPORTED PATIENT INJURY.
28-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PARTIAL HEPATECTOMY, THE STAPLER WAS UNABLE TO FIRE AND THE REPLACEMENT MAGAZINE COULD NOT BE USED. THE SURGEON WAS ALSO UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. THEY REPLACED THE STAPLER TO FIRE THE CLIPS NORMALLY AND TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | EEA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION PROCEDURE, AFTER LOADING THE RELOAD INTO THE STAPLER AND AFTER PRESSING THE GREEN BUTTON, THE RELOAD DID NOT FIRE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PARTIAL HEPATECTOMY, THE STAPLER WAS UNABLE TO FIRE AND THE REPLACEMENT MAGAZINE COULD NOT BE USED. THE SURGEON WAS ALSO UNABLE TO SQUEEZE THE HANDLE BUT HE WAS ABLE TO PRESS THE GREEN BUTTON. THEY REPLACED THE STAPLER TO FIRE THE CLIPS NORMALLY AND TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) RIGHT WEDGE RESECTION, WHEN RESECTING THE LUNG TISSUE FOR A BIOPSY, THE USER ASKED FOR A PURPLE CARTRIDGE TO BE USED WITH THE DEVICE. THE CARTRIDGE WAS LOADED PROPERLY, CYCLED ON THE DEVICE AND PASSED ONTO THE USER. WHEN THE USER PUT IT ONTO THE PART OF LUNG THAT NEEDED RESECTING AND ARTICULATED THE CARTRIDGE AND CLOSED THE JAWS ON TISSUE BUT THE GREEN LIGHT INDICATOR DID NOT COME ON. THE USER OPENED AND CLOSED THE JAWS TO COMPRESS THE TISSUE SO THE FLUIDS ARE DISPERSED BUT THE GREEN LIGHT STILL DID NOT COME ON WHICH WAS PARTLY DOWN TO THE TISSUE BEING TOO THICK. WHEN THE SCRUB NURSE TRIED TO REMOVE THE PURPLE CARTRIDGE SO THEY CAN REPLACE IT WITH THE BLACK CARTRIDGE, THE USER WAS NOT ABLE TO STRAIGHTEN THE CARTRIDGE. THE NURSE TOOK THE CARTRIDGE OFF SOME HOW AND THE USER FIRED TWO BLACK CARTRIDGES WHICH WORKED ABSOLUTELY FINE. WHEN THE USER ASKED FOR THE SAME PURPLE CARTRIDGE, IT DID NOT REGISTER WITH THE DEVICE. AS STATED, THERE WAS A PROBLEM WITH UNLOADING THE DEVICE, THE STAPLER DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS SQUEEZED. THE RELOAD WAS REPLACED BY ANOTHER CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APICAL BLEB RESECTION, WHILE BEING APPLIED ON THE LUNG, THE STAPLER WAS UNABLE TO FIRE BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE FIRST CARTRIDGE USED FIRED ROUGHLY HALF WAY AND THEN IT STOPPED. THE SURGEON REMOVED THE STAPLER. THE STAPLE LINE WAS ALSO FINE, THE DEVICE JUST DIDN'T FIRE ALL THE WAY. THE 2ND CARTRIDGE WAS LOADED AND IT COULD NOT BE CYCLED. LEADING TO AN INCOMPLETE STAPLE LINE. A NEW HANDLE AND CARTRIDGE WERE OPENED AND USED WITHOUT INCIDENT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, WHEN ON LOBECTOMY ON THE LUNG TUMOR, THE STAPLER AND LOAD DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | RELIATACK | Display or Visual Feedback Problem | A0902 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VENTRAL HERNIA, TOWARDS THE END OF THE PROCEDURE, WHEN THE MESH WAS FIXATED INTO THE ABDOMINAL WALL, A TRIGGER PULL CAUSED A DOUBLE FIRING. THEREFORE THE SURGEON REMOVED THE DEVICE AND OPENED A SECOND DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | EEA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC STAPLER HEMORRHOIDECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO INTERMEDIATE LAYER BETWEEN MUCOSA AND SUB-MUCOSA, THE DEVICE DID NOT FIRE. THE SURGEON USED ONE MORE HEMORRHOID STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | IDRIVE ULTRA | Retraction Problem | A0510 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY/GASTRIC TUBE PROCEDURE, AT THE SECOND FIRING, DUE TO ADHESION OD ESOPHAGUS AND LUNG, THE SURGEON RESECTED THE PULMONARY PARENCHYMA WITH THE RELOAD. THE SURGEON INITIALLY FIRED THE DEVICE AND PUSHED THE SILVER BUTTON, HOWEVER, THE DEVICE COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION AND COULD NOT OPEN THE JAWS. THEN THE BLUE BUTTON AND SILVER BUTTON WERE PUSHED AT THE SAME TIME, HOWEVER, THE JAWS COULD NOT RETURN TO THE STRAIGHT POSITION AND COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION. RE-INSERTED THE BATTERY TO THE HANDLE (ADAPTER WAS STILL CONNECTED) AND THE KNIFE COULD RETURN TO THE INITIAL POSITION AND COULD OPEN THE JAWS BUT COULD NOT RETURN THE JAWS TO THE STRAIGHT POSITION. PUSHED THE TOGGLE, AND THE JAWS RETURNED TO ITS STRAIGHT POSITION AND COULD REMOVE THE CARTRIDGE FROM THE ADAPTER. AFTER THAT, THE DEVICE WORKED FINE AND FOUR FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | IDRIVE ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY/GASTRIC TUBE PROCEDURE, AT THE SECOND FIRING, DUE TO ADHESION OD ESOPHAGUS AND LUNG, THE SURGEON RESECTED THE PULMONARY PARENCHYMA WITH THE RELOAD. THE SURGEON INITIALLY FIRED THE DEVICE AND PUSHED THE SILVER BUTTON, HOWEVER, THE DEVICE COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION AND COULD NOT OPEN THE JAWS. THEN THE BLUE BUTTON AND SILVER BUTTON WERE PUSHED AT THE SAME TIME, HOWEVER, THE JAWS COULD NOT RETURN TO THE STRAIGHT POSITION AND COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION. RE-INSERTED THE BATTERY TO THE HANDLE (ADAPTER WAS STILL CONNECTED) AND THE KNIFE COULD RETURN TO THE INITIAL POSITION AND COULD OPEN THE JAWS BUT COULD NOT RETURN THE JAWS TO THE STRAIGHT POSITION. PUSHED THE TOGGLE, AND THE JAWS RETURNED TO ITS STRAIGHT POSITION AND COULD REMOVE THE CARTRIDGE FROM THE ADAPTER. AFTER THAT, THE DEVICE WORKED FINE AND FOUR FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | IDRIVE ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY/GASTRIC TUBE PROCEDURE, AT THE SECOND FIRING, DUE TO ADHESION OD ESOPHAGUS AND LUNG, THE SURGEON RESECTED THE PULMONARY PARENCHYMA WITH THE RELOAD. THE SURGEON INITIALLY FIRED THE DEVICE AND PUSHED THE SILVER BUTTON, HOWEVER, THE DEVICE COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION AND COULD NOT OPEN THE JAWS. THEN THE BLUE BUTTON AND SILVER BUTTON WERE PUSHED AT THE SAME TIME, HOWEVER, THE JAWS COULD NOT RETURN TO THE STRAIGHT POSITION AND COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION. RE-INSERTED THE BATTERY TO THE HANDLE (ADAPTER WAS STILL CONNECTED) AND THE KNIFE COULD RETURN TO THE INITIAL POSITION AND COULD OPEN THE JAWS BUT COULD NOT RETURN THE JAWS TO THE STRAIGHT POSITION. PUSHED THE TOGGLE, AND THE JAWS RETURNED TO ITS STRAIGHT POSITION AND COULD REMOVE THE CARTRIDGE FROM THE ADAPTER. AFTER THAT, THE DEVICE WORKED FINE AND FOUR FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APICAL BLEB RESECTION, WHILE BEING APPLIED ON THE LUNG, THE STAPLER WAS UNABLE TO FIRE BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE FIRST CARTRIDGE USED FIRED ROUGHLY HALF WAY AND THEN IT STOPPED. THE SURGEON REMOVED THE STAPLER. THE STAPLE LINE WAS ALSO FINE, THE DEVICE JUST DIDN'T FIRE ALL THE WAY. THE 2ND CARTRIDGE WAS LOADED AND IT COULD NOT BE CYCLED. LEADING TO AN INCOMPLETE STAPLE LINE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | EEA | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY PROCEDURE/ROUX-EN-Y, THE BLUNT TROCAR COULD NOT BE ATTACHED TO THE ANVIL CENTER ROD. THE ROOT OF THE TROCAR WAS SLIGHTLY THINNER THAN USUAL. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | EEA | Product Quality Problem | A0201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY PROCEDURE/ROUX-EN-Y, THE BLUNT TROCAR COULD NOT BE ATTACHED TO THE ANVIL CENTER ROD. THE ROOT OF THE TROCAR WAS SLIGHTLY THINNER THAN USUAL. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | SIGNIA | Overheating of Device | A1005 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE, THE STATUS INDICATOR LIGHTS, THE FIVE WHITE LIGHTS, THE HANDLE INDICATOR, THE ADAPTER INDICATOR AND THE RELOAD INDICATOR WERE OFF. AS STATED, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD; THE STAPLER DID NOT WAKE UP WHEN ONE RELOAD WAS NEEDED. THE HANDLE WAS CHANGED AND IT REMAINED HOT FOR THREE TO FOUR HOURS. THERE WAS NO PATIENT INVOLVEMENT.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | SIGNIA | Failure to Sense | A070908 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE, THE STATUS INDICATOR LIGHTS, THE FIVE WHITE LIGHTS, THE HANDLE INDICATOR, THE ADAPTER INDICATOR AND THE RELOAD INDICATOR WERE OFF. AS STATED, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD; THE STAPLER DID NOT WAKE UP WHEN ONE RELOAD WAS NEEDED. THE HANDLE WAS CHANGED AND IT REMAINED HOT FOR THREE TO FOUR HOURS. THERE WAS NO PATIENT INVOLVEMENT.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE, THE STATUS INDICATOR LIGHTS, THE FIVE WHITE LIGHTS, THE HANDLE INDICATOR, THE ADAPTER INDICATOR AND THE RELOAD INDICATOR WERE OFF. AS STATED, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD; THE STAPLER DID NOT WAKE UP WHEN ONE RELOAD WAS NEEDED. THE HANDLE WAS CHANGED AND IT REMAINED HOT FOR THREE TO FOUR HOURS. THERE WAS NO PATIENT INVOLVEMENT.
15-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/LOBECTOMY, DURING FIRING, THE SURGEON PUSHED THE GREEN BUTTON AND TRIED TO FIRE THE DEVICE, HOWEVER, IT WOULD NOT FIRE. RE-CONNECTED THE SAME CARTRIDGE TO THE ADAPTER AND RE-TRIED THE FIRING BUT THE SAME EVENT OCCURRED. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT UPPER LOBE RESECTION, WHILE RESECTING, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE COULD NOT FIRE. THEY REPLACED THE HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT UPPER LOBE RESECTION, WHILE RESECTING, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE COULD NOT FIRE. THEY REPLACED THE HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Retraction Problem | A0510 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY/GASTRIC TUBE PROCEDURE, AT THE SECOND FIRING, DUE TO ADHESION OD ESOPHAGUS AND LUNG, THE SURGEON RESECTED THE PULMONARY PARENCHYMA WITH THE RELOAD. THE SURGEON INITIALLY FIRED THE DEVICE AND PUSHED THE SILVER BUTTON, HOWEVER, THE DEVICE COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION AND COULD NOT OPEN THE JAWS. THEN THE BLUE BUTTON AND SILVER BUTTON WERE PUSHED AT THE SAME TIME, HOWEVER, THE JAWS COULD NOT RETURN TO THE STRAIGHT POSITION AND COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION. RE-INSERTED THE BATTERY TO THE HANDLE (ADAPTER WAS STILL CONNECTED) AND THE KNIFE COULD RETURN TO THE INITIAL POSITION AND COULD OPEN THE JAWS BUT COULD NOT RETURN THE JAWS TO THE STRAIGHT POSITION. PUSHED THE TOGGLE, AND THE JAWS RETURNED TO ITS STRAIGHT POSITION AND COULD REMOVE THE CARTRIDGE FROM THE ADAPTER. AFTER THAT, THE DEVICE WORKED FINE AND FOUR FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY/GASTRIC TUBE PROCEDURE, AT THE SECOND FIRING, DUE TO ADHESION OD ESOPHAGUS AND LUNG, THE SURGEON RESECTED THE PULMONARY PARENCHYMA WITH THE RELOAD. THE SURGEON INITIALLY FIRED THE DEVICE AND PUSHED THE SILVER BUTTON, HOWEVER, THE DEVICE COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION AND COULD NOT OPEN THE JAWS. THEN THE BLUE BUTTON AND SILVER BUTTON WERE PUSHED AT THE SAME TIME, HOWEVER, THE JAWS COULD NOT RETURN TO THE STRAIGHT POSITION AND COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION. RE-INSERTED THE BATTERY TO THE HANDLE (ADAPTER WAS STILL CONNECTED) AND THE KNIFE COULD RETURN TO THE INITIAL POSITION AND COULD OPEN THE JAWS BUT COULD NOT RETURN THE JAWS TO THE STRAIGHT POSITION. PUSHED THE TOGGLE, AND THE JAWS RETURNED TO ITS STRAIGHT POSITION AND COULD REMOVE THE CARTRIDGE FROM THE ADAPTER. AFTER THAT, THE DEVICE WORKED FINE AND FOUR FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY/GASTRIC TUBE PROCEDURE, AT THE SECOND FIRING, DUE TO ADHESION OD ESOPHAGUS AND LUNG, THE SURGEON RESECTED THE PULMONARY PARENCHYMA WITH THE RELOAD. THE SURGEON INITIALLY FIRED THE DEVICE AND PUSHED THE SILVER BUTTON, HOWEVER, THE DEVICE COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION AND COULD NOT OPEN THE JAWS. THEN THE BLUE BUTTON AND SILVER BUTTON WERE PUSHED AT THE SAME TIME, HOWEVER, THE JAWS COULD NOT RETURN TO THE STRAIGHT POSITION AND COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION. RE-INSERTED THE BATTERY TO THE HANDLE (ADAPTER WAS STILL CONNECTED) AND THE KNIFE COULD RETURN TO THE INITIAL POSITION AND COULD OPEN THE JAWS BUT COULD NOT RETURN THE JAWS TO THE STRAIGHT POSITION. PUSHED THE TOGGLE, AND THE JAWS RETURNED TO ITS STRAIGHT POSITION AND COULD REMOVE THE CARTRIDGE FROM THE ADAPTER. AFTER THAT, THE DEVICE WORKED FINE AND FOUR FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Injury | us surgical puerto rico | 17-04-2018 | GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE SURGEON FOUND A LEAK ON THE PATIENT¿S LINEAR ANASTOMOTIC STAPLE SITE FOLLOWING A HARTMANN¿S COLOSTOMY CLOSURE TWO DAYS POST PROCEDURE. THE SURGEON TOOK THE PATIENT BACK TO THE OPERATING ROOM TO FIX THE LEAK, AND HE FOUND THAT THE PATIENT HAS FECAL LEAKAGE FROM THE SITE AND HAS FECES IN THE ABDOMINAL CAVITY. THE NURSE THINKS THE STAPLE LINE WAS COMPROMISED IN THE MIDDLE. THE PATIENT HAD AN EXTENDED HOSPITALIZATION BUT IT WAS UNKNOWN HOW LONG. THE PATIENT WAS BROUGHT TO THE ICU FOLLOWING THE SECOND PROCEDURE.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | TRI-STAPLE 2.0 | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION WITH NORMAL THICKNESS TISSUE, THE STAPLER WAS NOT ABLE TO FIRE AND HAS PRODUCED AN INCOMPLETE STAPLE LINE. THE LITTLE BLACK PARTS OF THE RELOAD WERE THOUGHT TO HAVE BROKEN OFF AND GONE INTO THE PATIENT¿S CAVITY, BUT THEY ARE UNSURE WHERE IT HAS COME FROM. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION WITH NORMAL THICKNESS TISSUE, THE STAPLER WAS NOT ABLE TO FIRE AND HAS PRODUCED AN INCOMPLETE STAPLE LINE. THE LITTLE BLACK PARTS OF THE RELOAD WERE THOUGHT TO HAVE BROKEN OFF AND GONE INTO THE PATIENT¿S CAVITY, BUT THEY ARE UNSURE WHERE IT HAS COME FROM. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Retraction Problem | A0510 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY/GASTRIC TUBE PROCEDURE, AT THE SECOND FIRING, DUE TO ADHESION OF ESOPHAGUS AND LUNG, THE SURGEON RESECTED THE PULMONARY PARENCHYMA WITH THE RELOAD. THE SURGEON INITIALLY FIRED THE DEVICE AND PUSHED THE SILVER BUTTON, HOWEVER, THE DEVICE COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION AND COULD NOT OPEN THE JAWS. THEN THE BLUE BUTTON AND SILVER BUTTON WERE PUSHED AT THE SAME TIME, HOWEVER, THE JAWS COULD NOT RETURN TO THE STRAIGHT POSITION AND COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION. RE-INSERTED THE BATTERY TO THE HANDLE (ADAPTER WAS STILL CONNECTED) AND THE KNIFE COULD RETURN TO THE INITIAL POSITION AND COULD OPEN THE JAWS BUT COULD NOT RETURN THE JAWS TO THE STRAIGHT POSITION. PUSHED THE TOGGLE, AND THE JAWS RETURNED TO ITS STRAIGHT POSITION AND COULD REMOVE THE CARTRIDGE FROM THE ADAPTER. AFTER THAT, THE DEVICE WORKED FINE AND FOUR FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY/GASTRIC TUBE PROCEDURE, AT THE SECOND FIRING, DUE TO ADHESION OF ESOPHAGUS AND LUNG, THE SURGEON RESECTED THE PULMONARY PARENCHYMA WITH THE RELOAD. THE SURGEON INITIALLY FIRED THE DEVICE AND PUSHED THE SILVER BUTTON, HOWEVER, THE DEVICE COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION AND COULD NOT OPEN THE JAWS. THEN THE BLUE BUTTON AND SILVER BUTTON WERE PUSHED AT THE SAME TIME, HOWEVER, THE JAWS COULD NOT RETURN TO THE STRAIGHT POSITION AND COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION. RE-INSERTED THE BATTERY TO THE HANDLE (ADAPTER WAS STILL CONNECTED) AND THE KNIFE COULD RETURN TO THE INITIAL POSITION AND COULD OPEN THE JAWS BUT COULD NOT RETURN THE JAWS TO THE STRAIGHT POSITION. PUSHED THE TOGGLE, AND THE JAWS RETURNED TO ITS STRAIGHT POSITION AND COULD REMOVE THE CARTRIDGE FROM THE ADAPTER. AFTER THAT, THE DEVICE WORKED FINE AND FOUR FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY/GASTRIC TUBE PROCEDURE, AT THE SECOND FIRING, DUE TO ADHESION OF ESOPHAGUS AND LUNG, THE SURGEON RESECTED THE PULMONARY PARENCHYMA WITH THE RELOAD. THE SURGEON INITIALLY FIRED THE DEVICE AND PUSHED THE SILVER BUTTON, HOWEVER, THE DEVICE COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION AND COULD NOT OPEN THE JAWS. THEN THE BLUE BUTTON AND SILVER BUTTON WERE PUSHED AT THE SAME TIME, HOWEVER, THE JAWS COULD NOT RETURN TO THE STRAIGHT POSITION AND COULD NOT RETURN THE KNIFE TO THE INITIAL POSITION. RE-INSERTED THE BATTERY TO THE HANDLE (ADAPTER WAS STILL CONNECTED) AND THE KNIFE COULD RETURN TO THE INITIAL POSITION AND COULD OPEN THE JAWS BUT COULD NOT RETURN THE JAWS TO THE STRAIGHT POSITION. PUSHED THE TOGGLE, AND THE JAWS RETURNED TO ITS STRAIGHT POSITION AND COULD REMOVE THE CARTRIDGE FROM THE ADAPTER. AFTER THAT, THE DEVICE WORKED FINE AND FOUR FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE, ON THE FIRST FIRING TO BILE DUCT DURING A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE, THE SURGEON TRIED TO FIRE THE DEVICE, BUT COULD NOT SQUEEZE THE HANDLE. THE SALES REP CONFIRMED THAT STAPLES AT THE ROOT OF THE JAWS WERE COMING UP. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE, ON THE FIRST FIRING TO BILE DUCT DURING A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE, THE SURGEON TRIED TO FIRE THE DEVICE, BUT COULD NOT SQUEEZE THE HANDLE. THE SALES REP CONFIRMED THAT STAPLES AT THE ROOT OF THE JAWS WERE COMING UP. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, THE SURGEON APPROACHED A BLOOD VESSEL WITH THE DEVICE ONCE, BUT REMOVED IT WITHOUT FIRING DUE TO INSUFFICIENT DISSECTION OF THE VESSEL. THE DEVICE WAS LEFT IN THE CLEAN FIELD FOR LATER FIRING. AFTER THAT, THE RELOAD WAS LOADED ONTO THE HANDLE WITHOUT PROBLEM, BUT WHEN SURGEON TRIED TO FIRE THE DEVICE TO THE BLOOD VESSEL, IT COULD NOT FIRE. THE STATUS OF THE PATIENT WAS FINE.
30-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, WHEN ON LOBECTOMY ON THE LUNG TUMOR, THE STAPLER AND LOAD DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. AS STATED, THEY REPLACED A NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE STAPLING, THE STAPLER, WITH ITS RELOAD, WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN TRANSECTING THE STOMACH TISSUE AND CREATING A GASTRIC TUBE, WHEN THE SURGEON FIRED THE RELOAD, THERE WAS A RESOUNDING CRACK SOUND AND THE RELOAD WAS UNABLE TO FIRE SINCE. THERE WAS A POOR STAPLE FORMATION WHEREIN SOME OF THE STAPLES WERE C-SHAPED AND DID NOT FORM WELL. USED ANOTHER STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION WITH NORMAL THICKNESS TISSUE, THE STAPLER WAS NOT ABLE TO FIRE AND HAS PRODUCED AN INCOMPLETE STAPLE LINE. THE LITTLE BLACK PARTS OF THE RELOAD WERE THOUGHT TO HAVE BROKEN OFF AND GONE INTO THE PATIENT¿S CAVITY, BUT THEY ARE UNSURE WHERE IT HAS COME FROM. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE PERFORMING LUNG RESECTION IN BREAKING AWAY FROM THE LUNG TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE COULD NOT BE SQUEEZE. ON THE FIRST FIRING THE STAPLES HAVE BEEN FIRED BUT DID NOT TOUCH THE LUNG TISSUE. TO COMPLETE THE OPERATION, THEY OPENED ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, THE SURGEON APPROACHED A BLOOD VESSEL WITH THE DEVICE ONCE, BUT REMOVED IT WITHOUT FIRING DUE TO INSUFFICIENT DISSECTION OF THE VESSEL. THE DEVICE WAS LEFT IN THE CLEAN FIELD FOR LATER FIRING. AFTER THAT, THE RELOAD WAS LOADED ONTO THE HANDLE WITHOUT PROBLEM, BUT WHEN SURGEON TRIED TO FIRE THE DEVICE TO THE BLOOD VESSEL, IT COULD NOT FIRE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE STAPLING, THE STAPLER, WITH ITS RELOAD, WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER STAPLER AND RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-07-2017 | Malfunction | ethicon endo-surgery | 17-04-2018 | ECHELON FLEX 60 LONG | Material Deformation | A0406 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN THE SLEEVE GASTRECTOMY PROCEDURE, THE STAPLER WAS WANTED TO BE USED FOR 3RD FIRING AND A NEW CARTRIDGE WAS WANTED TO BE INSERTED. HOWEVER INSERTION COULDN'T BE DONE. IT WAS NOTED THAT THE CARTRIDGE INSERTION AREA WAS DEFORMED.
19-12-2017 | Malfunction | ethicon endo-surgery | 17-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Device Inoperable |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE DIDN'T WORK. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THERE IS NO OTHER INFORMATION KNOWN OR AVAILABLE.
19-03-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | ECHELON FLEX 60 LONG | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: THE PHYSICIAN REPORTED THAT PRODUCT, AFTER PUT ON FROM ORIGINAL PACKAGE WAS MARKED AS CLOSED (PADLOCK). PHYSICIAN TRIED TO UNLOCKED THE PRODUCT WITHOUT ANY RESULT. TO COMPLETE PROCEDURE, PHYSICIAN USED ANOTHER ONE OF THE PRODUCT OF THE SAME TYPE. NO CONSEQUENCES FOR THE PATIENT REPORTED.
20-03-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Break | A0401 | No Consequences Or Impact To Patient | Event Description: A VERTICAL GASTRECTOMY WAS PERFORMED BY VIDEOLAPAROSCOPY, WHEN ASSEMBLING THE STAPLER OF THE KIT CODE BRKPW05GSTV2 THE GREEN CHARGE CODE GST60G, WAS DAMAGED NOT ALLOWING THE FIT IN THE STAPLER.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | SIGNIA | Failure to Sense | A070908 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) RIGHT WEDGE RESECTION, WHEN RESECTING THE LUNG TISSUE FOR A BIOPSY, THE USER ASKED FOR A PURPLE CARTRIDGE TO BE USED WITH THE DEVICE. THE CARTRIDGE WAS LOADED PROPERLY, CYCLED ON THE DEVICE AND PASSED ONTO THE USER. WHEN THE USER PUT IT ONTO THE PART OF LUNG THAT NEEDED RESECTING AND ARTICULATED THE CARTRIDGE AND CLOSED THE JAWS ON TISSUE BUT THE GREEN LIGHT INDICATOR DID NOT COME ON. THE USER OPENED AND CLOSED THE JAWS TO COMPRESS THE TISSUE SO THE FLUIDS ARE DISPERSED BUT THE GREEN LIGHT STILL DID NOT COME ON WHICH WAS PARTLY DOWN TO THE TISSUE BEING TOO THICK. WHEN THE SCRUB NURSE TRIED TO REMOVE THE PURPLE CARTRIDGE SO THEY CAN REPLACE IT WITH THE BLACK CARTRIDGE, THE USER WAS NOT ABLE TO STRAIGHTEN THE CARTRIDGE. THE NURSE TOOK THE CARTRIDGE OFF SOME HOW AND THE USER FIRED TWO BLACK CARTRIDGES WHICH WORKED ABSOLUTELY FINE. WHEN THE USER ASKED FOR THE SAME PURPLE CARTRIDGE, IT DID NOT REGISTER WITH THE DEVICE. AS STATED, THERE WAS A PROBLEM WITH UNLOADING THE DEVICE, THE STAPLER DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS SQUEEZED. THE RELOAD WAS REPLACED BY ANOTHER CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | SIGNIA | Calibration Problem | A08 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) RIGHT WEDGE RESECTION, WHEN RESECTING THE LUNG TISSUE FOR A BIOPSY, THE USER ASKED FOR A PURPLE CARTRIDGE TO BE USED WITH THE DEVICE. THE CARTRIDGE WAS LOADED PROPERLY, CYCLED ON THE DEVICE AND PASSED ONTO THE USER. WHEN THE USER PUT IT ONTO THE PART OF LUNG THAT NEEDED RESECTING AND ARTICULATED THE CARTRIDGE AND CLOSED THE JAWS ON TISSUE BUT THE GREEN LIGHT INDICATOR DID NOT COME ON. THE USER OPENED AND CLOSED THE JAWS TO COMPRESS THE TISSUE SO THE FLUIDS ARE DISPERSED BUT THE GREEN LIGHT STILL DID NOT COME ON WHICH WAS PARTLY DOWN TO THE TISSUE BEING TOO THICK. WHEN THE SCRUB NURSE TRIED TO REMOVE THE PURPLE CARTRIDGE SO THEY CAN REPLACE IT WITH THE BLACK CARTRIDGE, THE USER WAS NOT ABLE TO STRAIGHTEN THE CARTRIDGE. THE NURSE TOOK THE CARTRIDGE OFF SOME HOW AND THE USER FIRED TWO BLACK CARTRIDGES WHICH WORKED ABSOLUTELY FINE. WHEN THE USER ASKED FOR THE SAME PURPLE CARTRIDGE, IT DID NOT REGISTER WITH THE DEVICE. AS STATED, THERE WAS A PROBLEM WITH UNLOADING THE DEVICE, THE STAPLER DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS SQUEEZED. THE RELOAD WAS REPLACED BY ANOTHER CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | SIGNIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) RIGHT WEDGE RESECTION, WHEN RESECTING THE LUNG TISSUE FOR A BIOPSY, THE USER ASKED FOR A PURPLE CARTRIDGE TO BE USED WITH THE DEVICE. THE CARTRIDGE WAS LOADED PROPERLY, CYCLED ON THE DEVICE AND PASSED ONTO THE USER. WHEN THE USER PUT IT ONTO THE PART OF LUNG THAT NEEDED RESECTING AND ARTICULATED THE CARTRIDGE AND CLOSED THE JAWS ON TISSUE BUT THE GREEN LIGHT INDICATOR DID NOT COME ON. THE USER OPENED AND CLOSED THE JAWS TO COMPRESS THE TISSUE SO THE FLUIDS ARE DISPERSED BUT THE GREEN LIGHT STILL DID NOT COME ON WHICH WAS PARTLY DOWN TO THE TISSUE BEING TOO THICK. WHEN THE SCRUB NURSE TRIED TO REMOVE THE PURPLE CARTRIDGE SO THEY CAN REPLACE IT WITH THE BLACK CARTRIDGE, THE USER WAS NOT ABLE TO STRAIGHTEN THE CARTRIDGE. THE NURSE TOOK THE CARTRIDGE OFF SOME HOW AND THE USER FIRED TWO BLACK CARTRIDGES WHICH WORKED ABSOLUTELY FINE. WHEN THE USER ASKED FOR THE SAME PURPLE CARTRIDGE, IT DID NOT REGISTER WITH THE DEVICE. AS STATED, THERE WAS A PROBLEM WITH UNLOADING THE DEVICE, THE STAPLER DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS SQUEEZED. THE RELOAD WAS REPLACED BY ANOTHER CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien north haven-mfg | 17-04-2018 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE POSITIONING THE STAPLER FOR LUNG LOBE ACCESS IN A LUNG LOBECTOMY PROCEDURE, THE STAPLER BROKE AT THE RELOAD AND SHAFT CONNECTION POINT, AND IT REMAINED ATTACHED. THIS HAPPENED WHILE THE SURGEON WAS MANIPULATING THE STAPLER THROUGH THE PATIENT¿S THORACIC CAVITY, WITH NO TROCAR, AND HIS OTHER HAND WAS ATTEMPTING TO BEND THE SHAFT UPWARD TO ACCESS A PORTION OF THE LUNG THAT WAS TOWARD THE APEX OF THE THORACIC WALL. THERE WAS NO PATIENT INJURY.
02-03-2018 | Malfunction | covidien north haven-mfg | 17-04-2018 | ABSORBATACK | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA REPAIR THE INSTRUMENT SNAPPED IN HALF UPON USING THE DEVICE. THERE WAS NO REPORTED PATIENT INVOLVEMENT AND OUTCOME.
22-03-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, WHILE PERFORMING A GASTRIC JEJUNECTOMY WITH A SIDE-TO-SIDE ANASTOMOSIS, SOME OF THE STAPLES DIDN'T FORM AND SOME STAPLES WERE MALFORMED, AFTER FIRING THE RELOAD. A SECOND LIKE RELOAD WAS USED TO CORRECT THE ISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
30-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE PERFORMING LUNG RESECTION IN BREAKING AWAY FROM THE LUNG TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE COULD NOT BE SQUEEZE. ON THE FIRST FIRING THE STAPLES HAVE BEEN FIRED BUT DID NOT TOUCH THE LUNG TISSUE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | us surgical puerto rico | 17-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE, ON THE FIRST FIRING TO BILE DUCT DURING A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE, THE SURGEON TRIED TO FIRE THE DEVICE, BUT COULD NOT SQUEEZE THE HANDLE. THE SALES REP CONFIRMED THAT STAPLES AT THE ROOT OF THE JAWS WERE COMING UP. THERE WAS NO PATIENT INJURY.
20-02-2018 | Malfunction | ethicon endo-surgery | 17-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE, THE DOCTOR WAS USING THE INSTRUMENT WHEN RESECTING A COLON. THE HANDLE JUST BROKE AND HE WAS UNABLE TO PERFORM THE RESECTION. HE TOOK ANOTHER LIKE DEVICE AND WAS ABLE TO PERFORM THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-06-2017 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING A LAPAROSCOPIC PROCEDURE. THE STAPLING DEVICE WAS BEING USED DURING THE SURGICAL CLIPPING OF THE TISSUE. THE MANUAL STAPLING HANDLE STOPPED DURING SURGERY. THE TRIGGER WAS MOLDED AND IT WAS NOT POSSIBLE TO MAKE THE SHOOTING OF THE OTHER LOADS. THE STAPLER WAS ABLE TO FIRE. THERE WAS NO PATIENT HARM. THE PATIENT HAS NO INJURY.
18-01-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | EEA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS PROCEDURE THE DEVICE DID NOT FIRE AND COMPLETELY MALFUNCTIONED. THEY HAVE TO REDO THE POUCH TO RESOLVE THE CASE. THERE WAS A 30 MINUTES DELAY. THERE WAS NO PATIENT INJURY.
07-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WITH A HANDLE AND RELOAD, THE STAPLER WAS ABLE TO FIRE, BUT THE CARTRIDGE COULD NOT FULLY CLOSE, IT STAYED OPEN ABOUT 3-5 MM WHEN TRIED TO INSERT INTO THE TROCAR PORT. ANOTHER LOAD WAS USED BUT SAME ISSUE OCCURRED. IN ORDER TO RESOLVE THE ISSUE, A NEW TRI-STAPLE CARTRIDGE WAS USED WITH THE SAME HANDLE. NO PATIENT INJURY.
29-03-2018 | Malfunction | covidien lp llc north haven | 17-04-2018 | EGIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) RIGHT WEDGE RESECTION, WHEN RESECTING THE LUNG TISSUE FOR A BIOPSY, THE USER ASKED FOR A PURPLE CARTRIDGE TO BE USED WITH THE DEVICE. THE CARTRIDGE WAS LOADED PROPERLY, CYCLED ON THE DEVICE AND PASSED ONTO THE USER. WHEN THE USER PUT IT ONTO THE PART OF LUNG THAT NEEDED RESECTING AND ARTICULATED THE CARTRIDGE AND CLOSED THE JAWS ON TISSUE BUT THE GREEN LIGHT INDICATOR DID NOT COME ON. THE USER OPENED AND CLOSED THE JAWS TO COMPRESS THE TISSUE SO THE FLUIDS ARE DISPERSED BUT THE GREEN LIGHT STILL DID NOT COME ON WHICH WAS PARTLY DOWN TO THE TISSUE BEING TOO THICK. WHEN THE SCRUB NURSE TRIED TO REMOVE THE PURPLE CARTRIDGE SO THEY CAN REPLACE IT WITH THE BLACK CARTRIDGE, THE USER WAS NOT ABLE TO STRAIGHTEN THE CARTRIDGE. THE NURSE TOOK THE CARTRIDGE OFF SOME HOW AND THE USER FIRED TWO BLACK CARTRIDGES WHICH WORKED ABSOLUTELY FINE. WHEN THE USER ASKED FOR THE SAME PURPLE CARTRIDGE, IT DID NOT REGISTER WITH THE DEVICE. AS STATED, THERE WAS A PROBLEM WITH UNLOADING THE DEVICE, THE STAPLER DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS SQUEEZED. THE RELOAD WAS REPLACED BY ANOTHER CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | ENDO ILS 29MM, CURVED | Device Displays Incorrect Message | A090201 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN ANTERIOR RESECTION PROCEDURE USING THE DEVICE, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE DID NOT PROCEED AND STAPLES ALSO WEREN'T FIRED. AS AN EMERGENCY MEASURE, THE SURGEON DETACHED IT FROM THE TISSUE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES FOR THE PATIENT.
23-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | PROXIMATE*PPH PROCEDURE SET | Activation, Positioning or Separation Problem | A15 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A STAPLER HEMORRHOIDECTOMY, AFTER THE DEVICE WAS FIRED THE SURGEON FOUND THAT NOT ALL THE STAPLES WERE DEPLOYED. HENCE HAD TO STITCH UP SUBSEQUENTLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
10-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THERE WAS AN INCOMPLETE CUT LINE. DURING THE RESECTION OF THE LESION OF THE ESOPHAGUS, COULD NOT CUT THE TISSUE ONLY PROXIMAL TISSUE WITH CUT LINE, AND STAPLES MALFORMED ON THE 2ND FIRING, WHEN SEVERING THE STOMACH. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
10-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | FLEX 60 ARTICULATING | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THERE WAS AN INCOMPLETE CUT LINE. DURING THE RESECTION OF THE LESION OF THE ESOPHAGUS, COULD NOT CUT THE TISSUE ONLY PROXIMAL TISSUE WITH CUT LINE, AND STAPLES MALFORMED ON THE 2ND FIRING, WHEN SEVERING THE STOMACH. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE WAS NO ABLE TO FIRE. THE GREEN BUTTON WAS PRESSED HOWEVER THE HANDLE COULD NO BE SQUEEZED.
07-03-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST FIRING OF THE DEVICE IN LAPAROSCOPIC NEPHRECTOMY PROCEDURE, THEY CONNECTED RELOAD TO THE HANDLE. THE SURGEON CLAMPED THE RENAL VEIN AND PUSHED THE GREEN BUTTON AND TRIED TO SQUEEZE THE HANDLE ,HOWEVER COULD NOT SQUEEZE IT. THEY REPLACED IT WITH A NEW RELOAD, AND THE SURGEON CLAMPED THE RENAL VEIN, HOWEVER THE SAME EVENT OCCURRED. THE SALES REP NOTED SOME STAPLES OF THE PROXIMAL END WERE PRE-FIRED. THERE WAS NO PATIENT INJURY.
21-07-2017 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY IN A LAPAROSCOPIC NEPHRECTOMY, THE DEVICE COULD NOT BE LOADED. THEY USED A NEW ADDITIONAL DEVICE FOR FIRING. NO PATIENT INVOLVED.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC THORACOSCOPY SEGMENTAL LOBECTOMY, THE HANDLE DID NOT ADVANCE THE RELOAD. THEY WERE ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE RELOAD TO A LARGER SIZE BUT AFTER OPERATING TWICE, IT DOES NOT ADVANCE. THE DEVICE WAS REMOVED FROM THE CAVITY AND THEY TRIED TO DISMANTLE IT BUT NOTED TO BE IMPOSSIBLE. REPORT ALSO INDICATED THAT JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY FORCE. A COMPONENT OF THE DEVICE BREAKS OFF WHERE THE RELOAD WAS ARTICULATED. THE PROCEDURE WAS COMPLETED BY A COMPETITOR¿S DEVICE.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC THORACOSCOPY SEGMENTAL LOBECTOMY, THE HANDLE DID NOT ADVANCE THE RELOAD. THEY WERE ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE RELOAD TO A LARGER SIZE BUT AFTER OPERATING TWICE, IT DOES NOT ADVANCE. THE DEVICE WAS REMOVED FROM THE CAVITY AND THEY TRIED TO DISMANTLE IT BUT NOTED TO BE IMPOSSIBLE. REPORT ALSO INDICATED THAT JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY FORCE. A COMPONENT OF THE DEVICE BREAKS OFF WHERE THE RELOAD WAS ARTICULATED. THE PROCEDURE WAS COMPLETED BY A COMPETITOR¿S DEVICE.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON FIRED THE DEVICE TO INTER LOBAR TISSUE WITH NO PROBLEM. AT THE THIRD FIRING TO BRONCHUS, THE HANDLE BECAME STIFF AND A CRACK SOUND WAS HEARD WHEN THE SURGEON SQUEEZED THE HANDLE. THEY SUSPECTED THE TISSUE WAS HARD AT FIRST ,BUT THE CARTRIDGE DID NOT SEEM TO HAVE BEEN PRE-FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON FIRED THE DEVICE TO INTER LOBAR TISSUE WITH NO PROBLEM. AT THE THIRD FIRING TO BRONCHUS, THE HANDLE BECAME STIFF AND A CRACK SOUND WAS HEARD WHEN THE SURGEON SQUEEZED THE HANDLE. THEY SUSPECTED THE TISSUE WAS HARD AT FIRST ,BUT THE CARTRIDGE DID NOT SEEM TO HAVE BEEN PRE-FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE STAPLING THE TISSUE, THE STAPLER WOULD NOT FIRE ONCE THE GREEN BUTTON/SAFETY WAS PRESSED AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Injury | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure To Adhere Or Bond |  | Blood Loss | Event Description: ACCORDING TO THE REPORTER, AFTER A LOBE RESECTION PROCEDURE, THE PATIENT SHOWED SIGNS OF BLEEDING. THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM AND IT WAS SEEN THAT THE STAPLE LINE HAD BROKEN APART. MORE THAN 500 CC OF BLOOD LOSS WAS REPORTED HOWEVER, BLOOD TRANSFUSION WAS NOT REQUIRED. THE STAPLE LINE WAS OVERSEWED TO RESOLVE THE ISSUE.
02-04-2018 | Injury | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure To Adhere Or Bond |  | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER A LOBE RESECTION PROCEDURE, THE PATIENT SHOWED SIGNS OF BLEEDING. THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM AND IT WAS SEEN THAT THE STAPLE LINE HAD BROKEN APART. MORE THAN 500 CC OF BLOOD LOSS WAS REPORTED HOWEVER, BLOOD TRANSFUSION WAS NOT REQUIRED. THE STAPLE LINE WAS OVERSEWED TO RESOLVE THE ISSUE.
01-12-2014 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER WOULD NOT FIRE CORRECTLY AND THE HANDLE WOULD NOT CLOSE THE CARTRIDGE OR STAPLE. THEY OBTAINED ANOTHER DEVICE TO COMPLETE THE CASE.
01-12-2014 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER WOULD NOT FIRE CORRECTLY AND THE HANDLE WOULD NOT CLOSE THE CARTRIDGE OR STAPLE. THEY OBTAINED ANOTHER DEVICE TO COMPLETE THE CASE.
29-03-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THEY USED THE DEVICE AT THE FIRST FIRING OF ARTERY RESECTION IN SEGMENTAL RESECTION OF LUNG AND THORACOSCOPY. THE SURGEON CLAMPED THE TISSUE ,PUSHED THE GREEN BUTTON AND TRIED TO SQUEEZE THE HANDLE ,HOWEVER IT WAS STIFF AND COULD NOT SQUEEZE THEN COULD NOT FIRE THE DEVICE. IT WAS ALSO NOTED THAT THERE WERE NO STAPLES DEPLOYED. REPLACED BY A NEW CARTRIDGE AND THE DEVICE BECAME USABLE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | MULTIFIRE ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, THE DEVICE FAILED BEFORE USING TO THE PATIENT. THERE WAS A PROBLEM LOADING THE DEVICE, THEY COULD NOT OPEN THE JAWS WITH THE LEVER OPENING. THEY OPENED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. CHANGED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE STAPLER DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE COULD NOT BE SQUEEZE. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT WE WAS ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC THORACOSCOPY SEGMENTAL LOBECTOMY, THE HANDLE DID NOT ADVANCE THE RELOAD. THEY WERE ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE RELOAD TO A LARGER SIZE BUT AFTER OPERATING TWICE, IT DOES NOT ADVANCE. THE DEVICE WAS REMOVED FROM THE CAVITY AND THEY TRIED TO DISMANTLE IT BUT NOTED TO BE IMPOSSIBLE. REPORT ALSO INDICATED THAT JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY FORCE. A COMPONENT OF THE DEVICE BREAKS OFF WHERE THE RELOAD WAS ARTICULATED. THE PROCEDURE WAS COMPLETED BY A COMPETITOR¿S DEVICE.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC THORACOSCOPY SEGMENTAL LOBECTOMY, THE HANDLE DID NOT ADVANCE THE RELOAD. THEY WERE ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE RELOAD TO A LARGER SIZE BUT AFTER OPERATING TWICE, IT DOES NOT ADVANCE. THE DEVICE WAS REMOVED FROM THE CAVITY AND THEY TRIED TO DISMANTLE IT BUT NOTED TO BE IMPOSSIBLE. REPORT ALSO INDICATED THAT JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY FORCE. A COMPONENT OF THE DEVICE BREAKS OFF WHERE THE RELOAD WAS ARTICULATED. THE PROCEDURE WAS COMPLETED BY A COMPETITOR¿S DEVICE.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC THORACOSCOPY SEGMENTAL LOBECTOMY, THE HANDLE DID NOT ADVANCE THE RELOAD. THEY WERE ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE RELOAD TO A LARGER SIZE BUT AFTER OPERATING TWICE, IT DOES NOT ADVANCE. THE DEVICE WAS REMOVED FROM THE CAVITY AND THEY TRIED TO DISMANTLE IT BUT NOTED TO BE IMPOSSIBLE. REPORT ALSO INDICATED THAT JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY FORCE. A COMPONENT OF THE DEVICE BREAKS OFF WHERE THE RELOAD WAS ARTICULATED. THE PROCEDURE WAS COMPLETED BY A COMPETITOR¿S DEVICE.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC THORACOSCOPY SEGMENTAL LOBECTOMY, THE HANDLE DID NOT ADVANCE THE RELOAD. THEY WERE ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE RELOAD TO A LARGER SIZE BUT AFTER OPERATING TWICE, IT DOES NOT ADVANCE. THE DEVICE WAS REMOVED FROM THE CAVITY AND THEY TRIED TO DISMANTLE IT BUT NOTED TO BE IMPOSSIBLE. REPORT ALSO INDICATED THAT JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY FORCE. A COMPONENT OF THE DEVICE BREAKS OFF WHERE THE RELOAD WAS ARTICULATED. THE PROCEDURE WAS COMPLETED BY A COMPETITOR¿S DEVICE.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE STAPLING OF THE ZENKER'S DIVERTICULUM, THE SURGEON WAS ABLE PRESS THE GREEN BUTTON, BUT HE WAS NOT ABLE TO SQUEEZE THE HANDLE AND DEVICE DID NOT FIRE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO REPORTED PATIENT OUTCOME.
30-03-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE CARTRIDGE DID NOT SEND STAPLES. A NEW CARTRIDGE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THE GREEN BUTTON WAS PRESSED HOWEVER THE HANDLE COULD NO BE SQUEEZED.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC WEDGE RESECTION, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE SECOND RELOAD HAD A LOADING PROBLEM. WHEN THEY PUT THE CARTRIDGE ON THE HANDLE, THE NECK AREA ARTICULATED. THE DEVICE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A BOWEL RESECTION PROCEDURE, THE DEVICE WAS FOUND TO BE BROKEN WHILE USING THE 2ND RELOAD DURING THE SURGERY. THERE WAS RESISTANCE WHEN FIRING THE 2ND RELOAD AND THE SURGEON COMPLETED THE FIRING WITH SOME DIFFICULTY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | 75MM SELECTABLE NEW TLC | Physical Resistance |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A BOWEL RESECTION PROCEDURE, THE DEVICE WAS FOUND TO BE BROKEN WHILE USING THE 2ND RELOAD DURING THE SURGERY. THERE WAS RESISTANCE WHEN FIRING THE 2ND RELOAD AND THE SURGEON COMPLETED THE FIRING WITH SOME DIFFICULTY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
29-03-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE DEVICE WAS OPENED FOR THE FIRST USAGE. THERE WAS NO PROBLEM USING THE BLUE AND WHITE CARTRIDGE. THE CARTRIDGE WAS PLACED TO THE DEVICE AND WAS CHECKED. THE TRIGGER MOVEMENT OCCURRED ONCE BUT IT COULD NOT CONTINUE. THE TRIGGER PART OF THE DEVICE FAILED. THE DEVICE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT INJURY.
03-08-2017 | Malfunction | ethicon endo-surgery | 18-04-2018 | ECHELON FLEX 60 LONG | Difficult to Insert | A150206 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, 4TH CARTRIDGE WAS WANTED TO BE INSERTED INTO THE PRODUCT, HOWEVER IT COULDN'T BE INSERTED. IT WAS INFORMED THAT THE CARTRIDGE AREA WAS UNEASY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
03-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | SIGNIA | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE SHAFT OF THE ADAPTER TOOL WAS CRACKED. THERE WAS NO PATIENT INVOLVEMENT.
02-04-2018 | Injury | covidien lp llc north haven | 18-04-2018 | EGIA | Failure To Adhere Or Bond |  | Blood Loss | Event Description: ACCORDING TO THE REPORTER, AFTER A LOBE RESECTION PROCEDURE, THE PATIENT SHOWED SIGNS OF BLEEDING. THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM AND IT WAS SEEN THAT THE STAPLE LINE HAD BROKEN APART. MORE THAN 500 CC OF BLOOD LOSS WAS REPORTED HOWEVER, BLOOD TRANSFUSION WAS NOT REQUIRED. THE STAPLE LINE WAS OVERSEWED TO RESOLVE THE ISSUE.
02-04-2018 | Injury | covidien lp llc north haven | 18-04-2018 | EGIA | Failure To Adhere Or Bond |  | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER A LOBE RESECTION PROCEDURE, THE PATIENT SHOWED SIGNS OF BLEEDING. THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM AND IT WAS SEEN THAT THE STAPLE LINE HAD BROKEN APART. MORE THAN 500 CC OF BLOOD LOSS WAS REPORTED HOWEVER, BLOOD TRANSFUSION WAS NOT REQUIRED. THE STAPLE LINE WAS OVERSEWED TO RESOLVE THE ISSUE.
29-03-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THEY USED THE DEVICE AT THE FIRST FIRING OF ARTERY RESECTION IN SEGMENTAL RESECTION OF LUNG AND THORACOSCOPY. THE SURGEON CLAMPED THE TISSUE, PUSHED THE GREEN BUTTON AND TRIED TO SQUEEZE THE HANDLE, HOWEVER IT WAS STIFF AND COULD NOT SQUEEZE THEN COULD NOT FIRE THE DEVICE. IT WAS ALSO NOTED THAT THERE WERE NO STAPLES DEPLOYED. REPLACED BY A NEW CARTRIDGE AND THE DEVICE BECAME USABLE. THE STATUS OF THE PATIENT WAS REPORTED AS NO PROBLEM.
29-03-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | EGIA SULU | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE DEVICE WAS OPENED FOR THE FIRST USAGE. THERE WAS NO PROBLEM USING THE BLUE AND WHITE CARTRIDGE. THE CARTRIDGE WAS PLACED TO THE DEVICE AND WAS CHECKED. THE TRIGGER MOVEMENT OCCURRED ONCE BUT IT COULD NOT CONTINUE. THE TRIGGER PART OF THE DEVICE FAILED. THE DEVICE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC THORACOSCOPY SEGMENTAL LOBECTOMY, THE HANDLE DID NOT ADVANCE THE RELOAD. THEY WERE ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE RELOAD TO A LARGER SIZE BUT AFTER OPERATING TWICE, IT DOES NOT ADVANCE. THE DEVICE WAS REMOVED FROM THE CAVITY AND THEY TRIED TO DISMANTLE IT BUT NOTED TO BE IMPOSSIBLE. REPORT ALSO INDICATED THAT JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY FORCE. A COMPONENT OF THE DEVICE BREAKS OFF WHERE THE RELOAD WAS ARTICULATED.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC THORACOSCOPY SEGMENTAL LOBECTOMY, THE HANDLE DID NOT ADVANCE THE RELOAD. THEY WERE ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE RELOAD TO A LARGER SIZE BUT AFTER OPERATING TWICE, IT DOES NOT ADVANCE. THE DEVICE WAS REMOVED FROM THE CAVITY AND THEY TRIED TO DISMANTLE IT BUT NOTED TO BE IMPOSSIBLE. REPORT ALSO INDICATED THAT JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY FORCE. A COMPONENT OF THE DEVICE BREAKS OFF WHERE THE RELOAD WAS ARTICULATED.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | MULTIFIRE ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, THE DEVICE FAILED BEFORE USING TO THE PATIENT. THERE WAS A PROBLEM LOADING THE DEVICE, THEY COULD NOT OPEN THE JAWS WITH THE LEVER OPENING. THEY OPENED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA SULU | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, THE STAPLER DID NOT CLOSE THE JAWS OF RELOADS PROPERLY. THEY REPLACED THE STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE STAPLING OF THE ZENKER'S DIVERTICULUM, THE SURGEON WAS ABLE PRESS THE GREEN BUTTON, BUT HE WAS NOT ABLE TO SQUEEZE THE HANDLE AND DEVICE DID NOT FIRE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-08-2017 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE WAS UNABLE TO FIRE. THEY USED ANOTHER RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDOGIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A COLON DISSECTION, WHILE THE EQUIPMENT WAS CLAMPED DOWN ON TISSUE, THE DEVICE STOPPED WORKING. THE DEVICE WOULD NOT OPEN, AND THEY THOUGHT THE BATTERY DIED. THEY CHANGED THE BATTERY AND RESET THE DEVICE, AND IT FIRED. THEY WERE ABLE TO OPEN THE DEVICE AND TOOK OUT THE TISSUE. THEY CONTINUED WITH THE SURGERY. HOWEVER, THE PHYSICIAN STRUGGLED RELEASING THE COLON TISSUE FROM THE DEVICE. THERE WAS NO REPORTED PATIENT OUTCOME. .
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE STAPLING THE TISSUE, THE STAPLER WOULD NOT FIRE ONCE THE GREEN BUTTON/SAFETY WAS PRESSED AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE WAS NO ABLE TO FIRE. THE GREEN BUTTON WAS PRESSED HOWEVER THE HANDLE COULD NO BE SQUEEZED.
01-12-2014 | Malfunction | covidien lp llc north haven | 18-04-2018 | EGIA HANDLE | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER WOULD NOT FIRE CORRECTLY AND THE HANDLE WOULD NOT CLOSE THE CARTRIDGE OR STAPLE.
01-12-2014 | Malfunction | covidien lp llc north haven | 18-04-2018 | EGIA HANDLE | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER WOULD NOT FIRE CORRECTLY AND THE HANDLE WOULD NOT CLOSE THE CARTRIDGE OR STAPLE.
02-04-2018 | Injury | us surgical puerto rico | 18-04-2018 | TA | Mechanics Altered | A0507 | Bowel Perforation | Event Description: ACCORDING TO THE REPORTER, DURING EXPLORATORY LAPAROSCOPIC TOTAL COLECTOMY PROCEDURE, WHILE THE DEVICE IS BEING USED IN THE COLON, NO STAPLE LINE WAS DEPLOYED FROM THE DEVICE WHEN FIRED. FECAL MATERIAL WAS RELEASED INTO ABDOMINAL CAVITY. THE PROCEDURE WAS EXTENDED FOR FORTY FIVE MINUTES TO CLEAN FECAL MATERIAL AND USE ANOTHER STAPLER TO COMPLETE THE PROCEDURE.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, THE STAPLER DID NOT CLOSE THE JAWS OF RELOADS PROPERLY. THEY REPLACED THE STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-09-2017 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING A LAPAROSCOPIC PROCEDURE. THE STAPLER DID NOT FIRE. THE TRIGGER MADE A STRONG NOISE AND IT HAD GROWN, STOPPING TO RUN. IT WAS REPORTED THAT THE TRIGGER WAS MOLDED, HANDLE WAS FLOPPY AND HAD NO MORE POSSIBILITY THAT THE DEVICE WILL FIRE. THERE WAS NO INJURY CAUSED TO THE PATIENT.
14-09-2017 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING A LAPAROSCOPIC PROCEDURE. THE STAPLER DID NOT FIRE. THE TRIGGER MADE A STRONG NOISE AND IT HAD GROWN, STOPPING TO RUN. IT WAS REPORTED THAT THE TRIGGER WAS MOLDED, HANDLE WAS FLOPPY AND HAD NO MORE POSSIBILITY THAT THE DEVICE WILL FIRE. THERE WAS NO INJURY CAUSED TO THE PATIENT.
30-03-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE CARTRIDGE DID NOT SEND STAPLES. A NEW CARTRIDGE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | TA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN DEVICE WAS BEING APPLIED ON THE COLON IN AN OPEN COLECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THE DEVICE'S HANDLE WAS ONLY SQUEEZED ONCE INSTEAD OF TWICE TO FIRE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | covidien north haven-mfg | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE FIRST FIRING OF PARTIAL RESECTION OF THE SMALL INTESTINE, THE SURGEON COULD NOT FIRE THE DEVICE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | EGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE FIRST FIRING OF PARTIAL RESECTION OF THE SMALL INTESTINE, THE SURGEON COULD NOT FIRE THE DEVICE. THEY REPLACED IT WITH A NEW CARTRIDGE, AND THE DEVICE BECAME USABLE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Injury | us surgical puerto rico | 18-04-2018 | EEA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING A LAPAROSCOPIC SIGMA RESECTION PROCEDURE. THE CIRCULAR STAPLER WAS BEING USED FOR THE DESCENDORECTOSTOMY. THE STAPLER PARTIALLY FIRED. THE STAPLER DOES NOT STAPLE CORRECTLY, THE SUTURE TEARS OPEN. THE ANVIL COULD BE TILTED AFTER FIRING AND WAS NOT BENT. THERE WAS UNANTICIPATED TISSUE LOSS AS A RESULT OF THE PROBLEM. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN ORDER TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THERE WAS A REANASTOMOSIS. THE PATIENT HAD NO INJURY.
29-03-2018 | Injury | us surgical puerto rico | 18-04-2018 | EEA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING A LAPAROSCOPIC SIGMA RESECTION PROCEDURE. THE CIRCULAR STAPLER WAS BEING USED FOR THE DESCENDORECTOSTOMY. THE STAPLER PARTIALLY FIRED. THE STAPLER DOES NOT STAPLE CORRECTLY, THE SUTURE TEARS OPEN. THE ANVIL COULD BE TILTED AFTER FIRING AND WAS NOT BENT. THERE WAS UNANTICIPATED TISSUE LOSS AS A RESULT OF THE PROBLEM. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN ORDER TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THERE WAS A REANASTOMOSIS. THE PATIENT HAD NO INJURY.
29-03-2018 | Injury | us surgical puerto rico | 18-04-2018 | EEA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING A LAPAROSCOPIC SIGMA RESECTION PROCEDURE. THE CIRCULAR STAPLER WAS BEING USED FOR THE DESCENDORECTOSTOMY. THE STAPLER PARTIALLY FIRED. THE STAPLER DOES NOT STAPLE CORRECTLY, THE SUTURE TEARS OPEN. THE ANVIL COULD BE TILTED AFTER FIRING AND WAS NOT BENT. THERE WAS UNANTICIPATED TISSUE LOSS AS A RESULT OF THE PROBLEM. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN ORDER TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THERE WAS A REANASTOMOSIS. THE PATIENT HAD NO INJURY.
03-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT WE WAS ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
08-02-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | EEA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING IN A GASTRECTOMY, THE DEVICE HAD POOR STAPLE FORMATION. THERE WAS OPENING OF THE POINTS CAUSING THEM TO REPEAT THE CLAMPING. TO FIX THE STAPLE LINE THEY HAD TO RESECT AND RE-STAPLE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-02-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING IN A GASTRECTOMY, THE DEVICE HAD POOR STAPLE FORMATION. THERE WAS OPENING OF THE POINTS CAUSING THEM TO REPEAT THE CLAMPING. TO FIX THE STAPLE LINE THEY HAD TO RESECT AND RE-STAPLE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE STAPLER DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE COULD NOT BE SQUEEZE. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC WEDGE RESECTION, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE SECOND RELOAD HAD A LOADING PROBLEM. WHEN THEY PUT THE CARTRIDGE ON THE HANDLE, THE NECK AREA ARTICULATED. THE DEVICE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PROF INFORMED US THAT THE DEVICE DIDN'T FIRE DURING A OESOPHAGECTOMY. THE MANUAL OVERRIDE DIDN'T WORK. NO PATIENT CONSEQUENCE
20-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE WOULD NOT OPEN ONCE INTRODUCED THROUGH THE TROCAR. NO FURTHER INFORMATION IS AVAILABLE. NO PATIENT CONSEQUENCES WERE REPORTED.
02-04-2018 | Malfunction | covidien north haven-mfg | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN LOBECTOMY PROCEDURE, THEY SET THE DEVICE WITHOUT PROBLEM, THEN WHEN THE SURGEON CLAMPED THE DEVICE ON THE PULMONARY PARENCHYMA IT DID NOT FIRE, HOWEVER THE SURGEON REMOVED THE DEVICE FROM THE TISSUE WITH NO PROBLEM. THE PATIENT WAS ALIVE WITH NO INJURY.
13-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | PROXIMATE*RL LINEAR STAPLER | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A BOWEL PROCEDURE THE STAPLER DIDN'T FIRE CORRECTLY. THE STAPLER WAS RELOADED AND STILL DIDN'T FIRE COMPLETELY ACROSS THE BOWEL. THERE WAS A COMPLETE STAPLE LINE ACROSS THE BOWEL AS A PATIENT CONSEQUENCE. THE TECH IN THE ROOM WASN'T SURE IF IT WAS THE STAPLER OR HUMAN ERROR THAT CAUSED THE ISSUE.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT HEMICOLECTOMY, THE STAPLE DID NOT FIRE. THE DEVICE WAS REPLACED WITH A NEW STAPLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN LOBECTOMY PROCEDURE, THEY SET THE DEVICE WITHOUT PROBLEM, THEN WHEN THE SURGEON CLAMPED THE DEVICE ON THE PULMONARY PARENCHYMA IT DID NOT FIRE, HOWEVER THE SURGEON REMOVED THE DEVICE FROM THE TISSUE WITH NO PROBLEM. THE PATIENT WAS ALIVE WITH NO INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Unintended Movement | A0512 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC WEDGE RESECTION, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE SECOND RELOAD HAD A LOADING PROBLEM. WHEN THEY PUT THE CARTRIDGE ON THE HANDLE, THE NECK AREA ARTICULATED. THE DEVICE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | IDRIVE ULTRA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A COLON DISSECTION, WHILE THE EQUIPMENT WAS CLAMPED DOWN ON TISSUE, THE DEVICE STOPPED WORKING. THE DEVICE WOULD NOT OPEN, AND THEY THOUGHT THE BATTERY DIED. THEY CHANGED THE BATTERY AND RESET THE DEVICE, AND IT FIRED. THEY WERE ABLE TO OPEN THE DEVICE AND TOOK OUT THE TISSUE. THEY CONTINUED WITH THE SURGERY. HOWEVER, THE PHYSICIAN STRUGGLED RELEASING THE COLON TISSUE FROM THE DEVICE. THERE WAS NO REPORTED PATIENT OUTCOME.
01-01-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A COLON DISSECTION, WHILE THE EQUIPMENT WAS CLAMPED DOWN ON TISSUE, THE DEVICE STOPPED WORKING. THE DEVICE WOULD NOT OPEN, AND THEY THOUGHT THE BATTERY DIED. THEY CHANGED THE BATTERY AND RESET THE DEVICE, AND IT FIRED. THEY WERE ABLE TO OPEN THE DEVICE AND TOOK OUT THE TISSUE. THEY CONTINUED WITH THE SURGERY. HOWEVER, THE PHYSICIAN STRUGGLED RELEASING THE COLON TISSUE FROM THE DEVICE. THERE WAS NO REPORTED PATIENT OUTCOME. .
05-03-2018 | Malfunction | ethicon endo-surgery | 18-04-2018 | 75 MM NTLC SELECTABLE RELOAD | Device Contamination with Chemical or Other Material | A180104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTROSTOMY, OPENED THE PACKAGE AND FOUND NAILS ON BOTH SIDES OF THE DAMAGE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
14-08-2017 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE WAS UNABLE TO FIRE. THEY USED ANOTHER RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA | Failure To Adhere Or Bond |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE ,WHEN THE DEVICE WERE BEING REMOVED, IT WAS OBSERVED THAT THE STAPLES DID NOT SEAL PROPERLY CAUSING A TEAR IN THE STOMACH TISSUE. A NEW DEVICE WAS SUPPLIED DURING THE PROCEDURE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/ WEDGE/ SEGMENTAL RESECTION, ON THE FIRST FIRING FOR SEGMENTAL RESECTION OF THE LUNG, THE SURGEON CLAMPED THE PULMONARY PARENCHYMA AND TRIED TO FIRE THE DEVICE, HOWEVER, THE HANDLE COULD NOT BE SQUEEZED AND THE KNIFE DID NOT ADVANCE. THE CARTRIDGE WAS REPLACED BY A NEW ONE AND THE FIRING WAS COMPLETED. THE SURGEON AND THE NURSE DID NOT PRE-FIRE THE DEVICE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
02-04-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE CARTRIDGE WAS EXCHANGED FOR THE SECOND FIRING, THE SURGEON FIRED THE DEVICE HALFWAY OF THE STAPLE LINE, HOWEVER, THE FIRING STOPPED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
26-01-2018 | Malfunction | us surgical puerto rico | 18-04-2018 | GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR TO AN OPEN "COLORECTOMY" PROCEDURE, THE DEVICE¿S STAPLE CAME OUT FROM THE LOAD UNIT. A NEW DEVICE WAS USED IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-03-2018 | Malfunction | covidien lp | 19-04-2018 | TA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: STAPLER WAS BEING FIRED ACROSS THE BOWEL. WHEN MD BEGAN CUTTING THERE WAS BLEEDING AND THIS IS WHEN IT WAS NOTED THAT STAPLER DID NOT FIRE.
06-12-2017 | Malfunction | covidien lp | 19-04-2018 | TA | Misfire | A050502 | No Information | Event Description: PATIENT HAVING AN EXPLORATORY LAP AND THE STAPLER MISFIRED.
20-03-2018 | Malfunction | covidien lp | 19-04-2018 | TA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: STAPLER WAS BEING FIRED ACROSS THE BOWEL. WHEN MD BEGAN CUTTING THERE WAS BLEEDING AND THIS IS WHEN IT WAS NOTED THAT STAPLER DID NOT FIRE
02-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE CARTRIDGE WAS EXCHANGED FOR THE SECOND FIRING THEN THE SURGEON FIRED THE DEVICE HALFWAY OF THE STAPLE LINE, HOWEVER, THE FIRING STOPPED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON FIRED THE DEVICE TO INTER LOBAR TISSUE WITH NO PROBLEM. AT THE THIRD FIRING TO BRONCHUS, THE HANDLE BECAME STIFF AND A CRACK SOUND WAS HEARD WHEN THE SURGEON SQUEEZED THE HANDLE. THEY SUSPECTED THE TISSUE WAS HARD AT FIRST, BUT THE CARTRIDGE DID NOT SEEM TO HAVE BEEN PRE-FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON FIRED THE DEVICE TO INTER LOBAR TISSUE WITH NO PROBLEM. AT THE THIRD FIRING TO BRONCHUS, THE HANDLE BECAME STIFF AND A CRACK SOUND WAS HEARD WHEN THE SURGEON SQUEEZED THE HANDLE. THEY SUSPECTED THE TISSUE WAS HARD AT FIRST, BUT THE CARTRIDGE DID NOT SEEM TO HAVE BEEN PRE-FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-03-2018 | Injury | us surgical puerto rico | 19-04-2018 | GIA | Fluid/Blood Leak | A050401 | Sepsis | Event Description: ACCORDING TO THE REPORTER, 10 DAYS POST CLOSURE OF ENTEROSTOMY IN LAPAROSCOPIC RIGHT COLON RESECTION, THE PATIENT EXPERIENCED SEPSIS AND BOWEL LEAK AT STAPLE LINE. THEY HAD TO PERFORM RESECTION AND ENTEROSTOMY 14 DAYS POST PROCEDURE TO RESOLVE THE CASE. THE PATIENT HAS MOVED TO LONG TERM CARE, AND HER COLOSTOMY WILL BE PERMANENT.
13-03-2018 | Injury | us surgical puerto rico | 19-04-2018 | GIA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, 10 DAYS POST CLOSURE OF ENTEROSTOMY IN LAPAROSCOPIC RIGHT COLON RESECTION, THE PATIENT EXPERIENCED SEPSIS AND BOWEL LEAK AT STAPLE LINE. THEY HAD TO PERFORM RESECTION AND ENTEROSTOMY 14 DAYS POST PROCEDURE TO RESOLVE THE CASE. THE PATIENT HAS MOVED TO LONG TERM CARE, AND HER COLOSTOMY WILL BE PERMANENT.
13-03-2018 | Injury | us surgical puerto rico | 19-04-2018 | GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 10 DAYS POST CLOSURE OF ENTEROSTOMY IN LAPAROSCOPIC RIGHT COLON RESECTION, THE PATIENT EXPERIENCED SEPSIS AND BOWEL LEAK AT STAPLE LINE. THEY HAD TO PERFORM RESECTION AND ENTEROSTOMY 14 DAYS POST PROCEDURE TO RESOLVE THE CASE. THE PATIENT HAS MOVED TO LONG TERM CARE, AND HER COLOSTOMY WILL BE PERMANENT.
06-05-2016 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO CARDIO-THORACIC SURGERY, THE DEVICE WAS UNABLE TO FIRE AND IT WAS IMPOSSIBLE TO USE THE LOADING UNIT. THEY ALSO NOTED THAT THE TWO CLAMPS DID NOT WORK THAT USED THE FIRST RELOAD. THERE WAS NO PATIENT INVOLVEMENT.
19-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
19-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE LOAD DID NOT FIRE AND THERE WAS A PROBLEM UNLOADING THE DEVICE. PATIENT STATUS WAS ASKED BUT UNKNOWN.
19-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE LOAD DID NOT FIRE AND THERE WAS A PROBLEM UNLOADING THE DEVICE. PATIENT STATUS WAS ASKED BUT UNKNOWN.
28-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | IDRIVE ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATECTOMY, AFTER THE SURGEON CLAMPED THE DEVICE ON THE TISSUE, THE FIRING STOPPED ON THE HALFWAY. THEY THEN REPLACED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY IN THIS EVENT.
26-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TRANSECTION OF JEJUNOJEJUNOSTOMY IN A THE ROUX-EN-Y SUBTOTAL GASTRECTOMY, THE DEVICE FIRED CORRECTLY WITH THE FIRST THREE RELOADS, BUT DURING THE FOURTH FIRING, THE DEVICE ONLY PARTIALLY FIRED. IT WAS JAMMED AND WOULD NOT ADVANCE. THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END. IT WAS ALSO NOTED THAT THERE WAS A MINIMAL TISSUE LOSS AS A RESULT WHEN THE SURGEON STARTED A NEW STAPLE LINE IMMEDIATELY BESIDE THE PREVIOUSLY FIRED STAPLES. A NEW DEVICE¿S HANDLE WAS USED AND THE OPERATION CONTINUED AS PLANNED. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RETOSSIGMOIDECTOMY, WHEN ON TRANSECTION THE RECTUM, THE STAPLER AND TWO RELOADS WERE ABLE TO FIRE BUT THE STAPLE LINE WAS INCOMPLETE DISTALLY. THERE WAS NO PATIENT INJURY.
28-03-2018 | Injury | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE CREATING THE GASTRIC POUCH IN A LAPAROSCOPIC BYPASS, THE KNIFE BLADE ADVANCED AND CUT THE TISSUE BUT THE STAPLES WERE NOT DEPLOYED COMPLETELY. THE FIRST 10 MM OF THE STAPLES WERE PROPERLY DEPLOYED, AND THEN THE NEXT 20 TO 30 MM THE STAPLES WERE NOT DEPLOYED. HOWEVER, THE LAST REMAINING STAPLES WERE DEPLOYED. TO COMPLETE THE CASE THEY OPENED A NEW DEVICE. THEY RE-CREATED THE GASTRIC ANASTOMOSIS AND OVER SEWED THE ENTIRE STAPLE LINE TO FIX THE STAPLE LINE. THIS EVENT CAUSED THE SURGICAL TIME TO BE EXTENDED BY 30 MINUTES OR MORE.
04-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE CONVERTED TO OPEN UNRELATED TO DEVICE PROBLEM, WHEN THE SURGEON WAS FIRING ON BRONCHUS, THE DEVICE WAS UNABLE TO FIRE BUT THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED BY THE SURGEON. THE RELOAD WAS PARTIALLY FIRED, LATER THERE WAS A RESISTANCE TO FIRE, THEN, THE SURGEON DECIDED TO UNLOCK THE DEVICE BY RETRACTING THE KNOB. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE BEING APPLIED ON THE STOMACH, THE SURGEON WAS ABLE TO FIRE ON THE FIRST TWO RELOAD BUT WITH THE THIRD, THE HANDLE GOT STUCK AND THE BLACK HANDLE WAS INEFFECTIVE TO FIRE. THE SURGEON WAS ALSO UNABLE TO SQUEEZE THE HANDLE ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON. THE SURGEON TOOK ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE CONVERTED TO OPEN UNRELATED TO DEVICE PROBLEM, WHEN THE SURGEON WAS FIRING ON BRONCHUS, THE DEVICE WAS UNABLE TO FIRE BUT THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED BY THE SURGEON. THE RELOAD WAS PARTIALLY FIRED, LATER THERE WAS A RESISTANCE TO FIRE, THEN, THE SURGEON DECIDED TO UNLOCK THE DEVICE BY RETRACTING THE KNOB. THE SURGEON USED OPEN RELOADS TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: AT THE FIRST FIRING OF SIGMOIDECTOMY, THE SURGEON FELT THE JAWS OPENED WIDER THAN USUAL. HE/SHE CLAMPED THE TISSUE, HOWEVER, AS HE/SHE GOT WORRIED ABOUT A TROUBLE MIGHT HAPPEN, HE/SHE STOPPED USING THE FAULTY DEVICE. REPLACED BY A NEW CARTRIDGE AND THE FIRING WAS COMPLETED. NO PATIENT HARM.
28-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATECTOMY, AFTER THE SURGEON CLAMPED THE DEVICE ON THE TISSUE, THE FIRING STOPPED ON THE HALFWAY. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE ON TRANSVERSE COLON DISSECTION IN A TRANSVERSE COLON ILEOSTOMY OPEN PROCEDURE, THE FIRST PART OF THE STAPLER LINE DEPLOYMENT WAS COMPLETED. HOWEVER, HALFWAY THRU THE MARK THE STAPLER WAS NOT FORMING AND IT WAS ONLY CUTTING. THEY OVER SEWED THE SUTURE TO COMPLETE THE CASE.
03-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE ON TRANSVERSE COLON DISSECTION IN A TRANSVERSE COLON ILEOSTOMY OPEN PROCEDURE, THE FIRST PART OF THE STAPLER LINE DEPLOYMENT WAS COMPLETED. HOWEVER, HALFWAY THRU THE MARK THE STAPLER WAS NOT FORMING AND IT WAS ONLY CUTTING. THEY OVER SEWED THE SUTURE TO COMPLETE THE CASE.
28-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATECTOMY, AFTER THE SURGEON CLAMPED THE DEVICE ON THE TISSUE, THE FIRING STOPPED ON THE HALFWAY. THERE WAS NO PATIENT INJURY IN THIS EVENT.
05-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC PROCEDURE (LAPAROSCOPIC PROCEDURE CONVERTED TO OPEN UNRELATED TO DEVICE PROBLEM), WHILE BEING APPLIED AT THE RECTUM, THE STAPLER WAS ABLE TO FIRE AND ITS HANDLE WAS ABLE TO RECOGNIZE ITS RELOAD BUT THE DISTAL STAPLE LINE WAS INCOMPLETE. THE SURGEON PLACED THE STAPLER LOW IN THE PELVIS AND HE SAID THAT HE FIRED THE STAPLER COMPLETELY, BUT THEN SAID THAT THE STAPLER DID NOT PLACE STAPLES ON THE DISTAL LINE. THE RELOAD LOOKED AS IF ALL THE STAPLE PUSHERS WERE PUSHED UP. ANOTHER STRAIGHT PURPLE RELOAD WAS USED TO FIX THE LINE AND COMPLETE THE CASE. THE PATIENT HAS NO INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE CREATING THE SLEEVE, THE DEVICE WAS ABLE TO FIRE COMPLETELY BUT IT TORE THE SEROSA. THE SURGEON HAD TO OVERSEW THE DEFECT TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE CREATING THE SLEEVE, THE DEVICE WAS ABLE TO FIRE COMPLETELY BUT IT TORE THE SEROSA. THE SURGEON HAD TO OVERSEW THE DEFECT TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE CREATING THE SLEEVE, THE DEVICE WAS ABLE TO FIRE COMPLETELY BUT IT TORE THE SEROSA. THE SURGEON HAD TO OVERSEW THE DEFECT TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE CREATING THE SLEEVE, THE DEVICE WAS ABLE TO FIRE COMPLETELY BUT IT TORE THE SEROSA. THE SURGEON HAD TO OVERSEW THE DEFECT TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, WHEN THE DEVICE WAS CLAMPED ON SEVENTH AND EIGHTH TRANSECTION BRONCHUS, THE CARTRIDGE ON THE SIDE OF THE STAPLER MOVE OUT FROM THE DEVICE AND WAS UNABLE TO FIRE. THEY ALSO NOTED THAT THEY WERE ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY CHANGE TO A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, WHEN THE DEVICE WAS CLAMPED ON SEVENTH AND EIGHTH TRANSECTION BRONCHUS, THE CARTRIDGE ON THE SIDE OF THE STAPLER MOVE OUT FROM THE DEVICE AND WAS UNABLE TO FIRE. THEY ALSO NOTED THAT THEY WERE ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY CHANGE TO A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | nan | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT WORK WHEN THE CARTRIDGE WAS INSTALLED AND IT WAS NOT ABLE TO FIRE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE SAME CARTRIDGE WAS TRIED WITH A DIFFERENT DEVICE AND IT WORKED. THERE WAS NO PATIENT INJURY.
03-04-2018 | Injury | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON FIRED TO PULMONARY PARENCHYMA WITH STAPLER, HOWEVER THE STAPLING AND RESECTION WERE NOT PERFORMED AT ALL AND A PART LIKE A METAL PLATE APPEARED FROM THE ARTICULATED PART OF THE CARTRIDGE. THE LUNG WAS VERY HARD. THE SURGEON ARTICULATED THE JAWS WITH MAXIMUM ANGLE AND CLAMPED THE TISSUE ,THEN THE DISPLAY SCREEN SHOWED THE DEVICE WAS READY FOR USE. THE SURGEON CONSIDERED THE THICKNESS OF THE TISSUE AND WAITED FOR A WHILE ,THEN STARTED THE FIRING. THE FIRING SPEED SEEMED TO BE CONSTANT. AFTER THE FIRING ,THE SURGEON TRIED TO RETURN THE KNIFE TO THE INITIAL POSITION , HOWEVER COULD NOT RETURN THE KNIFE AND COULD NOT CLOSE THE JAWS. SURGEON RE-TRIED TO OPEN THE JAWS OVER AGAIN, THEN FINALLY COULD OPEN THEM, HOWEVER HE/SHE NOTED THE TISSUE WAS NOT STAPLED NOR RESECTED AT ALL. HE/SHE TRIED TO REMOVE THE CARTRIDGE OUT OF PATIENT' S BODY, HOWEVER COULD NOT REMOVE IT FROM THE TROCAR POINT. SURGEON NOTED A PART LIKE A METAL PLATE APPEARED FROM THE ARTICULATED PART OF THE CARTRIDGE AND IT STUCK TO THE TROCAR AND DISTURBED REMOVING THE CARTRIDGE OUT OF PATIENT'S BODY. HE/SHE SLIGHTLY EXTENDED INCISION OF THE TROCAR POINT , THEN COULD REMOVE THE CARTRIDGE OUT OF PATIENT' S BODY. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE DEVICE WAS REMOVED FROM THE TISSUE BY ADDITIONALLY RESECTING THE TISSUE. THE INCISION SITE WAS EXTENDED BY LESS THAN 3 CM. THERE WAS NO PATIENT INJURY.
03-04-2018 | Injury | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON FIRED TO PULMONARY PARENCHYMA WITH STAPLER, HOWEVER THE STAPLING AND RESECTION WERE NOT PERFORMED AT ALL AND A PART LIKE A METAL PLATE APPEARED FROM THE ARTICULATED PART OF THE CARTRIDGE. THE LUNG WAS VERY HARD. THE SURGEON ARTICULATED THE JAWS WITH MAXIMUM ANGLE AND CLAMPED THE TISSUE ,THEN THE DISPLAY SCREEN SHOWED THE DEVICE WAS READY FOR USE. THE SURGEON CONSIDERED THE THICKNESS OF THE TISSUE AND WAITED FOR A WHILE ,THEN STARTED THE FIRING. THE FIRING SPEED SEEMED TO BE CONSTANT. AFTER THE FIRING ,THE SURGEON TRIED TO RETURN THE KNIFE TO THE INITIAL POSITION , HOWEVER COULD NOT RETURN THE KNIFE AND COULD NOT CLOSE THE JAWS. SURGEON RE-TRIED TO OPEN THE JAWS OVER AGAIN, THEN FINALLY COULD OPEN THEM, HOWEVER HE/SHE NOTED THE TISSUE WAS NOT STAPLED NOR RESECTED AT ALL. HE/SHE TRIED TO REMOVE THE CARTRIDGE OUT OF PATIENT' S BODY, HOWEVER COULD NOT REMOVE IT FROM THE TROCAR POINT. SURGEON NOTED A PART LIKE A METAL PLATE APPEARED FROM THE ARTICULATED PART OF THE CARTRIDGE AND IT STUCK TO THE TROCAR AND DISTURBED REMOVING THE CARTRIDGE OUT OF PATIENT'S BODY. HE/SHE SLIGHTLY EXTENDED INCISION OF THE TROCAR POINT , THEN COULD REMOVE THE CARTRIDGE OUT OF PATIENT' S BODY. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE DEVICE WAS REMOVED FROM THE TISSUE BY ADDITIONALLY RESECTING THE TISSUE. THE INCISION SITE WAS EXTENDED BY LESS THAN 3 CM. THERE WAS NO PATIENT INJURY.
03-04-2018 | Injury | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Difficult to Remove | A150207 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON FIRED TO PULMONARY PARENCHYMA WITH STAPLER, HOWEVER THE STAPLING AND RESECTION WERE NOT PERFORMED AT ALL AND A PART LIKE A METAL PLATE APPEARED FROM THE ARTICULATED PART OF THE CARTRIDGE. THE LUNG WAS VERY HARD. THE SURGEON ARTICULATED THE JAWS WITH MAXIMUM ANGLE AND CLAMPED THE TISSUE ,THEN THE DISPLAY SCREEN SHOWED THE DEVICE WAS READY FOR USE. THE SURGEON CONSIDERED THE THICKNESS OF THE TISSUE AND WAITED FOR A WHILE ,THEN STARTED THE FIRING. THE FIRING SPEED SEEMED TO BE CONSTANT. AFTER THE FIRING ,THE SURGEON TRIED TO RETURN THE KNIFE TO THE INITIAL POSITION , HOWEVER COULD NOT RETURN THE KNIFE AND COULD NOT CLOSE THE JAWS. SURGEON RE-TRIED TO OPEN THE JAWS OVER AGAIN, THEN FINALLY COULD OPEN THEM, HOWEVER HE/SHE NOTED THE TISSUE WAS NOT STAPLED NOR RESECTED AT ALL. HE/SHE TRIED TO REMOVE THE CARTRIDGE OUT OF PATIENT' S BODY, HOWEVER COULD NOT REMOVE IT FROM THE TROCAR POINT. SURGEON NOTED A PART LIKE A METAL PLATE APPEARED FROM THE ARTICULATED PART OF THE CARTRIDGE AND IT STUCK TO THE TROCAR AND DISTURBED REMOVING THE CARTRIDGE OUT OF PATIENT'S BODY. HE/SHE SLIGHTLY EXTENDED INCISION OF THE TROCAR POINT , THEN COULD REMOVE THE CARTRIDGE OUT OF PATIENT' S BODY. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE DEVICE WAS REMOVED FROM THE TISSUE BY ADDITIONALLY RESECTING THE TISSUE. THE INCISION SITE WAS EXTENDED BY LESS THAN 3 CM. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING IN A LAPAROSCOPIC GASTRIC SEPTATION, THE LOAD GOT CAUGHT AND THE STAPLE COULD NOT BE OPENED THE DEVICE WAS NOT FIRING AT THAT TIME. THEY USED ANOTHER LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TRANSECTION OF JEJUNOJEJUNOSTOMY IN A THE ROUX-EN-Y SUBTOTAL GASTRECTOMY, THE DEVICE FIRED CORRECTLY WITH THE FIRST THREE RELOADS, BUT DURING THE FOURTH FIRING, THE DEVICE ONLY PARTIALLY FIRED. IT WAS JAMMED AND WOULD NOT ADVANCE. THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END. IT WAS ALSO NOTED THAT THERE WAS A MINIMAL TISSUE LOSS AS A RESULT WHEN THE SURGEON STARTED A NEW STAPLE LINE IMMEDIATELY BESIDE THE PREVIOUSLY FIRED STAPLES. A NEW RELOAD WAS TRIED, BUT IT STILL WOULD NOT FIRE. A NEW DEVICE¿S HANDLE WAS USED AND THE OPERATION CONTINUED AS PLANNED. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TRANSECTION OF JEJUNOJEJUNOSTOMY IN A THE ROUX-EN-Y SUBTOTAL GASTRECTOMY, THE DEVICE FIRED CORRECTLY WITH THE FIRST THREE RELOADS, BUT DURING THE FOURTH FIRING, THE DEVICE ONLY PARTIALLY FIRED. IT WAS JAMMED AND WOULD NOT ADVANCE. THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END. IT WAS ALSO NOTED THAT THERE WAS A MINIMAL TISSUE LOSS AS A RESULT WHEN THE SURGEON STARTED A NEW STAPLE LINE IMMEDIATELY BESIDE THE PREVIOUSLY FIRED STAPLES. A NEW RELOAD WAS TRIED, BUT IT STILL WOULD NOT FIRE. A NEW DEVICE¿S HANDLE WAS USED AND THE OPERATION CONTINUED AS PLANNED. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO SMALL BOWEL, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
31-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, WHILE BEING APPLIED AT THE BOWEL ANASTOMOSIS AND AFTER THE DEVICE WAS CLAMPED OVER TISSUE, THE JAWS WERE UNABLE TO BE OPENED OR REMOVED. THEY WERE HOWEVER, ABLE TO PULL IT APART AND OFF THE TISSUE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDOGIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING IN A LAPAROSCOPIC GASTRIC SEPTATION, THE LOAD GOT CAUGHT AND THE STAPLE COULD NOT BE OPENED THE DEVICE WAS NOT FIRING AT THAT TIME. THEY USED ANOTHER LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, BEFORE THE FIRST STAPLE, THE STAPLER'S BLACK HANDLE WAS IMPOSSIBLE TO PRESS AND FIRE THEN IT BREAK OFF ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. THE SURGEON TOOK ANOTHER STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY, WHEN ON TRANSECTION THE RECTUM, THE STAPLER AND TWO RELOADS WERE ABLE TO FIRE BUT THE STAPLE LINE WAS INCOMPLETE DISTALLY. AS STATED, THEY OPENED A NEW STAPLER AND TWO MORE RELOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, BEFORE THE FIRST STAPLE, THE STAPLER'S BLACK HANDLE WAS IMPOSSIBLE TO PRESS AND FIRE THEN IT BREAK OFF ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. THE SURGEON TOOK ANOTHER STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, BEFORE THE FIRST STAPLE, THE STAPLER'S BLACK HANDLE WAS IMPOSSIBLE TO PRESS AND FIRE THEN IT BREAK OFF ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. THE SURGEON TOOK ANOTHER STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO STOMACH, THE RELOAD WAS UNABLE TO FIRE BUT THE STAPLE PREFIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. A REPLACEMENT OF RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | SIGNIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO STOMACH, THE RELOAD WAS UNABLE TO FIRE BUT THE STAPLE PREFIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. A REPLACEMENT OF RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
03-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT WORK WHEN THE CARTRIDGE WAS INSTALLED AND IT WAS NOT ABLE TO FIRE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE SAME CARTRIDGE WAS TRIED WITH A DIFFERENT DEVICE AND IT WORKED. THE DEVICE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | covidien medical products | 19-04-2018 | PREMIUM PLUS CEEA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ROBOTIC LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED AT THE ANASTOMOSIS SITE, THE SAFETY WOULD NOT RELEASE AND THE ANVIL AND ROD WERE NOT LINING UP PROPERLY, SO THE STAPLER COULD NOT BE FIRED. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT OR UNABLE TO OPEN. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | PREMIUM PLUS CEEA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ROBOTIC LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED AT THE ANASTOMOSIS SITE, THE SAFETY WOULD NOT RELEASE AND THE ANVIL AND ROD WERE NOT LINING UP PROPERLY, SO THE STAPLER COULD NOT BE FIRED. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT OR UNABLE TO OPEN. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TRANSECTION OF JEJUNOJEJUNOSTOMY IN A THE ROUX-EN-Y SUBTOTAL GASTRECTOMY, THE DEVICE FIRED CORRECTLY WITH THE FIRST THREE RELOADS, BUT DURING THE FOURTH FIRING, THE DEVICE ONLY PARTIALLY FIRED. IT WAS JAMMED AND WOULD NOT ADVANCE. THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END. IT WAS ALSO NOTED THAT THERE WAS A MINIMAL TISSUE LOSS AS A RESULT WHEN THE SURGEON STARTED A NEW STAPLE LINE IMMEDIATELY BESIDE THE PREVIOUSLY FIRED STAPLES. A NEW RELOAD WAS TRIED, BUT IT STILL WOULD NOT FIRE. A NEW DEVICE¿S HANDLE WAS USED AND THE OPERATION CONTINUED AS PLANNED. THERE WAS NO PATIENT INJURY.
05-04-2018 | Malfunction | covidien north haven-mfg | 19-04-2018 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC PROCEDURE (LAPAROSCOPIC PROCEDURE CONVERTED TO OPEN UNRELATED TO DEVICE PROBLEM), WHILE BEING APPLIED AT THE RECTUM, THE STAPLER WAS ABLE TO FIRE AND ITS HANDLE WAS ABLE TO RECOGNIZE ITS RELOAD BUT THE DISTAL STAPLE LINE WAS INCOMPLETE. THE SURGEON PLACED THE STAPLER LOW IN THE PELVIS AND HE SAID THAT HE FIRED THE STAPLER COMPLETELY, BUT THEN SAID THAT THE STAPLER DID NOT PLACE STAPLES ON THE DISTAL LINE. THE RELOAD LOOKED AS IF ALL THE STAPLE PUSHERS WERE PUSHED UP. THE PATIENT HAS NO INJURY.
28-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PARTIAL LUNG RESECTION, THE DEVICE WAS PLACED ON THE TISSUE BUT AFTER PUSHING THE GREEN BUTTON, THE SURGEON COULD NOT FIRE THE INSTRUMENT. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO STOMACH, THE RELOAD WAS UNABLE TO FIRE BUT THE STAPLE PREFIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. A REPLACEMENT OF RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | SIGNIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO STOMACH, THE RELOAD WAS UNABLE TO FIRE BUT THE STAPLE PREFIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. A REPLACEMENT OF RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-05-2016 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO CARDIO-THORACIC SURGERY, THE DEVICE WAS UNABLE TO FIRE AND IT WAS IMPOSSIBLE TO USE THE LOADING UNIT. THEY ALSO NOTED THAT THE TWO CLAMPS DID NOT WORK THAT USED THE FIRST RELOAD. THERE WAS NO PATIENT INVOLVEMENT.
19-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE LOAD DID NOT FIRE AND THERE WAS A PROBLEM UNLOADING THE DEVICE. PATIENT STATUS WAS ASKED BUT UNKNOWN.
19-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE LOAD DID NOT FIRE AND THERE WAS A PROBLEM UNLOADING THE DEVICE. PATIENT STATUS WAS ASKED BUT UNKNOWN.
29-03-2018 | Malfunction | covidien north haven-mfg | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT HARM.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, THE STAPLER MISFIRED 2 GREEN STRAIGHT RELOADS. THEY REPLACED BOTH THE STAPLER AND RELOADS TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, THE STAPLER MISFIRED 2 GREEN STRAIGHT RELOADS. THEY REPLACED BOTH THE STAPLER AND RELOADS TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-02-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, THE STAPLER WAS UNABLE TO FIRE THE SECOND SHOT WHEREIN THEY WERE UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, WHILE FIRING THE DEVICE TO THE VESSEL, THE STAPLE LINE WAS INCOMPLETE. THE SURGEON MANUALLY SUTURED THE INCOMPLETE STAPLE LINE AND SECURED THE BLEEDING. MANUAL SUTURING TECHNIQUE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, WHILE FIRING THE DEVICE TO THE VESSEL, THE STAPLE LINE WAS INCOMPLETE. THE SURGEON MANUALLY SUTURED THE INCOMPLETE STAPLE LINE AND SECURED THE BLEEDING. MANUAL SUTURING TECHNIQUE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, WHILE FIRING THE DEVICE TO THE VESSEL, THE STAPLE LINE WAS INCOMPLETE. THE SURGEON MANUALLY SUTURED THE INCOMPLETE STAPLE LINE AND SECURED THE BLEEDING. MANUAL SUTURING TECHNIQUE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE BEING APPLIED ON THE STOMACH, THE SURGEON WAS ABLE TO FIRE ON THE FIRST TWO RELOAD BUT WITH THE THIRD, THE HANDLE GOT STUCK AND THE BLACK HANDLE WAS INEFFECTIVE TO FIRE. THE SURGEON WAS ALSO UNABLE TO SQUEEZE THE HANDLE ALTHOUGH HE WAS ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
28-03-2018 | Injury | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE CREATING THE GASTRIC POUCH IN A LAPAROSCOPIC BYPASS, THE KNIFE BLADE ADVANCED AND CUT THE TISSUE BUT THE STAPLES WERE NOT DEPLOYED COMPLETELY. THE FIRST 10MM OF THE STAPLES WERE PROPERLY DEPLOYED, AND THEN THE NEXT 20 TO 30MM THE STAPLES WERE NOT DEPLOYED. HOWEVER, THE LAST REMAINING STAPLES WERE DEPLOYED. TO COMPLETE THE CASE THEY OPENED A NEW DEVICE. THEY RE-CREATED THE GASTRIC ANASTOMOSIS AND OVER SEWED THE ENTIRE STAPLE LINE TO FIX THE STAPLE LINE. THIS EVENT CAUSED THE SURGICAL TIME TO BE EXTENDED BY 30 MINUTES OR MORE.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO SMALL BOWEL, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
19-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
03-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | nan | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: AT THE FIRST FIRING OF SIGMOIDECTOMY, THE SURGEON FELT THE JAWS OPENED WIDER THAN USUAL. HE/SHE CLAMPED THE TISSUE, HOWEVER, AS HE/SHE GOT WORRIED ABOUT A TROUBLE MIGHT HAPPEN, HE/SHE STOPPED USING THE FAULTY DEVICE. REPLACED BY A NEW CARTRIDGE AND THE FIRING WAS COMPLETED. NO PATIENT HARM.
15-03-2018 | Malfunction | ethicon endo-surgery | 19-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY WHEN THE DEVICE WAS FIRED, THREE ROWS OF STAPLES ON ONE SIDE DEPLOYED AND FORMED CORRECTLY THE STAPLES ON THE OTHER SIDE DID NOT. THERE WERE NO PATIENT CONSEQUENCES REPORTED AND THERE IS NO ADDITIONAL INFORMATION AVAILABLE.
04-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDOGIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE HOLE THORACOSCOPIC LEFT UPPER LOBE RESECTION, WHILE BEING APPLIED AT THE "FRACTURES", THE STAPLER WAS ABLE TO FIRE BUT THE STAPLE FORMATION WAS POOR. THEY REPLACED THE STAPLE CARTRIDGE WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDOGIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE HOLE THORACOSCOPIC LEFT UPPER LOBE RESECTION, WHILE BEING APPLIED AT THE "FRACTURES", THE STAPLER WAS ABLE TO FIRE BUT THE STAPLE FORMATION WAS POOR. THEY REPLACED THE STAPLE CARTRIDGE WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, THE STAPLER MISFIRED 2 GREEN STRAIGHT RELOADS. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, THE STAPLER MISFIRED 2 GREEN STRAIGHT RELOADS. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | ethicon endo-surgery | 19-04-2018 | 55MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC ENTEROSTOMY AND RECOVERY, COULD NOT FIRE FARTHER AFTER FIRED 2/3. CHANGED TO ANOTHER RELOAD, FOUND IT WAS DIFFICULT TO FIRE AND WHEN FIRED 3 MM, THE FIRING KNOB WAS BROKEN OFF. ALL THE PIECES WOULD BE RETURNED. SUTURE WAS USED TO SEW THE WOUND, THE SURGERY DELAYED ABOUT 30 MINUTES. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS STANDARD FOR THIS PROCEDURE AND IS NOT EXTENDED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-03-2018 | Malfunction | ethicon endo-surgery | 19-04-2018 | 55MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC ENTEROSTOMY AND RECOVERY, COULD NOT FIRE FARTHER AFTER FIRED 2/3. CHANGED TO ANOTHER RELOAD, FOUND IT WAS DIFFICULT TO FIRE AND WHEN FIRED 3 MM, THE FIRING KNOB WAS BROKEN OFF. ALL THE PIECES WOULD BE RETURNED. SUTURE WAS USED TO SEW THE WOUND, THE SURGERY DELAYED ABOUT 30 MINUTES. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS STANDARD FOR THIS PROCEDURE AND IS NOT EXTENDED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
02-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE CREATING THE SLEEVE, THE DEVICE WAS ABLE TO FIRE COMPLETELY BUT IT TORE THE SEROSA. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE CREATING THE SLEEVE, THE DEVICE WAS ABLE TO FIRE COMPLETELY BUT IT TORE THE SEROSA. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE CREATING THE SLEEVE, THE DEVICE WAS ABLE TO FIRE COMPLETELY BUT IT TORE THE SEROSA. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DUODENAL SWITCH, WHILE CREATING THE SLEEVE, THE DEVICE WAS ABLE TO FIRE COMPLETELY BUT IT TORE THE SEROSA. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | nan | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO STOMACH, THE RELOAD WAS UNABLE TO FIRE BUT THE STAPLE PREFIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
13-03-2018 | Injury | us surgical puerto rico | 19-04-2018 | GIA | Fluid/Blood Leak | A050401 | Sepsis | Event Description: ACCORDING TO THE REPORTER, 10 DAYS POST CLOSURE OF ENTEROSTOMY IN LAPAROSCOPIC RIGHT COLON RESECTION, THE PATIENT EXPERIENCED SEPSIS AND BOWEL LEAK AT STAPLE LINE. THEY HAD TO PERFORM RESECTION AND ENTEROSTOMY 14 DAYS POST PROCEDURE TO RESOLVE THE CASE. THE PATIENT HAS MOVED TO LONG TERM CARE, AND HER COLOSTOMY WILL BE PERMANENT.
13-03-2018 | Injury | us surgical puerto rico | 19-04-2018 | GIA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, 10 DAYS POST CLOSURE OF ENTEROSTOMY IN LAPAROSCOPIC RIGHT COLON RESECTION, THE PATIENT EXPERIENCED SEPSIS AND BOWEL LEAK AT STAPLE LINE. THEY HAD TO PERFORM RESECTION AND ENTEROSTOMY 14 DAYS POST PROCEDURE TO RESOLVE THE CASE. THE PATIENT HAS MOVED TO LONG TERM CARE, AND HER COLOSTOMY WILL BE PERMANENT.
13-03-2018 | Injury | us surgical puerto rico | 19-04-2018 | GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 10 DAYS POST CLOSURE OF ENTEROSTOMY IN LAPAROSCOPIC RIGHT COLON RESECTION, THE PATIENT EXPERIENCED SEPSIS AND BOWEL LEAK AT STAPLE LINE. THEY HAD TO PERFORM RESECTION AND ENTEROSTOMY 14 DAYS POST PROCEDURE TO RESOLVE THE CASE. THE PATIENT HAS MOVED TO LONG TERM CARE, AND HER COLOSTOMY WILL BE PERMANENT.
26-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE STAPLER WAS UNABLE TO FIRE. THE DEVICE WAS CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE STAPLER WAS UNABLE TO FIRE. THE DEVICE WAS CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Injury | ethicon endo-surgery | 19-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Loss of or Failure to Bond | A040102 | Blood Loss | Event Description: IT WAS REPORTED THAT THE PATIENT RETURNED TO THE SURGICAL CENTER 1 DAY AFTER BLEEDING SURGERY (BLEEDING HEAVILY ON THE STAPLE LINE) AND THE PATIENT HAD TO REOPERATE THE PATIENT, REMOVE ALL INTERNAL BLEEDING AND PERFORM A NEW SUTURE ON THE STAPLE LINE. PATIENT IS IN THE ICU AND NEEDED TO RECEIVE 2 BAGS OF BLOOD PER TRANSFUSION.
21-03-2018 | Injury | ethicon endo-surgery | 19-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: IT WAS REPORTED THAT THE PATIENT RETURNED TO THE SURGICAL CENTER 1 DAY AFTER BLEEDING SURGERY (BLEEDING HEAVILY ON THE STAPLE LINE) AND THE PATIENT HAD TO REOPERATE THE PATIENT, REMOVE ALL INTERNAL BLEEDING AND PERFORM A NEW SUTURE ON THE STAPLE LINE. PATIENT IS IN THE ICU AND NEEDED TO RECEIVE 2 BAGS OF BLOOD PER TRANSFUSION.
03-12-2003 | Malfunction | covidien lp llc north haven | 19-04-2018 | GIA | Misfire | A050502 | No Information | Event Description: ON THE WAY OF FIRING, THE HANDLE STUCK IN THE MIDDLE.
01-12-2003 | Malfunction | us surgical puerto rico | 19-04-2018 | ROTICULATOR | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE FIRED BUT THE STAPLE COULD NOT BE FORMED PROPERLY. EXCHANGED THE LOADING UNIT AND TRIED TO FIRE BUT THE JAW COULD NOT BE OPENED. THE AREA WAS MANUALLY SUTURED, AND NO PATIENT INJURY RESULTED FROM THE ISSUE.
20-03-2018 | Malfunction | ethicon endo-surgery | 19-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Activation, Positioning or Separation Problem | A15 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN GENERAL PROCEDURE THE DEVICE FIRED, CUT, BUT DID NOT DEPLOY STAPLES ON ONE SIDE. THEY DID STAPLE ON THE OTHER. A LIKE DEVICE AND RELOAD (GST60B) WAS USED TO COMPLETE THE PROCEDURE.
10-03-2018 | Injury | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure To Adhere Or Bond |  | Unspecified Infection | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, THEY OBSERVED THAT THE CENTRAL STAPLE LINE WAS INCOMPLETE. THIS CAUSED AN ANASTOMOTIC BREACH AND IT LED TO AN ABDOMINAL INFECTION. TO RESOLVE THE CASE, THE PATIENT WAS RE-OPERATED AND THEY PERFORMED AN ILEOSTOMY TWO DAYS POST ILEAL RESECTION FOR VASCULITIS. THIS EVENT HAD CAUSED THE PATIENT TO HAVE AN EXTENDED HOSPITALIZATION.
10-03-2018 | Injury | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure To Adhere Or Bond |  | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, THEY OBSERVED THAT THE CENTRAL STAPLE LINE WAS INCOMPLETE. THIS CAUSED AN ANASTOMOTIC BREACH AND IT LED TO AN ABDOMINAL INFECTION. TO RESOLVE THE CASE, THE PATIENT WAS RE-OPERATED AND THEY PERFORMED AN ILEOSTOMY TWO DAYS POST ILEAL RESECTION FOR VASCULITIS. THIS EVENT HAD CAUSED THE PATIENT TO HAVE AN EXTENDED HOSPITALIZATION.
10-03-2018 | Injury | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Infection | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, THEY OBSERVED THAT THE CENTRAL STAPLE LINE WAS INCOMPLETE. THIS CAUSED AN ANASTOMOTIC BREACH AND IT LED TO AN ABDOMINAL INFECTION. TO RESOLVE THE CASE, THE PATIENT WAS RE-OPERATED AND THEY PERFORMED AN ILEOSTOMY TWO DAYS POST ILEAL RESECTION FOR VASCULITIS. THIS EVENT HAD CAUSED THE PATIENT TO HAVE AN EXTENDED HOSPITALIZATION.
10-03-2018 | Injury | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, THEY OBSERVED THAT THE CENTRAL STAPLE LINE WAS INCOMPLETE. THIS CAUSED AN ANASTOMOTIC BREACH AND IT LED TO AN ABDOMINAL INFECTION. TO RESOLVE THE CASE, THE PATIENT WAS RE-OPERATED AND THEY PERFORMED AN ILEOSTOMY TWO DAYS POST ILEAL RESECTION FOR VASCULITIS. THIS EVENT HAD CAUSED THE PATIENT TO HAVE AN EXTENDED HOSPITALIZATION.
03-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. CHANGED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | EEA | Failure To Adhere Or Bond |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, THE ANVIL COULD NOT BE ATTACHED TO THE STAPLER. THE SURGERY WAS EXTENDED MORE THAN 30 MINUTES DUE TO PRODUCT PROBLEM. A NEW PRODUCT WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-01-2018 | Malfunction | ethicon endo-surgery | 19-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE INSTRUMENT HAS BEEN STAPLED BUT IT DID NOT CUT. PROCEDURE HAS BEEN FINISHED MANUALLY. NO HARM TO THE PATIENT.
30-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, WHILE FIRING THE DEVICE TO THE VESSEL, THE STAPLE LINE WAS INCOMPLETE. THE SURGEON MANUALLY SUTURED THE INCOMPLETE STAPLE LINE AND SECURED THE BLEEDING. MANUAL SUTURING TECHNIQUE WAS USED TO COMPLETE THE CASE. THE PATIENT HAS NO INJURY.
30-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, WHILE FIRING THE DEVICE TO THE VESSEL, THE STAPLE LINE WAS INCOMPLETE. THE SURGEON MANUALLY SUTURED THE INCOMPLETE STAPLE LINE AND SECURED THE BLEEDING. MANUAL SUTURING TECHNIQUE WAS USED TO COMPLETE THE CASE. THE PATIENT HAS NO INJURY.
30-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, WHILE FIRING THE DEVICE TO THE VESSEL, THE STAPLE LINE WAS INCOMPLETE. THE SURGEON MANUALLY SUTURED THE INCOMPLETE STAPLE LINE AND SECURED THE BLEEDING. MANUAL SUTURING TECHNIQUE WAS USED TO COMPLETE THE CASE. THE PATIENT HAS NO INJURY.
05-04-2018 | Injury | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | Blood Loss | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING A LAPAROSCOPIC VIDEO ASSISTED RIGHT MIDDLE AND LOWER LOBECTOMY. WHILE TRANSECTING AN ARTERY, THE STAPLER DID NOT DEPLOY ANY STAPLES BUT CUT THE VESSEL. TO RESOLVE THE ISSUE, THE SURGEON CONVERTED THE CASE TO OPEN THORACOTOMY AND HAND-SUTURED THE STAPLE LINE. THE ISSUE RESULTED IN GREATER THAN 500 CC OF BLOOD LOSS, AN EXTENDED INCISION, AND AN INCREASE IN SURGICAL TIME OF MORE THAN 30 MINUTES. THE PATIENT'S HOSPITAL STAY WAS EXTENDED DUE TO THE ISSUE. TO COMPLETE THE CASE, THE HANDLE WAS USED WITH ADDITIONAL LOADING UNITS.
05-04-2018 | Injury | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING A LAPAROSCOPIC VIDEO ASSISTED RIGHT MIDDLE AND LOWER LOBECTOMY. WHILE TRANSECTING AN ARTERY, THE STAPLER DID NOT DEPLOY ANY STAPLES BUT CUT THE VESSEL. TO RESOLVE THE ISSUE, THE SURGEON CONVERTED THE CASE TO OPEN THORACOTOMY AND HAND-SUTURED THE STAPLE LINE. THE ISSUE RESULTED IN GREATER THAN 500 CC OF BLOOD LOSS, AN EXTENDED INCISION, AND AN INCREASE IN SURGICAL TIME OF MORE THAN 30 MINUTES. THE PATIENT'S HOSPITAL STAY WAS EXTENDED DUE TO THE ISSUE. TO COMPLETE THE CASE, THE HANDLE WAS USED WITH ADDITIONAL LOADING UNITS.
05-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | TRI-STAPLE | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC PROCEDURE (LAPAROSCOPIC PROCEDURE CONVERTED TO OPEN UNRELATED TO DEVICE PROBLEM), WHILE BEING APPLIED AT THE RECTUM, THE STAPLER WAS ABLE TO FIRE AND ITS HANDLE WAS ABLE TO RECOGNIZE ITS RELOAD BUT THE DISTAL STAPLE LINE WAS INCOMPLETE. THE SURGEON PLACED THE STAPLER LOW IN THE PELVIS AND HE SAID THAT HE FIRED THE STAPLER COMPLETELY, BUT THEN SAID THAT THE STAPLER DID NOT PLACE STAPLES ON THE DISTAL LINE. THE RELOAD LOOKED AS IF ALL THE STAPLE PUSHERS WERE PUSHED UP. ANOTHER STRAIGHT PURPLE RELOAD WAS USED TO FIX THE LINE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | us surgical puerto rico | 19-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT HARM.
04-04-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | EEA ORVIL | Failure To Adhere Or Bond |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, THE ANVIL COULD NOT BE ATTACHED TO THE STAPLER. THE SURGERY WAS EXTENDED MORE THAN 30 MINUTES DUE TO PRODUCT PROBLEM. A NEW PRODUCT WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | covidien lp llc north haven | 19-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PARTIAL LUNG RESECTION, THE DEVICE WAS PLACED ON THE TISSUE BUT AFTER PUSHING THE GREEN BUTTON, THE SURGEON COULD NOT FIRE THE INSTRUMENT. THEY OPENED A NEW DEVICE AND A NEW LOADING UNIT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-04-2018 | Injury | covidien lp llc north haven | 19-04-2018 | EGIA45CTAV | Failure to Form Staple | A050704 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO ASSISTED RIGHT MIDDLE AND LOWER LOBECTOMY, WHILE TRANSECTING AN ARTERY, THE STAPLER DID NOT DEPLOY ANY STAPLES BUT CUT THE VESSEL. TO RESOLVE THE ISSUE, THE SURGEON CONVERTED THE CASE TO OPEN THORACOTOMY AND HAND-SUTURED THE STAPLE LINE. THE ISSUE RESULTED IN GREATER THAN 500CC OF BLOOD LOSS, AN EXTENDED INCISION, AND AN INCREASE IN SURGICAL TIME OF MORE THAN 30 MINUTES. THE PATIENT'S HOSPITAL STAY WAS EXTENDED DUE TO THE ISSUE. TO COMPLETE THE CASE, THE HANDLE WAS USED WITH ADDITIONAL LOADING UNITS.
05-04-2018 | Injury | covidien lp llc north haven | 19-04-2018 | EGIA45CTAV | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO ASSISTED RIGHT MIDDLE AND LOWER LOBECTOMY, WHILE TRANSECTING AN ARTERY, THE STAPLER DID NOT DEPLOY ANY STAPLES BUT CUT THE VESSEL. TO RESOLVE THE ISSUE, THE SURGEON CONVERTED THE CASE TO OPEN THORACOTOMY AND HAND-SUTURED THE STAPLE LINE. THE ISSUE RESULTED IN GREATER THAN 500CC OF BLOOD LOSS, AN EXTENDED INCISION, AND AN INCREASE IN SURGICAL TIME OF MORE THAN 30 MINUTES. THE PATIENT'S HOSPITAL STAY WAS EXTENDED DUE TO THE ISSUE. TO COMPLETE THE CASE, THE HANDLE WAS USED WITH ADDITIONAL LOADING UNITS.
05-04-2018 | Injury | covidien lp llc north haven | 19-04-2018 | EGIA45CTAV | Mechanics Altered | A0507 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO ASSISTED RIGHT MIDDLE AND LOWER LOBECTOMY, WHILE TRANSECTING AN ARTERY, THE STAPLER DID NOT DEPLOY ANY STAPLES BUT CUT THE VESSEL. TO RESOLVE THE ISSUE, THE SURGEON CONVERTED THE CASE TO OPEN THORACOTOMY AND HAND-SUTURED THE STAPLE LINE. THE ISSUE RESULTED IN GREATER THAN 500CC OF BLOOD LOSS, AN EXTENDED INCISION, AND AN INCREASE IN SURGICAL TIME OF MORE THAN 30 MINUTES. THE PATIENT'S HOSPITAL STAY WAS EXTENDED DUE TO THE ISSUE. TO COMPLETE THE CASE, THE HANDLE WAS USED WITH ADDITIONAL LOADING UNITS.
05-04-2018 | Injury | covidien lp llc north haven | 19-04-2018 | EGIA45CTAV | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO ASSISTED RIGHT MIDDLE AND LOWER LOBECTOMY, WHILE TRANSECTING AN ARTERY, THE STAPLER DID NOT DEPLOY ANY STAPLES BUT CUT THE VESSEL. TO RESOLVE THE ISSUE, THE SURGEON CONVERTED THE CASE TO OPEN THORACOTOMY AND HAND-SUTURED THE STAPLE LINE. THE ISSUE RESULTED IN GREATER THAN 500CC OF BLOOD LOSS, AN EXTENDED INCISION, AND AN INCREASE IN SURGICAL TIME OF MORE THAN 30 MINUTES. THE PATIENT'S HOSPITAL STAY WAS EXTENDED DUE TO THE ISSUE. TO COMPLETE THE CASE, THE HANDLE WAS USED WITH ADDITIONAL LOADING UNITS.
16-03-2018 | Injury | incisive surgical | 19-04-2018 | INSORB SUBCUTICUULAR STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: A PATIENT EXPERIENCED A WOUND INFECTION AND WOUND SEPARATION POST FEMORAL POPLITEAL BYPASS PROCEDURE. THE PATIENT WAS TREATED IN SURGERY WITH ANTIBIOTICS AND RE-CLOSED WITH SUTURE.
16-03-2018 | Injury | incisive surgical | 19-04-2018 | INSORB SUBCUTICUULAR STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Peripheral Vascular Disease | Event Description: A PATIENT EXPERIENCED A WOUND INFECTION AND WOUND SEPARATION POST FEMORAL POPLITEAL BYPASS PROCEDURE. THE PATIENT WAS TREATED IN SURGERY WITH ANTIBIOTICS AND RE-CLOSED WITH SUTURE.
16-03-2018 | Injury | incisive surgical | 19-04-2018 | INSORB SUBCUTICUULAR STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Post Operative Wound Infection | Event Description: A PATIENT EXPERIENCED A WOUND INFECTION AND WOUND SEPARATION POST FEMORAL POPLITEAL BYPASS PROCEDURE. THE PATIENT WAS TREATED IN SURGERY WITH ANTIBIOTICS AND RE-CLOSED WITH SUTURE.
01-01-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | No Display/Image | A090206 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLER FIRED PARTIALLY. DID NOT COMPLETE TRANSACTION. NO PATIENT CONSEQUENCE REPORTED.
24-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RESECTION OF THE LESION OF THE ESOPHAGUS, WHEN SEVERING THE STOMACH, THE FIRING KNOB WAS BROKEN OFF ON 3RD FIRING. ALL OF THE PIECES WILL BE RETURNED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
23-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | ILS 33MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE STAPLE HEIGHT INDICATOR WOULD NOT MOVE FREELY. WHEN THE DEVICE WAS FIRED THE STAPLES DID NOT FORMED CORRECTLY. THE SURGEON OVER SEWED THE STAPLE LINE TO ADDRESS THIS ISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A RECTOSIGMOIDECTOMY PROCEDURE, ACCORDING TO THE MEDICAL REPORT AND THE TECHNICAL ASSISTANT WHO WAS IN THE SURGICAL ROOM AT THE MOMENT THE DOCTOR WOULD MAKE THE TISSUES JUNCTION TO BE CLIPPED, THE DEVICE CUT BUT DID NOT CLIP, AS IT CAN BE SEEN IN THE PRODUCT ITSELF, WHERE THE CLAMPS WENT UP BUT DID NOT CLOSE. REPLACEMENT OF THE PRODUCT WAS PERFORMED, BY ANOTHER OF THE SAME REFERENCE, ALLOWING CONTINUITY TO THE SURGICAL PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON CLOSED 1ST ATS45 GUN WITH WHITE RELOAD OVER VESSELS - HARD TO CLOSE. ATTEMPTED TO FIRE OVER - MADE CRUNCHING NOISE. ATTEMPTED TO OPEN HAD TO FORCE HANDLE - NOTICES DEVICE HADN'T FULLY FIRED. RELOADED 1ST DEVICE CLOSED OVER VESSEL AND SAME SITUATION. LOADED GUN #2 - SAME SITUATION - RELOADED AND FIRED OUTSIDE THE BODY - SAME SITUATION (GUN STILL CONTAINS SECOND RELOAD FIRED OUTSIDE BODY- PARTIAL FIRING). CASE COMPLETED USING COMPETITOR STAPLER. DELAY IN CASE OF 20 - 30 MINS. PATIENT IS CURRENTLY OK.
22-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON CLOSED 1ST ATS45 GUN WITH WHITE RELOAD OVER VESSELS - HARD TO CLOSE. ATTEMPTED TO FIRE OVER - MADE CRUNCHING NOISE. ATTEMPTED TO OPEN HAD TO FORCE HANDLE - NOTICES DEVICE HADN'T FULLY FIRED. RELOADED 1ST DEVICE CLOSED OVER VESSEL AND SAME SITUATION. LOADED GUN #2 - SAME SITUATION - RELOADED AND FIRED OUTSIDE THE BODY - SAME SITUATION (GUN STILL CONTAINS SECOND RELOAD FIRED OUTSIDE BODY- PARTIAL FIRING). CASE COMPLETED USING COMPETITOR STAPLER. DELAY IN CASE OF 20 - 30 MINS. PATIENT IS CURRENTLY OK.
22-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Noise, Audible | A0508 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON CLOSED 1ST ATS45 GUN WITH WHITE RELOAD OVER VESSELS - HARD TO CLOSE. ATTEMPTED TO FIRE OVER - MADE CRUNCHING NOISE. ATTEMPTED TO OPEN HAD TO FORCE HANDLE - NOTICES DEVICE HADN'T FULLY FIRED. RELOADED 1ST DEVICE CLOSED OVER VESSEL AND SAME SITUATION. LOADED GUN #2 - SAME SITUATION - RELOADED AND FIRED OUTSIDE THE BODY - SAME SITUATION (GUN STILL CONTAINS SECOND RELOAD FIRED OUTSIDE BODY- PARTIAL FIRING). CASE COMPLETED USING COMPETITOR STAPLER. DELAY IN CASE OF 20 - 30 MINS. PATIENT IS CURRENTLY OK.
27-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | PROXIMATE*RL TITAN HEAVY WIRE | Adverse Event Without Identified Device or Use Problem | A24 | Seroma | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: THE ROLE OF MUSCLE FLAP IN PREVENTING BRONCHUS STUMP INSUFFICIENCY AFTER PNEUMONECTOMY FOR MALIGNANT PLEURAL MESOTHELIOMA IN HIGH-RISK PATIENT AUTHORS: MORRIS BESHAY, GIOVANNI CARBONI, BEATRIX HOKSCH, MARC A. REYMOND, RALPH A. SCHMID CITATION: INTERACTIVE CARDIOVASCULAR AND THORACIC SURGERY. 2008; 7: 621 625. DOI: 10.1510/ICVTS.2007.166546 THE AUTHORS RETROSPECTIVELY ANALYZED ALL PATIENTS WHO UNDERWENT EXTRA PLEURAL PNEUMONECTOMY (EPP) WITH PRIMARY MUSCLE FLAP COVER OF THE BRONCHUS STUMP. A TOTAL OF 30 PATIENTS (AGE RANGE: 40 TO 83 YEARS OLD) WITH MALIGNANT PLEURAL MESOTHELIOMA (MPM) WERE INCLUDED IN THE STUDY. DURING THE PROCEDURE, ALL PATIENTS RECEIVED A MUSCLE SPARING ANTERO-LATERAL THORACOTOMY. THE HEAVY WIRE LINEAR STAPLER TLH30 WAS USED FOR DIVISION OF THE BRONCHUS. THEREAFTER, LATISSIMUS DORSI MUSCLE OR SERRATUS ANTERIOR MUSCLE WAS USED TO COVER THE BRONCHUS STUMP IN ALL PATIENTS. THE MUSCLE FLAP (MF) IS FIXED ON THE BRONCHUS STUMP IN A SANDWICH TECHNIQUE USING 3 INTERRUPTED PDS II 3-0 SUTURES FROM BEHIND AND IN FRONT. REPORTED COMPLICATIONS INCLUDED HEMORRHAGIC SECRETION (N-2) WHICH REQUIRED THORACIC DRAINAGE TO BE RETAINED AND LOCAL WOUND SEROMA (N-1) WHICH REQUIRED DRAINAGE. THREE MONTHS POST-OPERATIVE, 1 PATIENT HAD LOCAL RECURRENCE AND DIED 2 MONTHS AFTER. SIX AND 12 MONTHS POST-OPERATIVE, 8 PATIENTS HAD LOCAL RECURRENCE. IT WAS CONCLUDED THAT SURGICAL TECHNIQUES USING MUSCLE FLAP SEEMS TO PLAY A MAJOR ROLE IN THE PREVENTION OF BRONCHUS STUMP INSUFFICIENCY ESPECIALLY AFTER NEO-ADJUVANT CHEMOTHERAPY.
27-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | PROXIMATE*RL TITAN HEAVY WIRE | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: THE ROLE OF MUSCLE FLAP IN PREVENTING BRONCHUS STUMP INSUFFICIENCY AFTER PNEUMONECTOMY FOR MALIGNANT PLEURAL MESOTHELIOMA IN HIGH-RISK PATIENT AUTHORS: MORRIS BESHAY, GIOVANNI CARBONI, BEATRIX HOKSCH, MARC A. REYMOND, RALPH A. SCHMID CITATION: INTERACTIVE CARDIOVASCULAR AND THORACIC SURGERY. 2008; 7: 621 625. DOI: 10.1510/ICVTS.2007.166546 THE AUTHORS RETROSPECTIVELY ANALYZED ALL PATIENTS WHO UNDERWENT EXTRA PLEURAL PNEUMONECTOMY (EPP) WITH PRIMARY MUSCLE FLAP COVER OF THE BRONCHUS STUMP. A TOTAL OF 30 PATIENTS (AGE RANGE: 40 TO 83 YEARS OLD) WITH MALIGNANT PLEURAL MESOTHELIOMA (MPM) WERE INCLUDED IN THE STUDY. DURING THE PROCEDURE, ALL PATIENTS RECEIVED A MUSCLE SPARING ANTERO-LATERAL THORACOTOMY. THE HEAVY WIRE LINEAR STAPLER TLH30 WAS USED FOR DIVISION OF THE BRONCHUS. THEREAFTER, LATISSIMUS DORSI MUSCLE OR SERRATUS ANTERIOR MUSCLE WAS USED TO COVER THE BRONCHUS STUMP IN ALL PATIENTS. THE MUSCLE FLAP (MF) IS FIXED ON THE BRONCHUS STUMP IN A SANDWICH TECHNIQUE USING 3 INTERRUPTED PDS II 3-0 SUTURES FROM BEHIND AND IN FRONT. REPORTED COMPLICATIONS INCLUDED HEMORRHAGIC SECRETION (N-2) WHICH REQUIRED THORACIC DRAINAGE TO BE RETAINED AND LOCAL WOUND SEROMA (N-1) WHICH REQUIRED DRAINAGE. THREE MONTHS POST-OPERATIVE, 1 PATIENT HAD LOCAL RECURRENCE AND DIED 2 MONTHS AFTER. SIX AND 12 MONTHS POST-OPERATIVE, 8 PATIENTS HAD LOCAL RECURRENCE. IT WAS CONCLUDED THAT SURGICAL TECHNIQUES USING MUSCLE FLAP SEEMS TO PLAY A MAJOR ROLE IN THE PREVENTION OF BRONCHUS STUMP INSUFFICIENCY ESPECIALLY AFTER NEO-ADJUVANT CHEMOTHERAPY.
22-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON CLOSED 1ST ATS45 GUN WITH WHITE RELOAD OVER VESSELS - HARD TO CLOSE. ATTEMPTED TO FIRE OVER - MADE CRUNCHING NOISE. ATTEMPTED TO OPEN HAD TO FORCE HANDLE - NOTICES DEVICE HADN'T FULLY FIRED. RELOADED 1ST DEVICE CLOSED OVER VESSEL AND SAME SITUATION. LOADED GUN #2 - SAME SITUATION - RELOADED AND FIRED OUTSIDE THE BODY - SAME SITUATION (GUN STILL CONTAINS SECOND RELOAD FIRED OUTSIDE BODY- PARTIAL FIRING). CASE COMPLETED USING COMPETITOR STAPLER. DELAY IN CASE OF 20 - 30 MINS. PATIENT IS CURRENTLY OK.
22-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON CLOSED 1ST ATS45 GUN WITH WHITE RELOAD OVER VESSELS - HARD TO CLOSE. ATTEMPTED TO FIRE OVER - MADE CRUNCHING NOISE. ATTEMPTED TO OPEN HAD TO FORCE HANDLE - NOTICES DEVICE HADN'T FULLY FIRED. RELOADED 1ST DEVICE CLOSED OVER VESSEL AND SAME SITUATION. LOADED GUN #2 - SAME SITUATION - RELOADED AND FIRED OUTSIDE THE BODY - SAME SITUATION (GUN STILL CONTAINS SECOND RELOAD FIRED OUTSIDE BODY- PARTIAL FIRING). CASE COMPLETED USING COMPETITOR STAPLER. DELAY IN CASE OF 20 - 30 MINS. PATIENT IS CURRENTLY OK.
22-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Noise, Audible | A0508 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON CLOSED 1ST ATS45 GUN WITH WHITE RELOAD OVER VESSELS - HARD TO CLOSE. ATTEMPTED TO FIRE OVER - MADE CRUNCHING NOISE. ATTEMPTED TO OPEN HAD TO FORCE HANDLE - NOTICES DEVICE HADN'T FULLY FIRED. RELOADED 1ST DEVICE CLOSED OVER VESSEL AND SAME SITUATION. LOADED GUN #2 - SAME SITUATION - RELOADED AND FIRED OUTSIDE THE BODY - SAME SITUATION (GUN STILL CONTAINS SECOND RELOAD FIRED OUTSIDE BODY- PARTIAL FIRING). CASE COMPLETED USING COMPETITOR STAPLER. DELAY IN CASE OF 20 - 30 MINS. PATIENT IS CURRENTLY OK.
20-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE ETHICON FLEX THAT JAMMED DURING SURGERY.
26-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, THE DEVICE COULD NOT BE FIRED. THE DEVICE WAS USED ON THE RECTUM. THEN, THE SURGEON LOOSENED AND TIGHTENED THE ADJUSTING KNOB OF THE SAME DEVICE, AND THE DEVICE COULD BE FIRED PROPERLY. THE SAME DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
15-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE FIRING KNOB BROKE OFF WHILE FIRING. ACCORDING TO THE OR TEAM THERE WASN'T STAPLING OVER ANOTHER STAPLE LINE. THE STAPLE LINE WAS FINISHED BY PUSHING THE KNIFE AND STAPLES MANUALLY AFTER THE FIRING KNOB BROKE OFF. AFTER THAT THE STAPLE LINE WAS OVERSTITCHED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-02-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLER FAILED DURING USE. STAPLER WAS NOT ACTIVATING THE RELOAD. DEVICE WAS REPLACED DURING SURGERY.
28-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Appropriate Term/Code Not Available |  | Post Operative Wound Infection | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: ANALYSIS OF THE INCIDENCE AND RISK FACTORS FOR WOUND INFECTIONS IN OPEN BARIATRIC SURGERY. AUTHORS: NICOLAS V. CHRISTOU, MD, PHD; JOYCE JARAND, RN; JEAN-LOUP SYLVESTRE, RN; AND ALEXANDER P. H. MCLEAN, MD. CITATION: OBESITY SURGERY, 14, 16-22. THE PURPOSE OF THIS STUDY WAS TO DETERMINE THE INCIDENCE AND RISK FACTORS FOR WOUND INFECTION IN OPEN BARIATRIC SURGERY. BETWEEN APRIL 1, 2001 TO MARCH 31, 2003, 269 PATIENTS, 219 WOMEN AND 50 MEN BETWEEN THE AGED OF 29 TO 49 WERE INCLUDED IN THE STUDY. DURING THE PROCEDURE, A PREVIOUSLY PLACED 32-FR BOUGIE WAS HELD AGAINST THE LESSER CURVE AND A 25-MM EEA® (ETHICON ENDOSURGERY) WAS USED TO CREATE A CIRCULAR OPENING INTO THE STOMACH. A ROUXLIMB OF JEJUNUM (100-300 CM LONG DEPENDING ON BMI) WAS BROUGHT UP IN A RETROCOLIC RETROGASTRIC FASHION, AND AN END-TO-SIDE GASTROJEJUNOSTOMY WAS FASHIONED USING A 3-0 PDS® (ETHICON ENDOSURGERY) SINGLE CONTINUOUS SUTURE AROUND AN 18-FR NASOGASTRIC TUBE. WOUND INFECTIONS (N=54) WAS REPORTED. GASTROJEJUNAL LEAKS (N=4) WERE ALSO REPORTED. TWO WERE DRAINED PERCUTANEOUSLY BY COMPUTED TOMOGRAPHY GUIDANCE. NO WOUND INFECTIONS WERE OBSERVED IN THESE TWO PATIENTS. THE OTHER TWO LEAKS ALSO SHOWED SEROSANGUINOUS DRAINAGE FROM THE WOUND AND WERE TREATED BY OPEN SURGICAL DRAINAGE. WHATEVER THE ACTUAL RATE OF WOUND INFECTION AND SUBSEQUENT HERNIA, THE COMPLICATION CAUSES SIGNIFICANT MORBIDITY AND SHOULD BE AVOIDED AT ALL COSTS. AS OF FEBRUARY 2002, THE AUTHORS HAVE ADOPTED A MINIMALLY INVASIVE APPROACH TO BARIATRIC SURGERY AND PERFORM LAPAROSCOPIC ISOLATED RYGBP (HAND-SEWN GASTROJEJUNOSTOMY) ALMOST EXCLUSIVELY, IN ORDER TO AVOID WOUND INFECTIONS AND SUBSEQUENT INCISIONAL HERNIA.
28-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Appropriate Term/Code Not Available |  | No Code Available | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: ANALYSIS OF THE INCIDENCE AND RISK FACTORS FOR WOUND INFECTIONS IN OPEN BARIATRIC SURGERY. AUTHORS: NICOLAS V. CHRISTOU, MD, PHD; JOYCE JARAND, RN; JEAN-LOUP SYLVESTRE, RN; AND ALEXANDER P. H. MCLEAN, MD. CITATION: OBESITY SURGERY, 14, 16-22. THE PURPOSE OF THIS STUDY WAS TO DETERMINE THE INCIDENCE AND RISK FACTORS FOR WOUND INFECTION IN OPEN BARIATRIC SURGERY. BETWEEN APRIL 1, 2001 TO MARCH 31, 2003, 269 PATIENTS, 219 WOMEN AND 50 MEN BETWEEN THE AGED OF 29 TO 49 WERE INCLUDED IN THE STUDY. DURING THE PROCEDURE, A PREVIOUSLY PLACED 32-FR BOUGIE WAS HELD AGAINST THE LESSER CURVE AND A 25-MM EEA® (ETHICON ENDOSURGERY) WAS USED TO CREATE A CIRCULAR OPENING INTO THE STOMACH. A ROUXLIMB OF JEJUNUM (100-300 CM LONG DEPENDING ON BMI) WAS BROUGHT UP IN A RETROCOLIC RETROGASTRIC FASHION, AND AN END-TO-SIDE GASTROJEJUNOSTOMY WAS FASHIONED USING A 3-0 PDS® (ETHICON ENDOSURGERY) SINGLE CONTINUOUS SUTURE AROUND AN 18-FR NASOGASTRIC TUBE. WOUND INFECTIONS (N=54) WAS REPORTED. GASTROJEJUNAL LEAKS (N=4) WERE ALSO REPORTED. TWO WERE DRAINED PERCUTANEOUSLY BY COMPUTED TOMOGRAPHY GUIDANCE. NO WOUND INFECTIONS WERE OBSERVED IN THESE TWO PATIENTS. THE OTHER TWO LEAKS ALSO SHOWED SEROSANGUINOUS DRAINAGE FROM THE WOUND AND WERE TREATED BY OPEN SURGICAL DRAINAGE. WHATEVER THE ACTUAL RATE OF WOUND INFECTION AND SUBSEQUENT HERNIA, THE COMPLICATION CAUSES SIGNIFICANT MORBIDITY AND SHOULD BE AVOIDED AT ALL COSTS. AS OF FEBRUARY 2002, THE AUTHORS HAVE ADOPTED A MINIMALLY INVASIVE APPROACH TO BARIATRIC SURGERY AND PERFORM LAPAROSCOPIC ISOLATED RYGBP (HAND-SEWN GASTROJEJUNOSTOMY) ALMOST EXCLUSIVELY, IN ORDER TO AVOID WOUND INFECTIONS AND SUBSEQUENT INCISIONAL HERNIA.
19-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE SUBTOTAL GASTRECTOMY, DURING 2ND FIRING, THE FIRING KNOB WAS BROKEN OFF. ALL THE PIECES WOULD BE RETURNED. SUTURE WAS USED TO SEW THE WOUND. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
19-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE COULD NOT BE FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
27-03-2018 | Injury | covidien north haven-mfg | 20-04-2018 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC PULMONARY BIOPSY PROCEDURE, WHILE DOING THE WEDGE RESECTION, THE DEVICE HAD LOCKED ON THE TISSUE. THE SURROUNDING TISSUE WAS RESECTED TO REMOVE THE DEVICE WHICH EXTENDED THE SURGICAL TIME FOR MORE THAN 30 MINUTES. THE RELOAD WAS ALSO STUCK ON THE STAPLER AND COULD NOT BE UNLOADED.
27-03-2018 | Injury | covidien north haven-mfg | 20-04-2018 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC PULMONARY BIOPSY PROCEDURE, WHILE DOING THE WEDGE RESECTION, THE DEVICE HAD LOCKED ON THE TISSUE. THE SURROUNDING TISSUE WAS RESECTED TO REMOVE THE DEVICE WHICH EXTENDED THE SURGICAL TIME FOR MORE THAN 30 MINUTES. THE RELOAD WAS ALSO STUCK ON THE STAPLER AND COULD NOT BE UNLOADED.
27-03-2018 | Injury | covidien north haven-mfg | 20-04-2018 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC PULMONARY BIOPSY PROCEDURE, WHILE DOING THE WEDGE RESECTION, THE DEVICE HAD LOCKED ON THE TISSUE. THE SURROUNDING TISSUE WAS RESECTED TO REMOVE THE DEVICE WHICH EXTENDED THE SURGICAL TIME FOR MORE THAN 30 MINUTES. THE RELOAD WAS ALSO STUCK ON THE STAPLER AND COULD NOT BE UNLOADED.
27-03-2018 | Injury | covidien north haven-mfg | 20-04-2018 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC PULMONARY BIOPSY PROCEDURE, WHILE DOING THE WEDGE RESECTION, THE DEVICE HAD LOCKED ON THE TISSUE. THE SURROUNDING TISSUE WAS RESECTED TO REMOVE THE DEVICE WHICH EXTENDED THE SURGICAL TIME FOR MORE THAN 30 MINUTES. THE RELOAD WAS ALSO STUCK ON THE STAPLER AND COULD NOT BE UNLOADED.
28-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | ENDO ILS 29MM, CURVED | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | nan
05-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | EEA | Bent |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING A LAPAROSCOPY TOTAL GASTRECTOMY PROCEDURE, WHEN THEY REMOVED THE DEVICE FROM THE PACKAGE, THEY FOUND OUT THAT THE ANVIL CENTER ROD WAS BENT TO INWARD SIDE. THEY THEN REPLACED A NEW DEVICE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Fluid/Blood Leak | A050401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE PATIENT HAD INCURRED AN ANASTOMOTIC LEAK DUE TO THE STAPLE LINE.
05-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER CLAMPING THE TISSUE, THE SURGEON TRIED TO FIRE THE BLADE AND IT DID NOT ACTIVATE.
06-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHILE THE DUODENUM WAS BEING TRANSECTED, THE JAW OF THE STAPLER WAS CLOSED ON THE TISSUE AFTER A RELOAD WAS LOADED. THE JAWS WERE OPENED AND THE STAPLER WAS REMOVED. AFTER DEPRESSING THE GREEN READY TO FIRE BUTTON, THE LOOP HANDLE ADVANCED FORWARD. ONCE THE DOCTOR STARTED TO SQUEEZE THE STAPLER HANDLE, IT DID NOT FIRE THE RELOAD/ACTIVATE THE RATCHET TO DEPLOY THE STAPLES. ANOTHER STAPLER HANDLE WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD A FIRING ISSUE. PATIENT STATUS WAS UNSPECIFIED.
20-04-2018 | Injury | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE PATIENT HAD INCURRED AN ANASTOMOTIC LEAK DUE TO THE STAPLE LINE.
09-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG CANCER RADICAL SURGERY, TO MAKE A PULMONARY LOBECTOMY, STAPLES DID NOT DEPLOY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS SQUEEZED. THEY REPLACED THE STAPLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Sticking |  | Cardiac Arrest | Event Description: IT WAS REPORTED THAT DURING A CASE FOR AN HCC LIVER TUMOUR REMOVAL, THE SURGEON WAS THE CONSULTANT WHO FIRED THE DEVICE, BUT WAS ALSO WITH CONSULTANT IN THEATRE. THEY MOBILIZED THE LIVER THAT WAS SLIGHTLY STUCK TO THE DIAPHRAGM, MOBILIZATION OF THE HILUM ALSO FINE AND OFF THE CAVA SUCCESSFULLY. THERE WAS A SMALL STITCH IN THE RIGHT HEPATIC VEIN, NO CLIPS WERE USED. SINGLE SLING ROUND THE VEIN. ATS45 WITH WHITE RELOAD WAS USED AND ARTICULATED FULLY AT THE TIME. SURGEON FIRED THE DEVICE ON THE RIGHT HEPATIC VEIN AND COULD TELL WITH THE TACTILE FEEDBACK THAT SOMETHING WAS WRONG. HE LEFT THE DEVICE CLOSED AND ASKED THE TEAM TO GET READY FOR SIGNIFICANT BLEEDING FROM THE RIGHT HEPATIC VEIN, ON RELEASE OF THE DEVICE IT WAS EVIDENT THAT THE STAPLE LINE HAD FAILED AND THE STAPLES HADN¿T FORMED CORRECTLY. THERE WERE SOME DEPLOYED STAPLES ON THE DISTAL PART OF THE STAPLE/CUT LINE (SUPERIOR END OF THE VEIN) BUT NO OTHERS WERE FORMED PROPERLY. CONSULTANT TRIED TO STEM TO THE BLEED INITIALLY WITH HIS FINGER BUT THE HOLE INCREASED IN SIZE. THE STAPLER WAS USED AGAIN FOR A SECOND TIME ON PART OF THE LIVER AND DEPLOYED OK. THE RELOAD IN THE DEVICE IS THE SECOND RELOAD THAT DEPLOYED. UNFORTUNATELY, THE RELOAD FROM THE FAILED STAPLE LINE WAS DISCARDED IN THE PRESSURIZED SITUATION AT THE TIME. THE PATIENT NEEDED 12 UNITS OF BLOOD. CLAMP CONTROL FAILED. LOST CONTROL OF BLEEDING TO A LIFE-THREATENING EXSANGUINATION MEASURE. HAD TO OPEN THE CHEST AND PERFORM CARDIAC MASSAGE. EVENTUALLY MANAGED TO GET CONTROL WITH SUTURES. PATIENT IS BACK IN THEATRE TODAY TO REMOVE PART OF THE LIVER THAT NOW HAS NO BLOOD SUPPLY. CHANCES OF SURVIVAL ARE SLIM.
23-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Sticking |  | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A CASE FOR AN HCC LIVER TUMOUR REMOVAL, THE SURGEON WAS THE CONSULTANT WHO FIRED THE DEVICE, BUT WAS ALSO WITH CONSULTANT IN THEATRE. THEY MOBILIZED THE LIVER THAT WAS SLIGHTLY STUCK TO THE DIAPHRAGM, MOBILIZATION OF THE HILUM ALSO FINE AND OFF THE CAVA SUCCESSFULLY. THERE WAS A SMALL STITCH IN THE RIGHT HEPATIC VEIN, NO CLIPS WERE USED. SINGLE SLING ROUND THE VEIN. ATS45 WITH WHITE RELOAD WAS USED AND ARTICULATED FULLY AT THE TIME. SURGEON FIRED THE DEVICE ON THE RIGHT HEPATIC VEIN AND COULD TELL WITH THE TACTILE FEEDBACK THAT SOMETHING WAS WRONG. HE LEFT THE DEVICE CLOSED AND ASKED THE TEAM TO GET READY FOR SIGNIFICANT BLEEDING FROM THE RIGHT HEPATIC VEIN, ON RELEASE OF THE DEVICE IT WAS EVIDENT THAT THE STAPLE LINE HAD FAILED AND THE STAPLES HADN¿T FORMED CORRECTLY. THERE WERE SOME DEPLOYED STAPLES ON THE DISTAL PART OF THE STAPLE/CUT LINE (SUPERIOR END OF THE VEIN) BUT NO OTHERS WERE FORMED PROPERLY. CONSULTANT TRIED TO STEM TO THE BLEED INITIALLY WITH HIS FINGER BUT THE HOLE INCREASED IN SIZE. THE STAPLER WAS USED AGAIN FOR A SECOND TIME ON PART OF THE LIVER AND DEPLOYED OK. THE RELOAD IN THE DEVICE IS THE SECOND RELOAD THAT DEPLOYED. UNFORTUNATELY, THE RELOAD FROM THE FAILED STAPLE LINE WAS DISCARDED IN THE PRESSURIZED SITUATION AT THE TIME. THE PATIENT NEEDED 12 UNITS OF BLOOD. CLAMP CONTROL FAILED. LOST CONTROL OF BLEEDING TO A LIFE-THREATENING EXSANGUINATION MEASURE. HAD TO OPEN THE CHEST AND PERFORM CARDIAC MASSAGE. EVENTUALLY MANAGED TO GET CONTROL WITH SUTURES. PATIENT IS BACK IN THEATRE TODAY TO REMOVE PART OF THE LIVER THAT NOW HAS NO BLOOD SUPPLY. CHANCES OF SURVIVAL ARE SLIM.
23-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Sticking |  | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A CASE FOR AN HCC LIVER TUMOUR REMOVAL, THE SURGEON WAS THE CONSULTANT WHO FIRED THE DEVICE, BUT WAS ALSO WITH CONSULTANT IN THEATRE. THEY MOBILIZED THE LIVER THAT WAS SLIGHTLY STUCK TO THE DIAPHRAGM, MOBILIZATION OF THE HILUM ALSO FINE AND OFF THE CAVA SUCCESSFULLY. THERE WAS A SMALL STITCH IN THE RIGHT HEPATIC VEIN, NO CLIPS WERE USED. SINGLE SLING ROUND THE VEIN. ATS45 WITH WHITE RELOAD WAS USED AND ARTICULATED FULLY AT THE TIME. SURGEON FIRED THE DEVICE ON THE RIGHT HEPATIC VEIN AND COULD TELL WITH THE TACTILE FEEDBACK THAT SOMETHING WAS WRONG. HE LEFT THE DEVICE CLOSED AND ASKED THE TEAM TO GET READY FOR SIGNIFICANT BLEEDING FROM THE RIGHT HEPATIC VEIN, ON RELEASE OF THE DEVICE IT WAS EVIDENT THAT THE STAPLE LINE HAD FAILED AND THE STAPLES HADN¿T FORMED CORRECTLY. THERE WERE SOME DEPLOYED STAPLES ON THE DISTAL PART OF THE STAPLE/CUT LINE (SUPERIOR END OF THE VEIN) BUT NO OTHERS WERE FORMED PROPERLY. CONSULTANT TRIED TO STEM TO THE BLEED INITIALLY WITH HIS FINGER BUT THE HOLE INCREASED IN SIZE. THE STAPLER WAS USED AGAIN FOR A SECOND TIME ON PART OF THE LIVER AND DEPLOYED OK. THE RELOAD IN THE DEVICE IS THE SECOND RELOAD THAT DEPLOYED. UNFORTUNATELY, THE RELOAD FROM THE FAILED STAPLE LINE WAS DISCARDED IN THE PRESSURIZED SITUATION AT THE TIME. THE PATIENT NEEDED 12 UNITS OF BLOOD. CLAMP CONTROL FAILED. LOST CONTROL OF BLEEDING TO A LIFE-THREATENING EXSANGUINATION MEASURE. HAD TO OPEN THE CHEST AND PERFORM CARDIAC MASSAGE. EVENTUALLY MANAGED TO GET CONTROL WITH SUTURES. PATIENT IS BACK IN THEATRE TODAY TO REMOVE PART OF THE LIVER THAT NOW HAS NO BLOOD SUPPLY. CHANCES OF SURVIVAL ARE SLIM.
23-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Sticking |  | No Code Available | Event Description: IT WAS REPORTED THAT DURING A CASE FOR AN HCC LIVER TUMOUR REMOVAL, THE SURGEON WAS THE CONSULTANT WHO FIRED THE DEVICE, BUT WAS ALSO WITH CONSULTANT IN THEATRE. THEY MOBILIZED THE LIVER THAT WAS SLIGHTLY STUCK TO THE DIAPHRAGM, MOBILIZATION OF THE HILUM ALSO FINE AND OFF THE CAVA SUCCESSFULLY. THERE WAS A SMALL STITCH IN THE RIGHT HEPATIC VEIN, NO CLIPS WERE USED. SINGLE SLING ROUND THE VEIN. ATS45 WITH WHITE RELOAD WAS USED AND ARTICULATED FULLY AT THE TIME. SURGEON FIRED THE DEVICE ON THE RIGHT HEPATIC VEIN AND COULD TELL WITH THE TACTILE FEEDBACK THAT SOMETHING WAS WRONG. HE LEFT THE DEVICE CLOSED AND ASKED THE TEAM TO GET READY FOR SIGNIFICANT BLEEDING FROM THE RIGHT HEPATIC VEIN, ON RELEASE OF THE DEVICE IT WAS EVIDENT THAT THE STAPLE LINE HAD FAILED AND THE STAPLES HADN¿T FORMED CORRECTLY. THERE WERE SOME DEPLOYED STAPLES ON THE DISTAL PART OF THE STAPLE/CUT LINE (SUPERIOR END OF THE VEIN) BUT NO OTHERS WERE FORMED PROPERLY. CONSULTANT TRIED TO STEM TO THE BLEED INITIALLY WITH HIS FINGER BUT THE HOLE INCREASED IN SIZE. THE STAPLER WAS USED AGAIN FOR A SECOND TIME ON PART OF THE LIVER AND DEPLOYED OK. THE RELOAD IN THE DEVICE IS THE SECOND RELOAD THAT DEPLOYED. UNFORTUNATELY, THE RELOAD FROM THE FAILED STAPLE LINE WAS DISCARDED IN THE PRESSURIZED SITUATION AT THE TIME. THE PATIENT NEEDED 12 UNITS OF BLOOD. CLAMP CONTROL FAILED. LOST CONTROL OF BLEEDING TO A LIFE-THREATENING EXSANGUINATION MEASURE. HAD TO OPEN THE CHEST AND PERFORM CARDIAC MASSAGE. EVENTUALLY MANAGED TO GET CONTROL WITH SUTURES. PATIENT IS BACK IN THEATRE TODAY TO REMOVE PART OF THE LIVER THAT NOW HAS NO BLOOD SUPPLY. CHANCES OF SURVIVAL ARE SLIM.
23-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | Cardiac Arrest | Event Description: IT WAS REPORTED THAT DURING A CASE FOR AN HCC LIVER TUMOUR REMOVAL, THE SURGEON WAS THE CONSULTANT WHO FIRED THE DEVICE, BUT WAS ALSO WITH CONSULTANT IN THEATRE. THEY MOBILIZED THE LIVER THAT WAS SLIGHTLY STUCK TO THE DIAPHRAGM, MOBILIZATION OF THE HILUM ALSO FINE AND OFF THE CAVA SUCCESSFULLY. THERE WAS A SMALL STITCH IN THE RIGHT HEPATIC VEIN, NO CLIPS WERE USED. SINGLE SLING ROUND THE VEIN. ATS45 WITH WHITE RELOAD WAS USED AND ARTICULATED FULLY AT THE TIME. SURGEON FIRED THE DEVICE ON THE RIGHT HEPATIC VEIN AND COULD TELL WITH THE TACTILE FEEDBACK THAT SOMETHING WAS WRONG. HE LEFT THE DEVICE CLOSED AND ASKED THE TEAM TO GET READY FOR SIGNIFICANT BLEEDING FROM THE RIGHT HEPATIC VEIN, ON RELEASE OF THE DEVICE IT WAS EVIDENT THAT THE STAPLE LINE HAD FAILED AND THE STAPLES HADN¿T FORMED CORRECTLY. THERE WERE SOME DEPLOYED STAPLES ON THE DISTAL PART OF THE STAPLE/CUT LINE (SUPERIOR END OF THE VEIN) BUT NO OTHERS WERE FORMED PROPERLY. CONSULTANT TRIED TO STEM TO THE BLEED INITIALLY WITH HIS FINGER BUT THE HOLE INCREASED IN SIZE. THE STAPLER WAS USED AGAIN FOR A SECOND TIME ON PART OF THE LIVER AND DEPLOYED OK. THE RELOAD IN THE DEVICE IS THE SECOND RELOAD THAT DEPLOYED. UNFORTUNATELY, THE RELOAD FROM THE FAILED STAPLE LINE WAS DISCARDED IN THE PRESSURIZED SITUATION AT THE TIME. THE PATIENT NEEDED 12 UNITS OF BLOOD. CLAMP CONTROL FAILED. LOST CONTROL OF BLEEDING TO A LIFE-THREATENING EXSANGUINATION MEASURE. HAD TO OPEN THE CHEST AND PERFORM CARDIAC MASSAGE. EVENTUALLY MANAGED TO GET CONTROL WITH SUTURES. PATIENT IS BACK IN THEATRE TODAY TO REMOVE PART OF THE LIVER THAT NOW HAS NO BLOOD SUPPLY. CHANCES OF SURVIVAL ARE SLIM.
23-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A CASE FOR AN HCC LIVER TUMOUR REMOVAL, THE SURGEON WAS THE CONSULTANT WHO FIRED THE DEVICE, BUT WAS ALSO WITH CONSULTANT IN THEATRE. THEY MOBILIZED THE LIVER THAT WAS SLIGHTLY STUCK TO THE DIAPHRAGM, MOBILIZATION OF THE HILUM ALSO FINE AND OFF THE CAVA SUCCESSFULLY. THERE WAS A SMALL STITCH IN THE RIGHT HEPATIC VEIN, NO CLIPS WERE USED. SINGLE SLING ROUND THE VEIN. ATS45 WITH WHITE RELOAD WAS USED AND ARTICULATED FULLY AT THE TIME. SURGEON FIRED THE DEVICE ON THE RIGHT HEPATIC VEIN AND COULD TELL WITH THE TACTILE FEEDBACK THAT SOMETHING WAS WRONG. HE LEFT THE DEVICE CLOSED AND ASKED THE TEAM TO GET READY FOR SIGNIFICANT BLEEDING FROM THE RIGHT HEPATIC VEIN, ON RELEASE OF THE DEVICE IT WAS EVIDENT THAT THE STAPLE LINE HAD FAILED AND THE STAPLES HADN¿T FORMED CORRECTLY. THERE WERE SOME DEPLOYED STAPLES ON THE DISTAL PART OF THE STAPLE/CUT LINE (SUPERIOR END OF THE VEIN) BUT NO OTHERS WERE FORMED PROPERLY. CONSULTANT TRIED TO STEM TO THE BLEED INITIALLY WITH HIS FINGER BUT THE HOLE INCREASED IN SIZE. THE STAPLER WAS USED AGAIN FOR A SECOND TIME ON PART OF THE LIVER AND DEPLOYED OK. THE RELOAD IN THE DEVICE IS THE SECOND RELOAD THAT DEPLOYED. UNFORTUNATELY, THE RELOAD FROM THE FAILED STAPLE LINE WAS DISCARDED IN THE PRESSURIZED SITUATION AT THE TIME. THE PATIENT NEEDED 12 UNITS OF BLOOD. CLAMP CONTROL FAILED. LOST CONTROL OF BLEEDING TO A LIFE-THREATENING EXSANGUINATION MEASURE. HAD TO OPEN THE CHEST AND PERFORM CARDIAC MASSAGE. EVENTUALLY MANAGED TO GET CONTROL WITH SUTURES. PATIENT IS BACK IN THEATRE TODAY TO REMOVE PART OF THE LIVER THAT NOW HAS NO BLOOD SUPPLY. CHANCES OF SURVIVAL ARE SLIM.
23-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A CASE FOR AN HCC LIVER TUMOUR REMOVAL, THE SURGEON WAS THE CONSULTANT WHO FIRED THE DEVICE, BUT WAS ALSO WITH CONSULTANT IN THEATRE. THEY MOBILIZED THE LIVER THAT WAS SLIGHTLY STUCK TO THE DIAPHRAGM, MOBILIZATION OF THE HILUM ALSO FINE AND OFF THE CAVA SUCCESSFULLY. THERE WAS A SMALL STITCH IN THE RIGHT HEPATIC VEIN, NO CLIPS WERE USED. SINGLE SLING ROUND THE VEIN. ATS45 WITH WHITE RELOAD WAS USED AND ARTICULATED FULLY AT THE TIME. SURGEON FIRED THE DEVICE ON THE RIGHT HEPATIC VEIN AND COULD TELL WITH THE TACTILE FEEDBACK THAT SOMETHING WAS WRONG. HE LEFT THE DEVICE CLOSED AND ASKED THE TEAM TO GET READY FOR SIGNIFICANT BLEEDING FROM THE RIGHT HEPATIC VEIN, ON RELEASE OF THE DEVICE IT WAS EVIDENT THAT THE STAPLE LINE HAD FAILED AND THE STAPLES HADN¿T FORMED CORRECTLY. THERE WERE SOME DEPLOYED STAPLES ON THE DISTAL PART OF THE STAPLE/CUT LINE (SUPERIOR END OF THE VEIN) BUT NO OTHERS WERE FORMED PROPERLY. CONSULTANT TRIED TO STEM TO THE BLEED INITIALLY WITH HIS FINGER BUT THE HOLE INCREASED IN SIZE. THE STAPLER WAS USED AGAIN FOR A SECOND TIME ON PART OF THE LIVER AND DEPLOYED OK. THE RELOAD IN THE DEVICE IS THE SECOND RELOAD THAT DEPLOYED. UNFORTUNATELY, THE RELOAD FROM THE FAILED STAPLE LINE WAS DISCARDED IN THE PRESSURIZED SITUATION AT THE TIME. THE PATIENT NEEDED 12 UNITS OF BLOOD. CLAMP CONTROL FAILED. LOST CONTROL OF BLEEDING TO A LIFE-THREATENING EXSANGUINATION MEASURE. HAD TO OPEN THE CHEST AND PERFORM CARDIAC MASSAGE. EVENTUALLY MANAGED TO GET CONTROL WITH SUTURES. PATIENT IS BACK IN THEATRE TODAY TO REMOVE PART OF THE LIVER THAT NOW HAS NO BLOOD SUPPLY. CHANCES OF SURVIVAL ARE SLIM.
23-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A CASE FOR AN HCC LIVER TUMOUR REMOVAL, THE SURGEON WAS THE CONSULTANT WHO FIRED THE DEVICE, BUT WAS ALSO WITH CONSULTANT IN THEATRE. THEY MOBILIZED THE LIVER THAT WAS SLIGHTLY STUCK TO THE DIAPHRAGM, MOBILIZATION OF THE HILUM ALSO FINE AND OFF THE CAVA SUCCESSFULLY. THERE WAS A SMALL STITCH IN THE RIGHT HEPATIC VEIN, NO CLIPS WERE USED. SINGLE SLING ROUND THE VEIN. ATS45 WITH WHITE RELOAD WAS USED AND ARTICULATED FULLY AT THE TIME. SURGEON FIRED THE DEVICE ON THE RIGHT HEPATIC VEIN AND COULD TELL WITH THE TACTILE FEEDBACK THAT SOMETHING WAS WRONG. HE LEFT THE DEVICE CLOSED AND ASKED THE TEAM TO GET READY FOR SIGNIFICANT BLEEDING FROM THE RIGHT HEPATIC VEIN, ON RELEASE OF THE DEVICE IT WAS EVIDENT THAT THE STAPLE LINE HAD FAILED AND THE STAPLES HADN¿T FORMED CORRECTLY. THERE WERE SOME DEPLOYED STAPLES ON THE DISTAL PART OF THE STAPLE/CUT LINE (SUPERIOR END OF THE VEIN) BUT NO OTHERS WERE FORMED PROPERLY. CONSULTANT TRIED TO STEM TO THE BLEED INITIALLY WITH HIS FINGER BUT THE HOLE INCREASED IN SIZE. THE STAPLER WAS USED AGAIN FOR A SECOND TIME ON PART OF THE LIVER AND DEPLOYED OK. THE RELOAD IN THE DEVICE IS THE SECOND RELOAD THAT DEPLOYED. UNFORTUNATELY, THE RELOAD FROM THE FAILED STAPLE LINE WAS DISCARDED IN THE PRESSURIZED SITUATION AT THE TIME. THE PATIENT NEEDED 12 UNITS OF BLOOD. CLAMP CONTROL FAILED. LOST CONTROL OF BLEEDING TO A LIFE-THREATENING EXSANGUINATION MEASURE. HAD TO OPEN THE CHEST AND PERFORM CARDIAC MASSAGE. EVENTUALLY MANAGED TO GET CONTROL WITH SUTURES. PATIENT IS BACK IN THEATRE TODAY TO REMOVE PART OF THE LIVER THAT NOW HAS NO BLOOD SUPPLY. CHANCES OF SURVIVAL ARE SLIM.
15-09-2016 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Physical Resistance |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE THE STAPLER WAS UNABLE TO FIRE, THE STAPLER HANDLE REMAINED STUCK WHILE USING THE 1ST LOADING UNIT MAKING THE LOADING OF THE 2ND RELOAD IMPOSSIBLE. THEY OPENED A SECOND STAPLER AND THE HANDLE IS ALSO STUCK. THERE WAS NO REPORTED PATIENT OUTCOME.
15-09-2016 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE THE STAPLER WAS UNABLE TO FIRE, THE STAPLER HANDLE REMAINED STUCK WHILE USING THE 1ST LOADING UNIT MAKING THE LOADING OF THE 2ND RELOAD IMPOSSIBLE. THEY OPENED A SECOND STAPLER AND THE HANDLE IS ALSO STUCK. THERE WAS NO REPORTED PATIENT OUTCOME.
15-09-2016 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE THE STAPLER WAS UNABLE TO FIRE, THE STAPLER HANDLE REMAINED STUCK WHILE USING THE 1ST LOADING UNIT MAKING THE LOADING OF THE 2ND RELOAD IMPOSSIBLE. THEY OPENED A SECOND STAPLER AND THE HANDLE IS ALSO STUCK. THERE WAS NO REPORTED PATIENT OUTCOME.
15-09-2016 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE THE STAPLER WAS UNABLE TO FIRE, THE STAPLER HANDLE REMAINED STUCK WHILE USING THE 1ST LOADING UNIT MAKING THE LOADING OF THE 2ND RELOAD IMPOSSIBLE. THEY OPENED A SECOND STAPLER AND THE HANDLE IS ALSO STUCK. THERE WAS NO REPORTED PATIENT OUTCOME.
15-09-2016 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDOGIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE THE STAPLER WAS UNABLE TO FIRE, THE STAPLER HANDLE REMAINED STUCK WHILE USING THE 1ST LOADING UNIT MAKING THE LOADING OF THE 2ND RELOAD IMPOSSIBLE. THEY OPENED A SECOND STAPLER AND THE HANDLE IS ALSO STUCK. THERE WAS NO REPORTED PATIENT OUTCOME.
09-04-2018 | Injury | covidien lp llc north haven | 20-04-2018 | TRI-STAPLE 2.0 | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MINI BYPASS PROCEDURE, WHILE FIRING ,THE DEVICE'S KNIFE STOPPED IN THE MIDDLE AND THE SURGEON COULDN'T CONTINUE. THEY HAD TO CHANGE THE RELOAD WITH ANOTHER ONE WHICH WAS WORKING AND THE OPERATION CONTINUED PROPERLY. NO CONVERSION TO OPEN PROCEDURE WAS MADE. RELOAD WAS REMOVED BY RETRACTING THE KNIFE AND OPENING THE JAWS BY PRESSING THE GREY BUTTON THEN PULLING OUT THE DEVICE FROM THE ABDOMINAL CAVITY
01-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | PROXIMATE*RL LINEAR STAPLER | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE; THE DEVICE MISFIRED. ONLY HALF OF THE STAPLES FIRED. THEY THEN HAD TO BE REMOVED BECAUSE THE STAPLE LEGS DID NOT BEND TO COMPLETE THE STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | PROXIMATE*RL LINEAR STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE; THE DEVICE MISFIRED. ONLY HALF OF THE STAPLES FIRED. THEY THEN HAD TO BE REMOVED BECAUSE THE STAPLE LEGS DID NOT BEND TO COMPLETE THE STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
23-03-2018 | Injury | ethicon endo-surgery | 20-04-2018 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT THE STAPLER TRIGGERED AND CUT NORMALLY BUT DID NOT STAPLE.
07-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, ON VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), WHEN CUTTING THE LUNG TISSUE, THE STAPLER WAS ABLE TO FIRE AND THE STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL AREA. AS STATED. THE RELOAD WAS NOT SMOOTH DURING FIRING, AND A SHORT STAPLE LINE WAS PERFORMED ON THE TISSUE BUT CANNOT CONTINUE FULL STAPLE LINE DUE TO DIFFICULTIES TO PUSH ANYMORE. THE DOCTOR THEN OPENED A NEW RELOAD IN THE SAME BATCH AND WAS OKAY, TO COMPLETE THE STAPLE LINE SMOOTHLY. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A PANCREATECTOMY, PERFORMED BY THE SURGEON, THE HAND GRIP BROKE OFF. NO PATIENT CONSEQUENCE REPORTED.
28-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | 75MM SELECTABLE NEW TLC | Detachment Of Device Component |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A PANCREATECTOMY, PERFORMED BY THE SURGEON, THE HAND GRIP BROKE OFF. NO PATIENT CONSEQUENCE REPORTED.
26-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN LOBECTOMY PROCEDURE THE DEVICE FIRED ACROSS THE PULMONARY ARTERY, THE BLADE RETRACTED, BUT THE REVERSE BUTTON DID NOT WORK TO OPEN THE JAWS. THE MANUAL OVER RIDE WAS TRIED, BUT IT DID NOT WORK EITHER. AFTER A COUPLE OF MINUTES THE JAWS JUST RANDOMLY POPPED OPEN ON THEIR OWN. THE PROCEDURE WAS COMPLETE AT THIS POINT. THE CUT AND STAPLE LINE WERE GOOD. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
26-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN LOBECTOMY PROCEDURE THE DEVICE FIRED ACROSS THE PULMONARY ARTERY, THE BLADE RETRACTED, BUT THE REVERSE BUTTON DID NOT WORK TO OPEN THE JAWS. THE MANUAL OVER RIDE WAS TRIED, BUT IT DID NOT WORK EITHER. AFTER A COUPLE OF MINUTES THE JAWS JUST RANDOMLY POPPED OPEN ON THEIR OWN. THE PROCEDURE WAS COMPLETE AT THIS POINT. THE CUT AND STAPLE LINE WERE GOOD. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
05-02-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC GASTROPLASTY, IN AN ATTEMPT TO FIRE THE GREEN LOAD, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE NOR PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
26-01-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | CONTOUR CURVED CUTTER STAPLER | Connection Problem | A12 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE PHYSICIAN REPORTED THAT HE COULD NOT CONNECT THE LOAD WITH STAPLER. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | GIA | Fluid/Blood Leak | A050401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE PATIENT HAD INCURRED AN ANASTOMOTIC LEAK DUE TO THE STAPLE LINE.
06-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THERE WAS A DIFFICULTY IN LOADING AND UNLOADING OF THE DEVICE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WASN'T ABLE TO SQUEEZE THE HANDLE.
06-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THERE WAS A DIFFICULTY IN LOADING AND UNLOADING OF THE DEVICE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WASN'T ABLE TO SQUEEZE THE HANDLE.
08-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LATERAL LOBECTOMY OF LIVER, WHILE DISCONNECTING LEFT HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE WAS UNABLE TO FIRE AND SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED WITH A NEW ONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT INJURY.
06-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THERE WAS A DIFFICULTY IN LOADING AND UNLOADING OF THE DEVICE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WASN'T ABLE TO SQUEEZE THE HANDLE.
06-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THERE WAS A DIFFICULTY IN LOADING AND UNLOADING OF THE DEVICE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WASN'T ABLE TO SQUEEZE THE HANDLE.
20-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | GIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A AN OPEN PROCEDURE, THE PATIENT HAD INCURRED AN ANASTOMOTIC LEAK DUE TO THE STAPLE LINE.
20-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | GIA | Fluid/Blood Leak | A050401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A AN OPEN PROCEDURE, THE PATIENT HAD INCURRED AN ANASTOMOTIC LEAK DUE TO THE STAPLE LINE.
06-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THERE WAS A DIFFICULTY IN LOADING AND UNLOADING OF THE DEVICE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WASN'T ABLE TO SQUEEZE THE HANDLE.
06-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THERE WAS A DIFFICULTY IN LOADING AND UNLOADING OF THE DEVICE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WASN'T ABLE TO SQUEEZE THE HANDLE.
20-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | EGIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE PATIENT HAD INCURRED AN ANASTOMOTIC LEAK DUE TO THE STAPLE LINE.
01-01-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | SGNIA | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, THE HANDLE WAS DROPPED BY AN UNKNOWN PERSON TO THE GROUND. THE HANDLE'S CHARGING TRAY HAD NO SAFE STAND, THEREFORE IT CAN BE TILTED OUT EVEN WITH LIGHT CONTACT AND FALL TO THE GROUND. THE DISPLAY WAS ALSO SILVERED. THERE WAS NO PATIENT INVOLVED.
02-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE FIRST FIRING DURING A WEDGE/SEGMENTAL RESECTION OF LUNG, THEY CONNECTED THE CARTRIDGE TO THE HANDLE AND CHECKED THE JAWS OPENING/CLOSING WITH NO PROBLEM. THE SURGEON CLAMPED THE TISSUE AND TRIED TO FIRE THE DEVICE, HOWEVER COULD NOT SQUEEZE THE HANDLE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
06-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHILE THE DUODENUM WAS BEING TRANSECTED, THE JAW OF THE STAPLER WAS CLOSED ON THE TISSUE AFTER A RELOAD WAS LOADED. THE JAWS WERE OPENED AND THE STAPLER WAS REMOVED. AFTER DEPRESSING THE GREEN READY TO FIRE BUTTON, THE LOOP HANDLE ADVANCED FORWARD. ONCE THE DOCTOR STARTED TO SQUEEZE THE STAPLER HANDLE, IT DID NOT FIRE THE RELOAD/ACTIVATE THE RATCHET TO DEPLOY THE STAPLES. THERE WAS NO PATIENT INJURY.
08-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA | Difficult or Delayed Positioning | A150203 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VATS RADICAL RESECTION OF LUNG CANCER, WHILE THE DEVICE WAS BEING APPLIED TO DISCONNECTED BRONCHIAL TUBE, THE HANDLE NEEDED TO BE SQUEEZED VERY HARD AND AFTER SOME STAPLES WERE FIRED, IT COULD NOT BE FIRED CONTINUOUSLY. IT ALSO COULD NOT BE FIRED AFTER REPLACING WITH TWO RELOADS. THE HANDLE SOUNDED WHEN FIRING THE THIRD RELOAD AND COULD NOT CONTINUE FIRING. THERE WAS NO PATIENT INJURY.
08-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VATS RADICAL RESECTION OF LUNG CANCER, WHILE THE DEVICE WAS BEING APPLIED TO DISCONNECTED BRONCHIAL TUBE, THE HANDLE NEEDED TO BE SQUEEZED VERY HARD AND AFTER SOME STAPLES WERE FIRED, IT COULD NOT BE FIRED CONTINUOUSLY. IT ALSO COULD NOT BE FIRED AFTER REPLACING WITH TWO RELOADS. THE HANDLE SOUNDED WHEN FIRING THE THIRD RELOAD AND COULD NOT CONTINUE FIRING. THERE WAS NO PATIENT INJURY.
08-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VATS RADICAL RESECTION OF LUNG CANCER, WHILE THE DEVICE WAS BEING APPLIED TO DISCONNECTED BRONCHIAL TUBE, THE HANDLE NEEDED TO BE SQUEEZED VERY HARD AND AFTER SOME STAPLES WERE FIRED, IT COULD NOT BE FIRED CONTINUOUSLY. IT ALSO COULD NOT BE FIRED AFTER REPLACING WITH TWO RELOADS. THE HANDLE SOUNDED WHEN FIRING THE THIRD RELOAD AND COULD NOT CONTINUE FIRING. THERE WAS NO PATIENT INJURY.
07-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | EGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD A FIRING ISSUE. PATIENT STATUS WAS UNSPECIFIED.
09-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | EGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG CANCER RADICAL SURGERY, TO MAKE A PULMONARY LOBECTOMY, STAPLES DID NOT DEPLOY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS SQUEEZED. THERE WAS NO PATIENT INJURY.
05-02-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC GASTROPLASTY, IN AN ATTEMPT TO FIRE THE GREEN LOAD, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE NOR PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
05-04-2018 | Injury | us surgical puerto rico | 20-04-2018 | EEA | Failure To Adhere Or Bond |  | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR FINAL ANASTOMOSIS, THE DEVICE WAS DIFFICULT OR UNABLE TO LOAD. DURING THE MATING OF THE PRE TILTED ANVIL OF THE DEVICE AND THE STAPLER HANDLE OF THE SURGEON STRUGGLED TO CONNECT THE 2 COMPONENTS. THIS LED TO A LOT OF STRESS ON THE TISSUE THAT WAS TO BE ANASTOMOSED. AFTER MANY ATTEMPTS, THE COMPONENTS WERE SUCCESSFULLY MATED AND THE DEVICE FUNCTIONED AS NORMAL AND THE ANASTOMOSIS WAS CREATED. THE PATIENT IS STILL IN HOSPITAL AND BEING OBSERVED TO MAKE SURE THE TISSUE HEALS AND IS NOT TORN IN ANY WAY.
05-04-2018 | Injury | us surgical puerto rico | 20-04-2018 | EEA | Failure To Adhere Or Bond |  | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR FINAL ANASTOMOSIS, THE DEVICE WAS DIFFICULT OR UNABLE TO LOAD. DURING THE MATING OF THE PRE TILTED ANVIL OF THE DEVICE AND THE STAPLER HANDLE OF THE SURGEON STRUGGLED TO CONNECT THE 2 COMPONENTS. THIS LED TO A LOT OF STRESS ON THE TISSUE THAT WAS TO BE ANASTOMOSED. AFTER MANY ATTEMPTS, THE COMPONENTS WERE SUCCESSFULLY MATED AND THE DEVICE FUNCTIONED AS NORMAL AND THE ANASTOMOSIS WAS CREATED. THE PATIENT IS STILL IN HOSPITAL AND BEING OBSERVED TO MAKE SURE THE TISSUE HEALS AND IS NOT TORN IN ANY WAY.
06-04-2018 | Injury | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Mechanics Altered | A0507 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT SUPERIOR LOBECTOMY, WHILE BEING APPLIED AT THE SECTION OF THE SUPERIOR PULMONARY VEIN, AFTER ACTIVATION OF THE STAPLER, NO STAPLE CAME OUT, BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE PROCEDURE WAS CONVERTED TO A THORACOTOMY LEADING TO AN EXTENDED SURGICAL TIME OF MORE THAN THIRTY MINUTES. THEY NOTED A BLOOD LOSS OF MORE THAN 500 CC AND THE INCISION WAS ALSO EXTENDED MORE THAN ONE INCH. THEY COMPLETED THE CASE BY SETTING UP THE CLIP.
06-04-2018 | Injury | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT SUPERIOR LOBECTOMY, WHILE BEING APPLIED AT THE SECTION OF THE SUPERIOR PULMONARY VEIN, AFTER ACTIVATION OF THE STAPLER, NO STAPLE CAME OUT, BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE PROCEDURE WAS CONVERTED TO A THORACOTOMY LEADING TO AN EXTENDED SURGICAL TIME OF MORE THAN THIRTY MINUTES. THEY NOTED A BLOOD LOSS OF MORE THAN 500 CC AND THE INCISION WAS ALSO EXTENDED MORE THAN ONE INCH. THEY COMPLETED THE CASE BY SETTING UP THE CLIP.
09-04-2018 | Injury | covidien lp llc north haven | 20-04-2018 | IDRIVE ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MINI BYPASS PROCEDURE, WHILE FIRING, THE DEVICE'S KNIFE STOPPED IN THE MIDDLE AND THE SURGEON COULDN'T CONTINUE. THEY HAD TO CHANGE THE RELOAD WITH ANOTHER ONE WHICH WAS WORKING AND THE OPERATION CONTINUED PROPERLY. NO CONVERSION TO OPEN PROCEDURE WAS MADE. RELOAD WAS REMOVED BY RETRACTING THE KNIFE AND OPENING THE JAWS BY PRESSING THE GREY BUTTON THEN PULLING OUT THE DEVICE FROM THE ABDOMINAL CAVITY.
28-03-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY FOR THE 1ST FIRING, THE SURGEON COULD NOT FIRE THE DEVICE. THERE WAS NO PATIENT INJURY.
06-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER RESECTION PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE AND THE STAPLE WAS NOT ABLE TO FIRE. THE RELOAD WENT ON THE DEVICE AT THE BACK OF THE SURGICAL TABLE BUT THE HANDLE OF THE DEVICE COULD NOT BE SQUEEZED EVEN IF THE GREEN SAFETY BUTTON WAS PRESSED. IT WAS TO BE THE LAST FIRING OF THE CASE AS IT WORKED FINE ON THE SUBSEQUENT TWO FIRINGS. ANOTHER HANDLE WAS OPENED AND USED TO RESOLVE THE ISSUE.
06-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER RESECTION PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE AND THE STAPLE WAS NOT ABLE TO FIRE. THE RELOAD WENT ON THE DEVICE AT THE BACK OF THE SURGICAL TABLE BUT THE HANDLE OF THE DEVICE COULD NOT BE SQUEEZED EVEN IF THE GREEN SAFETY BUTTON WAS PRESSED. IT WAS TO BE THE LAST FIRING OF THE CASE AS IT WORKED FINE ON THE SUBSEQUENT TWO FIRINGS. ANOTHER HANDLE WAS OPENED AND USED TO RESOLVE THE ISSUE.
01-03-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ABSORBATACK | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC HERNIA REPAIR, WHEN THE SURGEON INSERTED THE INSTRUMENT AND TRIED TO PLACE THE TACK, WHILE POSITIONING, THE SHAFT CRACKED. NO TACKS WERE FIRED FROM THE DEVICE. THERE WAS NO PATIENT HARM.
26-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RESECTION OF ESOPHAGEAL CARCINOMA, WHEN FIRING, THE FIRING KNOB WAS BROKEN OFF. ALL THE PIECES WOULD BE RETURNED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
26-03-2018 | Malfunction | ethicon endo-surgery | 20-04-2018 | 75MM SELECTABLE NEW TLC | Detachment Of Device Component |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RESECTION OF ESOPHAGEAL CARCINOMA, WHEN FIRING, THE FIRING KNOB WAS BROKEN OFF. ALL THE PIECES WOULD BE RETURNED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
20-03-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO LOBECTOMY PROCEDURE, THE SURGEON COULD NOT FIRE THE DEVICE. THERE WAS NO PATIENT INJURY.
07-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, ON VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), WHEN CUTTING THE LUNG TISSUE, THE STAPLER WAS ABLE TO FIRE AND THE STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL AREA. AS STATED. THE RELOAD WAS NOT SMOOTH DURING FIRING, AND A SHORT STAPLE LINE WAS PERFORMED ON THE TISSUE BUT CANNOT CONTINUE FULL STAPLE LINE DUE TO DIFFICULTIES TO PUSH ANYMORE. THE DOCTOR THEN OPENED A NEW RELOAD IN THE SAME BATCH AND WAS OKAY, TO COMPLETE THE STAPLE LINE SMOOTHLY. THERE WAS NO PATIENT INJURY.
08-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Difficult or Delayed Positioning | A150203 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VATS RADICAL RESECTION OF LUNG CANCER, WHILE THE DEVICE WAS BEING APPLIED TO DISCONNECTED BRONCHIAL TUBE, THE HANDLE NEEDED TO BE SQUEEZED VERY HARD AND AFTER SOME STAPLES WERE FIRED, IT COULD NOT BE FIRED CONTINUOUSLY. IT ALSO COULD NOT BE FIRED AFTER REPLACING WITH TWO RELOADS. THE HANDLE SOUNDED WHEN FIRING THE THIRD RELOAD AND COULD NOT CONTINUE FIRING. THEY REPLACED A NEW DEVICE TO COMPLETE THE OPERATION. NO PATIENT INJURY.
08-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VATS RADICAL RESECTION OF LUNG CANCER, WHILE THE DEVICE WAS BEING APPLIED TO DISCONNECTED BRONCHIAL TUBE, THE HANDLE NEEDED TO BE SQUEEZED VERY HARD AND AFTER SOME STAPLES WERE FIRED, IT COULD NOT BE FIRED CONTINUOUSLY. IT ALSO COULD NOT BE FIRED AFTER REPLACING WITH TWO RELOADS. THE HANDLE SOUNDED WHEN FIRING THE THIRD RELOAD AND COULD NOT CONTINUE FIRING. THEY REPLACED A NEW DEVICE TO COMPLETE THE OPERATION. NO PATIENT INJURY.
08-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VATS RADICAL RESECTION OF LUNG CANCER, WHILE THE DEVICE WAS BEING APPLIED TO DISCONNECTED BRONCHIAL TUBE, THE HANDLE NEEDED TO BE SQUEEZED VERY HARD AND AFTER SOME STAPLES WERE FIRED, IT COULD NOT BE FIRED CONTINUOUSLY. IT ALSO COULD NOT BE FIRED AFTER REPLACING WITH TWO RELOADS. THE HANDLE SOUNDED WHEN FIRING THE THIRD RELOAD AND COULD NOT CONTINUE FIRING. THEY REPLACED A NEW DEVICE TO COMPLETE THE OPERATION. NO PATIENT INJURY.
09-04-2018 | Injury | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE DISSECTION OF PANCREAS ON A LAPAROSCOPIC PANCREATECTOMY, THE DOCTOR TRIED TO FIRE STAPLER BUT HE WAS UNABLE TO SQUEEZE HANDLE. SOME BREAKING SOUND CAME FROM THE STAPLER AND THEN HANDLE WAS SQUEEZED BUT BLADE WAS NOT ADVANCING. THE INCISION WAS EXTENDED TO MORE THAN ONE INCH. THE LAPAROSCOPIC CASE WAS CONVERTED INTO OPEN PROCEDURE. THE PANCREAS WAS CUT USING A BLADE AND HAND SEWED TO COMPLETE THE CASE.
08-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LATERAL LOBECTOMY OF LIVER, WHILE DISCONNECTING LEFT HEPATIC VEIN, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE WAS UNABLE TO FIRE AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
06-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER RESECTION PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE AND THE STAPLE WAS NOT ABLE TO FIRE. THE RELOAD WENT ON THE DEVICE AT THE BACK OF THE SURGICAL TABLE BUT THE HANDLE OF THE DEVICE COULD NOT BE SQUEEZED EVEN IF THE GREEN SAFETY BUTTON WAS PRESSED. IT WAS TO BE THE LAST FIRING OF THE CASE AS IT WORKED FINE ON THE SUBSEQUENT TWO FIRINGS. ANOTHER HANDLE WAS OPENED AND USED TO RESOLVE THE ISSUE.
06-04-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDOGIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER RESECTION PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE AND THE STAPLE WAS NOT ABLE TO FIRE. THE RELOAD WENT ON THE DEVICE AT THE BACK OF THE SURGICAL TABLE BUT THE HANDLE OF THE DEVICE COULD NOT BE SQUEEZED EVEN IF THE GREEN SAFETY BUTTON WAS PRESSED. IT WAS TO BE THE LAST FIRING OF THE CASE AS IT WORKED FINE ON THE SUBSEQUENT TWO FIRINGS. ANOTHER HANDLE WAS OPENED AND USED TO RESOLVE THE ISSUE.
28-03-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY FOR THE 1ST FIRING, THE SURGEON COULD NOT FIRE THE DEVICE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE FIRST FIRING DURING A WEDGE/SEGMENTAL RESECTION OF LUNG, THEY CONNECTED THE CARTRIDGE TO THE HANDLE AND CHECKED THE JAWS OPENING/CLOSING WITH NO PROBLEM. THE SURGEON CLAMPED THE TISSUE AND TRIED TO FIRE THE DEVICE, HOWEVER COULD NOT SQUEEZE THE HANDLE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
05-04-2018 | Injury | covidien north haven-mfg | 20-04-2018 | EEA ORVIL | Failure To Adhere Or Bond |  | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR FINAL ANASTOMOSIS, THE DEVICE WAS DIFFICULT OR UNABLE TO LOAD. DURING THE MATING OF THE PRE TILTED ANVIL OF THE DEVICE AND THE STAPLER HANDLE OF THE SURGEON STRUGGLED TO CONNECT THE 2 COMPONENTS. THIS LED TO A LOT OF STRESS ON THE TISSUE THAT WAS TO BE ANASTOMOSED. AFTER MANY ATTEMPTS, THE COMPONENTS WERE SUCCESSFULLY MATED AND THE DEVICE FUNCTIONED AS NORMAL AND THE ANASTOMOSIS WAS CREATED. THE PATIENT IS STILL IN HOSPITAL AND BEING OBSERVED TO MAKE SURE THE TISSUE HEALS AND IS NOT TORN IN ANY WAY.
05-04-2018 | Injury | covidien north haven-mfg | 20-04-2018 | EEA ORVIL | Failure To Adhere Or Bond |  | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR FINAL ANASTOMOSIS, THE DEVICE WAS DIFFICULT OR UNABLE TO LOAD. DURING THE MATING OF THE PRE TILTED ANVIL OF THE DEVICE AND THE STAPLER HANDLE OF THE SURGEON STRUGGLED TO CONNECT THE 2 COMPONENTS. THIS LED TO A LOT OF STRESS ON THE TISSUE THAT WAS TO BE ANASTOMOSED. AFTER MANY ATTEMPTS, THE COMPONENTS WERE SUCCESSFULLY MATED AND THE DEVICE FUNCTIONED AS NORMAL AND THE ANASTOMOSIS WAS CREATED. THE PATIENT IS STILL IN HOSPITAL AND BEING OBSERVED TO MAKE SURE THE TISSUE HEALS AND IS NOT TORN IN ANY WAY.
09-04-2018 | Injury | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MINI BYPASS PROCEDURE, WHILE FIRING ,THE DEVICE'S KNIFE STOPPED IN THE MIDDLE AND THE SURGEON COULDN'T CONTINUE. THEY HAD TO CHANGE THE RELOAD WITH ANOTHER ONE WHICH WAS WORKING AND THE OPERATION CONTINUED PROPERLY. NO CONVERSION TO OPEN PROCEDURE WAS MADE. RELOAD WAS REMOVED BY RETRACTING THE KNIFE AND OPENING THE JAWS BY PRESSING THE GREY BUTTON THEN PULLING OUT THE DEVICE FROM THE ABDOMINAL CAVITY.
20-03-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO LOBECTOMY PROCEDURE, THE SURGEON COULD NOT FIRE THE DEVICE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
08-12-2017 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PARTIAL GASTRECTOMY, WHILE TRANSECTING THE FUNDUS OF STOMACH, THE DEVICE HAD POOR STAPLE FORMATION. WHEN SQUEEZING THE HANDLE AND FIRING ONE OF THE RELOADS, THERE WAS A CRACK SOUND AND SOME OF THE STAPLES DID NOT FORM WELL, SOME REMAINED C-SHAPED. WHEN ATTACHED WITH ANOTHER RELOAD, THE SITUATION STAYED THE SAME. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO INJURY CAUSED TO THE PATIENT AND NO MEDICAL INTERVENTION WAS REQUIRED.
26-03-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, WHEN RECHARGING THE STAPLER, THE SURGERY WAS NOT POSSIBLE TO SHOOT SINCE THE CABLE WAS LOCKED. IT WAS INDICATED THAT THE GREEN BUTTON WAS ABLE TO BE PRESSED BUT THE HANDLE WAS NOT ABLE TO BE SQUEEZED. THE LOAD WAS REPLACED AND THE PROBLEM PERSISTED. IT WAS ALSO NOTED THAT THE HANDLE BROKE OFF. THE STAPLER WAS REPLACED AND THE SURGERY GAVE CONTINUITY. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, WHEN RECHARGING THE STAPLER, THE SURGERY WAS NOT POSSIBLE TO SHOOT SINCE THE CABLE WAS LOCKED. IT WAS INDICATED THAT THE GREEN BUTTON WAS ABLE TO BE PRESSED BUT THE HANDLE WAS NOT ABLE TO BE SQUEEZED. THE LOAD WAS REPLACED AND THE PROBLEM PERSISTED. IT WAS ALSO NOTED THAT THE HANDLE BROKE OFF. THE STAPLER WAS REPLACED AND THE SURGERY GAVE CONTINUITY. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, WHEN RECHARGING THE STAPLER, THE SURGERY WAS NOT POSSIBLE TO SHOOT SINCE THE CABLE WAS LOCKED. IT WAS INDICATED THAT THE GREEN BUTTON WAS ABLE TO BE PRESSED BUT THE HANDLE WAS NOT ABLE TO BE SQUEEZED. THE LOAD WAS REPLACED AND THE PROBLEM PERSISTED. IT WAS ALSO NOTED THAT THE HANDLE BROKE OFF. THE STAPLER WAS REPLACED AND THE SURGERY GAVE CONTINUITY. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | covidien north haven-mfg | 20-04-2018 | ENDO GIA | Fluid/Blood Leak | A050401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE STAPLING OF THE DUODENUM WITH A 45MM STAPLE WAS SUCCESSFUL. THE SURGEON THEN FIRED TO THE STOMACH OF MOUTH SIDE A 60MM RELOAD. DURING THE FOLLOWING FIRING WITH THE 60MM STAPLE, THE SURGEON NOTED STOMACH CONTENT LEAKED FROM THE CUT END OF THE PREVIOUS FIRING. THE SURGEON NOTED SOME STAPLES WERE U-FORMED AND THE STAPLING WAS INCOMPLETE. THE SURGEON ADDITIONALLY RESECTED BY 1 OR 2 CM OF THE STOMACH OF MOUTH SIDE WITH 2 UNITS OF THE 60MM STAPLE. THE STAPLING WERE COMPLETED WITH NO PROBLEM. THE SURGEON CONTINUED TO USE THE SAME DEVICE AND COMPLETED THE ANASTOMOSIS WITH A 45MM STAPLE AND CLOSED THE ENTRY HOLE WITH 60MM STAPLE. THERE WAS NO PATIENT HARM.
09-04-2018 | Malfunction | covidien north haven-mfg | 20-04-2018 | ENDO GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE STAPLING OF THE DUODENUM WITH A 45MM STAPLE WAS SUCCESSFUL. THE SURGEON THEN FIRED TO THE STOMACH OF MOUTH SIDE A 60MM RELOAD. DURING THE FOLLOWING FIRING WITH THE 60MM STAPLE, THE SURGEON NOTED STOMACH CONTENT LEAKED FROM THE CUT END OF THE PREVIOUS FIRING. THE SURGEON NOTED SOME STAPLES WERE U-FORMED AND THE STAPLING WAS INCOMPLETE. THE SURGEON ADDITIONALLY RESECTED BY 1 OR 2 CM OF THE STOMACH OF MOUTH SIDE WITH 2 UNITS OF THE 60MM STAPLE. THE STAPLING WERE COMPLETED WITH NO PROBLEM. THE SURGEON CONTINUED TO USE THE SAME DEVICE AND COMPLETED THE ANASTOMOSIS WITH A 45MM STAPLE AND CLOSED THE ENTRY HOLE WITH 60MM STAPLE. THERE WAS NO PATIENT HARM.
09-04-2018 | Malfunction | covidien north haven-mfg | 20-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE STAPLING OF THE DUODENUM WITH A 45MM STAPLE WAS SUCCESSFUL. THE SURGEON THEN FIRED TO THE STOMACH OF MOUTH SIDE A 60MM RELOAD. DURING THE FOLLOWING FIRING WITH THE 60MM STAPLE, THE SURGEON NOTED STOMACH CONTENT LEAKED FROM THE CUT END OF THE PREVIOUS FIRING. THE SURGEON NOTED SOME STAPLES WERE U-FORMED AND THE STAPLING WAS INCOMPLETE. THE SURGEON ADDITIONALLY RESECTED BY 1 OR 2 CM OF THE STOMACH OF MOUTH SIDE WITH 2 UNITS OF THE 60MM STAPLE. THE STAPLING WERE COMPLETED WITH NO PROBLEM. THE SURGEON CONTINUED TO USE THE SAME DEVICE AND COMPLETED THE ANASTOMOSIS WITH A 45MM STAPLE AND CLOSED THE ENTRY HOLE WITH 60MM STAPLE. THERE WAS NO PATIENT HARM.
09-04-2018 | Malfunction | covidien north haven-mfg | 20-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE STAPLING OF THE DUODENUM WITH A 45MM STAPLE WAS SUCCESSFUL. THE SURGEON THEN FIRED TO THE STOMACH OF MOUTH SIDE A 60MM RELOAD. DURING THE FOLLOWING FIRING WITH THE 60MM STAPLE, THE SURGEON NOTED STOMACH CONTENT LEAKED FROM THE CUT END OF THE PREVIOUS FIRING. THE SURGEON NOTED SOME STAPLES WERE U-FORMED AND THE STAPLING WAS INCOMPLETE. THE SURGEON ADDITIONALLY RESECTED BY 1 OR 2 CM OF THE STOMACH OF MOUTH SIDE WITH 2 UNITS OF THE 60MM STAPLE. THE STAPLING WERE COMPLETED WITH NO PROBLEM. THE SURGEON CONTINUED TO USE THE SAME DEVICE AND COMPLETED THE ANASTOMOSIS WITH A 45MM STAPLE AND CLOSED THE ENTRY HOLE WITH 60MM STAPLE. THERE WAS NO PATIENT HARM.
09-02-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC GASTROPLASTY, AFTER ONE SHOT, THE STAPLER BROKE, IT WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE NOR PRESS THE GREEN BUTTON. THEY OPENED A NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-04-2018 | Malfunction | covidien medical products | 20-04-2018 | PREMIUM PLUS CEEA | Failure To Adhere Or Bond |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE ANVIL DID NOT CONNECT WITH THE STAPLER PROPERLY.
30-03-2018 | Malfunction | us surgical puerto rico | 20-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACIC LOBECTOMY, THE STAPLER WAS ABLE TO FIRE BUT AT THE THIRD FEEDER, IT JAMMED ON THE LOBAR BRONCHUS. THE STAPLER WAS TESTED AND SEEING THAT IT STOPPED RESPONDING, IT WAS IMMEDIATELY CHANGED WITH ANOTHER DEVICE. THE INSTRUMENT COULD BE REOPENED BUT THE STAPLER WAS BLOCKED DURING STAPLING OF THE BRONCHUS. THERE WAS NO PATIENT INJURY.
09-02-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | EGIAUXL | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC GASTROPLASTY, AFTER ONE SHOT, THE STAPLER BROKE, THE SURGEON WAS UNABLE TO FIRE AND SQUEEZE THE HANDLE NOR PRESS THE GREEN BUTTON. THEY OPENED A NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | covidien lp llc north haven | 20-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACIC LOBECTOMY, THE STAPLER WAS ABLE TO FIRE BUT AT THE THIRD FEEDER IT JAMMED ON THE LOBAR BRONCHUS. THE STAPLER WAS TESTED AND SEEING THAT IT STOPPED RESPONDING, IT WAS IMMEDIATELY CHANGED WITH ANOTHER DEVICE. THE INSTRUMENT COULD BE REOPENED BUT THE STAPLER WAS BLOCKED DURING STAPLING OF THE BRONCHUS. THERE WAS NO PATIENT INJURY.
07-03-2018 | Malfunction | ethicon | 23-04-2018 | SECURESTRAP UNKNOWN PRODUCT | Failure To Adhere Or Bond |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND ABSORBABLE STRAPS WERE USED. DURING THE PROCEDURE THE STRAPS DID NOT SECURE TO TISSUE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
14-03-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THEY TRIED TO STAPLE THE BASE OF THE APPENDIX ON A LAPAROSCOPIC APPENDECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE AND THE STAPLER WAS UNABLE TO FIRE. THERE WAS NO PATIENT INJURY.
23-04-2018 | Injury | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, THE DEVICE¿S FAILURE WAS NOTICED AFTER A LAPAROSCOPIC PROCEDURE. THIS RESULTED TO DAMAGED TISSUE. THE PATIENT HAD AN EXTENDED HOSPITALIZATION FOR 12 DAYS.
14-03-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO APPENDIX ON A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON OF THE DEVICE BUT THE STAPLE WAS NOT ABLE TO FIRE AND THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE JAW DID NOT OPEN. NO PATIENT CONSEQUENCES WERE REPORTED.
05-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT JEJUNUM RESECTION OF ILEAL CONDUIT IN CYSTECTOMY, THE SURGEON COULD NOT FIRE THE DEVICE WHEREIN THE HANDLE COULD NOT BE SQUEEZED EVEN AFTER PRESSING THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LOBECTOMY, THE STAPLER WAS ABLE TO FIRE THE FIRST TIME BUT NOT ANYMORE AFTER THE SECOND FIRING. SOME PARTS OF THE STAPLE FORMATION WERE ALSO POORLY FORMED. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LOBECTOMY, THE STAPLER WAS ABLE TO FIRE THE FIRST TIME BUT NOT ANYMORE AFTER THE SECOND FIRING. SOME PARTS OF THE STAPLE FORMATION WERE ALSO POORLY FORMED. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC PROCEDURE, THE STAPLE LINE WAS POOR AND THE STAPLES WERE FALLING OUT OF THE RELOAD. A NEW STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC PROCEDURE, THE STAPLE LINE WAS POOR AND THE STAPLES WERE FALLING OUT OF THE RELOAD. A NEW STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC PROCEDURE, THE STAPLE LINE WAS POOR AND THE STAPLES WERE FALLING OUT OF THE RELOAD. A NEW STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT UPPER LOBE RESECTION, WHILE THE DEVICE WAS BEING APPLIED NOT ON TISSUE OR BLOOD VESSELS, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. TO RESOLVE THE ISSUE, THEY REPLACED ANOTHER STAPLE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT UPPER LOBE RESECTION, WHILE THE DEVICE WAS BEING APPLIED NOT ON TISSUE OR BLOOD VESSELS, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. TO RESOLVE THE ISSUE, THEY REPLACED ANOTHER STAPLE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Unintended System Motion | A051206 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTAL RESECTION, WHEN THE DEVICE APPROACHED THROUGH THE HOLE OF WOUND AND WENT INTO THE CAVITY SMOOTHLY, A CRACK WAS HEARD. THE SURGEON COULD OPEN AND CLOSE THE JAWS, THEN THE GREEN BUTTONS ILLUMINATED AND HE/SHE STARTED FIRING THE DEVICE. DURING THE FIRING, THE CARTRIDGE ARTICULATED ON THEIR OWN AND COULD NOT FIRE TO THE PLANNED STAPLE LINE. THEY CHANGED THE ADAPTER AND THE SURGEON MADE ANOTHER RESECTION DUE TO THE ISSUE. THERE WAS NO TISSUE DAMAGE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTAL RESECTION, WHEN THE DEVICE APPROACHED THROUGH THE HOLE OF WOUND AND WENT INTO THE CAVITY SMOOTHLY, A CRACK WAS HEARD. THE SURGEON COULD OPEN AND CLOSE THE JAWS, THEN THE GREEN BUTTONS ILLUMINATED AND HE/SHE STARTED FIRING THE DEVICE. DURING THE FIRING, THE CARTRIDGE ARTICULATED ON THEIR OWN AND COULD NOT FIRE TO THE PLANNED STAPLE LINE. THEY CHANGED THE ADAPTER AND THE SURGEON MADE ANOTHER RESECTION DUE TO THE ISSUE. THERE WAS NO TISSUE DAMAGE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Injury | ethicon endo-surgery | 23-04-2018 | ILS 29MM, CURVED | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC REVERSAL OF HARTMANN¿S PROCEDURE THERE WAS AN ISSUE WITH THE DEVICE. THEREFORE, THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN TO COMPLETE THE ANASTOMOSIS AND ANOTHER LIKE DEVICE WAS USED. THE ANVIL WAS IN THE COLON, SPIKE OUT OF RECTUM ¿ UNITED (BUT NO CLICK DESPITE CLEARLY UNITED). CLOSED DOWN INTO GREEN ZONE. WAITED, SAFETY CATCH OFF, FIRED AND THERE WAS NO ¿CRUNCH¿. GUN OPENED ¿ HEAD OF GUN VISIBLE IE HOLE IN RECTAL STUMP. ANASTOMOSIS NOT SALVAGEABLE LAPAROSCOPICALLY. CANNOT DISENGAGE ANVIL. GUN CLOSED AGAIN AND 2ND ATTEMPT MADE TO FIRE IN ORDER TO TRY AND MINIMIZE THE ENTEROTOMY ( IE CONTAMINATION) AND TO PERMIT REMOVAL OF DEVICE. THIS TIME THERE WAS A CRUNCH ON FIRING. GUN OPENED. ANASTOMOSIS DISRUPTED AS EXPECTED. ANVIL DISENGAGED AND REMOVED PER ABDOMEN.
27-03-2018 | Injury | ethicon endo-surgery | 23-04-2018 | ILS 29MM, CURVED | Mechanical Jam | A0506 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC REVERSAL OF HARTMANN¿S PROCEDURE THERE WAS AN ISSUE WITH THE DEVICE. THEREFORE, THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN TO COMPLETE THE ANASTOMOSIS AND ANOTHER LIKE DEVICE WAS USED. THE ANVIL WAS IN THE COLON, SPIKE OUT OF RECTUM ¿ UNITED (BUT NO CLICK DESPITE CLEARLY UNITED). CLOSED DOWN INTO GREEN ZONE. WAITED, SAFETY CATCH OFF, FIRED AND THERE WAS NO ¿CRUNCH¿. GUN OPENED ¿ HEAD OF GUN VISIBLE IE HOLE IN RECTAL STUMP. ANASTOMOSIS NOT SALVAGEABLE LAPAROSCOPICALLY. CANNOT DISENGAGE ANVIL. GUN CLOSED AGAIN AND 2ND ATTEMPT MADE TO FIRE IN ORDER TO TRY AND MINIMIZE THE ENTEROTOMY ( IE CONTAMINATION) AND TO PERMIT REMOVAL OF DEVICE. THIS TIME THERE WAS A CRUNCH ON FIRING. GUN OPENED. ANASTOMOSIS DISRUPTED AS EXPECTED. ANVIL DISENGAGED AND REMOVED PER ABDOMEN.
05-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER FIRED TWICE AND, AFTER THE SECOND FIRING, THE DEVICE BROKE. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
09-04-2018 | Injury | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE DISSECTION OF PANCREAS ON A LAPAROSCOPIC PANCREATECTOMY, THE DOCTOR TRIED TO FIRE STAPLER BUT HE WAS UNABLE TO SQUEEZE HANDLE. SOME BREAKING SOUND CAME FROM THE STAPLER AND THEN HANDLE WAS SQUEEZED BUT BLADE WAS NOT ADVANCING. THE INCISION WAS EXTENDED TO MORE THAN ONE INCH. THE LAPAROSCOPIC CASE WAS CONVERTED INTO OPEN PROCEDURE. THE PANCREAS WAS CUT USING A BLADE AND HAND SEWED TO COMPLETE THE CASE.
06-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A THORACOTOMY, THE SURGEON STARTED THE 1ST FIRING TO PULMONARY PARENCHYMA (CLOSE TO THORACIC DIAPHRAGM). THE SURGEON THOUGHT THE TISSUE WAS NOT THICK, HOWEVER THE FIRING STOPPED ON THE HALFWAY. THE SURGEON THEN PUSHED THE BUTTON A FEW TIMES, HOWEVER THE KNIFE DID NOT ADVANCE, THE KNIFE THEN RETURNED TO THE INITIAL POSITION WITH NO PROBLEM. THE SURGEON RE-STARTED THE FIRING OVER THE ORIGINAL STAPLE LINE WITH A NEW REINFORCE CARTRIDGE, AND THE FIRING WAS COMPLETED. THERE WAS NO PATIENT HARM.
09-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC WHIPPLE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS NOT ABLE TO FIRE AND STAPLE LINE WAS INCOMPLETE. THE SURGEON FIRED THE RELOAD THE SECOND TIME AND THE LIGHTS ON THE STAPLER ALL LOOKED LIKE IT WAS GOOD TO GO. WHEN THEY PUSHED THE BLUE FIRE BUTTON IT MADE A SOUND LIKE THE KNIFE BLADE BEGAN TO ADVANCE FOR A COUPLE MILLIMETERS BUT NO STAPLES WERE PUT DOWN AND THEN STOPPED ON ITS OWN. THE STATUS INDICATOR LIGHTS WENT FROM FLASHING GREEN AFTER THEY PUSHED THE GREEN SAFETY BUTTON TO SOLID GREEN. PUSHED THE GREEN SAFETY BUTTON AGAIN, BUT THAT DID NOT CORRECT THE PROBLEM. THEY BACKED THE RELOAD OFF AND GOT A NEW STAPLER. THE FIRST TIME THIS HAPPENED THEY GOT A NEW RELOAD AND CONTINUED WITHOUT INCIDENT. THE SECOND TIME IT HAPPENED THEY SWITCHED TO THE ROBOTIC STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER ABLE WAS UNABLE TO FIRE AND WAS UNABLE TO BE SQUEEZED BY THE SURGEON. THE SURGEON TOOK ANOTHER STAPLER TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-03-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO APPENDIX ON A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON OF THE DEVICE BUT THE STAPLE WAS NOT ABLE TO FIRE AND THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC WHIPPLE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS NOT ABLE TO FIRE AND STAPLE LINE WAS INCOMPLETE. THE SURGEON FIRED THE RELOAD THE SECOND TIME AND THE LIGHTS ON THE STAPLER ALL LOOKED LIKE IT WAS GOOD TO GO. WHEN THEY PUSHED THE BLUE FIRE BUTTON IT MADE A SOUND LIKE THE KNIFE BLADE BEGAN TO ADVANCE FOR A COUPLE MILLIMETERS BUT NO STAPLES WERE PUT DOWN AND THEN STOPPED ON ITS OWN. THE STATUS INDICATOR LIGHTS WENT FROM FLASHING GREEN AFTER THEY PUSHED THE GREEN SAFETY BUTTON TO SOLID GREEN. PUSHED THE GREEN SAFETY BUTTON AGAIN, BUT THAT DID NOT CORRECT THE PROBLEM. THEY BACKED THE RELOAD OFF AND GOT A NEW STAPLER. THE FIRST TIME THIS HAPPENED THEY GOT A NEW RELOAD AND CONTINUED WITHOUT INCIDENT. THE SECOND TIME IT HAPPENED THEY SWITCHED TO THE ROBOTIC STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | TRI-STAPLE 2.0 | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC WHIPPLE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS NOT ABLE TO FIRE AND STAPLE LINE WAS INCOMPLETE. THE SURGEON FIRED THE RELOAD THE SECOND TIME AND THE LIGHTS ON THE STAPLER ALL LOOKED LIKE IT WAS GOOD TO GO. WHEN THEY PUSHED THE BLUE FIRE BUTTON IT MADE A SOUND LIKE THE KNIFE BLADE BEGAN TO ADVANCE FOR A COUPLE MILLIMETERS BUT NO STAPLES WERE PUT DOWN AND THEN STOPPED ON ITS OWN. THE STATUS INDICATOR LIGHTS WENT FROM FLASHING GREEN AFTER THEY PUSHED THE GREEN SAFETY BUTTON TO SOLID GREEN. PUSHED THE GREEN SAFETY BUTTON AGAIN, BUT THAT DID NOT CORRECT THE PROBLEM. THEY BACKED THE RELOAD OFF AND GOT A NEW STAPLER. THE FIRST TIME THIS HAPPENED THEY GOT A NEW RELOAD AND CONTINUED WITHOUT INCIDENT. THE SECOND TIME IT HAPPENED THEY SWITCHED TO THE ROBOTIC STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON RE-CLAMPED THE BRONCHUS TWICE AND TRIED TO FIRE THE DEVICE FOR THE FIRST TIME HOWEVER, THE HANDLE WAS LOCKED AND COULD NOT BE FIRED. THE SURGEON DID NOT PUSHED THE GREEN BUTTON WHEN THEY RE-CLAMPED THE TISSUE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. PATIENT HAS NO INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/ESOPHAGUS PROCEDURE, WHEN THE DEVICE WAS FIRED, A STRANGE SOUND WAS HEARD. THE RESECTED TISSUE WAS CHECKED AND IT WAS SEEN THAT THE STAPLING WAS MALFORMED. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/ESOPHAGUS PROCEDURE, WHEN THE DEVICE WAS FIRED, A STRANGE SOUND WAS HEARD. THE RESECTED TISSUE WAS CHECKED AND IT WAS SEEN THAT THE STAPLING WAS MALFORMED. THERE WAS NO PATIENT INJURY.
07-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, PRIOR TO STAPLE FIRING, THE SURGEON COULD NOT COMPRESS THE HANDLE TO FIRE THE STAPLE LOAD. ANOTHER HANDLE WAS OPENED AND THE RELOAD FIRED. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | SIGNIA | Unintended System Motion | A051206 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTAL RESECTION, WHEN THE DEVICE APPROACHED THROUGH THE HOLE OF WOUND AND WENT INTO THE CAVITY SMOOTHLY, A CRACK WAS HEARD. THE SURGEON COULD OPEN AND CLOSE THE JAWS, THEN THE GREEN BUTTONS ILLUMINATED AND HE/SHE STARTED FIRING THE DEVICE. DURING THE FIRING, THE CARTRIDGE ARTICULATED ON THEIR OWN AND COULD NOT FIRE TO THE PLANNED STAPLE LINE. THEY CHANGED THE ADAPTER AND THE SURGEON MADE ANOTHER RESECTION DUE TO THE ISSUE. THERE WAS NO TISSUE DAMAGE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | SIGNIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTAL RESECTION, WHEN THE DEVICE APPROACHED THROUGH THE HOLE OF WOUND AND WENT INTO THE CAVITY SMOOTHLY, A CRACK WAS HEARD. THE SURGEON COULD OPEN AND CLOSE THE JAWS, THEN THE GREEN BUTTONS ILLUMINATED AND HE/SHE STARTED FIRING THE DEVICE. DURING THE FIRING, THE CARTRIDGE ARTICULATED ON THEIR OWN AND COULD NOT FIRE TO THE PLANNED STAPLE LINE. THEY CHANGED THE ADAPTER AND THE SURGEON MADE ANOTHER RESECTION DUE TO THE ISSUE. THERE WAS NO TISSUE DAMAGE. THERE WAS NO PATIENT INJURY.
23-04-2018 | Injury | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, THE DEVICE¿S FAILURE WAS NOTICED AFTER A LAPAROSCOPIC PROCEDURE. THIS RESULTED TO DAMAGED TISSUE. THE PATIENT HAD AN EXTENDED HOSPITALIZATION FOR 12 DAYS.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | SIGNIA | Unintended System Motion | A051206 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTAL RESECTION, WHEN THE DEVICE APPROACHED THROUGH THE HOLE OF WOUND AND WENT INTO THE CAVITY SMOOTHLY, A CRACK WAS HEARD. THE SURGEON COULD OPEN AND CLOSE THE JAWS, THEN THE GREEN BUTTONS ILLUMINATED AND HE/SHE STARTED FIRING THE DEVICE. DURING THE FIRING, THE CARTRIDGE ARTICULATED ON THEIR OWN AND COULD NOT FIRE TO THE PLANNED STAPLE LINE. THEY CHANGED THE ADAPTER AND THE SURGEON MADE ANOTHER RESECTION DUE TO THE ISSUE. THERE WAS NO TISSUE DAMAGE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | SIGNIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTAL RESECTION, WHEN THE DEVICE APPROACHED THROUGH THE HOLE OF WOUND AND WENT INTO THE CAVITY SMOOTHLY, A CRACK WAS HEARD. THE SURGEON COULD OPEN AND CLOSE THE JAWS, THEN THE GREEN BUTTONS ILLUMINATED AND HE/SHE STARTED FIRING THE DEVICE. DURING THE FIRING, THE CARTRIDGE ARTICULATED ON THEIR OWN AND COULD NOT FIRE TO THE PLANNED STAPLE LINE. THEY CHANGED THE ADAPTER AND THE SURGEON MADE ANOTHER RESECTION DUE TO THE ISSUE. THERE WAS NO TISSUE DAMAGE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | SIGNIA | Unintended System Motion | A051206 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTAL RESECTION, WHEN THE DEVICE APPROACHED THROUGH THE HOLE OF WOUND AND WENT INTO THE CAVITY SMOOTHLY, A CRACK WAS HEARD. THE SURGEON COULD OPEN AND CLOSE THE JAWS ,THEN THE GREEN BUTTONS ILLUMINATED AND HE/SHE STARTED FIRING THE DEVICE. DURING THE FIRING ,THE CARTRIDGE ARTICULATED ON THEIR OWN AND COULD NOT FIRE TO THE PLANNED STAPLE LINE. THEY CHANGED THE ADAPTER AND THE SURGEON MADE ANOTHER RESECTION DUE TO THE ISSUE. THERE WAS NO TISSUE DAMAGE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | SIGNIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTAL RESECTION, WHEN THE DEVICE APPROACHED THROUGH THE HOLE OF WOUND AND WENT INTO THE CAVITY SMOOTHLY, A CRACK WAS HEARD. THE SURGEON COULD OPEN AND CLOSE THE JAWS ,THEN THE GREEN BUTTONS ILLUMINATED AND HE/SHE STARTED FIRING THE DEVICE. DURING THE FIRING ,THE CARTRIDGE ARTICULATED ON THEIR OWN AND COULD NOT FIRE TO THE PLANNED STAPLE LINE. THEY CHANGED THE ADAPTER AND THE SURGEON MADE ANOTHER RESECTION DUE TO THE ISSUE. THERE WAS NO TISSUE DAMAGE. THERE WAS NO PATIENT INJURY.
07-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SIDE TO SIDE AND END TO END ANASTOMOSIS PROCEDURE, THEY USED THE SAME HANDLE AND ADAPTER THAT THEY BEEN USED ON THE OTHER CASES. BUT IN THIS CASE WHEN THE SURGEON CLAMP THE TISSUE, HOWEVER THE TISSUE SLIPPED OUT FROM THE JAW, FOR THE REASON THAT THE FORCE OF THE CLAMP WAS VERY WEAK, THEY REPLACED WITH A NEW CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE. DURING THE LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, COULD NOT FIRE ON THE 1ST FIRING WITH ECR45D WHEN SEVERING THE RECTUM. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
10-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE DISCONNECTION OF THE STOMACH TISSUE ON A LAPAROSCOPIC RADICAL GASTRECTOMY FOR GASTRIC CANCER, WHEN CLOSING THE STAPLES, THE JAWS COULD NOT BE COMPLETELY CLOSED. THERE WAS AN INCOMPLETE LINE IN THE NEAR END-REMOTE CENTRAL. THEY PRESSED THE INSURANCE BUTTON AND THE BLACK GRIP COULD NOT BE GRIPPED UNDER NORMAL FORCE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ILS 21MM, CURVED | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE MISFIRED. "THE COLOSTOMY WAS INCOMPLETE". IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED AND THERE WERE NO PATIENT CONSEQUENCES REPORTED. THE DEVICE WAS DISCARDED. THERE IS NO ADDITIONAL INFORMATION.
25-02-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ECHELON 60 ENDOPATH STAPLER | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE FARTHER. DURING THE LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, COULD NOT FIRE FARTHER AFTER FIRED 1/3. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
05-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT JEJUNUM RESECTION OF ILEAL CONDUIT IN CYSTECTOMY, THE SURGEON COULD NOT FIRE THE DEVICE WHEREIN THE HANDLE COULD NOT BE SQUEEZED EVEN AFTER PRESSING THE GREEN BUTTON. THEY REPLACED IT WITH A NEW DEVICE, AND THE FIRING WAS COMPLETED. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING USE IN AN ULTRA LOW ANTERIOR RESECTION, THE GUN WAS FIRED AND STAPLES APPEARED NOT TO FORM. RESOLVED BY DOING A HAND SEWN ANASTOMOSIS. NO REPORTED ADVERSE EVENT TO PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
05-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE. DURING THE SINGLE HOLE THORACOSCOPIC LUNG SURGERY, WHEN SEVERING PULMONARY SEGMENT, COULD NOT FIRE WITH ECR60B, AND HEARD ABNORMAL SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
06-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | IDRIVE ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A THORACOTOMY, THE SURGEON STARTED THE 1ST FIRING TO PULMONARY PARENCHYMA (CLOSE TO THORACIC DIAPHRAGM). THE SURGEON THOUGHT THE TISSUE WAS NOT THICK, HOWEVER THE FIRING STOPPED ON THE HALFWAY. THE SURGEON THEN PUSHED THE BUTTON A FEW TIMES, HOWEVER THE KNIFE DID NOT ADVANCE, THE KNIFE THEN RETURNED TO THE INITIAL POSITION WITH NO PROBLEM. THE SURGEON RE-STARTED THE FIRING OVER THE ORIGINAL STAPLE LINE WITH A NEW REINFORCE CARTRIDGE , AND THE FIRING WAS COMPLETED. THERE WAS NO PATIENT HARM.
05-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Bent |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE COMPLETING THE INCOMPLETE FISSURE IN A LOBECTOMY, THE RELOAD WAS BENT AT THE JUNCTION OF THE RELOAD AND ADAPTER. THEY ADDED A LITTLE FORCE TO TWIST FOR THEM TO UNLOAD THE HANDLE. THEY OPENED A SECOND RELOAD, AND ATTACHED IT TO THE HANDLE, AND SAME RESULT. THE BLACK PLASTIC WHERE THE RELOAD WAS ATTACHED TO AN ADAPTER BROKE OFF. THE STAPLER WAS UNABLE TO FIRE. THEY OPENED ANOTHER RELOAD AND ATTACHED IT TO A DIFFERENT HANDLE TO COMPLETE THE CASE. THEY ALSO NOTED THAT THE STATUS INDICATOR LIGHTS, FIVE WHITE LIGHTS, HANDLE INDICATOR, ADAPTER INDICATOR, AND RELOAD INDICATOR WERE SOLID. THEY WERE ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
05-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE COMPLETING THE INCOMPLETE FISSURE IN A LOBECTOMY, THE RELOAD WAS BENT AT THE JUNCTION OF THE RELOAD AND ADAPTER. THEY ADDED A LITTLE FORCE TO TWIST FOR THEM TO UNLOAD THE HANDLE. THEY OPENED A SECOND RELOAD, AND ATTACHED IT TO THE HANDLE, AND SAME RESULT. THE BLACK PLASTIC WHERE THE RELOAD WAS ATTACHED TO AN ADAPTER BROKE OFF. THE STAPLER WAS UNABLE TO FIRE. THEY OPENED ANOTHER RELOAD AND ATTACHED IT TO A DIFFERENT HANDLE TO COMPLETE THE CASE. THEY ALSO NOTED THAT THE STATUS INDICATOR LIGHTS, FIVE WHITE LIGHTS, HANDLE INDICATOR, ADAPTER INDICATOR, AND RELOAD INDICATOR WERE SOLID. THEY WERE ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
05-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE COMPLETING THE INCOMPLETE FISSURE IN A LOBECTOMY, THE RELOAD WAS BENT AT THE JUNCTION OF THE RELOAD AND ADAPTER. THEY ADDED A LITTLE FORCE TO TWIST FOR THEM TO UNLOAD THE HANDLE. THEY OPENED A SECOND RELOAD, AND ATTACHED IT TO THE HANDLE, AND SAME RESULT. THE BLACK PLASTIC WHERE THE RELOAD WAS ATTACHED TO AN ADAPTER BROKE OFF. THE STAPLER WAS UNABLE TO FIRE. THEY OPENED ANOTHER RELOAD AND ATTACHED IT TO A DIFFERENT HANDLE TO COMPLETE THE CASE. THEY ALSO NOTED THAT THE STATUS INDICATOR LIGHTS, FIVE WHITE LIGHTS, HANDLE INDICATOR, ADAPTER INDICATOR, AND RELOAD INDICATOR WERE SOLID. THEY WERE ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
05-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE COMPLETING THE INCOMPLETE FISSURE IN A LOBECTOMY, THE RELOAD WAS BENT AT THE JUNCTION OF THE RELOAD AND ADAPTER. THEY ADDED A LITTLE FORCE TO TWIST FOR THEM TO UNLOAD THE HANDLE. THEY OPENED A SECOND RELOAD, AND ATTACHED IT TO THE HANDLE, AND SAME RESULT. THE BLACK PLASTIC WHERE THE RELOAD WAS ATTACHED TO AN ADAPTER BROKE OFF. THE STAPLER WAS UNABLE TO FIRE. THEY OPENED ANOTHER RELOAD AND ATTACHED IT TO A DIFFERENT HANDLE TO COMPLETE THE CASE. THEY ALSO NOTED THAT THE STATUS INDICATOR LIGHTS, FIVE WHITE LIGHTS, HANDLE INDICATOR, ADAPTER INDICATOR, AND RELOAD INDICATOR WERE SOLID. THEY WERE ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
05-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE COMPLETING THE INCOMPLETE FISSURE IN A LOBECTOMY, THE RELOAD WAS BENT AT THE JUNCTION OF THE RELOAD AND ADAPTER. THEY ADDED A LITTLE FORCE TO TWIST FOR THEM TO UNLOAD THE HANDLE. THEY OPENED A SECOND RELOAD, AND ATTACHED IT TO THE HANDLE, AND SAME RESULT. THE BLACK PLASTIC WHERE THE RELOAD WAS ATTACHED TO AN ADAPTER BROKE OFF. THE STAPLER WAS UNABLE TO FIRE. THEY OPENED ANOTHER RELOAD AND ATTACHED IT TO A DIFFERENT HANDLE TO COMPLETE THE CASE. THEY ALSO NOTED THAT THE STATUS INDICATOR LIGHTS, FIVE WHITE LIGHTS, HANDLE INDICATOR, ADAPTER INDICATOR, AND RELOAD INDICATOR WERE SOLID. THEY WERE ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
27-03-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER ABLE WAS UNABLE TO FIRE AND WAS UNABLE TO BE SQUEEZED BY THE SURGEON. THE SURGEON TOOK ANOTHER STAPLER TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC DUODENAL SWITCH, WHILE THE DEVICE WAS BEING APPLIED TO STOMACH, THE DEVICE WAS UNABLE TO FIRE. THERE WAS NO PATIENT INJURY.
14-03-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THEY TRIED TO STAPLE THE BASE OF THE APPENDIX ON A LAPAROSCOPIC APPENDECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE AND THE STAPLER WAS UNABLE TO FIRE. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS UNABLE TO FIRE AND THE RELOAD WOULD NOT CLOSE. THERE WAS NO PATIENT INVOLVEMENT.
20-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DOCTOR WAS USING THE DEVICE TO STAPLE THE URETER AND RENAL ARTERY. AFTER SUCCESSFULLY STAPLING THE URETER THE STAPLER WAS HANDED TO THE SCRUB TECH WITH THE JAWS CLOSED. THE TECH ATTEMPTED TO OPEN THE JAWS WITH THE GRAY BUTTON. AFTER MULTIPLE ATTEMPTS TO OPEN THE JAW WITH NO SUCCESS A SECOND STAPLER WAS TRIED AND SUCCESSFULLY USED TO STAPLE THE RENAL VEIN. NO PATIENT CONSEQUENCES REPORTED.
20-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DOCTOR WAS USING THE DEVICE TO STAPLE THE URETER AND RENAL ARTERY. AFTER SUCCESSFULLY STAPLING THE URETER THE STAPLER WAS HANDED TO THE SCRUB TECH WITH THE JAWS CLOSED. THE TECH ATTEMPTED TO OPEN THE JAWS WITH THE GRAY BUTTON. AFTER MULTIPLE ATTEMPTS TO OPEN THE JAW WITH NO SUCCESS A SECOND STAPLER WAS TRIED AND SUCCESSFULLY USED TO STAPLE THE RENAL VEIN. NO PATIENT CONSEQUENCES REPORTED.
20-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DOCTOR WAS USING THE DEVICE TO STAPLE THE URETER AND RENAL ARTERY. THE STAPLER WAS TRIED AND SUCCESSFULLY USED TO STAPLE THE RENAL VEIN, HOWEVER THE DEVICE ONLY STAPLED HALF OF THE WAY. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCES REPORTED.
01-02-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE MISFIRED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION.
16-03-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | EEA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SIGMOIDECTOMY DIVERTICULITIS PROCEDURE, THE STAPLE MISFIRED. IT WAS REPORTED THAT AFTER RESECTION THE STAPLER WAS PLACED IN THE STOMACH. THE STAPLER PASSED IN THE RECTO STOP. THE SHARP POINT WAS DEPLOYED EASILY, AND IT WAS CONNECTED IN THE ANVIL. THERE WAS NO PATIENT HARM.
05-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | IDRIVE ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE COMPLETING THE INCOMPLETE FISSURE IN A LOBECTOMY, THE RELOAD WAS BENT AT THE JUNCTION OF THE RELOAD AND ADAPTER. THEY ADDED A LITTLE FORCE TO TWIST FOR THEM TO UNLOAD THE HANDLE. THE OPENED A SECOND RELOAD, AND ATTACHED IT TO THE HANDLE, AND SAME RESULT. THE BLACK PLASTIC WHERE THE RELOAD WAS ATTACHED TO AN ADAPTER BROKE OFF. THE STAPLER WAS UNABLE TO FIRE. THEY OPENED ANOTHER RELOAD AND ATTACHED IT TO A DIFFERENT HANDLE TO COMPLETE THE CASE. THEY ALSO NOTED THAT THE STATUS INDICATOR LIGHTS, FIVE WHITE LIGHTS, HANDLE INDICATOR, ADAPTER INDICATOR, AND RELOAD INDICATOR WERE SOLID. THEY WERE ABLE TO ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE DISCONNECTION OF THE STOMACH TISSUE ON A LAPAROSCOPIC RADICAL GASTRECTOMY FOR CANCER, THE STAPLES WERE CLOSED BUT THE JAWS WERE NOT COMPLETELY CLOSED. THERE WAS A MAIN LINE INCOMPLETE IN THE NEAR END-REMOTE CENTRAL. THEY PRESSED THE INSURANCE BUTTON AND THE BLACK GRIP COULD NOT BE GRIPPED UNDER NORMAL FORCE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON RE-CLAMPED THE BRONCHUS TWICE AND TRIED TO FIRE THE DEVICE FOR THE FIRST TIME HOWEVER, THE HANDLE WAS LOCKED AND COULD NOT BE FIRED. THE SURGEON DID NOT PUSHED THE GREEN BUTTON WHEN THEY RE-CLAMPED THE TISSUE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. PATIENT HAS NO INJURY.
09-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | IDRIVE ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC WHIPPLE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS NOT ABLE TO FIRE AND STAPLE LINE WAS INCOMPLETE. THE SURGEON FIRED THE RELOAD THE SECOND TIME AND THE LIGHTS ON THE STAPLER ALL LOOKED LIKE IT WAS GOOD TO GO. WHEN THEY PUSHED THE BLUE FIRE BUTTON IT MADE A SOUND LIKE THE KNIFE BLADE BEGAN TO ADVANCE FOR A COUPLE MILLIMETERS BUT NO STAPLES WERE PUT DOWN AND THEN STOPPED ON ITS OWN. THE STATUS INDICATOR LIGHTS WENT FROM FLASHING GREEN AFTER THEY PUSHED THE GREEN SAFETY BUTTON TO SOLID GREEN. PUSHED THE GREEN SAFETY BUTTON AGAIN, BUT THAT DID NOT CORRECT THE PROBLEM. THEY BACKED THE RELOAD OFF AND GOT A NEW STAPLER. THE FIRST TIME THIS HAPPENED THEY GOT A NEW RELOAD AND CONTINUED WITHOUT INCIDENT. THE SECOND TIME IT HAPPENED THEY SWITCHED TO THE ROBOTIC STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | TA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE PROCEDURE, DURING PARAESOPHAGEAL HERNIA REPAIR FOR NISSEN, THE STAPLER WAS FIRED PROPERLY BUT THE STAPLES DID NOT DEPLOY. ANOTHER STAPLER WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, PRIOR TO STAPLE FIRING, THE SURGEON COULD NOT COMPRESS THE HANDLE TO FIRE THE STAPLE LOAD. ANOTHER HANDLE WAS OPENED AND THE RELOAD FIRED. THERE WAS NO PATIENT INJURY.
05-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | GST BLUE RELOAD, 60MM, 6 ROW | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE CARTRIDGE BROKE APART AND THE CARTRIDGE PAN WAS LEFT INSIDE THE JAWS WHEN THE CARTRIDGE WAS REMOVED FROM THE DEVICE AFTER THE FIRING. THE LEFT PIECE WAS REMOVED FROM THE JAWS BY HAND, AND THE SAME DEVICE LOADING ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
06-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A THORACOTOMY, THE SURGEON STARTED THE 1ST FIRING TO PULMONARY PARENCHYMA (CLOSE TO THORACIC DIAPHRAGM). THE SURGEON THOUGHT THE TISSUE WAS NOT THICK, HOWEVER THE FIRING STOPPED ON THE HALFWAY. THE SURGEON THEN PUSHED THE BUTTON A FEW TIMES, HOWEVER THE KNIFE DID NOT ADVANCE, THE KNIFE THEN RETURNED TO THE INITIAL POSITION WITH NO PROBLEM. THE SURGEON RE-STARTED THE FIRING OVER THE ORIGINAL STAPLE LINE WITH A NEW REINFORCE CARTRIDGE, AND THE FIRING WAS COMPLETED. THERE WAS NO PATIENT HARM.
08-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PEDIATRIC LAPAROSCOPIC APPENDECTOMY, THE STAPLER DID NOT FIRE. WHILE REMOVING THE STAPLER FROM THE ABDOMEN TO BE CHECKED, THE RELOAD CAME LOOSE FROM THE HANDLE. A RETRIEVAL BAG AND HAND INSTRUMENT WAS USED TO SECURE THE RELOAD AND REMOVE IT FROM THE ABDOMEN. REPLACED BY A NEW DEVICE TO COMPLETE THE CASE. NO PATIENT INJURY.
08-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PEDIATRIC LAPAROSCOPIC APPENDECTOMY, THE STAPLER DID NOT FIRE. WHILE REMOVING THE STAPLER FROM THE ABDOMEN TO BE CHECKED, THE RELOAD CAME LOOSE FROM THE HANDLE. A RETRIEVAL BAG AND HAND INSTRUMENT WAS USED TO SECURE THE RELOAD AND REMOVE IT FROM THE ABDOMEN. REPLACED BY A NEW DEVICE TO COMPLETE THE CASE. NO PATIENT INJURY.
05-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE COMPLETING THE INCOMPLETE FISSURE IN A LOBECTOMY, THE RELOAD WAS BENT AT THE JUNCTION OF THE RELOAD AND ADAPTER. THEY ADDED A LITTLE FORCE TO TWIST FOR THEM TO UNLOAD THE HANDLE. THE OPENED A SECOND RELOAD, AND ATTACHED IT TO THE HANDLE, AND SAME RESULT. THE BLACK PLASTIC WHERE THE RELOAD WAS ATTACHED TO AN ADAPTER BROKE OFF. THE STAPLER WAS UNABLE TO FIRE. THEY OPENED ANOTHER RELOAD AND ATTACHED IT TO A DIFFERENT HANDLE TO COMPLETE THE CASE. THEY ALSO NOTED THAT THE STATUS INDICATOR LIGHTS, FIVE WHITE LIGHTS, HANDLE INDICATOR, ADAPTER INDICATOR, AND RELOAD INDICATOR WERE SOLID. THEY WERE ABLE TO ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
05-04-2018 | Malfunction | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE COMPLETING THE INCOMPLETE FISSURE IN A LOBECTOMY, THE RELOAD WAS BENT AT THE JUNCTION OF THE RELOAD AND ADAPTER. THEY ADDED A LITTLE FORCE TO TWIST FOR THEM TO UNLOAD THE HANDLE. THE OPENED A SECOND RELOAD, AND ATTACHED IT TO THE HANDLE, AND SAME RESULT. THE BLACK PLASTIC WHERE THE RELOAD WAS ATTACHED TO AN ADAPTER BROKE OFF. THE STAPLER WAS UNABLE TO FIRE. THEY OPENED ANOTHER RELOAD AND ATTACHED IT TO A DIFFERENT HANDLE TO COMPLETE THE CASE. THEY ALSO NOTED THAT THE STATUS INDICATOR LIGHTS, FIVE WHITE LIGHTS, HANDLE INDICATOR, ADAPTER INDICATOR, AND RELOAD INDICATOR WERE SOLID. THEY WERE ABLE TO ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INJURY.
08-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER MALFUNCTIONED. A SECOND LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE IS NO OTHER INFORMATION KNOWN ABOUT THE ISSUE AND NO ADDITIONAL CONTACT THAT WOULD KNOW MORE ABOUT THE ISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
09-04-2018 | Malfunction | davis & geck caribe | 23-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC DUODENAL SWITCH, WHILE THE DEVICE WAS BEING APPLIED TO STOMACH, THE DEVICE'S HANDLE WAS NOT ABLE TO SQUEEZED AND COULD NOT FIRE . THE OPENED A NEW STAPLE LOAD TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. FROM (B)(4) - ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC DUODENAL SWITCH, WHILE THE DEVICE WAS BEING APPLIED TO STOMACH, THE DEVICE WAS UNABLE TO FIRE. THE OPENED A NEW STAPLE LOAD TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2019 | Malfunction | us surgical puerto rico | 23-04-2018 | MULTIFIRE ENDO GIA | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS UNABLE TO FIRE AND THE RELOAD WOULD NOT CLOSE. THERE WAS NO PATIENT INVOLVEMENT.
09-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ILA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, FOR THE THIRD FIRING OF COLECTOMY PROCEDURE, THEY CONNECTED THE CARTRIDGE TO INSTRUMENT; HOWEVER, A WHITE PART OF THE CARTRIDGE WAS UNSTABLE. THEY CONSIDERED THE CARTRIDGE HAD NO PROBLEM AND CONNECTED IT TO THE INSTRUMENT. AT FUNCTIONAL SIDE TO SIDE ANASTOMOSIS, THE SURGEON INSERTED ANVIL FORK AND THE CARTRIDGE FROM SMALL OPENING AT INTESTINAL CANAL, AND THEN CLOSED THE INSTRUMENT. HE TRIED TO FIRE THE DEVICE, HOWEVER HE COULD NOT. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE STATUS OF THE PATIENT WAS REPORTED NO INJURY.
09-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ILA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, FOR THE THIRD FIRING OF COLECTOMY PROCEDURE, THEY CONNECTED THE CARTRIDGE TO INSTRUMENT; HOWEVER, A WHITE PART OF THE CARTRIDGE WAS UNSTABLE. THEY CONSIDERED THE CARTRIDGE HAD NO PROBLEM AND CONNECTED IT TO THE INSTRUMENT. AT FUNCTIONAL SIDE TO SIDE ANASTOMOSIS, THE SURGEON INSERTED ANVIL FORK AND THE CARTRIDGE FROM SMALL OPENING AT INTESTINAL CANAL, AND THEN CLOSED THE INSTRUMENT. HE TRIED TO FIRE THE DEVICE, HOWEVER HE COULD NOT. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE STATUS OF THE PATIENT WAS REPORTED AS NO PROBLEM.
04-04-2018 | Injury | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Patient-Device Incompatibility | A0101 | Reaction | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC COLON SURGERY, FIVE DAYS BEFORE THE PROCEDURE, THE SURGEON PERFORMED FUNCTIONAL END-TO-END ANASTOMOSIS USING OTHER DEVICE. THE SURGEON ANASTOMOSED THE LARGE INTESTINE AND SMALL INTESTINE. THE PATIENT WAS NOT IN BAD CONDITION SPECIALLY AT THAT TIME, EVEN THOUGH STEROID WAS BEING USED. DURING THE SURGERY, THE TISSUE AROUND END POINT OF SIDE TO SIDE ANASTOMOSIS BECAME RED AND BLACK LIKE NECROTIC TISSUE. FIVE DAYS LATER, AS THE ANASTOMOSED SITE WAS STILL IN BAD CONDITION, THE SURGEON PERFORMED OPEN RE-SURGERY WITH HANDLE AND POWER RELOAD. THEN THE SURGEON PERFORMED THE SECOND RE-SURGERY AND SUTURE THE TISSUE MANUALLY. THE SURGEON SAID TITANIUM MIGHT HAVE CAUSED THESE EVENTS AND REQUESTED US TO PROVIDE THE INFORMATION OF TAKING ACTION AND THE CHECKING OF TITANIUM ALLERGY.
04-04-2018 | Injury | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Patient-Device Incompatibility | A0101 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC COLON SURGERY, FIVE DAYS BEFORE THE PROCEDURE, THE SURGEON PERFORMED FUNCTIONAL END-TO-END ANASTOMOSIS USING OTHER DEVICE. THE SURGEON ANASTOMOSED THE LARGE INTESTINE AND SMALL INTESTINE. THE PATIENT WAS NOT IN BAD CONDITION SPECIALLY AT THAT TIME, EVEN THOUGH STEROID WAS BEING USED. DURING THE SURGERY, THE TISSUE AROUND END POINT OF SIDE TO SIDE ANASTOMOSIS BECAME RED AND BLACK LIKE NECROTIC TISSUE. FIVE DAYS LATER, AS THE ANASTOMOSED SITE WAS STILL IN BAD CONDITION, THE SURGEON PERFORMED OPEN RE-SURGERY WITH HANDLE AND POWER RELOAD. THEN THE SURGEON PERFORMED THE SECOND RE-SURGERY AND SUTURE THE TISSUE MANUALLY. THE SURGEON SAID TITANIUM MIGHT HAVE CAUSED THESE EVENTS AND REQUESTED US TO PROVIDE THE INFORMATION OF TAKING ACTION AND THE CHECKING OF TITANIUM ALLERGY.
04-04-2018 | Injury | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Reaction | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC COLON SURGERY, FIVE DAYS BEFORE THE PROCEDURE, THE SURGEON PERFORMED FUNCTIONAL END-TO-END ANASTOMOSIS USING OTHER DEVICE. THE SURGEON ANASTOMOSED THE LARGE INTESTINE AND SMALL INTESTINE. THE PATIENT WAS NOT IN BAD CONDITION SPECIALLY AT THAT TIME, EVEN THOUGH STEROID WAS BEING USED. DURING THE SURGERY, THE TISSUE AROUND END POINT OF SIDE TO SIDE ANASTOMOSIS BECAME RED AND BLACK LIKE NECROTIC TISSUE. FIVE DAYS LATER, AS THE ANASTOMOSED SITE WAS STILL IN BAD CONDITION, THE SURGEON PERFORMED OPEN RE-SURGERY WITH HANDLE AND POWER RELOAD. THEN THE SURGEON PERFORMED THE SECOND RE-SURGERY AND SUTURE THE TISSUE MANUALLY. THE SURGEON SAID TITANIUM MIGHT HAVE CAUSED THESE EVENTS AND REQUESTED US TO PROVIDE THE INFORMATION OF TAKING ACTION AND THE CHECKING OF TITANIUM ALLERGY.
04-04-2018 | Injury | covidien lp llc north haven | 23-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC COLON SURGERY, FIVE DAYS BEFORE THE PROCEDURE, THE SURGEON PERFORMED FUNCTIONAL END-TO-END ANASTOMOSIS USING OTHER DEVICE. THE SURGEON ANASTOMOSED THE LARGE INTESTINE AND SMALL INTESTINE. THE PATIENT WAS NOT IN BAD CONDITION SPECIALLY AT THAT TIME, EVEN THOUGH STEROID WAS BEING USED. DURING THE SURGERY, THE TISSUE AROUND END POINT OF SIDE TO SIDE ANASTOMOSIS BECAME RED AND BLACK LIKE NECROTIC TISSUE. FIVE DAYS LATER, AS THE ANASTOMOSED SITE WAS STILL IN BAD CONDITION, THE SURGEON PERFORMED OPEN RE-SURGERY WITH HANDLE AND POWER RELOAD. THEN THE SURGEON PERFORMED THE SECOND RE-SURGERY AND SUTURE THE TISSUE MANUALLY. THE SURGEON SAID TITANIUM MIGHT HAVE CAUSED THESE EVENTS AND REQUESTED US TO PROVIDE THE INFORMATION OF TAKING ACTION AND THE CHECKING OF TITANIUM ALLERGY.
27-03-2018 | Malfunction | ethicon endo-surgery | 23-04-2018 | GST BLACK RELOAD, 60MM, 6 ROW | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE VERTICAL GASTRECTOMY PROCEDURE, THE FIRST SHOT WAS PERFORMED ON THE STAPLER WITH THE BLACK LOAD. IT WAS DETACHED FROM THE METAL STRUCTURE THAT PROVIDES SUPPORT AND "DISASSEMBLED" INTO THE ABDOMINAL CAVITY OF THE PATIENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
04-04-2018 | Injury | us surgical puerto rico | 23-04-2018 | ENDOGIA | Patient-Device Incompatibility | A0101 | Reaction | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC COLON SURGERY, FIVE DAYS BEFORE THE PROCEDURE, THE SURGEON PERFORMED FUNCTIONAL END-TO-END ANASTOMOSIS USING OTHER DEVICE. THE SURGEON ANASTOMOSED THE LARGE INTESTINE AND SMALL INTESTINE. THE PATIENT WAS NOT IN BAD CONDITION SPECIALLY AT THAT TIME, EVEN THOUGH STEROID WAS BEING USED. DURING THE SURGERY, THE TISSUE AROUND END POINT OF SIDE TO SIDE ANASTOMOSIS BECAME RED AND BLACK LIKE NECROTIC TISSUE. FIVE DAYS LATER, AS THE ANASTOMOSED SITE WAS STILL IN BAD CONDITION, THE SURGEON PERFORMED OPEN RE-SURGERY WITH HANDLE AND POWER RELOAD. THEN THE SURGEON PERFORMED THE SECOND RE-SURGERY, AND SUTURE THE TISSUE MANUALLY. THE SURGEON SAID TITANIUM MIGHT HAVE CAUSED THESE EVENTS AND REQUESTED US TO PROVIDE THE INFORMATION OF TAKING ACTION AND THE CHECKING OF TITANIUM ALLERGY.
04-04-2018 | Injury | us surgical puerto rico | 23-04-2018 | ENDOGIA | Patient-Device Incompatibility | A0101 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC COLON SURGERY, FIVE DAYS BEFORE THE PROCEDURE, THE SURGEON PERFORMED FUNCTIONAL END-TO-END ANASTOMOSIS USING OTHER DEVICE. THE SURGEON ANASTOMOSED THE LARGE INTESTINE AND SMALL INTESTINE. THE PATIENT WAS NOT IN BAD CONDITION SPECIALLY AT THAT TIME, EVEN THOUGH STEROID WAS BEING USED. DURING THE SURGERY, THE TISSUE AROUND END POINT OF SIDE TO SIDE ANASTOMOSIS BECAME RED AND BLACK LIKE NECROTIC TISSUE. FIVE DAYS LATER, AS THE ANASTOMOSED SITE WAS STILL IN BAD CONDITION, THE SURGEON PERFORMED OPEN RE-SURGERY WITH HANDLE AND POWER RELOAD. THEN THE SURGEON PERFORMED THE SECOND RE-SURGERY, AND SUTURE THE TISSUE MANUALLY. THE SURGEON SAID TITANIUM MIGHT HAVE CAUSED THESE EVENTS AND REQUESTED US TO PROVIDE THE INFORMATION OF TAKING ACTION AND THE CHECKING OF TITANIUM ALLERGY.
04-04-2018 | Injury | us surgical puerto rico | 23-04-2018 | ENDOGIA | Adverse Event Without Identified Device or Use Problem | A24 | Reaction | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC COLON SURGERY, FIVE DAYS BEFORE THE PROCEDURE, THE SURGEON PERFORMED FUNCTIONAL END-TO-END ANASTOMOSIS USING OTHER DEVICE. THE SURGEON ANASTOMOSED THE LARGE INTESTINE AND SMALL INTESTINE. THE PATIENT WAS NOT IN BAD CONDITION SPECIALLY AT THAT TIME, EVEN THOUGH STEROID WAS BEING USED. DURING THE SURGERY, THE TISSUE AROUND END POINT OF SIDE TO SIDE ANASTOMOSIS BECAME RED AND BLACK LIKE NECROTIC TISSUE. FIVE DAYS LATER, AS THE ANASTOMOSED SITE WAS STILL IN BAD CONDITION, THE SURGEON PERFORMED OPEN RE-SURGERY WITH HANDLE AND POWER RELOAD. THEN THE SURGEON PERFORMED THE SECOND RE-SURGERY, AND SUTURE THE TISSUE MANUALLY. THE SURGEON SAID TITANIUM MIGHT HAVE CAUSED THESE EVENTS AND REQUESTED US TO PROVIDE THE INFORMATION OF TAKING ACTION AND THE CHECKING OF TITANIUM ALLERGY.
04-04-2018 | Injury | us surgical puerto rico | 23-04-2018 | ENDOGIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC COLON SURGERY, FIVE DAYS BEFORE THE PROCEDURE, THE SURGEON PERFORMED FUNCTIONAL END-TO-END ANASTOMOSIS USING OTHER DEVICE. THE SURGEON ANASTOMOSED THE LARGE INTESTINE AND SMALL INTESTINE. THE PATIENT WAS NOT IN BAD CONDITION SPECIALLY AT THAT TIME, EVEN THOUGH STEROID WAS BEING USED. DURING THE SURGERY, THE TISSUE AROUND END POINT OF SIDE TO SIDE ANASTOMOSIS BECAME RED AND BLACK LIKE NECROTIC TISSUE. FIVE DAYS LATER, AS THE ANASTOMOSED SITE WAS STILL IN BAD CONDITION, THE SURGEON PERFORMED OPEN RE-SURGERY WITH HANDLE AND POWER RELOAD. THEN THE SURGEON PERFORMED THE SECOND RE-SURGERY, AND SUTURE THE TISSUE MANUALLY. THE SURGEON SAID TITANIUM MIGHT HAVE CAUSED THESE EVENTS AND REQUESTED US TO PROVIDE THE INFORMATION OF TAKING ACTION AND THE CHECKING OF TITANIUM ALLERGY.
01-02-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | TA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY WITH A SMALL BOWEL RESECTION, THE STAPLER DID NOT STAPLE WHEN THE SURGEON WAS USING THE DEVICE TO CLOSE THE ANASTOMOSIS. THEY CLOSED THE ANASTOMOSIS WITH HAND SUTURING. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LOBECTOMY, THE STAPLER WAS ABLE TO FIRE THE FIRST TIME BUT NOT ANYMORE AFTER THE SECOND FIRING. SOME PARTS OF THE STAPLE FORMATION WERE ALSO POORLY FORMED. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 23-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LOBECTOMY, THE STAPLER WAS ABLE TO FIRE THE FIRST TIME BUT NOT ANYMORE AFTER THE SECOND FIRING. SOME PARTS OF THE STAPLE FORMATION WERE ALSO POORLY FORMED. THERE WAS NO PATIENT INJURY.
01-03-2018 | Malfunction | ethicon endo-surgery | 24-04-2018 | PROXIMATE*PPH PROCEDURE SET | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY, THE SURGEON COMMENTED THAT THERE WAS NO 'CLUTCH' SOUND AFTER FIRING. IT STAPLED BUT DIDN'T CUT. HE HAD TO DO CUTTING AND SUTURE IT BACK MANUALLY. HE CHECKED AND THERE WAS NO WASHER IN THE ANVIL HEAD AND IN THE STAPLER. THERE WAS LONGER OT TIME. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-03-2018 | Malfunction | ethicon endo-surgery | 24-04-2018 | PROXIMATE*PPH PROCEDURE SET | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY, THE SURGEON COMMENTED THAT THERE WAS NO 'CLUTCH' SOUND AFTER FIRING. IT STAPLED BUT DIDN'T CUT. HE HAD TO DO CUTTING AND SUTURE IT BACK MANUALLY. HE CHECKED AND THERE WAS NO WASHER IN THE ANVIL HEAD AND IN THE STAPLER. THERE WAS LONGER OT TIME. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
30-03-2018 | Malfunction | ethicon endo-surgery | 24-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER GOT STUCK AND WITH APPLYING FORCE THE HANDLE BROKE. A SECOND STAPLER WAS USED TO FINISH THE ANASTOMOSIS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
30-03-2018 | Malfunction | ethicon endo-surgery | 24-04-2018 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER GOT STUCK AND WITH APPLYING FORCE THE HANDLE BROKE. A SECOND STAPLER WAS USED TO FINISH THE ANASTOMOSIS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
26-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ECHELON FLEX 45 STANDARD | Device Operates Differently Than Expected |  | No Code Available | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, WHEN THE DOCTOR PLACED THE STAPLER WITH THE RECHARGE ON THE PATIENT'S LUNG AND MADE THE SHOT, THE STAPLING DID NOT OCCUR. THE DEVICE LOCKED. ANOTHER STAPLER WAS ORDERED, THE SAME THING HAPPENED, AND THE SURGERY WAS CONVERTED TO OPEN.
26-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ECHELON FLEX 45 STANDARD | Mechanical Jam | A0506 | No Code Available | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, WHEN THE DOCTOR PLACED THE STAPLER WITH THE RECHARGE ON THE PATIENT'S LUNG AND MADE THE SHOT, THE STAPLING DID NOT OCCUR. THE DEVICE LOCKED. ANOTHER STAPLER WAS ORDERED, THE SAME THING HAPPENED, AND THE SURGERY WAS CONVERTED TO OPEN.
26-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ECHELON FLEX 45 STANDARD | Device Operates Differently Than Expected |  | No Code Available | Event Description: DURING THE SURGERY, WHEN THE DOCTOR PLACED THE STAPLER WITH THE RECHARGE ON THE PATIENT'S LUNG AND MADE THE SHOT, THE STAPLING DID NOT OCCUR. THE DEVICE LOCKED. ANOTHER STAPLER WAS ORDERED, THE SAME THING HAPPENED, AND THE SURGERY WAS CONVERTED TO OPEN.
26-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ECHELON FLEX 45 STANDARD | Mechanical Jam | A0506 | No Code Available | Event Description: DURING THE SURGERY, WHEN THE DOCTOR PLACED THE STAPLER WITH THE RECHARGE ON THE PATIENT'S LUNG AND MADE THE SHOT, THE STAPLING DID NOT OCCUR. THE DEVICE LOCKED. ANOTHER STAPLER WAS ORDERED, THE SAME THING HAPPENED, AND THE SURGERY WAS CONVERTED TO OPEN.
01-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE THE PATIENT MAY HAVE HAD A BLEED OR LEAK. THE DOCTOR STATED THAT THERE WAS MORE BLEEDING WHEN STAPLING WITH THE DEVICE. THE DOCTOR IS STILL TREATING THE PATIENT. THE DOCTOR DID NOT OFFER ANY ADDITIONAL INFORMATION. NO OTHER INFORMATION IS KNOWN AT THIS TIME.
01-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE THE PATIENT MAY HAVE HAD A BLEED OR LEAK. THE DOCTOR STATED THAT THERE WAS MORE BLEEDING WHEN STAPLING WITH THE DEVICE. THE DOCTOR IS STILL TREATING THE PATIENT. THE DOCTOR DID NOT OFFER ANY ADDITIONAL INFORMATION. NO OTHER INFORMATION IS KNOWN AT THIS TIME.
01-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE THE PATIENT MAY HAVE HAD A BLEED OR LEAK. THE DOCTOR STATED THAT THERE WAS MORE BLEEDING WHEN STAPLING WITH THE DEVICE. THE DOCTOR IS STILL TREATING THE PATIENT. THE DOCTOR DID NOT OFFER ANY ADDITIONAL INFORMATION. NO OTHER INFORMATION IS KNOWN AT THIS TIME.
01-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE THE PATIENT MAY HAVE HAD A BLEED OR LEAK. THE DOCTOR STATED THAT THERE WAS MORE BLEEDING WHEN STAPLING WITH THE DEVICE. THE DOCTOR IS STILL TREATING THE PATIENT. THE DOCTOR DID NOT OFFER ANY ADDITIONAL INFORMATION. NO OTHER INFORMATION IS KNOWN AT THIS TIME.
01-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE THE PATIENT MAY HAVE HAD A BLEED OR LEAK. THE DOCTOR STATED THAT THERE WAS MORE BLEEDING WHEN STAPLING WITH THE DEVICE. THE DOCTOR IS STILL TREATING THE PATIENT. THE DOCTOR DID NOT OFFER ANY ADDITIONAL INFORMATION. NO OTHER INFORMATION IS KNOWN AT THIS TIME.
01-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Not Applicable | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE THE PATIENT MAY HAVE HAD A BLEED OR LEAK. THE DOCTOR STATED THAT THERE WAS MORE BLEEDING WHEN STAPLING WITH THE DEVICE. THE DOCTOR IS STILL TREATING THE PATIENT. THE DOCTOR DID NOT OFFER ANY ADDITIONAL INFORMATION. NO OTHER INFORMATION IS KNOWN AT THIS TIME.
01-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE THE PATIENT MAY HAVE HAD A BLEED OR LEAK. THE DOCTOR STATED THAT THERE WAS MORE BLEEDING WHEN STAPLING WITH THE DEVICE. THE DOCTOR IS STILL TREATING THE PATIENT. THE DOCTOR DID NOT OFFER ANY ADDITIONAL INFORMATION. NO OTHER INFORMATION IS KNOWN AT THIS TIME.
24-04-2018 | Malfunction | davol inc sub c.r bard | 24-04-2018 | OPTIFIX 30 CT STRAIGHT 39CM | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE OPTIFIX FIXATION DEVICE STOPPED FIRING AFTER THE FIRST FASTENER WAS DEPLOYED. THERE WAS NO PATIENT INJURY REPORTED. THE SUBJECT PRODUCT WAS RETURNED AND IT WAS NOTED DURING THE PRELIMINARY EVALUATION THAT THE BARREL INSERT HAD DETACHED AND WAS NOT RETURNED WITH THE DEVICE. A FOLLOW UP WITH THE USER FACILITY WAS MADE. IT WAS FURTHER REPORTED THAT THE DEVICE WOULD NOT FIRE CORRECTLY SO THE SURGEON TRIED TO RESET THE DEVICE BY FIRING OUTSIDE OF THE PATIENT. AT THAT TIME THE TIP BECAME DETACHED FROM THE DEVICE. THE OR TECH THREW THE DETACHED PIECE AWAY FOLLOWING THE CASE.
24-04-2018 | Malfunction | davol inc sub c.r bard | 24-04-2018 | OPTIFIX 30 CT STRAIGHT 39CM | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE OPTIFIX FIXATION DEVICE STOPPED FIRING AFTER THE FIRST FASTENER WAS DEPLOYED. THERE WAS NO PATIENT INJURY REPORTED. THE SUBJECT PRODUCT WAS RETURNED AND IT WAS NOTED DURING THE PRELIMINARY EVALUATION THAT THE BARREL INSERT HAD DETACHED AND WAS NOT RETURNED WITH THE DEVICE. A FOLLOW UP WITH THE USER FACILITY WAS MADE. IT WAS FURTHER REPORTED THAT THE DEVICE WOULD NOT FIRE CORRECTLY SO THE SURGEON TRIED TO RESET THE DEVICE BY FIRING OUTSIDE OF THE PATIENT. AT THAT TIME THE TIP BECAME DETACHED FROM THE DEVICE. THE OR TECH THREW THE DETACHED PIECE AWAY FOLLOWING THE CASE.
24-04-2018 | Malfunction | davol inc sub c.r bard | 24-04-2018 | OPTIFIX 30 CT STRAIGHT 39CM | Material Deformation | A0406 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE OPTIFIX FIXATION DEVICE STOPPED FIRING AFTER THE FIRST FASTENER WAS DEPLOYED. THERE WAS NO PATIENT INJURY REPORTED. THE SUBJECT PRODUCT WAS RETURNED AND IT WAS NOTED DURING THE PRELIMINARY EVALUATION THAT THE BARREL INSERT HAD DETACHED AND WAS NOT RETURNED WITH THE DEVICE. A FOLLOW UP WITH THE USER FACILITY WAS MADE. IT WAS FURTHER REPORTED THAT THE DEVICE WOULD NOT FIRE CORRECTLY SO THE SURGEON TRIED TO RESET THE DEVICE BY FIRING OUTSIDE OF THE PATIENT. AT THAT TIME THE TIP BECAME DETACHED FROM THE DEVICE. THE OR TECH THREW THE DETACHED PIECE AWAY FOLLOWING THE CASE.
13-03-2018 | Death | ethicon endo-surgery | 24-04-2018 | PWRD ECH FLEX 45MM 34CM SHAFT | Material Rupture | A0412 | Death | Event Description: IT WAS REPORTED THAT WE HAVE A CASE WHERE THERE IS SUSPECTED OF RUPTURE OF THE CHAIN OF STAPLES OF A PATIENT WHO DIED YESTERDAY. IN THE SURGICAL ACT EVERYTHING WAS VERY SMOOTH (CALM) AND IN ANY MOMENT WAS NOTICED THE DEFICIENCY OF THE LINE OF STAPLES, HOWEVER THE SAME HAS BROKEN DAYS LATER. THE PATIENT HAD AN ABDOMINAL PAIN AFTER HAVING BEEN DISCHARGED, INFECTION WAS FOUND IN THE ABDOMINAL CAVITY, WHICH PROBABLY CAME FROM THE RUPTURE OF THE EXCLUDED STOMACH. THE RUPTURE OF THE CHAIN OF STAPLES IS A SUSPECT, WE ARE NOT SURE OF WHAT ACTUALLY OCCURRED. PRODUCT WAS DISCARDED BECAUSE THERE WAS NO PROBLEM DURING SURGERY. DATE OF THE FIRST SURGERY: (B)(6) 2018 . DATE OF THE RE-OPERATION: (B)(6) 2018.
13-03-2018 | Death | ethicon endo-surgery | 24-04-2018 | PWRD ECH FLEX 45MM 34CM SHAFT | Material Rupture | A0412 | Unspecified Infection | Event Description: IT WAS REPORTED THAT WE HAVE A CASE WHERE THERE IS SUSPECTED OF RUPTURE OF THE CHAIN OF STAPLES OF A PATIENT WHO DIED YESTERDAY. IN THE SURGICAL ACT EVERYTHING WAS VERY SMOOTH (CALM) AND IN ANY MOMENT WAS NOTICED THE DEFICIENCY OF THE LINE OF STAPLES, HOWEVER THE SAME HAS BROKEN DAYS LATER. THE PATIENT HAD AN ABDOMINAL PAIN AFTER HAVING BEEN DISCHARGED, INFECTION WAS FOUND IN THE ABDOMINAL CAVITY, WHICH PROBABLY CAME FROM THE RUPTURE OF THE EXCLUDED STOMACH. THE RUPTURE OF THE CHAIN OF STAPLES IS A SUSPECT, WE ARE NOT SURE OF WHAT ACTUALLY OCCURRED. PRODUCT WAS DISCARDED BECAUSE THERE WAS NO PROBLEM DURING SURGERY. DATE OF THE FIRST SURGERY: (B)(6) 2018 . DATE OF THE RE-OPERATION: (B)(6) 2018.
13-03-2018 | Death | ethicon endo-surgery | 24-04-2018 | PWRD ECH FLEX 45MM 34CM SHAFT | Material Rupture | A0412 | No Code Available | Event Description: IT WAS REPORTED THAT WE HAVE A CASE WHERE THERE IS SUSPECTED OF RUPTURE OF THE CHAIN OF STAPLES OF A PATIENT WHO DIED YESTERDAY. IN THE SURGICAL ACT EVERYTHING WAS VERY SMOOTH (CALM) AND IN ANY MOMENT WAS NOTICED THE DEFICIENCY OF THE LINE OF STAPLES, HOWEVER THE SAME HAS BROKEN DAYS LATER. THE PATIENT HAD AN ABDOMINAL PAIN AFTER HAVING BEEN DISCHARGED, INFECTION WAS FOUND IN THE ABDOMINAL CAVITY, WHICH PROBABLY CAME FROM THE RUPTURE OF THE EXCLUDED STOMACH. THE RUPTURE OF THE CHAIN OF STAPLES IS A SUSPECT, WE ARE NOT SURE OF WHAT ACTUALLY OCCURRED. PRODUCT WAS DISCARDED BECAUSE THERE WAS NO PROBLEM DURING SURGERY. DATE OF THE FIRST SURGERY: (B)(6) 2018 . DATE OF THE RE-OPERATION: (B)(6) 2018.
05-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS, THE SURGEON TRIED TO FIRE THE DEVICE BUT COULD NOT SQUEEZE THE HANDLE. THE PROCEDURE WAS COMPLETED WITH THE USE OF ANOTHER DEVICE.
11-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | Fever | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, THE SURGEON OPENED THE PACKAGE OF A NEW DEVICE. APPROXIMATED SPECIMEN AND TRIED FIRING. THE HANDLE DID NOT MOVE AND STAPLER DID NOT FIRE. PROCEEDED TO OPEN A CONTOUR CURVED STAPLER. CASE WAS CONCLUDED WITHOUT COMPLICATIONS TO THE PATIENT. THE SALE REP WAS NOT PRESENT.
11-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, THE SURGEON OPENED THE PACKAGE OF A NEW DEVICE. APPROXIMATED SPECIMEN AND TRIED FIRING. THE HANDLE DID NOT MOVE AND STAPLER DID NOT FIRE. PROCEEDED TO OPEN A CONTOUR CURVED STAPLER. CASE WAS CONCLUDED WITHOUT COMPLICATIONS TO THE PATIENT. THE SALE REP WAS NOT PRESENT.
11-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | Sepsis | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, THE SURGEON OPENED THE PACKAGE OF A NEW DEVICE. APPROXIMATED SPECIMEN AND TRIED FIRING. THE HANDLE DID NOT MOVE AND STAPLER DID NOT FIRE. PROCEEDED TO OPEN A CONTOUR CURVED STAPLER. CASE WAS CONCLUDED WITHOUT COMPLICATIONS TO THE PATIENT. THE SALE REP WAS NOT PRESENT.
11-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | Tachycardia | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, THE SURGEON OPENED THE PACKAGE OF A NEW DEVICE. APPROXIMATED SPECIMEN AND TRIED FIRING. THE HANDLE DID NOT MOVE AND STAPLER DID NOT FIRE. PROCEEDED TO OPEN A CONTOUR CURVED STAPLER. CASE WAS CONCLUDED WITHOUT COMPLICATIONS TO THE PATIENT. THE SALE REP WAS NOT PRESENT.
11-01-2018 | Injury | ethicon endo-surgery | 24-04-2018 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, THE SURGEON OPENED THE PACKAGE OF A NEW DEVICE. APPROXIMATED SPECIMEN AND TRIED FIRING. THE HANDLE DID NOT MOVE AND STAPLER DID NOT FIRE. PROCEEDED TO OPEN A CONTOUR CURVED STAPLER. CASE WAS CONCLUDED WITHOUT COMPLICATIONS TO THE PATIENT. THE SALE REP WAS NOT PRESENT.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, WHILE GASTRIC STAPLING, THE SURGEON COULD NOT FIRE THE LOAD AFTER DEPRESSING THE GREEN FIRING BUTTONS, THE LOAD THEN PRE-FIRED. THE SURGEON WAS ALSO UNABLE TO SQUEEZE THE HANDLE. A NEW LOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE, THE STAPLER WAS UNABLE TO FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY TOOK A NEW MAGAZINE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-03-2018 | Injury | covidien lp llc north haven | 24-04-2018 | EEA | Failure To Adhere Or Bond |  | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, FOUR DAYS POST A LOW ANTERIOR RESECTION SURGERY FOR RECTAL CANCER, A TOTAL DISRUPTION OF ANASTOMOSES OCCURRED WITH FECAL PERITONITIS AND SEPTICEMIA. THE DEVICE HAD DEHISCENCE WITH NO LOOSE STAPLE FOUND. IT WAS NOTED THAT THE DEVICE WAS FULLY SQUEEZED WHILE FIRING. RE-EXPLORATION OF THE PATIENT WITH EXTENDED RECOVERY, AND ILEOSTOMY WITH COLOSTOMY WAS DONE AS THE RESULT OF A DELAY IN SURGERY. THE SURGERY WAS EXTENDED FOR MORE THAN 30 MINUTES AND THE PATIENT¿S HOSPITALIZATION STAY WAS EXTENDED FOR 7 DAYS.
24-03-2018 | Injury | covidien lp llc north haven | 24-04-2018 | EEA | Failure To Adhere Or Bond |  | Sepsis | Event Description: ACCORDING TO THE REPORTER, FOUR DAYS POST A LOW ANTERIOR RESECTION SURGERY FOR RECTAL CANCER, A TOTAL DISRUPTION OF ANASTOMOSES OCCURRED WITH FECAL PERITONITIS AND SEPTICEMIA. THE DEVICE HAD DEHISCENCE WITH NO LOOSE STAPLE FOUND. IT WAS NOTED THAT THE DEVICE WAS FULLY SQUEEZED WHILE FIRING. RE-EXPLORATION OF THE PATIENT WITH EXTENDED RECOVERY, AND ILEOSTOMY WITH COLOSTOMY WAS DONE AS THE RESULT OF A DELAY IN SURGERY. THE SURGERY WAS EXTENDED FOR MORE THAN 30 MINUTES AND THE PATIENT¿S HOSPITALIZATION STAY WAS EXTENDED FOR 7 DAYS.
24-03-2018 | Injury | covidien lp llc north haven | 24-04-2018 | EEA | Failure To Adhere Or Bond |  | Peritonitis | Event Description: ACCORDING TO THE REPORTER, FOUR DAYS POST A LOW ANTERIOR RESECTION SURGERY FOR RECTAL CANCER, A TOTAL DISRUPTION OF ANASTOMOSES OCCURRED WITH FECAL PERITONITIS AND SEPTICEMIA. THE DEVICE HAD DEHISCENCE WITH NO LOOSE STAPLE FOUND. IT WAS NOTED THAT THE DEVICE WAS FULLY SQUEEZED WHILE FIRING. RE-EXPLORATION OF THE PATIENT WITH EXTENDED RECOVERY, AND ILEOSTOMY WITH COLOSTOMY WAS DONE AS THE RESULT OF A DELAY IN SURGERY. THE SURGERY WAS EXTENDED FOR MORE THAN 30 MINUTES AND THE PATIENT¿S HOSPITALIZATION STAY WAS EXTENDED FOR 7 DAYS.
24-03-2018 | Injury | covidien lp llc north haven | 24-04-2018 | EEA | Failure To Adhere Or Bond |  | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOUR DAYS POST A LOW ANTERIOR RESECTION SURGERY FOR RECTAL CANCER, A TOTAL DISRUPTION OF ANASTOMOSES OCCURRED WITH FECAL PERITONITIS AND SEPTICEMIA. THE DEVICE HAD DEHISCENCE WITH NO LOOSE STAPLE FOUND. IT WAS NOTED THAT THE DEVICE WAS FULLY SQUEEZED WHILE FIRING. RE-EXPLORATION OF THE PATIENT WITH EXTENDED RECOVERY, AND ILEOSTOMY WITH COLOSTOMY WAS DONE AS THE RESULT OF A DELAY IN SURGERY. THE SURGERY WAS EXTENDED FOR MORE THAN 30 MINUTES AND THE PATIENT¿S HOSPITALIZATION STAY WAS EXTENDED FOR 7 DAYS.
24-03-2018 | Injury | covidien lp llc north haven | 24-04-2018 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, FOUR DAYS POST A LOW ANTERIOR RESECTION SURGERY FOR RECTAL CANCER, A TOTAL DISRUPTION OF ANASTOMOSES OCCURRED WITH FECAL PERITONITIS AND SEPTICEMIA. THE DEVICE HAD DEHISCENCE WITH NO LOOSE STAPLE FOUND. IT WAS NOTED THAT THE DEVICE WAS FULLY SQUEEZED WHILE FIRING. RE-EXPLORATION OF THE PATIENT WITH EXTENDED RECOVERY, AND ILEOSTOMY WITH COLOSTOMY WAS DONE AS THE RESULT OF A DELAY IN SURGERY. THE SURGERY WAS EXTENDED FOR MORE THAN 30 MINUTES AND THE PATIENT¿S HOSPITALIZATION STAY WAS EXTENDED FOR 7 DAYS.
24-03-2018 | Injury | covidien lp llc north haven | 24-04-2018 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE REPORTER, FOUR DAYS POST A LOW ANTERIOR RESECTION SURGERY FOR RECTAL CANCER, A TOTAL DISRUPTION OF ANASTOMOSES OCCURRED WITH FECAL PERITONITIS AND SEPTICEMIA. THE DEVICE HAD DEHISCENCE WITH NO LOOSE STAPLE FOUND. IT WAS NOTED THAT THE DEVICE WAS FULLY SQUEEZED WHILE FIRING. RE-EXPLORATION OF THE PATIENT WITH EXTENDED RECOVERY, AND ILEOSTOMY WITH COLOSTOMY WAS DONE AS THE RESULT OF A DELAY IN SURGERY. THE SURGERY WAS EXTENDED FOR MORE THAN 30 MINUTES AND THE PATIENT¿S HOSPITALIZATION STAY WAS EXTENDED FOR 7 DAYS.
24-03-2018 | Injury | covidien lp llc north haven | 24-04-2018 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Peritonitis | Event Description: ACCORDING TO THE REPORTER, FOUR DAYS POST A LOW ANTERIOR RESECTION SURGERY FOR RECTAL CANCER, A TOTAL DISRUPTION OF ANASTOMOSES OCCURRED WITH FECAL PERITONITIS AND SEPTICEMIA. THE DEVICE HAD DEHISCENCE WITH NO LOOSE STAPLE FOUND. IT WAS NOTED THAT THE DEVICE WAS FULLY SQUEEZED WHILE FIRING. RE-EXPLORATION OF THE PATIENT WITH EXTENDED RECOVERY, AND ILEOSTOMY WITH COLOSTOMY WAS DONE AS THE RESULT OF A DELAY IN SURGERY. THE SURGERY WAS EXTENDED FOR MORE THAN 30 MINUTES AND THE PATIENT¿S HOSPITALIZATION STAY WAS EXTENDED FOR 7 DAYS.
24-03-2018 | Injury | covidien lp llc north haven | 24-04-2018 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOUR DAYS POST A LOW ANTERIOR RESECTION SURGERY FOR RECTAL CANCER, A TOTAL DISRUPTION OF ANASTOMOSES OCCURRED WITH FECAL PERITONITIS AND SEPTICEMIA. THE DEVICE HAD DEHISCENCE WITH NO LOOSE STAPLE FOUND. IT WAS NOTED THAT THE DEVICE WAS FULLY SQUEEZED WHILE FIRING. RE-EXPLORATION OF THE PATIENT WITH EXTENDED RECOVERY, AND ILEOSTOMY WITH COLOSTOMY WAS DONE AS THE RESULT OF A DELAY IN SURGERY. THE SURGERY WAS EXTENDED FOR MORE THAN 30 MINUTES AND THE PATIENT¿S HOSPITALIZATION STAY WAS EXTENDED FOR 7 DAYS.
19-06-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THEY WERE CLOSING ON THE TISSUE BUT THE STAPLER WOULD NOT FIRE. THEY OPENED A NEW ONE AND COMPLETED THE PROCEDURE. THERE WAS NO REPORTED PATIENT INVOLVEMENT AND OUTCOME.
13-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON LUNG CANCER RADICAL SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AS IT WAS POORLY CLOSED, THE STAPLE LINE WAS INCOMPLETE THAT WAS FIXED BY HAND SEWING. BLUE RELOADS WERE USED REMOVE THE MALOCCLUSION OF THE INTERLOBULAR LOBE AND REPLACED IT WITH A GREEN ONE. SPUTUM SOUND AND POSTERIOR HEMORRHAGE WAS NOTED WHICH THEY HAND-SEWN THAT THE OPERATION TIME WAS EXTENDED BY TWENTY MINUTES TO COMPLETE THE CASE.
13-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON LUNG CANCER RADICAL SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AS IT WAS POORLY CLOSED, THE STAPLE LINE WAS INCOMPLETE THAT WAS FIXED BY HAND SEWING. BLUE RELOADS WERE USED REMOVE THE MALOCCLUSION OF THE INTERLOBULAR LOBE AND REPLACED IT WITH A GREEN ONE. SPUTUM SOUND AND POSTERIOR HEMORRHAGE WAS NOTED WHICH THEY HAND-SEWN THAT THE OPERATION TIME WAS EXTENDED BY TWENTY MINUTES TO COMPLETE THE CASE.
10-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | EEA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMORRHOIDECTOMY PROCEDURE, THE PRODUCT WAS BROKEN WHEN IT WAS REMOVED FROM THE ORIGINAL STERILE PACKAGING. THE ANVIL OF THE PRODUCT WAS SEPARATED FROM THE TROCAR. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | EEA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMORRHOIDECTOMY PROCEDURE, THE PRODUCT WAS BROKEN WHEN IT WAS REMOVED FROM THE ORIGINAL STERILE PACKAGING. THE ANVIL OF THE PRODUCT WAS SEPARATED FROM THE TROCAR. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | EEA | Device Packaging Compromised |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMORRHOIDECTOMY PROCEDURE, THE PRODUCT WAS BROKEN WHEN IT WAS REMOVED FROM THE ORIGINAL STERILE PACKAGING. THE ANVIL OF THE PRODUCT WAS SEPARATED FROM THE TROCAR. THERE WAS NO PATIENT INJURY.
24-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THEY WERE CLOSING ON THE TISSUE BUT THE STAPLER WOULD NOT FIRE. THEY OPENED A NEW ONE AND COMPLETED THE PROCEDURE. THERE WAS NO REPORTED PATIENT INVOLVEMENT AND OUTCOME.
26-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY BY THORACOSCOPE PROCEDURE, AT THE MOMENT OF THE SECTION OF THE PULMONARY VEIN, THE DEVICE DID NOT TOTALLY STAPLED THE VEIN WHICH LED TO A BLEEDING AND CHANGE OF PROCEDURE. A THORACOTOMIE WAS QUICKLY PERFORMED IN ORDER TO STOP THE BLEEDING.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, AT RE-CONSTRUCTION OF ILEUM AND COLON, FOR THE FIRING TO BLIND END OF ESOPHAGUS, THEY CONNECTED THE CARTRIDGE TO DEVICE HOWEVER COULD NOT CLOSE THE JAWS. REPLACED BY NEW DEVICE, AND THE PROCEDURE WAS COMPLETED. AFTER THAT ,THE SALES REP RETURN THE DEVICE TO THE NEUTRAL POSITION FORCIBLY AND THE DEVICE REACTED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
25-03-2018 | Malfunction | ethicon endo-surgery | 24-04-2018 | ECHELON FLEX 45 STANDARD | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLES MALFORMED AFTER FIRING. DURING THE LEFT UPPER LOBE EXCISION, AFTER FIRED WITH ECR45W, FOUND THE PROXIMAL AND DISTAL STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | Fistula | Event Description: IT WAS REPORTED THAT DURING A CARCINOLOGIC RESECTION OF RECTUM, THE ANVIL GOT DETACHED FROM THE TROCAR AND IT STAYED IN THE COLON DURING THE REMOVAL. THE SURGEON HAS RECOVERED THE ANVIL WITH A GRASPING CLAMP AND HE HAS PERFORMED A MANUAL SUTURE ON THE STAPLES LINE BY SAFETY. HOWEVER THE STAPLES LINE DID NOT HOLD AND ANOTHER DEVICE HAS BEEN USED TO PERFORM THE ANASTAMOSIS ON FRAGILE AND RESTRICTED AREA.
12-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A CARCINOLOGIC RESECTION OF RECTUM, THE ANVIL GOT DETACHED FROM THE TROCAR AND IT STAYED IN THE COLON DURING THE REMOVAL. THE SURGEON HAS RECOVERED THE ANVIL WITH A GRASPING CLAMP AND HE HAS PERFORMED A MANUAL SUTURE ON THE STAPLES LINE BY SAFETY. HOWEVER THE STAPLES LINE DID NOT HOLD AND ANOTHER DEVICE HAS BEEN USED TO PERFORM THE ANASTAMOSIS ON FRAGILE AND RESTRICTED AREA.
12-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A CARCINOLOGIC RESECTION OF RECTUM, THE ANVIL GOT DETACHED FROM THE TROCAR AND IT STAYED IN THE COLON DURING THE REMOVAL. THE SURGEON HAS RECOVERED THE ANVIL WITH A GRASPING CLAMP AND HE HAS PERFORMED A MANUAL SUTURE ON THE STAPLES LINE BY SAFETY. HOWEVER THE STAPLES LINE DID NOT HOLD AND ANOTHER DEVICE HAS BEEN USED TO PERFORM THE ANASTAMOSIS ON FRAGILE AND RESTRICTED AREA.
12-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ILS 29MM, CURVED | Detachment of Device or Device Component | A0501 | Fistula | Event Description: IT WAS REPORTED THAT DURING A CARCINOLOGIC RESECTION OF RECTUM, THE ANVIL GOT DETACHED FROM THE TROCAR AND IT STAYED IN THE COLON DURING THE REMOVAL. THE SURGEON HAS RECOVERED THE ANVIL WITH A GRASPING CLAMP AND HE HAS PERFORMED A MANUAL SUTURE ON THE STAPLES LINE BY SAFETY. HOWEVER THE STAPLES LINE DID NOT HOLD AND ANOTHER DEVICE HAS BEEN USED TO PERFORM THE ANASTAMOSIS ON FRAGILE AND RESTRICTED AREA.
12-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ILS 29MM, CURVED | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A CARCINOLOGIC RESECTION OF RECTUM, THE ANVIL GOT DETACHED FROM THE TROCAR AND IT STAYED IN THE COLON DURING THE REMOVAL. THE SURGEON HAS RECOVERED THE ANVIL WITH A GRASPING CLAMP AND HE HAS PERFORMED A MANUAL SUTURE ON THE STAPLES LINE BY SAFETY. HOWEVER THE STAPLES LINE DID NOT HOLD AND ANOTHER DEVICE HAS BEEN USED TO PERFORM THE ANASTAMOSIS ON FRAGILE AND RESTRICTED AREA.
12-03-2018 | Injury | ethicon endo-surgery | 24-04-2018 | ILS 29MM, CURVED | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A CARCINOLOGIC RESECTION OF RECTUM, THE ANVIL GOT DETACHED FROM THE TROCAR AND IT STAYED IN THE COLON DURING THE REMOVAL. THE SURGEON HAS RECOVERED THE ANVIL WITH A GRASPING CLAMP AND HE HAS PERFORMED A MANUAL SUTURE ON THE STAPLES LINE BY SAFETY. HOWEVER THE STAPLES LINE DID NOT HOLD AND ANOTHER DEVICE HAS BEEN USED TO PERFORM THE ANASTAMOSIS ON FRAGILE AND RESTRICTED AREA.
09-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE JAWS GOT LOCK ON TISSUE. THE DEVICE WAS REMOVED BY OPENING IT. THERE WAS NO PATIENT INJURY.
05-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY PROCEDURE, THE FIRING FOR ESOPHAGUS RESECTION AND 4 FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE SUCCESSFUL BUT AT THE 5TH FIRING OF THE DEVICE, THE SURGEON NOTED THE GASTRIC WALL WAS SLIGHTLY THICK AND HE STARTED THE FIRING THE DEVICE SLOWLY, HOWEVER ,ON THE HALFWAY, THE HANDLE BECAME STIFF AND HE COULD NOT SQUEEZE IT. SURGEON OPENED THE JAWS AND CONNECTED A NEW CARTRIDGE TO A NEW HANDLE, THE FIRING WAS SUCCESSFUL. THERE WAS NO PATIENT INJURY.
22-03-2018 | Malfunction | ethicon endo-surgery | 24-04-2018 | ECHELON 60 ENDOPATH STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLES MALFORMED AFTER FIRING. DURING THE VATS LOBECTOMY, AFTER FIRED WITH ECR60B, WHEN SEVERING THE LOBE OF LUNG, FOUND THE STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | SIGNIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, AT RE-CONSTRUCTION OF ILEUM AND COLON, FOR THE FIRING TO BLIND END OF ESOPHAGUS, THEY CONNECTED THE CARTRIDGE TO DEVICE HOWEVER COULD NOT CLOSE THE JAWS. REPLACED BY NEW DEVICE ,AND THE PROCEDURE WAS COMPLETED. AFTER THAT ,THE SALES REP RETURN THE DEVICE TO THE NEUTRAL POSITION FORCIBLY AND THE DEVICE REACTED WITH NO PROBLEM. NO PATIENT INJURY.
05-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY PROCEDURE, THE FIRING FOR ESOPHAGUS RESECTION AND 4 FIRINGS FOR GASTRIC TUBE CONSTRUCTION WERE SUCCESSFUL BUT AT THE 5TH FIRING OF THE DEVICE, THE SURGEON NOTED THE GASTRIC WALL WAS SLIGHTLY THICK AND HE STARTED THE FIRING THE DEVICE SLOWLY, HOWEVER, ON THE HALFWAY, THE HANDLE BECAME STIFF AND HE COULD NOT SQUEEZE IT. SURGEON OPENED THE JAWS AND CONNECTED A NEW CARTRIDGE TO A NEW HANDLE, THE FIRING WAS SUCCESSFUL. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON LUNG CANCER RADICAL SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AS IT WAS POORLY CLOSED, THE STAPLE LINE WAS INCOMPLETE THAT WAS FIXED BY HAND SEWING. BLUE RELOADS WERE USED REMOVE THE MALOCCLUSION OF THE INTERLOBULAR LOBE AND REPLACED IT WITH A GREEN ONE. SPUTUM SOUND AND POSTERIOR HEMORRHAGE WAS NOTED WHICH THEY HAND-SEWN THAT THE OPERATION TIME WAS EXTENDED BY TWENTY MINUTES TO COMPLETE THE CASE.
13-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON LUNG CANCER RADICAL SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AS IT WAS POORLY CLOSED, THE STAPLE LINE WAS INCOMPLETE THAT WAS FIXED BY HAND SEWING. BLUE RELOADS WERE USED REMOVE THE MALOCCLUSION OF THE INTERLOBULAR LOBE AND REPLACED IT WITH A GREEN ONE. SPUTUM SOUND AND POSTERIOR HEMORRHAGE WAS NOTED WHICH THEY HAND-SEWN THAT THE OPERATION TIME WAS EXTENDED BY TWENTY MINUTES TO COMPLETE THE CASE.
08-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PEDIATRIC LAPAROSCOPIC APPENDECTOMY, THE STAPLER DID NOT FIRE. WHILE REMOVING THE STAPLER FROM THE ABDOMEN TO BE CHECKED, THE RELOAD CAME LOOSE FROM THE HANDLE. A RETRIEVAL BAG AND HAND INSTRUMENT WAS USED TO SECURE THE RELOAD AND REMOVE IT FROM THE ABDOMEN. THERE WAS NO PATIENT INJURY.
08-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PEDIATRIC LAPAROSCOPIC APPENDECTOMY, THE STAPLER DID NOT FIRE. WHILE REMOVING THE STAPLER FROM THE ABDOMEN TO BE CHECKED, THE RELOAD CAME LOOSE FROM THE HANDLE. A RETRIEVAL BAG AND HAND INSTRUMENT WAS USED TO SECURE THE RELOAD AND REMOVE IT FROM THE ABDOMEN. THERE WAS NO PATIENT INJURY.
12-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRECTOMY, AFTER FIRST FIRING ON THE STOMACH, THE FOLLOWING RELOAD WAS ATTEMPTED FOR A REMAINING RESECTION. AS STATED, THE COMPLAINT WAS THAT HALF OF STAPLE LINE'S DISTAL SIDE WAS NOT PROPERLY FORMED. THE MALFORMED STAPLE LINE WAS LIGATED USING MANUAL SUTURE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | GIA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN COLECTOMY, WHILE TRANSECTION OF THE BOWEL, THE BLACK FIRING KNOB FELL OFF THE STAPLER WHEN THE DEVICE WAS OPENED FROM ITS PACKAGING. THE KNOB WAS ATTACHED BACK ON THE HANDLE, AND IT WAS FIRED ACROSS THE BOWEL. AT THAT POINT, THE TEAM GOT A NEW STAPLER FOR FURTHER TRANSACTIONS. THERE WERE NO ISSUES WITH THE FIRING AND DEPLOYMENT OF THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
12-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS UNABLE TO FIRE AND ITS HANDLE WAS UNABLE TO BE SQUEEZED. THERE WAS NO PATIENT INVOLVEMENT.
13-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON LUNG CANCER RADICAL SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AS IT WAS POORLY CLOSED, THE STAPLE LINE WAS INCOMPLETE THAT WAS FIXED BY HAND SEWING. BLUE RELOADS WERE USED REMOVE THE MALOCCLUSION OF THE INTERLOBULAR LOBE AND REPLACED IT WITH A GREEN ONE. SPUTUM SOUND AND POSTERIOR HEMORRHAGE WAS NOTED WHICH THEY HAND-SEWN THAT THE OPERATION TIME WAS EXTENDED BY TWENTY MINUTES TO COMPLETE THE CASE.
13-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE ON LUNG CANCER RADICAL SURGERY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AS IT WAS POORLY CLOSED, THE STAPLE LINE WAS INCOMPLETE THAT WAS FIXED BY HAND SEWING. BLUE RELOADS WERE USED REMOVE THE MALOCCLUSION OF THE INTERLOBULAR LOBE AND REPLACED IT WITH A GREEN ONE. SPUTUM SOUND AND POSTERIOR HEMORRHAGE WAS NOTED WHICH THEY HAND-SEWN THAT THE OPERATION TIME WAS EXTENDED BY TWENTY MINUTES TO COMPLETE THE CASE.
12-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS UNABLE TO FIRE AND ITS HANDLE WAS UNABLE TO BE SQUEEZED. THERE WAS NO PATIENT INVOLVEMENT.
09-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE JAWS GOT LOCK ON TISSUE. THE DEVICE WAS REMOVED BY OPENING IT. THEY CHANGED THE DEVICE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHILE THE ASSEMBLY OF THE SECOND CHARGE OF 45 MM PURPLE, THE DEVICE DID NOT FIRE. THERE WAS AN INVOLUNTARY DRIVE, WHERE THE SAME CAME TO START THE DISPARATE AND CONSEQUENTLY LOCKING OF THE STAPLER. TO THE ACTUATED, THERE WAS A SNAP ON THE GAUNTLET LEAVING IT UNUSABLE. THERE WAS NO PATIENT INJURY.
06-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHILE THE ASSEMBLY OF THE SECOND CHARGE OF 45 MM PURPLE, THE DEVICE DID NOT FIRE. THERE WAS AN INVOLUNTARY DRIVE, WHERE THE SAME CAME TO START THE DISPARATE AND CONSEQUENTLY LOCKING OF THE STAPLER. TO THE ACTUATED, THERE WAS A SNAP ON THE GAUNTLET LEAVING IT UNUSABLE. THERE WAS NO PATIENT INJURY.
06-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | nan | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHILE THE ASSEMBLY OF THE SECOND CHARGE OF 45 MM PURPLE, THE DEVICE DID NOT FIRE. THERE WAS AN INVOLUNTARY DRIVE, WHERE THE SAME CAME TO START THE DISPARATE AND CONSEQUENTLY LOCKING OF THE STAPLER. TO THE ACTUATED, THERE WAS A SNAP ON THE GAUNTLET LEAVING IT UNUSABLE. THERE WAS NO PATIENT INJURY.
06-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | nan | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHILE THE ASSEMBLY OF THE SECOND CHARGE OF 45 MM PURPLE, THE DEVICE DID NOT FIRE. THERE WAS AN INVOLUNTARY DRIVE, WHERE THE SAME CAME TO START THE DISPARATE AND CONSEQUENTLY LOCKING OF THE STAPLER. TO THE ACTUATED, THERE WAS A SNAP ON THE GAUNTLET LEAVING IT UNUSABLE. THERE WAS NO PATIENT INJURY.
05-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS, THE SURGEON TRIED TO FIRE THE DEVICE BUT COULD NOT SQUEEZE THE HANDLE. THE PROCEDURE WAS COMPLETED WITH THE USE OF ANOTHER DEVICE.
06-04-2018 | Death | us surgical puerto rico | 24-04-2018 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, ON THE FOURTH POST-OPERATIVE DAY, FOLLOWING A RECTUM RESECTION, AFTER THEY HAD BUILT AN ANASTOMOSE, AN INSUFFICIENCY OCCURRED AND TWO STAPLES DID NOT PROPERLY TAKE SHAPE. IT WAS FOUND IN A SUBSEQUENT ENDOSCOPIC EXAMINATION. AT 6 O'CLOCK A SMALL HOLE OF LESS THAN A FOUR MILLIMETER WAS FOUND.
06-04-2018 | Death | us surgical puerto rico | 24-04-2018 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, ON THE FOURTH POST-OPERATIVE DAY, FOLLOWING A RECTUM RESECTION, AFTER THEY HAD BUILT AN ANASTOMOSE, AN INSUFFICIENCY OCCURRED AND TWO STAPLES DID NOT PROPERLY TAKE SHAPE. IT WAS FOUND IN A SUBSEQUENT ENDOSCOPIC EXAMINATION. AT 6 O'CLOCK A SMALL HOLE OF LESS THAN A FOUR MILLIMETER WAS FOUND.
10-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | nan | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, PROCEDURE, WHILE BEING APPLIED TO THE LUNG, THE SURGEON REMARKED THAT THE STAPLER FELT OFF AND WOULD NOT CLICK CLOSED ON THE TISSUE. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN COLON SURGERY, WHILE INTESTINE CUTTING, THE DEVICE CAN BE USED AT FIRST AND SECOND SUTURE, BUT WHEN RELOAD ON THE THIRD AND THE FOURTH STAPLER, IT STUCK AND CANNOT BE PUSH. TO RESOLVE THE ISSUE, THE SURGEON NEEDED TO SUTURE IT BY HAND. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: DURING WEDGE RESECTION, AND AFTER THE FIRING, THE CARTRIDGE WAS CHANGED, HOWEVER COULD NOT CLOSE THE JAWS. REPLACED A NEW CARTRIDGE, AND THE FIRING WAS COMPLETED. NO PATIENT INJURY.
11-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLON SURGERY, THE DEVICE DID NOT FIRE, THEY THEN REPLACED IT BY ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE, THE STAPLER WAS UNABLE TO FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY TOOK A NEW MAGAZINE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/WEDGE/SEGMENTAL RESECTION OF THE LUNG, AFTER FIRING THE DEVICE, THE SURGEON COULD NOT PULL THE BLACK RETURN KNOB. SINCE THE JAWS COULD NOT BE OPENED, THE TISSUE WAS SLID LITTLE BY LITTLE TO REMOVE THE CARTRIDGE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, PROCEDURE, WHILE BEING APPLIED TO THE LUNG, THE SURGEON REMARKED THAT THE STAPLER FELT OFF AND WOULD NOT CLICK CLOSED ON THE TISSUE. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE AND AFTER SWITCHING FROM ARTICULATE STRAIGHT RELOAD AND BACK TO ARTICULATE FOR THE SECOND FIRING, THE RELOAD, WHEN TESTED, DIDN'T OPEN AND THEY NEEDED TO MANUALLY RETRACT THE KNIFE. THEN THE RELOAD WAS PUT ON TISSUE, WHEN ARTICULATED AND TRIED TO FIRE, IT NO LONGER FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE.
11-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/WEDGE/SEGMENTAL RESECTION OF THE LUNG, AFTER FIRING THE DEVICE, THE SURGEON COULD NOT PULL THE BLACK RETURN KNOB. SINCE THE JAWS COULD NOT BE OPENED, THE TISSUE WAS SLID LITTLE BY LITTLE TO REMOVE THE CARTRIDGE. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN COLON SURGERY, WHILE CUTTING THE INTESTINE, THE DEVICE CAN BE USED ON THE FIRST AND SECOND SUTURE, BUT WHEN RELOAD ON THE THIRD AND THE FOURTH STAPLER, IT WAS STUCK AND CANNOT BE PUSHED. TO RESOLVE THE ISSUE, THE SURGEON NEEDED TO SUTURE IT BY HAND. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORASCOPIC WEDGE SEGMENTAL RESECTION, AT THE FIRST FIRING OF SEGMENTAL RESECTION OF LUNG, THE SURGEON STARTED SQUEEZING THE HANDLE, HOWEVER COULD NOT FULLY SQUEEZE IT AND COULD NOT FULLY FIRE THE DEVICE TO PULMONARY PARENCHYMA . THEY DID NOT CLAMP ANY CLIPS NOR FORCEPS. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT INJURY.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE AND AFTER SWITCHING FROM ARTICULATE STRAIGHT RELOAD AND BACK TO ARTICULATE FOR THE SECOND FIRING, THE RELOAD, WHEN TESTED, DIDN'T OPEN AND THEY NEEDED TO MANUALLY RETRACT THE KNIFE. WHEN THE RELOAD WAS PUT ON TISSUE, ARTICULATED AND TRIED TO FIRE, IT NO LONGER FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE DEVICE'S FIRING STOPPED ON THE HALFWAY. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLON SURGERY, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: DURING WEDGE RESECTION, AND AFTER THE FIRING, THE CARTRIDGE WAS CHANGED, HOWEVER COULD NOT CLOSE THE JAWS. REPLACED A NEW CARTRIDGE, AND THE FIRING WAS COMPLETED. NO PATIENT INJURY.
12-04-2018 | Injury | us surgical puerto rico | 24-04-2018 | EEA | Failure To Adhere Or Bond |  | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC IVOR LEWIS ESOPHAGECTOMY, THE ANVIL HAD BEEN INTRODUCED TRANSORALLY INTO ESOPHAGUS, AND THE STAPLER WAS INTRODUCED THROUGH THORACOTOMY INTO THE GASTRIC POUCH. THE TWO WERE TO BE JOINED TOGETHER IN ORDER TO CREATE ANASTOMOSIS. THE SURGEON FOLLOWED THE SAME STEPS AS PRIOR SIMILAR PROCEDURES AND HAD VERY GOOD ACCESS AND ANGLE FOR CONNECTING. IT ALSO DID NOT APPEAR THAT ANY DAMAGE OCCURRED TO THE DEVICE DURING HANDLING, HOWEVER, THE TWO PARTS COULD NOT BE MATED TOGETHER. A SECOND STAPLER WAS OPENED TO SEE IF THEY COULD BE CONNECTED. WHEN THAT FAILED, THE ANVIL WAS REMOVED TRANSORALLY AND A SECOND ANVIL WAS INTRODUCED. THE DEVICES STILL COULD NOT BE MATED USING A REASONABLE FORCE. THE CASE WAS CONVERTED TO AN OPEN PROCEDURE SO THAT THE ANVIL AND THE STAPLER COULD BE JOINED USING HANDS INSTEAD OF ENDOSCOPICALLY.
12-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | EEA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF STAPLING ON A RECTOSIGMOIDECTOMY, THE DOCTOR FAILED TO FIRE BECAUSE THE SAFETY CATCH WAS BROKEN, PREVENTED THE STAPLER FROM CLOSING. THERE WAS NO PATIENT INJURY.
04-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN COLON SURGERY, WHILE CUTTING THE INTESTINE, THE DEVICE CAN BE USED AT FIRST AND SECOND SUTURE, BUT WHEN RELOAD ON THE THIRD AND FOURTH STAPLER, IT GOT STUCK AND CANNOT BE PUSHED. TO RESOLVE THE ISSUE, THE SURGEON NEEDED TO SUTURE IT BY HAND. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORASCOPIC WEDGE SEGMENTAL RESECTION, AT THE FIRST FIRING OF SEGMENTAL RESECTION OF LUNG, THE SURGEON STARTED SQUEEZING THE HANDLE, HOWEVER COULD NOT FULLY SQUEEZE IT AND COULD NOT FULLY FIRE THE DEVICE TO PULMONARY PARENCHYMA. THEY DID NOT CLAMP ANY CLIPS NOR FORCEPS. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | EEA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO ANASTOMOSES ON A LAPAROSCOPIC PROCEDURE, THEY WERE ABLE TO LOAD THE DEVICE PROPERLY BUT COMPONENT DISENGAGED. THE STAPLE CARTRIDGE DISENGAGED AND STAYED IN THE PARIETAL MUSCLE BUT WAS RETRIEVED. THEY CHANGED THE DEVICE TO COMPLETE THE CASE.
10-04-2018 | Injury | us surgical puerto rico | 24-04-2018 | EEA | Mechanics Altered | A0507 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, WHILE THE ANASTOMOSIS OF THE RECTUM TO THE COLON, THE STAPLER CUT THE TISSUE, BUT IT DID NOT STAPLE DUE TO THE LACK OF STAPLES IN THE DEVICE. A SURGICAL INTERVENTION WAS NEEDED TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION AND THE SURGERY WAS MADE TO GIVE BACK A GOOD HEALTH STATE TO THE PATIENT. DUE TO ALREADY REDUCED SPACE FOR PERFORMING THE STAPLING, THE PATIENT'S RECTUM WAS COMMITTED. A LAPAROSCOPIC PROCEDURE CONVERTED TO OPEN DUE TO DEVICE PROBLEM. IT TOOK OVER 1 HOUR AND A HALF FOR THE SURGEON TO SOLVE THE PROBLEM. AT FIRST, THE SURGEON TRIED TO PERFORM THE ANASTOMOSIS THROUGH VIDEO LAPAROSCOPY, SINCE IT WAS NOT POSSIBLE, THE SURGEON DISSECTED THE RECTUM IN ORDER TO PERFORM A DEEPER STAPLING. FINALLY, THE SURGEON WAS FORCED TO CONVERT THE MINIMALLY INVASIVE PROCEDURE TO AN OPEN SURGERY AND IT TOOK A LITTLE MORE TIME TO MAKE THE ANASTOMOSIS. THE INCISION WAS EXTENDED TO 2.54 CM DUE TO THE PRODUCT PROBLEM.
10-04-2018 | Injury | us surgical puerto rico | 24-04-2018 | EEA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, WHILE THE ANASTOMOSIS OF THE RECTUM TO THE COLON, THE STAPLER CUT THE TISSUE, BUT IT DID NOT STAPLE DUE TO THE LACK OF STAPLES IN THE DEVICE. A SURGICAL INTERVENTION WAS NEEDED TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION AND THE SURGERY WAS MADE TO GIVE BACK A GOOD HEALTH STATE TO THE PATIENT. DUE TO ALREADY REDUCED SPACE FOR PERFORMING THE STAPLING, THE PATIENT'S RECTUM WAS COMMITTED. A LAPAROSCOPIC PROCEDURE CONVERTED TO OPEN DUE TO DEVICE PROBLEM. IT TOOK OVER 1 HOUR AND A HALF FOR THE SURGEON TO SOLVE THE PROBLEM. AT FIRST, THE SURGEON TRIED TO PERFORM THE ANASTOMOSIS THROUGH VIDEO LAPAROSCOPY, SINCE IT WAS NOT POSSIBLE, THE SURGEON DISSECTED THE RECTUM IN ORDER TO PERFORM A DEEPER STAPLING. FINALLY, THE SURGEON WAS FORCED TO CONVERT THE MINIMALLY INVASIVE PROCEDURE TO AN OPEN SURGERY AND IT TOOK A LITTLE MORE TIME TO MAKE THE ANASTOMOSIS. THE INCISION WAS EXTENDED TO 2.54 CM DUE TO THE PRODUCT PROBLEM.
10-04-2018 | Injury | us surgical puerto rico | 24-04-2018 | EEA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, WHILE THE ANASTOMOSIS OF THE RECTUM TO THE COLON, THE STAPLER CUT THE TISSUE, BUT IT DID NOT STAPLE DUE TO THE LACK OF STAPLES IN THE DEVICE. A SURGICAL INTERVENTION WAS NEEDED TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION AND THE SURGERY WAS MADE TO GIVE BACK A GOOD HEALTH STATE TO THE PATIENT. DUE TO ALREADY REDUCED SPACE FOR PERFORMING THE STAPLING, THE PATIENT'S RECTUM WAS COMMITTED. A LAPAROSCOPIC PROCEDURE CONVERTED TO OPEN DUE TO DEVICE PROBLEM. IT TOOK OVER 1 HOUR AND A HALF FOR THE SURGEON TO SOLVE THE PROBLEM. AT FIRST, THE SURGEON TRIED TO PERFORM THE ANASTOMOSIS THROUGH VIDEO LAPAROSCOPY, SINCE IT WAS NOT POSSIBLE, THE SURGEON DISSECTED THE RECTUM IN ORDER TO PERFORM A DEEPER STAPLING. FINALLY, THE SURGEON WAS FORCED TO CONVERT THE MINIMALLY INVASIVE PROCEDURE TO AN OPEN SURGERY AND IT TOOK A LITTLE MORE TIME TO MAKE THE ANASTOMOSIS. THE INCISION WAS EXTENDED TO 2.54 CM DUE TO THE PRODUCT PROBLEM.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | EEA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LONGO HAEMORRHOIDOPEXY PROCEDURE, A COMPONENT DISENGAGED AND THE DEVICE WAS NOT PROPERLY LOADED. THE INSTRUMENT WAS NOT APPLIED TO THE PATIENT AND NO COMPONENTS WERE LOST. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | EEA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LONGO HAEMORRHOIDOPEXY PROCEDURE, A COMPONENT DISENGAGED AND THE DEVICE WAS NOT PROPERLY LOADED. THE INSTRUMENT WAS NOT APPLIED TO THE PATIENT AND NO COMPONENTS WERE LOST. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | Endo GIA Ultra | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE. ANOTHER STAPLER DEVICE AND NEW RELOAD WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-04-2018 | Injury | covidien lp llc north haven | 24-04-2018 | EEA ORVIL | Failure To Adhere Or Bond |  | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC IVOR LEWIS ESOPHAGECTOMY, THE ANVIL HAD BEEN INTRODUCED TRANSORALLY INTO ESOPHAGUS, AND THE STAPLER WAS INTRODUCED THROUGH THORACOTOMY INTO THE GASTRIC POUCH. THE TWO WERE TO BE JOINED TOGETHER IN ORDER TO CREATE ANASTOMOSIS. THE SURGEON FOLLOWED THE SAME STEPS AS PRIOR SIMILAR PROCEDURES AND HAD VERY GOOD ACCESS AND ANGLE FOR CONNECTING. IT ALSO DID NOT APPEAR THAT ANY DAMAGE OCCURRED TO THE DEVICE DURING HANDLING, HOWEVER, THE TWO PARTS COULD NOT BE MATED TOGETHER. A SECOND STAPLER WAS OPENED TO SEE IF THEY COULD BE CONNECTED. WHEN THAT FAILED, THE ANVIL WAS REMOVED TRANSORALLY AND A SECOND ANVIL WAS INTRODUCED. THE DEVICES STILL COULD NOT BE MATED USING A REASONABLE FORCE. THE CASE WAS CONVERTED TO AN OPEN PROCEDURE SO THAT THE ANVIL AND THE STAPLER COULD BE JOINED USING HANDS INSTEAD OF DOING IT ENDOSCOPICALLY. THE PATIENT HAD AN ENLARGED THORACOTOMY INCISION. THE EVENT LEAD TO EXTENDED HOSPITALIZATION.
01-01-2018 | Malfunction | ethicon endo-surgery | 24-04-2018 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, JUST RECEIVED A PRODUCT COMPLAINT FOR THE DEVICE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-04-2018 | Malfunction | covidien lp llc north haven | 24-04-2018 | SIGNIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC SLEEVE GASTRECTOMY, THE TECHNICIAN TRIED TO LOAD THE STAPLE BACKWARDS AND BROKE THE ADAPTER SHAFT. THE SHAFT WAS LOCKED IN THE ¿CLOSE RELOAD¿ PHASE HOWEVER THERE WAS NO RELOAD ON THE SHAFT. THE FIRST DEVICE WAS ALSO UNABLE TO FIRE. THEY TRIED TO REMOVE THE ADAPTER AND RE-ATTACH BUT IT GAVE FAILURE INDICATOR. THEY THE OPENED ANOTHER STAPLER AND FINISHED THE CASE. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | us surgical puerto rico | 24-04-2018 | nan | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE DEVICE'S FIRING STOPPED ON THE HALFWAY. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | PWRD ECH FLEX 60MM 34CM SHAFT | Overheating of Device | A1005 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE. DURING THE VATS LOBECTOMY, COULD NOT FIRE WITHOUT MOTOR SOUND, THE BATTERY WAS HEATED WITH HIGH TEMPERATURE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE.
26-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE. DURING THE VATS LOBECTOMY, COULD NOT FIRE WITHOUT MOTOR SOUND, THE BATTERY WAS HEATED WITH HIGH TEMPERATURE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE.
27-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE AFTER SUCCESSFULLY COMPLETING A FIRING, THE GREEN PLASTIC OF THE RELOAD BROKE IN TWO PLACES AND THE PIECES FELL IN THE PATIENT. THE PIECES WERE RETRIEVED LAPAROSCOPICALLY. PROCEDURE WAS COMPLETED WITH THE SAME STAPLER. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE AFTER SUCCESSFULLY COMPLETING A FIRING, THE GREEN PLASTIC OF THE RELOAD BROKE IN TWO PLACES AND THE PIECES FELL IN THE PATIENT. THE PIECES WERE RETRIEVED LAPAROSCOPICALLY. PROCEDURE WAS COMPLETED WITH THE SAME STAPLER. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Component Falling |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE AFTER SUCCESSFULLY COMPLETING A FIRING, THE GREEN PLASTIC OF THE RELOAD BROKE IN TWO PLACES AND THE PIECES FELL IN THE PATIENT. THE PIECES WERE RETRIEVED LAPAROSCOPICALLY. PROCEDURE WAS COMPLETED WITH THE SAME STAPLER. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE THE DEVICE WENT THROUGH THE FIRING SEQUENCE, BOTH CUTTING AND STAPLING OCCURRED, HOWEVER THE BLADE WOULD NOT RETURN TO HOME POSITION. THE REVERSE BUTTON WAS TRIED AND IT DID NOT WORK. THE MANUAL OVER RIDE WAS USED TO BRING IT BACK. THE PROCEDURE WAS CONTINUED WITH AN ALTERNATE UNLIKE DEVICE. THE PROCEDURE WAS COMPLETED BY OVER SEWING THE INCISION. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
28-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Fitting Problem | A1208 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE THE DEVICE FIRED ON THICK TISSUE USING A GST60T RELOAD. AFTER RELOADING THE DEVICE AND TRYING TO PUT IT BACK DOWN THROUGH THE TROCAR IT WOULD NOT FIT DUE TO THE ANVIL BEING BENT. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
13-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Unintended Movement | A0512 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC BRICKER CYSTECTOMY, THE DEVICE WITH THE ADAPTER MADE UNPREDICTABLE MOTION TO THE LEFT AND RIGHT BEFORE CLAMPING THE VESSEL. THERE WAS NO PATIENT INJURY.
28-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | ILS 25MM, CURVED | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC COLECTOMY, FOUND IT WAS DIFFICULT TO DETACH THE ANVIL AND TROCAR, DETACHED THEM WITH BIG FORCE, BUT COULD NOT ASSEMBLE THEM. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
10-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, FOR THE FIRST FIRING A CRACK SOUND WAS HEARD FROM THE DEVICE AND IT COULD NOT BE FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, FOR THE FIRST FIRING A CRACK SOUND WAS HEARD FROM THE DEVICE AND IT COULD NOT BE FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE STAPLER DID NOT FIRE AND THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INVOLVEMENT.
11-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE STAPLER DID NOT FIRE AND THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INVOLVEMENT.
11-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: DURING LAP ASSISTED DISTAL GASTRECTOMY, A RELOAD WAS CONNECTED TO A HANDLE, HOWEVER COULD NOT OPEN/CLOSE THE JAWS. THE RELOAD WAS RECONNECTED BUT TO NO AVAIL. REPLACED BY A NEW RELOAD, THEY COULD OPEN/CLOSE THE JAWS. NO PATIENT INJURY.
28-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC DISTAL RADICAL GASTRECTOMY FOR GASTRIC CANCER, AFTER FIRED, FOUND THE TISSUE ONLY WITH HALF STAPLE LINE. SUTURE WAS USED TO SEW THE WOUND. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Device Slipped | A051204 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE USING TWO HANDLES WITH NEW RELOADS, THE SURGEONS HEARD A ¿CRUNCH¿ WHEN THE HANDLES WERE CLOSED TO FIRE OVER THE APPENDIX; THE HANDLES CLOSED, BUT WITHOUT FIRING. THEY HAVE OBSERVED THE LOADING, THE CYCLING, THE PLACEMENT, PRESSING THE FIRING BUTTON, AND THE ¿CRUNCH,¿ AFTER WHICH, THE HANDLE HAD NO RESISTANCE, AND STILL DID NOT FIRE. THE CASE WAS EXTENDED BUT IT WAS UNKNOWN FOR HOW LONG. A NEW RELOAD FROM THE SAME BOX AS THE FIRST THREE WERE TRIED AND FIRED. TWO OF WHICH RESPONDED THE SAME, WITH NO RESISTANCE ON THE HANDLE, AND WITH NO ¿CATCH.¿ THEY WERE ONLY PULLED CLOSED AFTER THE FIRING BUTTON WAS DEPRESSED. ON A FOURTH ATTEMPT, AN XL HANDLE AND ANOTHER NEW RELOAD FROM THE SAME BOX WAS USED, AND THIS TIME, WAS SUCCESSFUL.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE USING TWO HANDLES WITH NEW RELOADS, THE SURGEONS HEARD A ¿CRUNCH¿ WHEN THE HANDLES WERE CLOSED TO FIRE OVER THE APPENDIX; THE HANDLES CLOSED, BUT WITHOUT FIRING. THEY HAVE OBSERVED THE LOADING, THE CYCLING, THE PLACEMENT, PRESSING THE FIRING BUTTON, AND THE ¿CRUNCH,¿ AFTER WHICH, THE HANDLE HAD NO RESISTANCE, AND STILL DID NOT FIRE. THE CASE WAS EXTENDED BUT IT WAS UNKNOWN FOR HOW LONG. A NEW RELOAD FROM THE SAME BOX AS THE FIRST THREE WERE TRIED AND FIRED. TWO OF WHICH RESPONDED THE SAME, WITH NO RESISTANCE ON THE HANDLE, AND WITH NO ¿CATCH.¿ THEY WERE ONLY PULLED CLOSED AFTER THE FIRING BUTTON WAS DEPRESSED. ON A FOURTH ATTEMPT, AN XL HANDLE AND ANOTHER NEW RELOAD FROM THE SAME BOX WAS USED, AND THIS TIME, WAS SUCCESSFUL.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE USING TWO HANDLES WITH NEW RELOADS, THE SURGEONS HEARD A ¿CRUNCH¿ WHEN THE HANDLES WERE CLOSED TO FIRE OVER THE APPENDIX; THE HANDLES CLOSED, BUT WITHOUT FIRING. THEY HAVE OBSERVED THE LOADING, THE CYCLING, THE PLACEMENT, PRESSING THE FIRING BUTTON, AND THE ¿CRUNCH,¿ AFTER WHICH, THE HANDLE HAD NO RESISTANCE, AND STILL DID NOT FIRE. THE CASE WAS EXTENDED BUT IT WAS UNKNOWN FOR HOW LONG. A NEW RELOAD FROM THE SAME BOX AS THE FIRST THREE WERE TRIED AND FIRED. TWO OF WHICH RESPONDED THE SAME, WITH NO RESISTANCE ON THE HANDLE, AND WITH NO ¿CATCH.¿ THEY WERE ONLY PULLED CLOSED AFTER THE FIRING BUTTON WAS DEPRESSED. ON A FOURTH ATTEMPT, AN XL HANDLE AND ANOTHER NEW RELOAD FROM THE SAME BOX WAS USED, AND THIS TIME, WAS SUCCESSFUL.
17-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE GASTROINTESTINAL SURGERY, AFTER FIRED, FOUND THE TISSUE WITH HALF STAPLE LINE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-03-2018 | Injury | ethicon endo-surgery | 25-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY CASE ON (B)(6). HE HAS BEEN USING OUR PVS (PVE35A/(B)(4)) FOR 1 YEAR BUT HAD ISSUE OF MASSIVE BLEEDING FROM PV (PULMONARY VEIN) AFTER FIRING. ACTUALLY THERE WAS NO BLEEDING RIGHT AFTER TRANSECTION BUT THEY REALIZED MASSIVE BLEEDING WHEN CLOSING SKIN. THERE WAS BLEEDING FROM PV STUMP AND STUMP WAS OPENING.
12-03-2018 | Injury | ethicon endo-surgery | 25-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY CASE ON (B)(6). HE HAS BEEN USING OUR PVS (PVE35A/(B)(4)) FOR 1 YEAR BUT HAD ISSUE OF MASSIVE BLEEDING FROM PV (PULMONARY VEIN) AFTER FIRING. ACTUALLY THERE WAS NO BLEEDING RIGHT AFTER TRANSECTION BUT THEY REALIZED MASSIVE BLEEDING WHEN CLOSING SKIN. THERE WAS BLEEDING FROM PV STUMP AND STUMP WAS OPENING.
12-03-2018 | Injury | ethicon endo-surgery | 25-04-2018 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY CASE ON (B)(6). HE HAS BEEN USING OUR PVS (PVE35A/(B)(4)) FOR 1 YEAR BUT HAD ISSUE OF MASSIVE BLEEDING FROM PV (PULMONARY VEIN) AFTER FIRING. ACTUALLY THERE WAS NO BLEEDING RIGHT AFTER TRANSECTION BUT THEY REALIZED MASSIVE BLEEDING WHEN CLOSING SKIN. THERE WAS BLEEDING FROM PV STUMP AND STUMP WAS OPENING.
11-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING INTESTINAL RESECTION, THE PLASTIC PART OF ONE SIDE OF THE SURGICAL ENRANGER WAS STRIPPED FROM THE METAL PART. THE SURGEON USED ANOTHER STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE STAPLER DID NOT FIRE AND THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. THERE WAS NO PATIENT INVOLVEMENT.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HE COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INVOLVEMENT.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LYMPHADENECTOMY, WHILE TRANSECTION OF TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE WAS UNABLE TO FIRE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Device Slipped | A051204 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE USING TWO HANDLES WITH NEW RELOADS, THE SURGEONS HEARD A ¿CRUNCH¿ WHEN THE HANDLES WERE CLOSED TO FIRE OVER THE APPENDIX; THE HANDLES CLOSED, BUT WITHOUT FIRING. THEY HAVE OBSERVED THE LOADING, THE CYCLING, THE PLACEMENT, PRESSING THE FIRING BUTTON, AND THE ¿CRUNCH,¿ AFTER WHICH, THE HANDLE HAD NO RESISTANCE, AND STILL DID NOT FIRE. THE CASE WAS EXTENDED BUT IT WAS UNKNOWN FOR HOW LONG. A NEW RELOAD FROM THE SAME BOX AS THE FIRST THREE WERE TRIED AND FIRED. TWO OF WHICH RESPONDED THE SAME, WITH NO RESISTANCE ON THE HANDLE, AND WITH NO ¿CATCH.¿ THEY WERE ONLY PULLED CLOSED AFTER THE FIRING BUTTON WAS DEPRESSED. ON A FOURTH ATTEMPT, AN XL HANDLE AND ANOTHER NEW RELOAD FROM THE SAME BOX WAS USED, AND THIS TIME, WAS SUCCESSFUL.
13-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE USING TWO HANDLES WITH NEW RELOADS, THE SURGEONS HEARD A ¿CRUNCH¿ WHEN THE HANDLES WERE CLOSED TO FIRE OVER THE APPENDIX; THE HANDLES CLOSED, BUT WITHOUT FIRING. THEY HAVE OBSERVED THE LOADING, THE CYCLING, THE PLACEMENT, PRESSING THE FIRING BUTTON, AND THE ¿CRUNCH,¿ AFTER WHICH, THE HANDLE HAD NO RESISTANCE, AND STILL DID NOT FIRE. THE CASE WAS EXTENDED BUT IT WAS UNKNOWN FOR HOW LONG. A NEW RELOAD FROM THE SAME BOX AS THE FIRST THREE WERE TRIED AND FIRED. TWO OF WHICH RESPONDED THE SAME, WITH NO RESISTANCE ON THE HANDLE, AND WITH NO ¿CATCH.¿ THEY WERE ONLY PULLED CLOSED AFTER THE FIRING BUTTON WAS DEPRESSED. ON A FOURTH ATTEMPT, AN XL HANDLE AND ANOTHER NEW RELOAD FROM THE SAME BOX WAS USED, AND THIS TIME, WAS SUCCESSFUL.
13-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE USING TWO HANDLES WITH NEW RELOADS, THE SURGEONS HEARD A ¿CRUNCH¿ WHEN THE HANDLES WERE CLOSED TO FIRE OVER THE APPENDIX; THE HANDLES CLOSED, BUT WITHOUT FIRING. THEY HAVE OBSERVED THE LOADING, THE CYCLING, THE PLACEMENT, PRESSING THE FIRING BUTTON, AND THE ¿CRUNCH,¿ AFTER WHICH, THE HANDLE HAD NO RESISTANCE, AND STILL DID NOT FIRE. THE CASE WAS EXTENDED BUT IT WAS UNKNOWN FOR HOW LONG. A NEW RELOAD FROM THE SAME BOX AS THE FIRST THREE WERE TRIED AND FIRED. TWO OF WHICH RESPONDED THE SAME, WITH NO RESISTANCE ON THE HANDLE, AND WITH NO ¿CATCH.¿ THEY WERE ONLY PULLED CLOSED AFTER THE FIRING BUTTON WAS DEPRESSED. ON A FOURTH ATTEMPT, AN XL HANDLE AND ANOTHER NEW RELOAD FROM THE SAME BOX WAS USED, AND THIS TIME, WAS SUCCESSFUL.
13-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | IDRIVE ULTRA | Unintended Movement | A0512 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC BRICKER CYSTECTOMY, THE DEVICE WITH THE ADAPTER MADE UNPREDICTABLE MOTION TO THE LEFT AND RIGHT BEFORE CLAMPING THE VESSEL. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-06-2015 | Malfunction | us surgical (inactive | 25-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN USING THE GREEN LOAD THE STAPLER LOCKED.
15-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY, AFTER LOADING THE SINGLE USE LOAD UNIT TO THE STAPLER TO MANIPULATE ON THE STOMACH, THE SINGLE USE LOAD UNIT RESISTED TO CLOSE ON THE TISSUE AND THE JAWS WERE OPENED. AFTER TRYING TO RELOAD IT AGAIN, IT STILL RESISTED TO CLOSE THE JAWS. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT OUTCOME.
29-04-2014 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER THE LOAD DID NOT LOCK. NO INJURY REPORTED.
13-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, AT THE FIRST FIRING, THE SURGEON TRIED TO FIRE THE DEVICE HOWEVER A CRACK SOUND WAS HEARD AND COULD NOT FIRE THE DEVICE. THE SURGEON SAID TISSUE WAS THICK. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, AT THE FIRST FIRING, THE SURGEON TRIED TO FIRE THE DEVICE HOWEVER A CRACK SOUND WAS HEARD AND COULD NOT FIRE THE DEVICE. THE SURGEON SAID TISSUE WAS THICK. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | nan | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HE COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INVOLVEMENT.
03-10-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | TA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOTOMY PROCEDURE, WHILE PERFORMING PULMONARY LOBAR RESECTION, THE DEVICE PARTIALLY CLOSED THE LOBAR BRONCHUS. THE UPPER PART WAS CLOSED BUT NOT THE LOWER PART OF THE BRONCHUS. THERE WAS OPENING AND THE SURGEON HAD TO SUTURE TO CLOSE. THERE WAS NO PATIENT INJURY.
03-10-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | TA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOTOMY PROCEDURE, WHILE PERFORMING PULMONARY LOBAR RESECTION, THE DEVICE PARTIALLY CLOSED THE LOBAR BRONCHUS. THE UPPER PART WAS CLOSED BUT NOT THE LOWER PART OF THE BRONCHUS. THERE WAS OPENING AND THE SURGEON HAD TO SUTURE TO CLOSE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, AT THE FIRST FIRING, THE SURGEON TRIED TO FIRE THE DEVICE HOWEVER A CRACK SOUND WAS HEARD AND COULD NOT FIRE THE DEVICE. THE SURGEON SAID TISSUE WAS THICK. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, AT THE FIRST FIRING, THE SURGEON TRIED TO FIRE THE DEVICE HOWEVER A CRACK SOUND WAS HEARD AND COULD NOT FIRE THE DEVICE. THE SURGEON SAID TISSUE WAS THICK. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
07-04-2018 | Injury | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING, STAPLES WERE LOOSE. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING, RESULTING TO EXTENDED HOSPITALIZATION.
07-04-2018 | Injury | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING, STAPLES WERE LOOSE. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING, RESULTING TO EXTENDED HOSPITALIZATION.
09-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC PROCEDURE, THE STAPLE LINE WAS POOR AND THE STAPLES WERE FALLING OUT OF THE RELOAD. A NEW STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC PROCEDURE, THE STAPLE LINE WAS POOR AND THE STAPLES WERE FALLING OUT OF THE RELOAD. A NEW STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | EEA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SIGMOIDECTOMY AFTER THE ANASTOMOSIS, THE SURGEON REMOVED STAPLER FROM THE PATIENT, HOWEVER THE ANVIL WAS DISENGAGED FROM THE DEVICE AND STAYED AT INTESTINAL CANAL. THE SURGEON REMOVED THE ANVIL BY ENDOSCOPE AND CONFIRMED ANASTOMOSED SITE WAS FINE. THE SALES REP VISITED THE SURGERY ROOM AND CONFIRMED THE DEVICE IN QUESTION WAS NO PROBLEM IN APPEARANCE. DONUTS WERE COMPLETED. NO DEFORMATION WAS NOTED IN ANVIL SIDE. AFTER THE SURGERY, THE SALES REP INTERVIEWED TO THE SURGEON. THE SURGEON OBSERVED THAT THERE WAS NO FOREIGN MATERIAL WHEN HE/SHE CONNECTED THE ANVIL TO THE TROCAR. THE SURGEON ROTATED TWIST KNOB ONLY TWICE. THE SURGEON ALSO OBSERVED A BIG RESISTANCE AT REMOVING THE DEVICE FROM THE PATIENT. THE PATIENT HAD HEAVY DIVERTICULAR DISEASE AND INTESTINAL CANAL WAS VERY THICKENED. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN.
10-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, FOR THE FIRST FIRING A CRACK SOUND WAS HEARD FROM THE DEVICE AND IT COULD NOT BE FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, FOR THE FIRST FIRING A CRACK SOUND WAS HEARD FROM THE DEVICE AND IT COULD NOT BE FIRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
05-03-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THEY HAD A STAPLER MISFIRE LEADING TO A TEAR IN THE INTESTINAL WALL. IN ORDER TO RESOLVE THE ISSUE AND COMPLETE THE CASE THE SURGEON SUTURED THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
05-03-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THEY HAD A STAPLER MISFIRE LEADING TO A TEAR IN THE INTESTINAL WALL. IN ORDER TO RESOLVE THE ISSUE AND COMPLETE THE CASE THE SURGEON SUTURED THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE RESECTION OF LEFT UPPER LOBE ARTERY, THE DEVICE WAS ABLE TO FIRE, BUT IT HAD AN INCOMPLETE STAPLE LINE AT THE DISTAL PART. TO RESOLVE THE ISSUE, THEY HAND STITCHED THE BLOOD VESSELS. THERE WAS NO PATIENT INJURY.
07-04-2018 | Injury | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, AFTER THE INTERVENTION FOR A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING POST-OPERATIVELY. THE PATIENT HAD BLEEDING FROM THE SECTION. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING RESULTING TO EXTENDED HOSPITALIZATION.
07-04-2018 | Injury | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER THE INTERVENTION FOR A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING POST-OPERATIVELY. THE PATIENT HAD BLEEDING FROM THE SECTION. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING RESULTING TO EXTENDED HOSPITALIZATION.
15-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY, AFTER LOADING THE SINGLE USE LOAD UNIT TO THE STAPLER TO MANIPULATE ON THE STOMACH, THE SINGLE USE LOAD UNIT RESISTED TO CLOSE ON THE TISSUE AND THE JAWS WERE OPENED. AFTER TRYING TO RELOAD IT AGAIN, IT STILL RESISTED TO CLOSE THE JAWS. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT OUTCOME.
08-02-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, SECTION CURSOR BROKE DURING STAPLING. NO FURTHER INFORMATION PROVIDED. NO INFORMATION ABOUT CONSEQUENCE FOR ANY PATIENT.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LOWER LOBE WEDGE RESECTION, WHILE THEY USED THE PRODUCT TO CUT LUNG TISSUE AND AFTER PRESSING THE GREEN BUTTON, IT WAS FOUND THAT THE STAPLER COULD NOT BE ACTIVATED AFTER CLOSING THE STAPLER. THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THEY REPLACED THE STAPLE CARTRIDGES OF THE SAME SIZE DURING SURGERY TO ENSURE CONTINUED OPERATION. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | MULTIFIRE ENDO GIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE WAS LOCKED UP WHEN THEY TRIED TO USE IT. THERE WAS NO PATIENT INVOLVEMENT.
25-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE WAS NOT ABLE TO FIRE AND ITS HANDLE COULD NOT BE SQUEEZED BY THE SURGEON.
07-04-2018 | Injury | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure To Adhere Or Bond |  | Blood Loss | Event Description: ACCORDING TO THE REPORTER, AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING, STAPLES WERE LOOSE. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING WAS DONE RESULTING TO EXTENDED HOSPITALIZATION.
07-04-2018 | Injury | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure To Adhere Or Bond |  | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING, STAPLES WERE LOOSE. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING WAS DONE RESULTING TO EXTENDED HOSPITALIZATION.
07-04-2018 | Injury | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING, STAPLES WERE LOOSE. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING WAS DONE RESULTING TO EXTENDED HOSPITALIZATION.
07-04-2018 | Injury | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING, STAPLES WERE LOOSE. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING WAS DONE RESULTING TO EXTENDED HOSPITALIZATION.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LOWER LOBE WEDGE RESECTION, WHILE THEY USED THE PRODUCT TO CUT LUNG TISSUE AND AFTER PRESSING THE GREEN BUTTON, IT WAS FOUND THAT THE STAPLER COULD NOT BE ACTIVATED AFTER CLOSING THE STAPLER. THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THEY REPLACED THE STAPLE CARTRIDGES OF THE SAME SIZE DURING SURGERY TO ENSURE CONTINUED OPERATION. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Failure to Fire | A050501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, AFTER FIRING THE DEVICE AT THE MOUTH SIDE OF COLON, THE CARTRIDGE COULD NOT BE REMOVED FROM THE STAPLER. THE CARTRIDGE WAS FINALLY REMOVED AFTER CONTINUOUS ATTEMPTS. A NEW CARTRIDGE WAS LOADED UNTO THE STAPLER HOWEVER, UPON APPLYING UNTO THE ANAL SIDE OF THE COLON, THE DEVICE COULD NOT FIRE. THERE WAS TISSUE DAMAGE. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS EXTENDED BY LESS THAN 30 MINUTES HOWEVER, THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, AFTER FIRING THE DEVICE AT THE MOUTH SIDE OF COLON, THE CARTRIDGE COULD NOT BE REMOVED FROM THE STAPLER. THE CARTRIDGE WAS FINALLY REMOVED AFTER CONTINUOUS ATTEMPTS. A NEW CARTRIDGE WAS LOADED UNTO THE STAPLER HOWEVER, UPON APPLYING UNTO THE ANAL SIDE OF THE COLON, THE DEVICE COULD NOT FIRE. THERE WAS TISSUE DAMAGE. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS EXTENDED BY LESS THAN 30 MINUTES HOWEVER, THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Human-Device Interface Problem | A22 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, AFTER FIRING THE DEVICE AT THE MOUTH SIDE OF COLON, THE CARTRIDGE COULD NOT BE REMOVED FROM THE STAPLER. THE CARTRIDGE WAS FINALLY REMOVED AFTER CONTINUOUS ATTEMPTS. A NEW CARTRIDGE WAS LOADED UNTO THE STAPLER HOWEVER, UPON APPLYING UNTO THE ANAL SIDE OF THE COLON, THE DEVICE COULD NOT FIRE. THERE WAS TISSUE DAMAGE. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS EXTENDED BY LESS THAN 30 MINUTES HOWEVER, THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, AFTER FIRING THE DEVICE AT THE MOUTH SIDE OF COLON, THE CARTRIDGE COULD NOT BE REMOVED FROM THE STAPLER. THE CARTRIDGE WAS FINALLY REMOVED AFTER CONTINUOUS ATTEMPTS. A NEW CARTRIDGE WAS LOADED UNTO THE STAPLER HOWEVER, UPON APPLYING UNTO THE ANAL SIDE OF THE COLON, THE DEVICE COULD NOT FIRE. THERE WAS TISSUE DAMAGE. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS EXTENDED BY LESS THAN 30 MINUTES HOWEVER, THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS PROCEDURE, THE DOCTOR LOADED 30 RELOAD AND IT WOULD NOT FIRE. IT WAS REMOVED FROM THE DEVICE¿S HANDLE AND ANOTHER RELOAD WAS INSERTED. THE SECOND DEVICE FUNCTIONED APPROPRIATELY. THE DEFECTIVE RELOAD WAS ISOLATED, CLEANED AND GIVEN TO THE OR COORDINATOR FOR FOLLOW UP. THERE WAS NO PATIENT INJURY.
12-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING ACROSS THE LUNG TISSUE ON A LAPAROSCOPIC WEDGE RESECTION (THORACIC), THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE. THEY HAD A 60MM BLACK RELOAD ON THE HANDLE AND THE RELOAD WOULD NOT FIRE ACROSS THE LUNG TISSUE. THEY OPENED ANOTHER RELOAD BUT THEY HAD THE SAME RESULT. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLES MALFORMED AFTER FIRING. DURING THE SLEEVE GASTRECTOMY, AFTER FIRED, FOUND THE STAPLES MALFORMED, USED SUTURE TO SEW THE WOUND. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
30-03-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RESECTION OF SIGMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE AND ITS HANDLE COULD NOT BE SQUEEZED. AFTER THEY CORRECTLY MOUNTED THE DEVICE, IT DID NOT WORK. AFTER REMOVING THE SAFETY, THE BLADE DID NOT ADVANCE, ALSO THE HANDLE DID NOT MOVE ON. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THEY CHANGED TO A NEW DEVICE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING RESECTION OF SIGMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE AND ITS HANDLE COULD NOT BE SQUEEZED. AFTER THEY CORRECTLY MOUNTED THE DEVICE, IT DID NOT WORK. AFTER REMOVING THE SAFETY, THE BLADE DID NOT ADVANCE, ALSO THE HANDLE DID NOT MOVE ON. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THEY CHANGED TO A NEW DEVICE TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AFTER FIRE, SURGEON FOUND STAPLES SHAPE WAS NOT COMPLETED FOR ALL STAPLE LINE.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, AFTER FIRING THE DEVICE AT THE MOUTH SIDE OF COLON, THE CARTRIDGE COULD NOT BE REMOVED FROM THE STAPLER. THE CARTRIDGE WAS FINALLY REMOVED AFTER CONTINUOUS ATTEMPTS. A NEW CARTRIDGE WAS LOADED UNTO THE STAPLER HOWEVER, UPON APPLYING UNTO THE ANAL SIDE OF THE COLON, THE DEVICE COULD NOT FIRE. THERE WAS TISSUE DAMAGE. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS EXTENDED BY LESS THAN 30 MINUTES HOWEVER, THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, AFTER FIRING THE DEVICE AT THE MOUTH SIDE OF COLON, THE CARTRIDGE COULD NOT BE REMOVED FROM THE STAPLER. THE CARTRIDGE WAS FINALLY REMOVED AFTER CONTINUOUS ATTEMPTS. A NEW CARTRIDGE WAS LOADED UNTO THE STAPLER HOWEVER, UPON APPLYING UNTO THE ANAL SIDE OF THE COLON, THE DEVICE COULD NOT FIRE. THERE WAS TISSUE DAMAGE. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS EXTENDED BY LESS THAN 30 MINUTES HOWEVER, THERE WAS NO PATIENT INJURY.
07-04-2018 | Injury | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, AFTER THE INTERVENTION FOR A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING. THE PATIENT HAD BLEEDING FROM THE SECTION. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING RESULTING TO EXTENDED HOSPITALIZATION.
07-04-2018 | Injury | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER THE INTERVENTION FOR A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON INTERVENED FOR BLEEDING. THE PATIENT HAD BLEEDING FROM THE SECTION. A NEW INTERVENTION WAS DONE TO STOP THE BLEEDING RESULTING TO EXTENDED HOSPITALIZATION.
11-04-2018 | Injury | us surgical puerto rico | 25-04-2018 | EEA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, WHEN PERFORMING END TO END ANASTOMOSIS, THE DEVICE PARTIALLY FIRED, THE ANVIL CAN BE TILTED AFTER FIRING AND WAS NOT BENT. AS STATED, THERE WAS AN UNANTICIPATED TISSUE LOSS AS THEY NEEDED TO DO ANOTHER ANASTOMOSIS. THEY SPENT MORE THAN ONE HOUR TO DO ANOTHER ANASTOMOSIS, SUTURES. THEY USED ANOTHER STAPLER TO COMPLETE THE CASE.
11-04-2018 | Injury | us surgical puerto rico | 25-04-2018 | EEA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, WHEN PERFORMING END TO END ANASTOMOSIS, THE DEVICE PARTIALLY FIRED, THE ANVIL CAN BE TILTED AFTER FIRING AND WAS NOT BENT. AS STATED, THERE WAS AN UNANTICIPATED TISSUE LOSS AS THEY NEEDED TO DO ANOTHER ANASTOMOSIS. THEY SPENT MORE THAN ONE HOUR TO DO ANOTHER ANASTOMOSIS, SUTURES. THEY USED ANOTHER STAPLER TO COMPLETE THE CASE.
11-04-2018 | Injury | us surgical puerto rico | 25-04-2018 | EEA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, WHEN PERFORMING END TO END ANASTOMOSIS, THE DEVICE PARTIALLY FIRED, THE ANVIL CAN BE TILTED AFTER FIRING AND WAS NOT BENT. AS STATED, THERE WAS AN UNANTICIPATED TISSUE LOSS AS THEY NEEDED TO DO ANOTHER ANASTOMOSIS. THEY SPENT MORE THAN ONE HOUR TO DO ANOTHER ANASTOMOSIS, SUTURES. THEY USED ANOTHER STAPLER TO COMPLETE THE CASE.
05-03-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THEY HAD A STAPLER MISFIRE LEADING TO A TEAR IN THE INTESTINAL WALL. IN ORDER TO RESOLVE THE ISSUE AND COMPLETE THE CASE THE SURGEON SUTURED THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
05-03-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THEY HAD A STAPLER MISFIRE LEADING TO A TEAR IN THE INTESTINAL WALL. IN ORDER TO RESOLVE THE ISSUE AND COMPLETE THE CASE THE SURGEON SUTURED THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS, THE DEVICE'S DID NOT FIRE, HOWEVER SHE NOTED STAPLES WERE ALREADY APPEARED FROM THE STAPLE POCKETS ACROSS 1CM FROM PROXIMAL END, BUT THEY COULD NOT SQUEEZE THE HANDLE, AND PRE-FIRED STAPLES WERE COMPLETELY B-FORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS, THE DEVICE'S DID NOT FIRE, HOWEVER SHE NOTED STAPLES WERE ALREADY APPEARED FROM THE STAPLE POCKETS ACROSS 1CM FROM PROXIMAL END, BUT THEY COULD NOT SQUEEZE THE HANDLE, AND PRE-FIRED STAPLES WERE COMPLETELY B-FORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS, THE DEVICE'S DID NOT FIRE, HOWEVER SHE NOTED STAPLES WERE ALREADY APPEARED FROM THE STAPLE POCKETS ACROSS 1CM FROM PROXIMAL END, BUT THEY COULD NOT SQUEEZE THE HANDLE, AND PRE-FIRED STAPLES WERE COMPLETELY B-FORMED. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS, THE DEVICE'S DID NOT FIRE, HOWEVER SHE NOTED STAPLES WERE ALREADY APPEARED FROM THE STAPLE POCKETS ACROSS 1CM FROM PROXIMAL END, BUT THEY COULD NOT SQUEEZE THE HANDLE, AND PRE-FIRED STAPLES WERE COMPLETELY B-FORMED. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | EGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHILE STAPLING THE INTESTINAL LOOP (BY-PASS), WHEN FIRING THE DEVICE, IT MADE A NOISE AND THE TRIGGER BECAME TOTALLY LIMP NOT FIRING THE LOAD. THE STAPLER WAS REPLACED BY A NEW ONE AND THE PROBLEM WAS SOLVED. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS, THE SURGEON CLAMPED THE COLON OF MOUTH SIDE AND TRIED TO PUSH THE FIRING KNOB, HOWEVER COULD NOT FIRE THE DEVICE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
29-03-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | GST BLACK RELOAD, 60MM, 6 ROW | Component Missing | A020602 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A TISSUE LAB, THE STAPLE DRIVERS WERE MISSING AND WHEN THE DEVICE WAS FIRED NO STAPLES DEPLOYED.
12-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE WAS NOT ABLE TO FIRE AND ITS HANDLE COULD NOT BE SQUEEZED BY THE SURGEON.
30-03-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RESECTION OF SIGMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE AND ITS HANDLE COULD NOT BE SQUEEZED BY THE SURGEON. AFTER THEY CORRECTLY MOUNTED THE DEVICE, IT DID NOT WORK. AFTER REMOVING THE SAFETY, THE BLADE DID NOT ADVANCE, ALSO THE HANDLE DID NOT MOVE ON. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THEY CHANGED A NEW DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-03-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDOGIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING RESECTION OF SIGMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE AND ITS HANDLE COULD NOT BE SQUEEZED BY THE SURGEON. AFTER THEY CORRECTLY MOUNTED THE DEVICE, IT DID NOT WORK. AFTER REMOVING THE SAFETY, THE BLADE DID NOT ADVANCE, ALSO THE HANDLE DID NOT MOVE ON. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THEY CHANGED A NEW DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-03-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | SIGNIA | Device Inoperable |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DEVICE FOLLOW-UP, THE DEVICE SHOWED THE SYMBOL HANDLE ERROR IN THE OLED. THE RED LED ALSO FLASHED IN THE CHARGING STATION. THE STAPLER WAS ALSO UNABLE TO FIRE. THEY TRIED USING DIFFERENT HANDLE. HOWEVER, THE SECOND HANDLE EVEN WHEN FULLY CHARGED, IT TURNS OFF AFTER THE CONNECTION WITH THE SHELL, THE ADAPTER AND THE RELOAD, THE DISPLAY TURNS OFF AND WAS BLACK. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE RESECTION OF LEFT UPPER LOBE ARTERY, THE DEVICE WAS ABLE TO FIRE, BUT IT HAD AN INCOMPLETE STAPLE LINE AT THE DISTAL PART. TO RESOLVE THE ISSUE, THEY HAND STITCHED THE BLOOD VESSELS. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHILE STAPLING THE INTESTINAL LOOP (BY-PASS), WHEN FIRING THE DEVICE, IT MADE A NOISE AND THE TRIGGER BECAME TOTALLY LIMP NOT FIRING THE LOAD. THE STAPLER WAS REPLACED BY A NEW ONE AND THE PROBLEM WAS SOLVED. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHILE STAPLING THE INTESTINAL LOOP (BY-PASS), WHEN FIRING THE DEVICE, IT MADE A NOISE AND THE TRIGGER BECAME TOTALLY LIMP NOT FIRING THE LOAD. THE STAPLER WAS REPLACED BY A NEW ONE AND THE PROBLEM WAS SOLVED. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY, THE STAPLER COULD NOT FIRE HOWEVER, THE BLACK PUSHER THUMB PIECE COULD BE MOVED. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | ECHELON FLEX 45 LONG | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC NEPHRECTOMY PROCEDURE THE DEVICE FULLY FIRED ON THE THIRD FIRING WITH A WHITE RELOAD BUT THE BLADE DID NOT FULLY RETURN AND THE LOCK APPEARED IN THE CYCLE WINDOW. ATTEMPTED TO ACTIVATE REVERSE SWITCH, AND NOTHING WOULD MOVE IN THE HANDLE. THEREFORE, JAWS COULD NOT BE OPENED. SURGEONS ENLARGED LAPAROTOMY SLIGHTLY TO CLAMP AND LIGATE VESSEL BY OTHER MEANS (DISTAL FIRING OF ANOTHER LIKE DEVICE AND RELOAD) AND THEN PRY THE JAWS OF THE FIRST ECHELON OFF OF THE TISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | ECHELON FLEX 45 LONG | Retraction Problem | A0510 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC NEPHRECTOMY PROCEDURE THE DEVICE FULLY FIRED ON THE THIRD FIRING WITH A WHITE RELOAD BUT THE BLADE DID NOT FULLY RETURN AND THE LOCK APPEARED IN THE CYCLE WINDOW. ATTEMPTED TO ACTIVATE REVERSE SWITCH, AND NOTHING WOULD MOVE IN THE HANDLE. THEREFORE, JAWS COULD NOT BE OPENED. SURGEONS ENLARGED LAPAROTOMY SLIGHTLY TO CLAMP AND LIGATE VESSEL BY OTHER MEANS (DISTAL FIRING OF ANOTHER LIKE DEVICE AND RELOAD) AND THEN PRY THE JAWS OF THE FIRST ECHELON OFF OF THE TISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | ECHELON FLEX 45 LONG | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC NEPHRECTOMY PROCEDURE THE DEVICE FULLY FIRED ON THE THIRD FIRING WITH A WHITE RELOAD BUT THE BLADE DID NOT FULLY RETURN AND THE LOCK APPEARED IN THE CYCLE WINDOW. ATTEMPTED TO ACTIVATE REVERSE SWITCH, AND NOTHING WOULD MOVE IN THE HANDLE. THEREFORE, JAWS COULD NOT BE OPENED. SURGEONS ENLARGED LAPAROTOMY SLIGHTLY TO CLAMP AND LIGATE VESSEL BY OTHER MEANS (DISTAL FIRING OF ANOTHER LIKE DEVICE AND RELOAD) AND THEN PRY THE JAWS OF THE FIRST ECHELON OFF OF THE TISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 25-04-2018 | ECHELON FLEX 45 LONG | Difficult to Open or Close | A051104 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC NEPHRECTOMY PROCEDURE THE DEVICE FULLY FIRED ON THE THIRD FIRING WITH A WHITE RELOAD BUT THE BLADE DID NOT FULLY RETURN AND THE LOCK APPEARED IN THE CYCLE WINDOW. ATTEMPTED TO ACTIVATE REVERSE SWITCH, AND NOTHING WOULD MOVE IN THE HANDLE. THEREFORE, JAWS COULD NOT BE OPENED. SURGEONS ENLARGED LAPAROTOMY SLIGHTLY TO CLAMP AND LIGATE VESSEL BY OTHER MEANS (DISTAL FIRING OF ANOTHER LIKE DEVICE AND RELOAD) AND THEN PRY THE JAWS OF THE FIRST ECHELON OFF OF THE TISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING ACROSS THE LUNG TISSUE ON A LAPAROSCOPIC WEDGE RESECTION (THORACIC), THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS UNABLE TO FIRE. THEY HAD A 60MM BLACK RELOAD ON THE HANDLE AND THE RELOAD WOULD NOT FIRE ACROSS THE LUNG TISSUE. THEY OPENED ANOTHER RELOAD BUT THEY HAD THE SAME RESULT. OPENED ANOTHER DEVICE AND WAS ABLE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-01-2015 | Malfunction | us surgical puerto rico | 25-04-2018 | EEA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE PROCEDURE (RECTOSIGMOIDECTOMY) WHEN THE COLON AND RECTUM WERE CONNECTED AND THE GREEN SIGN WAS RELEASED IT WAS DECIDED TO REGRESS BECAUSE IT WAS VERIFIED ISCHEMIA OF THE COLON. SO WHEN THE REGRESSION WAS PERFORMED THE WARHEAD ROTATED AND IT WAS IMPOSSIBLE TO PERFORM THE STAPLING. IT WAS NECESSARY TO USE A NEW STAPLER TO FINISH THE PROCEDURE.
16-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LYMPHADENECTOMY, WHILE TRANSECTION OF TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE WAS UNABLE TO FIRE. THEY OPENED A NEW STAPLER AND RELOAD TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-05-2014 | Malfunction | us surgical puerto rico | 25-04-2018 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE STAPLES DID NOT CLOSE PROPERLY ON THE PATIENTS STOMACH, LETTING A PIECE TOTALLY OPENED AND IT WAS NECESSARY TO PERFORM MANUAL SUTURE. NO INJURY REPORTED.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | TA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A KIDNEY TRANSPLANT PROCEDURE, WHILE ON THE LIGATION OF RENAL VEIN, THE OPERATOR PERFORMED FULL AND COMPLETE SQUEEZES OF THE HANDLE, THE HANDLE RETURNED TO THE OPEN POSITION FOLLOWING THE FIRST FULL SQUEEZE OF THE HANDLE. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED FOR THE SECOND TIME TO FIRE. A NEW OPEN STAPLER WAS OPENED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC TUMOR RESECTION, PRE-OPERATIVELY THERE WAS A PROBLEM LOADING THE DEVICE, RECHARGING OF THE DEVICE DID NOT WORK. WHEN THEY ASSEMBLED THE INSTRUMENT, IT DID NOT OPEN OR CLOSE THE JAWS, FOR SUCH REASON, IT COULD NOT BE USED. THE RELOAD ALSO COULD NOT BE CLOSED BEFORE ENTERING THE TROCAR. THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
02-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC TUMOR RESECTION, PRE-OPERATIVELY THERE WAS A PROBLEM LOADING THE DEVICE, RECHARGING OF THE DEVICE DID NOT WORK. WHEN THEY ASSEMBLED THE INSTRUMENT, IT DID NOT OPEN OR CLOSE THE JAWS, FOR SUCH REASON, IT COULD NOT BE USED. THE RELOAD ALSO COULD NOT BE CLOSED BEFORE ENTERING THE TROCAR. THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
02-04-2018 | Malfunction | covidien lp llc north haven | 25-04-2018 | IDRIVE ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC TUMOR RESECTION, PRE-OPERATIVELY THERE WAS A PROBLEM LOADING THE DEVICE, RECHARGING OF THE DEVICE DID NOT WORK. WHEN THEY ASSEMBLED THE INSTRUMENT, IT DID NOT OPEN OR CLOSE THE JAWS, FOR SUCH REASON, IT COULD NOT BE USED. THE RELOAD ALSO COULD NOT BE CLOSED BEFORE ENTERING THE TROCAR. THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
11-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE STAPLING OF THE STOMACH ON A BARIATRIC SURGERY, THE DEVICE DID NOT FIRE. THEY REPLACED THE DEVICE WITH A NEW STAPLER AND NEW LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | us surgical puerto rico | 25-04-2018 | GIA | Detachment Of Device Component |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A DISTAL GASTRECTOMY, THEY OPENED THE PACKAGE AND TOOK THE DEVICE OUT, HOWEVER, THE FIRING KNOB WAS BROKEN AND PART CAME OFF THE DEVICE. THEY HAD TO REPLACE WITH A NEW DEVICE FOR THEM TO COMPLETE THE CASE. THE PATIENT HAS NO INJURY.
27-03-2018 | Malfunction | ethicon endo-surgery | 26-04-2018 | POWERED 60 ECHELON +, 280MM SHAFT | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN PNEUMONECTOMY, THE ARTICULATION JOINT COVER WAS BROKEN OFF WHEN THE DEVICE WAS USED ON THE LUNG. IT WAS TOO HARD TO APPROACH THE TARGET TISSUE SINCE THE INTERCOSTAL SPACE WAS NARROW. NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-01-2018 | Malfunction | ethicon endo-surgery | 26-04-2018 | POWERED 60 ECHELON +, 440MM SHAFT | Device Operates Differently Than Expected |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STOMACH TISSUE SLIPPED OUT OF THE STAPLER DURING FIRING SEQUENCE SO TISSUE WAS NOT STAPLED WELL.
27-03-2018 | Malfunction | ethicon endo-surgery | 26-04-2018 | ECHELON FLEX | Improper or Incorrect Procedure or Method | A2303 | No Consequences Or Impact To Patient | Event Description: DURING THE CASE, THE 45 MM ECHELON FLEX POWERED PLUS STAPLER ARTICULATING ENDOSCOPIC LINEAR CUTTER WITH 45 MM GREEN LOAD MISFIRED DURING THE CASE. MD HANDED THE STAPLER OFF THE FIELD AND IT WAS BAGGED IN A RED BAG AND GAVE IT TO THE CONTROL DESK WHO WILL GET IT TO THE SAFETY OFFICER. NO HARM CAME TO THE PATIENT AT THIS TIME. NEW STAPLER WAS OPENED UP TO THE FIELD AND CASE WAS COMPLETED WITHOUT COMPLICATIONS. A FEW WEEKS LATER THE SALES REP. VISITED THE SAFETY OFFICE AND LOOKED OVER DEVICE AND TOLD ME THAT IT LOOKED LIKE THE DEVICES STAPLES WERE LOADED WRONG AND POSSIBLY CAUSING THE MISFIRE.
27-03-2018 | Malfunction | ethicon endo-surgery | 26-04-2018 | ECHELON FLEX | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: DURING THE CASE, THE 45 MM ECHELON FLEX POWERED PLUS STAPLER ARTICULATING ENDOSCOPIC LINEAR CUTTER WITH 45 MM GREEN LOAD MISFIRED DURING THE CASE. MD HANDED THE STAPLER OFF THE FIELD AND IT WAS BAGGED IN A RED BAG AND GAVE IT TO THE CONTROL DESK WHO WILL GET IT TO THE SAFETY OFFICER. NO HARM CAME TO THE PATIENT AT THIS TIME. NEW STAPLER WAS OPENED UP TO THE FIELD AND CASE WAS COMPLETED WITHOUT COMPLICATIONS. A FEW WEEKS LATER THE SALES REP. VISITED THE SAFETY OFFICE AND LOOKED OVER DEVICE AND TOLD ME THAT IT LOOKED LIKE THE DEVICES STAPLES WERE LOADED WRONG AND POSSIBLY CAUSING THE MISFIRE.
26-03-2018 | Injury | incisive surgical | 26-04-2018 | INSORB SUBCUTICULAR STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: A PATIENT EXPERIENCED A WOUND SEPARATION POST CESAREAN SECTION. THE PATIENT WAS TREATED WITH SUTURE IN SURGERY/OPERATING ROOM.
11-04-2018 | Malfunction | ethicon endo-surgery | 26-04-2018 | ECHELON FLEX | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: DURING THE ROBOT PROSTATECTOMY ECHELON FLEX 45 MM STAPLER WITH ECR45G GREEN RELOAD WAS USED. THE STAPLER MISFIRED. MISFIRING DID NOT CAUSE ANY HARM TO THE PATIENT. SALES REPRESENTATIVE FOR ETHICON WAS CALLED IN TO REVIEW AND BROUGHT DEVICE TO SAFETY. SALES REP AND SAFETY OFFICER LOOKED OVER OBTAINED DEVICE AND SALES REP. OPINION WAS THAT DEVICE DID MALFUNCTION AND FAILED TO FIRE.
03-04-2018 | Malfunction | ethicon endo-surgery | 26-04-2018 | PROXIMATE | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: DURING USE OF DEVICE BY SURGEON FOR AN EXPLORATORY LAPAROTOMY, THE DEVICE STAPLES WERE NOT FORMING CORRECTLY. REMOVED DEVICE FROM THE FIELD AND REPLACED WITH SAME DEVICE, DIFFERENT LOT NUMBER AND NO ISSUES. NO PATIENT HARM.
26-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | EEA | Failure To Adhere Or Bond |  | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS ON A LAPAROSCOPIC ESOPHAGUS PROCEDURE, THE DEVICE WAS DIFFICULT OR UNABLE TO CLOSE AND THE HEAD OF THE DEVICE WAS NOT ABLE TO CONNECT WITH THE STAPLER. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | EGIAUSTND | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON PROCEDURE, AFTER ONE NORMAL FIRING, DEVICE REFUSED TO FIRE. A NEW INSTRUMENT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL GASTRECTOMY PROCEDURE, ON THE 4TH FIRINGS WITH RELOAD WERE SUCCESSFUL. FOR THE 5TH FIRING, THEY CONNECTED RELOAD TO THE SAME HANDLE AND CHECKED THE JAWS OPENING/CLOSING. THE SURGEON CLAMPED THE TISSUE AND TRIED TO FIRE THE DEVICE, HOWEVER THE KNIFE DID NOT ADVANCE. THE SURGEON DID NOT MAKE AN IMPACT TO THE DEVICE BEFORE THE FIRING WAS STARTED AND DID NOT TOUCH THE HANDLE UNNECESSARILY AFTER THE GREEN BUTTON WAS PUSHED. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON PROCEDURE, WHILE CUTTING THE COLON, THE STAPLER WAS UNABLE TO FIRE BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. WHEN THE SURGEON USED THE LOAD, THERE WAS ONLY A PARTIAL STAPLE OF THREE OF THE RELOADS, SO THE TISSUE WAS A LITTLE DAMAGED. ALSO FOUND OUT THAT THE PROXIMAL STAPLE FORMATION WAS POOR AND THE DISTAL STAPLE LINE WAS INCOMPLETE. FIRST THEY USED A STAPLER AND RELOAD BUT IT DID NOT WORK, SO HE TOOK AN OTHER LOAD AND CHANGED TO ANOTHER DEVICE BUT THEY FORGOT THAT THEY NEEDED A 15MM. TROCAR, SO THEY USED ANOTHER STAPLER AND IT WAS OKAY.
13-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIGESTIVE PROCEDURE, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, ALTHOUGH HE WAS ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
13-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIGESTIVE PROCEDURE, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, ALTHOUGH HE WAS ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
05-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, WHEN THE VASCULAR LOAD WAS USED TO CLOSE OFF RIGHT LOBE, THE DEVICE DID NOT CLOSE THE SURGICAL SITE PROPERLY. THERE WAS NO PATIENT INJURY.
10-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON PROCEDURE, WHILE CUTTING THE COLON, THE STAPLER WAS UNABLE TO FIRE BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. WHEN THE SURGEON USED THE LOAD, THERE WAS ONLY A PARTIAL STAPLE OF THREE OF THE RELOADS, SO THE TISSUE WAS A LITTLE DAMAGED. ALSO FOUND OUT THAT THE PROXIMAL STAPLE FORMATION WAS POOR AND THE DISTAL STAPLE LINE WAS INCOMPLETE. FIRST THEY USED A STAPLER AND RELOAD BUT IT DID NOT WORK, SO HE TOOK AN OTHER LOAD AND CHANGED TO ANOTHER DEVICE BUT THEY FORGOT THAT THEY NEEDED A 15MM. TROCAR, SO THEY USED ANOTHER STAPLER AND IT WAS OKAY.
10-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON PROCEDURE, WHILE CUTTING THE COLON, THE STAPLER WAS UNABLE TO FIRE BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. WHEN THE SURGEON USED THE LOAD, THERE WAS ONLY A PARTIAL STAPLE OF THREE OF THE RELOADS, SO THE TISSUE WAS A LITTLE DAMAGED. ALSO FOUND OUT THAT THE PROXIMAL STAPLE FORMATION WAS POOR AND THE DISTAL STAPLE LINE WAS INCOMPLETE. FIRST THEY USED A STAPLER AND RELOAD BUT IT DID NOT WORK, SO HE TOOK AN OTHER LOAD AND CHANGED TO ANOTHER DEVICE BUT THEY FORGOT THAT THEY NEEDED A 15MM. TROCAR, SO THEY USED ANOTHER STAPLER AND IT WAS OKAY.
10-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | nan | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WITH A FLASHING INDICATOR LIGHTS WAS UNABLE TO FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. ALSO FOUND THAT THE DISTAL STAPLE LINE WAS INCOMPLETE. STAPLE FIRING ONLY OCCURRED AT THE PROXIMAL TISSUE, AND AFTER STOP THE PROCESS OF FIRING. WHEN THE SURGEON OPENED THE JAW OF THE DEVICE, THE BLADE DID NOT ACTIVATE AND NO CUTTING OF TISSUE OCCURRED.
16-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDOGIA | Break | A0401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, WHILE DOING A LUNG RESECTION, WHEN THE SURGEON FIRED THE HANDLE, RELOAD STARTED TO CUT AND LAY STAPLES AND THEN NOTHING CAME OUT/CUT; AT THE SAME TIME THE ELBOW OF THE SHAFT BROKE AND TONS OF METAL CAME OUT OF THE ELBOW. THE SURGEON SAID IT WAS ALMOST AS IF THE DRIVE MECHANISM WAS HITTING RESISTANCE AND THE METAL FATIGUED AT THE JOINT. THERE WAS TRAUMA TO THE TISSUE DUE TO THE CLAMP AND FAILURE TO FIRE OR CUT BUT THERE WAS NO PERMANENT DAMAGE. A NEW HANDLE AND RELOAD USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDOGIA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, WHILE DOING A LUNG RESECTION, WHEN THE SURGEON FIRED THE HANDLE, RELOAD STARTED TO CUT AND LAY STAPLES AND THEN NOTHING CAME OUT/CUT; AT THE SAME TIME THE ELBOW OF THE SHAFT BROKE AND TONS OF METAL CAME OUT OF THE ELBOW. THE SURGEON SAID IT WAS ALMOST AS IF THE DRIVE MECHANISM WAS HITTING RESISTANCE AND THE METAL FATIGUED AT THE JOINT. THERE WAS TRAUMA TO THE TISSUE DUE TO THE CLAMP AND FAILURE TO FIRE OR CUT BUT THERE WAS NO PERMANENT DAMAGE. A NEW HANDLE AND RELOAD USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-09-2017 | Injury | us surgical puerto rico | 26-04-2018 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN UPPER AND LOWER LOBE WEDGES FOR PULMONARY METASTASES FROM BREAST CANCER, THE PATIENT PRESENTED A BRUTAL DISCHARGE OF 1.3 LITERS OF BLOOD INTO THE CHEST TUBE AND WAS URGENTLY SENT TO SURGERY DUE TO TEARING OF THE LUNG PARENCHYMA. ALSO LEAD TO AN EXTENDED PATIENT HOSPITALIZATION.
25-09-2017 | Injury | us surgical puerto rico | 26-04-2018 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING AN UPPER AND LOWER LOBE WEDGES FOR PULMONARY METASTASES FROM BREAST CANCER, THE PATIENT PRESENTED A BRUTAL DISCHARGE OF 1.3 LITERS OF BLOOD INTO THE CHEST TUBE AND WAS URGENTLY SENT TO SURGERY DUE TO TEARING OF THE LUNG PARENCHYMA. ALSO LEAD TO AN EXTENDED PATIENT HOSPITALIZATION.
25-09-2017 | Injury | us surgical puerto rico | 26-04-2018 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN UPPER AND LOWER LOBE WEDGES FOR PULMONARY METASTASES FROM BREAST CANCER, THE PATIENT PRESENTED A BRUTAL DISCHARGE OF 1.3 LITERS OF BLOOD INTO THE CHEST TUBE AND WAS URGENTLY SENT TO SURGERY DUE TO TEARING OF THE LUNG PARENCHYMA. ALSO LEAD TO AN EXTENDED PATIENT HOSPITALIZATION.
25-09-2017 | Injury | us surgical puerto rico | 26-04-2018 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN UPPER AND LOWER LOBE WEDGES FOR PULMONARY METASTASES FROM BREAST CANCER, THE PATIENT PRESENTED A BRUTAL DISCHARGE OF 1.3 LITERS OF BLOOD INTO THE CHEST TUBE AND WAS URGENTLY SENT TO SURGERY DUE TO TEARING OF THE LUNG PARENCHYMA. ALSO LEAD TO AN EXTENDED PATIENT HOSPITALIZATION.
25-09-2017 | Injury | us surgical puerto rico | 26-04-2018 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING AN UPPER AND LOWER LOBE WEDGES FOR PULMONARY METASTASES FROM BREAST CANCER, THE PATIENT PRESENTED A BRUTAL DISCHARGE OF 1.3 LITERS OF BLOOD INTO THE CHEST TUBE AND WAS URGENTLY SENT TO SURGERY DUE TO TEARING OF THE LUNG PARENCHYMA. ALSO LEAD TO AN EXTENDED PATIENT HOSPITALIZATION.
25-09-2017 | Injury | us surgical puerto rico | 26-04-2018 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN UPPER AND LOWER LOBE WEDGES FOR PULMONARY METASTASES FROM BREAST CANCER, THE PATIENT PRESENTED A BRUTAL DISCHARGE OF 1.3 LITERS OF BLOOD INTO THE CHEST TUBE AND WAS URGENTLY SENT TO SURGERY DUE TO TEARING OF THE LUNG PARENCHYMA. ALSO LEAD TO AN EXTENDED PATIENT HOSPITALIZATION.
17-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | nan | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ROBOTIC LAPAROSCOPIC BYPASS, THROUGHOUT THE DIVIDED STEP PROCEDURE, THE DEVICE WAS ABLE TO FIRE, BUT FAILED TO RELOAD. TO RESOLVE THE ISSUE, THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
17-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | EEA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOIDECTOMY, WHILE BEING APPLIED AT THE ANUS, THE STAPLER DID NOT FIRE ALTHOUGH THE SPACE BETWEEN THE CARTRIDGE AND ANVIL WAS CLOSED AND THE GREEN BAR WAS VISIBLE IN THE INDICATOR WINDOW. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A WHIPPLE PROCEDURE, PRIOR TO THE USE OF THE DEVICE, THEY CONNECTED THE CARTRIDGE TO THE HANDLE, THEY TRIED TO CLOSE THE JAW FOR CHECKING HOWEVER IT COULD NOT OPEN. THEY DID NOT PUSH THE GREEN BUTTON AT ALL WHEN THEY SQUEEZED THE HANDLE. THEY PULLED THE BLACK RELEASE KNOB, THEN THE HANDLE RETURNED TO THE INITIAL POSITION AND THE JAWS OPENED. THEN THEY NOTICED THE CARTRIDGE WAS PRE-FIRED AND IT BECAME UNUSABLE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT HARM.
17-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA ULTRA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A WHIPPLE PROCEDURE, PRIOR TO THE USE OF THE DEVICE, THEY CONNECTED THE CARTRIDGE TO THE HANDLE, THEY TRIED TO CLOSE THE JAW FOR CHECKING HOWEVER IT COULD NOT OPEN. THEY DID NOT PUSH THE GREEN BUTTON AT ALL WHEN THEY SQUEEZED THE HANDLE. THEY PULLED THE BLACK RELEASE KNOB, THEN THE HANDLE RETURNED TO THE INITIAL POSITION AND THE JAWS OPENED. THEN THEY NOTICED THE CARTRIDGE WAS PRE-FIRED AND IT BECAME UNUSABLE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT HARM.
31-03-2018 | Injury | ethicon endo-surgery | 26-04-2018 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Code Available | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT FIRE. DURING THE ENDOSCOPIC LOW RECTAL ANTERIOR RESECTION, COULD NOT FIRE. USED KNIFE REVERSE BUTTON TO OPEN THE JAW. AS THIS HOSPITAL DID NO HAVE BACK-UP DEVICE, CHANGED THE SURGERY TO OPEN SURGERY, USED SCALPEL AND SUTURE TO CUT AND SEW THE WOUND. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
17-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC WITH LYSIS OF ADHESIONS WITH APPENDECTOMY, ON THE 3RD STAPLE FIRING, THE STAPLER RELOAD GOT STUCK ON TISSUE ACROSS THE APPENDIX. ALSO, AFTER 2-3 SQUEEZES OF THE HANDLE, A CRACK WAS HEARD AND THE DISTAL END OF THE HANDLE (BLACK TIP) BROKE OFF AND SHATTERED AND FELL INTO THE PATIENT CAVITY, WHICH WAS REMOVED BY USING ANOTHER INSTRUMENT. THEY STATED THEY COULD NOT GET THE STAPLE RELOAD OFF INITIALLY, THEY THEN WERE ABLE TO GET THE STAPLE RELOAD DETACHED FROM THE CURRENT HANDLE (WHILE STILL ATTACHED TO THE TISSUE) AND TRIED REATTACHING THE STAPLE RELOAD TO A NEW HANDLE BUT STILL COULD NOT GET THE RELOAD JAWS TO OPEN. THEY REPLACED THE DEVICE AND RELOAD AND PLACED IT PARALLEL TO THE OTHER STAPLER RELOAD AND FINISHED THE TRANSECTION OF THE TISSUE.
17-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC WITH LYSIS OF ADHESIONS WITH APPENDECTOMY, ON THE 3RD STAPLE FIRING, THE STAPLER RELOAD GOT STUCK ON TISSUE ACROSS THE APPENDIX. ALSO, AFTER 2-3 SQUEEZES OF THE HANDLE, A CRACK WAS HEARD AND THE DISTAL END OF THE HANDLE (BLACK TIP) BROKE OFF AND SHATTERED AND FELL INTO THE PATIENT CAVITY, WHICH WAS REMOVED BY USING ANOTHER INSTRUMENT. THEY STATED THEY COULD NOT GET THE STAPLE RELOAD OFF INITIALLY, THEY THEN WERE ABLE TO GET THE STAPLE RELOAD DETACHED FROM THE CURRENT HANDLE (WHILE STILL ATTACHED TO THE TISSUE) AND TRIED REATTACHING THE STAPLE RELOAD TO A NEW HANDLE BUT STILL COULD NOT GET THE RELOAD JAWS TO OPEN. THEY REPLACED THE DEVICE AND RELOAD AND PLACED IT PARALLEL TO THE OTHER STAPLER RELOAD AND FINISHED THE TRANSECTION OF THE TISSUE.
10-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WITH A FLASHING INDICATOR LIGHTS WAS UNABLE TO FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. ALSO FOUND THAT THE DISTAL STAPLE LINE WAS INCOMPLETE. STAPLE FIRING ONLY OCCURRED AT THE PROXIMAL TISSUE, AND AFTER STOP THE PROCESS OF FIRING. WHEN THE SURGEON OPENED THE JAW OF THE DEVICE, THE BLADE DID NOT ACTIVATE AND NO CUTTING OF TISSUE OCCURRED.
10-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST FIRING OF COLECTOMY, THEY TOOK THE DEVICE FROM THE PACKAGE AND TRIED TO FIRE, HOWEVER, THE FIRING KNOB DID NOT ADVANCE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE PATIENT HAS NO INJURY.
02-12-2015 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER AFTER THE STAPLER WAS POSITIONED WITH THE PURPLE LOAD AT THE STOMACH, ONLY THE FIRST CLAMPS HAVE SHOT. REPLACED BY A SECOND CHARGE (GREEN) DIDN'T WORK. IT WAS NECESSARY REPLACE THE STAPLER.
01-06-2015 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE LOAD DID NOT FIRE. NO INJURY REPORTED.
19-11-2015 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER THE LOADS PRESENTED MALFUNCTION.
17-05-2016 | Malfunction | covidien lp llc north haven | 26-04-2018 | ABSORBATACK | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE PRODUCT'S STEM PRESENTED A PROBLEM AND BROKE. NO INJURY REPORTED.
27-03-2018 | Malfunction | ethicon endo-surgery | 26-04-2018 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLES MALFORMED AFTER FIRING. DURING THE ENDOSCOPIC LOBECTOMY, AFTER FIRED, FOUND THE STAPLES MALFORMED WITH IRREGULAR SHAPE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
16-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA ULTRA | Break | A0401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, WHILE DOING A LUNG RESECTION, WHEN THE SURGEON FIRED THE HANDLE, THE RELOAD STARTED TO CUT AND LAY STAPLES AND THEN NOTHING CAME OUT OR CUT, AT THE SAME TIME THE ELBOW OF THE SHAFT BROKE AND TONS OF METAL CAME OUT OF THE ELBOW. THE SURGEON SAID IT WAS ALMOST AS IF THE DRIVE MECHANISM WAS HITTING RESISTANCE AND THE METAL FATIGUED AT THE JOINT. THERE WAS TRAUMA TO THE TISSUE DUE TO CLAMP AND THE FAILURE TO FIRE OR CUT THE DAMAGE WAS NOT PERMANENT. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, WHILE DOING A LUNG RESECTION, WHEN THE SURGEON FIRED THE HANDLE, THE RELOAD STARTED TO CUT AND LAY STAPLES AND THEN NOTHING CAME OUT OR CUT, AT THE SAME TIME THE ELBOW OF THE SHAFT BROKE AND TONS OF METAL CAME OUT OF THE ELBOW. THE SURGEON SAID IT WAS ALMOST AS IF THE DRIVE MECHANISM WAS HITTING RESISTANCE AND THE METAL FATIGUED AT THE JOINT. THERE WAS TRAUMA TO THE TISSUE DUE TO CLAMP AND THE FAILURE TO FIRE OR CUT THE DAMAGE WAS NOT PERMANENT. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-07-2015 | Malfunction | covidien lp llc north haven | 26-04-2018 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PLEURECTOMY PROCEDURE OCCURRED THE LOAD LOCK. NO INJURY REPORTED.
11-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON PROCEDURE, AFTER ONE NORMAL FIRING OF THE DEVICE, IT REFUSED TO FIRE. A NEW INSTRUMENT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | MULTIFIRE ENDO TA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC DONOR NEPHRECTOMY (LEFT SIDE), THE DEVICE WAS ABLE TO FIRE, BUT THE STAPLES ON THE RELOAD AND TISSUE ARE MALFORMED. THE DOCTOR FIRED THE DEVICE ON THE RENAL VEIN (3RD FIRING). ON FIRST FIRING, URETER, SECOND, RENAL ARTERY, THIRD: RENTAL VEIN. UPON COMPLETION OF FIRING, THE DOCTOR ATTEMPTED TO OPEN THE STAPLER BUT THE TISSUE/STAPLES WERE STUCK TO THE RELOAD. THE DOCTOR WAS ABLE TO FIRE TWO CLIPS BELOW THE JAWS OF THE RELOAD AND HAD TO CUT THE TISSUE TO REMOVE THE TISSUE FROM THE STAPLER. THERE WAS UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE A RESULT OF THE ISSUE. A LOSS OF 1-2MM TISSUE WAS NOTED. THERE WAS NO PATIENT INJURY.
17-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | MULTIFIRE ENDO TA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC DONOR NEPHRECTOMY (LEFT SIDE), THE DEVICE WAS ABLE TO FIRE, BUT THE STAPLES ON THE RELOAD AND TISSUE ARE MALFORMED. THE DOCTOR FIRED THE DEVICE ON THE RENAL VEIN (3RD FIRING). ON FIRST FIRING, URETER, SECOND, RENAL ARTERY, THIRD: RENTAL VEIN. UPON COMPLETION OF FIRING, THE DOCTOR ATTEMPTED TO OPEN THE STAPLER BUT THE TISSUE/STAPLES WERE STUCK TO THE RELOAD. THE DOCTOR WAS ABLE TO FIRE TWO CLIPS BELOW THE JAWS OF THE RELOAD AND HAD TO CUT THE TISSUE TO REMOVE THE TISSUE FROM THE STAPLER. THERE WAS UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE A RESULT OF THE ISSUE. A LOSS OF 1-2MM TISSUE WAS NOTED. THERE WAS NO PATIENT INJURY.
26-03-2018 | Malfunction | ethicon endo-surgery | 26-04-2018 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLES MALFORMED AFTER FIRING. DURING THE ENDOSCOPIC LOBECTOMY, WHEN SEVERING THE PULMONARY FISSURE,AFTER FIRED, FOUND THE STAPLES MALFORMED WITH IRREGULAR SHAPE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
13-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | TA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLES DID NOT DEPLOY. TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THE SURGEON SUTURED IT BY HAND. THERE WAS NO PATIENT INJURY.
16-04-2018 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE INSTALLING THE DEVICE, IT COULD NOT BE CLAMPED. THEY TRIED TO UNLOAD THE "NAIL" AGAIN AND THEN INSTALLING BUT IT STILL COULD NOT BE CLAMPED. THEY REPLACED THE DEVICE AND THE OPERATION CAN CONTINUE.
03-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | EGIA | Retraction Problem | A0510 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON FIRED TO PULMONARY PARENCHYMA WITH STAPLER, HOWEVER THE STAPLING AND RESECTION WERE NOT PERFORMED AT ALL AND A PART LIKE A METAL PLATE APPEARED FROM THE ARTICULATED PART OF THE CARTRIDGE. THE LUNG WAS VERY HARD. THE SURGEON ARTICULATED THE JAWS WITH MAXIMUM ANGLE AND CLAMPED THE TISSUE,THEN THE DISPLAY SCREEN SHOWED THE DEVICE WAS READY FOR USE. THE SURGEON CONSIDERED THE THICKNESS OF THE TISSUE AND WAITED FOR A WHILE ,THEN STARTED THE FIRING. THE FIRING SPEED SEEMED TO BE CONSTANT. AFTER THE FIRING ,THE SURGEON TRIED TO RETURN THE KNIFE TO THE INITIAL POSITION, HOWEVER COULD NOT RETURN THE KNIFE AND COULD NOT CLOSE THE JAWS. SURGEON RE-TRIED TO OPEN THE JAWS AGAIN, THEN FINALLY COULD OPEN THEM ,HOWEVER HE/SHE NOTED THE TISSUE WAS NOT STAPLED NOR RESECTED AT ALL. HE/SHE TRIED TO REMOVE THE CARTRIDGE OUT OF PATIENT'S BODY, HOWEVER COULD NOT REMOVE IT FROM THE TROCAR POINT. HE/SHE SLIGHTLY EXTENDED INCISION OF THE TROCAR POINT JUST TO REMOVE THE CARTRIDGE OUT OF PATIENT' S BODY. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE DEVICE WAS REMOVED FROM THE TISSUE BY ADDITIONALLY RESECTING THE TISSUE. THE INCISION SITE WAS EXTENDED BY LESS THAN 3 CM. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
03-04-2018 | Malfunction | covidien lp llc north haven | 26-04-2018 | EGIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY PROCEDURE, THE SURGEON FIRED TO PULMONARY PARENCHYMA WITH STAPLER, HOWEVER THE STAPLING AND RESECTION WERE NOT PERFORMED AT ALL AND A PART LIKE A METAL PLATE APPEARED FROM THE ARTICULATED PART OF THE CARTRIDGE. THE LUNG WAS VERY HARD. THE SURGEON ARTICULATED THE JAWS WITH MAXIMUM ANGLE AND CLAMPED THE TISSUE,THEN THE DISPLAY SCREEN SHOWED THE DEVICE WAS READY FOR USE. THE SURGEON CONSIDERED THE THICKNESS OF THE TISSUE AND WAITED FOR A WHILE ,THEN STARTED THE FIRING. THE FIRING SPEED SEEMED TO BE CONSTANT. AFTER THE FIRING ,THE SURGEON TRIED TO RETURN THE KNIFE TO THE INITIAL POSITION, HOWEVER COULD NOT RETURN THE KNIFE AND COULD NOT CLOSE THE JAWS. SURGEON RE-TRIED TO OPEN THE JAWS AGAIN, THEN FINALLY COULD OPEN THEM ,HOWEVER HE/SHE NOTED THE TISSUE WAS NOT STAPLED NOR RESECTED AT ALL. HE/SHE TRIED TO REMOVE THE CARTRIDGE OUT OF PATIENT'S BODY, HOWEVER COULD NOT REMOVE IT FROM THE TROCAR POINT. HE/SHE SLIGHTLY EXTENDED INCISION OF THE TROCAR POINT JUST TO REMOVE THE CARTRIDGE OUT OF PATIENT' S BODY. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE DEVICE WAS REMOVED FROM THE TISSUE BY ADDITIONALLY RESECTING THE TISSUE. THE INCISION SITE WAS EXTENDED BY LESS THAN 3 CM. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
17-03-2017 | Malfunction | us surgical puerto rico | 26-04-2018 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION PROCEDURE, THE DEVICE DID NOT FIRE WHEN SQUEEZING THE LEVER. IT WAS STATED THAT THE DEVICE CANNOT BE FIRED SO THEY CAN'T JUDGE WHETHER IT CAN BE CUT. ANOTHER RELOAD WAS USED TO CONTINUE THE PROCEDURE.
05-12-2018 | Death | ethicon endo-surgery | 02-01-2019 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Difficult to Open or Close | A051104 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE DOCTOR PERFORMED A LAPAROSCOPIC SIGMOID COLECTOMY ON A PATIENT USING AN UNKNOWN STAPLER AND THE STAPLER LOCKED AND COULDN'T BE UNLOCKED FROM THE PATIENT. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED AT THE TIME OF THE PROCEDURE. THE PATIENT WAS MOVED TO THE RECOVERY FLOOR. ON (B)(6) 2018, THE PATIENT DEVELOPED UNKNOWN SYMPTOMS THAT REQUIRED SURGICAL INTERVENTION. A SECOND DOCTOR PERFORMED SURGERY ON THE PATIENT, DISCOVERED A LEAK, SEPSIS, AND A HOLE IN THE COLON. DETAILS OF THE SECOND PROCEDURE WERE UNKNOWN. THE PATIENT WAS MOVED BACK TO THE RECOVERY FLOOR AND EXPIRED ON (B)(6) 2018.
05-12-2018 | Death | ethicon endo-surgery | 02-01-2019 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Difficult to Open or Close | A051104 | Death | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE DOCTOR PERFORMED A LAPAROSCOPIC SIGMOID COLECTOMY ON A PATIENT USING AN UNKNOWN STAPLER AND THE STAPLER LOCKED AND COULDN'T BE UNLOCKED FROM THE PATIENT. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED AT THE TIME OF THE PROCEDURE. THE PATIENT WAS MOVED TO THE RECOVERY FLOOR. ON (B)(6) 2018, THE PATIENT DEVELOPED UNKNOWN SYMPTOMS THAT REQUIRED SURGICAL INTERVENTION. A SECOND DOCTOR PERFORMED SURGERY ON THE PATIENT, DISCOVERED A LEAK, SEPSIS, AND A HOLE IN THE COLON. DETAILS OF THE SECOND PROCEDURE WERE UNKNOWN. THE PATIENT WAS MOVED BACK TO THE RECOVERY FLOOR AND EXPIRED ON (B)(6) 2018.
05-12-2018 | Death | ethicon endo-surgery | 02-01-2019 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Difficult to Open or Close | A051104 | Sepsis | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE DOCTOR PERFORMED A LAPAROSCOPIC SIGMOID COLECTOMY ON A PATIENT USING AN UNKNOWN STAPLER AND THE STAPLER LOCKED AND COULDN'T BE UNLOCKED FROM THE PATIENT. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED AT THE TIME OF THE PROCEDURE. THE PATIENT WAS MOVED TO THE RECOVERY FLOOR. ON (B)(6) 2018, THE PATIENT DEVELOPED UNKNOWN SYMPTOMS THAT REQUIRED SURGICAL INTERVENTION. A SECOND DOCTOR PERFORMED SURGERY ON THE PATIENT, DISCOVERED A LEAK, SEPSIS, AND A HOLE IN THE COLON. DETAILS OF THE SECOND PROCEDURE WERE UNKNOWN. THE PATIENT WAS MOVED BACK TO THE RECOVERY FLOOR AND EXPIRED ON (B)(6) 2018.
05-12-2018 | Death | ethicon endo-surgery | 02-01-2019 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Difficult to Open or Close | A051104 | No Code Available | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE DOCTOR PERFORMED A LAPAROSCOPIC SIGMOID COLECTOMY ON A PATIENT USING AN UNKNOWN STAPLER AND THE STAPLER LOCKED AND COULDN'T BE UNLOCKED FROM THE PATIENT. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED AT THE TIME OF THE PROCEDURE. THE PATIENT WAS MOVED TO THE RECOVERY FLOOR. ON (B)(6) 2018, THE PATIENT DEVELOPED UNKNOWN SYMPTOMS THAT REQUIRED SURGICAL INTERVENTION. A SECOND DOCTOR PERFORMED SURGERY ON THE PATIENT, DISCOVERED A LEAK, SEPSIS, AND A HOLE IN THE COLON. DETAILS OF THE SECOND PROCEDURE WERE UNKNOWN. THE PATIENT WAS MOVED BACK TO THE RECOVERY FLOOR AND EXPIRED ON (B)(6) 2018.
05-12-2018 | Death | ethicon endo-surgery | 02-01-2019 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Physical Resistance/Sticking | A0509 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE DOCTOR PERFORMED A LAPAROSCOPIC SIGMOID COLECTOMY ON A PATIENT USING AN UNKNOWN STAPLER AND THE STAPLER LOCKED AND COULDN'T BE UNLOCKED FROM THE PATIENT. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED AT THE TIME OF THE PROCEDURE. THE PATIENT WAS MOVED TO THE RECOVERY FLOOR. ON (B)(6) 2018, THE PATIENT DEVELOPED UNKNOWN SYMPTOMS THAT REQUIRED SURGICAL INTERVENTION. A SECOND DOCTOR PERFORMED SURGERY ON THE PATIENT, DISCOVERED A LEAK, SEPSIS, AND A HOLE IN THE COLON. DETAILS OF THE SECOND PROCEDURE WERE UNKNOWN. THE PATIENT WAS MOVED BACK TO THE RECOVERY FLOOR AND EXPIRED ON (B)(6) 2018.
05-12-2018 | Death | ethicon endo-surgery | 02-01-2019 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Physical Resistance/Sticking | A0509 | Death | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE DOCTOR PERFORMED A LAPAROSCOPIC SIGMOID COLECTOMY ON A PATIENT USING AN UNKNOWN STAPLER AND THE STAPLER LOCKED AND COULDN'T BE UNLOCKED FROM THE PATIENT. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED AT THE TIME OF THE PROCEDURE. THE PATIENT WAS MOVED TO THE RECOVERY FLOOR. ON (B)(6) 2018, THE PATIENT DEVELOPED UNKNOWN SYMPTOMS THAT REQUIRED SURGICAL INTERVENTION. A SECOND DOCTOR PERFORMED SURGERY ON THE PATIENT, DISCOVERED A LEAK, SEPSIS, AND A HOLE IN THE COLON. DETAILS OF THE SECOND PROCEDURE WERE UNKNOWN. THE PATIENT WAS MOVED BACK TO THE RECOVERY FLOOR AND EXPIRED ON (B)(6) 2018.
05-12-2018 | Death | ethicon endo-surgery | 02-01-2019 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Physical Resistance/Sticking | A0509 | Sepsis | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE DOCTOR PERFORMED A LAPAROSCOPIC SIGMOID COLECTOMY ON A PATIENT USING AN UNKNOWN STAPLER AND THE STAPLER LOCKED AND COULDN'T BE UNLOCKED FROM THE PATIENT. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED AT THE TIME OF THE PROCEDURE. THE PATIENT WAS MOVED TO THE RECOVERY FLOOR. ON (B)(6) 2018, THE PATIENT DEVELOPED UNKNOWN SYMPTOMS THAT REQUIRED SURGICAL INTERVENTION. A SECOND DOCTOR PERFORMED SURGERY ON THE PATIENT, DISCOVERED A LEAK, SEPSIS, AND A HOLE IN THE COLON. DETAILS OF THE SECOND PROCEDURE WERE UNKNOWN. THE PATIENT WAS MOVED BACK TO THE RECOVERY FLOOR AND EXPIRED ON (B)(6) 2018.
05-12-2018 | Death | ethicon endo-surgery | 02-01-2019 | ENDOPATH*STEALTH CIR STAPLER (EXACT CODE UNKNOWN) | Physical Resistance/Sticking | A0509 | No Code Available | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE DOCTOR PERFORMED A LAPAROSCOPIC SIGMOID COLECTOMY ON A PATIENT USING AN UNKNOWN STAPLER AND THE STAPLER LOCKED AND COULDN'T BE UNLOCKED FROM THE PATIENT. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED AT THE TIME OF THE PROCEDURE. THE PATIENT WAS MOVED TO THE RECOVERY FLOOR. ON (B)(6) 2018, THE PATIENT DEVELOPED UNKNOWN SYMPTOMS THAT REQUIRED SURGICAL INTERVENTION. A SECOND DOCTOR PERFORMED SURGERY ON THE PATIENT, DISCOVERED A LEAK, SEPSIS, AND A HOLE IN THE COLON. DETAILS OF THE SECOND PROCEDURE WERE UNKNOWN. THE PATIENT WAS MOVED BACK TO THE RECOVERY FLOOR AND EXPIRED ON (B)(6) 2018.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE STAPLING OF THE STOMACH, AT THE SECOND, THIRD FIRING, THE SURGEON PRESSED THE GREEN SAFETY BUTTON AND THE LIGHTS FLASHED. IT WAS NOTED THAT THE STAPLER WAS NOT ABLE TO FIRE, BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. WHEN THE DOWN FIRE BUTTON WAS PRESSED ON THE DEVICE, IT WOULD NOT ADVANCE AT ALL. ANOTHER RELOAD WAS OPENED AND WAS FIRED SUCCESSFULLY WITH THE SAME STAPLE HANDLE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE STAPLING OF THE STOMACH, AT THE SECOND, THIRD FIRING, THE SURGEON PRESSED THE GREEN SAFETY BUTTON AND THE LIGHTS FLASHED. IT WAS NOTED THAT THE STAPLER WAS NOT ABLE TO FIRE, BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. WHEN THE DOWN FIRE BUTTON WAS PRESSED ON THE DEVICE, IT WOULD NOT ADVANCE AT ALL. ANOTHER RELOAD WAS OPENED AND WAS FIRED SUCCESSFULLY WITH THE SAME STAPLE HANDLE. THERE WAS NO PATIENT INJURY.
01-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY PROCEDURE, AFTER FIRING THE FIRST SHOT WITH THE GREEN CHARGE, THE STAPLER MADE A NOISE IN THE TRIGGER AND THE FIRING WAS NOT POSSIBLE TO CONTINUE. IT WAS ALSO STATED THAT THE TRIGGER WAS BROKEN. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY PROCEDURE, AFTER FIRING THE FIRST SHOT WITH THE GREEN CHARGE, THE STAPLER MADE A NOISE IN THE TRIGGER AND THE FIRING WAS NOT POSSIBLE TO CONTINUE. IT WAS ALSO STATED THAT THE TRIGGER WAS BROKEN. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY PROCEDURE, AFTER FIRING THE FIRST SHOT WITH THE GREEN CHARGE, THE STAPLER MADE A NOISE IN THE TRIGGER AND THE FIRING WAS NOT POSSIBLE TO CONTINUE. IT WAS ALSO STATED THAT THE TRIGGER WAS BROKEN. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-11-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY PROCEDURE, AFTER FIRING THE FIRST SHOT WITH THE GREEN CHARGE, THE STAPLER MADE A NOISE IN THE TRIGGER AND THE FIRING WAS NOT POSSIBLE TO CONTINUE. IT WAS ALSO STATED THAT THE TRIGGER WAS BROKEN. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, DURING LUNG RESECTION, AFTER THE TISSUE WAS CLAMPED, THE FIRING WAS NOT ABLE TO BE PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, DURING LUNG RESECTION, AFTER THE TISSUE WAS CLAMPED, THE FIRING WAS NOT ABLE TO BE PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 02-01-2019 | 75 MM NTLC SELECTABLE RELOAD | Device Displays Incorrect Message | A090201 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN THE COURSE OF A PROCEDURE (RT.HEMI-COLOECTOMY) USING THE NTLC , AN ERROR OCCURRED IN THE PROCESS OF FIRING. THE BLADE SUDDENLY STOPPED IN THE MIDDLE OF THE DEVICE, AND THE BLADE DIDN'T MOVE FORWARD. AS AN EMERGENCY MEASURE, THE SURGEON BACKED OFF THE FIRING KNOB AND DETACH IT FROM THE TISSUE. HE CHANGED A NEW CARTRIDGE, HOWEVER, THE SAME ISSUE WAS OCCURRED 3 TIMES. AFTER CHANGING A NEW CARTRIDGE, HE COULD FINISHED THE PROCEDURE IN FOUR TIMES. THERE WERE NO REPORTED ADVERSE CONSEQUENCES FOR THE PATIENT SO FAR.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, DURING LUNG RESECTION, AFTER THE TISSUE WAS CLAMPED, THE FIRING WAS NOT ABLE TO BE PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING AN ENDOSCOPIC RADICAL GASTRECTOMY, BEFORE DISCONNECTING THE DUODENUM, THE DEVICE DID NOT FIRE, ALTHOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. RELOAD WAS REPLACE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING AN ENDOSCOPIC RADICAL GASTRECTOMY, BEFORE DISCONNECTING THE DUODENUM, THE DEVICE DID NOT FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. RELOAD WAS REPLACE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 02-01-2019 | 75MM SELECTABLE NEW TLC | Failure to Advance | A150204 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING PPPD PROCEDURE, THE BLADE STOPPED IN THE JAW AND COULD NOT PROCEED TOWARD DISTAL SIDE, SO SURGEON CHANGED ANOTHER DEVICE AND FINISH THE OPERATION. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 02-01-2019 | 75MM SELECTABLE NEW TLC | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING PPPD PROCEDURE, THE BLADE STOPPED IN THE JAW AND COULD NOT PROCEED TOWARD DISTAL SIDE, SO SURGEON CHANGED ANOTHER DEVICE AND FINISH THE OPERATION. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | Abscess | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE: TITLE: "INTENDED COST REDUCTION IN LAPAROSCOPIC APPENDECTOMY BY INTRODUCING THE ENDOLOOP: A SINGLE CENTER EXPERIENCE." AUTHORS: MATTHIAS MEHDORN, OLAF SCHÜRMANN, H. MAXIMILIAN MEHDORN, INES GOCKEL. CITATION: BMC SURGERY. 2017; 17: 80. DOI: 10.1186/S12893-017-0277-Z. COST REDUCTION MEASURES IN MEDICINE ARE GAINING GREATER IMPORTANCE NOWADAYS, ESPECIALLY IN HIGH VOLUME PROCEDURES SUCH AS LAPAROSCOPIC APPENDECTOMY (LAE). THE GOAL OF THIS STUDY WAS TO RETROSPECTIVELY EVALUATE POSSIBLE COST REDUCTION DUE TO INCREASED USE OF ENDOLOOP IN LAE IN OUR GENERAL SURGERY DEPARTMENT OF A TERTIARY REFERRAL UNIVERSITY HOSPITAL. THE AUTHORS PREVIOUSLY USED THE STAPLER FOR APPENDIX DISSECTION IN LAE AS THE LOCAL PROTOCOL BUT INTRODUCED THE ENDOLOOP (ETHICON) AS STANDARD METHOD IN 2015 TO REDUCE INTRAOPERATIVE COSTS. THE AUTHORS CONDUCTED A RETROSPECTIVE ANALYSIS OF 177 PATIENTS (77 MALE AND 100 FEMALE PATIENTS; AGE: 27.8 ± 15.5 YEARS) WHO UNDERWENT LAE BETWEEN JUNE 2014 AND OCTOBER 2015 IN THE DEPARTMENT. IN PHASE I, THE STAPLER WAS THE STANDARD APPROACH; IN PHASE II THE ENDOLOOP PDS II 0 SUTURE (ETHICON) WAS THE STANDARD DEVICE BUT IF AN ADVANCED APPENDICITIS WITH INFLAMMATION OF THE PROXIMAL APPENDIX, THE CECAL POLE OR PERFORATION WAS ENCOUNTERED, RESECTION WAS PERFORMED WITH THE STAPLER. THE AUTHORS USED AN ETS ARTICULATING LINEAR CUTTER (ETHICON). FOR PHASE I AND PHASE II COHORT, REPORTED COMPLICATIONS INCLUDED SUPERFICIAL WOUND SITE INFECTION (N-1), PERSISTENT INTRAABDOMINAL ABSCESS WHICH REQUIRED REINTERVENTION, "INTRAABDOMMINAL" ABSCESS FORMATION, PERITYPHLITIC ABSCESS WHICH WERE TREATED WITH POST-OPERATIVE ABDOMINAL LAVAGE VIA INTRA-OPERATIVELY PLACED DRAINAGE CATHETERS AND INTRAVENOUS ANTIBIOTICS, AND SEVERE SEPSIS (N-1) AND CONSEQUENTLY FULLY RECOVERED. ALTHOUGH THE ENDOLOOP (ETHICON) IS A PRACTICAL AND CHEAP DEVICE FOR RESECTING THE APPENDIX IN LAPAROSCOPIC APPENDECTOMY, THE ¿TRUE¿ ECONOMIC BENEFIT OF THE ENDOLOOP AS STANDARD METHOD FOR LAE IS NOT AS HIGH AS EXPECTED A PRIORITY WHEN REGARDING COSTS OF EACH DEVICE. IN THE SERIES, COSTS BY PROLONGED DO NEGATED SOME OF THE EXPECTED SAVINGS BY ENDOLOOP FOR VARIOUS POSSIBLE REASONS, E.G. SURGEON¿S LEARNING CURVES, SURGEONS AT AN EARLY STAGE OF TRAINING. GROWING EXPERIENCE WITH THE ENDOLOOP MIGHT REVEAL MORE PRONOUNCED COST-REDUCTION POSSIBILITIES.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | Unspecified Infection | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE: TITLE: "INTENDED COST REDUCTION IN LAPAROSCOPIC APPENDECTOMY BY INTRODUCING THE ENDOLOOP: A SINGLE CENTER EXPERIENCE." AUTHORS: MATTHIAS MEHDORN, OLAF SCHÜRMANN, H. MAXIMILIAN MEHDORN, INES GOCKEL. CITATION: BMC SURGERY. 2017; 17: 80. DOI: 10.1186/S12893-017-0277-Z. COST REDUCTION MEASURES IN MEDICINE ARE GAINING GREATER IMPORTANCE NOWADAYS, ESPECIALLY IN HIGH VOLUME PROCEDURES SUCH AS LAPAROSCOPIC APPENDECTOMY (LAE). THE GOAL OF THIS STUDY WAS TO RETROSPECTIVELY EVALUATE POSSIBLE COST REDUCTION DUE TO INCREASED USE OF ENDOLOOP IN LAE IN OUR GENERAL SURGERY DEPARTMENT OF A TERTIARY REFERRAL UNIVERSITY HOSPITAL. THE AUTHORS PREVIOUSLY USED THE STAPLER FOR APPENDIX DISSECTION IN LAE AS THE LOCAL PROTOCOL BUT INTRODUCED THE ENDOLOOP (ETHICON) AS STANDARD METHOD IN 2015 TO REDUCE INTRAOPERATIVE COSTS. THE AUTHORS CONDUCTED A RETROSPECTIVE ANALYSIS OF 177 PATIENTS (77 MALE AND 100 FEMALE PATIENTS; AGE: 27.8 ± 15.5 YEARS) WHO UNDERWENT LAE BETWEEN JUNE 2014 AND OCTOBER 2015 IN THE DEPARTMENT. IN PHASE I, THE STAPLER WAS THE STANDARD APPROACH; IN PHASE II THE ENDOLOOP PDS II 0 SUTURE (ETHICON) WAS THE STANDARD DEVICE BUT IF AN ADVANCED APPENDICITIS WITH INFLAMMATION OF THE PROXIMAL APPENDIX, THE CECAL POLE OR PERFORATION WAS ENCOUNTERED, RESECTION WAS PERFORMED WITH THE STAPLER. THE AUTHORS USED AN ETS ARTICULATING LINEAR CUTTER (ETHICON). FOR PHASE I AND PHASE II COHORT, REPORTED COMPLICATIONS INCLUDED SUPERFICIAL WOUND SITE INFECTION (N-1), PERSISTENT INTRAABDOMINAL ABSCESS WHICH REQUIRED REINTERVENTION, "INTRAABDOMMINAL" ABSCESS FORMATION, PERITYPHLITIC ABSCESS WHICH WERE TREATED WITH POST-OPERATIVE ABDOMINAL LAVAGE VIA INTRA-OPERATIVELY PLACED DRAINAGE CATHETERS AND INTRAVENOUS ANTIBIOTICS, AND SEVERE SEPSIS (N-1) AND CONSEQUENTLY FULLY RECOVERED. ALTHOUGH THE ENDOLOOP (ETHICON) IS A PRACTICAL AND CHEAP DEVICE FOR RESECTING THE APPENDIX IN LAPAROSCOPIC APPENDECTOMY, THE ¿TRUE¿ ECONOMIC BENEFIT OF THE ENDOLOOP AS STANDARD METHOD FOR LAE IS NOT AS HIGH AS EXPECTED A PRIORITY WHEN REGARDING COSTS OF EACH DEVICE. IN THE SERIES, COSTS BY PROLONGED DO NEGATED SOME OF THE EXPECTED SAVINGS BY ENDOLOOP FOR VARIOUS POSSIBLE REASONS, E.G. SURGEON¿S LEARNING CURVES, SURGEONS AT AN EARLY STAGE OF TRAINING. GROWING EXPERIENCE WITH THE ENDOLOOP MIGHT REVEAL MORE PRONOUNCED COST-REDUCTION POSSIBILITIES.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | Sepsis | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE: TITLE: "INTENDED COST REDUCTION IN LAPAROSCOPIC APPENDECTOMY BY INTRODUCING THE ENDOLOOP: A SINGLE CENTER EXPERIENCE." AUTHORS: MATTHIAS MEHDORN, OLAF SCHÜRMANN, H. MAXIMILIAN MEHDORN, INES GOCKEL. CITATION: BMC SURGERY. 2017; 17: 80. DOI: 10.1186/S12893-017-0277-Z. COST REDUCTION MEASURES IN MEDICINE ARE GAINING GREATER IMPORTANCE NOWADAYS, ESPECIALLY IN HIGH VOLUME PROCEDURES SUCH AS LAPAROSCOPIC APPENDECTOMY (LAE). THE GOAL OF THIS STUDY WAS TO RETROSPECTIVELY EVALUATE POSSIBLE COST REDUCTION DUE TO INCREASED USE OF ENDOLOOP IN LAE IN OUR GENERAL SURGERY DEPARTMENT OF A TERTIARY REFERRAL UNIVERSITY HOSPITAL. THE AUTHORS PREVIOUSLY USED THE STAPLER FOR APPENDIX DISSECTION IN LAE AS THE LOCAL PROTOCOL BUT INTRODUCED THE ENDOLOOP (ETHICON) AS STANDARD METHOD IN 2015 TO REDUCE INTRAOPERATIVE COSTS. THE AUTHORS CONDUCTED A RETROSPECTIVE ANALYSIS OF 177 PATIENTS (77 MALE AND 100 FEMALE PATIENTS; AGE: 27.8 ± 15.5 YEARS) WHO UNDERWENT LAE BETWEEN JUNE 2014 AND OCTOBER 2015 IN THE DEPARTMENT. IN PHASE I, THE STAPLER WAS THE STANDARD APPROACH; IN PHASE II THE ENDOLOOP PDS II 0 SUTURE (ETHICON) WAS THE STANDARD DEVICE BUT IF AN ADVANCED APPENDICITIS WITH INFLAMMATION OF THE PROXIMAL APPENDIX, THE CECAL POLE OR PERFORATION WAS ENCOUNTERED, RESECTION WAS PERFORMED WITH THE STAPLER. THE AUTHORS USED AN ETS ARTICULATING LINEAR CUTTER (ETHICON). FOR PHASE I AND PHASE II COHORT, REPORTED COMPLICATIONS INCLUDED SUPERFICIAL WOUND SITE INFECTION (N-1), PERSISTENT INTRAABDOMINAL ABSCESS WHICH REQUIRED REINTERVENTION, "INTRAABDOMMINAL" ABSCESS FORMATION, PERITYPHLITIC ABSCESS WHICH WERE TREATED WITH POST-OPERATIVE ABDOMINAL LAVAGE VIA INTRA-OPERATIVELY PLACED DRAINAGE CATHETERS AND INTRAVENOUS ANTIBIOTICS, AND SEVERE SEPSIS (N-1) AND CONSEQUENTLY FULLY RECOVERED. ALTHOUGH THE ENDOLOOP (ETHICON) IS A PRACTICAL AND CHEAP DEVICE FOR RESECTING THE APPENDIX IN LAPAROSCOPIC APPENDECTOMY, THE ¿TRUE¿ ECONOMIC BENEFIT OF THE ENDOLOOP AS STANDARD METHOD FOR LAE IS NOT AS HIGH AS EXPECTED A PRIORITY WHEN REGARDING COSTS OF EACH DEVICE. IN THE SERIES, COSTS BY PROLONGED DO NEGATED SOME OF THE EXPECTED SAVINGS BY ENDOLOOP FOR VARIOUS POSSIBLE REASONS, E.G. SURGEON¿S LEARNING CURVES, SURGEONS AT AN EARLY STAGE OF TRAINING. GROWING EXPERIENCE WITH THE ENDOLOOP MIGHT REVEAL MORE PRONOUNCED COST-REDUCTION POSSIBILITIES.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | No Code Available | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE: TITLE: "INTENDED COST REDUCTION IN LAPAROSCOPIC APPENDECTOMY BY INTRODUCING THE ENDOLOOP: A SINGLE CENTER EXPERIENCE." AUTHORS: MATTHIAS MEHDORN, OLAF SCHÜRMANN, H. MAXIMILIAN MEHDORN, INES GOCKEL. CITATION: BMC SURGERY. 2017; 17: 80. DOI: 10.1186/S12893-017-0277-Z. COST REDUCTION MEASURES IN MEDICINE ARE GAINING GREATER IMPORTANCE NOWADAYS, ESPECIALLY IN HIGH VOLUME PROCEDURES SUCH AS LAPAROSCOPIC APPENDECTOMY (LAE). THE GOAL OF THIS STUDY WAS TO RETROSPECTIVELY EVALUATE POSSIBLE COST REDUCTION DUE TO INCREASED USE OF ENDOLOOP IN LAE IN OUR GENERAL SURGERY DEPARTMENT OF A TERTIARY REFERRAL UNIVERSITY HOSPITAL. THE AUTHORS PREVIOUSLY USED THE STAPLER FOR APPENDIX DISSECTION IN LAE AS THE LOCAL PROTOCOL BUT INTRODUCED THE ENDOLOOP (ETHICON) AS STANDARD METHOD IN 2015 TO REDUCE INTRAOPERATIVE COSTS. THE AUTHORS CONDUCTED A RETROSPECTIVE ANALYSIS OF 177 PATIENTS (77 MALE AND 100 FEMALE PATIENTS; AGE: 27.8 ± 15.5 YEARS) WHO UNDERWENT LAE BETWEEN JUNE 2014 AND OCTOBER 2015 IN THE DEPARTMENT. IN PHASE I, THE STAPLER WAS THE STANDARD APPROACH; IN PHASE II THE ENDOLOOP PDS II 0 SUTURE (ETHICON) WAS THE STANDARD DEVICE BUT IF AN ADVANCED APPENDICITIS WITH INFLAMMATION OF THE PROXIMAL APPENDIX, THE CECAL POLE OR PERFORATION WAS ENCOUNTERED, RESECTION WAS PERFORMED WITH THE STAPLER. THE AUTHORS USED AN ETS ARTICULATING LINEAR CUTTER (ETHICON). FOR PHASE I AND PHASE II COHORT, REPORTED COMPLICATIONS INCLUDED SUPERFICIAL WOUND SITE INFECTION (N-1), PERSISTENT INTRAABDOMINAL ABSCESS WHICH REQUIRED REINTERVENTION, "INTRAABDOMMINAL" ABSCESS FORMATION, PERITYPHLITIC ABSCESS WHICH WERE TREATED WITH POST-OPERATIVE ABDOMINAL LAVAGE VIA INTRA-OPERATIVELY PLACED DRAINAGE CATHETERS AND INTRAVENOUS ANTIBIOTICS, AND SEVERE SEPSIS (N-1) AND CONSEQUENTLY FULLY RECOVERED. ALTHOUGH THE ENDOLOOP (ETHICON) IS A PRACTICAL AND CHEAP DEVICE FOR RESECTING THE APPENDIX IN LAPAROSCOPIC APPENDECTOMY, THE ¿TRUE¿ ECONOMIC BENEFIT OF THE ENDOLOOP AS STANDARD METHOD FOR LAE IS NOT AS HIGH AS EXPECTED A PRIORITY WHEN REGARDING COSTS OF EACH DEVICE. IN THE SERIES, COSTS BY PROLONGED DO NEGATED SOME OF THE EXPECTED SAVINGS BY ENDOLOOP FOR VARIOUS POSSIBLE REASONS, E.G. SURGEON¿S LEARNING CURVES, SURGEONS AT AN EARLY STAGE OF TRAINING. GROWING EXPERIENCE WITH THE ENDOLOOP MIGHT REVEAL MORE PRONOUNCED COST-REDUCTION POSSIBILITIES.
04-11-2018 | Injury | us surgical puerto rico | 02-01-2019 | GIA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, 10 DAYS AFTER AN OPEN EXTENDED PANCREATIC RESECTION, ON ESOPHAGUS STAPLING, A HOLE IN THE ANASTOMOSE (LEAKAGE) WAS FOUND. REOPERATION WAS PERFORMED TO CLOSE THE ANASTOMOSIS USING SUTURES.
04-11-2018 | Injury | us surgical puerto rico | 02-01-2019 | GIA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 10 DAYS AFTER AN OPEN EXTENDED PANCREATIC RESECTION, ON ESOPHAGUS STAPLING, A HOLE IN THE ANASTOMOSE (LEAKAGE) WAS FOUND. REOPERATION WAS PERFORMED TO CLOSE THE ANASTOMOSIS USING SUTURES.
04-11-2018 | Injury | us surgical puerto rico | 02-01-2019 | GIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, 10 DAYS AFTER AN OPEN EXTENDED PANCREATIC RESECTION, ON ESOPHAGUS STAPLING, A HOLE IN THE ANASTOMOSE (LEAKAGE) WAS FOUND. REOPERATION WAS PERFORMED TO CLOSE THE ANASTOMOSIS USING SUTURES.
04-11-2018 | Injury | us surgical puerto rico | 02-01-2019 | GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 10 DAYS AFTER AN OPEN EXTENDED PANCREATIC RESECTION, ON ESOPHAGUS STAPLING, A HOLE IN THE ANASTOMOSE (LEAKAGE) WAS FOUND. REOPERATION WAS PERFORMED TO CLOSE THE ANASTOMOSIS USING SUTURES.
04-11-2018 | Injury | us surgical puerto rico | 02-01-2019 | EEA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, 10 DAYS AFTER AN OPEN EXTENDED PANCREATIC RESECTION, ON ESOPHAGUS STAPLING, A HOLE IN THE ANASTOMOSE (LEAKAGE) WAS FOUND. REOPERATION WAS PERFORMED TO CLOSE THE ANASTOMOSIS USING SUTURES.
04-11-2018 | Injury | us surgical puerto rico | 02-01-2019 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 10 DAYS AFTER AN OPEN EXTENDED PANCREATIC RESECTION, ON ESOPHAGUS STAPLING, A HOLE IN THE ANASTOMOSE (LEAKAGE) WAS FOUND. REOPERATION WAS PERFORMED TO CLOSE THE ANASTOMOSIS USING SUTURES.
04-11-2018 | Injury | us surgical puerto rico | 02-01-2019 | EEA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, 10 DAYS AFTER AN OPEN EXTENDED PANCREATIC RESECTION, ON ESOPHAGUS STAPLING, A HOLE IN THE ANASTOMOSE (LEAKAGE) WAS FOUND. REOPERATION WAS PERFORMED TO CLOSE THE ANASTOMOSIS USING SUTURES.
04-11-2018 | Injury | us surgical puerto rico | 02-01-2019 | EEA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 10 DAYS AFTER AN OPEN EXTENDED PANCREATIC RESECTION, ON ESOPHAGUS STAPLING, A HOLE IN THE ANASTOMOSE (LEAKAGE) WAS FOUND. REOPERATION WAS PERFORMED TO CLOSE THE ANASTOMOSIS USING SUTURES.
06-12-2018 | Injury | us surgical puerto rico | 02-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, ONE ON LAPAROSCOPIC RESECTION, THE NAILS WERE BROKEN WHEN THE LUNG WAS BEING DISCONNECTED, AND COULD NOT BE CLOSED NORMALLY AFTER BEING REPLACED WITH THE 60-STRAIGHT NAIL. THE THIRD STAPLES WERE POORLY FORMED, HAD AN INCOMPLETE STAPLE LINE DISTALLY AND THEY SUTURED AGAIN TO FIX THE STAPLE LINE. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Injury | us surgical puerto rico | 02-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, ONE ON LAPAROSCOPIC RESECTION, THE NAILS WERE BROKEN WHEN THE LUNG WAS BEING DISCONNECTED, AND COULD NOT BE CLOSED NORMALLY AFTER BEING REPLACED WITH THE 60-STRAIGHT NAIL. THE THIRD STAPLES WERE POORLY FORMED, HAD AN INCOMPLETE STAPLE LINE DISTALLY AND THEY SUTURED AGAIN TO FIX THE STAPLE LINE. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Injury | us surgical puerto rico | 02-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, ONE ON LAPAROSCOPIC RESECTION, THE NAILS WERE BROKEN WHEN THE LUNG WAS BEING DISCONNECTED, AND COULD NOT BE CLOSED NORMALLY AFTER BEING REPLACED WITH THE 60-STRAIGHT NAIL. THE THIRD STAPLES WERE POORLY FORMED, HAD AN INCOMPLETE STAPLE LINE DISTALLY AND THEY SUTURED AGAIN TO FIX THE STAPLE LINE. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | Abdominal Pain | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE TITLE: "LOOPED SUTURE VERSUS STAPLER DEVICE IN PEDIATRIC LAPAROSCOPIC APPENDECTOMY: A COMPARATIVE OUTCOMES AND INTRAOPERATIVE COST ANALYSIS." AUTHOR(S): PUNAM P. PARIKH, JUN TASHIRO, AMY E.WAGENAAR, MIOSOTYS CURBELO, EDUARDO A. PEREZ, HOLLY L. NEVILLE, ANTHONY R. HOGAN, JUAN E. SOLA. CITATION: JOURNAL OF PEDIATRIC SURGERY 53 (2018) 616¿619; DOI: HTTP://DX.DOI.ORG/10.1016/J.JPEDSURG.2017.05.016. THE PURPOSE OF THIS STUDY WAS TO COMPARE CLINICAL OUTCOMES IN ADDITION TO INTRAOPERATIVE COST AFTER APPLICATION OF LOOPED SUTURE VERSUS STAPLER FOR STUMP CLOSURE IN PEDIATRIC PATIENTS UNDERGOING LAPAROSCOPIC APPENDECTOMY (LA) FOR APPENDICITIS. FROM FISCAL YEARS 2013 TO 2014, 238 PATIENTS UNDERWENT LA AND WERE STRATIFIED BY TYPE OF APPENDICEAL LIGATION: LOOP SUTURE (N=110) AND STAPLER LIGATION (N=128). IN THE PROCEDURE, THE APPENDIX WAS DIVIDED BETWEEN DISTAL GRASPER AND TWO ENDOLOOP PDS II WAS PLACED AT THE BASE OF THE APPENDIX OR WITH AN ENDOPATH ETS STAPLER/CUTTER, ETHICON (WHITE LOAD FOR THE MESOAPPENDIX AND BLUE LOAD FOR THE APPENDIX). THE MESOAPPENDIX IN THIS GROUP OF PATIENTS WAS DIVIDED WITH THE USE OF CAUTERY. IN HOSPITAL COMPLICATIONS INCLUDED RIGHT LOWER QUADRANT PELVIC ABSCESS (N=2) IN ENDOLOOP GROUP AND INTRA-ABDOMINAL ABSCESS (N=1) IN STAPLER GROUP. IN A 30-DAY RETURN TO ER ON ENDOLOOP GROUP, OUTCOMES INCLUDED NON-SPECIFIC ABDOMINAL PAIN (N=2), AND INTRA-ABDOMINAL ABSCESS (N=2) TO WHICH PATIENTS WERE READMITTED FOR DRAINAGE OF INTRA-ABDOMINAL ABSCESS AND PARENTERAL ANTIBIOTIC THERAPY. WHILE IN A 30-DAY RETURN TO ER ON STAPLER GROUP, OUTCOMES INCLUDED NON-SPECIFIC ABDOMINAL PAIN (N=2) OF WHICH ONE WAS TREATED CONSERVATIVELY, CONSTIPATION/ILEUS (N=3) WHICH WERE MANAGED CONSERVATIVELY, ABDOMINAL WALL ABSCESS (N=2) TREATED WITH DEBRIDEMENT, AND INTRA-ABDOMINAL ABSCESS (N=1) TREATED WITH DRAINAGE. THERE IS A HIGHER INCIDENCE OF INTRA-ABDOMINAL INFECTION AND ABSCESS FORMATION WHEN LOOPED SUTURE IS USED IN LIEU OF STAPLER DEVICE, EVEN IN PATIENTS WITH ACUTE NON PERFORATED APPENDICITIS. THIS ASSOCIATION MAY BE ATTRIBUTED TO INSUFFICIENT CLOSURE OF THE APPENDICEAL STUMP WITH LOOPED SUTURE, EXPOSURE OF CONTAMINATED MUCOSA AT THE CUT EDGES TO THE ABDOMINAL CAVITY, OR MUCOSAL NECROSIS IF LOOPS LOOSEN. A COMPARATIVE ANALYSIS OF LOOPED SUTURE VERSUS STAPLER DEVICE DURING LA FOR PEDIATRIC APPENDICITIS REVEALED THAT POSTOPERATIVE COMPLICATIONS, LENGTH OF STAY, ER VISITS AND READMISSIONS WERE NOT SIGNIFICANTLY DIFFERENT. LOOPED SUTURE LA WAS SIGNIFICANTLY MORE COST EFFICIENT THAN STAPLER LA.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | Abscess | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE TITLE: "LOOPED SUTURE VERSUS STAPLER DEVICE IN PEDIATRIC LAPAROSCOPIC APPENDECTOMY: A COMPARATIVE OUTCOMES AND INTRAOPERATIVE COST ANALYSIS." AUTHOR(S): PUNAM P. PARIKH, JUN TASHIRO, AMY E.WAGENAAR, MIOSOTYS CURBELO, EDUARDO A. PEREZ, HOLLY L. NEVILLE, ANTHONY R. HOGAN, JUAN E. SOLA. CITATION: JOURNAL OF PEDIATRIC SURGERY 53 (2018) 616¿619; DOI: HTTP://DX.DOI.ORG/10.1016/J.JPEDSURG.2017.05.016. THE PURPOSE OF THIS STUDY WAS TO COMPARE CLINICAL OUTCOMES IN ADDITION TO INTRAOPERATIVE COST AFTER APPLICATION OF LOOPED SUTURE VERSUS STAPLER FOR STUMP CLOSURE IN PEDIATRIC PATIENTS UNDERGOING LAPAROSCOPIC APPENDECTOMY (LA) FOR APPENDICITIS. FROM FISCAL YEARS 2013 TO 2014, 238 PATIENTS UNDERWENT LA AND WERE STRATIFIED BY TYPE OF APPENDICEAL LIGATION: LOOP SUTURE (N=110) AND STAPLER LIGATION (N=128). IN THE PROCEDURE, THE APPENDIX WAS DIVIDED BETWEEN DISTAL GRASPER AND TWO ENDOLOOP PDS II WAS PLACED AT THE BASE OF THE APPENDIX OR WITH AN ENDOPATH ETS STAPLER/CUTTER, ETHICON (WHITE LOAD FOR THE MESOAPPENDIX AND BLUE LOAD FOR THE APPENDIX). THE MESOAPPENDIX IN THIS GROUP OF PATIENTS WAS DIVIDED WITH THE USE OF CAUTERY. IN HOSPITAL COMPLICATIONS INCLUDED RIGHT LOWER QUADRANT PELVIC ABSCESS (N=2) IN ENDOLOOP GROUP AND INTRA-ABDOMINAL ABSCESS (N=1) IN STAPLER GROUP. IN A 30-DAY RETURN TO ER ON ENDOLOOP GROUP, OUTCOMES INCLUDED NON-SPECIFIC ABDOMINAL PAIN (N=2), AND INTRA-ABDOMINAL ABSCESS (N=2) TO WHICH PATIENTS WERE READMITTED FOR DRAINAGE OF INTRA-ABDOMINAL ABSCESS AND PARENTERAL ANTIBIOTIC THERAPY. WHILE IN A 30-DAY RETURN TO ER ON STAPLER GROUP, OUTCOMES INCLUDED NON-SPECIFIC ABDOMINAL PAIN (N=2) OF WHICH ONE WAS TREATED CONSERVATIVELY, CONSTIPATION/ILEUS (N=3) WHICH WERE MANAGED CONSERVATIVELY, ABDOMINAL WALL ABSCESS (N=2) TREATED WITH DEBRIDEMENT, AND INTRA-ABDOMINAL ABSCESS (N=1) TREATED WITH DRAINAGE. THERE IS A HIGHER INCIDENCE OF INTRA-ABDOMINAL INFECTION AND ABSCESS FORMATION WHEN LOOPED SUTURE IS USED IN LIEU OF STAPLER DEVICE, EVEN IN PATIENTS WITH ACUTE NON PERFORATED APPENDICITIS. THIS ASSOCIATION MAY BE ATTRIBUTED TO INSUFFICIENT CLOSURE OF THE APPENDICEAL STUMP WITH LOOPED SUTURE, EXPOSURE OF CONTAMINATED MUCOSA AT THE CUT EDGES TO THE ABDOMINAL CAVITY, OR MUCOSAL NECROSIS IF LOOPS LOOSEN. A COMPARATIVE ANALYSIS OF LOOPED SUTURE VERSUS STAPLER DEVICE DURING LA FOR PEDIATRIC APPENDICITIS REVEALED THAT POSTOPERATIVE COMPLICATIONS, LENGTH OF STAY, ER VISITS AND READMISSIONS WERE NOT SIGNIFICANTLY DIFFERENT. LOOPED SUTURE LA WAS SIGNIFICANTLY MORE COST EFFICIENT THAN STAPLER LA.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | No Code Available | Event Description: IT WAS REPORTED IN JOURNAL ARTICLE TITLE: "LOOPED SUTURE VERSUS STAPLER DEVICE IN PEDIATRIC LAPAROSCOPIC APPENDECTOMY: A COMPARATIVE OUTCOMES AND INTRAOPERATIVE COST ANALYSIS." AUTHOR(S): PUNAM P. PARIKH, JUN TASHIRO, AMY E.WAGENAAR, MIOSOTYS CURBELO, EDUARDO A. PEREZ, HOLLY L. NEVILLE, ANTHONY R. HOGAN, JUAN E. SOLA. CITATION: JOURNAL OF PEDIATRIC SURGERY 53 (2018) 616¿619; DOI: HTTP://DX.DOI.ORG/10.1016/J.JPEDSURG.2017.05.016. THE PURPOSE OF THIS STUDY WAS TO COMPARE CLINICAL OUTCOMES IN ADDITION TO INTRAOPERATIVE COST AFTER APPLICATION OF LOOPED SUTURE VERSUS STAPLER FOR STUMP CLOSURE IN PEDIATRIC PATIENTS UNDERGOING LAPAROSCOPIC APPENDECTOMY (LA) FOR APPENDICITIS. FROM FISCAL YEARS 2013 TO 2014, 238 PATIENTS UNDERWENT LA AND WERE STRATIFIED BY TYPE OF APPENDICEAL LIGATION: LOOP SUTURE (N=110) AND STAPLER LIGATION (N=128). IN THE PROCEDURE, THE APPENDIX WAS DIVIDED BETWEEN DISTAL GRASPER AND TWO ENDOLOOP PDS II WAS PLACED AT THE BASE OF THE APPENDIX OR WITH AN ENDOPATH ETS STAPLER/CUTTER, ETHICON (WHITE LOAD FOR THE MESOAPPENDIX AND BLUE LOAD FOR THE APPENDIX). THE MESOAPPENDIX IN THIS GROUP OF PATIENTS WAS DIVIDED WITH THE USE OF CAUTERY. IN HOSPITAL COMPLICATIONS INCLUDED RIGHT LOWER QUADRANT PELVIC ABSCESS (N=2) IN ENDOLOOP GROUP AND INTRA-ABDOMINAL ABSCESS (N=1) IN STAPLER GROUP. IN A 30-DAY RETURN TO ER ON ENDOLOOP GROUP, OUTCOMES INCLUDED NON-SPECIFIC ABDOMINAL PAIN (N=2), AND INTRA-ABDOMINAL ABSCESS (N=2) TO WHICH PATIENTS WERE READMITTED FOR DRAINAGE OF INTRA-ABDOMINAL ABSCESS AND PARENTERAL ANTIBIOTIC THERAPY. WHILE IN A 30-DAY RETURN TO ER ON STAPLER GROUP, OUTCOMES INCLUDED NON-SPECIFIC ABDOMINAL PAIN (N=2) OF WHICH ONE WAS TREATED CONSERVATIVELY, CONSTIPATION/ILEUS (N=3) WHICH WERE MANAGED CONSERVATIVELY, ABDOMINAL WALL ABSCESS (N=2) TREATED WITH DEBRIDEMENT, AND INTRA-ABDOMINAL ABSCESS (N=1) TREATED WITH DRAINAGE. THERE IS A HIGHER INCIDENCE OF INTRA-ABDOMINAL INFECTION AND ABSCESS FORMATION WHEN LOOPED SUTURE IS USED IN LIEU OF STAPLER DEVICE, EVEN IN PATIENTS WITH ACUTE NON PERFORATED APPENDICITIS. THIS ASSOCIATION MAY BE ATTRIBUTED TO INSUFFICIENT CLOSURE OF THE APPENDICEAL STUMP WITH LOOPED SUTURE, EXPOSURE OF CONTAMINATED MUCOSA AT THE CUT EDGES TO THE ABDOMINAL CAVITY, OR MUCOSAL NECROSIS IF LOOPS LOOSEN. A COMPARATIVE ANALYSIS OF LOOPED SUTURE VERSUS STAPLER DEVICE DURING LA FOR PEDIATRIC APPENDICITIS REVEALED THAT POSTOPERATIVE COMPLICATIONS, LENGTH OF STAY, ER VISITS AND READMISSIONS WERE NOT SIGNIFICANTLY DIFFERENT. LOOPED SUTURE LA WAS SIGNIFICANTLY MORE COST EFFICIENT THAN STAPLER LA.
05-12-2018 | Injury | covidien lp llc north haven | 02-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC RIGHT UPPER LOBE RESECTION, WHEN THE DEVICE WAS BEING APPLIED TO THE PULMONARY FISSURE AND BLOOD VESSELS, THE GREEN RELOAD HAD POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE AT THE CENTRAL PART. SO A SECOND RELOAD WAS USED, BUT IT DID NOT FIRE. THE STAPLE LINE WAS FIXED BY SUTURING. A NEW DEVICE WAS USED TO COMPLETE THE CASE.
05-12-2018 | Injury | covidien lp llc north haven | 02-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC RIGHT UPPER LOBE RESECTION, WHEN THE DEVICE WAS BEING APPLIED TO THE PULMONARY FISSURE AND BLOOD VESSELS, THE GREEN RELOAD HAD POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE AT THE CENTRAL PART. SO A SECOND RELOAD WAS USED, BUT IT DID NOT FIRE. THE STAPLE LINE WAS FIXED BY SUTURING. A NEW DEVICE WAS USED TO COMPLETE THE CASE.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | 45MM ARTICNG LNR CUTR/GRN RLD | Patient-Device Incompatibility | A0101 | Abscess | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE TITLE: ANALYSIS OF ENDOLOOPS AND ENDOSTAPLES FOR CLOSING THE APPENDICEAL STUMP DURING LAPAROSCOPIC APPENDECTOMY AUTHORS: MISLAV RAKIC, MIRO JUKIC, ZENON POGORELIC, IVANA MRKLIC, ROBERT KLICEK, NIKICA DRUZIJANIC, ZDRAVKO PERKO AND LEONARDO PATRLJ CITATION: SURG TODAY (2014) 44:1716¿1722; DOI 10.1007/S00595-013-0818-8; PUBLISHED ONLINE: 12 DECEMBER 2013. THE AIM OF THIS STUDY WAS TO ASSESS THE EFFICIENCY OF EACH TECHNIQUE BY FOCUSING ON THE INTRAOPERATIVE AND POSTOPERATIVE COMPLICATIONS VERSUS COST. BETWEEN JUNE 2011 AND JUNE 2013, A TOTAL OF 333 PATIENTS (188 MALES AND 155 FEMALES WITH AGE RANGE FROM 5 TO 80 YEARS OLD) WHO UNDERWENT LAPAROSCOPIC APPENDECTOMY (LA) WERE INCLUDED. THE DATA WERE GROUPED ACCORDING TO THE STUMP CLOSURE TECHNIQUE; A STAPLER GROUP (ENDOPATH-ENDOCUTTER ATG45, ETHICON ENDO-SURGERY, CINCINNATI, OH, USA) (104 PATIENTS; 31.2 %) AND AN ENDOLOOP GROUP (VICRYL-ENDOLOOP 0; ETHICON ENDO-SURGERY, (B)(6) USA) (229 PATIENTS; 68.8 %). THE MESOAPPENDIX WAS DISSECTED, DEPENDING ON THE SURGEONS¿ PREFERENCES, WITH EITHER A HARMONIC SCALPEL (ULTRACISIONTM, ETHICON ENDO-SURGERY, (B)(6) USA), A BIPOLAR TISSUE SEALING SYSTEM (LIGASURETM, VALLEYLAB, BOULDER, CO, USA) OR USING THERMAL FUSION TECHNOLOGY (MISEALTM, MICROLINE). AMONG THE 333 PATIENTS WHO UNDERWENT LA, THERE WAS ONE CASE OF BLEEDING FROM THE MESOAPPENDIX). POSTOPERATIVE COMPLICATIONS WERE SEVEN (2.1 %) WOUND INFECTIONS (FOUR IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), SEVEN (2.1 %) INTRAABDOMINAL ABSCESSES (THREE IN THE STAPLER GROUP, FOUR IN THE LOOP GROUP), FIVE (1.5 %) SMALL BOWEL OBSTRUCTIONS (TWO IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), FIVE (1.5 %) SMALL BOWEL OBSTRUCTIONS (TWO IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), AND ONE (0.3 %) APPENDIX STUMP INSUFFICIENCY IN THE LOOP GROUP. THE AUTHORS PERFORMED A REOPERATION IN ONLY THREE PATIENTS. IN CONCLUSION, THERE IS NO CLINICAL EVIDENCE SUPPORTING THE ROUTINE USE OF ENDOSCOPIC STAPLERS. THE APPENDICEAL STUMP CAN BE SECURED SAFELY WITH THE USE OF ENDOLOOPS IN THE MAJORITY OF PATIENTS. SURGEONS HAVE TO BE MORE SELECTIVE WHEN CHOOSING HOW TO PERFORM CLOSURE, AND AN ENDOSTAPLER SHOULD BE USED ONLY IN CASES WHERE IT IS CLINICALLY INDICATED.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | 45MM ARTICNG LNR CUTR/GRN RLD | Patient-Device Incompatibility | A0101 | Obstruction/Occlusion | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE TITLE: ANALYSIS OF ENDOLOOPS AND ENDOSTAPLES FOR CLOSING THE APPENDICEAL STUMP DURING LAPAROSCOPIC APPENDECTOMY AUTHORS: MISLAV RAKIC, MIRO JUKIC, ZENON POGORELIC, IVANA MRKLIC, ROBERT KLICEK, NIKICA DRUZIJANIC, ZDRAVKO PERKO AND LEONARDO PATRLJ CITATION: SURG TODAY (2014) 44:1716¿1722; DOI 10.1007/S00595-013-0818-8; PUBLISHED ONLINE: 12 DECEMBER 2013. THE AIM OF THIS STUDY WAS TO ASSESS THE EFFICIENCY OF EACH TECHNIQUE BY FOCUSING ON THE INTRAOPERATIVE AND POSTOPERATIVE COMPLICATIONS VERSUS COST. BETWEEN JUNE 2011 AND JUNE 2013, A TOTAL OF 333 PATIENTS (188 MALES AND 155 FEMALES WITH AGE RANGE FROM 5 TO 80 YEARS OLD) WHO UNDERWENT LAPAROSCOPIC APPENDECTOMY (LA) WERE INCLUDED. THE DATA WERE GROUPED ACCORDING TO THE STUMP CLOSURE TECHNIQUE; A STAPLER GROUP (ENDOPATH-ENDOCUTTER ATG45, ETHICON ENDO-SURGERY, CINCINNATI, OH, USA) (104 PATIENTS; 31.2 %) AND AN ENDOLOOP GROUP (VICRYL-ENDOLOOP 0; ETHICON ENDO-SURGERY, (B)(6) USA) (229 PATIENTS; 68.8 %). THE MESOAPPENDIX WAS DISSECTED, DEPENDING ON THE SURGEONS¿ PREFERENCES, WITH EITHER A HARMONIC SCALPEL (ULTRACISIONTM, ETHICON ENDO-SURGERY, (B)(6) USA), A BIPOLAR TISSUE SEALING SYSTEM (LIGASURETM, VALLEYLAB, BOULDER, CO, USA) OR USING THERMAL FUSION TECHNOLOGY (MISEALTM, MICROLINE). AMONG THE 333 PATIENTS WHO UNDERWENT LA, THERE WAS ONE CASE OF BLEEDING FROM THE MESOAPPENDIX). POSTOPERATIVE COMPLICATIONS WERE SEVEN (2.1 %) WOUND INFECTIONS (FOUR IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), SEVEN (2.1 %) INTRAABDOMINAL ABSCESSES (THREE IN THE STAPLER GROUP, FOUR IN THE LOOP GROUP), FIVE (1.5 %) SMALL BOWEL OBSTRUCTIONS (TWO IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), FIVE (1.5 %) SMALL BOWEL OBSTRUCTIONS (TWO IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), AND ONE (0.3 %) APPENDIX STUMP INSUFFICIENCY IN THE LOOP GROUP. THE AUTHORS PERFORMED A REOPERATION IN ONLY THREE PATIENTS. IN CONCLUSION, THERE IS NO CLINICAL EVIDENCE SUPPORTING THE ROUTINE USE OF ENDOSCOPIC STAPLERS. THE APPENDICEAL STUMP CAN BE SECURED SAFELY WITH THE USE OF ENDOLOOPS IN THE MAJORITY OF PATIENTS. SURGEONS HAVE TO BE MORE SELECTIVE WHEN CHOOSING HOW TO PERFORM CLOSURE, AND AN ENDOSTAPLER SHOULD BE USED ONLY IN CASES WHERE IT IS CLINICALLY INDICATED.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | 45MM ARTICNG LNR CUTR/GRN RLD | Patient-Device Incompatibility | A0101 | Blood Loss | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE TITLE: ANALYSIS OF ENDOLOOPS AND ENDOSTAPLES FOR CLOSING THE APPENDICEAL STUMP DURING LAPAROSCOPIC APPENDECTOMY AUTHORS: MISLAV RAKIC, MIRO JUKIC, ZENON POGORELIC, IVANA MRKLIC, ROBERT KLICEK, NIKICA DRUZIJANIC, ZDRAVKO PERKO AND LEONARDO PATRLJ CITATION: SURG TODAY (2014) 44:1716¿1722; DOI 10.1007/S00595-013-0818-8; PUBLISHED ONLINE: 12 DECEMBER 2013. THE AIM OF THIS STUDY WAS TO ASSESS THE EFFICIENCY OF EACH TECHNIQUE BY FOCUSING ON THE INTRAOPERATIVE AND POSTOPERATIVE COMPLICATIONS VERSUS COST. BETWEEN JUNE 2011 AND JUNE 2013, A TOTAL OF 333 PATIENTS (188 MALES AND 155 FEMALES WITH AGE RANGE FROM 5 TO 80 YEARS OLD) WHO UNDERWENT LAPAROSCOPIC APPENDECTOMY (LA) WERE INCLUDED. THE DATA WERE GROUPED ACCORDING TO THE STUMP CLOSURE TECHNIQUE; A STAPLER GROUP (ENDOPATH-ENDOCUTTER ATG45, ETHICON ENDO-SURGERY, CINCINNATI, OH, USA) (104 PATIENTS; 31.2 %) AND AN ENDOLOOP GROUP (VICRYL-ENDOLOOP 0; ETHICON ENDO-SURGERY, (B)(6) USA) (229 PATIENTS; 68.8 %). THE MESOAPPENDIX WAS DISSECTED, DEPENDING ON THE SURGEONS¿ PREFERENCES, WITH EITHER A HARMONIC SCALPEL (ULTRACISIONTM, ETHICON ENDO-SURGERY, (B)(6) USA), A BIPOLAR TISSUE SEALING SYSTEM (LIGASURETM, VALLEYLAB, BOULDER, CO, USA) OR USING THERMAL FUSION TECHNOLOGY (MISEALTM, MICROLINE). AMONG THE 333 PATIENTS WHO UNDERWENT LA, THERE WAS ONE CASE OF BLEEDING FROM THE MESOAPPENDIX). POSTOPERATIVE COMPLICATIONS WERE SEVEN (2.1 %) WOUND INFECTIONS (FOUR IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), SEVEN (2.1 %) INTRAABDOMINAL ABSCESSES (THREE IN THE STAPLER GROUP, FOUR IN THE LOOP GROUP), FIVE (1.5 %) SMALL BOWEL OBSTRUCTIONS (TWO IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), FIVE (1.5 %) SMALL BOWEL OBSTRUCTIONS (TWO IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), AND ONE (0.3 %) APPENDIX STUMP INSUFFICIENCY IN THE LOOP GROUP. THE AUTHORS PERFORMED A REOPERATION IN ONLY THREE PATIENTS. IN CONCLUSION, THERE IS NO CLINICAL EVIDENCE SUPPORTING THE ROUTINE USE OF ENDOSCOPIC STAPLERS. THE APPENDICEAL STUMP CAN BE SECURED SAFELY WITH THE USE OF ENDOLOOPS IN THE MAJORITY OF PATIENTS. SURGEONS HAVE TO BE MORE SELECTIVE WHEN CHOOSING HOW TO PERFORM CLOSURE, AND AN ENDOSTAPLER SHOULD BE USED ONLY IN CASES WHERE IT IS CLINICALLY INDICATED.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | 45MM ARTICNG LNR CUTR/GRN RLD | Patient-Device Incompatibility | A0101 | No Code Available | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE TITLE: ANALYSIS OF ENDOLOOPS AND ENDOSTAPLES FOR CLOSING THE APPENDICEAL STUMP DURING LAPAROSCOPIC APPENDECTOMY AUTHORS: MISLAV RAKIC, MIRO JUKIC, ZENON POGORELIC, IVANA MRKLIC, ROBERT KLICEK, NIKICA DRUZIJANIC, ZDRAVKO PERKO AND LEONARDO PATRLJ CITATION: SURG TODAY (2014) 44:1716¿1722; DOI 10.1007/S00595-013-0818-8; PUBLISHED ONLINE: 12 DECEMBER 2013. THE AIM OF THIS STUDY WAS TO ASSESS THE EFFICIENCY OF EACH TECHNIQUE BY FOCUSING ON THE INTRAOPERATIVE AND POSTOPERATIVE COMPLICATIONS VERSUS COST. BETWEEN JUNE 2011 AND JUNE 2013, A TOTAL OF 333 PATIENTS (188 MALES AND 155 FEMALES WITH AGE RANGE FROM 5 TO 80 YEARS OLD) WHO UNDERWENT LAPAROSCOPIC APPENDECTOMY (LA) WERE INCLUDED. THE DATA WERE GROUPED ACCORDING TO THE STUMP CLOSURE TECHNIQUE; A STAPLER GROUP (ENDOPATH-ENDOCUTTER ATG45, ETHICON ENDO-SURGERY, CINCINNATI, OH, USA) (104 PATIENTS; 31.2 %) AND AN ENDOLOOP GROUP (VICRYL-ENDOLOOP 0; ETHICON ENDO-SURGERY, (B)(6) USA) (229 PATIENTS; 68.8 %). THE MESOAPPENDIX WAS DISSECTED, DEPENDING ON THE SURGEONS¿ PREFERENCES, WITH EITHER A HARMONIC SCALPEL (ULTRACISIONTM, ETHICON ENDO-SURGERY, (B)(6) USA), A BIPOLAR TISSUE SEALING SYSTEM (LIGASURETM, VALLEYLAB, BOULDER, CO, USA) OR USING THERMAL FUSION TECHNOLOGY (MISEALTM, MICROLINE). AMONG THE 333 PATIENTS WHO UNDERWENT LA, THERE WAS ONE CASE OF BLEEDING FROM THE MESOAPPENDIX). POSTOPERATIVE COMPLICATIONS WERE SEVEN (2.1 %) WOUND INFECTIONS (FOUR IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), SEVEN (2.1 %) INTRAABDOMINAL ABSCESSES (THREE IN THE STAPLER GROUP, FOUR IN THE LOOP GROUP), FIVE (1.5 %) SMALL BOWEL OBSTRUCTIONS (TWO IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), FIVE (1.5 %) SMALL BOWEL OBSTRUCTIONS (TWO IN THE STAPLER GROUP, THREE IN THE LOOP GROUP), AND ONE (0.3 %) APPENDIX STUMP INSUFFICIENCY IN THE LOOP GROUP. THE AUTHORS PERFORMED A REOPERATION IN ONLY THREE PATIENTS. IN CONCLUSION, THERE IS NO CLINICAL EVIDENCE SUPPORTING THE ROUTINE USE OF ENDOSCOPIC STAPLERS. THE APPENDICEAL STUMP CAN BE SECURED SAFELY WITH THE USE OF ENDOLOOPS IN THE MAJORITY OF PATIENTS. SURGEONS HAVE TO BE MORE SELECTIVE WHEN CHOOSING HOW TO PERFORM CLOSURE, AND AN ENDOSTAPLER SHOULD BE USED ONLY IN CASES WHERE IT IS CLINICALLY INDICATED.
05-12-2018 | Injury | us surgical puerto rico | 02-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC RIGHT UPPER LOBE RESECTION, WHEN THE DEVICE WAS BEING APPLIED TO THE PULMONARY FISSURE AND BLOOD VESSELS, THE GREEN RELOAD HAD POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE AT THE CENTRAL PART. SO A SECOND RELOAD WAS USED, BUT IT DID NOT FIRE. THE STAPLE LINE WAS FIXED BY SUTURING. A NEW DEVICE WAS USED TO COMPLETE THE CASE.
05-12-2018 | Injury | us surgical puerto rico | 02-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORASCOPIC RIGHT UPPER LOBE RESECTION, WHEN THE DEVICE WAS BEING APPLIED TO THE PULMONARY FISSURE AND BLOOD VESSELS, THE GREEN RELOAD HAD POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE AT THE CENTRAL PART. SO A SECOND RELOAD WAS USED, BUT IT DID NOT FIRE. THE STAPLE LINE WAS FIXED BY SUTURING. A NEW DEVICE WAS USED TO COMPLETE THE CASE.
05-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PULMONARY SEGMENTECTOMY, DURING STAPLING OF FIBROSIS AREA WITH BLACK LOAD AND AFTER THE CLOSING AND BEGINNING OF THE FIRING, THERE WAS A CRACK WITH LOSS OF THE CAPACITY OF THE STAPLER. THE DEVICE WAS REPLACED WITH ANOTHER LOT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PULMONARY SEGMENTECTOMY, DURING STAPLING OF FIBROSIS AREA WITH BLACK LOAD AND AFTER THE CLOSING AND BEGINNING OF THE FIRING, THERE WAS A CRACK WITH LOSS OF THE CAPACITY OF THE STAPLER. THE DEVICE WAS REPLACED WITH ANOTHER LOT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PULMONARY SEGMENTECTOMY, DURING STAPLING OF FIBROSIS AREA WITH BLACK LOAD AND AFTER THE CLOSING AND BEGINNING OF THE FIRING, THERE WAS A CRACK WITH LOSS OF THE CAPACITY OF THE STAPLER. THE DEVICE WAS REPLACED WITH ANOTHER LOT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PULMONARY SEGMENTECTOMY, DURING STAPLING OF FIBROSIS AREA WITH BLACK LOAD AND AFTER THE CLOSING AND BEGINNING OF THE FIRING, THERE WAS A CRACK WITH LOSS OF THE CAPACITY OF THE STAPLER. THE DEVICE WAS REPLACED WITH ANOTHER LOT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PULMONARY SEGMENTECTOMY, DURING STAPLING OF FIBROSIS AREA WITH BLACK LOAD AND AFTER THE CLOSING AND BEGINNING OF THE FIRING, THERE WAS A CRACK WITH LOSS OF THE CAPACITY OF THE STAPLER. THE DEVICE WAS REPLACED WITH ANOTHER LOT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PULMONARY SEGMENTECTOMY, DURING STAPLING OF FIBROSIS AREA WITH BLACK LOAD AND AFTER THE CLOSING AND BEGINNING OF THE FIRING, THERE WAS A CRACK WITH LOSS OF THE CAPACITY OF THE STAPLER. THE DEVICE WAS REPLACED WITH ANOTHER LOT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | UNKNOWN ENDO GIA SULU | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PULMONARY SEGMENTECTOMY, DURING STAPLING OF FIBROSIS AREA WITH BLACK LOAD AND AFTER THE CLOSING AND BEGINNING OF THE FIRING, THERE WAS A CRACK WITH LOSS OF THE CAPACITY OF THE STAPLER. THE DEVICE WAS REPLACED WITH ANOTHER LOT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | UNKNOWN ENDO GIA SULU | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PULMONARY SEGMENTECTOMY, DURING STAPLING OF FIBROSIS AREA WITH BLACK LOAD AND AFTER THE CLOSING AND BEGINNING OF THE FIRING, THERE WAS A CRACK WITH LOSS OF THE CAPACITY OF THE STAPLER. THE DEVICE WAS REPLACED WITH ANOTHER LOT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | GREEN RELOAD FOR ECHELON 60 | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING LAPAROSCOPIC LOBECTOMY, IT WAS SEEN THAT AFTER THE CARTRIDGE WAS FIRED, THE STAPLES DIDN'T FORM DECENT B FORMATION, RESULTING A LEAKAGE FROM THE PARENCHYMA. THE SURGERY WAS SWITCHED TO OPEN AND THE LEAKAGE WAS SUTURED.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | GREEN RELOAD FOR ECHELON 60 | Fluid/Blood Leak | A050401 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LAPAROSCOPIC LOBECTOMY, IT WAS SEEN THAT AFTER THE CARTRIDGE WAS FIRED, THE STAPLES DIDN'T FORM DECENT B FORMATION, RESULTING A LEAKAGE FROM THE PARENCHYMA. THE SURGERY WAS SWITCHED TO OPEN AND THE LEAKAGE WAS SUTURED.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING LAPAROSCOPIC LOBECTOMY, IT WAS SEEN THAT AFTER THE CARTRIDGE WAS FIRED, THE STAPLES DIDN'T FORM DECENT B FORMATION, RESULTING A LEAKAGE FROM THE PARENCHYMA. THE SURGERY WAS SWITCHED TO OPEN AND THE LEAKAGE WAS SUTURED.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LAPAROSCOPIC LOBECTOMY, IT WAS SEEN THAT AFTER THE CARTRIDGE WAS FIRED, THE STAPLES DIDN'T FORM DECENT B FORMATION, RESULTING A LEAKAGE FROM THE PARENCHYMA. THE SURGERY WAS SWITCHED TO OPEN AND THE LEAKAGE WAS SUTURED.
06-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Contamination | A1801 | Abscess | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE: TITLE: "CURRENT ANALYSIS OF ENDOLOOPS IN APPENDICEAL STUMP CLOSURE." AUTHORS: MAIK SAHM, RAINER KUBE, SYBILLE SCHMIDT, CHRISTINA RITTER, MATTHIAS PROSS, HANS LIPPERT. CITATION: SURG ENDOSC. 2011; 25: 124¿129. DOI: 10.1007/S00464-010-1144-5. AN INADEQUATE CLOSURE OF THE APPENDIX STUMP LEADS TO INTRA-ABDOMINAL SURGICAL SITE INFECTION. THE EFFECTIVENESS OF VARIOUS APPENDICEAL STUMP CLOSURE METHODS, FOR INSTANCE, STAPLERS OR ENDOLOOPS, WAS EVALUATED. THE AIM OF THIS STUDY WAS TO EVALUATE THE COMPLICATION AFTER USING ENDOLOOPS IN A HIGH-VOLUME CENTER. LAPAROSCOPIC APPENDECTOMY WAS PERFORMED IN 1,790 PATIENTS AND OPEN APPENDECTOMY IN 425 PATIENTS. IN LAPAROSCOPIC APPENDECTOMY, STUMP CLOSURE USING ENDOLOOPS (ETHICON) WAS PERFORMED IN 1670 PATIENTS (690 MALE AND 980 FEMALE PATIENTS; AGE RANGE: 8 TO 95 YEARS OLD; BMI 12 TO 50) AND STUMP CLOSURE USING STAPLER WAS USED IN 46 PATIENTS (31 MALE AND 15 FEMALE PATIENTS; AGE RANGE: 11 TO 79 YEARS OLD; BMI 18 TO 36). IN THE LAPAROSCOPIC APPROACH, THE MESOAPPENDIX WAS DISSECTED USING BIPOLAR COAGULATION AND LAPAROSCOPIC SCISSORS. THE AUTHORS PLACED 1 ENDOLOOP LIGATURE PDS II (ETHICON) TO THE APPENDIX BASE. AN ADDITIONAL LOOP WAS PLACED 10 MM DISTALLY. THE BASE WAS CUT WITH LAPAROSCOPIC SCISSORS BETWEEN THE LOOPS. WE DISINFECTED THE APPENDICEAL STUMP WITH POLYVIDONE-IODINE. WHEN THE AUTHORS USED A STAPLER FOR APPENDICEAL STUMP CLOSURE, THE AUTHORS TRANSECTED THE APPENDIX BASE WITH THE ENDOSCOPIC LINEAR CUTTER, WITH ENDOSCOPIC LINEAR CUTTER (ETHICON). IN THE ENDOLOOP GROUP, REPORTED COMPLICATIONS INCLUDED SEROMA (N-13), ABSCESS (N-21), INTRAABDOMINAL ABSCESS (N-27) WHICH REQUIRED RELAPAROSCOPY, RELAPAROTOMY, OR CT DRAINAGE, AND ANTIBIOTIC THERAPY, INSUFFICIENT STUMP CLOSURE (N-2), POSTOPERATIVE BLEEDING (N-3), POSTOPERATIVE PERITONITIS (N-3), MECHANICAL SMALL BOWEL OBSTRUCTION (N-13), AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-23). IN THE STAPLER GROUP, REPORTED COMPLICATIONS INCLUDED INTRAABDOMINAL ABSCESS (N-1) WHICH REQUIRED RELAPAROSCOPY, AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-3). THE STUDY SHOWED THAT THE TREATMENT OF THE APPENDICEAL BASE CAN BE DONE SAFELY WITH THE USE OF ENDOLOOPS (ETHICON) IN THE MAJORITY OF PATIENTS AND A SELECTIVE USE OF A STAPLER IN A SMALL MINORITY OF PATIENTS. APPENDICEAL STUMP CLOSURE USING AN ENDOLOOP IS AN EASY, SAFE, AND COST EFFECTIVE PROCEDURE.
06-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Contamination | A1801 | Obstruction/Occlusion | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE: TITLE: "CURRENT ANALYSIS OF ENDOLOOPS IN APPENDICEAL STUMP CLOSURE." AUTHORS: MAIK SAHM, RAINER KUBE, SYBILLE SCHMIDT, CHRISTINA RITTER, MATTHIAS PROSS, HANS LIPPERT. CITATION: SURG ENDOSC. 2011; 25: 124¿129. DOI: 10.1007/S00464-010-1144-5. AN INADEQUATE CLOSURE OF THE APPENDIX STUMP LEADS TO INTRA-ABDOMINAL SURGICAL SITE INFECTION. THE EFFECTIVENESS OF VARIOUS APPENDICEAL STUMP CLOSURE METHODS, FOR INSTANCE, STAPLERS OR ENDOLOOPS, WAS EVALUATED. THE AIM OF THIS STUDY WAS TO EVALUATE THE COMPLICATION AFTER USING ENDOLOOPS IN A HIGH-VOLUME CENTER. LAPAROSCOPIC APPENDECTOMY WAS PERFORMED IN 1,790 PATIENTS AND OPEN APPENDECTOMY IN 425 PATIENTS. IN LAPAROSCOPIC APPENDECTOMY, STUMP CLOSURE USING ENDOLOOPS (ETHICON) WAS PERFORMED IN 1670 PATIENTS (690 MALE AND 980 FEMALE PATIENTS; AGE RANGE: 8 TO 95 YEARS OLD; BMI 12 TO 50) AND STUMP CLOSURE USING STAPLER WAS USED IN 46 PATIENTS (31 MALE AND 15 FEMALE PATIENTS; AGE RANGE: 11 TO 79 YEARS OLD; BMI 18 TO 36). IN THE LAPAROSCOPIC APPROACH, THE MESOAPPENDIX WAS DISSECTED USING BIPOLAR COAGULATION AND LAPAROSCOPIC SCISSORS. THE AUTHORS PLACED 1 ENDOLOOP LIGATURE PDS II (ETHICON) TO THE APPENDIX BASE. AN ADDITIONAL LOOP WAS PLACED 10 MM DISTALLY. THE BASE WAS CUT WITH LAPAROSCOPIC SCISSORS BETWEEN THE LOOPS. WE DISINFECTED THE APPENDICEAL STUMP WITH POLYVIDONE-IODINE. WHEN THE AUTHORS USED A STAPLER FOR APPENDICEAL STUMP CLOSURE, THE AUTHORS TRANSECTED THE APPENDIX BASE WITH THE ENDOSCOPIC LINEAR CUTTER, WITH ENDOSCOPIC LINEAR CUTTER (ETHICON). IN THE ENDOLOOP GROUP, REPORTED COMPLICATIONS INCLUDED SEROMA (N-13), ABSCESS (N-21), INTRAABDOMINAL ABSCESS (N-27) WHICH REQUIRED RELAPAROSCOPY, RELAPAROTOMY, OR CT DRAINAGE, AND ANTIBIOTIC THERAPY, INSUFFICIENT STUMP CLOSURE (N-2), POSTOPERATIVE BLEEDING (N-3), POSTOPERATIVE PERITONITIS (N-3), MECHANICAL SMALL BOWEL OBSTRUCTION (N-13), AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-23). IN THE STAPLER GROUP, REPORTED COMPLICATIONS INCLUDED INTRAABDOMINAL ABSCESS (N-1) WHICH REQUIRED RELAPAROSCOPY, AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-3). THE STUDY SHOWED THAT THE TREATMENT OF THE APPENDICEAL BASE CAN BE DONE SAFELY WITH THE USE OF ENDOLOOPS (ETHICON) IN THE MAJORITY OF PATIENTS AND A SELECTIVE USE OF A STAPLER IN A SMALL MINORITY OF PATIENTS. APPENDICEAL STUMP CLOSURE USING AN ENDOLOOP IS AN EASY, SAFE, AND COST EFFECTIVE PROCEDURE.
06-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Contamination | A1801 | No Code Available | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE: TITLE: "CURRENT ANALYSIS OF ENDOLOOPS IN APPENDICEAL STUMP CLOSURE." AUTHORS: MAIK SAHM, RAINER KUBE, SYBILLE SCHMIDT, CHRISTINA RITTER, MATTHIAS PROSS, HANS LIPPERT. CITATION: SURG ENDOSC. 2011; 25: 124¿129. DOI: 10.1007/S00464-010-1144-5. AN INADEQUATE CLOSURE OF THE APPENDIX STUMP LEADS TO INTRA-ABDOMINAL SURGICAL SITE INFECTION. THE EFFECTIVENESS OF VARIOUS APPENDICEAL STUMP CLOSURE METHODS, FOR INSTANCE, STAPLERS OR ENDOLOOPS, WAS EVALUATED. THE AIM OF THIS STUDY WAS TO EVALUATE THE COMPLICATION AFTER USING ENDOLOOPS IN A HIGH-VOLUME CENTER. LAPAROSCOPIC APPENDECTOMY WAS PERFORMED IN 1,790 PATIENTS AND OPEN APPENDECTOMY IN 425 PATIENTS. IN LAPAROSCOPIC APPENDECTOMY, STUMP CLOSURE USING ENDOLOOPS (ETHICON) WAS PERFORMED IN 1670 PATIENTS (690 MALE AND 980 FEMALE PATIENTS; AGE RANGE: 8 TO 95 YEARS OLD; BMI 12 TO 50) AND STUMP CLOSURE USING STAPLER WAS USED IN 46 PATIENTS (31 MALE AND 15 FEMALE PATIENTS; AGE RANGE: 11 TO 79 YEARS OLD; BMI 18 TO 36). IN THE LAPAROSCOPIC APPROACH, THE MESOAPPENDIX WAS DISSECTED USING BIPOLAR COAGULATION AND LAPAROSCOPIC SCISSORS. THE AUTHORS PLACED 1 ENDOLOOP LIGATURE PDS II (ETHICON) TO THE APPENDIX BASE. AN ADDITIONAL LOOP WAS PLACED 10 MM DISTALLY. THE BASE WAS CUT WITH LAPAROSCOPIC SCISSORS BETWEEN THE LOOPS. WE DISINFECTED THE APPENDICEAL STUMP WITH POLYVIDONE-IODINE. WHEN THE AUTHORS USED A STAPLER FOR APPENDICEAL STUMP CLOSURE, THE AUTHORS TRANSECTED THE APPENDIX BASE WITH THE ENDOSCOPIC LINEAR CUTTER, WITH ENDOSCOPIC LINEAR CUTTER (ETHICON). IN THE ENDOLOOP GROUP, REPORTED COMPLICATIONS INCLUDED SEROMA (N-13), ABSCESS (N-21), INTRAABDOMINAL ABSCESS (N-27) WHICH REQUIRED RELAPAROSCOPY, RELAPAROTOMY, OR CT DRAINAGE, AND ANTIBIOTIC THERAPY, INSUFFICIENT STUMP CLOSURE (N-2), POSTOPERATIVE BLEEDING (N-3), POSTOPERATIVE PERITONITIS (N-3), MECHANICAL SMALL BOWEL OBSTRUCTION (N-13), AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-23). IN THE STAPLER GROUP, REPORTED COMPLICATIONS INCLUDED INTRAABDOMINAL ABSCESS (N-1) WHICH REQUIRED RELAPAROSCOPY, AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-3). THE STUDY SHOWED THAT THE TREATMENT OF THE APPENDICEAL BASE CAN BE DONE SAFELY WITH THE USE OF ENDOLOOPS (ETHICON) IN THE MAJORITY OF PATIENTS AND A SELECTIVE USE OF A STAPLER IN A SMALL MINORITY OF PATIENTS. APPENDICEAL STUMP CLOSURE USING AN ENDOLOOP IS AN EASY, SAFE, AND COST EFFECTIVE PROCEDURE.
06-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | Abscess | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE: TITLE: "CURRENT ANALYSIS OF ENDOLOOPS IN APPENDICEAL STUMP CLOSURE." AUTHORS: MAIK SAHM, RAINER KUBE, SYBILLE SCHMIDT, CHRISTINA RITTER, MATTHIAS PROSS, HANS LIPPERT. CITATION: SURG ENDOSC. 2011; 25: 124¿129. DOI: 10.1007/S00464-010-1144-5. AN INADEQUATE CLOSURE OF THE APPENDIX STUMP LEADS TO INTRA-ABDOMINAL SURGICAL SITE INFECTION. THE EFFECTIVENESS OF VARIOUS APPENDICEAL STUMP CLOSURE METHODS, FOR INSTANCE, STAPLERS OR ENDOLOOPS, WAS EVALUATED. THE AIM OF THIS STUDY WAS TO EVALUATE THE COMPLICATION AFTER USING ENDOLOOPS IN A HIGH-VOLUME CENTER. LAPAROSCOPIC APPENDECTOMY WAS PERFORMED IN 1,790 PATIENTS AND OPEN APPENDECTOMY IN 425 PATIENTS. IN LAPAROSCOPIC APPENDECTOMY, STUMP CLOSURE USING ENDOLOOPS (ETHICON) WAS PERFORMED IN 1670 PATIENTS (690 MALE AND 980 FEMALE PATIENTS; AGE RANGE: 8 TO 95 YEARS OLD; BMI 12 TO 50) AND STUMP CLOSURE USING STAPLER WAS USED IN 46 PATIENTS (31 MALE AND 15 FEMALE PATIENTS; AGE RANGE: 11 TO 79 YEARS OLD; BMI 18 TO 36). IN THE LAPAROSCOPIC APPROACH, THE MESOAPPENDIX WAS DISSECTED USING BIPOLAR COAGULATION AND LAPAROSCOPIC SCISSORS. THE AUTHORS PLACED 1 ENDOLOOP LIGATURE PDS II (ETHICON) TO THE APPENDIX BASE. AN ADDITIONAL LOOP WAS PLACED 10 MM DISTALLY. THE BASE WAS CUT WITH LAPAROSCOPIC SCISSORS BETWEEN THE LOOPS. WE DISINFECTED THE APPENDICEAL STUMP WITH POLYVIDONE-IODINE. WHEN THE AUTHORS USED A STAPLER FOR APPENDICEAL STUMP CLOSURE, THE AUTHORS TRANSECTED THE APPENDIX BASE WITH THE ENDOSCOPIC LINEAR CUTTER, WITH ENDOSCOPIC LINEAR CUTTER (ETHICON). IN THE ENDOLOOP GROUP, REPORTED COMPLICATIONS INCLUDED SEROMA (N-13), ABSCESS (N-21), INTRAABDOMINAL ABSCESS (N-27) WHICH REQUIRED RELAPAROSCOPY, RELAPAROTOMY, OR CT DRAINAGE, AND ANTIBIOTIC THERAPY, INSUFFICIENT STUMP CLOSURE (N-2), POSTOPERATIVE BLEEDING (N-3), POSTOPERATIVE PERITONITIS (N-3), MECHANICAL SMALL BOWEL OBSTRUCTION (N-13), AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-23). IN THE STAPLER GROUP, REPORTED COMPLICATIONS INCLUDED INTRAABDOMINAL ABSCESS (N-1) WHICH REQUIRED RELAPAROSCOPY, AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-3). THE STUDY SHOWED THAT THE TREATMENT OF THE APPENDICEAL BASE CAN BE DONE SAFELY WITH THE USE OF ENDOLOOPS (ETHICON) IN THE MAJORITY OF PATIENTS AND A SELECTIVE USE OF A STAPLER IN A SMALL MINORITY OF PATIENTS. APPENDICEAL STUMP CLOSURE USING AN ENDOLOOP IS AN EASY, SAFE, AND COST EFFECTIVE PROCEDURE.
06-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | Obstruction/Occlusion | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE: TITLE: "CURRENT ANALYSIS OF ENDOLOOPS IN APPENDICEAL STUMP CLOSURE." AUTHORS: MAIK SAHM, RAINER KUBE, SYBILLE SCHMIDT, CHRISTINA RITTER, MATTHIAS PROSS, HANS LIPPERT. CITATION: SURG ENDOSC. 2011; 25: 124¿129. DOI: 10.1007/S00464-010-1144-5. AN INADEQUATE CLOSURE OF THE APPENDIX STUMP LEADS TO INTRA-ABDOMINAL SURGICAL SITE INFECTION. THE EFFECTIVENESS OF VARIOUS APPENDICEAL STUMP CLOSURE METHODS, FOR INSTANCE, STAPLERS OR ENDOLOOPS, WAS EVALUATED. THE AIM OF THIS STUDY WAS TO EVALUATE THE COMPLICATION AFTER USING ENDOLOOPS IN A HIGH-VOLUME CENTER. LAPAROSCOPIC APPENDECTOMY WAS PERFORMED IN 1,790 PATIENTS AND OPEN APPENDECTOMY IN 425 PATIENTS. IN LAPAROSCOPIC APPENDECTOMY, STUMP CLOSURE USING ENDOLOOPS (ETHICON) WAS PERFORMED IN 1670 PATIENTS (690 MALE AND 980 FEMALE PATIENTS; AGE RANGE: 8 TO 95 YEARS OLD; BMI 12 TO 50) AND STUMP CLOSURE USING STAPLER WAS USED IN 46 PATIENTS (31 MALE AND 15 FEMALE PATIENTS; AGE RANGE: 11 TO 79 YEARS OLD; BMI 18 TO 36). IN THE LAPAROSCOPIC APPROACH, THE MESOAPPENDIX WAS DISSECTED USING BIPOLAR COAGULATION AND LAPAROSCOPIC SCISSORS. THE AUTHORS PLACED 1 ENDOLOOP LIGATURE PDS II (ETHICON) TO THE APPENDIX BASE. AN ADDITIONAL LOOP WAS PLACED 10 MM DISTALLY. THE BASE WAS CUT WITH LAPAROSCOPIC SCISSORS BETWEEN THE LOOPS. WE DISINFECTED THE APPENDICEAL STUMP WITH POLYVIDONE-IODINE. WHEN THE AUTHORS USED A STAPLER FOR APPENDICEAL STUMP CLOSURE, THE AUTHORS TRANSECTED THE APPENDIX BASE WITH THE ENDOSCOPIC LINEAR CUTTER, WITH ENDOSCOPIC LINEAR CUTTER (ETHICON). IN THE ENDOLOOP GROUP, REPORTED COMPLICATIONS INCLUDED SEROMA (N-13), ABSCESS (N-21), INTRAABDOMINAL ABSCESS (N-27) WHICH REQUIRED RELAPAROSCOPY, RELAPAROTOMY, OR CT DRAINAGE, AND ANTIBIOTIC THERAPY, INSUFFICIENT STUMP CLOSURE (N-2), POSTOPERATIVE BLEEDING (N-3), POSTOPERATIVE PERITONITIS (N-3), MECHANICAL SMALL BOWEL OBSTRUCTION (N-13), AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-23). IN THE STAPLER GROUP, REPORTED COMPLICATIONS INCLUDED INTRAABDOMINAL ABSCESS (N-1) WHICH REQUIRED RELAPAROSCOPY, AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-3). THE STUDY SHOWED THAT THE TREATMENT OF THE APPENDICEAL BASE CAN BE DONE SAFELY WITH THE USE OF ENDOLOOPS (ETHICON) IN THE MAJORITY OF PATIENTS AND A SELECTIVE USE OF A STAPLER IN A SMALL MINORITY OF PATIENTS. APPENDICEAL STUMP CLOSURE USING AN ENDOLOOP IS AN EASY, SAFE, AND COST EFFECTIVE PROCEDURE.
06-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ENDOSCOPIC LNR CUTR (EXACT CODE UNKNOWN) | Patient-Device Incompatibility | A0101 | No Code Available | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE: TITLE: "CURRENT ANALYSIS OF ENDOLOOPS IN APPENDICEAL STUMP CLOSURE." AUTHORS: MAIK SAHM, RAINER KUBE, SYBILLE SCHMIDT, CHRISTINA RITTER, MATTHIAS PROSS, HANS LIPPERT. CITATION: SURG ENDOSC. 2011; 25: 124¿129. DOI: 10.1007/S00464-010-1144-5. AN INADEQUATE CLOSURE OF THE APPENDIX STUMP LEADS TO INTRA-ABDOMINAL SURGICAL SITE INFECTION. THE EFFECTIVENESS OF VARIOUS APPENDICEAL STUMP CLOSURE METHODS, FOR INSTANCE, STAPLERS OR ENDOLOOPS, WAS EVALUATED. THE AIM OF THIS STUDY WAS TO EVALUATE THE COMPLICATION AFTER USING ENDOLOOPS IN A HIGH-VOLUME CENTER. LAPAROSCOPIC APPENDECTOMY WAS PERFORMED IN 1,790 PATIENTS AND OPEN APPENDECTOMY IN 425 PATIENTS. IN LAPAROSCOPIC APPENDECTOMY, STUMP CLOSURE USING ENDOLOOPS (ETHICON) WAS PERFORMED IN 1670 PATIENTS (690 MALE AND 980 FEMALE PATIENTS; AGE RANGE: 8 TO 95 YEARS OLD; BMI 12 TO 50) AND STUMP CLOSURE USING STAPLER WAS USED IN 46 PATIENTS (31 MALE AND 15 FEMALE PATIENTS; AGE RANGE: 11 TO 79 YEARS OLD; BMI 18 TO 36). IN THE LAPAROSCOPIC APPROACH, THE MESOAPPENDIX WAS DISSECTED USING BIPOLAR COAGULATION AND LAPAROSCOPIC SCISSORS. THE AUTHORS PLACED 1 ENDOLOOP LIGATURE PDS II (ETHICON) TO THE APPENDIX BASE. AN ADDITIONAL LOOP WAS PLACED 10 MM DISTALLY. THE BASE WAS CUT WITH LAPAROSCOPIC SCISSORS BETWEEN THE LOOPS. WE DISINFECTED THE APPENDICEAL STUMP WITH POLYVIDONE-IODINE. WHEN THE AUTHORS USED A STAPLER FOR APPENDICEAL STUMP CLOSURE, THE AUTHORS TRANSECTED THE APPENDIX BASE WITH THE ENDOSCOPIC LINEAR CUTTER, WITH ENDOSCOPIC LINEAR CUTTER (ETHICON). IN THE ENDOLOOP GROUP, REPORTED COMPLICATIONS INCLUDED SEROMA (N-13), ABSCESS (N-21), INTRAABDOMINAL ABSCESS (N-27) WHICH REQUIRED RELAPAROSCOPY, RELAPAROTOMY, OR CT DRAINAGE, AND ANTIBIOTIC THERAPY, INSUFFICIENT STUMP CLOSURE (N-2), POSTOPERATIVE BLEEDING (N-3), POSTOPERATIVE PERITONITIS (N-3), MECHANICAL SMALL BOWEL OBSTRUCTION (N-13), AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-23). IN THE STAPLER GROUP, REPORTED COMPLICATIONS INCLUDED INTRAABDOMINAL ABSCESS (N-1) WHICH REQUIRED RELAPAROSCOPY, AND PARALYTIC SMALL BOWEL OBSTRUCTION (N-3). THE STUDY SHOWED THAT THE TREATMENT OF THE APPENDICEAL BASE CAN BE DONE SAFELY WITH THE USE OF ENDOLOOPS (ETHICON) IN THE MAJORITY OF PATIENTS AND A SELECTIVE USE OF A STAPLER IN A SMALL MINORITY OF PATIENTS. APPENDICEAL STUMP CLOSURE USING AN ENDOLOOP IS AN EASY, SAFE, AND COST EFFECTIVE PROCEDURE.
05-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE COLECTOMY, DURING SURGICAL STAPLING, THE DEVICE DID NOT FIRE. IT WAS REPORTED THAT THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE COLECTOMY, DURING SURGICAL STAPLING, THE DEVICE DID NOT FIRE. IT WAS REPORTED THAT THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | UNKNOWN GIA PRODUCT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE COLECTOMY, DURING SURGICAL STAPLING, THE DEVICE DID NOT FIRE. IT WAS REPORTED THAT THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. THERE WAS NO PATIENT INJURY.
10-10-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ROTICULATOR | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCTECTOMY PROCEDURE, ON STAPLING PHASE, THE STAPLER DID NOT FIRE. IT WAS ENABLE TO STAPLE ON THE RECTUM. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Injury | covidien lp llc north haven | 02-01-2019 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE OF HEMORRHOIDOPEXY, THE DEVICE FIRED PARTIALLY, LESS THAN A QUARTER OF THE STAPLING WAS DONE. THE SURGEON HAND SEW THE PART THAT WAS NOT STAPLED.
06-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ETS*35 ENDO LINEAR CUTTER BL | Insufficient Information | E2401 | Pleural Effusion | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE: TITLE: "VIDEOTHORACOSCOPIC TREATMENT OF PRIMARY SPONTANEOUS PNEUMOTHORAX: A 6-YEAR EXPERIENCE." AUTHORS: GIUSEPPE CARDILLO, MD, FRANCESCO FACCIOLO, MD, ROBERTO GIUNTI, MD, ROBERTO GASPARRI, MD, MICHELE LOPERGOLO, MD, REMO ORSETTI, MD, AND MASSIMO MARTELLI, MD. CITATION: ANN THORAC SURG 2000; 69:357¿ 62. FROM JULY 1, 1992, TO MAY 1, 1998, OF 1,307 VIDEO-ASSISTED THORACIC SURGERY (VATS) PROCEDURES PERFORMED AT THE DEPARTMENT OF THORACIC SURGERY, FORLANINI HOSPITAL IN ROME, 432 PATIENTS (33.1%) UNDERWENT VATS TREATMENT OF PRIMARY SPONTANEOUS PNEUMOTHORAX (PSP). ALL PROCEDURES WERE PERFORMED WITH THE PATIENT UNDER GENERAL ANESTHESIA AND ONE LUNG VENTILATION. THE PATIENT WAS PLACED IN THE LATERAL POSITION WITH THE ARM HELD ABDUCTED TO ALLOW MAXIMUM SUPERIOR DISPLACEMENT OF THE SCAPULA. DURING THE TALC POUDRAGE PROCEDURE, WITH REGARDS THE TREATMENT OF THE BULLAE, IT INCLUDED ALWAYS GRASPING AT THE BASE TOGETHER WITH HEALTHY TISSUE; IN STAGE III, BULLAE WERE LIGATED WITH A 0 OR 00 COATED VICRYL ROEDER LOOP (ENDOLOOP EJ 10G, ETHICON INC, SOMERVILLE, NJ) OR TREATED BY MINIMAL WEDGE RESECTION WITH THE ENDOSCOPIC STAPLER (TSB 35, ETHICON ENDOSURGERY, CINCINNATI, OH). AT THE END OF THE SURGICAL PROCEDURE, USUALLY ONE DRAIN (28 FRENCH GAUGE) WAS PLACED THROUGH THE EXISTING INCISIONS UNDER ENDOSCOPIC CONTROL WITH A BLUNT GUIDE. IF THE WATER-SEAL REVEALED A MAJOR AIR LEAK, THE LUNG IS ROUTINELY REINSPECTED BY TELESCOPE (WATER-SEAL TEST). POSTOPERATIVE COMPLICATION RATE WAS 4.16% (18 OF 432); IT INCLUDED 4 CASES OF SUBCUTANEOUS EMPHYSEMA, 5 CASES OF LOCALIZED PLEURAL EFFUSIONS, 6 PROLONGED AIR LEAKS (.5 DAYS), 1 MINIMAL PLEURAL DETACHMENT, 1 TRANSIENT BERNARD-HORNER SYNDROME. ONE LOCALIZED PLEURAL EFFUSION OCCURRED AFTER ISOLATE TALC POUDRAGE. PLEURAL EFFUSION WAS TREATED BY THORACENTESIS. THE TRANSIENT BERNARD-HORNER SYNDROME OCCURRED AFTER MONOPOLAR COAGULATION OF APICAL PLEUROPULMONARY ADHESIONS. PROLONGED AIR LEAKS OCCURRED AFTER LIGATION OF THE BULLAE (4 CASES) AND AFTER STAPLING OF THE BULLAE (2 CASES, STAGE IV). THEY RESOLVED SPONTANEOUSLY. SUBCUTANEOUS EMPHYSEMA OCCURRED AT THE BEGINNING OF OUR EXPERIENCE: 3 CASES FOLLOWING LIGATION OF THE BULLAE AND 1 CASE FOLLOWING STAPLING OF THE BULLAE. MINIMAL APICAL PLEURAL DETACHMENT OCCURRED 3 DAYS AFTER ISOLATE TALC POUDRAGE AND RESOLVED SPONTANEOUSLY. RECURRENCES OCCURRED IN 10 CASES (52.6%) WITHIN 1 YEAR FROM SURGERY. AT 2 YEARS FROM SURGERY, 84.2% OF RECURRENCES (16 CASES) WERE PRESENT. RECURRENCES WERE TREATED BY RE-VATS (12 PATIENTS), STANDARD THORACOTOMY (3 PATIENTS), CHEST DRAINAGE (3 PATIENTS), AND OBSERVATION (1 PATIENT). IN CONCLUSION, IN A FIRST EPISODE OF PSP, PLEURAL DRAINAGE REPRESENTS THE STANDARD TREATMENT; IN CASE OF BILATERAL SIMULTANEOUS PNEUMOTHORAX, WE SUGGEST PLEURAL DRAINAGE AND CONTRALATERAL VATS TREATMENT (BECAUSE OF THE HIGH MORBIDITY OF BILATERAL SIMULTANEOUS RECURRENCE). VATS BECOMES OUR TREATMENT OF CHOICE IN PATIENTS WITH SPONTANEOUS HEMO-PNEUMOTHORAX, IN PATIENTS WITH FIRST EPISODE OF PSP AND PREVIOUS CONTRALATERAL PNEUMOTHORAX, IN PATIENTS WHOSE PROFESSIONS PUT THEM AT RISK (PILOTS, SCUBA DIVERS, SPORTSMEN), IN PATIENTS WITH TREATMENT FAILURE OF A FIRST EPISODE OF PSP, AND IN PATIENTS WITH RECURRENT SP. IN RECURRENCES FOLLOWING VATS, WE BELIEVE THAT REDO-VATS WILL PROBABLY REPRESENT THE TREATMENT OF CHOICE, EVEN IF, UP TO THE PRESENT TIME, THERE HAS BEEN VERY LITTLE EXPERIENCE WITH THIS TREATMENT. A MINIMUM 2-YEAR FOLLOW-UP IS NECESSARY FOR THE EVALUATION OF THE RESULTS.
06-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ETS*35 ENDO LINEAR CUTTER BL | Insufficient Information | E2401 | No Code Available | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE: TITLE: "VIDEOTHORACOSCOPIC TREATMENT OF PRIMARY SPONTANEOUS PNEUMOTHORAX: A 6-YEAR EXPERIENCE." AUTHORS: GIUSEPPE CARDILLO, MD, FRANCESCO FACCIOLO, MD, ROBERTO GIUNTI, MD, ROBERTO GASPARRI, MD, MICHELE LOPERGOLO, MD, REMO ORSETTI, MD, AND MASSIMO MARTELLI, MD. CITATION: ANN THORAC SURG 2000; 69:357¿ 62. FROM JULY 1, 1992, TO MAY 1, 1998, OF 1,307 VIDEO-ASSISTED THORACIC SURGERY (VATS) PROCEDURES PERFORMED AT THE DEPARTMENT OF THORACIC SURGERY, FORLANINI HOSPITAL IN ROME, 432 PATIENTS (33.1%) UNDERWENT VATS TREATMENT OF PRIMARY SPONTANEOUS PNEUMOTHORAX (PSP). ALL PROCEDURES WERE PERFORMED WITH THE PATIENT UNDER GENERAL ANESTHESIA AND ONE LUNG VENTILATION. THE PATIENT WAS PLACED IN THE LATERAL POSITION WITH THE ARM HELD ABDUCTED TO ALLOW MAXIMUM SUPERIOR DISPLACEMENT OF THE SCAPULA. DURING THE TALC POUDRAGE PROCEDURE, WITH REGARDS THE TREATMENT OF THE BULLAE, IT INCLUDED ALWAYS GRASPING AT THE BASE TOGETHER WITH HEALTHY TISSUE; IN STAGE III, BULLAE WERE LIGATED WITH A 0 OR 00 COATED VICRYL ROEDER LOOP (ENDOLOOP EJ 10G, ETHICON INC, SOMERVILLE, NJ) OR TREATED BY MINIMAL WEDGE RESECTION WITH THE ENDOSCOPIC STAPLER (TSB 35, ETHICON ENDOSURGERY, CINCINNATI, OH). AT THE END OF THE SURGICAL PROCEDURE, USUALLY ONE DRAIN (28 FRENCH GAUGE) WAS PLACED THROUGH THE EXISTING INCISIONS UNDER ENDOSCOPIC CONTROL WITH A BLUNT GUIDE. IF THE WATER-SEAL REVEALED A MAJOR AIR LEAK, THE LUNG IS ROUTINELY REINSPECTED BY TELESCOPE (WATER-SEAL TEST). POSTOPERATIVE COMPLICATION RATE WAS 4.16% (18 OF 432); IT INCLUDED 4 CASES OF SUBCUTANEOUS EMPHYSEMA, 5 CASES OF LOCALIZED PLEURAL EFFUSIONS, 6 PROLONGED AIR LEAKS (.5 DAYS), 1 MINIMAL PLEURAL DETACHMENT, 1 TRANSIENT BERNARD-HORNER SYNDROME. ONE LOCALIZED PLEURAL EFFUSION OCCURRED AFTER ISOLATE TALC POUDRAGE. PLEURAL EFFUSION WAS TREATED BY THORACENTESIS. THE TRANSIENT BERNARD-HORNER SYNDROME OCCURRED AFTER MONOPOLAR COAGULATION OF APICAL PLEUROPULMONARY ADHESIONS. PROLONGED AIR LEAKS OCCURRED AFTER LIGATION OF THE BULLAE (4 CASES) AND AFTER STAPLING OF THE BULLAE (2 CASES, STAGE IV). THEY RESOLVED SPONTANEOUSLY. SUBCUTANEOUS EMPHYSEMA OCCURRED AT THE BEGINNING OF OUR EXPERIENCE: 3 CASES FOLLOWING LIGATION OF THE BULLAE AND 1 CASE FOLLOWING STAPLING OF THE BULLAE. MINIMAL APICAL PLEURAL DETACHMENT OCCURRED 3 DAYS AFTER ISOLATE TALC POUDRAGE AND RESOLVED SPONTANEOUSLY. RECURRENCES OCCURRED IN 10 CASES (52.6%) WITHIN 1 YEAR FROM SURGERY. AT 2 YEARS FROM SURGERY, 84.2% OF RECURRENCES (16 CASES) WERE PRESENT. RECURRENCES WERE TREATED BY RE-VATS (12 PATIENTS), STANDARD THORACOTOMY (3 PATIENTS), CHEST DRAINAGE (3 PATIENTS), AND OBSERVATION (1 PATIENT). IN CONCLUSION, IN A FIRST EPISODE OF PSP, PLEURAL DRAINAGE REPRESENTS THE STANDARD TREATMENT; IN CASE OF BILATERAL SIMULTANEOUS PNEUMOTHORAX, WE SUGGEST PLEURAL DRAINAGE AND CONTRALATERAL VATS TREATMENT (BECAUSE OF THE HIGH MORBIDITY OF BILATERAL SIMULTANEOUS RECURRENCE). VATS BECOMES OUR TREATMENT OF CHOICE IN PATIENTS WITH SPONTANEOUS HEMO-PNEUMOTHORAX, IN PATIENTS WITH FIRST EPISODE OF PSP AND PREVIOUS CONTRALATERAL PNEUMOTHORAX, IN PATIENTS WHOSE PROFESSIONS PUT THEM AT RISK (PILOTS, SCUBA DIVERS, SPORTSMEN), IN PATIENTS WITH TREATMENT FAILURE OF A FIRST EPISODE OF PSP, AND IN PATIENTS WITH RECURRENT SP. IN RECURRENCES FOLLOWING VATS, WE BELIEVE THAT REDO-VATS WILL PROBABLY REPRESENT THE TREATMENT OF CHOICE, EVEN IF, UP TO THE PRESENT TIME, THERE HAS BEEN VERY LITTLE EXPERIENCE WITH THIS TREATMENT. A MINIMUM 2-YEAR FOLLOW-UP IS NECESSARY FOR THE EVALUATION OF THE RESULTS.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Cut | A050702 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY WHEN THE SURGEON WAS STAPLING ACROSS THE PULMONARY VESSELS, THE STAPLE FIRED BUT DID NOT CUT THE TISSUE. AN ADDITIONAL RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY WHEN THE SURGEON WAS STAPLING ACROSS THE PULMONARY VESSELS, THE STAPLE FIRED BUT DID NOT CUT THE TISSUE. AN ADDITIONAL RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
01-11-2018 | Malfunction | ethicon endo-surgery | 02-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: SEE USER FACILITY MEDWATCH FORM, # (B)(4).
01-11-2018 | Malfunction | ethicon endo-surgery | 02-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: SEE USER FACILITY MEDWATCH FORM, # (B)(4).
04-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | EEA | Difficult to Remove | A150207 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOIDECTOMY, THERE WAS DIFFICULTY REMOVING THE STAPLER FROM THE CAVITY.
16-10-2018 | Injury | ethicon endo-surgery | 02-01-2019 | 75MM SELECTABLE NEW TLC | Firing Problem | A0505 | Fistula | Event Description: IT WAS REPORTED THAT DURING A GASTROPLASTY PROCEDURE, DURING THE SURGICAL PROCEDURE OF THE PATIENT, THE HOSPITAL EQUIPMENT (STAPLER) WAS BLOCKED, WHICH MADE THE SAFE AND EFFECTIVE SUTURES / CLOSURES OF THE ORGANS (STOMACH) OF THE PATIENT IMPOSSIBLE. THE PATIENT HAD A POSTOPERATIVE FISTULA.
16-10-2018 | Injury | ethicon endo-surgery | 02-01-2019 | 75MM SELECTABLE NEW TLC | Complete Blockage | A140901 | Fistula | Event Description: IT WAS REPORTED THAT DURING A GASTROPLASTY PROCEDURE, DURING THE SURGICAL PROCEDURE OF THE PATIENT, THE HOSPITAL EQUIPMENT (STAPLER) WAS BLOCKED, WHICH MADE THE SAFE AND EFFECTIVE SUTURES / CLOSURES OF THE ORGANS (STOMACH) OF THE PATIENT IMPOSSIBLE. THE PATIENT HAD A POSTOPERATIVE FISTULA.
03-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY ON A LAPAROSCOPY ON THE LOW RECTUM, THE STAPLES DID NOT STIMULATE, THE DEVICE WAS NOT USED IN THE PROCEDURE. THEY OPENED A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY ON A LAPAROSCOPY ON THE LOW RECTUM, THE STAPLES DID NOT STIMULATE, THE DEVICE WAS NOT USED IN THE PROCEDURE. THEY OPENED A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY ON A LAPAROSCOPY ON THE LOW RECTUM, THE STAPLES DID NOT STIMULATE, THE DEVICE WAS NOT USED IN THE PROCEDURE. THEY OPENED A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Misfire | A050502 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING AN DONOR NEPHRECTOMY PROCEDURE THAT THE STAPLER MIS-FIRED OVER THE RENAL ARTERY. UNSURE AS TO EXACTLY HOW THE DEVICE MIS-FIRED. EXTENSIVE BLEEDING CAUSING THE PROCEDURE TO CONVERT FROM LAPAROSCOPIC TO OPEN.
05-12-2018 | Injury | ethicon endo-surgery | 02-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Misfire | A050502 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN DONOR NEPHRECTOMY PROCEDURE THAT THE STAPLER MIS-FIRED OVER THE RENAL ARTERY. UNSURE AS TO EXACTLY HOW THE DEVICE MIS-FIRED. EXTENSIVE BLEEDING CAUSING THE PROCEDURE TO CONVERT FROM LAPAROSCOPIC TO OPEN.
01-01-2018 | Injury | ethicon endo-surgery | 02-01-2019 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A CASE OF OPEN TOTAL PANCREATECTOMY. THE CHIEF COMPLAINT IS BLEEDING FROM STAPLE LINE. PATIENT HAD NGT INSERTED AND DRAINAGE POD 1 WAS FRESH BLOOD. BLEEDING STAPLE LINE WAS AT GASTRO JEJUNOSTOMY ANASTOMOSIS. PATIENT WENT FOR GASTROSCOPY AND ENDOSCOPIC CLIP X 3 WAS APPLIED ON BLEEDING SITE. THERE WERE NO GAPPING OF ANASTOMOSIS. DURING OPERATION, ECR60G WAS USED FOR ANASTOMOSIS. PATIENT NOT ON ANTI COAGULANT DRUGS. PATIENT IS NOW WELL POST ENDOSCOPIC CLIPPING.
01-01-2018 | Injury | ethicon endo-surgery | 02-01-2019 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A CASE OF OPEN TOTAL PANCREATECTOMY. THE CHIEF COMPLAINT IS BLEEDING FROM STAPLE LINE. PATIENT HAD NGT INSERTED AND DRAINAGE POD 1 WAS FRESH BLOOD. BLEEDING STAPLE LINE WAS AT GASTRO JEJUNOSTOMY ANASTOMOSIS. PATIENT WENT FOR GASTROSCOPY AND ENDOSCOPIC CLIP X 3 WAS APPLIED ON BLEEDING SITE. THERE WERE NO GAPPING OF ANASTOMOSIS. DURING OPERATION, ECR60G WAS USED FOR ANASTOMOSIS. PATIENT NOT ON ANTI COAGULANT DRUGS. PATIENT IS NOW WELL POST ENDOSCOPIC CLIPPING.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS NOT ABLE TO FIRE. THE REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER WAS NOT ABLE TO FIRE. THE REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-11-2018 | Malfunction | ethicon endo-surgery | 02-01-2019 | ETS*45 ENDO LIN CUT RELOAD WH | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DOCTOR REPORTED THAT AFTER FIRING THE STAPLER IN THE APPENDIX, IT DID NOT CLOSE IN THE MIDDLE PART OF THE TISSUE. THE DOCTOR CLOSED THE TISSUE WITH STRATAFIX SUTURE. NO PATIENT CONSEQUENCE REPORTED.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE ONLY FIRED HALFWAY. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE ONLY FIRED HALFWAY. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE ONLY FIRED HALFWAY. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ENDOSCOPIC LIVER PROCEDURE, DURING THE CUTTING OF THE LIVER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE SURGEON COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. THEY REPLACED THE RELOAD TO COMPLETE THE CASE. NO PATIENT HARM.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ENDOSCOPIC LIVER PROCEDURE, DURING THE CUTTING OF THE LIVER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE SURGEON COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. THEY REPLACED THE RELOAD TO COMPLETE THE CASE. NO PATIENT HARM.
05-12-2018 | Injury | us surgical puerto rico | 02-01-2019 | EEA | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SIGMOIDECTOMY PROCEDURE, ON THE FIRST FIRING OF THE STAPLER, THE STAPLER WAS DIFFICULT TO REMOVE FROM ANASTOMOSIS AND UPON THE REMOVAL STAPLE LINE WAS INCOMPLETE. TWO COMPLETE DONUTS, BUT NEEDED TO DO HAND SEWN ANASTAMOSIS. SURGICAL TIME WAS EXTENDED TO 45 MINUTES TO HAND SEW THE ANASTOMOSIS.
05-12-2018 | Injury | us surgical puerto rico | 02-01-2019 | EEA | Difficult to Remove | A150207 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SIGMOIDECTOMY PROCEDURE, ON THE FIRST FIRING OF THE STAPLER, THE STAPLER WAS DIFFICULT TO REMOVE FROM ANASTOMOSIS AND UPON THE REMOVAL STAPLE LINE WAS INCOMPLETE. TWO COMPLETE DONUTS, BUT NEEDED TO DO HAND SEWN ANASTAMOSIS. SURGICAL TIME WAS EXTENDED TO 45 MINUTES TO HAND SEW THE ANASTOMOSIS.
05-12-2018 | Injury | us surgical puerto rico | 02-01-2019 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SIGMOIDECTOMY PROCEDURE, ON THE FIRST FIRING OF THE STAPLER, THE STAPLER WAS DIFFICULT TO REMOVE FROM ANASTOMOSIS AND UPON THE REMOVAL STAPLE LINE WAS INCOMPLETE. TWO COMPLETE DONUTS, BUT NEEDED TO DO HAND SEWN ANASTAMOSIS. SURGICAL TIME WAS EXTENDED TO 45 MINUTES TO HAND SEW THE ANASTOMOSIS.
05-12-2018 | Injury | us surgical puerto rico | 02-01-2019 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SIGMOIDECTOMY PROCEDURE, ON THE FIRST FIRING OF THE STAPLER, THE STAPLER WAS DIFFICULT TO REMOVE FROM ANASTOMOSIS AND UPON THE REMOVAL STAPLE LINE WAS INCOMPLETE. TWO COMPLETE DONUTS, BUT NEEDED TO DO HAND SEWN ANASTAMOSIS. SURGICAL TIME WAS EXTENDED TO 45 MINUTES TO HAND SEW THE ANASTOMOSIS.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE RESECTION OF BULLAE ON A VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE ONLY FIRED 2 OR 3 PINS AND FORMED 1 OR TWO PINS RESULTING IN AN INCOMPLETE AND POOR STAPLE LINE. IT WAS ALSO STATED THAT THE ARTICULATING KNOB AS BROKEN. THEY USED NEW CARTRIDGES AND STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE RESECTION OF BULLAE ON A VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE ONLY FIRED 2 OR 3 PINS AND FORMED 1 OR TWO PINS RESULTING IN AN INCOMPLETE AND POOR STAPLE LINE. IT WAS ALSO STATED THAT THE ARTICULATING KNOB AS BROKEN. THEY USED NEW CARTRIDGES AND STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE RESECTION OF BULLAE ON A VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT, THE DEVICE ONLY FIRED 2 OR 3 PINS AND FORMED 1 OR TWO PINS RESULTING IN AN INCOMPLETE AND POOR STAPLE LINE. IT WAS ALSO STATED THAT THE ARTICULATING KNOB AS BROKEN. THEY USED NEW CARTRIDGES AND STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 02-01-2019 | ENDO GIA | Material Separation | A0413 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROINTESTINAL PROCEDURE, WHILE CUTTING THE STOMACH, DURING FIRING THEY NOTICED THAT SOME STAPLES WERE BROKEN. THEY REPLACED THE RELOAD AND THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT HARM.
03-12-2018 | Malfunction | covidien lp llc north haven | 02-01-2019 | PI | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCTECTOMY AND J POUCH, WHEN THE STAPLER WAS POSITIONED TO STAPLE THE RECTUM, THE BLACK BUTTON ON THE SIDE OF THE HANDLE FELL INTO THE PATIENT CAVITY. THE SURGEON EXPLORED THE PELVIS AND FOUND THE PIECE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
05-12-2018 | Malfunction | covidien medical products | 02-01-2019 | PREMIUM PLUS CEEA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A JEJUNAL ANASTOMOSIS, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | EEA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED LEFT COLECTOMY PROCEDURE, THE LEAK TEST SHOWED A LEAK AND INSPECTION UNDER THE STAPLE LINE WAS INCOMPLETE ON THE ANTERIOR SIDE. TO RESOLVE THE ISSUE, ADDITIONAL TRANSECTION WAS DONE LOWER ON THE RECTUM AND FIRED A SECOND DEVICE TO CREATE AN ANASTOMOSIS. THE EVENT LED TO AN UNANTICIPATED TISSUE LOSS, EXTENDED SURGICAL TIME FOR 30 MINUTES OR MORE DUE TO LONGER ADMINISTRATION OF ANESTHESIA.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED LEFT COLECTOMY PROCEDURE, THE LEAK TEST SHOWED A LEAK AND INSPECTION UNDER THE STAPLE LINE WAS INCOMPLETE ON THE ANTERIOR SIDE. TO RESOLVE THE ISSUE, ADDITIONAL TRANSECTION WAS DONE LOWER ON THE RECTUM AND FIRED A SECOND DEVICE TO CREATE AN ANASTOMOSIS. THE EVENT LED TO AN UNANTICIPATED TISSUE LOSS, EXTENDED SURGICAL TIME FOR 30 MINUTES OR MORE DUE TO LONGER ADMINISTRATION OF ANESTHESIA.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | EEA | Gas/Air Leak | A050402 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED LEFT COLECTOMY PROCEDURE, THE LEAK TEST SHOWED A LEAK AND INSPECTION UNDER THE STAPLE LINE WAS INCOMPLETE ON THE ANTERIOR SIDE. TO RESOLVE THE ISSUE, ADDITIONAL TRANSECTION WAS DONE LOWER ON THE RECTUM AND FIRED A SECOND DEVICE TO CREATE AN ANASTOMOSIS. THE EVENT LED TO AN UNANTICIPATED TISSUE LOSS, EXTENDED SURGICAL TIME FOR 30 MINUTES OR MORE DUE TO LONGER ADMINISTRATION OF ANESTHESIA.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | EEA | Gas/Air Leak | A050402 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED LEFT COLECTOMY PROCEDURE, THE LEAK TEST SHOWED A LEAK AND INSPECTION UNDER THE STAPLE LINE WAS INCOMPLETE ON THE ANTERIOR SIDE. TO RESOLVE THE ISSUE, ADDITIONAL TRANSECTION WAS DONE LOWER ON THE RECTUM AND FIRED A SECOND DEVICE TO CREATE AN ANASTOMOSIS. THE EVENT LED TO AN UNANTICIPATED TISSUE LOSS, EXTENDED SURGICAL TIME FOR 30 MINUTES OR MORE DUE TO LONGER ADMINISTRATION OF ANESTHESIA.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY PROCEDURE, WHILE USING THE DEVICE ON THE SMALL INTESTINE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE COULD NOT BE FIRED. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. NO PATIENT HARM.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY PROCEDURE, WHILE USING THE DEVICE ON THE SMALL INTESTINE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE COULD NOT BE FIRED. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2018 | Injury | ethicon endo-surgery | 03-01-2019 | ILS 33MM, CURVED | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN BOWEL RESECTION, THEY HAD TWO GOOD DONUTS WHEN THEY WERE PUTTING THE RECTUM AND THE COLON TOGETHER THEY TIGHTENED IT DOWN UNTIL THE INDICATOR WAS IN THE GREEN EVEN THE LOWER END OF THE GREEN, THEY FIRED, AND AS THEY WENT ON TO THE NEXT STEPS THEY REALIZED THAT THERE WAS ONLY ONE STAPLE, AND THEY SMELLED STOOL AND SAW STOOL. THE WHOLE ANASTOMISE FELL APART. THERE WAS STOOL IN THE PATIENTS ABDOMEN. AFTER THE FIRST ONE FAILED THEY WENT AND GOT A SECOND STAPLER. THEY SAID THEY HAD ONE GOOD DONUT AND ONE THAT WAS NOT AS GREAT, ONE SIDE SEEMED TO STAPLE WHILE THE OTHER SIDE SEEMED TO FAIL. CASE COMPLETED BY DOING A LOOP ILEOSTOMY, CASE EXTENDED BY TWO HOURS. PATIENT POST OPERATIVELY HAD SEPTIC SHOCK DUE TO STOOL CONTAMINATION.
08-12-2018 | Injury | ethicon endo-surgery | 03-01-2019 | ILS 33MM, CURVED | Failure to Form Staple | A050704 | Septic Shock | Event Description: IT WAS REPORTED THAT DURING AN BOWEL RESECTION, THEY HAD TWO GOOD DONUTS WHEN THEY WERE PUTTING THE RECTUM AND THE COLON TOGETHER THEY TIGHTENED IT DOWN UNTIL THE INDICATOR WAS IN THE GREEN EVEN THE LOWER END OF THE GREEN, THEY FIRED, AND AS THEY WENT ON TO THE NEXT STEPS THEY REALIZED THAT THERE WAS ONLY ONE STAPLE, AND THEY SMELLED STOOL AND SAW STOOL. THE WHOLE ANASTOMISE FELL APART. THERE WAS STOOL IN THE PATIENTS ABDOMEN. AFTER THE FIRST ONE FAILED THEY WENT AND GOT A SECOND STAPLER. THEY SAID THEY HAD ONE GOOD DONUT AND ONE THAT WAS NOT AS GREAT, ONE SIDE SEEMED TO STAPLE WHILE THE OTHER SIDE SEEMED TO FAIL. CASE COMPLETED BY DOING A LOOP ILEOSTOMY, CASE EXTENDED BY TWO HOURS. PATIENT POST OPERATIVELY HAD SEPTIC SHOCK DUE TO STOOL CONTAMINATION.
08-12-2018 | Injury | ethicon endo-surgery | 03-01-2019 | ILS 33MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN BOWEL RESECTION, THEY HAD TWO GOOD DONUTS WHEN THEY WERE PUTTING THE RECTUM AND THE COLON TOGETHER THEY TIGHTENED IT DOWN UNTIL THE INDICATOR WAS IN THE GREEN EVEN THE LOWER END OF THE GREEN, THEY FIRED, AND AS THEY WENT ON TO THE NEXT STEPS THEY REALIZED THAT THERE WAS ONLY ONE STAPLE, AND THEY SMELLED STOOL AND SAW STOOL. THE WHOLE ANASTOMISE FELL APART. THERE WAS STOOL IN THE PATIENTS ABDOMEN. AFTER THE FIRST ONE FAILED THEY WENT AND GOT A SECOND STAPLER. THEY SAID THEY HAD ONE GOOD DONUT AND ONE THAT WAS NOT AS GREAT, ONE SIDE SEEMED TO STAPLE WHILE THE OTHER SIDE SEEMED TO FAIL. CASE COMPLETED BY DOING A LOOP ILEOSTOMY, CASE EXTENDED BY TWO HOURS. PATIENT POST OPERATIVELY HAD SEPTIC SHOCK DUE TO STOOL CONTAMINATION.
08-12-2018 | Injury | ethicon endo-surgery | 03-01-2019 | ILS 33MM, CURVED | Defective Device | A0203 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN BOWEL RESECTION, THEY HAD TWO GOOD DONUTS WHEN THEY WERE PUTTING THE RECTUM AND THE COLON TOGETHER THEY TIGHTENED IT DOWN UNTIL THE INDICATOR WAS IN THE GREEN EVEN THE LOWER END OF THE GREEN, THEY FIRED, AND AS THEY WENT ON TO THE NEXT STEPS THEY REALIZED THAT THERE WAS ONLY ONE STAPLE, AND THEY SMELLED STOOL AND SAW STOOL. THE WHOLE ANASTOMISE FELL APART. THERE WAS STOOL IN THE PATIENT'S ABDOMEN. AFTER THE FIRST ONE FAILED THEY WENT AND GOT A SECOND STAPLER. THEY SAID THEY HAD ONE GOOD DONUT AND ONE THAT WAS NOT AS GREAT, ONE SIDE SEEMED TO STAPLE WHILE THE OTHER SIDE SEEMED TO FAIL. CASE COMPLETED BY DOING A LOOP ILEOSTOMY, CASE EXTENDED BY TWO HOURS. PATIENT POST OPERATIVELY HAD SEPTIC SHOCK DUE TO STOOL CONTAMINATION.
08-12-2018 | Injury | ethicon endo-surgery | 03-01-2019 | ILS 33MM, CURVED | Defective Device | A0203 | Septic Shock | Event Description: IT WAS REPORTED THAT DURING AN BOWEL RESECTION, THEY HAD TWO GOOD DONUTS WHEN THEY WERE PUTTING THE RECTUM AND THE COLON TOGETHER THEY TIGHTENED IT DOWN UNTIL THE INDICATOR WAS IN THE GREEN EVEN THE LOWER END OF THE GREEN, THEY FIRED, AND AS THEY WENT ON TO THE NEXT STEPS THEY REALIZED THAT THERE WAS ONLY ONE STAPLE, AND THEY SMELLED STOOL AND SAW STOOL. THE WHOLE ANASTOMISE FELL APART. THERE WAS STOOL IN THE PATIENT'S ABDOMEN. AFTER THE FIRST ONE FAILED THEY WENT AND GOT A SECOND STAPLER. THEY SAID THEY HAD ONE GOOD DONUT AND ONE THAT WAS NOT AS GREAT, ONE SIDE SEEMED TO STAPLE WHILE THE OTHER SIDE SEEMED TO FAIL. CASE COMPLETED BY DOING A LOOP ILEOSTOMY, CASE EXTENDED BY TWO HOURS. PATIENT POST OPERATIVELY HAD SEPTIC SHOCK DUE TO STOOL CONTAMINATION.
08-12-2018 | Injury | ethicon endo-surgery | 03-01-2019 | ILS 33MM, CURVED | Defective Device | A0203 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN BOWEL RESECTION, THEY HAD TWO GOOD DONUTS WHEN THEY WERE PUTTING THE RECTUM AND THE COLON TOGETHER THEY TIGHTENED IT DOWN UNTIL THE INDICATOR WAS IN THE GREEN EVEN THE LOWER END OF THE GREEN, THEY FIRED, AND AS THEY WENT ON TO THE NEXT STEPS THEY REALIZED THAT THERE WAS ONLY ONE STAPLE, AND THEY SMELLED STOOL AND SAW STOOL. THE WHOLE ANASTOMISE FELL APART. THERE WAS STOOL IN THE PATIENT'S ABDOMEN. AFTER THE FIRST ONE FAILED THEY WENT AND GOT A SECOND STAPLER. THEY SAID THEY HAD ONE GOOD DONUT AND ONE THAT WAS NOT AS GREAT, ONE SIDE SEEMED TO STAPLE WHILE THE OTHER SIDE SEEMED TO FAIL. CASE COMPLETED BY DOING A LOOP ILEOSTOMY, CASE EXTENDED BY TWO HOURS. PATIENT POST OPERATIVELY HAD SEPTIC SHOCK DUE TO STOOL CONTAMINATION.
04-11-2015 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON ABLE TO PRESSED THE GREEN BUTTON AND WAS ABLE TO SQUEEZED THE HANDLE BUT THE DEVICE DID NOT FIRE. THEY REPLACED THE PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-11-2015 | Malfunction | covidien lp llc north haven | 03-01-2019 | UNKNOWN ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON ABLE TO PRESSED THE GREEN BUTTON AND WAS ABLE TO SQUEEZED THE HANDLE BUT THE DEVICE DID NOT FIRE. THEY REPLACED THE PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-11-2018 | Injury | ethicon endo-surgery | 03-01-2019 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN OPERATION REQUIRING COLORECTAL ANASTOMOSIS, THE SURGEON WAS UNABLE TO STAPLE WITH THE DEVICE BECAUSE STAPLES WERE NOT WELL POSITIONED. THE CUTTING OF THE TISSUES WAS EFFECTIVE. THE PROCEDURE TURNED INTO A COLOSTOMY WITH RELEASE OF FAECES INTO THE PERITONEUM. PATIENT PUT ON ANTIBIOTIC.
08-11-2018 | Injury | ethicon endo-surgery | 03-01-2019 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING AN OPERATION REQUIRING COLORECTAL ANASTOMOSIS, THE SURGEON WAS UNABLE TO STAPLE WITH THE DEVICE BECAUSE STAPLES WERE NOT WELL POSITIONED. THE CUTTING OF THE TISSUES WAS EFFECTIVE. THE PROCEDURE TURNED INTO A COLOSTOMY WITH RELEASE OF FAECES INTO THE PERITONEUM. PATIENT PUT ON ANTIBIOTIC.
08-11-2018 | Injury | ethicon endo-surgery | 03-01-2019 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN OPERATION REQUIRING COLORECTAL ANASTOMOSIS, THE SURGEON WAS UNABLE TO STAPLE WITH THE DEVICE BECAUSE STAPLES WERE NOT WELL POSITIONED. THE CUTTING OF THE TISSUES WAS EFFECTIVE. THE PROCEDURE TURNED INTO A COLOSTOMY WITH RELEASE OF FAECES INTO THE PERITONEUM. PATIENT PUT ON ANTIBIOTIC.
17-08-2017 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE AND THE HANDLE WAS JAMMED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-08-2017 | Malfunction | covidien lp llc north haven | 03-01-2019 | UNKNOWN TRI STAPLE PRODUCT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE AND THE HANDLE WAS JAMMED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-10-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A LAPAROSCOPIC APPENDECTOMY, THE STAPLE RELOAD WOULD NOT LOCK INTO THE STAPLER AND A PIECE OF PLASTIC WAS BROKEN. THEY USED ANOTHER RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
26-10-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A LAPAROSCOPIC APPENDECTOMY, THE STAPLE RELOAD WOULD NOT LOCK INTO THE STAPLER AND A PIECE OF PLASTIC WAS BROKEN. THEY USED ANOTHER RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
01-01-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | PROXIMATE*PPH PROCEDURE SET | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A PPH PROCEDURE, THE FIRING LEVER ON THE PPH03 REFUSED TO RELEASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THERE IS NO FURTHER INFORMATION AVAILABLE.
01-03-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING, THE SURGEON ABLE TO PRESSED THE GREEN BUTTON, BUT CANNOT SQUEEZE THE HANDLE AND THE DEVICE WAS NOT ABLE TO FIRE. THE LIGASURE DEVICE'S HANDLE WAS UNABLE TO BE CLOSED AND THE KNIFE BLADE WAS DIFFICULT TO ADVANCE. USED A NEW PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-03-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING, THE SURGEON ABLE TO PRESSED THE GREEN BUTTON, BUT CANNOT SQUEEZE THE HANDLE AND THE DEVICE WAS NOT ABLE TO FIRE. THE LIGASURE DEVICE'S HANDLE WAS UNABLE TO BE CLOSED AND THE KNIFE BLADE WAS DIFFICULT TO ADVANCE. USED A NEW PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | PROXIMATE*PPH PROCEDURE SET | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE STAPLING PROCEDURE, THE STAPLE HAS ONLY STAPLED PART OF THE TISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
10-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, WHEN TRIGGERED THE DEVICE, THERE WAS NOT THE FAMILIAR SOUND WHILE DISRUPT THE PLASTIC RING AND THE DEVICE DID ONLY STAPLED BUT DID NOT CUT. ONLY AFTER OPENING AND CLOSING COULD BE CUT, BUT IT MADE OVER SEWING NECESSARY. NO HARM TO THE PATIENT.
05-05-2016 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLER BROKE AND WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
05-05-2016 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLER BROKE AND WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
05-05-2016 | Malfunction | covidien lp llc north haven | 03-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLER BROKE AND WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE MAGAZINE COULD BE TRIGGERED, THEN THE STAPLER COULD NOT BE OPENED. KNIFE DID NOT GO BACK. BY MEANS OF THE ACCESS FLAP, THE KNIFE COULD BE PULLED BACK MANUALLY. THE SURGERY COULD BE COMPLETED WITH A NEW DEVICE AND A NEW MAGAZINE. NO HARM TO THE PATIENT. TIME LOST UNKNOWN.
21-01-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AT THE 1ST FIRING THE DEVICE DID NOT FIRE. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | PWRD ECH FLEX 60MM 34CM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN ANTERIOR RESECTION PROCEDURE, COULD NOT RETURN BLADE. IT WAS REPORTED THAT DURING THE DIXON, COULD NOT RETURN THE BLADE AFTER FIRED. USED MANUAL OVERRIDE LEVER TO RETURN THE BLADE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
21-01-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AT THE 1ST FIRING THE DEVICE DID NOT FIRE. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ETS*45 ENDO LIN CUT RELOAD WH | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLE RELOAD WAS PARTIALLY FIRED AND WAS NOT FIRED ANY MORE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
11-05-2016 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE INSTRUMENT LOCKED EVEN WITHOUT PRESSING THE GREEN BUTTON AND WITH NORMAL OPENING WHEN FIRING WAS ATTEMPTED ON THE TWO RELOADS, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
11-05-2016 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE INSTRUMENT LOCKED EVEN WITHOUT PRESSING THE GREEN BUTTON AND WITH NORMAL OPENING WHEN FIRING WAS ATTEMPTED ON THE TWO RELOADS, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
11-05-2016 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE INSTRUMENT LOCKED EVEN WITHOUT PRESSING THE GREEN BUTTON AND WITH NORMAL OPENING WHEN FIRING WAS ATTEMPTED ON THE TWO RELOADS, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
10-01-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS PROCEDURE, AFTER FIRING CAN NOT OPEN JAW IN CASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | POWERED PLUS LONG 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, ACCORDING TO THE SURGICAL ASSIST, ON THE FIFTH AND SIXTH FIRINGS OF A LAPAROSCOPIC GASTRIC BYPASS, TWO BLUE LOADS WERE FIRED AND STAPLES ¿PULLED OUT¿ OF THE TISSUE AND APPEARED TO BE INCOMPLETELY FORMED. A SECOND STAPLER AND RELOAD WAS OPENED AND THE CASE WAS CONTINUED. ADDITIONAL TISSUE HAD TO BE RESECTED ON THE GASTRIC POUCH, BUT THE SURGEON DID NOT HAVE TO COMPROMISE THE ANATOMICAL CONSTRUCT THAT WAS INTENDED.
20-11-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE RELEASED THE CLIPS, BUT DIDN'T SUTURE. THE DEVICE WAS REPLACED. NO AES REPORTED.
02-01-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER DID NOT FIRE. ANOTHER STAPLER HAD TO BE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE LAST FIRING OF THE STAPLER, WITH A WHITE RELOAD, NO BUTTRESSING MATERIAL, THE STAPLER FIRED BUT THE JAWS WOULDN'T OPEN. SINCE THIS WAS THE LAST FIRING, THE DOCTOR WAS ABLE TO REMOVED THE STAPLER FROM TISSUE WITHOUT ANY ISSUES. NO ADDITIONAL DEVICES WERE NEEDED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
10-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING SETUP FOR A SLEEVE GASTRECTOMY THE SCRUB TECHNICIAN WAS HAVING PROBLEMS WITH THE CLOSING HANDLE. IT BECAME STUCK AND WOULD NOT OPEN. THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED.
06-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | FLEX 60 ARTICULATING | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: COULD NOT FIRE FARTHER. IT WAS REPORTED THAT DURING THE THORACOSCOPIC COMBINED WITH ESOPHAGECTOMY FOR ESOPHAGEAL CANCER, COULD NOT FIRE FARTHER AFTER FIRED A HALF WHEN SEVERING THE ESOPHAGUS AND COULD NOT RETURN THE BLADE. USED ANOTHER DEVICE TO PRIZE THE JAW AND MAKE BIGGER GAP TO REMOVE OUT THE DEVICE. THE RETURNED DEVICE WITH CLOSED THE JAW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, DURING THE PROCEDURE, THE STAPLING WAS NOT COMPLETE. ON ONE SIDE THE STAPLING WAS GOOD, BUT NOT ON THE OTHER SIDE. PART OF THE STOMACH WAS NOT STAPLED CAUSING THE OPENING OF THE GASTRIC CAVITY IN THE ABDOMINAL CAVITY. WHEN REMOVING THE REFILL, IT WAS RELEASED. TO FINISH THE PROCEDURE, USE OF SUTURE.
05-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT THE FIRST FIRING, THE STAPLE WAS PRE-FIRED AND WAS NOT ABLE TO BE USED. ANOTHER RELOAD WAS USED, HOWEVER IT WAS TOO STIFF AND JUST FIRED HALFWAY, THE STAPLES THAT WERE NOT ABLE TO BE FIRED FELL OFF INTO THE CAVITY AND WAS ABLE TO BE RETRIEVED. OOZING BLEEDING OCCURRED AS A RESULT OF THE ISSUE. THEY THEN USED ANOTHER RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT THE FIRST FIRING, THE STAPLE WAS PRE-FIRED AND WAS NOT ABLE TO BE USED. ANOTHER RELOAD WAS USED, HOWEVER IT WAS TOO STIFF AND JUST FIRED HALFWAY, THE STAPLES THAT WERE NOT ABLE TO BE FIRED FELL OFF INTO THE CAVITY AND WAS ABLE TO BE RETRIEVED. OOZING BLEEDING OCCURRED AS A RESULT OF THE ISSUE. THEY THEN USED ANOTHER RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, EC60A, THEY FIRED IT TWICE, THEN AFTER THAT THEY WEREN'T ABLE TO OPEN, CLOSE OR ROTATE THE STAPLER. THERE IS NO FURTHER INFORMATION AVAILABLE ABOUT THE EVENT. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
07-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | FLEX 60 ARTICULATING | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, EC60A, THEY FIRED IT TWICE, THEN AFTER THAT THEY WEREN'T ABLE TO OPEN, CLOSE OR ROTATE THE STAPLER. THERE IS NO FURTHER INFORMATION AVAILABLE ABOUT THE EVENT. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
05-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT THE FIRST FIRING, THE STAPLE WAS PRE-FIRED AND WAS NOT ABLE TO BE USED. ANOTHER RELOAD WAS USED, HOWEVER IT WAS TOO STIFF AND JUST FIRED HALFWAY, THE STAPLES THAT WERE NOT ABLE TO BE FIRED FELL OFF INTO THE CAVITY AND WAS ABLE TO BE RETRIEVED. OOZING BLEEDING OCCURRED AS A RESULT OF THE ISSUE. THEY THEN USED ANOTHER RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT THE FIRST FIRING, THE STAPLE WAS PRE-FIRED AND WAS NOT ABLE TO BE USED. ANOTHER RELOAD WAS USED, HOWEVER IT WAS TOO STIFF AND JUST FIRED HALFWAY, THE STAPLES THAT WERE NOT ABLE TO BE FIRED FELL OFF INTO THE CAVITY AND WAS ABLE TO BE RETRIEVED. OOZING BLEEDING OCCURRED AS A RESULT OF THE ISSUE. THEY THEN USED ANOTHER RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT THE FIRST FIRING, THE STAPLE WAS PRE-FIRED AND WAS NOT ABLE TO BE USED. ANOTHER RELOAD WAS USED, HOWEVER IT WAS TOO STIFF AND JUST FIRED HALFWAY, THE STAPLES THAT WERE NOT ABLE TO BE FIRED FELL OFF INTO THE CAVITY AND WAS ABLE TO BE RETRIEVED. OOZING BLEEDING OCCURRED AS A RESULT OF THE ISSUE. THEY THEN USED ANOTHER RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER DID NOT FIRE. ANOTHER STAPLER HAD TO BE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Material Twisted/Bent | A040609 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, ON THE FIRST FIRING IN BLACK THE TISSUE WAS VERY THICK. AND THE SURGEON FIRED IT AND EVERYTHING WAS FINE, AND WHEN THEY RELOADED THE STAPLER THE ANVIL WAS BENT AND WOULD NOT RE-FIT DOWN THE TROCAR. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
08-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY: A GAP WAS NOTICED BETWEEN THE RELOAD AND ANVIL JAW, THIS HAPPENED AFTER THE SECOND FIRE AND A NEW GUN WAS OPENED TO COMPLETE REST OF THE PROCEDURE. IT WAS FEMALE PATIENT AGE (B)(6), SEAMGAURD WAS USED WITH BLACK RELOADS ON BOTH SIDES OF THE RELOAD, BOUGIE SIZE 34FR. REP WAS PRESENT AT THE CASE. NO AE TO THE PATIENT. 10 MINUTES DELAY IN THE SURGERY.
10-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ECHELON ENDOPATH STAPLER | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, CLAMP LINE WITHOUT CONSISTENCY IN THE SECOND SHOT WITH GREEN LOAD.
04-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A DISTAL PANCREATECTOMY, WHILE PERFORMING A PANCREATIC RESECTION, THE FIRST FIRING FAILED. IT WAS NOTED THAT A TRACE THAT THE PUSHER MOVED WAS FOUND. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
04-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A DISTAL PANCREATECTOMY, WHILE PERFORMING A PANCREATIC RESECTION, THE FIRST FIRING FAILED. IT WAS NOTED THAT A TRACE THAT THE PUSHER MOVED WAS FOUND. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
04-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A DISTAL PANCREATECTOMY, WHILE PERFORMING A PANCREATIC RESECTION, THE FIRST FIRING FAILED. IT WAS NOTED THAT A TRACE THAT THE PUSHER MOVED WAS FOUND. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
02-10-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED DURING A LAPAROSCOPIC RADICAL NEPHRECTOMY, THE CONSULTANT UROLOGY SURGEON ATTEMPTED TO USE AN ETHICON ENDOPATH ETS-FLEX ARTICULATING ENDOSCOPIC LINEAR CUTTER TO DIVIDE THE RENAL VEIN. THE CONSULTANT UROLOGIST CALLED IN A CONSULTANT GENERAL SURGEON WHO CONFIRMED THE STAPLER WAS NOT WORKING CORRECTLY THEN PROCEEDED TO OCCLUDE VEIN USING ONE GOLD HEM-O-LOK CLIP. THE UROLOGIST CONSULTANT TOOK OVER AND PUT TWO FURTHER GOLD HEM-O-LOK CLIPS UNDER AND PROXIMAL TO THE ETS STAPLING DEVICE AND THE ETS WAS THEN REMOVED AND THE VEIN WAS SAFELY DIVIDED. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-09-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE RELEASE LEVER IS BROKEN. DEVICE COULD BE OPENED AND REMOVED FROM THE PATIENT WITHOUT FURTHER PROBLEMS. NO HARM TO THE PATIENT.
05-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY, AT THE FIRST FIRING, THE SURGEON ATTEMPTED TO RESECT THE ANAL SIDE COLON, BUT WAS UNABLE TO FIRE THE DEVICE. IT IS POSSIBLE THAT THE DEVICE WAS PREFIRED FOR UNKNOWN REASON BECAUSE THE STAPLE AT THE PROXIMAL END HAD BEEN PUSHED UP. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ECHELON FLEX 45 STANDARD | Dull, Blunt | A020101 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, THE DOCTOR USED EC45A IN SEGMENTECTOMY SURGERY, AND AFTER HE FIRED 6 TIMES, HE FELT THE KNIFE IS NOT SHARP, SO HE REPORTED THE COMPLAINT. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
05-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY, AT THE FIRST FIRING, THE SURGEON ATTEMPTED TO RESECT THE ANAL SIDE COLON, BUT WAS UNABLE TO FIRE THE DEVICE. IT IS POSSIBLE THAT THE DEVICE WAS PREFIRED FOR UNKNOWN REASON BECAUSE THE STAPLE AT THE PROXIMAL END HAD BEEN PUSHED UP. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY, AT THE FIRST FIRING, THE SURGEON ATTEMPTED TO RESECT THE ANAL SIDE COLON, BUT WAS UNABLE TO FIRE THE DEVICE. IT IS POSSIBLE THAT THE DEVICE WAS PREFIRED FOR UNKNOWN REASON BECAUSE THE STAPLE AT THE PROXIMAL END HAD BEEN PUSHED UP. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
21-01-1988 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Mechanical Jam | A0506 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE TISSUE, THE DEVICE HANDLE WAS JAMMED AND THERE WAS RESISTANCE. THEY USED A NEW STAPLER TO COMPLETE THE CASE.
21-01-1988 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE TISSUE, THE DEVICE HANDLE WAS JAMMED AND THERE WAS RESISTANCE. THEY USED A NEW STAPLER TO COMPLETE THE CASE.
21-01-1988 | Malfunction | covidien lp llc north haven | 03-01-2019 | UNKNOWN ENDO GIA SULU | Mechanical Jam | A0506 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE TISSUE, THE DEVICE HANDLE WAS JAMMED AND THERE WAS RESISTANCE. THEY USED A NEW STAPLER TO COMPLETE THE CASE.
21-01-1988 | Malfunction | covidien lp llc north haven | 03-01-2019 | UNKNOWN ENDO GIA SULU | Physical Resistance/Sticking | A0509 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE TISSUE, THE DEVICE HANDLE WAS JAMMED AND THERE WAS RESISTANCE. THEY USED A NEW STAPLER TO COMPLETE THE CASE.
26-10-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT USING THE STAPLER IN A LAP CHOLE FIRING OVER THE BILE DUCT, DOCTOR HAD FELT THE DUCT AND DIDN¿T FEEL ANY STONES. 1ST ONE - COULDN¿T CLOSE THE CLOSING TRIGGER ON FIRST, DIDN¿T FIRE, REMOVED AND GOT NEW STAPLER 2ND - TOOK SAME CARTRIDGE FROM FIRST STAPLER, CLOSED (AND WAS ABLE TO CLOSE) WENT TO FIRE AND FIRING TRIGGER BROKE, REMOVED THAT STAPLER - FIRING NOT COMPLETED OPENED NEW 3RD STAPLER AND NEW CARTRIDGE AND FIRED NO PROBLEM THERE WERE CLIPS IN THE AREA BUT DOCTOR DIDN¿T THINK SHE WAS FIRING OVER ONE. NO PATIENT CONSEQUENCE REPORTED.
26-10-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT USING THE STAPLER IN A LAP CHOLE FIRING OVER THE BILE DUCT, DOCTOR HAD FELT THE DUCT AND DIDN¿T FEEL ANY STONES. 1ST ONE - COULDN¿T CLOSE THE CLOSING TRIGGER ON FIRST, DIDN¿T FIRE, REMOVED AND GOT NEW STAPLER 2ND - TOOK SAME CARTRIDGE FROM FIRST STAPLER, CLOSED (AND WAS ABLE TO CLOSE) WENT TO FIRE AND FIRING TRIGGER BROKE, REMOVED THAT STAPLER - FIRING NOT COMPLETED OPENED NEW 3RD STAPLER AND NEW CARTRIDGE AND FIRED NO PROBLEM THERE WERE CLIPS IN THE AREA BUT DOCTOR DIDN¿T THINK SHE WAS FIRING OVER ONE. NO PATIENT CONSEQUENCE REPORTED.
26-10-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT USING THE STAPLER IN A LAP CHOLE FIRING OVER THE BILE DUCT, DOCTOR HAD FELT THE DUCT AND DIDN¿T FEEL ANY STONES. 1ST ONE - COULDN¿T CLOSE THE CLOSING TRIGGER ON FIRST, DIDN¿T FIRE, REMOVED AND GOT NEW STAPLER 2ND - TOOK SAME CARTRIDGE FROM FIRST STAPLER, CLOSED (AND WAS ABLE TO CLOSE) WENT TO FIRE AND FIRING TRIGGER BROKE, REMOVED THAT STAPLER - FIRING NOT COMPLETED OPENED NEW 3RD STAPLER AND NEW CARTRIDGE AND FIRED NO PROBLEM THERE WERE CLIPS IN THE AREA BUT DOCTOR DIDN¿T THINK SHE WAS FIRING OVER ONE. NO PATIENT CONSEQUENCE REPORTED.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN EXCISION OF CHEST WALL LESION AND THORACOSCOPY PROCEDURE, AT THE FIRST FIRING, WHEN THE DIAPHRAGM WAS BEING RESECTED, THE HANDLE HAD STRONG RESISTANCE AND COULD NOT BE SQUEEZED. A NEW PURPLE CARTRIDGE WAS USED AND THE DEVICE WAS ABLE TO FIRE WITHOUT PROBLEMS. NO INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, THE STAPLER BLOCKED AFTER INTRODUCING THE TISSUE BETWEEN THE JAWS. AFTER FEW ATTEMPTS THE STAPLER UNJAMMED, NOTING THAT SOME STAPLES ESCAPED WITHOUT CONSEQUENCES. THE ISSUE WAS SOLVED USING ANOTHER STAPLER.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN EXCISION OF CHEST WALL LESION AND THORACOSCOPY PROCEDURE, AT THE FIRST FIRING, WHEN THE DIAPHRAGM WAS BEING RESECTED, THE HANDLE HAD STRONG RESISTANCE AND COULD NOT BE SQUEEZED. A NEW PURPLE CARTRIDGE WAS USED AND THE DEVICE WAS ABLE TO FIRE WITHOUT PROBLEMS. NO INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, WHILE THE DEVICE WAS BEING APPLIED TO THE TISSUE, THE SURGEON COULD NOT FIRE COMPLETELY AND A NOISE WAS HEARD. IT WAS NOTED THAT SEVERAL STAPLES DID NOT FORM COMPLETELY AND THAT THE STAPLE LINE WAS INCOMPLETELY DISTALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, WHILE THE DEVICE WAS BEING APPLIED TO THE TISSUE, THE SURGEON COULD NOT FIRE COMPLETELY AND A NOISE WAS HEARD. IT WAS NOTED THAT SEVERAL STAPLES DID NOT FORM COMPLETELY AND THAT THE STAPLE LINE WAS INCOMPLETELY DISTALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, WHILE THE DEVICE WAS BEING APPLIED TO THE TISSUE, THE SURGEON COULD NOT FIRE COMPLETELY AND A NOISE WAS HEARD. IT WAS NOTED THAT SEVERAL STAPLES DID NOT FORM COMPLETELY AND THAT THE STAPLE LINE WAS INCOMPLETELY DISTALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, WHILE THE DEVICE WAS BEING APPLIED TO THE TISSUE, THE SURGEON COULD NOT FIRE COMPLETELY AND A NOISE WAS HEARD. IT WAS NOTED THAT SEVERAL STAPLES DID NOT FORM COMPLETELY AND THAT THE STAPLE LINE WAS INCOMPLETELY DISTALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, WHILE THE DEVICE WAS BEING APPLIED TO THE TISSUE, THE SURGEON COULD NOT FIRE COMPLETELY AND A NOISE WAS HEARD. IT WAS NOTED THAT SEVERAL STAPLES DID NOT FORM COMPLETELY AND THAT THE STAPLE LINE WAS INCOMPLETELY DISTALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, WHILE THE DEVICE WAS BEING APPLIED TO THE TISSUE, THE SURGEON COULD NOT FIRE COMPLETELY AND A NOISE WAS HEARD. IT WAS NOTED THAT SEVERAL STAPLES DID NOT FORM COMPLETELY AND THAT THE STAPLE LINE WAS INCOMPLETELY DISTALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-11-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC ANTERIOR RESECTION, DURING THE DIVISION OF LARGE BOWEL, THE RELOAD WAS PLACED ONTO THE TISSUE AND WOULD NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-11-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC ANTERIOR RESECTION, DURING THE DIVISION OF LARGE BOWEL, THE RELOAD WAS PLACED ONTO THE TISSUE AND WOULD NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY WHEN THE DEVICE WAS SET, THEY ATTEMPTED TO STAPLE THE VESSEL, THE HANDLE WAS STIFF AND COULDN'T FIRE. THE OTHER DOCTOR SET THE DEVICE TO VESSEL AGAIN AND FIRED, IT WORKED WITH NO PROBLEM AND THE STAPLES WERE FORMED AS USUAL. HOWEVER, IN CASE OF THE EVENT OCCURRED AGAIN, AFTER THE SECOND FIRING THEY REPLACED WITH A NEW HANDLE. THE NEW HANDLE FIRED WITH NO PROBLEM. THERE WAS NO PATIENT HARM.
16-11-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | CONTOUR CURVED CUTTER STAPLER | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN COLECTOMY THE SURGEON TRIED TO FIRE THE DEVICE AND IT WOULD NOT CLOSE PROPERLY ON TISSUE. THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED.
07-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING ON BLOOD VESSELS DURING AN ENDOSCOPIC LOBECTOMY WITH THORACIC SURGERY, THEY COULD FIRE HALF OF THE SHOTS.
27-11-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE JAWS DID NOT OPEN AFTER THE FIRING ON THE COLON. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-01-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | PROXIMATE* LINEAR CUTTER | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, NO KNIFE WORKING DUE TO HARD AND THICK TISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | ENDO GIA | Break | A0401 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHEN ON LAPAROSCOPIC RESECTION, THE NAILS WERE BROKEN WHEN THE LUNG WAS BEING DISCONNECTED, AND COULD NOT BE CLOSED NORMALLY AFTER BEING REPLACED WITH THE 60-STRAIGHT NAIL. THE THIRD STAPLES WERE POORLY FORMED, HAD AN INCOMPLETE STAPLE LINE DISTALLY AND THEY SUTURED AGAIN TO FIX THE STAPLE LINE. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHEN ON LAPAROSCOPIC RESECTION, THE NAILS WERE BROKEN WHEN THE LUNG WAS BEING DISCONNECTED, AND COULD NOT BE CLOSED NORMALLY AFTER BEING REPLACED WITH THE 60-STRAIGHT NAIL. THE THIRD STAPLES WERE POORLY FORMED, HAD AN INCOMPLETE STAPLE LINE DISTALLY AND THEY SUTURED AGAIN TO FIX THE STAPLE LINE. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | ENDO GIA | Break | A0401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHEN ON LAPAROSCOPIC RESECTION, THE NAILS WERE BROKEN WHEN THE LUNG WAS BEING DISCONNECTED, AND COULD NOT BE CLOSED NORMALLY AFTER BEING REPLACED WITH THE 60-STRAIGHT NAIL. THE THIRD STAPLES WERE POORLY FORMED, HAD AN INCOMPLETE STAPLE LINE DISTALLY AND THEY SUTURED AGAIN TO FIX THE STAPLE LINE. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHEN ON LAPAROSCOPIC RESECTION, THE NAILS WERE BROKEN WHEN THE LUNG WAS BEING DISCONNECTED, AND COULD NOT BE CLOSED NORMALLY AFTER BEING REPLACED WITH THE 60-STRAIGHT NAIL. THE THIRD STAPLES WERE POORLY FORMED, HAD AN INCOMPLETE STAPLE LINE DISTALLY AND THEY SUTURED AGAIN TO FIX THE STAPLE LINE. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHEN ON LAPAROSCOPIC RESECTION, THE NAILS WERE BROKEN WHEN THE LUNG WAS BEING DISCONNECTED, AND COULD NOT BE CLOSED NORMALLY AFTER BEING REPLACED WITH THE 60-STRAIGHT NAIL. THE THIRD STAPLES WERE POORLY FORMED, HAD AN INCOMPLETE STAPLE LINE DISTALLY AND THEY SUTURED AGAIN TO FIX THE STAPLE LINE. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Injury | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHEN ON LAPAROSCOPIC RESECTION, THE NAILS WERE BROKEN WHEN THE LUNG WAS BEING DISCONNECTED, AND COULD NOT BE CLOSED NORMALLY AFTER BEING REPLACED WITH THE 60-STRAIGHT NAIL. THE THIRD STAPLES WERE POORLY FORMED, HAD AN INCOMPLETE STAPLE LINE DISTALLY AND THEY SUTURED AGAIN TO FIX THE STAPLE LINE. THEY REPLACED WITH NEW PRODUCTS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | GST BLUE RELOAD, 60MM, 6 ROW | Component Missing | A020602 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IT WAS NOTICED UPON OPENING THE GST60B LOAD THAT THE METAL PAN (TRAY) WAS DISLODGED OR OUT OF PLACE. DEVICE WAS REMOVED FROM STERILE FIELD AND REPLACED WITH NEW CARTRIDGE. UPON FURTHER INSPECTION, IT APPEARS THAT MULTIPLE DRIVERS ARE MISSING FROM THE RELOAD. DRIVERS WERE NOT FOUND LOSE IN PACKING OR IN THE STERILE FIELD. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | GST BLUE RELOAD, 60MM, 6 ROW | Device Dislodged or Dislocated | A051201 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IT WAS NOTICED UPON OPENING THE GST60B LOAD THAT THE METAL PAN (TRAY) WAS DISLODGED OR OUT OF PLACE. DEVICE WAS REMOVED FROM STERILE FIELD AND REPLACED WITH NEW CARTRIDGE. UPON FURTHER INSPECTION, IT APPEARS THAT MULTIPLE DRIVERS ARE MISSING FROM THE RELOAD. DRIVERS WERE NOT FOUND LOSE IN PACKING OR IN THE STERILE FIELD. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
14-08-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT HEMICOLECTOMY PROCEDURE, THEY WERE UNABLE TO GET THE RELOAD OUT OF THE DEVICE. IT WAS STUCK IN THE STAPLER. THERE WAS NO PATIENT INJURY.
01-11-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION PROCEDURE, WHILE STAPLING THE SIGMOID COLON, TWO HANDLES POPPED DURING THE FIRST TWO FIRINGS. THE SURGEON STOPPED FIRING WHEN THE HANDLES POPPED AND OPENED ANOTHER HANDLE AND RELOAD TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP SIGMOIDECTOMY, ¿STAPLES WENT OUT BUT NEVER SEALED. RECTUM WAS PARTIALLY CUT.¿ A SECOND ECS29A WAS USED TO FINISH THE SURGERY AND SIGNIFICANT OR TIME WAS ADDED. IT IS UNKNOWN IF THERE WERE ANY ADVERSE CONSEQUENCES TO THE PATIENT.
11-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | CONTOUR CURVED CUTTER STAPLER | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PATIENT WAS BOOKED FOR LOW ANTERIOR RESECTION BY THE SURGEON. ON FIRING CONTOUR STAPLER IN PELVIS THE SURGEON STATED THAT THE STAPLER HAD MISFIRED. THE SURGEON USED ANOTHER STAPLER FOR A LOWER BOWEL ANASTOMOSIS THAN WAS PLANNED. THE SURGEON DID DISCLOSE TO PATIENT AND FAMILY AND HAS ASKED FOR STAPLER TO BE RETURNED TO ETHICON FOR INVESTIGATION. PATIENT TRANSFERRED TO PACU IN STABLE CONDITION. STAPLER ISOLATED, AEMS COMPLETED AND STAPLER LOCKED IN QUARANTINE CAGE IN MDR. THE SURGEON CLARIFIED THAT THE STAPLER CUT BUT DID NOT STAPLE RESULTING IN 2 OPEN SEGMENTS ON THE SPECIMEN AND THE PATIENT SIDE. ACTION TAKEN FOR PROCEDURE: SUTURE CLOSURE OF BOWEL SEGMENTS, LOOP THROUGH BOWEL, FINISHED ANASTOMOSIS TRANS-ANAL WITH COMPETITIVE HEMORRHOID STAPLER. THERE WERE NO PATIENT CONSEQUENCES REPORTED. PATIENT RELEASED DOING OK.
12-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: EMERGENCY HIGH ANTERIOR RESECTION (B)(6) 2018. ISSUE: DURING THE USE OF THE ECS29A, THE ANVIL OF THE DEVICE WAS SECURED TO THE PROXIMALLY TRANSECTED TRANSVERSE COLON VIA A HAND SEWN PURSE STRING SUTURE, THE TRANS-ANAL PORTION OF THE STAPLER WAS APPROPRIATELY POSITIONED AND OPENED AS PER THE J&J OPTIMAL DEVICE PERFORMANCE INSTRUCTIONS, WITH THE TROCAR OF THE STAPLER PIERCING THROUGH THE STAPLED SIGMOID STUMP - ANTERIOR TO THE STAPLE LINE. THE SUTURE WAS NOT INTERFERING WITH THE DOCKING SPRING AND IT APPEARED TO DOCK WITH THE TROCAR OF THE STAPLER AS NORMAL, THE STAPLER WAS THEN CLOSED ONTO THE SELECTED TISSUE AND CLOSED UNTIL AN APPROPRIATE TISSUE RESISTANCE WAS ACHIEVED WITHIN THE GREEN RANGE OF THE TISSUE COMPRESSION SCALE, WAITED 15 SECONDS, RE-TENSIONED AND FIRED. I DID NOT HEAR THE AUDIBLE FEEDBACK ¿CRUNCH¿ ALTHOUGH IT WAS REPORTED BY DR THAT THE TACTILE AND AUDIBLE FEEDBACK OCCURRED. DR OPENED THE DEVICE A HALF TO ¾ TURNS, BUT WAS HAVING TROUBLE TRYING TO REMOVE THE STAPLER FROM THE ANASTOMOSIS, IT WAS THEN REALIZED THAT THE ANTERIOR PORTION OF THE STAPLE LINE HAD NOT FORMED AND THAT THE BOWEL WAS LEFT OPEN ON THE ANTERIOR SIDE OF THE ANASTOMOSIS. ONCE THE DEVICE WAS SUCCESSFULLY REMOVED FROM THE PARTIALLY FORMED ANASTOMOSIS, IT WAS DECIDED THAT THE BEST FORM OF ACTION WAS TO TRANSECT A FURTHER 3CM FROM THE DISTAL STUMP AND 1CM FROM THE PROXIMAL STUMP. THE SAME TECHNIQUES WERE USED FOR THE 2ND ATTEMPT AT AN ANASTOMOSIS AND THE 2ND FIRING OF A NEW ECS29A WAS SUCCESSFUL, WITH THE BALANCE OF THE CASE CONTINUING AS NORMAL AND DUE TO BEING A HIGH ANTERIOR RESECTION, THE PATIENT AVOIDED A TEMPORARY ILEOSTOMY OR COLOSTOMY. DELAY TO THE PROCEDURE: 1 HOUR, WITH ANOTHER ECS29A BEING USED TO COMPLETE THE ANASTOMOSIS. PATIENT OUTCOME: THE CASE COMPLETED AS EXPECTED AND WAS SUCCESSFULLY LEAK TESTED. A CLEAR PATIENT OUTCOME WILL ONLY BE AVAILABLE OVER THE NEXT FEW DAYS. POST PROCEDURE: DR AND JNJ REPRESENTATIVE INSPECTED THE STAPLER AND THE RESECTED, FAILED ANASTOMOSIS. THERE WAS ONLY 1 APPARENTLY COMPLETE DONUT CONTAINED WITHIN THE STAPLER, ALL STAPLE DRIVERS APPEARED TO DEPLOY AS EXPECTED AND THE TACTILE FEEDBACK DISC WAS CUT THROUGH AS PER A NORMAL & SUCCESSFUL FIRING. IT WAS NOTED THAT STAPLES APPEARED TO BE MISSING ALTOGETHER FROM THE ANTERIOR SIDE OF THE FAILED ANASTOMOSIS, THERE WERE NO MALFORMED STAPLES PRESENT.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY ON THE EXCISION OF LOBES, THE STAPLER WAS NOT ABLE TO FIRE AND SQUEEZE THE HANDLE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY ON THE EXCISION OF LOBES, THE STAPLER WAS NOT ABLE TO FIRE AND SQUEEZE THE HANDLE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY ON THE EXCISION OF LOBES, THE STAPLER WAS NOT ABLE TO FIRE AND SQUEEZE THE HANDLE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY ON THE EXCISION OF LOBES, THE STAPLER WAS NOT ABLE TO FIRE AND SQUEEZE THE HANDLE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY ON THE EXCISION OF LOBES, THE STAPLER WAS NOT ABLE TO FIRE AND SQUEEZE THE HANDLE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Physical Resistance/Sticking | A0509 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY ON THE EXCISION OF LOBES, THE STAPLER WAS NOT ABLE TO FIRE AND SQUEEZE THE HANDLE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY ON THE EXCISION OF LOBES, THE STAPLER WAS NOT ABLE TO FIRE AND SQUEEZE THE HANDLE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY ON THE EXCISION OF LOBES, THE STAPLER WAS NOT ABLE TO FIRE AND SQUEEZE THE HANDLE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, THE RELOAD WAS ATTACHED TO THE FOURTH FIRING. WHEN THE CHECK OF OPENING AND CLOSING WAS PERFORMED, THE JAWS WERE NOT CLOSED. THEY WERE IMPOSSIBLE TO BE INSERTED THROUGH THE TROCAR UNLESS THE JAWS WERE CLOSED BY HAND, THEREFORE, THE OPERATION WAS PERFORMED BY TAKING OUT A NEW PRODUCT. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, THE RELOAD WAS ATTACHED TO THE FOURTH FIRING. WHEN THE C HECK OF OPENING AND CLOSING WAS PERFORMED, THE JAWS WERE NOT CLOSED. THEY WERE IMPOSSIBLE TO BE INSERTED THROUGH THE TROCAR UNLESS THE JAWS WERE CLOSED BY HAND, THEREFORE, THE OPERATION WAS PERFORMED BY TAKING OUT A NEW PRODUCT. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN COLECTOMY, THE FIRING KNOB WAS BROKEN OFF AT THE 2ND FIRING ON THE COLON. THE DEVICE WAS USED TO CLOSE THE INSERTION SITE FOR FUNCTIONAL END TO END ANASTOMOSIS (FEEA). NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
05-12-2018 | Malfunction | ethicon endo-surgery | 03-01-2019 | 75MM SELECTABLE NEW TLC | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN COLECTOMY, THE FIRING KNOB WAS BROKEN OFF AT THE 2ND FIRING ON THE COLON. THE DEVICE WAS USED TO CLOSE THE INSERTION SITE FOR FUNCTIONAL END TO END ANASTOMOSIS (FEEA). NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED ON A DISCONNECTED BRONCHUS DURING AN ENDOSCOPIC RIGHT MIDDLE LOBE RESECTION, THE DEVICE WAS FOUND THAT IT COULD NOT BE COMPLETELY CLOSED. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED ON A DISCONNECTED BRONCHUS DURING AN ENDOSCOPIC RIGHT MIDDLE LOBE RESECTION, THE DEVICE WAS FOUND THAT IT COULD NOT BE COMPLETELY CLOSED. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-05-2016 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE AND THE RELOAD LOCKED DURING THE STAPLING PROCESS. IT WAS STATED THAT IT WAS NOT POSSIBLE TO PROGRESS THROUGH THE PROCEDURE AND IT WAS NECESSARY TO OPEN ANOTHER STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-05-2016 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE AND THE RELOAD LOCKED DURING THE STAPLING PROCESS. IT WAS STATED THAT IT WAS NOT POSSIBLE TO PROGRESS THROUGH THE PROCEDURE AND IT WAS NECESSARY TO OPEN ANOTHER STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-11-2018 | Malfunction | ethicon | 03-01-2019 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN INGUINAL HERNIA REPAIR ON (B)(6) 2018 AND THE ABSORBABLE STRAPS WERE USED. DURING THE PROCEDURE, THE DEVICE DID NOT WORK CORRECTLY. IT WAS REPORTED THAT THE ANCHORS WERE NOT FIXED WELL. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
27-11-2018 | Injury | us surgical puerto rico | 03-01-2019 | EEA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, 8 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC CARCINOMA RESECTION PROCEDURE, THE ANASTOMOSIS WAS OK DURING THE PROCEDURE, HOWEVER, THEY FOUND OUT THAT THERE WAS A HOLE IN THE ANASTOMOSIS WHEN THE PATIENT BACK FOR CHECK UP. TO RESOLVE THE ISSUE THEY OVERSAW THE STAPLE LINE.
27-11-2018 | Injury | us surgical puerto rico | 03-01-2019 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 8 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC CARCINOMA RESECTION PROCEDURE, THE ANASTOMOSIS WAS OK DURING THE PROCEDURE, HOWEVER, THEY FOUND OUT THAT THERE WAS A HOLE IN THE ANASTOMOSIS WHEN THE PATIENT BACK FOR CHECK UP. TO RESOLVE THE ISSUE THEY OVERSAW THE STAPLE LINE.
22-11-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DISTAL PANCREATECTOMY, WHEN FIRING OVER THE PANCREAS, THERE WAS POOR STAPLE FORMATION AND THE CENTRAL PART OF THE STAPLE LINE IS INCOMPLETE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
22-11-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DISTAL PANCREATECTOMY, WHEN FIRING OVER THE PANCREAS, THERE WAS POOR STAPLE FORMATION AND THE CENTRAL PART OF THE STAPLE LINE IS INCOMPLETE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC WEDGE RESECTION, WHEN THE OPERATOR FIRED THE SECOND RELOAD, A SQUEAKING SOUND WAS HEARD. AFTER THE FIRING, THE ANVIL WAS BENT, AND THE NAIL WAS BADLY FORMED. THE FIRST TWO CARTRIDGES WERE POORLY FORMED. WHEN THE THIRD NAIL WAS FIRED, AFTER THE SAFETY WAS ACTIVATED , BUT IT COULDN'T BE FIRED, AND THEN REPLACED WITH ANOTHER ONE, THE SAME PROBLEM OCCURRED; FINALLY, A NEW STAPLER WAS REPLACED AND REUSED THE RELOAD, THE DEVICE WAS USED NORMALLY AND THE OPERATION WAS SUCCESSFULLY COMPLETED.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC WEDGE RESECTION, WHEN THE OPERATOR FIRED THE SECOND RELOAD, A SQUEAKING SOUND WAS HEARD. AFTER THE FIRING, THE ANVIL WAS BENT, AND THE NAIL WAS BADLY FORMED. THE FIRST TWO CARTRIDGES WERE POORLY FORMED. WHEN THE THIRD NAIL WAS FIRED, AFTER THE SAFETY WAS ACTIVATED , BUT IT COULDN'T BE FIRED, AND THEN REPLACED WITH ANOTHER ONE, THE SAME PROBLEM OCCURRED; FINALLY, A NEW STAPLER WAS REPLACED AND REUSED THE RELOAD, THE DEVICE WAS USED NORMALLY AND THE OPERATION WAS SUCCESSFULLY COMPLETED.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC WEDGE RESECTION, WHEN THE OPERATOR FIRED THE SECOND RELOAD, A SQUEAKING SOUND WAS HEARD. AFTER THE FIRING, THE ANVIL WAS BENT, AND THE NAIL WAS BADLY FORMED. THE FIRST TWO CARTRIDGES WERE POORLY FORMED. WHEN THE THIRD NAIL WAS FIRED, AFTER THE SAFETY WAS ACTIVATED , BUT IT COULDN'T BE FIRED, AND THEN REPLACED WITH ANOTHER ONE, THE SAME PROBLEM OCCURRED; FINALLY, A NEW STAPLER WAS REPLACED AND REUSED THE RELOAD, THE DEVICE WAS USED NORMALLY AND THE OPERATION WAS SUCCESSFULLY COMPLETED.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC WEDGE RESECTION, WHEN THE OPERATOR FIRED THE SECOND RELOAD, A SQUEAKING SOUND WAS HEARD. AFTER THE FIRING, THE ANVIL WAS BENT, AND THE NAIL WAS BADLY FORMED. THE FIRST TWO CARTRIDGES WERE POORLY FORMED. WHEN THE THIRD NAIL WAS FIRED, AFTER THE SAFETY WAS ACTIVATED , BUT IT COULDN'T BE FIRED, AND THEN REPLACED WITH ANOTHER ONE, THE SAME PROBLEM OCCURRED; FINALLY, A NEW STAPLER WAS REPLACED AND REUSED THE RELOAD, THE DEVICE WAS USED NORMALLY AND THE OPERATION WAS SUCCESSFULLY COMPLETED.
06-12-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC WEDGE RESECTION, WHEN THE OPERATOR FIRED THE SECOND RELOAD, A SQUEAKING SOUND WAS HEARD. AFTER THE FIRING, THE ANVIL WAS BENT, AND THE NAIL WAS BADLY FORMED. THE FIRST TWO CARTRIDGES WERE POORLY FORMED. WHEN THE THIRD NAIL WAS FIRED, AFTER THE SAFETY WAS ACTIVATED , BUT IT COULDN'T BE FIRED, AND THEN REPLACED WITH ANOTHER ONE, THE SAME PROBLEM OCCURRED; FINALLY, A NEW STAPLER WAS REPLACED AND REUSED THE RELOAD, THE DEVICE WAS USED NORMALLY AND THE OPERATION WAS SUCCESSFULLY COMPLETED.
19-11-2018 | Malfunction | us surgical puerto rico | 03-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, A FIRST RELOAD WAS USED WITH THE HANDLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT IT DID NOT FIRE, ANOTHER RELOAD WAS USED AND IT STILL WASN'T ABLE TO FIRE. A NEW HANDLE WAS USED AND IT WORKED FINE. THERE WAS NO PATIENT INJURY.
19-11-2018 | Malfunction | covidien lp llc north haven | 03-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, A FIRST RELOAD WAS USED WITH THE HANDLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT IT DID NOT FIRE, ANOTHER RELOAD WAS USED AND IT STILL WASN'T ABLE TO FIRE. A NEW HANDLE WAS USED AND IT WORKED FINE. THERE WAS NO PATIENT INJURY.
29-11-2018 | Malfunction | ethicon | 03-01-2019 | ETHICON SECURESTRAP | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC HERNIA REPAIR PROCEDURE ON (B)(6) 2018 AND THE ABSORBABLE STRAPS WERE USED. IT WAS ALSO REPORTED THAT DURING THE PROCEDURE, THE HOOKS DO NOT STAPLE, THEY ARE RELEASED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
17-10-2018 | Malfunction | ethicon | 03-01-2019 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A UMBILICAL HERNIA REPAIR PROCEDURE ON (B)(6) 2018 AND THE ABSORBABLE STRAPS WERE USED. DURING THE USE, THE DEVICE DID NOT WORK AS INTENDED, THE CLIPS/STRAPS DID NOT ADHERE TO THE ABDOMINAL WALL. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
17-10-2018 | Malfunction | ethicon | 03-01-2019 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A UMBILICAL HERNIA REPAIR PROCEDURE ON (B)(6) 2018 AND THE ABSORBABLE STRAPS WERE USED. DURING THE USE, THE DEVICE DID NOT WORK AS INTENDED, THE CLIPS/STRAPS DID NOT ADHERE TO THE ABDOMINAL WALL. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
30-11-2018 | Injury | intuitive surgical | 04-01-2019 | ENDOWRIST | Misfire | A050502 | Fistula | Event Description: IT WAS INITIALLY REPORTED THAT DURING AN UNSPECIFIED DA VINCI-ASSISTED THORACIC PROCEDURE, THE SURGEON ATTEMPTED UNSUCCESSFULLY TO FIRE TWO GREEN STAPLER 30 RELOADS ON BRONCHUS TISSUE USING A CURVED-TIP STAPLER 30 INSTRUMENT. AS A RESULT OF THE ALLEGED ISSUE, THE CASE WAS CONVERTED TO OPEN SURGERY. ON (B)(6) 2018, INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: WHEN THE SURGEON FIRED THE TWO GREEN STAPLER 30 RELOADS, MALFORMED STAPLES WERE IDENTIFIED AND THE STAPLE LINE/CUT WAS NOTED TO BE ¿POOR.¿ AFTER FIRING THE TWO GREEN STAPLER 30 RELOADS, THE SURGEON ELECTED TO PROCEED WITH THE SURGICAL TASK OF TRANSECTING THE BRONCHUS TISSUE BY CONVERTING THE PROCEDURE TO OPEN SURGERY. THE SURGEON DID NOT ATTEMPT TO USE A LAPAROSCOPIC STAPLER INSTRUMENT. THE CSR INDICATED THAT THE REPORTED STAPLING ISSUE MAY HAVE OCCURRED DUE TO THE PATIENT¿S ANATOMY SINCE THE SURGEON INDICATED THAT THE STAPLE LINE MIGHT HAVE FAILED WITH ANOTHER TYPE OF STAPLER INSTRUMENT. THE SURGEON MADE THE DECISION TO MANUALLY TRANSECT AND SEW THE TARGETED BRONCHUS TISSUE IN ORDER TO STAY WITHIN THE PLANNED MARGINS. THE SURGEON SUCCESSFULLY COMPLETED THE TASK AT HAND AND THE PROCEDURE AS A WHOLE VIA OPEN SURGERY. THE CSR CONFIRMED THAT THE SURGEON REPORTED NO INJURY OR HARM TO THE PATIENT. THE PATIENT WAS REPORTEDLY STABLE AND IN NORMAL CONDITION. THE CUSTOMER CONFIRMED THAT THERE WAS NO ABERRANT TISSUE, CARTILAGE, OR CALCIFICATION NEAR THE TARGET TISSUE. THE CSR CONFIRMED THAT THE CURVED-TIP STAPLER 30 INSTRUMENT WORKED AS EXPECTED FOR VASCULAR TISSUE DURING THE CASE. THE CSR WAS NOT CERTAIN IF THE CUSTOMER PLANNED ON RETURNING THE TWO GREEN STAPLER 30 RELOADS INVOLVED WITH THE REPORTED EVENT. ON (B)(6) 2018, ISI ALSO CONTACTED THE SITE¿S ROBOTICS COORDINATOR AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: DUE TO THE ALLEGED PARTIAL STAPLING FIRES AND LIMITED MARGINS, THE SURGEON REMOVED ADDITIONAL TISSUE THAT WAS PLANNED AS PART OF THE ROBOTIC SURGICAL PROCEDURE. THE ROBOTICS COORDINATOR EXPLAINED THAT THE SURGEON ¿CUT THE TISSUE THAT WAS TORN¿ AS A RESULT OF THE ALLEGED PARTIAL FIRES. THE SURGEON PROCEEDED WITH A BRONCHIAL SLEEVE VIA OPEN SURGERY. ON (B)(6) 2018, THE ROBOTICS COORDINATOR CONTACTED ISI AND INDICATED THAT THE PATIENT HAD EXPERIENCED A BRONCHIAL FISTULA.
30-11-2018 | Injury | intuitive surgical | 04-01-2019 | ENDOWRIST | Misfire | A050502 | Tissue Damage | Event Description: IT WAS INITIALLY REPORTED THAT DURING AN UNSPECIFIED DA VINCI-ASSISTED THORACIC PROCEDURE, THE SURGEON ATTEMPTED UNSUCCESSFULLY TO FIRE TWO GREEN STAPLER 30 RELOADS ON BRONCHUS TISSUE USING A CURVED-TIP STAPLER 30 INSTRUMENT. AS A RESULT OF THE ALLEGED ISSUE, THE CASE WAS CONVERTED TO OPEN SURGERY. ON (B)(6) 2018, INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: WHEN THE SURGEON FIRED THE TWO GREEN STAPLER 30 RELOADS, MALFORMED STAPLES WERE IDENTIFIED AND THE STAPLE LINE/CUT WAS NOTED TO BE ¿POOR.¿ AFTER FIRING THE TWO GREEN STAPLER 30 RELOADS, THE SURGEON ELECTED TO PROCEED WITH THE SURGICAL TASK OF TRANSECTING THE BRONCHUS TISSUE BY CONVERTING THE PROCEDURE TO OPEN SURGERY. THE SURGEON DID NOT ATTEMPT TO USE A LAPAROSCOPIC STAPLER INSTRUMENT. THE CSR INDICATED THAT THE REPORTED STAPLING ISSUE MAY HAVE OCCURRED DUE TO THE PATIENT¿S ANATOMY SINCE THE SURGEON INDICATED THAT THE STAPLE LINE MIGHT HAVE FAILED WITH ANOTHER TYPE OF STAPLER INSTRUMENT. THE SURGEON MADE THE DECISION TO MANUALLY TRANSECT AND SEW THE TARGETED BRONCHUS TISSUE IN ORDER TO STAY WITHIN THE PLANNED MARGINS. THE SURGEON SUCCESSFULLY COMPLETED THE TASK AT HAND AND THE PROCEDURE AS A WHOLE VIA OPEN SURGERY. THE CSR CONFIRMED THAT THE SURGEON REPORTED NO INJURY OR HARM TO THE PATIENT. THE PATIENT WAS REPORTEDLY STABLE AND IN NORMAL CONDITION. THE CUSTOMER CONFIRMED THAT THERE WAS NO ABERRANT TISSUE, CARTILAGE, OR CALCIFICATION NEAR THE TARGET TISSUE. THE CSR CONFIRMED THAT THE CURVED-TIP STAPLER 30 INSTRUMENT WORKED AS EXPECTED FOR VASCULAR TISSUE DURING THE CASE. THE CSR WAS NOT CERTAIN IF THE CUSTOMER PLANNED ON RETURNING THE TWO GREEN STAPLER 30 RELOADS INVOLVED WITH THE REPORTED EVENT. ON (B)(6) 2018, ISI ALSO CONTACTED THE SITE¿S ROBOTICS COORDINATOR AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: DUE TO THE ALLEGED PARTIAL STAPLING FIRES AND LIMITED MARGINS, THE SURGEON REMOVED ADDITIONAL TISSUE THAT WAS PLANNED AS PART OF THE ROBOTIC SURGICAL PROCEDURE. THE ROBOTICS COORDINATOR EXPLAINED THAT THE SURGEON ¿CUT THE TISSUE THAT WAS TORN¿ AS A RESULT OF THE ALLEGED PARTIAL FIRES. THE SURGEON PROCEEDED WITH A BRONCHIAL SLEEVE VIA OPEN SURGERY. ON (B)(6) 2018, THE ROBOTICS COORDINATOR CONTACTED ISI AND INDICATED THAT THE PATIENT HAD EXPERIENCED A BRONCHIAL FISTULA.
17-12-2018 | Injury | intuitive surgical | 04-01-2019 | SUREFORM | Material Fragmentation | A040103 | Injury | Event Description: IT WAS INITIALLY REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, A FRAGMENT FROM A SUREFORM 60 RELOAD WAS IDENTIFIED ON A STAPLE LINE. THE SURGEON REPORTEDLY USED ANOTHER RELOAD TO CREATE A DIFFERENT STAPLE LINE AND ELIMINATE THE STAPLE LINE WITH THE FRAGMENT. THE RELOAD FRAGMENT WAS REMOVED DURING THE SAME SURGICAL PROCEDURE. ON (B)(4) 2018, INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: THE CSR WAS NOT PRESENT DURING THE SURGICAL PROCEDURE. AFTER THE SURGEON HAD TRANSECTED STOMACH TISSUE USING A SUREFORM STAPLER 60 INSTRUMENT, THE SURGICAL STAFF NOTICED GREEN MATERIAL ¿ENCAPSULATED¿ IN THE PATIENT¿S TISSUE. AT THE TIME THE TRANSECTION WAS PERFORMED, A GREEN SUREFORM 60 RELOAD WAS INSTALLED. THE SURGEON CREATED ANOTHER STAPLE LINE THAT WAS CLOSER TO THE PATIENT SIDE IN ORDER TO REMOVE THE GREEN FRAGMENT FORM THE PATIENT. IN DOING SO, THE SURGEON REMOVED ADDITIONAL TISSUE THAT WAS NOT INTENDED TO BE REMOVED AS PART OF THE PLANNED SURGICAL PROCEDURE. ACCORDING TO THE CSR, THE SURGEON REMOVED A LITTLE MORE THAN A QUARTER INCH OF TISSUE TO REMOVE THE FRAGMENT. THE CSR INSPECTED THE GREEN SUREFORM 60 RELOAD INVOLVED WITH THE REPORTED EVENT AND NOTICED A TINY PIECE MISSING. THE CSR SPOKE TO THE SURGEON REGARDING THE REPORTED EVENT. ACCORDING TO THE SURGEON, HE HAD NOT OBSERVED ANYTHING ABNORMAL PRIOR TO AND AFTER TRANSECTION OF THE PATIENT TISSUE. THE TISSUE WAS NOT TOO THICK AND THERE WAS NO TISSUE BUNCHING NOTED BY THE SURGICAL STAFF. THERE WERE POST-OPERATIVE COMPLICATIONS REPORTED AND THE SURGEON INDICATED THAT THE PATIENT WAS RECOVERING WELL.
17-12-2018 | Injury | intuitive surgical | 04-01-2019 | SUREFORM | Material Fragmentation | A040103 | Device Embedded In Tissue or Plaque | Event Description: IT WAS INITIALLY REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, A FRAGMENT FROM A SUREFORM 60 RELOAD WAS IDENTIFIED ON A STAPLE LINE. THE SURGEON REPORTEDLY USED ANOTHER RELOAD TO CREATE A DIFFERENT STAPLE LINE AND ELIMINATE THE STAPLE LINE WITH THE FRAGMENT. THE RELOAD FRAGMENT WAS REMOVED DURING THE SAME SURGICAL PROCEDURE. ON (B)(4) 2018, INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: THE CSR WAS NOT PRESENT DURING THE SURGICAL PROCEDURE. AFTER THE SURGEON HAD TRANSECTED STOMACH TISSUE USING A SUREFORM STAPLER 60 INSTRUMENT, THE SURGICAL STAFF NOTICED GREEN MATERIAL ¿ENCAPSULATED¿ IN THE PATIENT¿S TISSUE. AT THE TIME THE TRANSECTION WAS PERFORMED, A GREEN SUREFORM 60 RELOAD WAS INSTALLED. THE SURGEON CREATED ANOTHER STAPLE LINE THAT WAS CLOSER TO THE PATIENT SIDE IN ORDER TO REMOVE THE GREEN FRAGMENT FORM THE PATIENT. IN DOING SO, THE SURGEON REMOVED ADDITIONAL TISSUE THAT WAS NOT INTENDED TO BE REMOVED AS PART OF THE PLANNED SURGICAL PROCEDURE. ACCORDING TO THE CSR, THE SURGEON REMOVED A LITTLE MORE THAN A QUARTER INCH OF TISSUE TO REMOVE THE FRAGMENT. THE CSR INSPECTED THE GREEN SUREFORM 60 RELOAD INVOLVED WITH THE REPORTED EVENT AND NOTICED A TINY PIECE MISSING. THE CSR SPOKE TO THE SURGEON REGARDING THE REPORTED EVENT. ACCORDING TO THE SURGEON, HE HAD NOT OBSERVED ANYTHING ABNORMAL PRIOR TO AND AFTER TRANSECTION OF THE PATIENT TISSUE. THE TISSUE WAS NOT TOO THICK AND THERE WAS NO TISSUE BUNCHING NOTED BY THE SURGICAL STAFF. THERE WERE POST-OPERATIVE COMPLICATIONS REPORTED AND THE SURGEON INDICATED THAT THE PATIENT WAS RECOVERING WELL.
07-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING ON BLOOD VESSELS DURING AN ENDOSCOPIC LOBECTOMY WITH THORACIC SURGERY, THEY COULD FIRE HALF OF THE SHOTS.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED ON SMALL INTESTINE DURING A LAPAROSCOPIC GASTROPLASTY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED ON SMALL INTESTINE DURING A LAPAROSCOPIC GASTROPLASTY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL GASTRECTOMY, WHILE PERFORMING DUODENECTOMY, THE RELOAD COULDN'T BE CLOSED AND THE STAPLER WAS NOT ABLE TO FIRE. THEN A NEW PRODUCT WAS TAKEN OUT TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL GASTRECTOMY, WHILE PERFORMING DUODENECTOMY, THE RELOAD COULDN'T BE CLOSED AND THE STAPLER WAS NOT ABLE TO FIRE. THEN A NEW PRODUCT WAS TAKEN OUT TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
04-01-2019 | Malfunction | covidien lp llc north haven | 04-01-2019 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY AT THE 1ST FIRING THE HANDLE BROKE AND THE DEVICE WAS NOT ABLE TO FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING STAPLING OF THE COLON ON A LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE STAPLER WAS ABLE TO FIRE, BUT THE STAPLER POPPED AT THE TIME OF THE STAPLING AND GOT THE GAUNTLET LOOSE. IT WAS STATED THAT THE STAPLER WAS NO LONGER WORKING. THEY REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING OF THE LUNG FOR A BIOPSY DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT STAPLER WAS UNABLE TO FIRE. THE SURGEON USED ANOTHER STAPLER TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | UNKNOWN ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING OF THE LUNG FOR A BIOPSY DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT STAPLER WAS UNABLE TO FIRE. THE SURGEON USED ANOTHER STAPLER TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
04-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, WHEN PA WAS HANDLING WITH POWER HANDLE AT THE FOURTH FIRING, AFTER THE GREEN BUTTON WAS PRESSED, FIRING COULD NOT BE PERFORMED. THERE IS NO TOTALLY SUSPICIOUS OPERATION FROM LOADING TO STAPLING OPERATION. IT SEEMED THAT ABNORMAL SOUND WAS HEARD FROM THE INSTRUMENT AFTER THE GREEN BUTTON WAS PRESSED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
04-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | SIGNIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, WHEN PA WAS HANDLING WITH POWER HANDLE AT THE FOURTH FIRING, AFTER THE GREEN BUTTON WAS PRESSED, FIRING COULD NOT BE PERFORMED. THERE IS NO TOTALLY SUSPICIOUS OPERATION FROM LOADING TO STAPLING OPERATION. IT SEEMED THAT ABNORMAL SOUND WAS HEARD FROM THE INSTRUMENT AFTER THE GREEN BUTTON WAS PRESSED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL GASTRECTOMY, AT THE TIME OF RESECTION, THE FIRING STOPPED WHEN USING THE POWERED STAPLER. AND BECAUSE THE RESECTION WAS COMPLETED WITHOUT PROBLEMS WITH THE SECOND RELOAD, IT WAS ASKED IF THE CARTRIDGE WAS DEFECTIVE. WHETHER THE TISSUE WAS THICK OR THE BUTTON HAD NOT BEEN PRESSED TWICE WAS BEING CHECKED. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | IDRIVE ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL GASTRECTOMY, AT THE TIME OF RESECTION, THE FIRING STOPPED WHEN USING THE POWERED STAPLER. AND BECAUSE THE RESECTION WAS COMPLETED WITHOUT PROBLEMS WITH THE SECOND RELOAD, IT WAS ASKED IF THE CARTRIDGE WAS DEFECTIVE. WHETHER THE TISSUE WAS THICK OR THE BUTTON HAD NOT BEEN PRESSED TWICE WAS BEING CHECKED. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL GASTRECTOMY, AT THE TIME OF RESECTION, THE FIRING STOPPED WHEN USING THE POWERED STAPLER. AND BECAUSE THE RESECTION WAS COMPLETED WITHOUT PROBLEMS WITH THE SECOND RELOAD, IT WAS ASKED IF THE CARTRIDGE WAS DEFECTIVE. WHETHER THE TISSUE WAS THICK OR THE BUTTON HAD NOT BEEN PRESSED TWICE WAS BEING CHECKED. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-07-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, ON FIRING THE STAPLER, THE HANDLE WAS ¿STICKING OR STUCK¿ WHEN TRYING TO TOGGLE THE NEEDLE BACK AND FORTH. ANOTHER STAPLER AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY WHEN THE DEVICE WAS SET, THEY ATTEMPTED TO STAPLE THE VESSEL, THE HANDLE WAS STIFF AND COULDN'T FIRE. THE OTHER DOCTOR SET THE DEVICE TO VESSEL AGAIN AND FIRED, IT WORKED WITH NO PROBLEM AND THE STAPLES WERE FORMED AS USUAL. HOWEVER, IN CASE OF THE EVENT OCCURRED AGAIN, AFTER THE SECOND FIRING THEY REPLACED WITH A NEW HANDLE. THE NEW HANDLE FIRED WITH NO PROBLEM. THERE WAS NO PATIENT HARM.
08-12-2018 | Injury | covidien lp llc north haven | 04-01-2019 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING OF THE SECOND RELOAD DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON OBSERVED WHITISH COLOR NASAL PROBE. EVIDENTLY, THE NASAL PROBE WAS ADVANCED FAR AND THE SURGEON WAS ACCIDENTALLY STAPLED PROBE WHEN FORMING THE SLEEVE. THE SURGEON HAD TO CUT OUT PROBE WITH LAPAROSCOPIC SHEERS. THIS RESULTED TO UNANTICIPATED TISSUE LOSS, EXTEND PATIENT HOSPITALIZATION AND EXTEND SURGICAL TIME FOR ABOUT 2 HOURS AND 30 MINUTES.
30-11-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON STARTED TO FIRE BUT MIDWAY THRU THE FIRING SEQUENCE, THE STAPLER STOPPED AND WOULD NOT CONTINUE, LEAVING THE INITIAL STAPLES IN SITU. IT WAS REPORTED THAT THERE IS AN INCOMPLETE STAPLE LINE ON THE DISTAL PART. THE SURGEON THEN REMOVED THE RELOAD AND USED ANOTHER RELOAD TO COMPLETE THE PROCEDURE.
19-10-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | PROXIMATE*RL TITANIUM STAPLER | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, THE STAPLER IS VERY HARD AND AT THE TIME OF SHOOTING THE DEVICE WAS OPENED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
06-11-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | 75 MM NTLC SELECTABLE RELOAD | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE GRIPPING SURFACE OF THE CARTRIDGE WAS DAMAGED. NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
11-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DAMAGED FIRING KNOB. IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF CARDIAC CANCER, WHEN SEVERING THE STOMACH ON THE 2ND FIRING, THE FIRING KNOB WAS BROKEN OFF. ALL THE PIECES WOULD BE RETURNED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
30-10-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, ON THE THIRD FIRE ON THE PULMONARY ARTERY, THE STAPLER JAMMED AND THE BLADE DIDN'T ADVANCED.
07-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | ECHELON ENDOPATH STAPLER | Insufficient Information | E2401 | Post Operative Wound Infection | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE TITLE: "LAPAROSCOPIC LEFT LATERAL SECTIONECTOMY AS A TRAINING PROCEDURE FOR SURGEONS LEARNING LAPAROSCOPIC HEPATECTOMY." AUTHORS: YASUSHI HASEGAWA, HIROYUKI NITTA, AKIRA SASAKI, TAKESHI TAKAHARA, NAOKO ITO, TOMOHIRO FUJITA, SHOJI KANNO, SATOSHI NISHIZUKA, GO WAKABAYASHI. CITATION: J HEPATOBILIARY PANCREAT SCI. 2013; 20: 525¿530. DOI: 10.1007/S00534-012-0591-X. LAPAROSCOPIC LIVER RESECTION REMAINS LIMITED TO A RELATIVELY SMALL NUMBER OF INSTITUTIONS BECAUSE OF INSUFFICIENT HEPATIC AND LAPAROSCOPIC SURGICAL EXPERIENCE AND FEW TRAINING OPPORTUNITIES. THE AIM OF THIS STUDY WAS TO ASSESS THE FEASIBILITY AND SAFETY OF AN IMPROVED LAPAROSCOPIC LEFT LATERAL SECTIONECTOMY TECHNIQUE AS A TRAINING PROCEDURE FOR NEW SURGEONS. A TOTAL OF 24 LAPAROSCOPIC LEFT LATERAL SECTIONECTOMIES (LLLSS; 14 MALE AND 10 FEMALE PATIENTS; AGE RANGE: 39 TO 85 YEARS OLD) WERE RETROSPECTIVELY REVIEWED. PATIENTS WERE DIVIDED INTO 3 GROUPS WITH 8 PATIENTS IN EACH: THOSE UNDERGOING SURGERY BY EXPERT SURGEONS PRIOR TO 2008 (GROUP A); THOSE UNDERGOING SURGERY BY EXPERT SURGEONS AFTER 2008, WHEN A STANDARDIZED LLLS TECHNIQUE WAS ADOPTED (GROUP B); AND THOSE UNDERGOING LLLS BY JUNIOR SURGEONS BEING TRAINED (GROUP C). THE LLLS PROCEDURE, CURRENTLY USED AT THE HOSPITAL, IS CHARACTERIZED BY THE ABSENCE OF THE PRINGLE MANEUVER; TRACTION OF THE ROUND LIGAMENT WITH ENDOLOOP LIGATURES (ETHICON); MINIMAL LIVER PARENCHYMAL TRANSECTION WITHOUT EXPOSURE OF THE PORTAL PEDICLES OR HEPATIC VEIN; AND THE USE OF LINEAR STAPLERS TO DIVIDE THE PORTAL PEDICLES OR HEPATIC VEIN. THE ROUND LIGAMENT AND THE FALCIFORM LIGAMENT ARE DIVIDED WITH A HARMONIC SCALPEL THROUGH THE RIGHT PORT. THE ROUND LIGAMENT IS GRASPED WITH ENDOLOOP LIGATURES (ETHICON) AND PULLED TO THE RIGHT USING A SUTURING DEVICE. BOTH PORTAL PEDICLES, WITH THE SURROUNDING LIVER TISSUE, ARE DIVIDED USING AN ECHELON FLEX STAPLER (ETHICON) STAPLER WITH A 60-MM GOLD CARTRIDGE. THEREAFTER, THE LEFT HEPATIC VEIN, WITHIN THE LIVER PARENCHYMA, IS DIVIDED USING THE ECHELON FLEX STAPLER (ETHICON) WITH A 60-MM WHITE CARTRIDGE. THERE IS ALMOST NO BLEEDING AFTER DIVIDING THE LIVER WITH THE LINEAR STAPLERS. IN ALL GROUPS, REPORTED COMPLICATION INCLUDED WOUND INFECTION (N-1). IN CONCLUSION, THE PROCEDURE FOR PURE LLLS HAS BEEN STANDARDIZED AND IS REPRODUCIBLE. TRAINING NEW SURGEONS BY ALLOWING THEM TO GAIN SUPERVISED EXPERIENCE WITH LAPAROSCOPIC HEPATECTOMY IS SAFE AND FEASIBLE, AND LLLS IS A USEFUL OPERATION FOR TRAINING SURGEONS PRIOR TO THEM ATTEMPTING MORE COMPLEX LAPAROSCOPIC HEPATECTOMY PROCEDURES.
16-07-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE ITEM IS FAULTY AND KEEPS JAMMING. NO PATIENT CONSEQUENCE REPORTED.
16-07-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE ITEM IS FAULTY AND KEEPS JAMMING. NO PATIENT CONSEQUENCE REPORTED.
04-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE FOLLOWING WAS REPORTED BY S.K. THEATRE LEAD ¿ ¿I HAVE NOTICED THERE IS SOMETHING WRONG WITH A BATCH OF CDH 29 BOWEL ANASTOMOSIS CIRCULAR STAPLE WHICH WE USE FOR ANTERIOR RESECTIONS. PROBLEM WHEN HE DEPLOYED THE GUN IN WHICH PROLENE PURSE STRING WAS NOT CUTTING.¿ I DISCUSSED THE EVENT WITH THEATRE LEAD WHO ADVISED THAT THERE WAS A ¿CRUNCH¿ AND THAT A STAPLED ANASTOMOSIS WAS FORMED. SURGEON CUT THE PROLENE PURSE STRING SUTURE LAPAROSCOPICALLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
03-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE FOLLOWING WAS REPORTED BY (B)(6) (THEATRE LEAD) ¿ "I HAVE NOTICED THERE IS SOMETHING WRONG WITH A BATCH OF CDH 29 BOWEL ANASTOMOSIS CIRCULAR STAPLE WHICH WE USE FOR ANTERIOR RESECTIONS. ON (B)(6) 2018 SURGEON (B)(6) HAD THE PROBLEM WHEN HE DEPLOYED THE GUN IN WHICH PROLENE PURSE STRING WAS NOT CUTTING. I DISCUSSED THE EVENT WITH THEATRE LEAD WHO ADVISED THAT THERE WAS A ¿CRUNCH¿ AND THAT A STAPLED ANASTOMOSIS WAS FORMED. THEATRE LEAD ADVISED THAT (B)(6) HAD TO CUT THE PROLENE PURSE STRING SUTURE FROM THE BOTTOM END." THERE WERE NO PATIENT CONSEQUENCES REPORTED.
19-10-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT CLIP AND DID NOT CUT. A SECOND DEVICE WAS USED WITHOUT PROBLEMS.
19-10-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT CLIP AND DID NOT CUT. A SECOND DEVICE WAS USED WITHOUT PROBLEMS.
04-01-2019 | Malfunction | us surgical puerto rico | 04-01-2019 | UNKNOWN GIA PRODUCT | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY ON AN OPEN PROCEDURE, THE DEVICE'S FIRING KNOB BROKE AND COMPONENT DISENGAGED. THERE IS NO PATIENT INJURY.
04-01-2019 | Malfunction | covidien lp llc north haven | 04-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON STARTED TO FIRE BUT MIDWAY THRU THE FIRING SEQUENCE, THE STAPLER STOPPED AND WOULD NOT CONTINUE, LEAVING THE INITIAL STAPLES IN SITU. IT WAS REPORTED THAT THERE IS AN INCOMPLETE STAPLE LINE ON THE DISTAL PART. THE SURGEON THEN REMOVED THE RELOAD AND USED ANOTHER RELOAD TO COMPLETE THE PROCEDURE.
04-01-2019 | Malfunction | covidien lp llc north haven | 04-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON STARTED TO FIRE BUT MIDWAY THRU THE FIRING SEQUENCE, THE STAPLER STOPPED AND WOULD NOT CONTINUE, LEAVING THE INITIAL STAPLES IN SITU. IT WAS REPORTED THAT THERE IS AN INCOMPLETE STAPLE LINE ON THE DISTAL PART. THE SURGEON THEN REMOVED THE RELOAD AND USED ANOTHER RELOAD TO COMPLETE THE PROCEDURE.
21-11-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | GIA | Failure to Fire | A050501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING, BOWEL OBSTRUCTION LAPAROTOMY PROCEDURE, THE BLACK PUSHER THUMB PIECE ABLE TO BE MOVED HOWEVER DEVICE WAS STUCK AND COULDN¿T FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-11-2018 | Injury | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE RELOADS CUT ALONE AND THE BLADE OF THE RELOADS WAS SHOOTING ALONE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
27-11-2018 | Injury | us surgical puerto rico | 04-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE RELOADS CUT ALONE AND THE BLADE OF THE RELOADS WAS SHOOTING ALONE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
05-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WHEN THE HANDLE WAS SQUEEZED, THE MATERIAL INSIDE THE STAPLER'S SHAFT POPPED OUT FROM THE JOINT OF THE RELOAD AND NEITHER DID THE STAPLES NOR THE KNIFE COME OUT. THEY USED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WHEN THE HANDLE WAS SQUEEZED, THE MATERIAL INSIDE THE STAPLER'S SHAFT POPPED OUT FROM THE JOINT OF THE RELOAD AND NEITHER DID THE STAPLES NOR THE KNIFE COME OUT. THEY USED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WHEN THE HANDLE WAS SQUEEZED, THE MATERIAL INSIDE THE STAPLER'S SHAFT POPPED OUT FROM THE JOINT OF THE RELOAD AND NEITHER DID THE STAPLES NOR THE KNIFE COME OUT. THEY USED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WHEN THE HANDLE WAS SQUEEZED, THE MATERIAL INSIDE THE STAPLER'S SHAFT POPPED OUT FROM THE JOINT OF THE RELOAD AND NEITHER DID THE STAPLES NOR THE KNIFE COME OUT. THEY USED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Device Slipped | A051204 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING LUNG TISSUE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND WHEN THE HANDLE WAS SQUEEZED, THERE WAS A CRACK SOUND AND THE GEAR WAS BROKEN. THEY USED ANOTHER HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Device Slipped | A051204 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING LUNG TISSUE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND WHEN THE HANDLE WAS SQUEEZED, THERE WAS A CRACK SOUND AND THE GEAR WAS BROKEN. THEY USED ANOTHER HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING LUNG TISSUE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND WHEN THE HANDLE WAS SQUEEZED, THERE WAS A CRACK SOUND AND THE GEAR WAS BROKEN. THEY USED ANOTHER HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING LUNG TISSUE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND WHEN THE HANDLE WAS SQUEEZED, THERE WAS A CRACK SOUND AND THE GEAR WAS BROKEN. THEY USED ANOTHER HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING LUNG TISSUE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND WHEN THE HANDLE WAS SQUEEZED, THERE WAS A CRACK SOUND AND THE GEAR WAS BROKEN. THEY USED ANOTHER HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING LUNG TISSUE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND WHEN THE HANDLE WAS SQUEEZED, THERE WAS A CRACK SOUND AND THE GEAR WAS BROKEN. THEY USED ANOTHER HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-10-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN LOW ANTERIOR RESECTION, THE EC60A LOADING ECR60B STOPPED ON THE WAY OF THE FIRST FIRING ON THE RECTUM. THEN, THE SURGEON TRIED TO USE ATS45 LOADING 6R45B TO FIRE, BUT THE DEVICE COULD NOT CLAMP THE TARGET TISSUE. THE TARGET TISSUE WAS VERY THICK. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-10-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN LOW ANTERIOR RESECTION, THE EC60A LOADING ECR60B STOPPED ON THE WAY OF THE FIRST FIRING ON THE RECTUM. THEN, THE SURGEON TRIED TO USE ATS45 LOADING 6R45B TO FIRE, BUT THE DEVICE COULD NOT CLAMP THE TARGET TISSUE. THE TARGET TISSUE WAS VERY THICK. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-01-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN THE SURGEON FIRED THE GST60W TO TISSUE, THE RELOAD WAS STUCK AND THE GUN COULD NOT BE FIRED, THEN THEY PRESSED THE RELEASE BUTTON TO REMOVE THE RELOAD. THEY USE THE SAME EC60A WANTED TO FIRE THE SECOND AND THE THIRD GST60W, THE SITUATION STILL OCCURRED. FINALLY, THEY USED THE NEW EC60A TO COMPLETE THE PROCEDURE.
01-01-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN THE SURGEON FIRED THE GST60W TO TISSUE, THE RELOAD WAS STUCK AND THE GUN COULD NOT BE FIRED, THEN THEY PRESSED THE RELEASE BUTTON TO REMOVE THE RELOAD. THEY USE THE SAME EC60A WANTED TO FIRE THE SECOND AND THE THIRD GST60W, THE SITUATION STILL OCCURRED. FINALLY, THEY USED THE NEW EC60A TO COMPLETE THE PROCEDURE.
30-11-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ETS ARTIC LNR CUTR NO KNIFE | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT NO STAPLE DEPLOYED. IT WAS REPORTED THAT DURING THE UNCUT RY SURGERY, AFTER SQUEEZED DOWN THE FIRING TRIGGER, NO STAPLE DEPLOYED AND THE FIRING TRIGGER COULD NOT RETURN BACK AUTOMATICALLY. CHANGED ANOTHER RELOAD, THE EVENT RE-HAPPENED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
05-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | ILS 33MM, CURVED | Defective Device | A0203 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, CDH33A WAS USED TO PERFORM THE END TO END ANASTOMOSES. WHEN THE SURGEON MANUALLY ACTIVATED THE DEVICE HE SAID IT FELT DIFFERENT. BOTH PROXIMAL AND DISTAL DOUGHNUTS LOOKED GOOD IN HIS OPINION AND WERE INTACT. APPROXIMATELY TWO DAYS POST OP PATIENT PRESENTED WITH ABDOMINAL PAINS. WAS READMITTED AND OPERATED ON TO REMOVE FECAL MATTER FROM THE ABDOMINAL. THE SURGEON SAID IT APPEARED THE STAPLE LINE DID NOT HOLD.
05-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | ILS 33MM, CURVED | Defective Device | A0203 | Sepsis | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, CDH33A WAS USED TO PERFORM THE END TO END ANASTOMOSES. WHEN THE SURGEON MANUALLY ACTIVATED THE DEVICE HE SAID IT FELT DIFFERENT. BOTH PROXIMAL AND DISTAL DOUGHNUTS LOOKED GOOD IN HIS OPINION AND WERE INTACT. APPROXIMATELY TWO DAYS POST OP PATIENT PRESENTED WITH ABDOMINAL PAINS. WAS READMITTED AND OPERATED ON TO REMOVE FECAL MATTER FROM THE ABDOMINAL. THE SURGEON SAID IT APPEARED THE STAPLE LINE DID NOT HOLD.
05-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | ILS 33MM, CURVED | Defective Device | A0203 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, CDH33A WAS USED TO PERFORM THE END TO END ANASTOMOSES. WHEN THE SURGEON MANUALLY ACTIVATED THE DEVICE HE SAID IT FELT DIFFERENT. BOTH PROXIMAL AND DISTAL DOUGHNUTS LOOKED GOOD IN HIS OPINION AND WERE INTACT. APPROXIMATELY TWO DAYS POST OP PATIENT PRESENTED WITH ABDOMINAL PAINS. WAS READMITTED AND OPERATED ON TO REMOVE FECAL MATTER FROM THE ABDOMINAL. THE SURGEON SAID IT APPEARED THE STAPLE LINE DID NOT HOLD.
13-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | ENDOPATH*THOR ENDO LIN CUT STD | Patient-Device Incompatibility | A0101 | Failure to Anastomose | Event Description: IT WAS REPORTED VIA LITERATURE ARTICLE ENTITLED: CIRCULAR STAPLING VERSUS TRIANGULATING STAPLING FOR THE CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS AFTER ESOPHAGECTOMY IN PATIENTS WITH THORACIC ESOPHAGEAL CANCER: A PROSPECTIVE, RANDOMIZED, CONTROLLED TRIAL AUTHOR: KEIJI HAYATA, MD, PHD, MIKIHITO NAKAMORI, MD, PHD, MASAKI NAKAMURA, MD, PHD, TOSHIYASU OJIMA, MD, PHD, MAKOTO IWAHASHI, MD, PHD, MASAHIRO KATSUDA, MD, PHD, TOSHIAKI TSUJI, MD, PHD, TOMOYA KATO, MD, JYUNYA KITADANI, MD, AKIHIRO TAKEUCHI, MD, HIROTAKA TABATA, MD, AND HIROKI YAMAUE, MD, PHD.. CITATION: SURGERY. 2017; 162: 131-8. DOI: HTTP://DX.DOI.ORG/10.1016/J.SURG.2017.01.013. THE AUTHORS COMPARED TRIANGULATING STAPLING AND CIRCULAR STAPLING FOR CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS REGARDING THE DECREASE IN ANASTOMOTIC STRICTURE AFTER ESOPHAGECTOMY FOR THORACIC ESOPHAGEAL CANCER. BETWEEN AUGUST 2010 AND APRIL 2014, A TOTAL OF 100 PATIENTS ENROLLED IN THIS RANDOMIZED, CONTROLLED TRIAL AT THE WAKAYAMA MEDICAL UNIVERSITY HOSPITAL AND WERE ALLOCATED RANDOMLY TO EITHER THE CIRCULAR STAPLING GROUP (CS; N-49; 40 MALE AND 9 FEMALE PATIENTS¿ AGE: 68 ± 8; BMI: 22.7 ± 7.9) OR THE TRIANGULATING STAPLING GROUP (TS; N-51; 37 MALE AND 14 FEMALE PATIENTS; AGE: 65 ± 9; BMI: 21.4 ± 2.7). THE CS METHOD WAS AN END-TO-SIDE ANASTOMOSIS USING AN ILS-CDH25 CS STAPLING DEVICE (ETHICON). THE CIRCULAR STAPLER WAS INSERTED INTO THE TIP OF THE GASTRIC CONDUIT, AND THE ANASTOMOSIS WAS CREATED AT THE POSTERIOR WALL OF THE GASTRIC CONDUIT APPROXIMATELY 5 CM FROM THE DISTAL END OF THE GASTRIC CONDUIT. THE TS METHOD WAS AN END-TO-END ANASTOMOSIS USING A LINEAR STAPLING DEVICE, A PARALLEL CLOSING SYSTEM AND AN EZ45-3.5 CUTTING KNIFE (ETHICON). IN THE CS GROUP, REPORTED COMPLICATIONS INCLUDED ANASTOMOTIC STRICTURE (N-8), ANASTOMOTIC LEAKAGE (N-5), REFLUX ESOPHAGITIS (N-6), CHYLOTHORAX (N-1), AND CERVICAL LYMPHATIC LEAKAGE (N-1). IN THE TS GROUP, REPORTED COMPLICATIONS INCLUDED ANASTOMOTIC STRICTURE (N-9), ANASTOMOTIC LEAKAGE (N-1), REFLUX ESOPHAGITIS (N-6), CERVICAL LYMPHATIC LEAKAGE (N-2), INTRATHORACIC BLEEDING (N-1) WHICH REQUIRED REOPERATION, HERNIATION OF THE TRANSVERSE COLON INTO THE CHEST VIA THE ESOPHAGEAL HIATUS (N-1) WHICH REQUIRED REOPERATION, AND DISTURBANCES IN THE PASSAGE OF INGESTED FOODS IN THE GASTRIC CONDUIT AT THE ESOPHAGEAL HIATUS (N-1) WHICH REQUIRED REOPERATION. IN CONCLUSION, THIS STUDY COULD NOT SHOW THAT THE TS METHOD WAS SUPERIOR TO THE CS METHOD FOR CERVICAL ESOPHAGOGASTROSTOMY TO REDUCE ANASTOMOTIC STRICTURE. THE TS METHOD, HOWEVER, MIGHT HAVE THE POTENTIAL TO DECREASE THE RATE OF ANASTOMOTIC LEAKAGE; A LARGE-SCALE RCT WILL BE REQUIRED TO ASSESS THIS QUESTION.
13-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | ENDOPATH*THOR ENDO LIN CUT STD | Patient-Device Incompatibility | A0101 | Hernia | Event Description: IT WAS REPORTED VIA LITERATURE ARTICLE ENTITLED: CIRCULAR STAPLING VERSUS TRIANGULATING STAPLING FOR THE CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS AFTER ESOPHAGECTOMY IN PATIENTS WITH THORACIC ESOPHAGEAL CANCER: A PROSPECTIVE, RANDOMIZED, CONTROLLED TRIAL AUTHOR: KEIJI HAYATA, MD, PHD, MIKIHITO NAKAMORI, MD, PHD, MASAKI NAKAMURA, MD, PHD, TOSHIYASU OJIMA, MD, PHD, MAKOTO IWAHASHI, MD, PHD, MASAHIRO KATSUDA, MD, PHD, TOSHIAKI TSUJI, MD, PHD, TOMOYA KATO, MD, JYUNYA KITADANI, MD, AKIHIRO TAKEUCHI, MD, HIROTAKA TABATA, MD, AND HIROKI YAMAUE, MD, PHD.. CITATION: SURGERY. 2017; 162: 131-8. DOI: HTTP://DX.DOI.ORG/10.1016/J.SURG.2017.01.013. THE AUTHORS COMPARED TRIANGULATING STAPLING AND CIRCULAR STAPLING FOR CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS REGARDING THE DECREASE IN ANASTOMOTIC STRICTURE AFTER ESOPHAGECTOMY FOR THORACIC ESOPHAGEAL CANCER. BETWEEN AUGUST 2010 AND APRIL 2014, A TOTAL OF 100 PATIENTS ENROLLED IN THIS RANDOMIZED, CONTROLLED TRIAL AT THE WAKAYAMA MEDICAL UNIVERSITY HOSPITAL AND WERE ALLOCATED RANDOMLY TO EITHER THE CIRCULAR STAPLING GROUP (CS; N-49; 40 MALE AND 9 FEMALE PATIENTS¿ AGE: 68 ± 8; BMI: 22.7 ± 7.9) OR THE TRIANGULATING STAPLING GROUP (TS; N-51; 37 MALE AND 14 FEMALE PATIENTS; AGE: 65 ± 9; BMI: 21.4 ± 2.7). THE CS METHOD WAS AN END-TO-SIDE ANASTOMOSIS USING AN ILS-CDH25 CS STAPLING DEVICE (ETHICON). THE CIRCULAR STAPLER WAS INSERTED INTO THE TIP OF THE GASTRIC CONDUIT, AND THE ANASTOMOSIS WAS CREATED AT THE POSTERIOR WALL OF THE GASTRIC CONDUIT APPROXIMATELY 5 CM FROM THE DISTAL END OF THE GASTRIC CONDUIT. THE TS METHOD WAS AN END-TO-END ANASTOMOSIS USING A LINEAR STAPLING DEVICE, A PARALLEL CLOSING SYSTEM AND AN EZ45-3.5 CUTTING KNIFE (ETHICON). IN THE CS GROUP, REPORTED COMPLICATIONS INCLUDED ANASTOMOTIC STRICTURE (N-8), ANASTOMOTIC LEAKAGE (N-5), REFLUX ESOPHAGITIS (N-6), CHYLOTHORAX (N-1), AND CERVICAL LYMPHATIC LEAKAGE (N-1). IN THE TS GROUP, REPORTED COMPLICATIONS INCLUDED ANASTOMOTIC STRICTURE (N-9), ANASTOMOTIC LEAKAGE (N-1), REFLUX ESOPHAGITIS (N-6), CERVICAL LYMPHATIC LEAKAGE (N-2), INTRATHORACIC BLEEDING (N-1) WHICH REQUIRED REOPERATION, HERNIATION OF THE TRANSVERSE COLON INTO THE CHEST VIA THE ESOPHAGEAL HIATUS (N-1) WHICH REQUIRED REOPERATION, AND DISTURBANCES IN THE PASSAGE OF INGESTED FOODS IN THE GASTRIC CONDUIT AT THE ESOPHAGEAL HIATUS (N-1) WHICH REQUIRED REOPERATION. IN CONCLUSION, THIS STUDY COULD NOT SHOW THAT THE TS METHOD WAS SUPERIOR TO THE CS METHOD FOR CERVICAL ESOPHAGOGASTROSTOMY TO REDUCE ANASTOMOTIC STRICTURE. THE TS METHOD, HOWEVER, MIGHT HAVE THE POTENTIAL TO DECREASE THE RATE OF ANASTOMOTIC LEAKAGE; A LARGE-SCALE RCT WILL BE REQUIRED TO ASSESS THIS QUESTION.
13-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | ENDOPATH*THOR ENDO LIN CUT STD | Patient-Device Incompatibility | A0101 | Blood Loss | Event Description: IT WAS REPORTED VIA LITERATURE ARTICLE ENTITLED: CIRCULAR STAPLING VERSUS TRIANGULATING STAPLING FOR THE CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS AFTER ESOPHAGECTOMY IN PATIENTS WITH THORACIC ESOPHAGEAL CANCER: A PROSPECTIVE, RANDOMIZED, CONTROLLED TRIAL AUTHOR: KEIJI HAYATA, MD, PHD, MIKIHITO NAKAMORI, MD, PHD, MASAKI NAKAMURA, MD, PHD, TOSHIYASU OJIMA, MD, PHD, MAKOTO IWAHASHI, MD, PHD, MASAHIRO KATSUDA, MD, PHD, TOSHIAKI TSUJI, MD, PHD, TOMOYA KATO, MD, JYUNYA KITADANI, MD, AKIHIRO TAKEUCHI, MD, HIROTAKA TABATA, MD, AND HIROKI YAMAUE, MD, PHD.. CITATION: SURGERY. 2017; 162: 131-8. DOI: HTTP://DX.DOI.ORG/10.1016/J.SURG.2017.01.013. THE AUTHORS COMPARED TRIANGULATING STAPLING AND CIRCULAR STAPLING FOR CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS REGARDING THE DECREASE IN ANASTOMOTIC STRICTURE AFTER ESOPHAGECTOMY FOR THORACIC ESOPHAGEAL CANCER. BETWEEN AUGUST 2010 AND APRIL 2014, A TOTAL OF 100 PATIENTS ENROLLED IN THIS RANDOMIZED, CONTROLLED TRIAL AT THE WAKAYAMA MEDICAL UNIVERSITY HOSPITAL AND WERE ALLOCATED RANDOMLY TO EITHER THE CIRCULAR STAPLING GROUP (CS; N-49; 40 MALE AND 9 FEMALE PATIENTS¿ AGE: 68 ± 8; BMI: 22.7 ± 7.9) OR THE TRIANGULATING STAPLING GROUP (TS; N-51; 37 MALE AND 14 FEMALE PATIENTS; AGE: 65 ± 9; BMI: 21.4 ± 2.7). THE CS METHOD WAS AN END-TO-SIDE ANASTOMOSIS USING AN ILS-CDH25 CS STAPLING DEVICE (ETHICON). THE CIRCULAR STAPLER WAS INSERTED INTO THE TIP OF THE GASTRIC CONDUIT, AND THE ANASTOMOSIS WAS CREATED AT THE POSTERIOR WALL OF THE GASTRIC CONDUIT APPROXIMATELY 5 CM FROM THE DISTAL END OF THE GASTRIC CONDUIT. THE TS METHOD WAS AN END-TO-END ANASTOMOSIS USING A LINEAR STAPLING DEVICE, A PARALLEL CLOSING SYSTEM AND AN EZ45-3.5 CUTTING KNIFE (ETHICON). IN THE CS GROUP, REPORTED COMPLICATIONS INCLUDED ANASTOMOTIC STRICTURE (N-8), ANASTOMOTIC LEAKAGE (N-5), REFLUX ESOPHAGITIS (N-6), CHYLOTHORAX (N-1), AND CERVICAL LYMPHATIC LEAKAGE (N-1). IN THE TS GROUP, REPORTED COMPLICATIONS INCLUDED ANASTOMOTIC STRICTURE (N-9), ANASTOMOTIC LEAKAGE (N-1), REFLUX ESOPHAGITIS (N-6), CERVICAL LYMPHATIC LEAKAGE (N-2), INTRATHORACIC BLEEDING (N-1) WHICH REQUIRED REOPERATION, HERNIATION OF THE TRANSVERSE COLON INTO THE CHEST VIA THE ESOPHAGEAL HIATUS (N-1) WHICH REQUIRED REOPERATION, AND DISTURBANCES IN THE PASSAGE OF INGESTED FOODS IN THE GASTRIC CONDUIT AT THE ESOPHAGEAL HIATUS (N-1) WHICH REQUIRED REOPERATION. IN CONCLUSION, THIS STUDY COULD NOT SHOW THAT THE TS METHOD WAS SUPERIOR TO THE CS METHOD FOR CERVICAL ESOPHAGOGASTROSTOMY TO REDUCE ANASTOMOTIC STRICTURE. THE TS METHOD, HOWEVER, MIGHT HAVE THE POTENTIAL TO DECREASE THE RATE OF ANASTOMOTIC LEAKAGE; A LARGE-SCALE RCT WILL BE REQUIRED TO ASSESS THIS QUESTION.
01-01-2017 | Injury | ethicon endo-surgery | 04-01-2019 | PROXIMATE*ILS CURVED CIR STAPL | Patient-Device Incompatibility | A0101 | Failure to Anastomose | Event Description: IT WAS REPORTED VIA LITERATURE ARTICLE ENTITLED: CIRCULAR STAPLING VERSUS TRIANGULATING STAPLING FOR THE CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS AFTER ESOPHAGECTOMY IN PATIENTS WITH THORACIC ESOPHAGEAL CANCER: A PROSPECTIVE, RANDOMIZED, CONTROLLED TRIAL AUTHOR: KEIJI HAYATA, MD, PHD, MIKIHITO NAKAMORI, MD, PHD, MASAKI NAKAMURA, MD, PHD, TOSHIYASU OJIMA, MD, PHD, MAKOTO IWAHASHI, MD, PHD, MASAHIRO KATSUDA, MD, PHD, TOSHIAKI TSUJI, MD, PHD, TOMOYA KATO, MD, JYUNYA KITADANI, MD, AKIHIRO TAKEUCHI, MD, HIROTAKA TABATA, MD, AND HIROKI YAMAUE, MD, PHD.. CITATION: SURGERY. 2017; 162: 131-8. DOI: HTTP://DX.DOI.ORG/10.1016/J.SURG.2017.01.013. THE AUTHORS COMPARED TRIANGULATING STAPLING AND CIRCULAR STAPLING FOR CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS REGARDING THE DECREASE IN ANASTOMOTIC STRICTURE AFTER ESOPHAGECTOMY FOR THORACIC ESOPHAGEAL CANCER. BETWEEN AUGUST 2010 AND APRIL 2014, A TOTAL OF 100 PATIENTS ENROLLED IN THIS RANDOMIZED, CONTROLLED TRIAL AT (B)(6) HOSPITAL AND WERE ALLOCATED RANDOMLY TO EITHER THE CIRCULAR STAPLING GROUP (CS; N-49; 40 MALE AND 9 FEMALE PATIENTS¿ AGE: 68 ± 8; BMI: 22.7 ± 7.9) OR THE TRIANGULATING STAPLING GROUP (TS; N-51; 37 MALE AND 14 FEMALE PATIENTS; AGE: 65 ± 9; BMI: 21.4 ± 2.7). THE CS METHOD WAS AN END-TO-SIDE ANASTOMOSIS USING AN ILS-CDH25 CS STAPLING DEVICE (ETHICON). THE CIRCULAR STAPLER WAS INSERTED INTO THE TIP OF THE GASTRIC CONDUIT, AND THE ANASTOMOSIS WAS CREATED AT THE POSTERIOR WALL OF THE GASTRIC CONDUIT APPROXIMATELY 5 CM FROM THE DISTAL END OF THE GASTRIC CONDUIT. THE TS METHOD WAS AN END-TO-END ANASTOMOSIS USING A LINEAR STAPLING DEVICE, A PARALLEL CLOSING SYSTEM AND AN EZ45-3.5 CUTTING KNIFE (ETHICON). IN THE CS GROUP, REPORTED COMPLICATIONS INCLUDED ANASTOMOTIC STRICTURE (N-8), ANASTOMOTIC LEAKAGE (N-5), REFLUX ESOPHAGITIS (N-6), CHYLOTHORAX (N-1), AND CERVICAL LYMPHATIC LEAKAGE (N-1). IN THE TS GROUP, REPORTED COMPLICATIONS INCLUDED ANASTOMOTIC STRICTURE (N-9), ANASTOMOTIC LEAKAGE (N-1), REFLUX ESOPHAGITIS (N-6), CERVICAL LYMPHATIC LEAKAGE (N-2), INTRATHORACIC BLEEDING (N-1) WHICH REQUIRED REOPERATION, HERNIATION OF THE TRANSVERSE COLON INTO THE CHEST VIA THE ESOPHAGEAL HIATUS (N-1) WHICH REQUIRED REOPERATION, AND DISTURBANCES IN THE PASSAGE OF INGESTED FOODS IN THE GASTRIC CONDUIT AT THE ESOPHAGEAL HIATUS (N-1) WHICH REQUIRED REOPERATION. IN CONCLUSION, THIS STUDY COULD NOT SHOW THAT THE TS METHOD WAS SUPERIOR TO THE CS METHOD FOR CERVICAL ESOPHAGOGASTROSTOMY TO REDUCE ANASTOMOTIC STRICTURE. THE TS METHOD, HOWEVER, MIGHT HAVE THE POTENTIAL TO DECREASE THE RATE OF ANASTOMOTIC LEAKAGE; A LARGE-SCALE RCT WILL BE REQUIRED TO ASSESS THIS QUESTION.
06-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHEN TRANSECTING LUNG TISSUE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND WHEN THE HANDLE WAS SQUEEZED, THERE WAS A CRACK SOUND AND THE GEAR WAS BROKEN. THEY USED ANOTHER HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | FLEX 60 ARTICULATING | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE TITLE: LAPAROSCOPIC SLEEVE GASTRECTOMY AS FIRST STAGE OR DEFINITIVE INTENT IN 300 CONSECUTIVE CASES AUTHOR(S): N. BASSO ¿ G. CASELLA ¿ M. RIZZELLO ¿ F. ABBATINI ¿ E. SORICELLI ¿ G. ALESSANDRI ¿ C. MAGLIO ¿ A. FANTINI CITATION: SURG ENDOSC (2011) 25:444¿449; DOI 10.1007/S00464-010-1187-7. THIS STUDY WAS DESIGNED TO EVALUATE THE RESULTS OF LAPAROSCOPIC SLEEVE GASTRECTOMY (SG) ON COMPLICATIONS, BODY MASS INDEX (BMI), AND COMORBIDITIES RESOLUTION IN 300 CONSECUTIVE OBESE PATIENTS AND TO ANALYZE THE LESSON LEARNED FROM THIS EXPERIENCE. FROM OCT2002 TO NOV2009, 300 PATIENTS UNDERWENT SG FOR FIRST STAGE OF BILIOPANCREATOC DIVERSION WITH DUODENAL SWITCH (N=100; N=63 FEMALE AND N=27 MALE; MEAN BMI OF 54.4 KG/M2 [RANGED 39-81 KG/M2]; MEAN AGE OF 43.6 YEARS [RANGED 21-58 YEARS]) AND FOR DEFINITIVE BARIATRIC PROCEDURE (N=200; N=138 FEMALE AND N=62 MALE; MEAN BMI OF 45.5 KG/M2 [RANGED 31-70.5 KG/M2]; MEAN AGE OF 43.1 YEARS [16-67 YEARS]). IN THE PROCEDURE, HARMONIC SCALPEL WAS USED IN DIVISION OF THE VASCULAR SUPPLY OF THE GASTRIC GREATER CURVATURE UNTIL THE ANGLE OF HIS. IN SOME PATIENTS, SG WAS CREATED USING ECHELON ENDOPATH WITH GOLD CARTRIDGE. IN GROUP 1, THE RESECTED STOMACH WAS EXTRACTED IN A RETRIEVAL BAG USING AN ENDOLOOP AROUND ONE EXTREMITY TO FACILITATE REMOVAL. POSTOPERATIVE COMPLICATION INCLUDED POSTOPERATIVE BLEEDING (N=13) IN WHICH FOUR HAD MELENA AND ONE HAD HEMATEMESIS (SHOWED BLEEDING FROM THE INNER SIDE OF THE STAPLE LINE), AND WERE TREATED CONSERVATIVELY IN (N=11) AND 1-5 UNITS BLOOD TRANSFUSION WAS REQUIRED IN 9 PATIENTS. THE OTHER TWO CASES REQUIRED REOPERATION: IN ONE CASE A RETROGASTRIC VESSEL, RESPONSIBLE FOR THE BLEEDING WAS CLIPPED BY LAPAROSCOPY; IN THE OTHER CASE A LAPAROTOMY WAS PERFORMED. OTHER COMPLICATION INCLUDED STAPLER LINE LEAK (N=9) TREATED WITH BAND REVISION (N=1) AND NON-OPERATIVE TREATMENT OF PERCUTANEOUS DRAINAGE, ENDOSCOPY AND STENTS IN ALL PATIENTS. IN ONE CASE FOR PERSISTENCE OF THE LEAK AFTER CONSERVATIVE TREATMENT AND STENTS PLACEMENT, A RESCUE REOPERATION WAS PERFORMED (ESOPHAGO-JEJUNOSTOMY) 4 MONTHS AFTER SG. AFTER LAPAROTOMY, AN UPPER GASTRECTOMY SAVING THE GASTRIC ANTRUM WAS PERFORMED. THE ESOPHAGO-JEJUNOSTOMY WAS PERFORMED USING A 21-MM CIRCULAR STAPLER AND REINFORCED WITH SEPARATED STITCHES. THE OPERATIVE OPTION WAS CHOSEN MAINLY FOR THE ABSOLUTE NONCOMPLIANCE OF THE PATIENT TO MAINTAIN THE FAST. THE PATIENT DIED ON THIRD POSTOPERATIVE DAY AS A RESULT OF ACUTE CANDIDIASIS SEPSIS ON ICU. IN 3 PATIENTS, ABDOMINAL ABSCESS WAS FOUND. SG IS A SAFE AND EFFECTIVE TREATMENT FOR MORBID OBESITY AT MID-TERM FOLLOW-UP. SG IS EFFECTIVE ON COMORBIDITIES RESOLUTION, ESPECIALLY FOR THE TREATMENT OF DIABETES. LEARNING CURVE AND SURGICAL SKILL PLAY AN IMPORTANT ROLE IN THE INCIDENCE OF MAJOR COMPLICATIONS AND OPERATIVE TIME. SUTURE LINE REINFORCEMENT ALLOWS A SIGNIFICANT REDUCTION OF BLEEDING BUT DOES NOT DECREASE THE OCCURRENCE OF LEAKS WHICH REMAIN A MAJOR PROBLEM.
10-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | FLEX 60 ARTICULATING | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE TITLE: LAPAROSCOPIC SLEEVE GASTRECTOMY AS FIRST STAGE OR DEFINITIVE INTENT IN 300 CONSECUTIVE CASES AUTHOR(S): N. BASSO ¿ G. CASELLA ¿ M. RIZZELLO ¿ F. ABBATINI ¿ E. SORICELLI ¿ G. ALESSANDRI ¿ C. MAGLIO ¿ A. FANTINI CITATION: SURG ENDOSC (2011) 25:444¿449; DOI 10.1007/S00464-010-1187-7. THIS STUDY WAS DESIGNED TO EVALUATE THE RESULTS OF LAPAROSCOPIC SLEEVE GASTRECTOMY (SG) ON COMPLICATIONS, BODY MASS INDEX (BMI), AND COMORBIDITIES RESOLUTION IN 300 CONSECUTIVE OBESE PATIENTS AND TO ANALYZE THE LESSON LEARNED FROM THIS EXPERIENCE. FROM OCT2002 TO NOV2009, 300 PATIENTS UNDERWENT SG FOR FIRST STAGE OF BILIOPANCREATOC DIVERSION WITH DUODENAL SWITCH (N=100; N=63 FEMALE AND N=27 MALE; MEAN BMI OF 54.4 KG/M2 [RANGED 39-81 KG/M2]; MEAN AGE OF 43.6 YEARS [RANGED 21-58 YEARS]) AND FOR DEFINITIVE BARIATRIC PROCEDURE (N=200; N=138 FEMALE AND N=62 MALE; MEAN BMI OF 45.5 KG/M2 [RANGED 31-70.5 KG/M2]; MEAN AGE OF 43.1 YEARS [16-67 YEARS]). IN THE PROCEDURE, HARMONIC SCALPEL WAS USED IN DIVISION OF THE VASCULAR SUPPLY OF THE GASTRIC GREATER CURVATURE UNTIL THE ANGLE OF HIS. IN SOME PATIENTS, SG WAS CREATED USING ECHELON ENDOPATH WITH GOLD CARTRIDGE. IN GROUP 1, THE RESECTED STOMACH WAS EXTRACTED IN A RETRIEVAL BAG USING AN ENDOLOOP AROUND ONE EXTREMITY TO FACILITATE REMOVAL. POSTOPERATIVE COMPLICATION INCLUDED POSTOPERATIVE BLEEDING (N=13) IN WHICH FOUR HAD MELENA AND ONE HAD HEMATEMESIS (SHOWED BLEEDING FROM THE INNER SIDE OF THE STAPLE LINE), AND WERE TREATED CONSERVATIVELY IN (N=11) AND 1-5 UNITS BLOOD TRANSFUSION WAS REQUIRED IN 9 PATIENTS. THE OTHER TWO CASES REQUIRED REOPERATION: IN ONE CASE A RETROGASTRIC VESSEL, RESPONSIBLE FOR THE BLEEDING WAS CLIPPED BY LAPAROSCOPY; IN THE OTHER CASE A LAPAROTOMY WAS PERFORMED. OTHER COMPLICATION INCLUDED STAPLER LINE LEAK (N=9) TREATED WITH BAND REVISION (N=1) AND NON-OPERATIVE TREATMENT OF PERCUTANEOUS DRAINAGE, ENDOSCOPY AND STENTS IN ALL PATIENTS. IN ONE CASE FOR PERSISTENCE OF THE LEAK AFTER CONSERVATIVE TREATMENT AND STENTS PLACEMENT, A RESCUE REOPERATION WAS PERFORMED (ESOPHAGO-JEJUNOSTOMY) 4 MONTHS AFTER SG. AFTER LAPAROTOMY, AN UPPER GASTRECTOMY SAVING THE GASTRIC ANTRUM WAS PERFORMED. THE ESOPHAGO-JEJUNOSTOMY WAS PERFORMED USING A 21-MM CIRCULAR STAPLER AND REINFORCED WITH SEPARATED STITCHES. THE OPERATIVE OPTION WAS CHOSEN MAINLY FOR THE ABSOLUTE NONCOMPLIANCE OF THE PATIENT TO MAINTAIN THE FAST. THE PATIENT DIED ON THIRD POSTOPERATIVE DAY AS A RESULT OF ACUTE CANDIDIASIS SEPSIS ON ICU. IN 3 PATIENTS, ABDOMINAL ABSCESS WAS FOUND. SG IS A SAFE AND EFFECTIVE TREATMENT FOR MORBID OBESITY AT MID-TERM FOLLOW-UP. SG IS EFFECTIVE ON COMORBIDITIES RESOLUTION, ESPECIALLY FOR THE TREATMENT OF DIABETES. LEARNING CURVE AND SURGICAL SKILL PLAY AN IMPORTANT ROLE IN THE INCIDENCE OF MAJOR COMPLICATIONS AND OPERATIVE TIME. SUTURE LINE REINFORCEMENT ALLOWS A SIGNIFICANT REDUCTION OF BLEEDING BUT DOES NOT DECREASE THE OCCURRENCE OF LEAKS WHICH REMAIN A MAJOR PROBLEM.
10-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | FLEX 60 ARTICULATING | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THROUGH JOURNAL ARTICLE TITLE: LAPAROSCOPIC SLEEVE GASTRECTOMY AS FIRST STAGE OR DEFINITIVE INTENT IN 300 CONSECUTIVE CASES AUTHOR(S): N. BASSO ¿ G. CASELLA ¿ M. RIZZELLO ¿ F. ABBATINI ¿ E. SORICELLI ¿ G. ALESSANDRI ¿ C. MAGLIO ¿ A. FANTINI CITATION: SURG ENDOSC (2011) 25:444¿449; DOI 10.1007/S00464-010-1187-7. THIS STUDY WAS DESIGNED TO EVALUATE THE RESULTS OF LAPAROSCOPIC SLEEVE GASTRECTOMY (SG) ON COMPLICATIONS, BODY MASS INDEX (BMI), AND COMORBIDITIES RESOLUTION IN 300 CONSECUTIVE OBESE PATIENTS AND TO ANALYZE THE LESSON LEARNED FROM THIS EXPERIENCE. FROM OCT2002 TO NOV2009, 300 PATIENTS UNDERWENT SG FOR FIRST STAGE OF BILIOPANCREATOC DIVERSION WITH DUODENAL SWITCH (N=100; N=63 FEMALE AND N=27 MALE; MEAN BMI OF 54.4 KG/M2 [RANGED 39-81 KG/M2]; MEAN AGE OF 43.6 YEARS [RANGED 21-58 YEARS]) AND FOR DEFINITIVE BARIATRIC PROCEDURE (N=200; N=138 FEMALE AND N=62 MALE; MEAN BMI OF 45.5 KG/M2 [RANGED 31-70.5 KG/M2]; MEAN AGE OF 43.1 YEARS [16-67 YEARS]). IN THE PROCEDURE, HARMONIC SCALPEL WAS USED IN DIVISION OF THE VASCULAR SUPPLY OF THE GASTRIC GREATER CURVATURE UNTIL THE ANGLE OF HIS. IN SOME PATIENTS, SG WAS CREATED USING ECHELON ENDOPATH WITH GOLD CARTRIDGE. IN GROUP 1, THE RESECTED STOMACH WAS EXTRACTED IN A RETRIEVAL BAG USING AN ENDOLOOP AROUND ONE EXTREMITY TO FACILITATE REMOVAL. POSTOPERATIVE COMPLICATION INCLUDED POSTOPERATIVE BLEEDING (N=13) IN WHICH FOUR HAD MELENA AND ONE HAD HEMATEMESIS (SHOWED BLEEDING FROM THE INNER SIDE OF THE STAPLE LINE), AND WERE TREATED CONSERVATIVELY IN (N=11) AND 1-5 UNITS BLOOD TRANSFUSION WAS REQUIRED IN 9 PATIENTS. THE OTHER TWO CASES REQUIRED REOPERATION: IN ONE CASE A RETROGASTRIC VESSEL, RESPONSIBLE FOR THE BLEEDING WAS CLIPPED BY LAPAROSCOPY; IN THE OTHER CASE A LAPAROTOMY WAS PERFORMED. OTHER COMPLICATION INCLUDED STAPLER LINE LEAK (N=9) TREATED WITH BAND REVISION (N=1) AND NON-OPERATIVE TREATMENT OF PERCUTANEOUS DRAINAGE, ENDOSCOPY AND STENTS IN ALL PATIENTS. IN ONE CASE FOR PERSISTENCE OF THE LEAK AFTER CONSERVATIVE TREATMENT AND STENTS PLACEMENT, A RESCUE REOPERATION WAS PERFORMED (ESOPHAGO-JEJUNOSTOMY) 4 MONTHS AFTER SG. AFTER LAPAROTOMY, AN UPPER GASTRECTOMY SAVING THE GASTRIC ANTRUM WAS PERFORMED. THE ESOPHAGO-JEJUNOSTOMY WAS PERFORMED USING A 21-MM CIRCULAR STAPLER AND REINFORCED WITH SEPARATED STITCHES. THE OPERATIVE OPTION WAS CHOSEN MAINLY FOR THE ABSOLUTE NONCOMPLIANCE OF THE PATIENT TO MAINTAIN THE FAST. THE PATIENT DIED ON THIRD POSTOPERATIVE DAY AS A RESULT OF ACUTE CANDIDIASIS SEPSIS ON ICU. IN 3 PATIENTS, ABDOMINAL ABSCESS WAS FOUND. SG IS A SAFE AND EFFECTIVE TREATMENT FOR MORBID OBESITY AT MID-TERM FOLLOW-UP. SG IS EFFECTIVE ON COMORBIDITIES RESOLUTION, ESPECIALLY FOR THE TREATMENT OF DIABETES. LEARNING CURVE AND SURGICAL SKILL PLAY AN IMPORTANT ROLE IN THE INCIDENCE OF MAJOR COMPLICATIONS AND OPERATIVE TIME. SUTURE LINE REINFORCEMENT ALLOWS A SIGNIFICANT REDUCTION OF BLEEDING BUT DOES NOT DECREASE THE OCCURRENCE OF LEAKS WHICH REMAIN A MAJOR PROBLEM.
07-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY PROCEDURE, DURING THE FIRST FIRING, THE SURGEON CONNECTED THE HANDLE AND THE RELOAD AND CHECKED OPENING AND CLOSING. THE CLOSURE OF THE JAWS WAS WEAK, THE JAWS MOVED WITH CLIP-CLOP SOUND, SO A NEW DEVICE WAS USED CONSIDERING THE SAFETY. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT HARM.
10-12-2018 | Injury | ethicon endo-surgery | 04-01-2019 | ECHELON FLEX 45 STANDARD | Physical Resistance/Sticking | A0509 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A WEDGE RESECTION PROCEDURE, THE ECHELON EC45A WAS STUCK WITH THE KNIFE IN A PARTIALLY FIRED POSITION. GAVE INSTRUCTIONS TO SALES REP AND HE SAID THAT THEY HAD ALREADY PERFORMED THE RESCUE MANEUVERS. SURGEON MADE THE DECISION TO OPEN PATIENT. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE.
07-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY PROCEDURE, DURING THE FIRST FIRING, THE SURGEON CONNECTED THE HANDLE AND THE RELOAD AND CHECKED OPENING AND CLOSING. THE CLOSURE OF THE JAWS WAS WEAK, THE JAWS MOVED WITH CLIP-CLOP SOUND, SO A NEW DEVICE WAS USED CONSIDERING THE SAFETY. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT HARM.
07-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY PROCEDURE, DURING THE FIRST FIRING, THE SURGEON CONNECTED THE HANDLE AND THE RELOAD AND CHECKED OPENING AND CLOSING. THE CLOSURE OF THE JAWS WAS WEAK, THE JAWS MOVED WITH CLIP-CLOP SOUND, SO A NEW DEVICE WAS USED CONSIDERING THE SAFETY. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT HARM.
12-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THAT DURING A LIVER RESECTION, A PVE35A WAS USED TO TAKE THE HEPATIC VEIN. THE VASECR35 RELOAD USED FOR THE HEPATIC VEIN HAD TWO STAPLES IN THE MIDDLE OF THE RELOAD NOT FULLY FORM. THE SURGEON DESCRIBED THE STAPLES THAT WERE NOT FORMED TO LOOK LIKE GOAL POSTS. HE REPORTED THAT WHEN HE USED THE VASECR35 RELOAD THAT THERE WASN¿T ANY LIVER PARENCHYMA WITHIN THE STAPLER, JUST THE HEPATIC VEIN. HE DID NOT THINK THAT THIS HEPATIC VEIN WAS THICKER THAN ANY OTHER HEPATIC VEIN TO CAUSE THE STAPLES TO NOT FULLY FORM. THE SURGEON HAD TO OVERSEW THE STAPLE LINE WITH A PROLENE SUTURE AND USED SURGICEL POWDER AND SURGICEL SNOW TO ACHIEVE HEMOSTASIS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THAT DURING A LIVER RESECTION, A PVE35A WAS USED TO TAKE THE HEPATIC VEIN. THE VASECR35 RELOAD USED FOR THE HEPATIC VEIN HAD TWO STAPLES IN THE MIDDLE OF THE RELOAD NOT FULLY FORM. THE SURGEON DESCRIBED THE STAPLES THAT WERE NOT FORMED TO LOOK LIKE GOAL POSTS. HE REPORTED THAT WHEN HE USED THE VASECR35 RELOAD THAT THERE WASN¿T ANY LIVER PARENCHYMA WITHIN THE STAPLER, JUST THE HEPATIC VEIN. HE DID NOT THINK THAT THIS HEPATIC VEIN WAS THICKER THAN ANY OTHER HEPATIC VEIN TO CAUSE THE STAPLES TO NOT FULLY FORM. THE SURGEON HAD TO OVERSEW THE STAPLE LINE WITH A PROLENE SUTURE AND USED SURGICEL POWDER AND SURGICEL SNOW TO ACHIEVE HEMOSTASIS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
06-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE PULMONARY SEGMENTAL RESECTION, WHEN SEVERING THE POSTERIOR END ARTERY, AFTER FIRED, NOTED BLEEDING,LOST ABOUT 100ML BLOOD, USED HOMELOCK TO STOP BLEEDING, CHECKED THE STAPLE LINE, FOUND SOME STAPLES WITH STRAIGHT LEG, SOME WITH IRREGULAR SHAPE. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS NOT EXTENDED.
06-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE PULMONARY SEGMENTAL RESECTION, WHEN SEVERING THE POSTERIOR END ARTERY, AFTER FIRED, NOTED BLEEDING,LOST ABOUT 100ML BLOOD, USED HOMELOCK TO STOP BLEEDING, CHECKED THE STAPLE LINE, FOUND SOME STAPLES WITH STRAIGHT LEG, SOME WITH IRREGULAR SHAPE. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS NOT EXTENDED.
11-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN COLECTOMY, THE FIRING KNOB WAS BROKEN OFF AT THE FIRST FIRING ON THE COLON. THE FIRING FORCE WAS HIGHER THAN USUAL. NO PIECES WERE LEFT INSIDE THE PATIENT. THE STAPLE HEIGHT WAS BLUE. ANOTHER DEVICE LOADING ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
11-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | 75MM SELECTABLE NEW TLC | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN COLECTOMY, THE FIRING KNOB WAS BROKEN OFF AT THE FIRST FIRING ON THE COLON. THE FIRING FORCE WAS HIGHER THAN USUAL. NO PIECES WERE LEFT INSIDE THE PATIENT. THE STAPLE HEIGHT WAS BLUE. ANOTHER DEVICE LOADING ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
11-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DURING 2ND STAPLE WITH 75MM RELOAD THE STAPLER COULD ONLY RELEASED TILL THE ORIGIN STAPLE LINE (TILL THE HALF). THE TRIGGER MECHANISM (HANDLE) WAS BROKEN. DURING FIXING THE RELOAD THERE WAS NO RESISTANCE - STAPLER STAPLED WITHOUT PROBLEM. THERE UPON A NTLC55 WAS OPENED. DURING RELEASE THE RELOAD THE HANDLE IS ALSO BROKEN. LIKE THE FIRST TIME WITH THE NTLC75 A 1.5CM PLASTIC PART WAS BROKEN (SAME FRACTURE POINT). THE APPLY OF THE RELOAD (BLUE) BEVOR WAS ALSO WITHOUT A RESISTANCE. THE ADJUSTMENT FROM THE STAPLE LINE DEPTH WASN´T CHANGED DURING TRIGGERING. THERE WERE NO CONSEQUENCES FOR THE PATIENT.
11-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | 55MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DURING 2ND STAPLE WITH 75MM RELOAD THE STAPLER COULD ONLY RELEASED TILL THE ORIGIN STAPLE LINE (TILL THE HALF). THE TRIGGER MECHANISM (HANDLE) WAS BROKEN. DURING FIXING THE RELOAD THERE WAS NO RESISTANCE - STAPLER STAPLED WITHOUT PROBLEM. THEREUPON A NTLC55 WAS OPENED. DURING RELEASE THE RELOAD THE HANDLE IS ALSO BROKEN. LIKE THE FIRST TIME WITH THE NTLC75 A 1.5CM PLASTIC PART WAS BROKEN (SAME FRACTURE POINT). THE APPLY OF THE RELOAD (BLUE) BEVOR WAS ALSO WITHOUT A RESISTANCE. THE ADJUSTMENT FROM THE STAPLE LINE DEPTH WASN´T CHANGED DURING TRIGGERING. THERE WERE NO CONSEQUENCES FOR THE PATIENT.
11-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | 55MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DURING 2ND STAPLE WITH 75MM RELOAD THE STAPLER COULD ONLY RELEASED TILL THE ORIGIN STAPLE LINE (TILL THE HALF). THE TRIGGER MECHANISM (HANDLE) WAS BROKEN. DURING FIXING THE RELOAD THERE WAS NO RESISTANCE - STAPLER STAPLED WITHOUT PROBLEM. THEREUPON A NTLC55 WAS OPENED. DURING RELEASE THE RELOAD THE HANDLE IS ALSO BROKEN. LIKE THE FIRST TIME WITH THE NTLC75 A 1.5CM PLASTIC PART WAS BROKEN (SAME FRACTURE POINT). THE APPLY OF THE RELOAD (BLUE) BEVOR WAS ALSO WITHOUT A RESISTANCE. THE ADJUSTMENT FROM THE STAPLE LINE DEPTH WASN´T CHANGED DURING TRIGGERING. THERE WERE NO CONSEQUENCES FOR THE PATIENT.
11-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | 55MM SELECTABLE NEW TLC | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DURING 2ND STAPLE WITH 75MM RELOAD THE STAPLER COULD ONLY RELEASED TILL THE ORIGIN STAPLE LINE (TILL THE HALF). THE TRIGGER MECHANISM (HANDLE) WAS BROKEN. DURING FIXING THE RELOAD THERE WAS NO RESISTANCE - STAPLER STAPLED WITHOUT PROBLEM. THEREUPON A NTLC55 WAS OPENED. DURING RELEASE THE RELOAD THE HANDLE IS ALSO BROKEN. LIKE THE FIRST TIME WITH THE NTLC75 A 1.5CM PLASTIC PART WAS BROKEN (SAME FRACTURE POINT). THE APPLY OF THE RELOAD (BLUE) BEVOR WAS ALSO WITHOUT A RESISTANCE. THE ADJUSTMENT FROM THE STAPLE LINE DEPTH WASN´T CHANGED DURING TRIGGERING. THERE WERE NO CONSEQUENCES FOR THE PATIENT.
11-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | 55MM SELECTABLE NEW TLC | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DURING 2ND STAPLE WITH 75MM RELOAD THE STAPLER COULD ONLY RELEASED TILL THE ORIGIN STAPLE LINE (TILL THE HALF). THE TRIGGER MECHANISM (HANDLE) WAS BROKEN. DURING FIXING THE RELOAD THERE WAS NO RESISTANCE - STAPLER STAPLED WITHOUT PROBLEM. THEREUPON A NTLC55 WAS OPENED. DURING RELEASE THE RELOAD THE HANDLE IS ALSO BROKEN. LIKE THE FIRST TIME WITH THE NTLC75 A 1.5CM PLASTIC PART WAS BROKEN (SAME FRACTURE POINT). THE APPLY OF THE RELOAD (BLUE) BEVOR WAS ALSO WITHOUT A RESISTANCE. THE ADJUSTMENT FROM THE STAPLE LINE DEPTH WASN´T CHANGED DURING TRIGGERING. THERE WERE NO CONSEQUENCES FOR THE PATIENT.
12-11-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | EEA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON PROCEDURE, AFTER FIRING, THE KNOB WAS TURNED TWO TIMES AND THE ANVIL WAS TILTED. HOWEVER, UPON REMOVING THE DEVICE, THE ANVIL STAYED IN THE COLON. THE SURGEON WAS ABLE TO RETRIEVE THE ANVIL AND THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THEY OPENED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | ethicon endo-surgery | 04-01-2019 | ILS 29MM, CURVED | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A RESECTION OF ADNEXAL TUMOR BY LAPAROSCOPY, EVENT: FAILURE. WHEN THE SURGEON WAS SHOTTING WITH THE CIRCULAR STAPLER29 (CDH29A) THIS DID NOT PERFORM THE PROPER STAMPING. THE STAPLER ONLY CUT THE TISSUE THAT HELD UNDER PRESSURE. THE HEALTH STATUS OF THE PATIENT WILL BE ASSESSED BY THE TREATING DOCTOR BUT IT IS POSSIBLE THAT IT WILL BE REINCISEED TO IMPROVE ITS STATUS, AT A LATER DATE AND TO BE DEFINED.
03-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE TRANSECTING THE COLON DURING A LAPAROSCOPIC SIGMOID RESECTION, THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THEY HAD TO OPEN THE JAWS AND PUT ANOTHER RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE TRANSECTING THE COLON DURING A LAPAROSCOPIC SIGMOID RESECTION, THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THEY HAD TO OPEN THE JAWS AND PUT ANOTHER RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-10-2018 | Injury | ethicon endo-surgery | 04-01-2019 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | Fistula | Event Description: IT WAS REPORTED THAT DURING THE SURGICAL PROCEDURE THE (STAPLER) EQUIPMENT WAS LOCKED, WHICH MADE IT IMPOSSIBLE TO SECURELY AND EFFECTIVELY SUTURE / CLOSE THE INTERVENED ORGAN (STOMACH). THE PATIENT EXPERIENCED A POST OPERATIVE FISTULA.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | MULTIFIRE ENDO GIA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ZENKER'S DIVERTICULUM PROCEDURE, WHILE FIRING THE STAPLER ON THE POUCH, THE RELOAD CUT TISSUE, BUT DID NOT DEPLOY ANY STAPLES. ADDITIONAL RELOADS WERE OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 04-01-2019 | MULTIFIRE ENDO GIA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ZENKER'S DIVERTICULUM PROCEDURE, WHILE FIRING THE STAPLER ON THE POUCH, THE RELOAD CUT TISSUE, BUT DID NOT DEPLOY ANY STAPLES. ADDITIONAL RELOADS WERE OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Premature Activation | A150103 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR TRANSECTION OF PROXIMAL BOWEL AT THE SPLENIC FLEXURE, THE DEVICE FIRED BUT EJECTED THE BLUE RELOAD WITHOUT FIRING STAPLES OR TRANSECTING THE BOWEL. THE STAPLE CARTRIDGE MOVED OUT FROM ITS HOUSING, FELL TO THE CAVITY OF THE PATIENT AND WAS RETRIEVED. THE JAWS LOCK ON TISSUE BUT COULD BE UNLOCKED WITHOUT CUTTING OPEN THE BOWEL. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR TRANSECTION OF PROXIMAL BOWEL AT THE SPLENIC FLEXURE, THE DEVICE FIRED BUT EJECTED THE BLUE RELOAD WITHOUT FIRING STAPLES OR TRANSECTING THE BOWEL. THE STAPLE CARTRIDGE MOVED OUT FROM ITS HOUSING, FELL TO THE CAVITY OF THE PATIENT AND WAS RETRIEVED. THE JAWS LOCK ON TISSUE BUT COULD BE UNLOCKED WITHOUT CUTTING OPEN THE BOWEL. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR TRANSECTION OF PROXIMAL BOWEL AT THE SPLENIC FLEXURE, THE DEVICE FIRED BUT EJECTED THE BLUE RELOAD WITHOUT FIRING STAPLES OR TRANSECTING THE BOWEL. THE STAPLE CARTRIDGE MOVED OUT FROM ITS HOUSING, FELL TO THE CAVITY OF THE PATIENT AND WAS RETRIEVED. THE JAWS LOCK ON TISSUE BUT COULD BE UNLOCKED WITHOUT CUTTING OPEN THE BOWEL. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR TRANSECTION OF PROXIMAL BOWEL AT THE SPLENIC FLEXURE, THE DEVICE FIRED BUT EJECTED THE BLUE RELOAD WITHOUT FIRING STAPLES OR TRANSECTING THE BOWEL. THE STAPLE CARTRIDGE MOVED OUT FROM ITS HOUSING, FELL TO THE CAVITY OF THE PATIENT AND WAS RETRIEVED. THE JAWS LOCK ON TISSUE BUT COULD BE UNLOCKED WITHOUT CUTTING OPEN THE BOWEL. THERE WAS NO PATIENT INJURY.
01-11-2018 | Injury | covidien lp llc north haven | 04-01-2019 | UNKNOWN EGIA ADAPTER | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A ROBOTIC LAPAROSCOPIC PROSTATE PROCEDURE, WHEN FIRING ACROSS THE DORSAL VENOUS COMPLEX THE STAPLE LINE IS INCOMPLETE IN THE DISTAL PART. THE STAPLES WERE COMPLETE AND FORMED CORRECT OUT 40 MM AND IT DIDN'T JUST STAPLE THE LAST 5MM. THE SURGEON SUTURED THE LAST 5MM.
31-05-2019 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY AT THE 1ST FIRING THE HANDLE BROKE AND THE DEVICE WAS NOT ABLE TO FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-05-2019 | Malfunction | us surgical puerto rico | 04-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY AT THE 1ST FIRING THE HANDLE BROKE AND THE DEVICE WAS NOT ABLE TO FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Injury | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Failure to Cut | A050702 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING ABDOMINOPERINEAL PROCTECTOMY, DURING TRANSECTION OF THE RECTUM, THE GREEN BUTTON WAS PRESSED WITHOUT NOTICING, AND THE FIRST SQUEEZING WAS PERFORMED. THE KNIFE MOVED ABOUT 1 CM ON THE RIGHT SIDE OF THE RECTUM, SO THE ANAL SIDE WAS RESECTED ADDITIONALLY. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. IT WAS NOTED THAT THERE WAS TISSUE DAMAGE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Injury | covidien lp llc north haven | 04-01-2019 | ENDO GIA | Failure to Cut | A050702 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING ABDOMINOPERINEAL PROCTECTOMY, DURING TRANSECTION OF THE RECTUM, THE GREEN BUTTON WAS PRESSED WITHOUT NOTICING, AND THE FIRST SQUEEZING WAS PERFORMED. THE KNIFE MOVED ABOUT 1 CM ON THE RIGHT SIDE OF THE RECTUM, SO THE ANAL SIDE WAS RESECTED ADDITIONALLY. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. IT WAS NOTED THAT THERE WAS TISSUE DAMAGE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Injury | us surgical puerto rico | 04-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Cut | A050702 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING ABDOMINOPERINEAL PROCTECTOMY, DURING TRANSECTION OF THE RECTUM, THE GREEN BUTTON WAS PRESSED WITHOUT NOTICING, AND THE FIRST SQUEEZING WAS PERFORMED. THE KNIFE MOVED ABOUT 1 CM ON THE RIGHT SIDE OF THE RECTUM, SO THE ANAL SIDE WAS RESECTED ADDITIONALLY. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. IT WAS NOTED THAT THERE WAS TISSUE DAMAGE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Injury | us surgical puerto rico | 04-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Cut | A050702 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING ABDOMINOPERINEAL PROCTECTOMY, DURING TRANSECTION OF THE RECTUM, THE GREEN BUTTON WAS PRESSED WITHOUT NOTICING, AND THE FIRST SQUEEZING WAS PERFORMED. THE KNIFE MOVED ABOUT 1 CM ON THE RIGHT SIDE OF THE RECTUM, SO THE ANAL SIDE WAS RESECTED ADDITIONALLY. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. IT WAS NOTED THAT THERE WAS TISSUE DAMAGE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-11-2018 | Malfunction | covidien lp llc north haven | 04-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PARTIAL GASTRECTOMY WITH VAGOTOMY, THE STAPLER DID NOT FIRE THE LOAD WHICH MADE IT NECESSARY TO REPLACE THE MATERIAL. IT WAS REPORTED THAT THE OTHER DEVICE WAS REMOVED. THERE WAS NO PATIENT INJURY.
05-11-2018 | Injury | us surgical puerto rico | 04-01-2019 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DEVICES ANVIL BENT BUT WERE FIRED COMPLETELY, HOWEVER, SEVEN DAYS FOLLOWING A LAPAROSCOPIC ESOPHAGOGASTROSTOMY, A LEAK WAS OBSERVED. THEY IMMEDIATELY PERFORMED GASTROSCOPY AND SAW THE HOLE, AND EVENTUALLY USED SUTURES FOR THE CLOSING AND THE CASE WAS COMPLETED.
05-11-2018 | Injury | us surgical puerto rico | 04-01-2019 | EEA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DEVICES ANVIL BENT BUT WERE FIRED COMPLETELY, HOWEVER, SEVEN DAYS FOLLOWING A LAPAROSCOPIC ESOPHAGOGASTROSTOMY, A LEAK WAS OBSERVED. THEY IMMEDIATELY PERFORMED GASTROSCOPY AND SAW THE HOLE, AND EVENTUALLY USED SUTURES FOR THE CLOSING AND THE CASE WAS COMPLETED.
05-11-2018 | Injury | us surgical puerto rico | 04-01-2019 | EEA | Material Twisted/Bent | A040609 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DEVICES ANVIL BENT BUT WERE FIRED COMPLETELY, HOWEVER, SEVEN DAYS FOLLOWING A LAPAROSCOPIC ESOPHAGOGASTROSTOMY, A LEAK WAS OBSERVED. THEY IMMEDIATELY PERFORMED GASTROSCOPY AND SAW THE HOLE, AND EVENTUALLY USED SUTURES FOR THE CLOSING AND THE CASE WAS COMPLETED.
05-11-2018 | Injury | us surgical puerto rico | 04-01-2019 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DEVICES ANVIL BENT BUT WERE FIRED COMPLETELY, HOWEVER, SEVEN DAYS FOLLOWING A LAPAROSCOPIC ESOPHAGOGASTROSTOMY, A LEAK WAS OBSERVED. THEY IMMEDIATELY PERFORMED GASTROSCOPY AND SAW THE HOLE, AND EVENTUALLY USED SUTURES FOR THE CLOSING AND THE CASE WAS COMPLETED.
27-11-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, AFTER THE FIRST FIRING OF THE DEVICE IT WAS PULLED OUT OF THE TROCAR TO REPLACE THE RELOAD. THE JAWS WOULD NOT OPEN. THE MANUAL OVER RIDE WAS TRIED BUT THE JAWS STILL WOULD NOT OPEN. WHILE TRYING TO OPEN THE JAWS THEY BROKE INTO PIECES. THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED.
10-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | POWERED PLUS STANDARD 45 | Loose or Intermittent Connection | A1205 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AT THE TIME OF THE FIRST FIRING WITH THE STAPLER, IT ADVANCED THE BLADE TO THE END, BUT THE BLADE DID NOT RETURN AT THE MOMENT THE TRIGGER WAS RELEASED. WE TRIED TO USE THE REVERSE BUTTON, BUT IT WAS COMPLETELY "LOOSE". WE USED THE MANUAL REVERSE FUNCTION TO OPEN THE STAPLER.
10-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | POWERED PLUS STANDARD 45 | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AT THE TIME OF THE FIRST FIRING WITH THE STAPLER, IT ADVANCED THE BLADE TO THE END, BUT THE BLADE DID NOT RETURN AT THE MOMENT THE TRIGGER WAS RELEASED. WE TRIED TO USE THE REVERSE BUTTON, BUT IT WAS COMPLETELY "LOOSE". WE USED THE MANUAL REVERSE FUNCTION TO OPEN THE STAPLER.
10-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | POWERED PLUS STANDARD 45 | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AT THE TIME OF THE FIRST FIRING WITH THE STAPLER, IT ADVANCED THE BLADE TO THE END, BUT THE BLADE DID NOT RETURN AT THE MOMENT THE TRIGGER WAS RELEASED. WE TRIED TO USE THE REVERSE BUTTON, BUT IT WAS COMPLETELY "LOOSE". WE USED THE MANUAL REVERSE FUNCTION TO OPEN THE STAPLER.
10-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: THE SURGEON HAD TO FACE WITH AN IMPORTANT BLEEDING AFTER STAPLING AN ARTERY. ACCORDING TO HIM THE STAPLE LINE WAS NOT CONSISTENT ON THE DISTAL PART. A BLEEDING OCCURRED. HE HAD TO CONTROL THIS BLEEDING WITH CLIPS. THE SURGEON PLACED CLIPS TO MANAGE THE BLEEDING. THE LOCATION OF THE BLEEDING WAS BAD AS IT DISTURBED THE STAPLING SEQUENCE OF THE MEDIASTINAL ARTERY.
05-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR TRANSECTION OF PROXIMAL BOWEL AT THE SPLENIC FLEXURE, THE DEVICE FIRED BUT EJECTED THE BLUE RELOAD WITHOUT FIRING STAPLES OR TRANSECTING THE BOWEL. THE STAPLE CARTRIDGE MOVED OUT FROM ITS HOUSING, FELL TO THE CAVITY OF THE PATIENT AND WAS RETRIEVED. THE JAWS LOCK ON TISSUE BUT COULD BE UNLOCKED WITHOUT CUTTING OPEN THE BOWEL. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR TRANSECTION OF PROXIMAL BOWEL AT THE SPLENIC FLEXURE, THE DEVICE FIRED BUT EJECTED THE BLUE RELOAD WITHOUT FIRING STAPLES OR TRANSECTING THE BOWEL. THE STAPLE CARTRIDGE MOVED OUT FROM ITS HOUSING, FELL TO THE CAVITY OF THE PATIENT AND WAS RETRIEVED. THE JAWS LOCK ON TISSUE BUT COULD BE UNLOCKED WITHOUT CUTTING OPEN THE BOWEL. THERE WAS NO PATIENT INJURY.
17-11-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY THE DEVICE WOULD NOT RELEASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED AND THE DEVICE WILL BE RETURNED. THERE IS NO ADDITIONAL INFORMATION.
06-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, WHILE APPROACHING THE PULMONARY VASCULAR, THEY TRIED TO FIRE THE DEVICE BUT THE IT WAS UNABLE TO BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE. THERE WAS A PROBLEM WITH RELEASING. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT HARM.
06-12-2018 | Malfunction | covidienlp llc north haven | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, WHILE APPROACHING THE PULMONARY VASCULAR, THEY TRIED TO FIRE THE DEVICE BUT THE IT WAS UNABLE TO BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE. THERE WAS A PROBLEM WITH RELEASING. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT HARM.
11-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP RIGHT HEMICOLECTOMY THE MANUAL ECHELON 60MM STAPLER WAS USED TO TRANSECT MESENTERY. A WHITE RELOAD WAS USED SUCCESSFULLY BUT THE SECOND WHITE RELOAD FIRING JAMMED DURING THE FIRING SEQUENCE. IT STOPPED ON THE SECOND TRIGGER FIRING. THE SURGEON ATTEMPTED THE KNIFE REVERSE BUT STILL COULD NOT RETURN THE KNIFE BACK TO THE HOME POSITION. AT THIS TIME THE STAPLER WAS PARTIALLY FIRED AND CLOSED ON MESENTERY. THE SURGEON THEN PLACED CLIPS ON EITHER SIDE OF THE STAPLER AND CUT THE TISSUE. THE STAPLER WAS REMOVED FROM THE PATIENT AND SET ASIDE. NO ADVERSE EVENTS REPORTED AND THE CASE WAS SUCCESSFULLY COMPLETED WITH ANOTHER ECHELON 60MM STAPLER WITH NO ISSUES.
07-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, WHEN THE DOCTOR ATTEMPTED TO PERFORM THE FIRST FIRING, DEVICE DID NOT FIRE. NEW RELOAD WAS USED AND THE DEVICE FIRED WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, WHEN THE DOCTOR ATTEMPTED TO PERFORM THE FIRST FIRING, DEVICE DID NOT FIRE. NEW RELOAD WAS USED AND THE DEVICE FIRED WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT OPEN FOLLOWING THE STAPLING. THE DEVICE WAS INSIDE THE PATIENT AND CLOSED ON THE STOMACH. THE SECURITY SYSTEM HAS BEEN ACTIVATED BUT THE DEVICE DID NOT OPEN. THE SURGEON HAS MANAGED TO RELEASE THE STOMACH BY PULL ON THE HANDLE OF DEVICE. NO PATIENT CONSEQUENCE.
07-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATIC RESECTION, ON TRANSECTION OF A VESSEL, THE STAPLER WAS UNABLE TO FIRE AND CLOSE. A NEW DEVICE WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATIC RESECTION, ON TRANSECTION OF A VESSEL, THE STAPLER WAS UNABLE TO FIRE AND CLOSE. A NEW DEVICE WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATIC RESECTION, ON TRANSECTION OF A VESSEL, THE STAPLER WAS UNABLE TO FIRE AND CLOSE. A NEW DEVICE WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | PWRD ECH FLEX 60MM 34CM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF LUNG CANCER, AFTER FIRED WITH ECR60W, THE BLADE COULD NOT RETURN AUTOMATICALLY. USED KNIFE REVERSE BUTTON TO RETURN BLADE. COULD NOT FIRE WITH ECR60B, AND NOTED THE BATTERY WAS ABNORMAL HEATED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
10-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | PWRD ECH FLEX 60MM 34CM SHAFT | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF LUNG CANCER, AFTER FIRED WITH ECR60W, THE BLADE COULD NOT RETURN AUTOMATICALLY. USED KNIFE REVERSE BUTTON TO RETURN BLADE. COULD NOT FIRE WITH ECR60B, AND NOTED THE BATTERY WAS ABNORMAL HEATED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
28-06-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST FIRING OF LAPAROSCOPIC SIGMOIDECTOMY, THEY CONNECTED THE RELOAD TO THE HANDLE THEN THEY TRIED TO FIRE THE DEVICE, HOWEVER THEY COULD NOT. REPLACED BY NEW HANDLE, HOWEVER THE SAME EVENT OCCURRED AND THE SURGEON COULD NOTFIRE THE DEVICE. THEY REPLACED BY NEW RELOAD THE SAME EVENT OCCURRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT INJURY.
29-11-2018 | Injury | ethicon endo-surgery | 07-01-2019 | PWRD ECH FLEX 60MM 34CM SHAFT | Loose or Intermittent Connection | A1205 | Embolism | Event Description: IT WAS REPORTED ON NOVEMBER 29TH, A BARIATRIC SURGERY WAS PERFORMED. EVERYTHING PROCEED NORMALLY. AS A SAFETY PROCEDURE, THE SURGEON INFLATED THE PART OF THE REMOVED STOMACH TO VERIFY THE QUALITY OF THE STAPLE LINE, AND IT PASSED THE TEST. HOWEVER, ON THE SECOND POSTOPERATIVE DAY, THE PATIENT PRESENTED WITH AN ACCUMULATION OF FLUID. AN ENDOSCOPY WAS PERFORMED, WHERE IT WAS IDENTIFIED THAT A STAPLE WAS LOOSE. A NEW INTERVENTION WAS REQUIRED, NOW AN OPEN PROCEDURE, TO CORRECT THE STAPLE LINE THAT WAS COMPROMISED, AND A LINEAR SHEAR STAPLER - REF. TLC75 WAS USED. AFTER THIS INTERVENTION, THE PATIENT'S HEALTH STATUS STABILIZED, AND ON THE 8TH DAY AFTER, WHEN HE WAS ALREADY ORALLY FEEDING AND WITH HIS DISCHARGE PRESCRIBED, HE HAD AN EMBOLISM AND DIED. THE SURGEON DOES NOT ASSOCIATE THE DEATH OF THE PATIENT WITH OUR MATERIAL. HE DID NOT WANT TO MAKE A PRODUCT COMPLAINT, BUT A REPORT OF THE EVENT. THE PATIENT HAD PREVIOUS SURGERY OF HIATUS HERNIA, AND THE SURGEON BELIEVES THAT THE VALVE OF THIS PREVIOUS SURGERY SHIFTED AND THIS CAUSED A VERY STRONG PRESSURE IN THE STOMACH, WHICH CAUSED THE RELEASE OF THE STAPLE. THIS OCCURRENCE WAS REPAIRED BY THE OPEN INTERVENTION USING THE TLC75 , AS PREVIOUSLY STATED, AND WITH THIS THE PATIENT'S CONDITION STABILIZED. THE PATIENT'S DEATH IS ASSOCIATED WITH EMBOLISM.
29-11-2018 | Injury | ethicon endo-surgery | 07-01-2019 | PWRD ECH FLEX 60MM 34CM SHAFT | Loose or Intermittent Connection | A1205 | No Code Available | Event Description: IT WAS REPORTED ON NOVEMBER 29TH, A BARIATRIC SURGERY WAS PERFORMED. EVERYTHING PROCEED NORMALLY. AS A SAFETY PROCEDURE, THE SURGEON INFLATED THE PART OF THE REMOVED STOMACH TO VERIFY THE QUALITY OF THE STAPLE LINE, AND IT PASSED THE TEST. HOWEVER, ON THE SECOND POSTOPERATIVE DAY, THE PATIENT PRESENTED WITH AN ACCUMULATION OF FLUID. AN ENDOSCOPY WAS PERFORMED, WHERE IT WAS IDENTIFIED THAT A STAPLE WAS LOOSE. A NEW INTERVENTION WAS REQUIRED, NOW AN OPEN PROCEDURE, TO CORRECT THE STAPLE LINE THAT WAS COMPROMISED, AND A LINEAR SHEAR STAPLER - REF. TLC75 WAS USED. AFTER THIS INTERVENTION, THE PATIENT'S HEALTH STATUS STABILIZED, AND ON THE 8TH DAY AFTER, WHEN HE WAS ALREADY ORALLY FEEDING AND WITH HIS DISCHARGE PRESCRIBED, HE HAD AN EMBOLISM AND DIED. THE SURGEON DOES NOT ASSOCIATE THE DEATH OF THE PATIENT WITH OUR MATERIAL. HE DID NOT WANT TO MAKE A PRODUCT COMPLAINT, BUT A REPORT OF THE EVENT. THE PATIENT HAD PREVIOUS SURGERY OF HIATUS HERNIA, AND THE SURGEON BELIEVES THAT THE VALVE OF THIS PREVIOUS SURGERY SHIFTED AND THIS CAUSED A VERY STRONG PRESSURE IN THE STOMACH, WHICH CAUSED THE RELEASE OF THE STAPLE. THIS OCCURRENCE WAS REPAIRED BY THE OPEN INTERVENTION USING THE TLC75 , AS PREVIOUSLY STATED, AND WITH THIS THE PATIENT'S CONDITION STABILIZED. THE PATIENT'S DEATH IS ASSOCIATED WITH EMBOLISM.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC BULLOUS RESECTION, AFTER CLAMPING THE TISSUE, THE SAFETY BUTTON WAS PRESSED AND THE HANDLE WAS GRASPED BUT THE GRIP DOES NOT MOVE THE HANDLE, IT COULDN'T BE FIRED AND NO STAPLES APPEAR. WHEN CHECKING THE STAPLES, SOME OF THE STAPLE BOARD HAD BEEN REVEALED. NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC BULLOUS RESECTION, AFTER CLAMPING THE TISSUE, THE SAFETY BUTTON WAS PRESSED AND THE HANDLE WAS GRASPED BUT THE GRIP DOES NOT MOVE THE HANDLE, IT COULDN'T BE FIRED AND NO STAPLES APPEAR. WHEN CHECKING THE STAPLES, SOME OF THE STAPLE BOARD HAD BEEN REVEALED. NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC BULLOUS RESECTION, AFTER CLAMPING THE TISSUE, THE SAFETY BUTTON WAS PRESSED AND THE HANDLE WAS GRASPED BUT THE GRIP DOES NOT MOVE THE HANDLE, IT COULDN'T BE FIRED AND NO STAPLES APPEAR. WHEN CHECKING THE STAPLES, SOME OF THE STAPLE BOARD HAD BEEN REVEALED. NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON REPORTED BLEEDING EVENT DURING LAP APPENDECTOMY. ALL OF THE INCIDENTS INVOLVED A PVS STAPLER. THE SURGEON ALSO REPORTED STAPLER LOCKOUTS THAT COULD NOT BE FIXED WITH THE USE OF REVERSE BUTTON. ALL OF THE CASES INVOLVED PEDIATRIC PATIENTS. CAUTERY AND PRESSURE WERE USED TO STOP BLEEDING.
10-12-2018 | Injury | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT UPPER LOBECTOMY DONE VIA AXILLARY THORACOTOMY, THE STAPLER HAND PIECE DID NOT WORK CORRECTLY. THE STAPLER WORKED FINE TWICE WHEN BEING USED ON THE LUNG TISSUE BUT DID NOT WORK ON THE THIRD TIME WHEN IT WAS USED IN VASCULAR TISSUE. THE SURGEON CHANGED THE RELOAD TWICE AND THEN CHANGED THE STAPLER. IT WAS REPORTED THERE WAS AN EXTENSION OF THE INCISION BY MORE THAN 1 INCH SINCE THORACOTOMY WAS NEEDED TO PERFORM SAVE AND RADICAL LOBECTOMY FOR A CENTRAL LUNG CANCER AND SURGERY TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
10-12-2018 | Injury | covidien lp llc north haven | 07-01-2019 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT UPPER LOBECTOMY DONE VIA AXILLARY THORACOTOMY, THE STAPLER HAND PIECE DID NOT WORK CORRECTLY. THE STAPLER WORKED FINE TWICE WHEN BEING USED ON THE LUNG TISSUE BUT DID NOT WORK ON THE THIRD TIME WHEN IT WAS USED IN VASCULAR TISSUE. THE SURGEON CHANGED THE RELOAD TWICE AND THEN CHANGED THE STAPLER. IT WAS REPORTED THERE WAS AN EXTENSION OF THE INCISION BY MORE THAN 1 INCH SINCE THORACOTOMY WAS NEEDED TO PERFORM SAVE AND RADICAL LOBECTOMY FOR A CENTRAL LUNG CANCER AND SURGERY TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
14-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AND LOOP ILEOSTOMY: CDH29A - STAPLES DID NOT PROPERLY FORM AFTER FIRING - DID NOT SEEM TO CUT THE TISSUE AND STAPLES HAD OPEN LEGS - NOT FORMED INTO B SHAPES. ANASTAMOSIS NOT FORMED - SEE PHOTO. WASHER DOES NOT APPEAR TO BE BROKEN. BOTH DEVICE AND ANVIL RETURNED. TISSUE NOT TOO THICK AND DEVICE FIRED WITH NO STOPPING OF MOVEMENT DURING FIRING PROCEDURE. DID NOT IDENTIFY A CRUNCH OR AUDIBLE SOUND WHEN FIRED. THE OPERATION TOOK ANOTHER 30 MINS TO COMPLETE AS HAD TO REDO THE ANASTOMOSIS - HAD TO OPEN ANOTHER CDH29 AND CONTOUR TO COMPLETE THE PROCEDURE.
10-12-2018 | Injury | us surgical puerto rico | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS, WHEN THE RELOAD WAS LOADED, THE OPENING AND CLOSING CHECK OF THE JAWS DID NOT GO WELL AT THE BEGINNING. BUT WHEN THE RELOAD WAS LOADED AGAIN, THERE WAS NO PROBLEM WITH THE OPENING AND CLOSING OF THE JAWS. IT WAS NOTED THAT WHEN THE DEVICE WAS SET ON THE TISSUE AND THE DOCTOR SQUEEZED THE HANDLE IN ORDER TO FIRE, THE DEVICE WAS SO STIFF THAT IT COULD NOT BE SQUEEZED. THE PROCEDURE WAS COMPLETED WITH MANUAL SUTURING AND THERE WAS NO TISSUE DAMAGE. IT WAS REPORTED THAT THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES.
10-12-2018 | Injury | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS, WHEN THE RELOAD WAS LOADED, THE OPENING AND CLOSING CHECK OF THE JAWS DID NOT GO WELL AT THE BEGINNING. BUT WHEN THE RELOAD WAS LOADED AGAIN, THERE WAS NO PROBLEM WITH THE OPENING AND CLOSING OF THE JAWS. IT WAS NOTED THAT WHEN THE DEVICE WAS SET ON THE TISSUE AND THE DOCTOR SQUEEZED THE HANDLE IN ORDER TO FIRE, THE DEVICE WAS SO STIFF THAT IT COULD NOT BE SQUEEZED. THE PROCEDURE WAS COMPLETED WITH MANUAL SUTURING AND THERE WAS NO TISSUE DAMAGE. IT WAS REPORTED THAT THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES.
12-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ILS 33MM, CURVED | Output above Specifications | A090404 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE FIRING FORCE WAS HIGHER THAN EXPECTED AT THE FIRING. THE DEVICE WAS NOT USED FOR THE PATIENT FURTHER. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ONE DEVICE WILL BE RETURNING. ALL THE STAPLES WERE FELL OUT FROM THE STAPLE HOUSING THOUGH THE CARTRIDGE WAS NOT USED. THERE WAS A POSSIBILITY THAT THE STAPLES FELL OUT WHEN THE CARTRIDGE WAS WASHED AFTER THE EVENT.
12-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ILS 33MM, CURVED | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE FIRING FORCE WAS HIGHER THAN EXPECTED AT THE FIRING. THE DEVICE WAS NOT USED FOR THE PATIENT FURTHER. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ONE DEVICE WILL BE RETURNING. ALL THE STAPLES WERE FELL OUT FROM THE STAPLE HOUSING THOUGH THE CARTRIDGE WAS NOT USED. THERE WAS A POSSIBILITY THAT THE STAPLES FELL OUT WHEN THE CARTRIDGE WAS WASHED AFTER THE EVENT.
12-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ILS 33MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE FIRING FORCE WAS HIGHER THAN EXPECTED AT THE FIRING. THE DEVICE WAS NOT USED FOR THE PATIENT FURTHER. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ONE DEVICE WILL BE RETURNING. ALL THE STAPLES WERE FELL OUT FROM THE STAPLE HOUSING THOUGH THE CARTRIDGE WAS NOT USED. THERE WAS A POSSIBILITY THAT THE STAPLES FELL OUT WHEN THE CARTRIDGE WAS WASHED AFTER THE EVENT.
01-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON REPORTED BLEEDING EVENT DURING LAP APPENDECTOMY. ALL OF THE INCIDENTS INVOLVED A PVS STAPLER. THE SURGEON ALSO REPORTED STAPLER LOCKOUTS THAT COULD NOT BE FIXED WITH THE USE OF REVERSE BUTTON. ALL OF THE CASES INVOLVED PEDIATRIC PATIENTS. CAUTERY AND PRESSURE WERE USED TO STOP BLEEDING.
01-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON REPORTED BLEEDING EVENT DURING LAP APPENDECTOMY. ALL OF THE INCIDENTS INVOLVED A PVS STAPLER. THE SURGEON ALSO REPORTED STAPLER LOCKOUTS THAT COULD NOT BE FIXED WITH THE USE OF REVERSE BUTTON. ALL OF THE CASES INVOLVED PEDIATRIC PATIENTS. CAUTERY AND PRESSURE WERE USED TO STOP BLEEDING.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | EEA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOIDS, DURING ENGAGEMENT OF THE STAPLER TO ANVIL, THE ANVIL COULD NOT BE TILTED AFTER FIRING AND DISENGAGED. ANOTHER STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | SIGNIA | Complete Loss of Power | A070801 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SETTING BY THE NURSE FOR THE FIRST FIRING IN A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY, WHEN TRYING TO CONNECT THE CARTRIDGE, IT WAS NOTICED THAT THE DISPLAY HAD DISAPPEARED, ALL OPERATIONS WERE IMPOSSIBLE, SO ANOTHER HANDLE WAS PREPARED. WHEN THE HANDLE WAS INSERTED INTO THE CHARGER, THE INDICATOR FLASHED, WHICH SHOWED THE HANDLE WAS BEING CHARGED NORMALLY. IT DID NOT START UP ALTHOUGH IT WAS REMOVED FROM THE CHARGER. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE.
17-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LAP COLON PROCEDURE, THE CIRCULAR STAPLER FIRED AND PRODUCED STAPLES BUT DID NOT CUT TISSUE DURING A LAP SIGMOID PROCEDURE. THE PROCEDURE WAS DELAYED BY 10 MINUTES; CUT TISSUE AND OPENED UP A SECOND STAPLER. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
09-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLON PROCEDURE, WHILE DISCONNECTING THE RIGHT COLON, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, IN ORDER TO RESECT THE BRONCHI DURING A LOBECTOMY, THE SURGEON SWITCHED THE HANDLE TO FIRING MODE AND PRESSED THE GREEN BUTTON BUT DEVICE WAS NOT ABLE TO FIRE. THEY REPLACED THE HANDLE AND RELOAD AND IT WAS ABLE TO FIRE WITHOUT PROBLEMS. THERE WAS NO PATIENT INJURY.
09-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLON PROCEDURE, WHILE DISCONNECTING THE RIGHT COLON, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT STAPLES DID NOT COME OUT AND THE BLADE DID NOT ADVANCE. IT WAS ALSO STATED THAT THE DEVICE HAD DIFFICULTY UNLOADING. ANOTHER RELOAD HAD TO BE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT STAPLES DID NOT COME OUT AND THE BLADE DID NOT ADVANCE. IT WAS ALSO STATED THAT THE DEVICE HAD DIFFICULTY UNLOADING. ANOTHER RELOAD HAD TO BE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT STAPLES DID NOT COME OUT AND THE BLADE DID NOT ADVANCE. IT WAS ALSO STATED THAT THE DEVICE HAD DIFFICULTY UNLOADING. ANOTHER RELOAD HAD TO BE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT STAPLES DID NOT COME OUT AND THE BLADE DID NOT ADVANCE. IT WAS ALSO STATED THAT THE DEVICE HAD DIFFICULTY UNLOADING. ANOTHER RELOAD HAD TO BE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT STAPLES DID NOT COME OUT AND THE BLADE DID NOT ADVANCE. IT WAS ALSO STATED THAT THE DEVICE HAD DIFFICULTY UNLOADING. ANOTHER RELOAD HAD TO BE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT STAPLES DID NOT COME OUT AND THE BLADE DID NOT ADVANCE. IT WAS ALSO STATED THAT THE DEVICE HAD DIFFICULTY UNLOADING. ANOTHER RELOAD HAD TO BE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: ADJUSTING KNOB ISSUE. IT WAS REPORTED THAT DURING THE PANCREATODUODENUM, DURING FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY, WHICH RESULT STAPLES MALFORMED AND TISSUE BLEEDING. SUTURE WAS USED TO OVERSEW AND STOP BLEEDING. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
07-12-2018 | Injury | us surgical puerto rico | 07-01-2019 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A WHIPPLE PROCEDURE, WHILE ON THE HEAD OF PANCREAS, THERE WAS A STAPLE LINE BLEEDING AFTER THE FIRING. STAPLE LINE SEEMED CORRECTLY FORMED. A REINFORCEMENT USING A SUTURE ALONG THE STAPLE LINE WAS USED TO SOLVE THE BLEEDING.
10-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, IN ORDER TO RESECT THE BRONCHI DURING A LOBECTOMY, THE SURGEON SWITCHED THE HANDLE TO FIRING MODE AND PRESSED THE GREEN BUTTON BUT DEVICE WAS NOT ABLE TO FIRE. THEY REPLACED THE HANDLE AND RELOAD AND IT WAS ABLE TO FIRE WITHOUT PROBLEMS. THERE WAS NO PATIENT INJURY.
11-12-2018 | Injury | ethicon endo-surgery | 07-01-2019 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES & OOZING. IT WAS REPORTED THAT DURING THE TOTAL GASTRECTOMY, AFTER FIRED, NOTED THE STAPLES MALFORMED, USED ANOTHER CDH25A TO DO ANASTOMOSIS, AFTER FIRED,THE STAPLE LINE AND CUT LINE ARE GOOD, BUT NOTED OOZING. SUTURE WAS USED TO OVERSEW AND STOP BLEEDING. THE SURGERY DELAYED ABOUT 40 MINUTES. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
11-12-2018 | Injury | ethicon endo-surgery | 07-01-2019 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED: MALFORMED STAPLES & OOZING. IT WAS REPORTED THAT DURING THE TOTAL GASTRECTOMY, AFTER FIRED, NOTED THE STAPLES MALFORMED, USED ANOTHER CDH25A TO DO ANASTOMOSIS, AFTER FIRED,THE STAPLE LINE AND CUT LINE ARE GOOD, BUT NOTED OOZING. SUTURE WAS USED TO OVERSEW AND STOP BLEEDING. THE SURGERY DELAYED ABOUT 40 MINUTES. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
27-11-2018 | Injury | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE RELOADS CUT ALONE AND THE BLADE OF THE RELOADS WAS SHOOTING ALONE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT COLECTOMY AT THE RIGHT COLON, THE DEVICE FIRED CORRECTLY AT THE FIRST RELOAD BUT WHEN A NEW RELOAD WAS USED, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A NEW DEVICE WAS OPENED USING THE SAME LOADING UNIT AND IT WORKED PROPERLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY, WHEN THE CARTRIDGE WAS ATTEMPTED TO BE ATTACHED TO THE HANDLE, IT WAS UNABLE TO BE CONNECTED, AND THE CARTRIDGE WAS REPLACED. ANOTHER CARTRIDGE WAS ATTACHED, AND IT WAS ABLE TO BE USED WITHOUT PROBLEM. THE REPORTED CARTRIDGE WAS PUT ON THE UNCLEAN FIELD, AND IT WAS ATTEMPTED TO BE ATTACHED TO THE MANUAL HANDLE, IT WAS UNABLE TO BE ATTACHED. BECAUSE THE PROXIMAL END OF THE JAWS SEEMED TO BE BENT. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. THERE WAS NO PATIENT HARM.
07-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Material Twisted/Bent | A040609 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY, WHEN THE CARTRIDGE WAS ATTEMPTED TO BE ATTACHED TO THE HANDLE, IT WAS UNABLE TO BE CONNECTED, AND THE CARTRIDGE WAS REPLACED. ANOTHER CARTRIDGE WAS ATTACHED, AND IT WAS ABLE TO BE USED WITHOUT PROBLEM. THE REPORTED CARTRIDGE WAS PUT ON THE UNCLEAN FIELD, AND IT WAS ATTEMPTED TO BE ATTACHED TO THE MANUAL HANDLE, IT WAS UNABLE TO BE ATTACHED. BECAUSE THE PROXIMAL END OF THE JAWS SEEMED TO BE BENT. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. THERE WAS NO PATIENT HARM.
01-01-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Battery Problem | A0705 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, AFTER THE 3RD SHOT, THE BATTERY WAS EMPTY AND THE STAPLER DID NOT FIRED AT ALL. THEY CHANGED THE INSTRUMENT AND IT WORKED AGAIN. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, AFTER THE 3RD SHOT, THE BATTERY WAS EMPTY AND THE STAPLER DID NOT FIRED AT ALL. THEY CHANGED THE INSTRUMENT AND IT WORKED AGAIN. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
06-05-2016 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE LOAD HAD NOT BEEN OPENED AND ANOTHER LOAD HAD TO BE USED. THERE WAS NO PATIENT INJURY.
06-05-2016 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE LOAD HAD NOT BEEN OPENED AND ANOTHER LOAD HAD TO BE USED. THERE WAS NO PATIENT INJURY.
06-05-2016 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE LOAD HAD NOT BEEN OPENED AND ANOTHER LOAD HAD TO BE USED. THERE WAS NO PATIENT INJURY.
06-05-2016 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE LOAD HAD NOT BEEN OPENED AND ANOTHER LOAD HAD TO BE USED. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT COLECTOMY AT THE RIGHT COLON, THE DEVICE FIRED CORRECTLY AT THE FIRST RELOAD, BUT WHEN A NEW RELOAD WAS USED, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. A NEW DEVICE WAS OPENED USING THE SAME LOADING UNIT AND IT WORKED PROPERLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY, WHILE CUTTING THE STOMACH, IT WAS IMPOSSIBLE TO STAPLE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-11-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY, WHILE CUTTING THE STOMACH, IT WAS IMPOSSIBLE TO STAPLE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2019 | Injury | covidien lp llc north haven | 07-01-2019 | ABSORBATACK | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: ACCORDING TO THE REPORTER FOLLOWING A RIGHT DIRECT INGUINAL PROCEDURE THE PATIENT FELT LIKE THERE WAS RAZOR BLADES IN HIS GROIN AND THAT THE PATIENT FELT LIKE A FLAMMABLE LIQUID LIT ON FIRE. THE PATIENT CURRENTLY SPENDS THE TIME BEDRIDDEN AND NEEDS A CANE TO GET AROUND.
07-01-2019 | Injury | covidien lp llc north haven | 07-01-2019 | ABSORBATACK | Adverse Event Without Identified Device or Use Problem | A24 | Disability | Event Description: ACCORDING TO THE REPORTER FOLLOWING A RIGHT DIRECT INGUINAL PROCEDURE THE PATIENT FELT LIKE THERE WAS RAZOR BLADES IN HIS GROIN AND THAT THE PATIENT FELT LIKE A FLAMMABLE LIQUID LIT ON FIRE. THE PATIENT CURRENTLY SPENDS THE TIME BEDRIDDEN AND NEEDS A CANE TO GET AROUND.
05-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, THE STANDARD STAPLER WITH BRAND-NEW PACKAGING WAS USED IN THE OPERATION. THE FIRST FIRING WAS NORMAL WHEN THE WHITE RELOAD WAS USED TO CLOSE AND CUT THE BLOOD VESSEL. WHEN THE SECOND FIRING WAS CLOSED WITH THE GREEN STRAIGHT RELOAD, THE POSITION OF THE BRONCHUS WAS ADJUSTED AFTER THE BRONCHUS WAS CLAMPED. THE BLACK GRIP WAS PUSHED FORWARD TO OPEN THE FORCEPS, AND THE FORCEPS WERE CLAMPED AGAIN AFTER THE ADJUSTMENT. IT COULDN'T BE COMPLETELY CLOSED, THERE WERE STILL LARGE OPENINGS; THE SAME SITUATION HAPPENED WHEN REPLACING A WITH BRAND-NEW 45MM GREEN RELOAD; BECAUSE THE 45MM GREEN RELOAD PREPARED IN THE OPERATING ROOM ON THE SAME DAY HAD BEEN USED UP, A BRAND-NEW 60MM GREEN STAPLE WAS USED. AFTER THAT, THE BRONCHUS WAS CLAMPED, THE FORCEPS WERE OPENED TO ADJUST THE POSITION AND CLAMPED AGAIN. THE BACK-NAIL ANVIL AND RELOAD COULDN'T BE CLOSED, AND THERE WERE STILL LARGE OPENINGS. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, THE STANDARD STAPLER WITH BRAND-NEW PACKAGING WAS USED IN THE OPERATION. THE FIRST FIRING WAS NORMAL WHEN THE WHITE RELOAD WAS USED TO CLOSE AND CUT THE BLOOD VESSEL. WHEN THE SECOND FIRING WAS CLOSED WITH THE GREEN STRAIGHT RELOAD, THE POSITION OF THE BRONCHUS WAS ADJUSTED AFTER THE BRONCHUS WAS CLAMPED. THE BLACK GRIP WAS PUSHED FORWARD TO OPEN THE FORCEPS, AND THE FORCEPS WERE CLAMPED AGAIN AFTER THE ADJUSTMENT. IT COULDN'T BE COMPLETELY CLOSED, THERE WERE STILL LARGE OPENINGS; THE SAME SITUATION HAPPENED WHEN REPLACING A WITH BRAND-NEW 45MM GREEN RELOAD; BECAUSE THE 45MM GREEN RELOAD PREPARED IN THE OPERATING ROOM ON THE SAME DAY HAD BEEN USED UP, A BRAND-NEW 60MM GREEN STAPLE WAS USED. AFTER THAT, THE BRONCHUS WAS CLAMPED, THE FORCEPS WERE OPENED TO ADJUST THE POSITION AND CLAMPED AGAIN. THE BACK-NAIL ANVIL AND RELOAD COULDN'T BE CLOSED, AND THERE WERE STILL LARGE OPENINGS. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
19-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE FIRST RELOAD THAT WAS USED DID NOT STAPLE. THE SURGEON THEN CHANGED TO ANOTHER RELOAD WHICH WAS FROM THE SAME LOT NUMBER AS THE FIRST, HOWEVER, IT ALSO DID NOT STAPLE. A THIRD RELOAD FROM ANOTHER LOT WAS OPENED AND USED COMPLETE THE CASE. IT WAS NOTED THAT THERE WAS A LONGER DURATION OF ANESTHESIA.
07-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING TRANSECTION OF THE LUNG TISSUE, THERE WAS A CRACK SOUND WHEN THE SURGEON FIRED THE RELOAD AND IT APPEARED THAT THE GEAR WITHIN THE STAPLER WAS BROKEN AND NEITHER THE STAPLES NOR THE KNIFE CAME OUT. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. IT WAS REPORTED THAT THE DEVICE WAS NOT ABLE TO FIRE. THEY USED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING TRANSECTION OF THE LUNG TISSUE, THERE WAS A CRACK SOUND WHEN THE SURGEON FIRED THE RELOAD AND IT APPEARED THAT THE GEAR WITHIN THE STAPLER WAS BROKEN AND NEITHER THE STAPLES NOR THE KNIFE CAME OUT. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. IT WAS REPORTED THAT THE DEVICE WAS NOT ABLE TO FIRE. THEY USED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING TRANSECTION OF THE LUNG TISSUE, THERE WAS A CRACK SOUND WHEN THE SURGEON FIRED THE RELOAD AND IT APPEARED THAT THE GEAR WITHIN THE STAPLER WAS BROKEN AND NEITHER THE STAPLES NOR THE KNIFE CAME OUT. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. IT WAS REPORTED THAT THE DEVICE WAS NOT ABLE TO FIRE. THEY USED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING TRANSECTION OF THE LUNG TISSUE, THERE WAS A CRACK SOUND WHEN THE SURGEON FIRED THE RELOAD AND IT APPEARED THAT THE GEAR WITHIN THE STAPLER WAS BROKEN AND NEITHER THE STAPLES NOR THE KNIFE CAME OUT. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. IT WAS REPORTED THAT THE DEVICE WAS NOT ABLE TO FIRE. THEY USED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING TRANSECTION OF THE LUNG TISSUE, THERE WAS A CRACK SOUND WHEN THE SURGEON FIRED THE RELOAD AND IT APPEARED THAT THE GEAR WITHIN THE STAPLER WAS BROKEN AND NEITHER THE STAPLES NOR THE KNIFE CAME OUT. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. IT WAS REPORTED THAT THE DEVICE WAS NOT ABLE TO FIRE. THEY USED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-11-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE FIRST RELOAD THAT WAS USED DID NOT STAPLE. THE SURGEON THEN CHANGED TO ANOTHER RELOAD WHICH WAS FROM THE SAME LOT NUMBER AS THE FIRST, HOWEVER, IT ALSO DID NOT STAPLE. A THIRD RELOAD FROM ANOTHER LOT WAS OPENED AND USED COMPLETE THE CASE. IT WAS NOTED THAT THERE WAS A LONGER DURATION OF ANESTHESIA.
12-11-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | BLUE RELOAD FOR ECHELON 60 | Loose or Intermittent Connection | A1205 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT COULD NOT FIRE FARTHER. DURING THE LPD SURGERY, COULD NOT FIRE FARTHER AFTER FIRED 1/3, THE FIRING TRIGGER WAS LOOSE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-11-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | BLUE RELOAD FOR ECHELON 60 | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT COULD NOT FIRE FARTHER. DURING THE LPD SURGERY, COULD NOT FIRE FARTHER AFTER FIRED 1/3, THE FIRING TRIGGER WAS LOOSE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
27-11-2018 | Injury | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE RELOADS CUT ALONE AND THE BLADE OF THE RELOADS WAS SHOOTING ALONE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
27-11-2018 | Injury | us surgical puerto rico | 07-01-2019 | ENDO GIA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE RELOADS CUT ALONE AND THE BLADE OF THE RELOADS WAS SHOOTING ALONE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
05-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PULMONARY RESECTION, THE STAPLER WAS NOT ABLE TO FIRE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PULMONARY RESECTION, THE STAPLER WAS NOT ABLE TO FIRE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2019 | Malfunction | us surgical puerto rico | 07-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
12-02-2019 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
07-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | TA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, WHEN THE SURGEON WAS TRYING TO STAPLE AN INTESTINAL TISSUE, THE DEVICE ONLY STAPLED HALF WAY. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | TA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, WHEN THE SURGEON WAS TRYING TO STAPLE AN INTESTINAL TISSUE, THE DEVICE ONLY STAPLED HALF WAY. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-11-2018 | Injury | us surgical puerto rico | 07-01-2019 | GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DOCTOR PERFORMED A TOTAL GASTRECTOMY AND FIRED THE STAPLERS ON PYLORUS. THE DOCTOR FOUND THE TISSUE TOO THICK AND CANNOT FORMED A COMPLETE STAPLE LINE PROXIMALLY. THEN, THE DOCTOR SWITCHED TO USE ANOTHER DEVICE AND STILL CANNOT COMPLETE THE STAPLE LINE. THEN, THE DOCTOR SWITCHED AGAIN TO USE ANOTHER DEVICE, BUT STILL, THE TISSUE WAS TOO THICK. THE SURGEON PERFORMED MANUAL SUTURE TO CLOSE THE WOUND THAT EXTENDS THE SURGICAL TIME FOR MORE THAN 30 MINUTES AND ADDITIONAL OPERATION TO CORRECT THE ISSUE. THE SURGERY WAS ORIGINALLY LAPAROSCOPIC PROCEDURE AND THEN CONVERTED TO OPEN PROCEDURE.
13-11-2018 | Injury | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DOCTOR PERFORMED A TOTAL GASTRECTOMY AND FIRED THE STAPLERS ON PYLORUS. THE DOCTOR FOUND THE TISSUE TOO THICK AND CANNOT FORMED A COMPLETE STAPLE LINE PROXIMALLY. THEN, THE DOCTOR SWITCHED TO USE ANOTHER DEVICE AND STILL CANNOT COMPLETE THE STAPLE LINE. THEN, THE DOCTOR SWITCHED AGAIN TO USE ANOTHER DEVICE, BUT STILL, THE TISSUE WAS TOO THICK. THE SURGEON PERFORMED MANUAL SUTURE TO CLOSE THE WOUND THAT EXTENDS THE SURGICAL TIME FOR MORE THAN 30 MINUTES AND ADDITIONAL OPERATION TO CORRECT THE ISSUE. THE SURGERY WAS ORIGINALLY LAPAROSCOPIC PROCEDURE AND THEN CONVERTED TO OPEN PROCEDURE.
13-11-2018 | Injury | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DOCTOR PERFORMED A TOTAL GASTRECTOMY AND FIRED THE STAPLERS ON PYLORUS. THE DOCTOR FOUND THE TISSUE TOO THICK AND CANNOT FORMED A COMPLETE STAPLE LINE PROXIMALLY. THEN, THE DOCTOR SWITCHED TO USE ANOTHER DEVICE AND STILL CANNOT COMPLETE THE STAPLE LINE. THEN, THE DOCTOR SWITCHED AGAIN TO USE ANOTHER DEVICE, BUT STILL, THE TISSUE WAS TOO THICK. THE SURGEON PERFORMED MANUAL SUTURE TO CLOSE THE WOUND THAT EXTENDS THE SURGICAL TIME FOR MORE THAN 30 MINUTES AND ADDITIONAL OPERATION TO CORRECT THE ISSUE. THE SURGERY WAS ORIGINALLY LAPAROSCOPIC PROCEDURE AND THEN CONVERTED TO OPEN PROCEDURE.
13-11-2018 | Injury | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DOCTOR PERFORMED A TOTAL GASTRECTOMY AND FIRED THE STAPLERS ON PYLORUS. THE DOCTOR FOUND THE TISSUE TOO THICK AND CANNOT FORMED A COMPLETE STAPLE LINE PROXIMALLY. THEN, THE DOCTOR SWITCHED TO USE ANOTHER DEVICE AND STILL CANNOT COMPLETE THE STAPLE LINE. THEN, THE DOCTOR SWITCHED AGAIN TO USE ANOTHER DEVICE, BUT STILL, THE TISSUE WAS TOO THICK. THE SURGEON PERFORMED MANUAL SUTURE TO CLOSE THE WOUND THAT EXTENDS THE SURGICAL TIME FOR MORE THAN 30 MINUTES AND ADDITIONAL OPERATION TO CORRECT THE ISSUE. THE SURGERY WAS ORIGINALLY LAPAROSCOPIC PROCEDURE AND THEN CONVERTED TO OPEN PROCEDURE.
13-11-2018 | Injury | covidien lp llc north haven | 07-01-2019 | IDRIVE ULTRA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DOCTOR PERFORMED A TOTAL GASTRECTOMY AND FIRED THE STAPLERS ON PYLORUS. THE DOCTOR FOUND THE TISSUE TOO THICK AND CANNOT FORMED A COMPLETE STAPLE LINE PROXIMALLY. THEN, THE DOCTOR SWITCHED TO USE ANOTHER DEVICE AND STILL CANNOT COMPLETE THE STAPLE LINE. THEN, THE DOCTOR SWITCHED AGAIN TO USE ANOTHER DEVICE, BUT STILL, THE TISSUE WAS TOO THICK. THE SURGEON PERFORMED MANUAL SUTURE TO CLOSE THE WOUND THAT EXTENDS THE SURGICAL TIME FOR MORE THAN 30 MINUTES AND ADDITIONAL OPERATION TO CORRECT THE ISSUE. THE SURGERY WAS ORIGINALLY LAPAROSCOPIC PROCEDURE AND THEN CONVERTED TO OPEN PROCEDURE.
13-11-2018 | Injury | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THE DOCTOR PERFORMED A TOTAL GASTRECTOMY AND FIRED THE STAPLERS ON PYLORUS. THE DOCTOR FOUND THE TISSUE TOO THICK AND CANNOT FORMED A COMPLETE STAPLE LINE PROXIMALLY. THEN, THE DOCTOR SWITCHED TO USE ANOTHER DEVICE AND STILL CANNOT COMPLETE THE STAPLE LINE. THEN, THE DOCTOR SWITCHED AGAIN TO USE ANOTHER DEVICE, BUT STILL, THE TISSUE WAS TOO THICK. THE SURGEON PERFORMED MANUAL SUTURE TO CLOSE THE WOUND THAT EXTENDS THE SURGICAL TIME FOR MORE THAN 30 MINUTES AND ADDITIONAL OPERATION TO CORRECT THE ISSUE. THE SURGERY WAS ORIGINALLY LAPAROSCOPIC PROCEDURE AND THEN CONVERTED TO OPEN PROCEDURE.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | RELIATACK | Component or Accessory Incompatibility | A1701 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VENTRAL HERNIA PROCEDURE, WHILE ON THE FIXATION OF THE HERNIA MESH, THE 1ST TACKER WORKED FINE UNTIL THE GEARS STARTED GRINDING AND SLIPPING TO THE POINT THAT THE TACKS WOULD NOT ADVANCE ANYMORE. A STRANGE NOISE WAS HEARD. THE 2ND TACKER WORKED UNTIL THE BLUE TOGGLE WOULDN'T OPERATE AND THEREFORE NOT ALLOW THE RELOAD TO BE LOADED. AFTER THE LOADING ISSUES WERE EXPERIENCED, THE RELOAD WAS NOT USED IN THE PATIENT. IT WAS ALSO NOTED THAT THERE WERE DIFFICULTIES WITH LOADING THE RELOAD ONTO THE HANDLE, DIFFICULTY IN PRESSING THE PUSH TO RELOAD BUTTON AND NAVIGATING THE LOCK AND UNLOCK BUTTON. THE CASE WAS FINISHED WITH A DIFFERENT DEVICE WITHOUT FURTHER INCIDENT. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | RELIATACK | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VENTRAL HERNIA PROCEDURE, WHILE ON THE FIXATION OF THE HERNIA MESH, THE 1ST TACKER WORKED FINE UNTIL THE GEARS STARTED GRINDING AND SLIPPING TO THE POINT THAT THE TACKS WOULD NOT ADVANCE ANYMORE. A STRANGE NOISE WAS HEARD. THE 2ND TACKER WORKED UNTIL THE BLUE TOGGLE WOULDN'T OPERATE AND THEREFORE NOT ALLOW THE RELOAD TO BE LOADED. AFTER THE LOADING ISSUES WERE EXPERIENCED, THE RELOAD WAS NOT USED IN THE PATIENT. IT WAS ALSO NOTED THAT THERE WERE DIFFICULTIES WITH LOADING THE RELOAD ONTO THE HANDLE, DIFFICULTY IN PRESSING THE PUSH TO RELOAD BUTTON AND NAVIGATING THE LOCK AND UNLOCK BUTTON. THE CASE WAS FINISHED WITH A DIFFERENT DEVICE WITHOUT FURTHER INCIDENT. THERE WAS NO PATIENT INJURY.
09-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT LOWER LOBE RESECTION, WHILE IN THE MIDDLE OF THE SLOTTED LOWER LOBE TRACHEA, ONLY 1/3 OF THE RELOAD WAS ABLE TO ADVANCE WHEN THE HANDLE WAS FIRED. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE CARTRIDGE BURST BROKEN. IT WAS ALSO STATED THAT THERE WERE POOR STAPLE FORMATION. WHEN THE DEVICE WAS REMOVED, THEY WERE UNABLE TO SEPARATE THE RELOAD AND THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT LOWER LOBE RESECTION, WHILE IN THE MIDDLE OF THE SLOTTED LOWER LOBE TRACHEA, ONLY 1/3 OF THE RELOAD WAS ABLE TO ADVANCE WHEN THE HANDLE WAS FIRED. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE CARTRIDGE BURST BROKEN. IT WAS ALSO STATED THAT THERE WERE POOR STAPLE FORMATION. WHEN THE DEVICE WAS REMOVED, THEY WERE UNABLE TO SEPARATE THE RELOAD AND THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-11-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT LOWER LOBE RESECTION, WHILE IN THE MIDDLE OF THE SLOTTED LOWER LOBE TRACHEA, ONLY 1/3 OF THE RELOAD WAS ABLE TO ADVANCE WHEN THE HANDLE WAS FIRED. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE CARTRIDGE BURST BROKEN. IT WAS ALSO STATED THAT THERE WERE POOR STAPLE FORMATION. WHEN THE DEVICE WAS REMOVED, THEY WERE UNABLE TO SEPARATE THE RELOAD AND THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-11-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT LOWER LOBE RESECTION, WHILE IN THE MIDDLE OF THE SLOTTED LOWER LOBE TRACHEA, ONLY 1/3 OF THE RELOAD WAS ABLE TO ADVANCE WHEN THE HANDLE WAS FIRED. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE CARTRIDGE BURST BROKEN. IT WAS ALSO STATED THAT THERE WERE POOR STAPLE FORMATION. WHEN THE DEVICE WAS REMOVED, THEY WERE UNABLE TO SEPARATE THE RELOAD AND THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-11-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT LOWER LOBE RESECTION, WHILE IN THE MIDDLE OF THE SLOTTED LOWER LOBE TRACHEA, ONLY 1/3 OF THE RELOAD WAS ABLE TO ADVANCE WHEN THE HANDLE WAS FIRED. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE CARTRIDGE BURST BROKEN. IT WAS ALSO STATED THAT THERE WERE POOR STAPLE FORMATION. WHEN THE DEVICE WAS REMOVED, THEY WERE UNABLE TO SEPARATE THE RELOAD AND THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-11-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT LOWER LOBE RESECTION, WHILE IN THE MIDDLE OF THE SLOTTED LOWER LOBE TRACHEA, ONLY 1/3 OF THE RELOAD WAS ABLE TO ADVANCE WHEN THE HANDLE WAS FIRED. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE CARTRIDGE BURST BROKEN. IT WAS ALSO STATED THAT THERE WERE POOR STAPLE FORMATION. WHEN THE DEVICE WAS REMOVED, THEY WERE UNABLE TO SEPARATE THE RELOAD AND THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT FIRST FIRING ATTEMPT ON STOMACH TISSUE RESECTION DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE HAD STARTED FLASHING IN FIRING MODE WHEN THE SURGEON PRESSED THE GREEN FIRING BUTTON, BUT WHEN THE SURGEON PRESSED DOWN, THE DEVICE DID NOT FIRE. THEY TRIED TO PRESS DOWN AGAIN, BUT IT HAD NO SIGN OF FIRING OCCURRED. THEY THEN PUSHED UP ON DEVICE TOGGLE TO OPEN THE JAWS AND THE DISPLAY GAVE A WHITE BOX WITH ZERO LIVES ON THE RELOAD INDICATOR SCREEN, BUT THE RELOAD ITSELF SHOWED NO INDICATIONS OF STAPLES FIRING AT ALL. THEY OPENED AND USED A NEW RELOAD AND LOADED IT, AND IT FIRED PROPERLY ON THE SAME HANDLE AND ADAPTER, AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT FIRST FIRING ATTEMPT ON STOMACH TISSUE RESECTION DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE HAD STARTED FLASHING IN FIRING MODE WHEN THE SURGEON PRESSED THE GREEN FIRING BUTTON, BUT WHEN THE SURGEON PRESSED DOWN, THE DEVICE DID NOT FIRE. THEY TRIED TO PRESS DOWN AGAIN, BUT IT HAD NO SIGN OF FIRING OCCURRED. THEY THEN PUSHED UP ON DEVICE TOGGLE TO OPEN THE JAWS AND THE DISPLAY GAVE A WHITE BOX WITH ZERO LIVES ON THE RELOAD INDICATOR SCREEN, BUT THE RELOAD ITSELF SHOWED NO INDICATIONS OF STAPLES FIRING AT ALL. THEY OPENED AND USED A NEW RELOAD AND LOADED IT, AND IT FIRED PROPERLY ON THE SAME HANDLE AND ADAPTER, AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | SIGNIA | Failure to Charge | A070603 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/ LOBECTOMY, AT THE 6TH FIRING, WHEN THE SURGEON ATTEMPTED TO RESECT THE AZYGOS VEIN WITH THE RELOAD, THE SCREEN GOT DARK SUDDENLY AND THE SURGEON WAS NOT ABLE TO OPERATE THE DEVICE. A NEW HANDLE AND THE OPERATION WAS CONTINUED. WHEN THE REPORTED HANDLE WAS REMOVED FROM THE SHELL, IT WAS HOT. WHEN THE REPORTED HANDLE WAS COLLECTED THE HANDLE WAS WITH HEAT. THE HANDLE WAS INSERTED INTO A CHARGER, BUT THERE WAS NO RESPONSE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | SIGNIA | Overheating of Device | A1005 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/ LOBECTOMY, AT THE 6TH FIRING, WHEN THE SURGEON ATTEMPTED TO RESECT THE AZYGOS VEIN WITH THE RELOAD, THE SCREEN GOT DARK SUDDENLY AND THE SURGEON WAS NOT ABLE TO OPERATE THE DEVICE. A NEW HANDLE AND THE OPERATION WAS CONTINUED. WHEN THE REPORTED HANDLE WAS REMOVED FROM THE SHELL, IT WAS HOT. WHEN THE REPORTED HANDLE WAS COLLECTED THE HANDLE WAS WITH HEAT. THE HANDLE WAS INSERTED INTO A CHARGER, BUT THERE WAS NO RESPONSE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | SIGNIA | Complete Loss of Power | A070801 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/ LOBECTOMY, AT THE 6TH FIRING, WHEN THE SURGEON ATTEMPTED TO RESECT THE AZYGOS VEIN WITH THE RELOAD, THE SCREEN GOT DARK SUDDENLY AND THE SURGEON WAS NOT ABLE TO OPERATE THE DEVICE. A NEW HANDLE AND THE OPERATION WAS CONTINUED. WHEN THE REPORTED HANDLE WAS REMOVED FROM THE SHELL, IT WAS HOT. WHEN THE REPORTED HANDLE WAS COLLECTED THE HANDLE WAS WITH HEAT. THE HANDLE WAS INSERTED INTO A CHARGER, BUT THERE WAS NO RESPONSE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, THE STANDARD STAPLER WITH BRAND-NEW PACKAGING WAS USED IN THE OPERATION. THE FIRST FIRING WAS NORMAL WHEN THE WHITE RELOAD WAS USED TO CLOSE AND CUT THE BLOOD VESSEL. WHEN THE SECOND FIRING WAS CLOSED WITH THE GREEN STRAIGHT RELOAD, THE POSITION OF THE BRONCHUS WAS ADJUSTED AFTER THE BRONCHUS WAS CLAMPED. THE BLACK GRIP WAS PUSHED FORWARD TO OPEN THE FORCEPS, AND THE FORCEPS WERE CLAMPED AGAIN AFTER THE ADJUSTMENT. IT COULDN'T BE COMPLETELY CLOSED, THERE WERE STILL LARGE OPENINGS; THE SAME SITUATION HAPPENED WHEN REPLACING A WITH BRAND-NEW 45MM GREEN RELOAD; BECAUSE THE 45MM GREEN RELOAD PREPARED IN THE OPERATING ROOM ON THE SAME DAY HAD BEEN USED UP, A BRAND-NEW 60MM GREEN STAPLE WAS USED. AFTER THAT, THE BRONCHUS WAS CLAMPED, THE FORCEPS WERE OPENED TO ADJUST THE POSITION AND CLAMPED AGAIN. THE BACK NAIL ANVIL AND RELOAD COULDN'T BE CLOSED, AND THERE WERE STILL LARGE OPENINGS. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, THE STANDARD STAPLER WITH BRAND-NEW PACKAGING WAS USED IN THE OPERATION. THE FIRST FIRING WAS NORMAL WHEN THE WHITE RELOAD WAS USED TO CLOSE AND CUT THE BLOOD VESSEL. WHEN THE SECOND FIRING WAS CLOSED WITH THE GREEN STRAIGHT RELOAD, THE POSITION OF THE BRONCHUS WAS ADJUSTED AFTER THE BRONCHUS WAS CLAMPED. THE BLACK GRIP WAS PUSHED FORWARD TO OPEN THE FORCEPS, AND THE FORCEPS WERE CLAMPED AGAIN AFTER THE ADJUSTMENT. IT COULDN'T BE COMPLETELY CLOSED, THERE WERE STILL LARGE OPENINGS; THE SAME SITUATION HAPPENED WHEN REPLACING A WITH BRAND-NEW 45MM GREEN RELOAD; BECAUSE THE 45MM GREEN RELOAD PREPARED IN THE OPERATING ROOM ON THE SAME DAY HAD BEEN USED UP, A BRAND-NEW 60MM GREEN STAPLE WAS USED. AFTER THAT, THE BRONCHUS WAS CLAMPED, THE FORCEPS WERE OPENED TO ADJUST THE POSITION AND CLAMPED AGAIN. THE BACK NAIL ANVIL AND RELOAD COULDN'T BE CLOSED, AND THERE WERE STILL LARGE OPENINGS. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, THE STANDARD STAPLER WITH BRAND-NEW PACKAGING WAS USED IN THE OPERATION. THE FIRST FIRING WAS NORMAL WHEN THE WHITE RELOAD WAS USED TO CLOSE AND CUT THE BLOOD VESSEL. WHEN THE SECOND FIRING WAS CLOSED WITH THE GREEN STRAIGHT RELOAD, THE POSITION OF THE BRONCHUS WAS ADJUSTED AFTER THE BRONCHUS WAS CLAMPED. THE BLACK GRIP WAS PUSHED FORWARD TO OPEN THE FORCEPS, AND THE FORCEPS WERE CLAMPED AGAIN AFTER THE ADJUSTMENT. IT COULDN'T BE COMPLETELY CLOSED, THERE WERE STILL LARGE OPENINGS; THE SAME SITUATION HAPPENED WHEN REPLACING A WITH BRAND-NEW 45MM GREEN RELOAD; BECAUSE THE 45MM GREEN RELOAD PREPARED IN THE OPERATING ROOM ON THE SAME DAY HAD BEEN USED UP, A BRAND-NEW 60MM GREEN STAPLE WAS USED. AFTER THAT, THE BRONCHUS WAS CLAMPED, THE FORCEPS WERE OPENED TO ADJUST THE POSITION AND CLAMPED AGAIN. THE BACK NAIL ANVIL AND RELOAD COULDN'T BE CLOSED, AND THERE WERE STILL LARGE OPENINGS. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, THE STANDARD STAPLER WITH BRAND-NEW PACKAGING WAS USED IN THE OPERATION. THE FIRST FIRING WAS NORMAL WHEN THE WHITE RELOAD WAS USED TO CLOSE AND CUT THE BLOOD VESSEL. WHEN THE SECOND FIRING WAS CLOSED WITH THE GREEN STRAIGHT RELOAD, THE POSITION OF THE BRONCHUS WAS ADJUSTED AFTER THE BRONCHUS WAS CLAMPED. THE BLACK GRIP WAS PUSHED FORWARD TO OPEN THE FORCEPS, AND THE FORCEPS WERE CLAMPED AGAIN AFTER THE ADJUSTMENT. IT COULDN'T BE COMPLETELY CLOSED, THERE WERE STILL LARGE OPENINGS; THE SAME SITUATION HAPPENED WHEN REPLACING A WITH BRAND-NEW 45MM GREEN RELOAD; BECAUSE THE 45MM GREEN RELOAD PREPARED IN THE OPERATING ROOM ON THE SAME DAY HAD BEEN USED UP, A BRAND-NEW 60MM GREEN STAPLE WAS USED. AFTER THAT, THE BRONCHUS WAS CLAMPED, THE FORCEPS WERE OPENED TO ADJUST THE POSITION AND CLAMPED AGAIN. THE BACK NAIL ANVIL AND RELOAD COULDN'T BE CLOSED, AND THERE WERE STILL LARGE OPENINGS. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, THE STANDARD STAPLER WITH BRAND-NEW PACKAGING WAS USED IN THE OPERATION. THE FIRST FIRING WAS NORMAL WHEN THE WHITE RELOAD WAS USED TO CLOSE AND CUT THE BLOOD VESSEL. WHEN THE SECOND FIRING WAS CLOSED WITH THE GREEN STRAIGHT RELOAD, THE POSITION OF THE BRONCHUS WAS ADJUSTED AFTER THE BRONCHUS WAS CLAMPED. THE BLACK GRIP WAS PUSHED FORWARD TO OPEN THE FORCEPS, AND THE FORCEPS WERE CLAMPED AGAIN AFTER THE ADJUSTMENT. IT COULDN'T BE COMPLETELY CLOSED, THERE WERE STILL LARGE OPENINGS; THE SAME SITUATION HAPPENED WHEN REPLACING A WITH BRAND-NEW 45MM GREEN RELOAD; BECAUSE THE 45MM GREEN RELOAD PREPARED IN THE OPERATING ROOM ON THE SAME DAY HAD BEEN USED UP, A BRAND-NEW 60MM GREEN STAPLE WAS USED. AFTER THAT, THE BRONCHUS WAS CLAMPED, THE FORCEPS WERE OPENED TO ADJUST THE POSITION AND CLAMPED AGAIN. THE BACK NAIL ANVIL AND RELOAD COULDN'T BE CLOSED, AND THERE WERE STILL LARGE OPENINGS. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Injury | covidien lp llc north haven | 07-01-2019 | EEA ORVIL | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS INSERTED TRANSORALLY THROUGH THE MOUTH DURING A LAPAROSCOPIC GASTRIC BYPASS, THE DEVICE WAS SEATED IN POUCH, BUT WHEN IT WAS TIME TO RELEASE THE SUTURE, SUTURE WOULD NOT RELEASE AND FORCE WAS NEEDED TO RELEASE SUTURE. IT WAS STATED THAT THIS RESULTED TO A RIP IN THE STAPLE LINE AND LEAK IN THE ANASTOMOSIS NEEDING REPAIR. THE SURGEON SUTURED THE STAPLE LINE.
10-12-2018 | Injury | covidien lp llc north haven | 07-01-2019 | EEA ORVIL | Difficult to Remove | A150207 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS INSERTED TRANSORALLY THROUGH THE MOUTH DURING A LAPAROSCOPIC GASTRIC BYPASS, THE DEVICE WAS SEATED IN POUCH, BUT WHEN IT WAS TIME TO RELEASE THE SUTURE, SUTURE WOULD NOT RELEASE AND FORCE WAS NEEDED TO RELEASE SUTURE. IT WAS STATED THAT THIS RESULTED TO A RIP IN THE STAPLE LINE AND LEAK IN THE ANASTOMOSIS NEEDING REPAIR. THE SURGEON SUTURED THE STAPLE LINE.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | SIGNIA | No Display/Image | A090206 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, ON THE FIRST FIRING ACROSS THE STOMACH, END, THE DEVICE FIRED FINE ON FIRST RELOAD, BUT WHEN RETRACTING BLADE, IT STOPPED HALF WAY, THE SCREEN WENT BLANK, THEN VERY SLOWLY RETRACTED BLADE REST OF WAY TO PROXIMAL THE JAWS OPENED OKAY. THE SURGEON NOTICED CHANGE IN NOISE ON THE SECOND HALF OF RETRACTION. THE JAWS COULD NOT BE ARTICULATED BACK TO THE CENTER POSITION. TO REMOVE ARTICULATED RELOAD HE TOOK OUT THE TROCAR AND GAVE ADAPTER AND TROCAR BACK TO NURSE TO REMOVE. NURSE EVENTUALLY GOT THE HANDLE TO POWER UP AND THEN SEEMED TO BE WORKING FINE AS SCREEN CAME BACK ON. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | SIGNIA | Retraction Problem | A0510 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, ON THE FIRST FIRING ACROSS THE STOMACH, END, THE DEVICE FIRED FINE ON FIRST RELOAD, BUT WHEN RETRACTING BLADE, IT STOPPED HALF WAY, THE SCREEN WENT BLANK, THEN VERY SLOWLY RETRACTED BLADE REST OF WAY TO PROXIMAL THE JAWS OPENED OKAY. THE SURGEON NOTICED CHANGE IN NOISE ON THE SECOND HALF OF RETRACTION. THE JAWS COULD NOT BE ARTICULATED BACK TO THE CENTER POSITION. TO REMOVE ARTICULATED RELOAD HE TOOK OUT THE TROCAR AND GAVE ADAPTER AND TROCAR BACK TO NURSE TO REMOVE. NURSE EVENTUALLY GOT THE HANDLE TO POWER UP AND THEN SEEMED TO BE WORKING FINE AS SCREEN CAME BACK ON. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | SIGNIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, ON THE FIRST FIRING ACROSS THE STOMACH, END, THE DEVICE FIRED FINE ON FIRST RELOAD, BUT WHEN RETRACTING BLADE, IT STOPPED HALF WAY, THE SCREEN WENT BLANK, THEN VERY SLOWLY RETRACTED BLADE REST OF WAY TO PROXIMAL THE JAWS OPENED OKAY. THE SURGEON NOTICED CHANGE IN NOISE ON THE SECOND HALF OF RETRACTION. THE JAWS COULD NOT BE ARTICULATED BACK TO THE CENTER POSITION. TO REMOVE ARTICULATED RELOAD HE TOOK OUT THE TROCAR AND GAVE ADAPTER AND TROCAR BACK TO NURSE TO REMOVE. NURSE EVENTUALLY GOT THE HANDLE TO POWER UP AND THEN SEEMED TO BE WORKING FINE AS SCREEN CAME BACK ON. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | SIGNIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, ON THE FIRST FIRING ACROSS THE STOMACH, END, THE DEVICE FIRED FINE ON FIRST RELOAD, BUT WHEN RETRACTING BLADE, IT STOPPED HALF WAY, THE SCREEN WENT BLANK, THEN VERY SLOWLY RETRACTED BLADE REST OF WAY TO PROXIMAL THE JAWS OPENED OKAY. THE SURGEON NOTICED CHANGE IN NOISE ON THE SECOND HALF OF RETRACTION. THE JAWS COULD NOT BE ARTICULATED BACK TO THE CENTER POSITION. TO REMOVE ARTICULATED RELOAD HE TOOK OUT THE TROCAR AND GAVE ADAPTER AND TROCAR BACK TO NURSE TO REMOVE. NURSE EVENTUALLY GOT THE HANDLE TO POWER UP AND THEN SEEMED TO BE WORKING FINE AS SCREEN CAME BACK ON. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN EGIA SULU | Retraction Problem | A0510 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, ON THE FIRST FIRING ACROSS THE STOMACH, END, THE DEVICE FIRED FINE ON FIRST RELOAD, BUT WHEN RETRACTING BLADE, IT STOPPED HALF WAY, THE SCREEN WENT BLANK, THEN VERY SLOWLY RETRACTED BLADE REST OF WAY TO PROXIMAL THE JAWS OPENED OKAY. THE SURGEON NOTICED CHANGE IN NOISE ON THE SECOND HALF OF RETRACTION. THE JAWS COULD NOT BE ARTICULATED BACK TO THE CENTER POSITION. TO REMOVE ARTICULATED RELOAD HE TOOK OUT THE TROCAR AND GAVE ADAPTER AND TROCAR BACK TO NURSE TO REMOVE. NURSE EVENTUALLY GOT THE HANDLE TO POWER UP AND THEN SEEMED TO BE WORKING FINE AS SCREEN CAME BACK ON. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN EGIA SULU | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, ON THE FIRST FIRING ACROSS THE STOMACH, END, THE DEVICE FIRED FINE ON FIRST RELOAD, BUT WHEN RETRACTING BLADE, IT STOPPED HALF WAY, THE SCREEN WENT BLANK, THEN VERY SLOWLY RETRACTED BLADE REST OF WAY TO PROXIMAL THE JAWS OPENED OKAY. THE SURGEON NOTICED CHANGE IN NOISE ON THE SECOND HALF OF RETRACTION. THE JAWS COULD NOT BE ARTICULATED BACK TO THE CENTER POSITION. TO REMOVE ARTICULATED RELOAD HE TOOK OUT THE TROCAR AND GAVE ADAPTER AND TROCAR BACK TO NURSE TO REMOVE. NURSE EVENTUALLY GOT THE HANDLE TO POWER UP AND THEN SEEMED TO BE WORKING FINE AS SCREEN CAME BACK ON. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN SIGNIA EGIA ADAPTER | Retraction Problem | A0510 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, ON THE FIRST FIRING ACROSS THE STOMACH, END, THE DEVICE FIRED FINE ON FIRST RELOAD, BUT WHEN RETRACTING BLADE, IT STOPPED HALF WAY, THE SCREEN WENT BLANK, THEN VERY SLOWLY RETRACTED BLADE REST OF WAY TO PROXIMAL THE JAWS OPENED OKAY. THE SURGEON NOTICED CHANGE IN NOISE ON THE SECOND HALF OF RETRACTION. THE JAWS COULD NOT BE ARTICULATED BACK TO THE CENTER POSITION. TO REMOVE ARTICULATED RELOAD HE TOOK OUT THE TROCAR AND GAVE ADAPTER AND TROCAR BACK TO NURSE TO REMOVE. NURSE EVENTUALLY GOT THE HANDLE TO POWER UP AND THEN SEEMED TO BE WORKING FINE AS SCREEN CAME BACK ON. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN SIGNIA EGIA ADAPTER | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, ON THE FIRST FIRING ACROSS THE STOMACH, END, THE DEVICE FIRED FINE ON FIRST RELOAD, BUT WHEN RETRACTING BLADE, IT STOPPED HALF WAY, THE SCREEN WENT BLANK, THEN VERY SLOWLY RETRACTED BLADE REST OF WAY TO PROXIMAL THE JAWS OPENED OKAY. THE SURGEON NOTICED CHANGE IN NOISE ON THE SECOND HALF OF RETRACTION. THE JAWS COULD NOT BE ARTICULATED BACK TO THE CENTER POSITION. TO REMOVE ARTICULATED RELOAD HE TOOK OUT THE TROCAR AND GAVE ADAPTER AND TROCAR BACK TO NURSE TO REMOVE. NURSE EVENTUALLY GOT THE HANDLE TO POWER UP AND THEN SEEMED TO BE WORKING FINE AS SCREEN CAME BACK ON. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
25-02-2019 | Malfunction | us surgical puerto rico | 07-01-2019 | GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE BROKE.
11-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE RELOAD DID NOT FIRE WHEN THE SURGEON PRESSED THE FIRE BUTTON. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE RELOAD DID NOT FIRE WHEN THE SURGEON PRESSED THE FIRE BUTTON. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY, WHEN CUTTING THE LEFT LOWER LOBE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE, THEREFORE THE DEVICE WAS NOT ABLE TO FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY, WHEN CUTTING THE LEFT LOWER LOBE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE, THEREFORE THE DEVICE WAS NOT ABLE TO FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
11-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN EGIA ADAPTER | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE RELOAD DID NOT FIRE WHEN THE SURGEON PRESSED THE FIRE BUTTON. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCTECTOMY, AT THE FIRST FIRING , A STRANGE SOUND WAS HEARD FROM THE HANDLE WHEN THE DEVICE SLIGHTLY ADVANCED. THE DEVICE WAS CHECKED AND NOTED THAT THE FIRING WAS PERFORMED TO THE HALFWAY, BUT IT WAS NOT PERFORMED ON THE REST. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCTECTOMY, AT THE FIRST FIRING , A STRANGE SOUND WAS HEARD FROM THE HANDLE WHEN THE DEVICE SLIGHTLY ADVANCED. THE DEVICE WAS CHECKED AND NOTED THAT THE FIRING WAS PERFORMED TO THE HALFWAY, BUT IT WAS NOT PERFORMED ON THE REST. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
18-10-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A BOWEL SURGERY, THE SURGEON APPLIED THE DEVICE TO THE BOWEL TISSUE, BUT WAS NOT ABLE TO FIRE THE DEVICE, THE DEVICE LOCKED ON TISSUE, BUT WAS REMOVED NORMALLY WITHOUT RESECTION, A NEW RELOAD WAS APPLIED BUT DID NOT RESOLVE THE FIRING ISSUE. A NEW DEVICE WAS OPENED AND WAS ABLE TO USE WITHOUT ANY DIFFICULTY. THERE WAS NO PATIENT INJURY.
18-10-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A BOWEL SURGERY, THE SURGEON APPLIED THE DEVICE TO THE BOWEL TISSUE, BUT WAS NOT ABLE TO FIRE THE DEVICE, THE DEVICE LOCKED ON TISSUE, BUT WAS REMOVED NORMALLY WITHOUT RESECTION, A NEW RELOAD WAS APPLIED BUT DID NOT RESOLVE THE FIRING ISSUE. A NEW DEVICE WAS OPENED AND WAS ABLE TO USE WITHOUT ANY DIFFICULTY. THERE WAS NO PATIENT INJURY.
18-10-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A BOWEL SURGERY, THE SURGEON APPLIED THE DEVICE TO THE BOWEL TISSUE, BUT WAS NOT ABLE TO FIRE THE DEVICE, THE DEVICE LOCKED ON TISSUE, BUT WAS REMOVED NORMALLY WITHOUT RESECTION, A NEW RELOAD WAS APPLIED BUT DID NOT RESOLVE THE FIRING ISSUE. A NEW DEVICE WAS OPENED AND WAS ABLE TO USE WITHOUT ANY DIFFICULTY. THERE WAS NO PATIENT INJURY.
18-10-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | UNKNOWN EGIA SULU | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A BOWEL SURGERY, THE SURGEON APPLIED THE DEVICE TO THE BOWEL TISSUE, BUT WAS NOT ABLE TO FIRE THE DEVICE, THE DEVICE LOCKED ON TISSUE, BUT WAS REMOVED NORMALLY WITHOUT RESECTION, A NEW RELOAD WAS APPLIED BUT DID NOT RESOLVE THE FIRING ISSUE. A NEW DEVICE WAS OPENED AND WAS ABLE TO USE WITHOUT ANY DIFFICULTY. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ENDO ILS 29MM, CURVED | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, IN THE HOSPITAL, THERE ARE NO COLORECTAL FELLOW OR ANY AVAILABLE ASSISTANT FOR THE SURGEON. SO, A RESIDENT FIRED ECS29A. THE SURGEON ROUTINELY PREFERS CDH29A. HE ISN'T SURE WHY THE SCRUB NURSE OPENED THE ECS29A FOR THE CASE. AFTER FIRING, HE ROUTINELY DOES REFORCE SUTURING WITH V-LOC. WHILE HE WAS SUTURING FOR THIS CASE, HE FOUND OUT THE OTHER PART OF THE LUMEN WAS FULLY OPENED. THAT'S WHEN HE KNEW ONLY THE KNIFE WAS FIRED, WITHOUT THE STAPLES ON THE LUMEN. HE DISSECTED LOWER PART OF THE RECTUM AND RE-FIRED ANOTHER CDH29A TO CLOSE THE CASE. HE ISN'T SURE WHETHER IT HAPPENED FOR THE DEFECT OF THE DEVICE OR THE TECHNICAL ISSUES REGARDING FAULTS OR MISTAKES OF THE RESIDENT.
10-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY ROUX-EN-Y PROCEDURE, THE 45MM RELOAD WAS USED DURING DUODENECTOMY. THE NURSE SAID AFTER OPENING/CLOSING CHECK, HE/SHE HANDED THE HANDLE TO THE SURGEON. THE SURGEON SET THE JAW ON THE TISSUE AND PRESSED THE GREEN BUTTON IN ORDER TO FIRE. WHEN SQUEEZING THE HANDLE, A RASPING SOUND WAS GENERATED AND THEN THE SURGEON RELEASED THE CARTRIDGE. WHEN HE/SHE FIRED AGAIN IN THE SAME WAY, IT WAS ABLE TO STAPLE WITHOUT PROBLEM. THE 45MM RELOAD WAS ALSO USED IN ESOPHAGECTOMY. THIS TIME THE SURGEON CHECKED THE ACCURACY AND WHEN FIRING IN THE SAME WAY, THERE WAS ABNORMAL NOISE AND THE CARTRIDGE ADVANCED A LITTLE POSSIBLY OR NOT AT ALL. THE CARTRIDGE DID NOT ADVANCE LAST TIME BUT THIS TIME IT WAS NOT ABLE TO BE RELEASED. THE SURGEON USED ESOPHAGEAL FORCEPS AND RESECTED THE TISSUE WITH A SCALPEL. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY ROUX-EN-Y PROCEDURE, THE 45MM RELOAD WAS USED DURING DUODENECTOMY. THE NURSE SAID AFTER OPENING/CLOSING CHECK, HE/SHE HANDED THE HANDLE TO THE SURGEON. THE SURGEON SET THE JAW ON THE TISSUE AND PRESSED THE GREEN BUTTON IN ORDER TO FIRE. WHEN SQUEEZING THE HANDLE, A RASPING SOUND WAS GENERATED AND THEN THE SURGEON RELEASED THE CARTRIDGE. WHEN HE/SHE FIRED AGAIN IN THE SAME WAY, IT WAS ABLE TO STAPLE WITHOUT PROBLEM. THE 45MM RELOAD WAS ALSO USED IN ESOPHAGECTOMY. THIS TIME THE SURGEON CHECKED THE ACCURACY AND WHEN FIRING IN THE SAME WAY, THERE WAS ABNORMAL NOISE AND THE CARTRIDGE ADVANCED A LITTLE POSSIBLY OR NOT AT ALL. THE CARTRIDGE DID NOT ADVANCE LAST TIME BUT THIS TIME IT WAS NOT ABLE TO BE RELEASED. THE SURGEON USED ESOPHAGEAL FORCEPS AND RESECTED THE TISSUE WITH A SCALPEL. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY ROUX-EN-Y PROCEDURE, THE 45MM RELOAD WAS USED DURING DUODENECTOMY. THE NURSE SAID AFTER OPENING/CLOSING CHECK, HE/SHE HANDED THE HANDLE TO THE SURGEON. THE SURGEON SET THE JAW ON THE TISSUE AND PRESSED THE GREEN BUTTON IN ORDER TO FIRE. WHEN SQUEEZING THE HANDLE, A RASPING SOUND WAS GENERATED AND THEN THE SURGEON RELEASED THE CARTRIDGE. WHEN HE/SHE FIRED AGAIN IN THE SAME WAY, IT WAS ABLE TO STAPLE WITHOUT PROBLEM. THE 45MM RELOAD WAS ALSO USED IN ESOPHAGECTOMY. THIS TIME THE SURGEON CHECKED THE ACCURACY AND WHEN FIRING IN THE SAME WAY, THERE WAS ABNORMAL NOISE AND THE CARTRIDGE ADVANCED A LITTLE POSSIBLY OR NOT AT ALL. THE CARTRIDGE DID NOT ADVANCE LAST TIME BUT THIS TIME IT WAS NOT ABLE TO BE RELEASED. THE SURGEON USED ESOPHAGEAL FORCEPS AND RESECTED THE TISSUE WITH A SCALPEL. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | CONTOUR CURVED CUTTER STAPLER | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A EXPLORATORY LAPAROTOMY WITH A SIGMOID RESECTION, THE CONTOUR MALFUNCTIONED, AND THAT INJURED THE BOWEL. DOCTOR SAID WAS ABLE TO DEPLOY THE PIN, WHEN THEY WENT TO FIRE THE STAPLER, GOT JAMMED IN PLACE AND WOULDN¿T OPEN OR CLOSE. THE DOCTOR CUT THE DEVICE OUT, TO REMOVE THE DEVICE FROM THE PATIENT, BUT HAD TO SUTURE TO CONTROL BLEEDING. BLOOD PRODUCTS NOT USED. USED A GREEN CONTOUR TO COMPLETE CASE. NO CONSEQUENCE TO THE PATIENT REPORTED.
17-12-2018 | Malfunction | ethicon endo-surgery | 07-01-2019 | ILS 33MM, CURVED | Material Twisted/Bent | A040609 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP SIGMA PROCEDURE, WHEN PLUGGING THE INSTRUMENT (HEAD WITH INSTRUMENT) THE MANDREL WAS BENT. NO HARM TO THE PATIENT.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY ROUX-EN-Y PROCEDURE, THE 45MM RELOAD WAS USED DURING DUODENECTOMY. THE NURSE SAID AFTER OPENING/CLOSING CHECK, HE/SHE HANDED THE HANDLE TO THE SURGEON. THE SURGEON SET THE JAW ON THE TISSUE AND PRESSED THE GREEN BUTTON IN ORDER TO FIRE. WHEN SQUEEZING THE HANDLE, A RASPING SOUND WAS GENERATED AND THEN THE SURGEON RELEASED THE CARTRIDGE. WHEN HE/SHE FIRED AGAIN IN THE SAME WAY, IT WAS ABLE TO STAPLE WITHOUT PROBLEM. THE 45MM RELOAD WAS ALSO USED IN ESOPHAGECTOMY. THIS TIME THE SURGEON CHECKED THE ACCURACY AND WHEN FIRING IN THE SAME WAY, THERE WAS ABNORMAL NOISE AND THE CARTRIDGE ADVANCED A LITTLE POSSIBLY OR NOT AT ALL. THE CARTRIDGE DID NOT ADVANCE LAST TIME BUT THIS TIME IT WAS NOT ABLE TO BE RELEASED. THE SURGEON USED ESOPHAGEAL FORCEPS AND RESECTED THE TISSUE WITH A SCALPEL. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY ROUX-EN-Y PROCEDURE, THE 45MM RELOAD WAS USED DURING DUODENECTOMY. THE NURSE SAID AFTER OPENING/CLOSING CHECK, HE/SHE HANDED THE HANDLE TO THE SURGEON. THE SURGEON SET THE JAW ON THE TISSUE AND PRESSED THE GREEN BUTTON IN ORDER TO FIRE. WHEN SQUEEZING THE HANDLE, A RASPING SOUND WAS GENERATED AND THEN THE SURGEON RELEASED THE CARTRIDGE. WHEN HE/SHE FIRED AGAIN IN THE SAME WAY, IT WAS ABLE TO STAPLE WITHOUT PROBLEM. THE 45MM RELOAD WAS ALSO USED IN ESOPHAGECTOMY. THIS TIME THE SURGEON CHECKED THE ACCURACY AND WHEN FIRING IN THE SAME WAY, THERE WAS ABNORMAL NOISE AND THE CARTRIDGE ADVANCED A LITTLE POSSIBLY OR NOT AT ALL. THE CARTRIDGE DID NOT ADVANCE LAST TIME BUT THIS TIME IT WAS NOT ABLE TO BE RELEASED. THE SURGEON USED ESOPHAGEAL FORCEPS AND RESECTED THE TISSUE WITH A SCALPEL. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY ROUX-EN-Y PROCEDURE, THE 45MM RELOAD WAS USED DURING DUODENECTOMY. THE NURSE SAID AFTER OPENING/CLOSING CHECK, HE/SHE HANDED THE HANDLE TO THE SURGEON. THE SURGEON SET THE JAW ON THE TISSUE AND PRESSED THE GREEN BUTTON IN ORDER TO FIRE. WHEN SQUEEZING THE HANDLE, A RASPING SOUND WAS GENERATED AND THEN THE SURGEON RELEASED THE CARTRIDGE. WHEN HE/SHE FIRED AGAIN IN THE SAME WAY, IT WAS ABLE TO STAPLE WITHOUT PROBLEM. THE 45MM RELOAD WAS ALSO USED IN ESOPHAGECTOMY. THIS TIME THE SURGEON CHECKED THE ACCURACY AND WHEN FIRING IN THE SAME WAY, THERE WAS ABNORMAL NOISE AND THE CARTRIDGE ADVANCED A LITTLE POSSIBLY OR NOT AT ALL. THE CARTRIDGE DID NOT ADVANCE LAST TIME BUT THIS TIME IT WAS NOT ABLE TO BE RELEASED. THE SURGEON USED ESOPHAGEAL FORCEPS AND RESECTED THE TISSUE WITH A SCALPEL. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-09-2018 | Malfunction | us surgical puerto rico | 07-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SEGMENTECTOMY, WHEN STAPLING A SEGMENT OF THE LUNG WHEN THE DEVICE WAS ABOUT TO BE FIRED THE DEVICE LOCKED AND WAS NOT ABLE TO BE FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-09-2018 | Malfunction | covidien lp llc north haven | 07-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SEGMENTECTOMY, WHEN STAPLING A SEGMENT OF THE LUNG WHEN THE DEVICE WAS ABOUT TO BE FIRED THE DEVICE LOCKED AND WAS NOT ABLE TO BE FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2019 | Malfunction | ethicon endo-surgery | 07-01-2019 | ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER, LAPAROSCOPIC STAPLER | Firing Problem | A0505 | No Known Impact Or Consequence To Patient | Event Description: ETHICON ENDO-SURGERY ENDOPATH ETS FLEX 45 STAPLER, AS BEING UTILIZED PER MFR'S INSTRUCTIONS, WOULD NOT DEPLOY WHEN ATTEMPTED TO BE ACTIVATED. ENDOPATH ETS45 3.5 MM RELOAD IN PLACE.
10-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY SIDE TO SIDE ANASTOMOSIS, AT THE THIRD FIRING, THE SURGEON CLAMPED TISSUE AND WAS ABLE TO CLOSE THE HANDLE, BUT HE/SHE WAS NOT ABLE TO FIRE THE DEVICE. OPERATION WAS CONTINUED WITH A NEW HANDLE AND A NEW RELOAD. OPERATION WAS CONTINUED WITH A NEW DEVICE. NO PATIENT HARM.
10-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY SIDE TO SIDE ANASTOMOSIS, AT THE THIRD FIRING, THE SURGEON CLAMPED TISSUE AND WAS ABLE TO CLOSE THE HANDLE, BUT HE/SHE WAS NOT ABLE TO FIRE THE DEVICE. OPERATION WAS CONTINUED WITH A NEW HANDLE AND A NEW RELOAD. OPERATION WAS CONTINUED WITH A NEW DEVICE. NO PATIENT HARM.
08-01-2019 | Malfunction | covidien lp llc north haven | 08-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
11-02-2019 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Difficult to Insert | A150206 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RIGHT COLECTOMY ON THE TRAVERSE COLON, THE CARTRIDGE COULD NOT BE INSERTED EASILY IN A TROCAR AS IT'S SUPPOSED TO BE, AND ONCE INSERTED THE DEVICE COULD NOT BE FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RIGHT COLECTOMY ON THE TRAVERSE COLON, THE CARTRIDGE COULD NOT BE INSERTED EASILY IN A TROCAR AS IT'S SUPPOSED TO BE, AND ONCE INSERTED THE DEVICE COULD NOT BE FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Insert | A150206 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RIGHT COLECTOMY ON THE TRAVERSE COLON, THE CARTRIDGE COULD NOT BE INSERTED EASILY IN A TROCAR AS IT'S SUPPOSED TO BE, AND ONCE INSERTED THE DEVICE COULD NOT BE FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RIGHT COLECTOMY ON THE TRAVERSE COLON, THE CARTRIDGE COULD NOT BE INSERTED EASILY IN A TROCAR AS IT'S SUPPOSED TO BE, AND ONCE INSERTED THE DEVICE COULD NOT BE FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
29-11-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHEN THEY ARE ABOUT TO FIRE THE DEVICE, THE WAS BLOCKED AND THEY TOOK ANOTHER HANDLE BUT IT WAS ALSO DID NOT FIRE. THEY HAD TO USE A THIRD DEVICE TO COMPLETE THE CASE.
29-11-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHEN THEY ARE ABOUT TO FIRE THE DEVICE, THE WAS BLOCKED AND THEY TOOK ANOTHER HANDLE BUT IT WAS ALSO DID NOT FIRE. THEY HAD TO USE A THIRD DEVICE TO COMPLETE THE CASE.
06-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE FIRED BUT FAILED CAUSING POOR STAPLE FORMATION AND INTRA-OPERATIVE BLEEDING IN THE PATIENT.
06-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE FIRED BUT FAILED CAUSING POOR STAPLE FORMATION AND INTRA-OPERATIVE BLEEDING IN THE PATIENT.
29-11-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHEN THEY WERE ABOUT TO FIRE THE DEVICE, IT WAS BLOCKED AND THEY TOOK ANOTHER HANDLE BUT IT ALSO DID NOT FIRE. THEY HAD TO USE A THIRD DEVICE TO COMPLETE THE CASE.
29-11-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHEN THEY WERE ABOUT TO FIRE THE DEVICE, IT WAS BLOCKED AND THEY TOOK ANOTHER HANDLE BUT IT ALSO DID NOT FIRE. THEY HAD TO USE A THIRD DEVICE TO COMPLETE THE CASE.
05-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | TA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A PNEUMONECTOMY, THE DEVICE FIRED PARTIALLY IN THE PULMONARY ARTERY AND STAPLED ONLY HALF OF THE SECTION. ANOTHER RELOAD WAS USED TO COMPLETE THE STAPLE LINE. THE PATIENT EXPERIENCED BLEEDING OF NOT MORE THAN 500C DURING THE DEVICE CHANGE, BUT THERE WAS NO HEMODYNAMIC REPERCUSSION. THERE WAS NO PATIENT HARM.
05-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | TA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A PNEUMONECTOMY, THE DEVICE FIRED PARTIALLY IN THE PULMONARY ARTERY AND STAPLED ONLY HALF OF THE SECTION. ANOTHER RELOAD WAS USED TO COMPLETE THE STAPLE LINE. THE PATIENT EXPERIENCED BLEEDING OF NOT MORE THAN 500C DURING THE DEVICE CHANGE, BUT THERE WAS NO HEMODYNAMIC REPERCUSSION. THERE WAS NO PATIENT HARM.
26-11-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Gas/Air Leak | A050402 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION PROCEDURE, THERE WAS BUBBLING AT THE ANASTOMOSIS. THERE WERE HEALTHY DONUTS, STAPLE LINE WAS COMPLETE BUT WHEN THEY DID THE LEAK CHECK, THERE WAS BUBBLING AT THE ANASTOMOSIS. NO SIZERS WERE USED IN THIS EVENT. TO RESOLVE THE ISSUE, THEY OVERSEW THE STAPLE LINE/ANASTOMOSIS. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
26-11-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Gas/Air Leak | A050402 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION PROCEDURE, THERE WAS BUBBLING AT THE ANASTOMOSIS. THERE WERE HEALTHY DONUTS, STAPLE LINE WAS COMPLETE BUT WHEN THEY DID THE LEAK CHECK, THERE WAS BUBBLING AT THE ANASTOMOSIS. NO SIZERS WERE USED IN THIS EVENT. TO RESOLVE THE ISSUE, THEY OVERSEW THE STAPLE LINE/ANASTOMOSIS. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
08-01-2019 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Gas/Air Leak | A050402 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION, THERE WAS BUBBLING AT THE ANASTOMOSIS. THERE WERE HEALTHY DONUTS, STAPLE LINE WAS COMPLETE BUT WHEN THEY DID THE LEAK CHECK, THERE WAS BUBBLING AT THE ANASTOMOSIS. NO SIZERS WERE USED IN THIS EVENT. TO RESOLVE THE ISSUE, THEY OVERSEW THE STAPLE LINE/ANASTOMOSIS. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
08-01-2019 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Gas/Air Leak | A050402 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION, THERE WAS BUBBLING AT THE ANASTOMOSIS. THERE WERE HEALTHY DONUTS, STAPLE LINE WAS COMPLETE BUT WHEN THEY DID THE LEAK CHECK, THERE WAS BUBBLING AT THE ANASTOMOSIS. NO SIZERS WERE USED IN THIS EVENT. TO RESOLVE THE ISSUE, THEY OVERSEW THE STAPLE LINE/ANASTOMOSIS. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
06-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LUNG RESECTION, THE INSTRUMENT STOPPED WORKING AFTER 11 RELOADS. THEY OPENED ANOTHER HANDLE WITH THE SAME RELOAD USED ON FAULTY HANDLE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LUNG RESECTION, THE INSTRUMENT STOPPED WORKING AFTER 11 RELOADS. THEY OPENED ANOTHER HANDLE WITH THE SAME RELOAD USED ON FAULTY HANDLE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN DISCONNECTING INTERLOBULAR FISSURE DURING A LAPAROSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE WAS NOT ABLE TO FIRE. THEY REPLACED THE DEVICES TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN DISCONNECTING INTERLOBULAR FISSURE DURING A LAPAROSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE WAS NOT ABLE TO FIRE. THEY REPLACED THE DEVICES TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE STAPLERS' HANDLES COULD NOT BE SQUEEZED AND THE STAPLERS DID NOT FIRE. THE RELOADS WERE OBSERVED TO HAVE PRE-FIRED. A NEW DEVICE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE STAPLERS' HANDLES COULD NOT BE SQUEEZED AND THE STAPLERS DID NOT FIRE. THE RELOADS WERE OBSERVED TO HAVE PRE-FIRED. A NEW DEVICE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE STAPLERS' HANDLES COULD NOT BE SQUEEZED AND THE STAPLERS DID NOT FIRE. THE RELOADS WERE OBSERVED TO HAVE PRE-FIRED. A NEW DEVICE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FOURTH FIRING OF THE STAPLER IN A LAPAROSCOPIC LUNG RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE STOPPED FIRING. THEY CHANGED HANDLE FIRST BUT IT STILL DID NOT WORK WITH THE SAME RELOAD. THEY CHANGED A RELOAD AND THE NEW HANDLE WORKED WELL. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE STAPLERS' HANDLES COULD NOT BE SQUEEZED AND THE STAPLERS DID NOT FIRE. THE RELOADS WERE OBSERVED TO HAVE PRE-FIRED. A NEW DEVICE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FOURTH FIRING OF THE STAPLER IN A LAPAROSCOPIC LUNG RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE STOPPED FIRING. THEY CHANGED HANDLE FIRST BUT IT STILL DID NOT WORK WITH THE SAME RELOAD. THEY CHANGED A RELOAD AND THE NEW HANDLE WORKED WELL. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE, THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE, THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE FIRING ON THE STOMACH, ONE OF THE STAPLES WAS DEPLOYED FROM THE STAPLE BUCKET BUT WAS CAUGHT/PINCHED WITHIN THE STAPLE BUCKET BY THE PUSHER. THE RELOAD COULD NOT INITIALLY BE REMOVED FROM THE TISSUE. THE SURGEON REMOVED IT BY GENTLY GRABBING THE TISSUE WITH A GRASPER AND PULLING IT AWAY FROM THE STAPLER. THE STAPLE THAT WAS STUCK ON THE RELOAD REMAINED ON THE RELOAD AND WAS PULLED FROM THE TISSUE. THE SURGEON STITCHED PART OF THE OMENTUM TO THE STAPLE LINE AS PART OF THE STANDARD ROUTINE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE FIRING ON THE STOMACH, ONE OF THE STAPLES WAS DEPLOYED FROM THE STAPLE BUCKET BUT WAS CAUGHT/PINCHED WITHIN THE STAPLE BUCKET BY THE PUSHER. THE RELOAD COULD NOT INITIALLY BE REMOVED FROM THE TISSUE. THE SURGEON REMOVED IT BY GENTLY GRABBING THE TISSUE WITH A GRASPER AND PULLING IT AWAY FROM THE STAPLER. THE STAPLE THAT WAS STUCK ON THE RELOAD REMAINED ON THE RELOAD AND WAS PULLED FROM THE TISSUE. THE SURGEON STITCHED PART OF THE OMENTUM TO THE STAPLE LINE AS PART OF THE STANDARD ROUTINE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WOULD NOT LOAD TO THE HANDLE. THE PLASTIC PIECE ON THE RELOAD WAS BROKEN BUT WAS STILL ATTACHED. ANOTHER RELOAD WAS LOADED ON THE SAME HANDLE AND IT WORKED FINE FOR THE CASE. PROBLEM WAS ISOLATED TO THIS PARTICULAR RELOAD. NO PATIENT INVOLVEMENT.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WOULD NOT LOAD TO THE HANDLE. THE PLASTIC PIECE ON THE RELOAD WAS BROKEN BUT WAS STILL ATTACHED. ANOTHER RELOAD WAS LOADED ON THE SAME HANDLE AND IT WORKED FINE FOR THE CASE. PROBLEM WAS ISOLATED TO THIS PARTICULAR RELOAD. NO PATIENT INVOLVEMENT.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE FIRING ON THE STOMACH, ONE OF THE STAPLES WAS DEPLOYED FROM THE STAPLE BUCKET BUT WAS CAUGHT/PINCHED WITHIN THE STAPLE BUCKET BY THE PUSHER. THE RELOAD COULD NOT INITIALLY BE REMOVED FROM THE TISSUE. THE SURGEON REMOVED IT BY GENTLY GRABBING THE TISSUE WITH A GRASPER AND PULLING IT AWAY FROM THE STAPLER. THE STAPLE THAT WAS STUCK ON THE RELOAD REMAINED ON THE RELOAD AND WAS PULLED FROM THE TISSUE. THE SURGEON STITCHED PART OF THE OMENTUM TO THE STAPLE LINE AS PART OF THE STANDARD ROUTINE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Break | A0401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING REINSTALLATION OF COLON TRANSIT, THE DEVICE BROKE. A COMPONENT FELL INTO PATIENT CAVITY. THERE WAS UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE THAT LED TO ADDITIONAL 30 MINUTES OR MORE SURGICAL TIME. ANOTHER STAPLER WAS PROVIDED TO COMPLETE THE CASE.
06-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Break | A0401 | Device Embedded In Tissue or Plaque | Event Description: ACCORDING TO THE REPORTER, DURING REINSTALLATION OF COLON TRANSIT, THE DEVICE BROKE. A COMPONENT FELL INTO PATIENT CAVITY. THERE WAS UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE THAT LED TO ADDITIONAL 30 MINUTES OR MORE SURGICAL TIME. ANOTHER STAPLER WAS PROVIDED TO COMPLETE THE CASE.
06-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Detachment of Device or Device Component | A0501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING REINSTALLATION OF COLON TRANSIT, THE DEVICE BROKE. A COMPONENT FELL INTO PATIENT CAVITY. THERE WAS UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE THAT LED TO ADDITIONAL 30 MINUTES OR MORE SURGICAL TIME. ANOTHER STAPLER WAS PROVIDED TO COMPLETE THE CASE.
06-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Detachment of Device or Device Component | A0501 | Device Embedded In Tissue or Plaque | Event Description: ACCORDING TO THE REPORTER, DURING REINSTALLATION OF COLON TRANSIT, THE DEVICE BROKE. A COMPONENT FELL INTO PATIENT CAVITY. THERE WAS UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE THAT LED TO ADDITIONAL 30 MINUTES OR MORE SURGICAL TIME. ANOTHER STAPLER WAS PROVIDED TO COMPLETE THE CASE.
06-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE RESECTING THE ARTERY, THERE WAS A PROBLEM WITH UNLOADING THE DEVICE, THE STAPLES BLOCKED THE HANDLE AND THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON, THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE RESECTING THE ARTERY, THERE WAS A PROBLEM WITH UNLOADING THE DEVICE, THE STAPLES BLOCKED THE HANDLE AND THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON, THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
11-10-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC CANCER RESECTION PROCEDURE, THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-10-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC CANCER RESECTION PROCEDURE, THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE RESECTING THE ARTERY, THERE WAS A PROBLEM WITH UNLOADING THE DEVICE, THE STAPLES BLOCKED THE HANDLE AND THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON, THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE RESECTING THE ARTERY, THERE WAS A PROBLEM WITH UNLOADING THE DEVICE, THE STAPLES BLOCKED THE HANDLE AND THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON, THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE ON THE STOMACH, THE DOCTOR LOADED THE RELOAD AND POSITIONED IT ON THE TISSUE AND WHEN READY TO FIRE THE RELOAD WAS BLOCKED. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED HE HANDLE AND THE DEVICE DID NOT FIRE. THE CASE WAS COMPLETED USING ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
05-04-2017 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SURGICAL STAPLING IN A LAPAROSCOPIC PROCEDURE, THE HANDLE AND THE 2 RELOADS LOCKED. THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. WHEN THE STAPLER AND RELOADS WERE REPLACED WITH THE SAME LOT NUMBER, THE SURGERY WAS PERFORMED NORMALLY. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE ON THE STOMACH, THE DOCTOR LOADED THE RELOAD AND POSITIONED IT ON THE TISSUE AND WHEN READY TO FIRE THE RELOAD WAS BLOCKED. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED HE HANDLE AND THE DEVICE DID NOT FIRE. THE CASE WAS COMPLETED USING ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
05-04-2017 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SURGICAL STAPLING IN A LAPAROSCOPIC PROCEDURE, THE HANDLE AND THE 2 RELOADS LOCKED. THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. WHEN THE STAPLER AND RELOADS WERE REPLACED WITH THE SAME LOT NUMBER, THE SURGERY WAS PERFORMED NORMALLY. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Unintended Movement | A0512 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED TOTAL GASTRECTOMY PROCEDURE, THE CARTRIDGE UNINTENTIONALLY ARTICULATED TO LEFT DURING GASTROJEJUNOSTOMY. BECAUSE THERE WAS A PANCREAS AT THE END OF THE CARTRIDGE WHICH WAS ARTICULATED TO LEFT (THE CARTRIDGE ON THE SIDE OF THE ESOPHAGUS AND ANVIL ON THE JEJUNUM SIDE), IT WAS QUITE DANGEROUS. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
05-04-2017 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SURGICAL STAPLING IN A LAPAROSCOPIC PROCEDURE, THE HANDLE AND THE 2 RELOADS LOCKED. THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. WHEN THE STAPLER AND RELOADS WERE REPLACED WITH THE SAME LOT NUMBER, THE SURGERY WAS PERFORMED NORMALLY. THERE WAS NO PATIENT INJURY.
29-11-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LEFT LATERAL LIVER RESECTION, WHEN STAPLING THE PARENCHYMA OF THE LIVER, THERE WAS A PROBLEM UNLOADING THE DEVICE. THE RELOAD COULD NOT BE OPENED AFTER BEING FIRED. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
29-11-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LEFT LATERAL LIVER RESECTION, WHEN STAPLING THE PARENCHYMA OF THE LIVER, THERE WAS A PROBLEM UNLOADING THE DEVICE. THE RELOAD COULD NOT BE OPENED AFTER BEING FIRED. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
01-12-2017 | Malfunction | us surgical puerto rico | 08-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: OCCURRED DURING AN OPEN GASTRECTOMY WITH PANCREATECTOMY PROCEDURE. THE STAPLER WAS NOT ABLE TO FIRE. THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. THE STAPLES DID NOT DEPLOY. IN ORDER TO RESOLVE THE ISSUE AND COMPLETE THE CASE, REPLACED THE STAPLER. THERE WAS NO PATIENT HARM.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Unintended Movement | A0512 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED TOTAL GASTRECTOMY PROCEDURE, THE CARTRIDGE UNINTENTIONALLY ARTICULATED TO LEFT DURING GASTROJEJUNOSTOMY. BECAUSE THERE WAS A PANCREAS AT THE END OF THE CARTRIDGE WHICH WAS ARTICULATED TO LEFT (THE CARTRIDGE ON THE SIDE OF THE ESOPHAGUS AND ANVIL ON THE JEJUNUM SIDE), IT WAS QUITE DANGEROUS. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
01-11-2016 | Injury | us surgical puerto rico | 08-01-2019 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Prolapse | Event Description: ACCORDING TO THE REPORTER, AFTER COLORECTAL PROCEDURE, PATIENT SUFFERED MEDICAL COMPLICATIONS. POSTOPERATIVE BOWEL DYSFUNCTION WAS ASSESSED. 11 (23%) PATIENTS HAD SOME DEGREE OF RECURRENT PROLAPSE. PERSISTENCE OF ANAL BLEEDING WAS SIGNIFICANTLY HIGHER IN THE PPH GROUP (P = 0.04). SEVEN (15%) OF THE 48 PATIENTS IN THE GROUP COMPLAINED OF SOME EVACUATION DISTURBANCE REGARDING RECURRENCE OF PROLAPSE AND OTHER HAEMORRHOIDAL SYMPTOMS.
01-11-2016 | Injury | us surgical puerto rico | 08-01-2019 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, AFTER COLORECTAL PROCEDURE, PATIENT SUFFERED MEDICAL COMPLICATIONS. POSTOPERATIVE BOWEL DYSFUNCTION WAS ASSESSED. 11 (23%) PATIENTS HAD SOME DEGREE OF RECURRENT PROLAPSE. PERSISTENCE OF ANAL BLEEDING WAS SIGNIFICANTLY HIGHER IN THE PPH GROUP (P = 0.04). SEVEN (15%) OF THE 48 PATIENTS IN THE GROUP COMPLAINED OF SOME EVACUATION DISTURBANCE REGARDING RECURRENCE OF PROLAPSE AND OTHER HAEMORRHOIDAL SYMPTOMS.
04-10-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER A NUMBER OF FIRINGS DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, THE DEVICE FIRED ONCE AND THEY COULD NOT UNLOAD THE LOADING UNIT. THERE WAS NO PATIENT INJURY.
17-08-2017 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A BYPASS GASTRECTOMY, DURING THE SURGICAL STAPLING, THE DEVICE WAS NOT ABLE TO PERFORM THE SECOND FIRING. IT WAS REPORTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
17-08-2017 | Malfunction | us surgical puerto rico | 08-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A BYPASS GASTRECTOMY, DURING THE SURGICAL STAPLING, THE DEVICE WAS NOT ABLE TO PERFORM THE SECOND FIRING. IT WAS REPORTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Device Displays Incorrect Message | A090201 | No Patient Involvement | Event Description: ACCORDING TO A ESOPHAGECTOMY, AFTER THE CHARGE COMPLETION WAS CONFIRMED, WHEN THE SHELL AND ADAPTER WERE ATTACHED, THE SCREEN WAS BLACKOUT, AND THEN THERE WAS NO RESPONSE ON THE DEVICE. AFTER THAT, ALTHOUGH THE HANDLE WAS CHARGED, THE INDICATOR FLASHED RED, AND THE DEVICE DID NOT RETURNED TO NORMAL. THE OPERATION WAS PERFORMED WITH A MANUAL ADAPTER. THERE WAS NO PATIENT INVOLVED.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Complete Loss of Power | A070801 | No Patient Involvement | Event Description: ACCORDING TO A ESOPHAGECTOMY, AFTER THE CHARGE COMPLETION WAS CONFIRMED, WHEN THE SHELL AND ADAPTER WERE ATTACHED, THE SCREEN WAS BLACKOUT, AND THEN THERE WAS NO RESPONSE ON THE DEVICE. AFTER THAT, ALTHOUGH THE HANDLE WAS CHARGED, THE INDICATOR FLASHED RED, AND THE DEVICE DID NOT RETURNED TO NORMAL. THE OPERATION WAS PERFORMED WITH A MANUAL ADAPTER. THERE WAS NO PATIENT INVOLVED.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE USING THE DEVICE ON SMALL INTESTINE ANASTOMOSIS, IT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON HOWEVER THERE WAS NO WAY TO FIRE THE RELOAD EVEN THE 3 GREEN LIGHT WERE ON. THERE WAS A SMALL NOISE AND ALARM, NO ERROR CODES AND WHEN JAWS OPENED, RELOAD TURN TO 0. THE RELOADS IN QUESTION WAS TESTED WITH NEW SHELL. IT WAS NOTED THAT THE LOADING AND CALIBRATION WAS FINE WITH 3 GREEN LIGHTS BUT STILL NO FIRE AVAILABLE. A NEW STAPLING SYSTEM WAS TAKEN AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE USING THE DEVICE ON SMALL INTESTINE ANASTOMOSIS, IT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON HOWEVER THERE WAS NO WAY TO FIRE THE RELOAD EVEN THE 3 GREEN LIGHT WERE ON. THERE WAS A SMALL NOISE AND ALARM, NO ERROR CODES AND WHEN JAWS OPENED, RELOAD TURN TO 0. THE RELOADS IN QUESTION WAS TESTED WITH NEW SHELL. IT WAS NOTED THAT THE LOADING AND CALIBRATION WAS FINE WITH 3 GREEN LIGHTS BUT STILL NO FIRE AVAILABLE. A NEW STAPLING SYSTEM WAS TAKEN AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Device Displays Incorrect Message | A090201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE USING THE DEVICE ON SMALL INTESTINE ANASTOMOSIS, IT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON HOWEVER THERE WAS NO WAY TO FIRE THE RELOAD EVEN THE 3 GREEN LIGHT WERE ON. THERE WAS A SMALL NOISE AND ALARM, NO ERROR CODES AND WHEN JAWS OPENED, RELOAD TURN TO 0. THE RELOADS IN QUESTION WAS TESTED WITH NEW SHELL. IT WAS NOTED THAT THE LOADING AND CALIBRATION WAS FINE WITH 3 GREEN LIGHTS BUT STILL NO FIRE AVAILABLE. A NEW STAPLING SYSTEM WAS TAKEN AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE USING THE DEVICE ON SMALL INTESTINE ANASTOMOSIS, IT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON HOWEVER THERE WAS NO WAY TO FIRE THE RELOAD EVEN THE 3 GREEN LIGHT WERE ON. THERE WAS A SMALL NOISE AND ALARM, NO ERROR CODES AND WHEN JAWS OPENED, RELOAD TURN TO 0. THE RELOADS IN QUESTION WAS TESTED WITH NEW SHELL. IT WAS NOTED THAT THE LOADING AND CALIBRATION WAS FINE WITH 3 GREEN LIGHTS BUT STILL NO FIRE AVAILABLE. A NEW STAPLING SYSTEM WAS TAKEN AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE USING THE DEVICE ON SMALL INTESTINE ANASTOMOSIS, IT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON HOWEVER THERE WAS NO WAY TO FIRE THE RELOAD EVEN THE 3 GREEN LIGHT WERE ON. THERE WAS A SMALL NOISE AND ALARM, NO ERROR CODES AND WHEN JAWS OPENED, RELOAD TURN TO 0. THE RELOADS IN QUESTION WAS TESTED WITH NEW SHELL. IT WAS NOTED THAT THE LOADING AND CALIBRATION WAS FINE WITH 3 GREEN LIGHTS BUT STILL NO FIRE AVAILABLE. A NEW STAPLING SYSTEM WAS TAKEN AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-11-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Gas/Air Leak | A050402 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION, THERE WAS BUBBLING AT THE ANASTOMOSIS. THERE WERE HEALTHY DONUTS, STAPLE LINE WAS COMPLETE BUT WHEN THEY DID THE LEAK CHECK, THERE WAS BUBBLING AT THE ANASTOMOSIS. TO RESOLVE THE ISSUE, THEY OVERSEW THE STAPLE LINE/ANASTOMOSIS. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
20-11-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Gas/Air Leak | A050402 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION, THERE WAS BUBBLING AT THE ANASTOMOSIS. THERE WERE HEALTHY DONUTS, STAPLE LINE WAS COMPLETE BUT WHEN THEY DID THE LEAK CHECK, THERE WAS BUBBLING AT THE ANASTOMOSIS. TO RESOLVE THE ISSUE, THEY OVERSEW THE STAPLE LINE/ANASTOMOSIS. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
01-01-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | 55 MM NTLC SELECTABLE RELOAD | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN THE COURSE OF A PROCEDURE (RT.HEMI-COLECTOMY) USING THE NTLC, AN ERROR OCCURRED IN THE PROCESS OF FIRING. THE BLADE SUDDENLY STOPPED IN THE MIDDLE OF THE DEVICE, AND THE BLADE DIDN'T MOVE FORWARD. AS AN EMERGENCY MEASURE, THE SURGEON BACKED OFF THE FIRING KNOB AND DETACH IT FROM THE TISSUE. HE CHANGED A NEW CARTRIDGE AND FINISHED THE PROCEDURE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES FOR THE PATIENT SO FAR.
01-01-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | CONTOUR CURVED CUTTER STAPLER | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON COULD NOT CLOSE THE STAPLER ON TISSUE. IT WAS INITIAL FIRING OF THE STAPLER AND THE RELOAD WAS NEVER TAKEN OUT OF THE STAPLER. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2017 | Malfunction | ethicon endo-surgery | 08-01-2019 | ILS 29MM, CURVED | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE DIDN'T WORK PROPERLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
11-12-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | PWRD ECH FLEX 45MM 34CM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RADICAL RESECTION OF LEFT UPPER LUNG CANCER, AFTER FIRED WHEN SEVERING THE BRONCHUS, THE BLADE WAS DIFFICULT TO RETURN WITH ABNORMAL WEAK MOTOR SOUND, AT LAST, THE BLADE ONLY RETURNED 2/3 AS THE PHOTO SHOWS. THE KNIFE REVERSE BUTTON WAS LOOSE, OPENED ONE NEW GUN'S PACKING, USED THIS BATTERY WITH PREVIOUS GUN, STILL COULD NOT RETURN THE BLADE. OPENED THE MANUAL OVERRIDE DOOR, FOUND COULD NOT PUSH OR PULL THE LEVER AS THE PHOTO SHOWS. USED ANOTHER DEVICE TO EXTEND JAW'S GAP TO REMOVE THE DEVICE FORM PATIENT. CHECKED THE TISSUE WITH GOOD CUT LINE AND STAPLE LINE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-10-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: COULD NOT FIRE. DURING THE ESOPHAGUS SURGERY, COULD NOT FIRE ON THE 1ST FIRING WHEN SEVERING THE BOTTOM OF STOMACH. ANOTHER RELOAD WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION, WHILE PERFORMING ESOPHAGEAL ANASTOMOSIS, THE DEVICE DID NOT CREATE A PROPER ANASTOMOSIS. THE STAPLES DID NOT FIRE A FULL 360 DEGREES AND IT WAS FOUND OUT THAT ONLY 1 DONUT (PROXIMAL ESOPHAGUS) BUT NOT THE OTHER WAS CREATED. SURGEON HAD TO OPEN UP, RESECTED THE DAMAGED AREA AND CREATED ANOTHER ANASTOMOSIS WITH CIRCULAR STAPLER THEN SUTURE AND CLOSED EXISTING HOLE IN STOMACH. THE SURGICAL TIME EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
12-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION, WHILE PERFORMING ESOPHAGEAL ANASTOMOSIS, THE DEVICE DID NOT CREATE A PROPER ANASTOMOSIS. THE STAPLES DID NOT FIRE A FULL 360 DEGREES AND IT WAS FOUND OUT THAT ONLY 1 DONUT (PROXIMAL ESOPHAGUS) BUT NOT THE OTHER WAS CREATED. SURGEON HAD TO OPEN UP, RESECTED THE DAMAGED AREA AND CREATED ANOTHER ANASTOMOSIS WITH CIRCULAR STAPLER THEN SUTURE AND CLOSED EXISTING HOLE IN STOMACH. THE SURGICAL TIME EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
12-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION, WHILE PERFORMING ESOPHAGEAL ANASTOMOSIS, THE DEVICE DID NOT CREATE A PROPER ANASTOMOSIS. THE STAPLES DID NOT FIRE A FULL 360 DEGREES AND IT WAS FOUND OUT THAT ONLY 1 DONUT (PROXIMAL ESOPHAGUS) BUT NOT THE OTHER WAS CREATED. SURGEON HAD TO OPEN UP, RESECTED THE DAMAGED AREA AND CREATED ANOTHER ANASTOMOSIS WITH CIRCULAR STAPLER THEN SUTURE AND CLOSED EXISTING HOLE IN STOMACH. THE SURGICAL TIME EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
12-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION, WHILE PERFORMING ESOPHAGEAL ANASTOMOSIS, THE DEVICE DID NOT CREATE A PROPER ANASTOMOSIS. THE STAPLES DID NOT FIRE A FULL 360 DEGREES AND IT WAS FOUND OUT THAT ONLY 1 DONUT (PROXIMAL ESOPHAGUS) BUT NOT THE OTHER WAS CREATED. SURGEON HAD TO OPEN UP, RESECTED THE DAMAGED AREA AND CREATED ANOTHER ANASTOMOSIS WITH CIRCULAR STAPLER THEN SUTURE AND CLOSED EXISTING HOLE IN STOMACH. THE SURGICAL TIME EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
10-12-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | GST BLUE RELOAD, 60MM, 6 ROW | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN CHOLECYSTECTOMY, THE KNIFE WAS MOVED FORWARD BY MISTAKE WHEN A NURSE CHECKED TO OPEN/CLOSE OF THE JAWS AFTER THE CARTRIDGE WAS LOADED FOR THE FIRST FIRING. IT FOLLOWED THAT THE JAWS COULD NOT OPEN. THE NURSE DID NOT ACQUIRE HOW TO REVERSE THE KNIFE TO HOME POSITION. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
03-12-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE DEVICE WAS LOCKED OUT AT THE FIRST FIRING ON THE APPENDIX. IT WAS FOUND THAT THE STAPLES AT THE PROXIMAL END OF THE CARTRIDGE WERE FORMED IN LUMBRICOID SHAPE WHEN THE JAWS WERE OPENED TO CHECK THE TARGET TISSUE SINCE THE FIRING TRIGGER COULD NOT BE FULLY GRASPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
03-12-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE DEVICE WAS LOCKED OUT AT THE FIRST FIRING ON THE APPENDIX. IT WAS FOUND THAT THE STAPLES AT THE PROXIMAL END OF THE CARTRIDGE WERE FORMED IN LUMBRICOID SHAPE WHEN THE JAWS WERE OPENED TO CHECK THE TARGET TISSUE SINCE THE FIRING TRIGGER COULD NOT BE FULLY GRASPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
08-12-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC RIGHT LUNG WEDGE RESECTION, AFTER FIRED ON THE 3RD FIRING, THE BLADE COULD NOT RETURN BACK AUTOMATICALLY. KNIFE REVERSE BUTTON DID NOT WORK, USED MANUAL OVERRIDE LEVER TO RETURN BLADE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
17-12-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | CONTOUR CURVED CUTTER STAPLER | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN GASTRIC TUBE FORMATION, THE KNIFE MOVED FORWARD AND STOPPED ON THE WAY OF FIRST FIRING. THE DEVICE WAS USED ON THE GASTRIC BODY. OOZING OCCURRED SINCE THE TARGET TISSUE WAS CUT SLIGHTLY BUT THE STAPLES AT THE DISTAL PART OF THE CARTRIDGE WERE FORMED IN LUMBRICOID SHAPE. THE AMOUNT OF THE OOZING WAS UNKNOWN. SUTURING WAS PERFORMED TO STOP THE OOZING. NO BLOOD TRANSFUSION WAS REQUIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
27-11-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC NEPHRECTOMY PROCEDURE WHEN FIRING THE FIRST TIME, THE TRIGGER BROKE. THE DEVICE WAS UNABLE TO FIRE AGAIN. A NEW DEVICE WAS TAKEN TO FINISH THE PROCEDURE.
27-11-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC NEPHRECTOMY PROCEDURE WHEN FIRING THE FIRST TIME, THE TRIGGER BROKE. THE DEVICE WAS UNABLE TO FIRE AGAIN. A NEW DEVICE WAS TAKEN TO FINISH THE PROCEDURE.
01-01-2018 | Malfunction | ethicon endo-surgery | 08-01-2019 | ETS*45 ENDO LIN CUT RELOAD WH | Failure to Advance | A150204 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THERE WAS A CARTRIDGE ERROR EXPERIENCED. THE BLADE DID NOT GO FORWARD ALTHOUGH A SURGEON TRY TO FIRE. SO SURGEON HAS CHANGED CARTRIDGE, OPERATION WAS SUCCESSFUL.
10-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Overheating of Device | A1005 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, WHEN THE POWER HANDLE WAS REMOVED FROM THE BATTERY CHARGER AND STARTUP WAS CHECKED, THEN BLACKOUT OCCURRED WHEN POWER SHELL WAS CONNECTED TO THE HANDLE. THE DISPLAY WAS ILLUMINATED BUT THEN IT TURNED TO BLACK-OUT COMPLETELY, AND NOTHING WAS DISPLAYED AT ALL. THE HANDLE IN QUESTION WAS RE-CONNECTED TO THE BATTERY CHARGER, AND TRIED RESETTING THE DEVICE BUT THE STATUS WAS NOT CHANGED NOR WORKED AT ALL. THE INDICATORS OF THE BATTERY CHARGER ILLUMINATED BLUE AND FLASHED GREEN, BUT THEY DID NOT TURN TO RED FLASH. THE POWER HANDLE AND THE POWER SHELL BECAME UNUSABLE. IT WAS ALSO CONFIRMED THAT THE INSIDE OF THE POWER HANDLE OF THE REAR PART (NEAR THE DISPLAY ) WAS EXTRAORDINARILY HEATED ( IT WAS NOTICED THAT THE TEMPERATURE HAD RISEN CONSIDERABLY WITHIN ABOUT A FEW MINUTES FROM THE CHARGE STARTED ). SURGERY WAS PERFORMED AND COMPLETED WITH ANOTHER POWER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
10-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Complete Loss of Power | A070801 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, WHEN THE POWER HANDLE WAS REMOVED FROM THE BATTERY CHARGER AND STARTUP WAS CHECKED, THEN BLACKOUT OCCURRED WHEN POWER SHELL WAS CONNECTED TO THE HANDLE. THE DISPLAY WAS ILLUMINATED BUT THEN IT TURNED TO BLACK-OUT COMPLETELY, AND NOTHING WAS DISPLAYED AT ALL. THE HANDLE IN QUESTION WAS RE-CONNECTED TO THE BATTERY CHARGER, AND TRIED RESETTING THE DEVICE BUT THE STATUS WAS NOT CHANGED NOR WORKED AT ALL. THE INDICATORS OF THE BATTERY CHARGER ILLUMINATED BLUE AND FLASHED GREEN, BUT THEY DID NOT TURN TO RED FLASH. THE POWER HANDLE AND THE POWER SHELL BECAME UNUSABLE. IT WAS ALSO CONFIRMED THAT THE INSIDE OF THE POWER HANDLE OF THE REAR PART (NEAR THE DISPLAY ) WAS EXTRAORDINARILY HEATED ( IT WAS NOTICED THAT THE TEMPERATURE HAD RISEN CONSIDERABLY WITHIN ABOUT A FEW MINUTES FROM THE CHARGE STARTED ). SURGERY WAS PERFORMED AND COMPLETED WITH ANOTHER POWER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
11-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Difficult to Remove | A150207 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN DOING A GASTRIC JEJUNOSTOMY ANASTOMOSIS DURING A LAPAROSCOPIC GASTRIC BYPASS, THE TWIST KNOB TURNED TWO FULL TIMES TO ALLOW THE ANVIL TO TILT HOWEVER, THE ANVIL NEVER TILTED AFTER FIRING. IT WAS ALSO REPORTED THAT THERE WAS A POOR STAPLE FORMATION WHICH WAS FIXED THROUGH OVER SEWING AND DONUTS WERE NOTED TO BE COMPLETE. THE ANVIL NEVER TILTED AND WAS JUST PULLED OUT TO COMPLETE THE PROCEDURE. THE ANVIL NEVER TILTED AND WAS JUST PULLED OUT TO COMPLETE THE PROCEDURE.
11-12-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN DOING A GASTRIC JEJUNOSTOMY ANASTOMOSIS DURING A LAPAROSCOPIC GASTRIC BYPASS, THE TWIST KNOB TURNED TWO FULL TIMES TO ALLOW THE ANVIL TO TILT HOWEVER, THE ANVIL NEVER TILTED AFTER FIRING. IT WAS ALSO REPORTED THAT THERE WAS A POOR STAPLE FORMATION WHICH WAS FIXED THROUGH OVER SEWING AND DONUTS WERE NOTED TO BE COMPLETE. THE ANVIL NEVER TILTED AND WAS JUST PULLED OUT TO COMPLETE THE PROCEDURE. THE ANVIL NEVER TILTED AND WAS JUST PULLED OUT TO COMPLETE THE PROCEDURE.
11-12-2018 | Injury | covidien lp llc north haven | 08-01-2019 | EEA ORVIL | Difficult to Remove | A150207 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN DOING A GASTRIC JEJUNOSTOMY ANASTOMOSIS DURING A LAPAROSCOPIC GASTRIC BYPASS, THE TWIST KNOB TURNED TWO FULL TIMES TO ALLOW THE ANVIL TO TILT HOWEVER, THE ANVIL NEVER TILTED AFTER FIRING. IT WAS ALSO REPORTED THAT THERE WAS A POOR STAPLE FORMATION WHICH WAS FIXED THROUGH OVER SEWING AND DONUTS WERE NOTED TO BE COMPLETE. THE ANVIL NEVER TILTED AND WAS JUST PULLED OUT TO COMPLETE THE PROCEDURE. THE ANVIL NEVER TILTED AND WAS JUST PULLED OUT TO COMPLETE THE PROCEDURE.
11-12-2018 | Injury | covidien lp llc north haven | 08-01-2019 | EEA ORVIL | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN DOING A GASTRIC JEJUNOSTOMY ANASTOMOSIS DURING A LAPAROSCOPIC GASTRIC BYPASS, THE TWIST KNOB TURNED TWO FULL TIMES TO ALLOW THE ANVIL TO TILT HOWEVER, THE ANVIL NEVER TILTED AFTER FIRING. IT WAS ALSO REPORTED THAT THERE WAS A POOR STAPLE FORMATION WHICH WAS FIXED THROUGH OVER SEWING AND DONUTS WERE NOTED TO BE COMPLETE. THE ANVIL NEVER TILTED AND WAS JUST PULLED OUT TO COMPLETE THE PROCEDURE. THE ANVIL NEVER TILTED AND WAS JUST PULLED OUT TO COMPLETE THE PROCEDURE.
04-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, DURING THE FIRST FIRING ON THE STOMACH, THE SURGEON CLOSED THE STAPLER ON THE STOMACH, DEPRESSED THE GREEN BUTTON, AND ATTEMPTED TO FIRE BUT THE HANDLE COULD NOT BE SQUEEZED. NO STAPLES WEREFIRED ACCORDING TO THE SURGEON. BOTH A NEW STAPLER AND RELOAD WERE OPENED THAT FUNCTIONED NORMALLY. THERE WAS NO PATIENT INJURY.
04-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, DURING THE FIRST FIRING ON THE STOMACH, THE SURGEON CLOSED THE STAPLER ON THE STOMACH, DEPRESSED THE GREEN BUTTON, AND ATTEMPTED TO FIRE BUT THE HANDLE COULD NOT BE SQUEEZED. NO STAPLES WERE FIRED ACCORDING TO THE SURGEON. BOTH A NEW STAPLER AND RELOAD WERE OPENED THAT FUNCTIONED NORMALLY. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE SURGEON ATTEMPTED TO PERFORM THE SECOND FIRING DURING STOMACH RESECTION ON A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, AFTER THE GREEN BUTTON WAS PRESSED, THE DEVICE COULD NOT FIRE BUT THERE WERE TRACES OF PRE-FIRING COULD BE SEEN ON THE CARTRIDGE. IT WAS ALSO REPORTED THAT THE DEVICE¿S FEEDBACK SCREEN SHOWED A DISPLAY THAT FIRING WAS COMPLETED. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE SURGEON ATTEMPTED TO PERFORM THE SECOND FIRING DURING STOMACH RESECTION ON A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, AFTER THE GREEN BUTTON WAS PRESSED, THE DEVICE COULD NOT FIRE BUT THERE WERE TRACES OF PRE-FIRING COULD BE SEEN ON THE CARTRIDGE. IT WAS ALSO REPORTED THAT THE DEVICE¿S FEEDBACK SCREEN SHOWED A DISPLAY THAT FIRING WAS COMPLETED. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE SURGEON ATTEMPTED TO PERFORM THE SECOND FIRING DURING STOMACH RESECTION ON A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, AFTER THE GREEN BUTTON WAS PRESSED, THE DEVICE COULD NOT FIRE BUT THERE WERE TRACES OF PRE-FIRING COULD BE SEEN ON THE CARTRIDGE. IT WAS ALSO REPORTED THAT THE DEVICE¿S FEEDBACK SCREEN SHOWED A DISPLAY THAT FIRING WAS COMPLETED. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE SURGEON ATTEMPTED TO PERFORM THE SECOND FIRING DURING STOMACH RESECTION ON A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, AFTER THE GREEN BUTTON WAS PRESSED, THE DEVICE COULD NOT FIRE BUT THERE WERE TRACES OF PRE-FIRING COULD BE SEEN ON THE CARTRIDGE. IT WAS ALSO REPORTED THAT THE DEVICE¿S FEEDBACK SCREEN SHOWED A DISPLAY THAT FIRING WAS COMPLETED. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, THE SURGEON SQUEEZED THE HANDLE TO PERFORM THE FIRST FIRING ON THE STOMACH. AT THE FIRST SQUEEZE WHICH WAS TO START FIRES, A CLUNKY SOUND WAS GENERATED WHEN THE GEAR DISENGAGED AND THE HANDLE COULD NOT BE USED. BOTH THE HANDLE AND THE RELOAD WAS REPLACED. AFTER REPLACEMENT, THE DEVICE WAS ABLE TO PERFORM RESECTION NORMALLY AND SINCE THEN THE PROCEDURE WAS COMPLETED WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, THE SURGEON SQUEEZED THE HANDLE TO PERFORM THE FIRST FIRING ON THE STOMACH. AT THE FIRST SQUEEZE WHICH WAS TO START FIRES, A CLUNKY SOUND WAS GENERATED WHEN THE GEAR DISENGAGED AND THE HANDLE COULD NOT BE USED. BOTH THE HANDLE AND THE RELOAD WAS REPLACED. AFTER REPLACEMENT, THE DEVICE WAS ABLE TO PERFORM RESECTION NORMALLY AND SINCE THEN THE PROCEDURE WAS COMPLETED WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
07-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, THE SURGEON SQUEEZED THE HANDLE TO PERFORM THE FIRST FIRING ON THE STOMACH. AT THE FIRST SQUEEZE WHICH WAS TO START FIRES, A CLUNKY SOUND WAS GENERATED WHEN THE GEAR DISENGAGED AND THE HANDLE COULD NOT BE USED. BOTH THE HANDLE AND THE RELOAD WAS REPLACED. AFTER REPLACEMENT, THE DEVICE WAS ABLE TO PERFORM RESECTION NORMALLY AND SINCE THEN THE PROCEDURE WAS COMPLETED WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE USING STOMACH TRANSECTION, THE RELOAD COULD NOT BE OPENED AND CLOSED EASILY PRIOR TO FIRING. THERE WAS ALSO POOR STAPLE FORMATION REPORTED AT THE BEGINNING OF THE RELOAD AND STAPLE LINE INCOMPLETE AT THE PROXIMAL. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE USING STOMACH TRANSECTION, THE RELOAD COULD NOT BE OPENED AND CLOSED EASILY PRIOR TO FIRING. THERE WAS ALSO POOR STAPLE FORMATION REPORTED AT THE BEGINNING OF THE RELOAD AND STAPLE LINE INCOMPLETE AT THE PROXIMAL. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE USING STOMACH TRANSECTION, THE RELOAD COULD NOT BE OPENED AND CLOSED EASILY PRIOR TO FIRING. THERE WAS ALSO POOR STAPLE FORMATION REPORTED AT THE BEGINNING OF THE RELOAD AND STAPLE LINE INCOMPLETE AT THE PROXIMAL. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE USING STOMACH TRANSECTION, THE RELOAD COULD NOT BE OPENED AND CLOSED EASILY PRIOR TO FIRING. THERE WAS ALSO POOR STAPLE FORMATION REPORTED AT THE BEGINNING OF THE RELOAD AND STAPLE LINE INCOMPLETE AT THE PROXIMAL. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-11-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC BARIATRIC SURGERY, DURING THE STAPLING, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE OF THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-11-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | UNKNOWN TRI STAPLE PRODUCT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC BARIATRIC SURGERY, DURING THE STAPLING, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE OF THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY, UPON APPROACHING THE COLON, WHEN THEY FIRED THE DEVICE, THERE WAS A GOOD FORMATION OF STAPLES DURING THE 50% AND THE OTHER 50%, ONLY THE KNIFE CUT THE TISSUE AND THEY HAVEN'T STAPLE. THEY USED ANOTHER STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | UNKNOWN ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY, UPON APPROACHING THE COLON, WHEN THEY FIRED THE DEVICE, THERE WAS A GOOD FORMATION OF STAPLES DURING THE 50% AND THE OTHER 50%, ONLY THE KNIFE CUT THE TISSUE AND THEY HAVEN'T STAPLE. THEY USED ANOTHER STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-11-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN CYSTECTOMY PROCEDURE, WHEN THE SURGEON WAS DOING THE POUCH, THE SURGEON TRIED TO FIRE THE DEVICE, BUT THE HANDLE BROKE AND THE BLACK PUSHER THUMB PIECE DISENGAGED FROM THE DEVICE. THE FIRING WAS ABORTED, AND THE DEVICE WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-11-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | GIA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN CYSTECTOMY PROCEDURE, WHEN THE SURGEON WAS DOING THE POUCH, THE SURGEON TRIED TO FIRE THE DEVICE, BUT THE HANDLE BROKE AND THE BLACK PUSHER THUMB PIECE DISENGAGED FROM THE DEVICE. THE FIRING WAS ABORTED, AND THE DEVICE WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | GIA | Difficult or Delayed Positioning | A150203 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL CYSTECTOMY/ ILEAL CONDUIT CONSTRUCT, WHEN USING THE DEVICE AT THE SECOND FIRING, THE FIRING LEVER CAME OFF. THE SALES REP TOLD THE CUSTOMER THE REASON OF THE EVENT WAS THAT THE PART FIXING THE LEVER WAS DAMAGED SLIGHTLY. AT THE 1ST FIRING, THE LEVER COULD NOT BE PRESSED TO THE END AND SOMETIMES THE LEVER WAS VERY STIFF. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | GIA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL CYSTECTOMY/ ILEAL CONDUIT CONSTRUCT, WHEN USING THE DEVICE AT THE SECOND FIRING, THE FIRING LEVER CAME OFF. THE SALES REP TOLD THE CUSTOMER THE REASON OF THE EVENT WAS THAT THE PART FIXING THE LEVER WAS DAMAGED SLIGHTLY. AT THE 1ST FIRING, THE LEVER COULD NOT BE PRESSED TO THE END AND SOMETIMES THE LEVER WAS VERY STIFF. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Injury | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PNEUMONECTOMY, WHILE THE DEVICE WAS USED ON BRONCHUS, HANDLE MADE A CRACKING SOUND AND THE STAPLE LINE HAD SEPARATED. SURGEON OVERSEW THE STAPLE LINE TO SECURE THE BRONCHUS. TISSUE DAMAGE OCCURRED SINCE THERE WAS TISSUE SEPARATION AT THE STAPLE LINE. THERE WAS ALSO A PROBLEM UNLOADING THE DEVICE.
10-12-2018 | Injury | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Loss of or Failure to Bond | A040102 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PNEUMONECTOMY, WHILE THE DEVICE WAS USED ON BRONCHUS, HANDLE MADE A CRACKING SOUND AND THE STAPLE LINE HAD SEPARATED. SURGEON OVERSEW THE STAPLE LINE TO SECURE THE BRONCHUS. TISSUE DAMAGE OCCURRED SINCE THERE WAS TISSUE SEPARATION AT THE STAPLE LINE. THERE WAS ALSO A PROBLEM UNLOADING THE DEVICE.
10-12-2018 | Injury | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PNEUMONECTOMY, WHILE THE DEVICE WAS USED ON BRONCHUS, HANDLE MADE A CRACKING SOUND AND THE STAPLE LINE HAD SEPARATED. SURGEON OVERSEW THE STAPLE LINE TO SECURE THE BRONCHUS. TISSUE DAMAGE OCCURRED SINCE THERE WAS TISSUE SEPARATION AT THE STAPLE LINE. THERE WAS ALSO A PROBLEM UNLOADING THE DEVICE.
10-12-2018 | Injury | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PNEUMONECTOMY, WHILE THE DEVICE WAS USED ON BRONCHUS, HANDLE MADE A CRACKING SOUND AND THE STAPLE LINE HAD SEPARATED. SURGEON OVERSEW THE STAPLE LINE TO SECURE THE BRONCHUS. TISSUE DAMAGE OCCURRED SINCE THERE WAS TISSUE SEPARATION AT THE STAPLE LINE. THERE WAS ALSO A PROBLEM UNLOADING THE DEVICE.
10-12-2018 | Injury | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PNEUMONECTOMY, WHILE THE DEVICE WAS USED ON BRONCHUS, HANDLE MADE A CRACKING SOUND AND THE STAPLE LINE HAD SEPARATED. SURGEON OVERSEW THE STAPLE LINE TO SECURE THE BRONCHUS. TISSUE DAMAGE OCCURRED SINCE THERE WAS TISSUE SEPARATION AT THE STAPLE LINE. THERE WAS ALSO A PROBLEM UNLOADING THE DEVICE.
10-12-2018 | Injury | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PNEUMONECTOMY, WHILE THE DEVICE WAS USED ON BRONCHUS, HANDLE MADE A CRACKING SOUND AND THE STAPLE LINE HAD SEPARATED. SURGEON OVERSEW THE STAPLE LINE TO SECURE THE BRONCHUS. TISSUE DAMAGE OCCURRED SINCE THERE WAS TISSUE SEPARATION AT THE STAPLE LINE. THERE WAS ALSO A PROBLEM UNLOADING THE DEVICE.
11-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, WHEN THE DEVICE WAS BEING APPLIED TO THE GASTRIC TISSUE, THE SURGEON TRIED TO FIRE THE STAPLER BUT COULD NOT SQUEEZE THE RING HANDLE. IT WAS NOTED THAT THE STAPLER WAS NOT ABLE TO FIRE AND THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. UPON CLOSER INSPECTION, IT WAS OBSERVED THAT THE LOAD HAD BEEN PREFIRED. AT NO TIME WAS IT OBSERVED BEING ACCIDENTALLY FIRED. THE GUN WAS USED FOR ALL OTHER FIRINGS WITHOUT INCIDENT. ANOTHER LOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, WHEN THE DEVICE WAS BEING APPLIED TO THE GASTRIC TISSUE, THE SURGEON TRIED TO FIRE THE STAPLER BUT COULD NOT SQUEEZE THE RING HANDLE. IT WAS NOTED THAT THE STAPLER WAS NOT ABLE TO FIRE AND THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. UPON CLOSER INSPECTION, IT WAS OBSERVED THAT THE LOAD HAD BEEN PREFIRED. AT NO TIME WAS IT OBSERVED BEING ACCIDENTALLY FIRED. THE GUN WAS USED FOR ALL OTHER FIRINGS WITHOUT INCIDENT. ANOTHER LOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, WHEN THE DEVICE WAS BEING APPLIED TO THE GASTRIC TISSUE, THE SURGEON TRIED TO FIRE THE STAPLER BUT COULD NOT SQUEEZE THE RING HANDLE. IT WAS NOTED THAT THE STAPLER WAS NOT ABLE TO FIRE AND THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. UPON CLOSER INSPECTION, IT WAS OBSERVED THAT THE LOAD HAD BEEN PREFIRED. AT NO TIME WAS IT OBSERVED BEING ACCIDENTALLY FIRED. THE GUN WAS USED FOR ALL OTHER FIRINGS WITHOUT INCIDENT. ANOTHER LOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 08-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE IN THE STOMACH THE DEVICE MADE A PARTIAL SHOT AND THE STAPLE LINE WAS INCOMPLETE, IN THE DISTAL PART. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 08-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE IN THE STOMACH THE DEVICE MADE A PARTIAL SHOT AND THE STAPLE LINE WAS INCOMPLETE, IN THE DISTAL PART. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-11-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Gas/Air Leak | A050402 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION, THERE WAS BUBBLING AT THE ANASTOMOSIS. THERE WERE HEALTHY DONUTS, STAPLE LINE WAS COMPLETE, BUT WHEN THEY DID THE LEAK CHECK, THERE WAS BUBBLING AT THE ANASTOMOSIS. TO RESOLVE THE ISSUE, THEY OVERSEW THE STAPLE LINE/ANASTOMOSIS. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
20-11-2018 | Injury | us surgical puerto rico | 08-01-2019 | EEA | Gas/Air Leak | A050402 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION, THERE WAS BUBBLING AT THE ANASTOMOSIS. THERE WERE HEALTHY DONUTS, STAPLE LINE WAS COMPLETE, BUT WHEN THEY DID THE LEAK CHECK, THERE WAS BUBBLING AT THE ANASTOMOSIS. TO RESOLVE THE ISSUE, THEY OVERSEW THE STAPLE LINE/ANASTOMOSIS. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
07-01-2019 | Malfunction | ethicon endo-surgery | 08-01-2019 | ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER; ENDOPATH ETS45 2.5MM 45MM VASCULAR/ THIN RELOAD CARTRIDGE | Failure to Fire | A050501 | No Information | Event Description: WHILE USING ETHICON ENDO-SURGERY, ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER PER MANUFACTURER'S INSTRUCTIONS, IT WAS NOTED THAT THE FIRST FIRING OF THE DEVICE WAS UNEVENTFUL, HOWEVER WHEN ATTEMPTING TO LOAD A SUBSEQUENT ETHICON ENDO-SURGERY ENDOPATH ETS45 2.5 MM 45 MM VASCULAR/THIN CARTRIDGE, TR45W, THE STAPLING DEVICE WOULD NOT SECURE THE NEW CARTRIDGE. A NEW "LIKE" LINEAR CUTTER WAS THEN OPENED AND UTILIZED FOR THE REMAINDER OF THE PROCEDURE.
07-01-2019 | Malfunction | ethicon endo-surgery | 08-01-2019 | ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER; ENDOPATH ETS45 2.5MM 45MM VASCULAR/ THIN RELOAD CARTRIDGE | Appropriate Term/Code Not Available |  | No Information | Event Description: WHILE USING ETHICON ENDO-SURGERY, ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER PER MANUFACTURER'S INSTRUCTIONS, IT WAS NOTED THAT THE FIRST FIRING OF THE DEVICE WAS UNEVENTFUL, HOWEVER WHEN ATTEMPTING TO LOAD A SUBSEQUENT ETHICON ENDO-SURGERY ENDOPATH ETS45 2.5 MM 45 MM VASCULAR/THIN CARTRIDGE, TR45W, THE STAPLING DEVICE WOULD NOT SECURE THE NEW CARTRIDGE. A NEW "LIKE" LINEAR CUTTER WAS THEN OPENED AND UTILIZED FOR THE REMAINDER OF THE PROCEDURE.
23-10-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION PROCEDURE, THE INSTRUMENT DID NOT FIRE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-10-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | UNKNOWN ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION PROCEDURE, THE INSTRUMENT DID NOT FIRE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE/SEGMENTAL RESECTION, AFTER CHECKING THE OPENING AND CLOSING OF THE JAWS, IT WAS CONFIRMED THAT THE STATUS INDICATORS OF THE HANDLE, THE ADAPTER AND THE CARTRIDGE ILLUMINATED ALL GREEN IN THE DISPLAY SCREEN. THE DEVICE WAS USED FOR THE FIRING HOWEVER THE RELOAD FIRED ONLY THE PROXIMAL SITE. ANOTHER RELOAD WAS USED BUT THE SAME EVENT OCCURRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE/SEGMENTAL RESECTION, AFTER CHECKING THE OPENING AND CLOSING OF THE JAWS, IT WAS CONFIRMED THAT THE STATUS INDICATORS OF THE HANDLE, THE ADAPTER AND THE CARTRIDGE ILLUMINATED ALL GREEN IN THE DISPLAY SCREEN. THE DEVICE WAS USED FOR THE FIRING HOWEVER THE RELOAD FIRED ONLY THE PROXIMAL SITE. ANOTHER RELOAD WAS USED BUT THE SAME EVENT OCCURRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE/SEGMENTAL RESECTION, AFTER CHECKING THE OPENING AND CLOSING OF THE JAWS, IT WAS CONFIRMED THAT THE STATUS INDICATORS OF THE HANDLE, THE ADAPTER AND THE CARTRIDGE ILLUMINATED ALL GREEN IN THE DISPLAY SCREEN. THE DEVICE WAS USED FOR THE FIRING HOWEVER THE RELOAD FIRED ONLY THE PROXIMAL SITE. ANOTHER RELOAD WAS USED BUT THE SAME EVENT OCCURRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE/SEGMENTAL RESECTION, AFTER CHECKING THE OPENING AND CLOSING OF THE JAWS, IT WAS CONFIRMED THAT THE STATUS INDICATORS OF THE HANDLE, THE ADAPTER AND THE CARTRIDGE ILLUMINATED ALL GREEN IN THE DISPLAY SCREEN. THE DEVICE WAS USED FOR THE FIRING HOWEVER THE RELOAD FIRED ONLY THE PROXIMAL SITE. ANOTHER RELOAD WAS USED BUT THE SAME EVENT OCCURRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE/SEGMENTAL RESECTION, AFTER CHECKING THE OPENING AND CLOSING OF THE JAWS, IT WAS CONFIRMED THAT THE STATUS INDICATORS OF THE HANDLE, THE ADAPTER AND THE CARTRIDGE ILLUMINATED ALL GREEN IN THE DISPLAY SCREEN. THE DEVICE WAS USED FOR THE FIRING HOWEVER THE RELOAD FIRED ONLY THE PROXIMAL SITE. ANOTHER RELOAD WAS USED BUT THE SAME EVENT OCCURRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE/SEGMENTAL RESECTION, AFTER CHECKING THE OPENING AND CLOSING OF THE JAWS, IT WAS CONFIRMED THAT THE STATUS INDICATORS OF THE HANDLE, THE ADAPTER AND THE CARTRIDGE ILLUMINATED ALL GREEN IN THE DISPLAY SCREEN. THE DEVICE WAS USED FOR THE FIRING HOWEVER THE RELOAD FIRED ONLY THE PROXIMAL SITE. ANOTHER RELOAD WAS USED BUT THE SAME EVENT OCCURRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE SURGEON WAS USING THE THIRD CARTRIDGE TO FIRE DURING AN ABDOMINAL WALL INCISIONAL HERNIA, A CRACK SOUND WAS HEARD BUT FIRING WAS COMPLETED. WHEN CONNECTING THE FOURTH CARTRIDGE, THE CARTRIDGE COULD NOT BE CONNECTED BECAUSE THE PIN OF THE CONNECTION PART OF THE CARTRIDGE AND HANDLE DID NOT BECOME VERTICAL. IT WAS REPORTED THAT THERE WAS NO PROBLEM WITH THE PATIENT AND THAT THE DOCTOR MAY HAVE APPLIED TOO MUCH TENSION TO THE TIP OF THE TACKER. IT WAS STATED THAT SINCE THE DOCTOR¿S HANDS ARE SMALL AND MAY HAVE SQUEEZED THE HANDLE TOO STRONGLY. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | RELIATACK | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE SURGEON WAS USING THE THIRD CARTRIDGE TO FIRE DURING AN ABDOMINAL WALL INCISIONAL HERNIA, A CRACK SOUND WAS HEARD BUT FIRING WAS COMPLETED. WHEN CONNECTING THE FOURTH CARTRIDGE, THE CARTRIDGE COULD NOT BE CONNECTED BECAUSE THE PIN OF THE CONNECTION PART OF THE CARTRIDGE AND HANDLE DID NOT BECOME VERTICAL. IT WAS REPORTED THAT THERE WAS NO PROBLEM WITH THE PATIENT AND THAT THE DOCTOR MAY HAVE APPLIED TOO MUCH TENSION TO THE TIP OF THE TACKER. IT WAS STATED THAT SINCE THE DOCTOR¿S HANDS ARE SMALL AND MAY HAVE SQUEEZED THE HANDLE TOO STRONGLY. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
06-08-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE WAS IMPOSSIBLE TO STAPLE AS THE STAPLER REMAINED BLOCKED. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-08-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE WAS IMPOSSIBLE TO STAPLE AS THE STAPLER REMAINED BLOCKED. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Injury | ethicon endo-surgery | 09-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | Abscess | Event Description: IT WAS REPORTED THAT THE SURGEON REPORTED FOUR SEPARATE INSTANCES OF BLEEDING ON THE APPENDICULAR STUMP DURING LAP APPENDECTOMIES, ONE LEADING TO A POST-OP ABSCESS AND HEMATOMA FORMATION. ALL OF THE INCIDENTS INVOLVED A PVS STAPLER. THE SURGEON ALSO REPORTED STAPLER LOCKOUTS THAT COULD NOT BE FIXED WITH THE USE OF REVERSE BUTTON. ALL OF THE CASES INVOLVED PEDIATRIC PATIENTS. CAUTERY AND PRESSURE WERE USED TO STOP BLEEDING.
01-01-2018 | Injury | ethicon endo-surgery | 09-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | Hematoma | Event Description: IT WAS REPORTED THAT THE SURGEON REPORTED FOUR SEPARATE INSTANCES OF BLEEDING ON THE APPENDICULAR STUMP DURING LAP APPENDECTOMIES, ONE LEADING TO A POST-OP ABSCESS AND HEMATOMA FORMATION. ALL OF THE INCIDENTS INVOLVED A PVS STAPLER. THE SURGEON ALSO REPORTED STAPLER LOCKOUTS THAT COULD NOT BE FIXED WITH THE USE OF REVERSE BUTTON. ALL OF THE CASES INVOLVED PEDIATRIC PATIENTS. CAUTERY AND PRESSURE WERE USED TO STOP BLEEDING.
01-01-2018 | Injury | ethicon endo-surgery | 09-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | Blood Loss | Event Description: IT WAS REPORTED THAT THE SURGEON REPORTED FOUR SEPARATE INSTANCES OF BLEEDING ON THE APPENDICULAR STUMP DURING LAP APPENDECTOMIES, ONE LEADING TO A POST-OP ABSCESS AND HEMATOMA FORMATION. ALL OF THE INCIDENTS INVOLVED A PVS STAPLER. THE SURGEON ALSO REPORTED STAPLER LOCKOUTS THAT COULD NOT BE FIXED WITH THE USE OF REVERSE BUTTON. ALL OF THE CASES INVOLVED PEDIATRIC PATIENTS. CAUTERY AND PRESSURE WERE USED TO STOP BLEEDING.
01-01-2018 | Injury | ethicon endo-surgery | 09-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON REPORTED FOUR SEPARATE INSTANCES OF BLEEDING ON THE APPENDICULAR STUMP DURING LAP APPENDECTOMIES, ONE LEADING TO A POST-OP ABSCESS AND HEMATOMA FORMATION. ALL OF THE INCIDENTS INVOLVED A PVS STAPLER. THE SURGEON ALSO REPORTED STAPLER LOCKOUTS THAT COULD NOT BE FIXED WITH THE USE OF REVERSE BUTTON. ALL OF THE CASES INVOLVED PEDIATRIC PATIENTS. CAUTERY AND PRESSURE WERE USED TO STOP BLEEDING.
08-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE DEVICE STOPPED FIRING THE REINFORCED RELOAD AFTER ONCE CENTIMETER. THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. THE STAPLE LINE WAS FIXED NORMALLY USING NEW RELOAD. THEY CUT THE REINFORCEMENT RELOAD AND THEY USED MANUAL STAPLER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
08-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE DEVICE STOPPED FIRING THE REINFORCED RELOAD AFTER ONCE CENTIMETER. THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. THE STAPLE LINE WAS FIXED NORMALLY USING NEW RELOAD. THEY CUT THE REINFORCEMENT RELOAD AND THEY USED MANUAL STAPLER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
08-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE DEVICE STOPPED FIRING THE REINFORCED RELOAD AFTER ONCE CENTIMETER. THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. THE STAPLE LINE WAS FIXED NORMALLY USING NEW RELOAD. THEY CUT THE REINFORCEMENT RELOAD AND THEY USED MANUAL STAPLER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
08-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE DEVICE STOPPED FIRING THE REINFORCED RELOAD AFTER ONCE CENTIMETER. THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. THE STAPLE LINE WAS FIXED NORMALLY USING NEW RELOAD. THEY CUT THE REINFORCEMENT RELOAD AND THEY USED MANUAL STAPLER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
08-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | IDRIVE ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE DEVICE STOPPED FIRING THE REINFORCED RELOAD AFTER ONCE CENTIMETER. THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. THE STAPLE LINE WAS FIXED NORMALLY USING NEW RELOAD. THEY CUT THE REINFORCEMENT RELOAD AND THEY USED MANUAL STAPLER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
08-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | IDRIVE ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE DEVICE STOPPED FIRING THE REINFORCED RELOAD AFTER ONCE CENTIMETER. THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. THE STAPLE LINE WAS FIXED NORMALLY USING NEW RELOAD. THEY CUT THE REINFORCEMENT RELOAD AND THEY USED MANUAL STAPLER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
08-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE DEVICE STOPPED FIRING THE REINFORCED RELOAD AFTER ONCE CENTIMETER. THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. THE STAPLE LINE WAS FIXED NORMALLY USING NEW RELOAD. THEY CUT THE REINFORCEMENT RELOAD AND THEY USED MANUAL STAPLER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
08-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE STOMACH, THE DEVICE STOPPED FIRING THE REINFORCED RELOAD AFTER ONCE CENTIMETER. THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. THE STAPLE LINE WAS FIXED NORMALLY USING NEW RELOAD. THEY CUT THE REINFORCEMENT RELOAD AND THEY USED MANUAL STAPLER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
06-07-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY PROCEDURE, THE DEVICE DID NOT CLAMP, AND IT WAS ALSO DIFFICULT TO UNLOAD. THERE WAS NO PATIENT INJURY.
06-07-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY PROCEDURE, THE DEVICE DID NOT CLAMP, AND IT WAS ALSO DIFFICULT TO UNLOAD. THERE WAS NO PATIENT INJURY.
06-07-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | UNKNOWN ENDO GI | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY PROCEDURE, THE DEVICE DID NOT CLAMP, AND IT WAS ALSO DIFFICULT TO UNLOAD. THERE WAS NO PATIENT INJURY.
06-07-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | UNKNOWN ENDO GI | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY PROCEDURE, THE DEVICE DID NOT CLAMP, AND IT WAS ALSO DIFFICULT TO UNLOAD. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC LOW ANTERIOR RESECTION, THE DEVICE WOULDN'T FIRE EVEN IF THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY USED ANOTHER RELOAD TO COMPLETE THE CASE. NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC LOW ANTERIOR RESECTION, THE DEVICE WOULDN'T FIRE EVEN IF THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY USED ANOTHER RELOAD TO COMPLETE THE CASE. NO PATIENT INJURY.
03-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE TRANSECTING THE COLON DURING A LAPAROSCOPIC SIGMOID RESECTION, THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THEY HAD TO OPEN THE JAWS AND PUT ANOTHER RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE IS STAYED JAMMED ON THE COLON, IT WAS IMPOSSIBLE TO STAPLE NOR OPEN IT. THE SURGEON HAS USED ANOTHER STAPLER TO CUT AROUND THE DEVICE TO FREE IT. NO PATIENT CONSEQUENCE. THE TIME OF PROCEDURE WAS EXTENDED OF 15 MINS.
12-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, 2 DAYS FOLLOWING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS PROCEDURE, THE PATIENT PULLED AN 8-10 INCHES LONG SUTURE OUT OF HER ESOPHAGUS. THE SURGEON PLACED AN ADDITIONAL RETENTION SUTURE ONTO THE ANVIL OF THE ORVIL. WHEN SCOPING THE PATIENT INTRAOPERATIVELY, THE RESIDENT NOTICED SUTURE PRESENT IN THE ESOPHAGUS. HOWEVER, SHE WAS UNABLE TO RETAIN THE SUTURE DUE TO THE DISTANCE. 2 DAYS POST OP THE PATIENT WAS VOMITING AND FELT SOMETHING IN HER THROAT. THE PATIENT REACHED INTO HER THROAT AND PULLED OUT A STRAND OF SUTURE. THE ANASTOMOSIS IS IN TACT. IT WAS UNSURE IF THE SUTURE PULLED FROM THE PATIENT'S ESOPHAGUS CAME FROM THE ORVIL OR FROM THE SURGEONS SUTURE STRING THE SURGEON APPLIED TO THE ANVIL. THE ISSUE WAS NOT RELATED TO STAPLER. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE PROCEDURE, DURING LAPAROSCOPIC COLONIC ANASTOMOSIS, WHEN STAPLING THE COLONIC TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. RELOAD AND HANDLE WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. THE DEVICE WAS REMOVED AND REBOOTED AND WAS DETERMINED THAT THE ADAPTER WAS NOT FUNCTIONING. A MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
12-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. THE DEVICE WAS REMOVED AND REBOOTED AND WAS DETERMINED THAT THE ADAPTER WAS NOT FUNCTIONING. A MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
12-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE PROCEDURE, DURING LAPAROSCOPIC COLONIC ANASTOMOSIS, WHEN STAPLING THE COLONIC TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. RELOAD AND HANDLE WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
06-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
06-11-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
16-12-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A HIGH ANTERIOR RESECTION PROCEDURE, IT WAS FOUND THAT SOME STAPLES WERE UNFORMED WHEN THE TARGET TISSUE WAS CHECKED WITH ENDOSCOPE. THE DONUTS WERE CREATED PROPERLY. ADDITIONAL SUTURE WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. THE DEVICE WAS REMOVED AND REBOOTED AND WAS DETERMINED THAT THE ADAPTER WAS NOT FUNCTIONING. A MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
13-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORASCOPIC RIGHT UPPER LOBECTOMY PROCEDURE, THE GREEN BUTTON WAS PRESSED BUT WHEN THE HANDLE WAS BEING SQUEEZED, IT COULD NOT BE SQUEEZED FIRMLY. THE BLACK RETURN KNOB WAS PULLED AND THE DEVICE WAS RELEASED FROM THE TISSUE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Injury | us surgical puerto rico | 09-01-2019 | TA PREMIUM | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOLLOWING AN OPEN LOW ANTERIOR RESECTION, A SMALL BLEEDING WAS NOTICED. IT WAS UNSURE IF THE BLEEDING WAS FROM THE STAPLE LINE. THEY WANTED TO CHECK SOMEONE ELSE HAD SAME PROBLEM WITH THE SAME RELOAD AND SAME LOT NUMBER. DRAINAGE WAS DONE TO RESOLVE THE ISSUE.
01-11-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | GOLD RELOAD FOR ECHELON 60 | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON WAS PERFORMING A LAP GASTRIC SLEEVE, AND HE POSITIONED THE POWERED ECHELON TO MAKE THE FIRST SHOT, AFTER 15 SECONDS OF PRECOMPRESSION HE FIRED WITH MAINTAINING THE TRIGGER PUSHED AND THE BLADE STOPPED ADVANCING, HE RELEASED THE TRIGGER AND THE BLADE GOT STUCKED AND THE COULDNT OPEN THE JAWS, HE USED THE REVERSE BUTTON AND THE BLADE WENT TO ITS ORIGINAL POSITION, HE CHANGED THE CARTRIDGE AND HE REMAKE THE SHOT. THE PROCEDURE HAS GONE ON SMOOTHLY.
17-12-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | ILS 25MM, CURVED | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE ANVIL WAS DETACHED FROM THE TROCAR WHILE THE ADJUSTING KNOB WAS TIGHTENED, AND THE ANVIL COULD NOT BE ATTACHED TO THE TROCAR. THE DEVICE WAS USED BETWEEN THE COLON AND THE RECTUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
07-12-2018 | Death | us surgical puerto rico | 09-01-2019 | UNKNOWN ENDO GI | Failure to Form Staple | A050704 | Death | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE TRANSECTING THE DUODENUM, THE DISTAL STAPLE LINE WAS INCOMPLETE. THE PROBLEM RESULTED TO DUODENAL FISTULA WHICH HAD REQUIRED REOPERATION, BILIARY DRAINAGE, AND SUTURE OF THE DUODENAL HOLE TO BE RESOLVED. THE PATIENT'S INCISION WAS EXTENDED FOR MORE THAN AN INCH. THE HOSPITALIZATION OF THE PATIENT WAS ALSO EXTENDED. IT WAS NOTED THAT THERE WILL BE AN ADDITIONAL OPERATION TO CORRECT THE ISSUE.
07-12-2018 | Death | us surgical puerto rico | 09-01-2019 | UNKNOWN ENDO GI | Failure to Form Staple | A050704 | Fistula | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE TRANSECTING THE DUODENUM, THE DISTAL STAPLE LINE WAS INCOMPLETE. THE PROBLEM RESULTED TO DUODENAL FISTULA WHICH HAD REQUIRED REOPERATION, BILIARY DRAINAGE, AND SUTURE OF THE DUODENAL HOLE TO BE RESOLVED. THE PATIENT'S INCISION WAS EXTENDED FOR MORE THAN AN INCH. THE HOSPITALIZATION OF THE PATIENT WAS ALSO EXTENDED. IT WAS NOTED THAT THERE WILL BE AN ADDITIONAL OPERATION TO CORRECT THE ISSUE.
07-12-2018 | Death | us surgical puerto rico | 09-01-2019 | UNKNOWN ENDO GI | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE TRANSECTING THE DUODENUM, THE DISTAL STAPLE LINE WAS INCOMPLETE. THE PROBLEM RESULTED TO DUODENAL FISTULA WHICH HAD REQUIRED REOPERATION, BILIARY DRAINAGE, AND SUTURE OF THE DUODENAL HOLE TO BE RESOLVED. THE PATIENT'S INCISION WAS EXTENDED FOR MORE THAN AN INCH. THE HOSPITALIZATION OF THE PATIENT WAS ALSO EXTENDED. IT WAS NOTED THAT THERE WILL BE AN ADDITIONAL OPERATION TO CORRECT THE ISSUE.
07-12-2018 | Death | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | Death | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE TRANSECTING THE DUODENUM, THE DISTAL STAPLE LINE WAS INCOMPLETE. THE PROBLEM RESULTED TO DUODENAL FISTULA WHICH HAD REQUIRED REOPERATION, BILIARY DRAINAGE, AND SUTURE OF THE DUODENAL HOLE TO BE RESOLVED. THE PATIENT'S INCISION WAS EXTENDED FOR MORE THAN AN INCH. THE HOSPITALIZATION OF THE PATIENT WAS ALSO EXTENDED. IT WAS NOTED THAT THERE WILL BE AN ADDITIONAL OPERATION TO CORRECT THE ISSUE.
07-12-2018 | Death | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | Fistula | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE TRANSECTING THE DUODENUM, THE DISTAL STAPLE LINE WAS INCOMPLETE. THE PROBLEM RESULTED TO DUODENAL FISTULA WHICH HAD REQUIRED REOPERATION, BILIARY DRAINAGE, AND SUTURE OF THE DUODENAL HOLE TO BE RESOLVED. THE PATIENT'S INCISION WAS EXTENDED FOR MORE THAN AN INCH. THE HOSPITALIZATION OF THE PATIENT WAS ALSO EXTENDED. IT WAS NOTED THAT THERE WILL BE AN ADDITIONAL OPERATION TO CORRECT THE ISSUE.
07-12-2018 | Death | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE TRANSECTING THE DUODENUM, THE DISTAL STAPLE LINE WAS INCOMPLETE. THE PROBLEM RESULTED TO DUODENAL FISTULA WHICH HAD REQUIRED REOPERATION, BILIARY DRAINAGE, AND SUTURE OF THE DUODENAL HOLE TO BE RESOLVED. THE PATIENT'S INCISION WAS EXTENDED FOR MORE THAN AN INCH. THE HOSPITALIZATION OF THE PATIENT WAS ALSO EXTENDED. IT WAS NOTED THAT THERE WILL BE AN ADDITIONAL OPERATION TO CORRECT THE ISSUE.
13-12-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORASCOPIC RIGHT UPPER LOBECTOMY PROCEDURE, THE GREEN BUTTON WAS PRESSED BUT WHEN THE HANDLE WAS BEING SQUEEZED, IT COULD NOT BE SQUEEZED FIRMLY. THE BLACK RETURN KNOB WAS PULLED AND THE DEVICE WAS RELEASED FROM THE TISSUE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-10-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, ON THE THIRD FIRING, THE STAPLER STOPPED AND CLICKED, AND STAPLING CAN NOT BE PROGRAMMED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
22-10-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, ON THE THIRD FIRING, THE STAPLER STOPPED AND CLICKED, AND STAPLING CAN NOT BE PROGRAMMED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
21-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | UNK-SIGADAPT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORASCOPIC LOBECTOMY PROCEDURE, THE DEVICE WOULD NOT FIRE. THERE WAS NO PATIENT HARM.
22-10-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, ON THE THIRD FIRING, THE STAPLER STOPPED AND CLICKED, AND STAPLING CAN NOT BE PROGRAMMED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
22-10-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, ON THE THIRD FIRING, THE STAPLER STOPPED AND CLICKED, AND STAPLING CAN NOT BE PROGRAMMED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
10-09-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY PROCEDURE, THE CARTRIDGE STOPPED HALFWAY WHILE FIRING. THE SURGEON HAD TO OPEN TWO HANDLES SINCE THE CARTRIDGES WERE STUCK AND COULD NOT BE UNLOADED. THERE WAS NO PATIENT INJURY.
13-12-2018 | Injury | us surgical puerto rico | 09-01-2019 | TA PREMIUM | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOLLOWING AN OPEN LOW ANTERIOR RESECTION PROCEDURE, SMALL BLEEDING WAS OBSERVED BUT, IT WAS NOT SURE IF IT CAME FROM THE STAPLE LINE. A DRAINAGE WAS DONE TO RESOLVE THE ISSUE.
26-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE LUNG TISSUE, THE JAWS OF THE DEVICE LOCKED ON TISSUE AND THE "BUNNY EARS" OF THE STAPLER WERE PULLED BACK AND RELOAD OPENED TO RELEASE TISSUE. THE SURGEON CLAMPED DOWN ON TISSUE, DEPRESSED GREEN BUTTON, WENT TO SQUEEZE HANDLE TO FIRE BUT COULD NOT, GUN WAS LOCKED INTO POSITION. A NEW STAPLER AND RELOAD WAS OPENED TO COMPLETE THE CASE. NO INJURY.
26-11-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE LUNG TISSUE, THE JAWS OF THE DEVICE LOCKED ON TISSUE AND THE "BUNNY EARS" OF THE STAPLER WERE PULLED BACK AND RELOAD OPENED TO RELEASE TISSUE. THE SURGEON CLAMPED DOWN ON TISSUE, DEPRESSED GREEN BUTTON, WENT TO SQUEEZE HANDLE TO FIRE BUT COULD NOT, GUN WAS LOCKED INTO POSITION. A NEW STAPLER AND RELOAD WAS OPENED TO COMPLETE THE CASE. NO INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SIGMOIDECTOMY, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS NOTICED THAT PRE-FIRING HAD OCCURRED AND THE DEVICE WAS LOADED INCORRECTLY. A NEW RELOAD WAS OPENED AND THE PROCEDURE WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SIGMOIDECTOMY, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS NOTICED THAT PRE-FIRING HAD OCCURRED AND THE DEVICE WAS LOADED INCORRECTLY. A NEW RELOAD WAS OPENED AND THE PROCEDURE WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SIGMOIDECTOMY, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS NOTICED THAT PRE-FIRING HAD OCCURRED AND THE DEVICE WAS LOADED INCORRECTLY. A NEW RELOAD WAS OPENED AND THE PROCEDURE WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, EVEN THOUGH THE JAWS WERE ABLE TO BE OPENED AND CLOSED, THE DEVICE WAS UNABLE TO BE FIRED. A MANUAL HANDLE WAS USED. NO INJURY.
12-12-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, EVEN THOUGH THE JAWS WERE ABLE TO BE OPENED AND CLOSED, THE DEVICE WAS UNABLE TO BE FIRED. A MANUAL HANDLE WAS USED. NO INJURY.
13-12-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LIVER RESECTION PROCEDURE, THE DEVICE WOULD NOT RELEASE TISSUE. JAWS WOULD NOT OPEN. HAD TO CUT TISSUE OUT OF JAWS. NO AE TO THE PATIENT, NO DELAY IN THE SURGERY.
13-11-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE¿S LOAD DID NOT SHOOT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-11-2018 | Malfunction | us surgical puerto rico | 09-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE¿S LOAD DID NOT SHOOT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Device Dislodged or Dislocated | A051201 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LIVER RESECTION THE DEVICE WAS FIRED AND IT JAMMED. THE SLED WAS DISLODGED. THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED. THERE IS NO ADDITIONAL INFORMATION.
10-12-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LIVER RESECTION THE DEVICE WAS FIRED AND IT JAMMED. THE SLED WAS DISLODGED. THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED. THERE IS NO ADDITIONAL INFORMATION.
13-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ABSORBATACK | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA PROCEDURE, AFTER THE FIFTH SHOT, THE SURGEON NOTICED A GREAT RESISTANCE IN THE TACKER, SOON IDENTIFIED THAT IT WAS BROKEN. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ABSORBATACK | Failure to Advance | A150204 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA PROCEDURE, AFTER THE FIFTH SHOT, THE SURGEON NOTICED A GREAT RESISTANCE IN THE TACKER, SOON IDENTIFIED THAT IT WAS BROKEN. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 09-01-2019 | ABSORBATACK | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA PROCEDURE, AFTER THE FIFTH SHOT, THE SURGEON NOTICED A GREAT RESISTANCE IN THE TACKER, SOON IDENTIFIED THAT IT WAS BROKEN. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
15-11-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | ILS 29MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY THAT THE DEVICE ¿MIS-FIRED¿. THE ANVIL WAS PLACED IN THE PATIENT, THEN OUTSIDE OF THE PATIENT THE THIRD-YEAR RESIDENT REMOVED THE SAFETY AND FIRED THE DEVICE PRIOR TO INSERTING INTO THE PATIENT AND CONNECTING TO THE ANVIL TO PERFORM A ¿FORCE TO FIRE¿ TEST AND THE STAPLES ¿FELL OUT¿. A CDH29 WAS THEN OPENED AND THE ANVIL FROM THE ECS29A WAS USED TO FORM THE ANASTOMOSIS. THE DONUTS WERE INTACT AND LOOKED GOOD. A LEAK TEST WAS PERFORMED, AND A LEAK WAS DETECTED BUT STAFF WAS UNABLE TO IDENTIFY THE LOCATION OF THE LEAK. STAFF THEN PULLED ANOTHER CDH29 AND RESECTED THE AREA WHERE THE ANASTOMOSIS WAS CREATED LOWERED THE COLON A LITTLE MORE AND FIRED THE CDH29 CREATING A NEW ANASTOMOSIS. THE DONUTS WERE INTACT AND LOOKED GOOD. A LEAK CHECK WAS DONE, AND NO LEAKS WERE DETECTED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. THE SURGEON AND THE SALES REP LOOKED AT THE SPECIMEN FROM THE INITIAL FIRING AND IT WAS THE SURGEON¿S CONCLUSION FOR THE REASON OF THE LEAK WAS A POORLY FORMED PURSE STRING, MORE SPECIFICALLY THERE WAS WHAT APPEARED TO BE A 3MM GAP IN THE PURSE STRING THAT CREATED THE LEAK.
12-12-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THORACOSCOPIC COMBINED WITH RADICAL ESOPHAGECTOMY FOR ESOPHAGEAL CANCER, AFTER FIRED WHEN SEVERING THE VESSEL, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ECR60D, EVENT RE-HAPPENED. SUTURE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-12-2018 | Malfunction | ethicon endo-surgery | 09-01-2019 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THORACOSCOPIC COMBINED WITH RADICAL ESOPHAGECTOMY FOR ESOPHAGEAL CANCER, AFTER FIRED WHEN SEVERING THE VESSEL, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ECR60D, EVENT RE-HAPPENED. SUTURE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2017 | Malfunction | ethicon endo-surgery | 10-01-2019 | FLEX 60 ARTICULATING | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT WORK. NO OTHER INFORMATION IS AVAILABLE. NO PATIENT CONSEQUENCE REPORTED.
12-12-2018 | Injury | ethicon endo-surgery | 10-01-2019 | PROXIMATE*PPH PROCEDURE SET | Failure to Fire | A050501 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY "PRCEDURE", STAPLER DID NOT FIRE BETWEEN 7 AND 9 O'CLOCK. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED.
13-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | GIA | Failure to Cut | A050702 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING BARIATRIC SURGERY, IN THE EXCHANGE OF THE DEVICE IN THE THIRD USE, THERE WAS PARTIAL STAPLING WITHOUT CUTTING. THE SURGEON NEEDED TO SUBSTITUTE THE LOAD IN ORDER TO RUN THE STAPLER NORMALLY. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PANCREATO DUODENECTOMIA WITH LYMPHADENECTOMY, WHEN PERFORMING THE FIRST STAPLING, THE SURGEON FELT RESISTANCE AND THE STAPLER PRESENTED A DIFFERENT SOUND FROM THE NORMAL ONE WHEN BEING TRIGGERED. AT THE SECOND STAPLING, THE STAPLER NO LONGER WORKED. THE LOAD WAS CHANGED IN THE ATTEMPT TO PERFORM THE STAPLING, BUT WITHOUT SUCCESS. SINCE IT WAS VERIFIED THAT THE PROBLEM WAS IN THE STAPLER AND NOT IN THE LOAD, THE STAPLER WAS REPLACED BY ONE THAT WAS IN THE CONSIGNMENT SO THAT THE SURGEON COULD CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PANCREATO DUODENECTOMIA WITH LYMPHADENECTOMY, WHEN PERFORMING THE FIRST STAPLING, THE SURGEON FELT RESISTANCE AND THE STAPLER PRESENTED A DIFFERENT SOUND FROM THE NORMAL ONE WHEN BEING TRIGGERED. AT THE SECOND STAPLING, THE STAPLER NO LONGER WORKED. THE LOAD WAS CHANGED IN THE ATTEMPT TO PERFORM THE STAPLING, BUT WITHOUT SUCCESS. SINCE IT WAS VERIFIED THAT THE PROBLEM WAS IN THE STAPLER AND NOT IN THE LOAD, THE STAPLER WAS REPLACED BY ONE THAT WAS IN THE CONSIGNMENT SO THAT THE SURGEON COULD CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PANCREATO DUODENECTOMIA WITH LYMPHADENECTOMY, WHEN PERFORMING THE FIRST STAPLING, THE SURGEON FELT RESISTANCE AND THE STAPLER PRESENTED A DIFFERENT SOUND FROM THE NORMAL ONE WHEN BEING TRIGGERED. AT THE SECOND STAPLING, THE STAPLER NO LONGER WORKED. THE LOAD WAS CHANGED IN THE ATTEMPT TO PERFORM THE STAPLING, BUT WITHOUT SUCCESS. SINCE IT WAS VERIFIED THAT THE PROBLEM WAS IN THE STAPLER AND NOT IN THE LOAD, THE STAPLER WAS REPLACED BY ONE THAT WAS IN THE CONSIGNMENT SO THAT THE SURGEON COULD CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Injury | covidien lp llc north haven | 10-01-2019 | EEA ORVIL | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A CANCER GASTRECTOMY, ON THE ESOPHAGO-JEJUNAL ANASTOMOSIS, THERE WERE SIZES USED, AFTER STAPLING THE ESOPHAGUS AND CORRECT PLACEMENT IN SITE OF THE CIRCULAR STAPLER, AFTER NUMEROUS ATTEMPTS, NO INSERTION OF THE DEVICE WAS POSSIBLE BECAUSE THE TILT-TOP ANVIL REMAINED IN CLOSED POSITION. AFTER REPEATED ATTEMPTS AND ALSO CHANGING OPERATORS, THE SURGEONS WERE FORCED TO OPEN THE STAPLE LINE ON THE ESOPHAGUS AND POSITION A COMMON ANVIL IN THE STANDARD MANNER CLOSING THE ESOPHAGEAL STUMP WITH A PURSE STRING SUTURE. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM AND THE PATIENT WAS HOSPITALIZED. THEY PERFORMED FURTHER SURGICAL MANEUVERS ON THE ESOPHAGUS AND A NEW ANASTOMOSIS AND THE ANVIL WAS REMOVED AND A NEW ONE WAS USED TO COMPLETE THE CASE. THERE WAS A PATIENT INJURY.
11-12-2018 | Injury | covidien lp llc north haven | 10-01-2019 | EEA ORVIL | Difficult to Remove | A150207 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A CANCER GASTRECTOMY, ON THE ESOPHAGO-JEJUNAL ANASTOMOSIS, THERE WERE SIZES USED, AFTER STAPLING THE ESOPHAGUS AND CORRECT PLACEMENT IN SITE OF THE CIRCULAR STAPLER, AFTER NUMEROUS ATTEMPTS, NO INSERTION OF THE DEVICE WAS POSSIBLE BECAUSE THE TILT-TOP ANVIL REMAINED IN CLOSED POSITION. AFTER REPEATED ATTEMPTS AND ALSO CHANGING OPERATORS, THE SURGEONS WERE FORCED TO OPEN THE STAPLE LINE ON THE ESOPHAGUS AND POSITION A COMMON ANVIL IN THE STANDARD MANNER CLOSING THE ESOPHAGEAL STUMP WITH A PURSE STRING SUTURE. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM AND THE PATIENT WAS HOSPITALIZED. THEY PERFORMED FURTHER SURGICAL MANEUVERS ON THE ESOPHAGUS AND A NEW ANASTOMOSIS AND THE ANVIL WAS REMOVED AND A NEW ONE WAS USED TO COMPLETE THE CASE. THERE WAS A PATIENT INJURY.
11-12-2018 | Injury | covidien lp llc north haven | 10-01-2019 | EEA ORVIL | Difficult to Remove | A150207 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A CANCER GASTRECTOMY, ON THE ESOPHAGO-JEJUNAL ANASTOMOSIS, THERE WERE SIZES USED, AFTER STAPLING THE ESOPHAGUS AND CORRECT PLACEMENT IN SITE OF THE CIRCULAR STAPLER, AFTER NUMEROUS ATTEMPTS, NO INSERTION OF THE DEVICE WAS POSSIBLE BECAUSE THE TILT-TOP ANVIL REMAINED IN CLOSED POSITION. AFTER REPEATED ATTEMPTS AND ALSO CHANGING OPERATORS, THE SURGEONS WERE FORCED TO OPEN THE STAPLE LINE ON THE ESOPHAGUS AND POSITION A COMMON ANVIL IN THE STANDARD MANNER CLOSING THE ESOPHAGEAL STUMP WITH A PURSE STRING SUTURE. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM AND THE PATIENT WAS HOSPITALIZED. THEY PERFORMED FURTHER SURGICAL MANEUVERS ON THE ESOPHAGUS AND A NEW ANASTOMOSIS AND THE ANVIL WAS REMOVED AND A NEW ONE WAS USED TO COMPLETE THE CASE. THERE WAS A PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE GREEN BUTTON WAS PRESSED, BUT THE HANDLE WAS STIFF THAT THE FIRING COULD NOT BE PERFORMED. THE EVENT OCCURRED DURING THE PROCEDURE, BUT THE PRODUCT WAS NOT USED FOR PATIENT. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORASCOPIC LOBECTOMY PROCEDURE, AFTER SUCCESSFUL FIRING OF THE FIRST RELOAD, THE SECOND RELOAD COULD NOT BE FIRED. A THIRD RELOAD WAS USED, HOWEVER, THE SAME ISSUE HAPPENED. THE WHOLE POWER HANDLE SET WAS REPLACED WITH A NEW SET ALONG WITH NEW RELOADS TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. AFTER THE PROCEDURE, THE POWER HANDLE SET WAS CHECKED BY THE SALES REPRESENTATIVE, AND WHEN A TEST RELOAD WAS INSERTED, THE CYCLE TEST DID NOT APPEAR ON THE SCREEN AND DIRECTLY SHOWED SOLID GREEN LIGHT ON THE RELOAD INDICATOR OF THE SCREEN. ADDITIONALLY, AN ERROR SOUND WAS HEARD FROM THE HANDLE. THE HANDLE WAS REBOOTED, AND AFTER WHICH THE HANDLE WAS ABLE TO CYCLE TEST AS NORMAL.
11-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TO UNIPORTAL WEDGE RESECTION, BEFORE STAPLING, WAS ARTICULATING A RELOAD TO THE RIGHT BEFORE INSERTING INTO THE PATIENT, BUT THE STAPLER WOULD NOT ARTICULATE BACK TO THE CENTER AND WAS STUCK. TRIED MULTIPLE TIME TO STRAIGHTEN OUT BUT WOULD NOT. THIS STAPLER DID NOT TOUCH THE PATIENT. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE GREEN BUTTON WAS PRESSED, BUT THE HANDLE WAS STIFF THAT THE FIRING COULD NOT BE PERFORMED. THE EVENT OCCURRED DURING THE PROCEDURE, BUT THE PRODUCT WAS NOT USED FOR PATIENT. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | TA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A TERMINAL ILEUM RESECTION, THE STAPLER WAS FIRED EFFECTIVELY FOR THE FIRST LOAD, A SECOND LOAD WAS RELOADED AND THE STAPLER WAS CLOSED ON THE TISSUE. THE SURGEON WAS UNABLE TO FIRE THE STAPLER. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORASCOPIC LOBECTOMY PROCEDURE, AFTER SUCCESSFUL FIRING OF THE FIRST RELOAD, THE SECOND RELOAD COULD NOT BE FIRED. A THIRD RELOAD WAS USED, HOWEVER, THE SAME ISSUE HAPPENED. THE WHOLE POWER HANDLE SET WAS REPLACED WITH A NEW SET ALONG WITH NEW RELOADS TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. AFTER THE PROCEDURE, THE POWER HANDLE SET WAS CHECKED BY THE SALES REPRESENTATIVE, AND WHEN A TEST RELOAD WAS INSERTED, THE CYCLE TEST DID NOT APPEAR ON THE SCREEN AND DIRECTLY SHOWED SOLID GREEN LIGHT ON THE RELOAD INDICATOR OF THE SCREEN. ADDITIONALLY, AN ERROR SOUND WAS HEARD FROM THE HANDLE. THE HANDLE WAS REBOOTED, AND AFTER WHICH THE HANDLE WAS ABLE TO CYCLE TEST AS NORMAL.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORASCOPIC LOBECTOMY PROCEDURE, AFTER SUCCESSFUL FIRING OF THE FIRST RELOAD, THE SECOND RELOAD COULD NOT BE FIRED. A THIRD RELOAD WAS USED, HOWEVER, THE SAME ISSUE HAPPENED. THE WHOLE POWER HANDLE SET WAS REPLACED WITH A NEW SET ALONG WITH NEW RELOADS TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. AFTER THE PROCEDURE, THE POWER HANDLE SET WAS CHECKED BY THE SALES REPRESENTATIVE, AND WHEN A TEST RELOAD WAS INSERTED, THE CYCLE TEST DID NOT APPEAR ON THE SCREEN AND DIRECTLY SHOWED SOLID GREEN LIGHT ON THE RELOAD INDICATOR OF THE SCREEN. ADDITIONALLY, AN ERROR SOUND WAS HEARD FROM THE HANDLE. THE HANDLE WAS REBOOTED, AND AFTER WHICH THE HANDLE WAS ABLE TO CYCLE TEST AS NORMAL.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORASCOPIC LOBECTOMY PROCEDURE, AFTER SUCCESSFUL FIRING OF THE FIRST RELOAD, THE SECOND RELOAD COULD NOT BE FIRED. A THIRD RELOAD WAS USED, HOWEVER, THE SAME ISSUE HAPPENED. THE WHOLE POWER HANDLE SET WAS REPLACED WITH A NEW SET ALONG WITH NEW RELOADS TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. AFTER THE PROCEDURE, THE POWER HANDLE SET WAS CHECKED BY THE SALES REPRESENTATIVE, AND WHEN A TEST RELOAD WAS INSERTED, THE CYCLE TEST DID NOT APPEAR ON THE SCREEN AND DIRECTLY SHOWED SOLID GREEN LIGHT ON THE RELOAD INDICATOR OF THE SCREEN. ADDITIONALLY, AN ERROR SOUND WAS HEARD FROM THE HANDLE. THE HANDLE WAS REBOOTED, AND AFTER WHICH THE HANDLE WAS ABLE TO CYCLE TEST AS NORMAL.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORASCOPIC LOBECTOMY PROCEDURE, AFTER SUCCESSFUL FIRING OF THE FIRST RELOAD, THE SECOND RELOAD COULD NOT BE FIRED. A THIRD RELOAD WAS USED, HOWEVER, THE SAME ISSUE HAPPENED. THE WHOLE POWER HANDLE SET WAS REPLACED WITH A NEW SET ALONG WITH NEW RELOADS TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. AFTER THE PROCEDURE, THE POWER HANDLE SET WAS CHECKED BY THE SALES REPRESENTATIVE, AND WHEN A TEST RELOAD WAS INSERTED, THE CYCLE TEST DID NOT APPEAR ON THE SCREEN AND DIRECTLY SHOWED SOLID GREEN LIGHT ON THE RELOAD INDICATOR OF THE SCREEN. ADDITIONALLY, AN ERROR SOUND WAS HEARD FROM THE HANDLE. THE HANDLE WAS REBOOTED, AND AFTER WHICH THE HANDLE WAS ABLE TO CYCLE TEST AS NORMAL.
07-12-2018 | Injury | ethicon | 10-01-2019 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Not Applicable | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC INGUINAL HERNIA REPAIR PROCEDURE ON (B)(6) 2018 AND THE ABSORBABLE STRAPS WERE USED. IT WAS ALSO REPORTED THAT BECAUSE THE DIFFICULTY EXPERIENCED WITH THE DEVICE DURING THE PROCEDURE, THE PROCEDURE WAS CONVERTED TO OPEN. THE DOCTOR FINISHED THE SURGERY, FIXING THE SCREEN WITH SUTURES. CURRENTLY, THE PATIENT IS WITHOUT REPORTED COMPLICATIONS. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
11-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | UNKNOWN EGIA SULU | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TO UNIPORTAL WEDGE RESECTION, BEFORE STAPLING, WAS ARTICULATING A RELOAD TO THE RIGHT BEFORE INSERTING INTO THE PATIENT, BUT THE STAPLER WOULD NOT ARTICULATE BACK TO THE CENTER AND WAS STUCK. TRIED MULTIPLE TIME TO STRAIGHTEN OUT BUT WOULD NOT. THIS STAPLER DID NOT TOUCH THE PATIENT. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | UNKNOWN SIGNIA EGIA ADAPTER | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TO UNIPORTAL WEDGE RESECTION, BEFORE STAPLING, WAS ARTICULATING A RELOAD TO THE RIGHT BEFORE INSERTING INTO THE PATIENT, BUT THE STAPLER WOULD NOT ARTICULATE BACK TO THE CENTER AND WAS STUCK. TRIED MULTIPLE TIME TO STRAIGHTEN OUT BUT WOULD NOT. THIS STAPLER DID NOT TOUCH THE PATIENT. MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | EEA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ROUX-N-Y GASTRIC BYPASS, WHILE INSERTING THE ANVIL THROUGH THE TROCAR SITE, THE ANVIL WAS BROKEN. IT WAS REPORTED THAT THE SURGEON LOADED A SUTURE THROUGH THE HOLES ON THE BASE OF THE ANVIL ON THE BACK TABLE. THE SUTURE/NEEDLE WAS TOO BIG AT FIRST AND THEY USED ANOTHER ONE. THEN, ONCE THE SUTURE WAS LOADED, THE SURGEON USED A CLAMP TO GRASP THE TIP OF THE ANVIL THROUGH THE OPENING AND INSERTED THROUGH THE TROCAR DEFECT. WHEN TRYING TO PUSH IT THROUGH, THE ANVIL POPPED OFF THE CLAMP. WHEN IT WAS REMOVED, THE SURGEON NOTICED THAT A PIECE OF THE ANVIL STEM BROKE OFF. IT WAS NOTED THAT THE ANVIL WAS BENT. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | EEA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ROUX-N-Y GASTRIC BYPASS, WHILE INSERTING THE ANVIL THROUGH THE TROCAR SITE, THE ANVIL WAS BROKEN. IT WAS REPORTED THAT THE SURGEON LOADED A SUTURE THROUGH THE HOLES ON THE BASE OF THE ANVIL ON THE BACK TABLE. THE SUTURE/NEEDLE WAS TOO BIG AT FIRST AND THEY USED ANOTHER ONE. THEN, ONCE THE SUTURE WAS LOADED, THE SURGEON USED A CLAMP TO GRASP THE TIP OF THE ANVIL THROUGH THE OPENING AND INSERTED THROUGH THE TROCAR DEFECT. WHEN TRYING TO PUSH IT THROUGH, THE ANVIL POPPED OFF THE CLAMP. WHEN IT WAS REMOVED, THE SURGEON NOTICED THAT A PIECE OF THE ANVIL STEM BROKE OFF. IT WAS NOTED THAT THE ANVIL WAS BENT. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | EEA | Material Twisted/Bent | A040609 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ROUX-N-Y GASTRIC BYPASS, WHILE INSERTING THE ANVIL THROUGH THE TROCAR SITE, THE ANVIL WAS BROKEN. IT WAS REPORTED THAT THE SURGEON LOADED A SUTURE THROUGH THE HOLES ON THE BASE OF THE ANVIL ON THE BACK TABLE. THE SUTURE/NEEDLE WAS TOO BIG AT FIRST AND THEY USED ANOTHER ONE. THEN, ONCE THE SUTURE WAS LOADED, THE SURGEON USED A CLAMP TO GRASP THE TIP OF THE ANVIL THROUGH THE OPENING AND INSERTED THROUGH THE TROCAR DEFECT. WHEN TRYING TO PUSH IT THROUGH, THE ANVIL POPPED OFF THE CLAMP. WHEN IT WAS REMOVED, THE SURGEON NOTICED THAT A PIECE OF THE ANVIL STEM BROKE OFF. IT WAS NOTED THAT THE ANVIL WAS BENT. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | UNKNOWN TA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A TERMINAL ILEUM RESECTION, THE STAPLER WAS FIRED EFFECTIVELY FOR THE FIRST LOAD, A SECOND LOAD WAS RELOADED AND THE STAPLER WAS CLOSED ON THE TISSUE. THE SURGEON WAS UNABLE TO FIRE THE STAPLER. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-11-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA | Failure to Align | A050701 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER A SUCCESSFUL STAPLING BY FELLOW, ONE OF THE METAL WINGS OF THE LOADING UNIT GOT BENT UPWARDS AFTER STAPLING. UNABLE TO CLOSE THE LOADING UNIT.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ABSORBATACK | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA REPAIR PROCEDURE, WHILE ATTACHING THE MESH TO THE PERITONEUM, THE SHAFT OF THE DEVICE BROKE. A REPLACEMENT PRODUCT WAS USED TO WHICH RESULTED TO A 30 MINUTE DELAY IN THE COMPLETION OF THE PROCEDURE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ABSORBATACK | Break | A0401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA REPAIR PROCEDURE, WHILE ATTACHING THE MESH TO THE PERITONEUM, THE SHAFT OF THE DEVICE BROKE. A REPLACEMENT PRODUCT WAS USED TO WHICH RESULTED TO A 30 MINUTE DELAY IN THE COMPLETION OF THE PROCEDURE. THERE WAS NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED WITH A RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED WITH A RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | UNKNOWN ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED WITH A RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | UNKNOWN EGIA SU | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A WEDGE LAPAROSCOPIC PROCEDURE, WHILE THE SURGEON WAS CUTTING THE TISSUE, THE STAPLE LINE WAS INCOMPLETE DISTALLY. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A WEDGE LAPAROSCOPIC PROCEDURE, WHILE THE SURGEON WAS CUTTING THE TISSUE, THE STAPLE LINE WAS INCOMPLETE DISTALLY. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY WHEN RESECTING THE STOMACH, THE RELOAD OF THE DEVICE DID NOT FIRE, ANOTHER OF THE SAME PRODUCT WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING USED FOR INTERLOBAR RESECTION, AS THE SURGEON APPROACHED THE DEVICE TO THE TISSUE AFTER COMPLETING THE DEVICE'S SETTING, SURGEON CLOSED THE JAWS AND THE INDICATOR ILLUMINATED GREEN, SURGEON PRESSED THE BUTTON ATTEMPTING TO FIRE, BUT THE KNIFE DID NOT MOVE, BEEP SOUND WAS GENERATED AND THE DEVICE WAS NOT ABLE TO BE USED. THE PROCEDURE WAS COMPLETED WITH ANOTHER RELOAD. NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | UNKNOWN ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY WHEN RESECTING THE STOMACH, THE RELOAD OF THE DEVICE DID NOT FIRE, ANOTHER OF THE SAME PRODUCT WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING USED FOR INTERLOBAR RESECTION, AS THE SURGEON APPROACHED THE DEVICE TO THE TISSUE AFTER COMPLETING THE DEVICE'S SETTING, SURGEON CLOSED THE JAWS AND THE INDICATOR ILLUMINATED GREEN, SURGEON PRESSED THE BUTTON ATTEMPTING TO FIRE, BUT THE KNIFE DID NOT MOVE, BEEP SOUND WAS GENERATED AND THE DEVICE WAS NOT ABLE TO BE USED. THE PROCEDURE WAS COMPLETED WITH ANOTHER RELOAD. NO PATIENT INJURY.
14-08-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | FLEX 60 ARTICULATING | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC LEFT HEMIHEPATECTOMY, COULD NOT FIRE WITH EC60A AND ECR60D. ANOTHER TWO RELOADS WERE USED TO CONTINUE, BUT THE EVENT RE-HAPPENED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING USED FOR INTERLOBAR RESECTION, AS THE SURGEON APPROACHED THE DEVICE TO THE TISSUE AFTER COMPLETING THE DEVICE'S SETTING, SURGEON CLOSED THE JAWS AND THE INDICATOR ILLUMINATED GREEN, SURGEON PRESSED THE BUTTON ATTEMPTING TO FIRE, BUT THE KNIFE DID NOT MOVE, BEEP SOUND WAS GENERATED AND THE DEVICE WAS NOT ABLE TO BE USED. THE PROCEDURE WAS COMPLETED WITH ANOTHER RELOAD. NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO / LOBECTOMY PROCEDURE, AT THE FOURTH FIRING, THEY CHECKED THE OPENING AND CLOSING OF THE JAW AND APPROACHED IT TO THE TISSUE. WHEN THEY PRESSED THE FIRING BUTTON, THE FIRING STOPPED. THE INDICATOR SHOWED THE FIRING DISPLAY AS USUAL. AFTER BEING REPLACED WITH A NEW CARTRIDGE, IT WAS USED WITHOUT PROBLEM AND FIRED FOR 3 SHOTS.
12-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO / LOBECTOMY PROCEDURE, AT THE FOURTH FIRING, THEY CHECKED THE OPENING AND CLOSING OF THE JAW AND APPROACHED IT TO THE TISSUE. WHEN THEY PRESSED THE FIRING BUTTON, THE FIRING STOPPED. THE INDICATOR SHOWED THE FIRING DISPLAY AS USUAL. AFTER BEING REPLACED WITH A NEW CARTRIDGE, IT WAS USED WITHOUT PROBLEM AND FIRED FOR 3 SHOTS.
24-08-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY OF WEDGE RESECTION PROCEDURE, WHILE RESECTING A LUNG NODULE, IN VERY THICK TISSUE, THE SURGEON WAS UNABLE TO COMPLETE THE FIRING SEQUENCE ON TWO FIRINGS WITH HANDLES CRACKING THROUGHOUT THE FIRING SEQUENCE. IN ORDER TO COMPLETE THE CASE, ANOTHER STAPLER WAS OPENED AND USED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
14-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | UNKNOWN GIA PRODUCT | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE FIRING KNOB OF THE DEVICE BROKE AND A COMPONENT DISENGAGED BUT DID NOT FALL INTO THE PATIENT'S CAVITY. NO INJURY.
14-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | UNKNOWN GIA PRODUCT | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE FIRING KNOB OF THE DEVICE BROKE AND A COMPONENT DISENGAGED BUT DID NOT FALL INTO THE PATIENT'S CAVITY. NO INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AFTER FIRE WAS COMPLETED, SURGEON PUSHED THE JAW OPEN BUTTON BUT JAW WAS NOT OPENED AT ALL. SURGEON TRIED CUT THE JAW'S SIDE TISSUE SMOOTHLY THEN JAW COULD MOVED OUT FROM THE TISSUE.
14-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON ATTEMPTED TO FIRE BY USING RELOAD WHICH WAS CONNECTED TO THE HANDLE, A STRANGE SOUND WAS HEARD AND SURGEON COULD NOT FIRE THE DEVICE. WHEN THEY ATTACHED THE CARTRIDGE TO THE HANDLE AGAIN, THE DISPLAY SCREEN OF HANDLE WAS GREEN, BUT WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE DISPLAY SCREEN SHOWED THE INDICATOR OF COMPLETED FIRING. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON ATTEMPTED TO FIRE BY USING RELOAD WHICH WAS CONNECTED TO THE HANDLE, A STRANGE SOUND WAS HEARD AND SURGEON COULD NOT FIRE THE DEVICE. WHEN THEY ATTACHED THE CARTRIDGE TO THE HANDLE AGAIN, THE DISPLAY SCREEN OF HANDLE WAS GREEN, BUT WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE DISPLAY SCREEN SHOWED THE INDICATOR OF COMPLETED FIRING. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INJURY.
18-10-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, EVEN THOUGH THE TISSUE WAS CLAMPED, THE HANDLE STOPPED OPERATING. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING HEPATECTOMY, RELOAD WAS CONNECTED WITH THE POWERED STAPLER AND AN ATTEMPT WAS MADE TO PERFORM TREATMENT ON HEPATIC VEIN AND GLISSON TOGETHER. THE FIRING BUTTON WAS PRESSED BUT THE STAPLES DID NOT COME OUT. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Device Displays Incorrect Message | A090201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON ATTEMPTED TO FIRE BY USING RELOAD WHICH WAS CONNECTED TO THE HANDLE, A STRANGE SOUND WAS HEARD AND SURGEON COULD NOT FIRE THE DEVICE. WHEN THEY ATTACHED THE CARTRIDGE TO THE HANDLE AGAIN, THE DISPLAY SCREEN OF HANDLE WAS GREEN, BUT WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE DISPLAY SCREEN SHOWED THE INDICATOR OF COMPLETED FIRING. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON ATTEMPTED TO FIRE BY USING RELOAD WHICH WAS CONNECTED TO THE HANDLE, A STRANGE SOUND WAS HEARD AND SURGEON COULD NOT FIRE THE DEVICE. WHEN THEY ATTACHED THE CARTRIDGE TO THE HANDLE AGAIN, THE DISPLAY SCREEN OF HANDLE WAS GREEN, BUT WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE DISPLAY SCREEN SHOWED THE INDICATOR OF COMPLETED FIRING. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON ATTEMPTED TO FIRE BY USING RELOAD WHICH WAS CONNECTED TO THE HANDLE, A STRANGE SOUND WAS HEARD AND SURGEON COULD NOT FIRE THE DEVICE. WHEN THEY ATTACHED THE CARTRIDGE TO THE HANDLE AGAIN, THE DISPLAY SCREEN OF HANDLE WAS GREEN, BUT WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE DISPLAY SCREEN SHOWED THE INDICATOR OF COMPLETED FIRING. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING HEPATECTOMY, RELOAD WAS CONNECTED WITH THE POWERED STAPLER AND AN ATTEMPT WAS MADE TO PERFORM TREATMENT ON HEPATIC VEIN AND GLISSON TOGETHER. THE FIRING BUTTON WAS PRESSED BUT THE STAPLES DID NOT COME OUT. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN LAPAROSCOPIC SIGMOID RESECTION PROCEDURE THAT THE SURGEON EXPERIENCED ¿STAPLE LINE BREAKDOWN¿ WITH THE FIRST TWO STAPLERS. THERE WERE AIR BUBBLES IN THE MIDDLE OF THE STAPLE LINE. A THIRD DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
25-09-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | GREEN RELOAD FOR ECHELON 45 | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE USED RELOAD BROKE IN TWO (PLASTIC AND METAL PARTS) AT THE MOMENT OF CHANGING IT. NO PATIENT CONSEQUENCE REPORTED.
17-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN LAPAROSCOPIC SIGMOID RESECTION PROCEDURE THAT THE SURGEON EXPERIENCED ¿STAPLE LINE BREAKDOWN¿ WITH THE FIRST TWO STAPLERS. THERE WERE AIR BUBBLES IN THE MIDDLE OF THE STAPLE LINE. A THIRD DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
17-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | CONTOUR CURVED CUTTER STAPLER | Gas/Air Leak | A050402 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN LAPAROSCOPIC SIGMOID RESECTION PROCEDURE THAT THE SURGEON EXPERIENCED ¿STAPLE LINE BREAKDOWN¿ WITH THE FIRST TWO STAPLERS. THERE WERE AIR BUBBLES IN THE MIDDLE OF THE STAPLE LINE. A THIRD DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
12-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING HEPATECTOMY, RELOAD WAS CONNECTED WITH THE POWERED STAPLER AND AN ATTEMPT WAS MADE TO PERFORM TREATMENT ON HEPATIC VEIN AND GLISSON TOGETHER. THE FIRING BUTTON WAS PRESSED BUT THE STAPLES DID NOT COME OUT. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON ATTEMPTED TO FIRE BY USING RELOAD WHICH WAS CONNECTED TO THE HANDLE, A STRANGE SOUND WAS HEARD AND SURGEON COULD NOT FIRE THE DEVICE. WHEN THEY ATTACHED THE CARTRIDGE TO THE HANDLE AGAIN, THE DISPLAY SCREEN OF HANDLE WAS GREEN, BUT WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE DISPLAY SCREEN SHOWED THE INDICATOR OF COMPLETED FIRING. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INJURY.
13-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | POWERED PLUS STANDARD 45 | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER FIRED BUT WOULDN'T RETRACT. UNKNOWN FIRING, UNKNOWN RELOAD, UNKNOWN BUTTRESSING. IT WAS NOT REPORTED HOW THE DEVICE WAS REMOVED FROM TISSUE. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
27-08-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPY BULLECTOMY AND PLEURECTOMY, WHEN RESECTING THE LUNG TISSUE, THE LOADING UNIT WAS LOADED PROPERLY, PLACED ON TISSUE AND THE GREEN BUTTON COULD BE PRESSED PROPERLY. WHEN STARTING FIRING, THE DEVICE MADE A LOUD CRACKING NOISE AND AFTER THAT THE STAPLER COULD NOT FIRE ANYMORE. THE JAW OF THE STAPLER WAS ABLE TO OPEN NORMALLY AND BE REMOVED. A SECOND STAPLE WAS APPLIED BUT WAS NOT ABLE TO FIRE ON TISSUE. THEY UNWRAPPED NEW PRODUCTS AND APPLIED ON TISSUE WHILE HOLDING THE GREEN FIRING BUTTON PUSHED MANUALLY DURING FIRING. THE OPERATING TIME WAS EXTENDED FOR LESS THAN 20 MINUTES BECAUSE OF UNWRAPPING NEW PRODUCTS AND TESTING IT OUTSIDE OF THE PATIENT . THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | ECHELON FLEX 45 STANDARD | Manufacturing, Packaging or Shipping Problem | A02 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN VATS LOBECTOMY PROCEDURE THAT WHEN THE OUTER BOX EC60A (LOT NUMBER R94M8W) WAS OPENED THE DEVICE INSIDE WAS AN EC45A (LOT NUMBER R94K15). ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
01-01-2018 | Injury | ethicon endo-surgery | 10-01-2019 | ENDO ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST OP OF A LAPAROSCOPIC SIGMOID RESECTION PROCEDURE THAT PATIENT DEVELOPED A COLON LEAK. UNKNOWN AS TO HOW MANY DAYS POST OP THE LEAK WAS DISCOVERED. UNKNOWN AS TO HOW THE LEAK WAS REPAIRED. PATIENT CONDITION IS CURRENTLY STABLE.
12-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | UNKNOWN GIA PRODUCT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-02-2018 | Injury | covidien lp llc north haven | 10-01-2019 | UNKNOWN TRI STAPLE PRODUCT | Fluid/Blood Leak | A050401 | No Code Available | Event Description: (B)(4) - ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY SURGERY, THERE WERE 187 PATIENTS THAT THE STANDARD RELOAD GROUP WERE USED AND COMPLICATIONS INCLUDED RE-OPERATION FOR 2 PATIENTS DUE TO STAPLE LINE LEAKS WAS REPORTED. STAPLE LINE BLEEDS WERE ALSO NOTED IN 78 PATIENTS. (B)(4)-ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY SURGERY, THERE WERE 103 PATIENTS THAT THE REINFORCED RELOAD GROUP WERE USED AND STAPLE LINE BLEEDS WERE NOTED IN 23 PATIENTS.
01-02-2018 | Injury | covidien lp llc north haven | 10-01-2019 | UNKNOWN TRI STAPLE PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: (B)(4) - ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY SURGERY, THERE WERE 187 PATIENTS THAT THE STANDARD RELOAD GROUP WERE USED AND COMPLICATIONS INCLUDED RE-OPERATION FOR 2 PATIENTS DUE TO STAPLE LINE LEAKS WAS REPORTED. STAPLE LINE BLEEDS WERE ALSO NOTED IN 78 PATIENTS. (B)(4)-ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY SURGERY, THERE WERE 103 PATIENTS THAT THE REINFORCED RELOAD GROUP WERE USED AND STAPLE LINE BLEEDS WERE NOTED IN 23 PATIENTS.
17-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: SHORT CUT. IT WAS REPORTED THAT WHEN DID ESOPHAGEAL AND TUBULAR GASTRIC ANASTOMOSIS, AFTER FIRED, NOTED ONLY CUT 3/5 TISSUE AND STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
11-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE LAPAROSCOPIC SIGMOIDECTOMY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. CUSTOMER SUSPECT THE ANVIL AND RELOAD DECK ARE NOT CONNECT CLOSELY. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | RELIAMAX | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER STAPLER USED TO COMPLETE THE CAES. THERE WAS NO PATIENT INJURY.
01-02-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | UNKNOWN REINFORCED RELOAD | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: PLI10- ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY SURGERY, THERE WERE 187 PATIENTS THAT THE STANDARD RELOAD GROUP WERE USED AND COMPLICATIONS INCLUDED RE-OPERATION FOR 2 PATIENTS DUE TO STAPLE LINE LEAKS WAS REPORTED. STAPLE LINE BLEEDS WERE ALSO NOTED IN 78 PATIENTS. PLI20-ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY SURGERY, THERE WERE 103 PATIENTS THAT THE REINFORCED RELOAD GROUP WERE USED AND STAPLE LINE BLEEDS WERE NOTED IN 23 PATIENTS.
17-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | PROXIMATE* LINEAR CUTTER | Separation Failure | A150301 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN GASTRECTOMY, THE ANVIL HALF COULD NOT BE SEPARATED FROM THE CARTRIDGE HALF DURING USE. THE DEVICE WAS USED ON THE STOMACH. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF ESOPHAGEAL CANCER, AFTER FIRED, NOTED THE STAPLES MALFORMED WHEN SEVERING THE STOMACH AND ESOPHAGUS. NEW CDH25A WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA ULTRA | Difficult to Remove | A150207 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A DIGESTIVE SURGERY, AFTER STAPLING THE BOWEL, IT WAS IMPOSSIBLE TO REMOVE THE STAPLER FROM THE TROCAR. THE RELOAD WAS MISMATCHED FROM THE HANDLE BEFORE BEING REMOVED FROM THE PATIENT UNDER VISUAL CONTROL AND USING RETRACTORS. AFTER CHECKING THE RELOAD, WE CAN SEE A PIECE OF METAL TOTALLY RETURNED. NO CLINICAL CONSEQUENCE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Difficult to Remove | A150207 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A DIGESTIVE SURGERY, AFTER STAPLING THE BOWEL, IT WAS IMPOSSIBLE TO REMOVE THE STAPLER FROM THE TROCAR. THE RELOAD WAS MISMATCHED FROM THE HANDLE BEFORE BEING REMOVED FROM THE PATIENT UNDER VISUAL CONTROL AND USING RETRACTORS. AFTER CHECKING THE RELOAD, WE CAN SEE A PIECE OF METAL TOTALLY RETURNED. NO CLINICAL CONSEQUENCE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-11-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | ETS*45 ENDO LIN CUT RELOAD WH | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLES DID NOT FIRE PROPERLY. IT CAME TO BLEEDING, WHICH HAD TO BE FED WITH CLIPS.
15-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | UNKNOWN GIA PRODUCT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER STAPLER USED TO COMPLETE THE CAES. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RESECTION OF THORAX, WHILE THE DEVICE WAS BEING APPLIED FOR UPPER LOBE RESECTION, THE DEVICE WAS UNABLE TO FIRE. THEY USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RESECTION OF THORAX, WHILE THE DEVICE WAS BEING APPLIED FOR UPPER LOBE RESECTION, THE DEVICE WAS UNABLE TO FIRE. THEY USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A WHIPPLE PROCEDURE, USING A GREEN RELOAD WITH STERISTRIPS BUTTRESSING MATERIAL, AFTER THE FIRST FIRING, THE KNIFE BLADE WAS SLOW TO RETRACT. ON THE SECOND FIRING OF THE STAPLER, GREEN RELOAD, STERISTRIPS BUTTRESSING MATERIAL, THE DEVICE FIRED, BUT THE KNIFE BLADE WOULDN'T RETRACT. THE STAPLE LINE WAS COMPLETE. THE DOCTOR USED THE MANUAL OVERRIDE TO REMOVE THE STAPLER FROM TISSUE. A SECOND LIKE STAPLER AND GREEN RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
11-02-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT STAPLE PROPERLY AND IT WAS NECESSARY TO CHANGE THE RELOADS AND THEN THE STAPLER. THE DEVICE WAS REPLACED AND THE PROCEDURE CONTINUED. HAD NO HARM TO THE PATIENT.
16-12-2018 | Injury | us surgical puerto rico | 10-01-2019 | TA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SMALL INTESTINE ANASTOMOSIS PROCEDURE, THE DEVICE STAPLER DID NOT DEPLOY. TO RESOLVE THE ISSUE THE SURGEON HAD TO SUTURE THE ANASTOMOSIS.
15-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ROBOTIC GASTRECTOMY PROCEDURE, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION. NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
29-10-2018 | Injury | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHILE ON SECOND FIRING DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THEY DID NOT NOTICE THAT THE TISSUE WAS NOT VERY THICK. THE STAPLER STOPPED FIRING ON ABOUT ONE THIRD OF THE WAY AND SHOWED THE RED WEDGE INDICATING THAT THE TISSUE WAS OUT OF RANGE FOR THE SELECTED TISSUE. THEY OPENED THE STAPLER TO FIND THAT IT HAD NOT FULLY FIRED AND THAT THERE WERE PARTIALLY FIRED STAPLES PRESENT. THEY EVENTUALLY REMOVED THE UNFORMED STAPLES AND TRIMMED THE SEAM GUARD IN PREPARATION FOR THE NEXT FIRING. THE NEXT FIRING STOPPED ALSO SHOWING THE WEDGE ONCE AGAIN AFTER IT WAS PARTIALLY FIRED. THE SEAM GUARD WAS LAIRED INCREASING THE THICKNESS THAT THE STAPLER HAD TO GET THROUGH. THE THIRD LOAD WAS FIRED, AND THE SAME PROBLEM OCCURRED. THE SURGEON HAD PUT IN AN EXTRA 5MM PORT AS THERE WAS BLEEDING AFTER THE FIRST INCIDENT FORM THE INFERIOR ASPECT OF THE REMNANT JUST BELOW WHERE THE MALFORMED STAPLES HAD BEEN. A CLIP APPLIER WAS USED TO CONTROL THE BLEEDING. THEY ALSO USED 15MM PORT FOR SUBSEQUENT FIRINGS. THEN, EVENTUALLY A MANUAL HANDLE WAS OPENED AND ONLY PARTIALLY FIRED AND THEN A SECOND ONE WAS OPENED AND ONLY PARTIALLY FIRED, AND BOTH HANDLES MADE LOUD CRACKING SOUNDS AS THEY WERE BEING FIRED. THERE WERE FOUR LOUD CRACKED WERE HEARD ON EACH STAPLER. THEY OPENED AND USED ANOTHER STAPLER TO COMPLETE THE CASE. THE SURGEON ESTIMATED ABOUT 150 ML OF BLOOD LOSS. THIS ALSO RESULTED TO EXTEND SURGICAL TIME OF 40 MINUTES DELAY.
29-10-2018 | Injury | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE ON SECOND FIRING DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THEY DID NOT NOTICE THAT THE TISSUE WAS NOT VERY THICK. THE STAPLER STOPPED FIRING ON ABOUT ONE THIRD OF THE WAY AND SHOWED THE RED WEDGE INDICATING THAT THE TISSUE WAS OUT OF RANGE FOR THE SELECTED TISSUE. THEY OPENED THE STAPLER TO FIND THAT IT HAD NOT FULLY FIRED AND THAT THERE WERE PARTIALLY FIRED STAPLES PRESENT. THEY EVENTUALLY REMOVED THE UNFORMED STAPLES AND TRIMMED THE SEAM GUARD IN PREPARATION FOR THE NEXT FIRING. THE NEXT FIRING STOPPED ALSO SHOWING THE WEDGE ONCE AGAIN AFTER IT WAS PARTIALLY FIRED. THE SEAM GUARD WAS LAIRED INCREASING THE THICKNESS THAT THE STAPLER HAD TO GET THROUGH. THE THIRD LOAD WAS FIRED, AND THE SAME PROBLEM OCCURRED. THE SURGEON HAD PUT IN AN EXTRA 5MM PORT AS THERE WAS BLEEDING AFTER THE FIRST INCIDENT FORM THE INFERIOR ASPECT OF THE REMNANT JUST BELOW WHERE THE MALFORMED STAPLES HAD BEEN. A CLIP APPLIER WAS USED TO CONTROL THE BLEEDING. THEY ALSO USED 15MM PORT FOR SUBSEQUENT FIRINGS. THEN, EVENTUALLY A MANUAL HANDLE WAS OPENED AND ONLY PARTIALLY FIRED AND THEN A SECOND ONE WAS OPENED AND ONLY PARTIALLY FIRED, AND BOTH HANDLES MADE LOUD CRACKING SOUNDS AS THEY WERE BEING FIRED. THERE WERE FOUR LOUD CRACKED WERE HEARD ON EACH STAPLER. THEY OPENED AND USED ANOTHER STAPLER TO COMPLETE THE CASE. THE SURGEON ESTIMATED ABOUT 150 ML OF BLOOD LOSS. THIS ALSO RESULTED TO EXTEND SURGICAL TIME OF 40 MINUTES DELAY.
29-10-2018 | Injury | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHILE ON SECOND FIRING DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THEY DID NOT NOTICE THAT THE TISSUE WAS NOT VERY THICK. THE STAPLER STOPPED FIRING ON ABOUT ONE THIRD OF THE WAY AND SHOWED THE RED WEDGE INDICATING THAT THE TISSUE WAS OUT OF RANGE FOR THE SELECTED TISSUE. THEY OPENED THE STAPLER TO FIND THAT IT HAD NOT FULLY FIRED AND THAT THERE WERE PARTIALLY FIRED STAPLES PRESENT. THEY EVENTUALLY REMOVED THE UNFORMED STAPLES AND TRIMMED THE SEAM GUARD IN PREPARATION FOR THE NEXT FIRING. THE NEXT FIRING STOPPED ALSO SHOWING THE WEDGE ONCE AGAIN AFTER IT WAS PARTIALLY FIRED. THE SEAM GUARD WAS LAIRED INCREASING THE THICKNESS THAT THE STAPLER HAD TO GET THROUGH. THE THIRD LOAD WAS FIRED, AND THE SAME PROBLEM OCCURRED. THE SURGEON HAD PUT IN AN EXTRA 5MM PORT AS THERE WAS BLEEDING AFTER THE FIRST INCIDENT FORM THE INFERIOR ASPECT OF THE REMNANT JUST BELOW WHERE THE MALFORMED STAPLES HAD BEEN. A CLIP APPLIER WAS USED TO CONTROL THE BLEEDING. THEY ALSO USED 15MM PORT FOR SUBSEQUENT FIRINGS. THEN, EVENTUALLY A MANUAL HANDLE WAS OPENED AND ONLY PARTIALLY FIRED AND THEN A SECOND ONE WAS OPENED AND ONLY PARTIALLY FIRED, AND BOTH HANDLES MADE LOUD CRACKING SOUNDS AS THEY WERE BEING FIRED. THERE WERE FOUR LOUD CRACKED WERE HEARD ON EACH STAPLER. THEY OPENED AND USED ANOTHER STAPLER TO COMPLETE THE CASE. THE SURGEON ESTIMATED ABOUT 150 ML OF BLOOD LOSS. THIS ALSO RESULTED TO EXTEND SURGICAL TIME OF 40 MINUTES DELAY.
29-10-2018 | Injury | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE ON SECOND FIRING DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THEY DID NOT NOTICE THAT THE TISSUE WAS NOT VERY THICK. THE STAPLER STOPPED FIRING ON ABOUT ONE THIRD OF THE WAY AND SHOWED THE RED WEDGE INDICATING THAT THE TISSUE WAS OUT OF RANGE FOR THE SELECTED TISSUE. THEY OPENED THE STAPLER TO FIND THAT IT HAD NOT FULLY FIRED AND THAT THERE WERE PARTIALLY FIRED STAPLES PRESENT. THEY EVENTUALLY REMOVED THE UNFORMED STAPLES AND TRIMMED THE SEAM GUARD IN PREPARATION FOR THE NEXT FIRING. THE NEXT FIRING STOPPED ALSO SHOWING THE WEDGE ONCE AGAIN AFTER IT WAS PARTIALLY FIRED. THE SEAM GUARD WAS LAIRED INCREASING THE THICKNESS THAT THE STAPLER HAD TO GET THROUGH. THE THIRD LOAD WAS FIRED, AND THE SAME PROBLEM OCCURRED. THE SURGEON HAD PUT IN AN EXTRA 5MM PORT AS THERE WAS BLEEDING AFTER THE FIRST INCIDENT FORM THE INFERIOR ASPECT OF THE REMNANT JUST BELOW WHERE THE MALFORMED STAPLES HAD BEEN. A CLIP APPLIER WAS USED TO CONTROL THE BLEEDING. THEY ALSO USED 15MM PORT FOR SUBSEQUENT FIRINGS. THEN, EVENTUALLY A MANUAL HANDLE WAS OPENED AND ONLY PARTIALLY FIRED AND THEN A SECOND ONE WAS OPENED AND ONLY PARTIALLY FIRED, AND BOTH HANDLES MADE LOUD CRACKING SOUNDS AS THEY WERE BEING FIRED. THERE WERE FOUR LOUD CRACKED WERE HEARD ON EACH STAPLER. THEY OPENED AND USED ANOTHER STAPLER TO COMPLETE THE CASE. THE SURGEON ESTIMATED ABOUT 150 ML OF BLOOD LOSS. THIS ALSO RESULTED TO EXTEND SURGICAL TIME OF 40 MINUTES DELAY.
11-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE MEDIASTINAL LYMPHADENECTOMY PROCEDURE, AT THE TIME OF CLAMPING THE PULMONARY ARTERY INTERLOBAR, THE CURVED RELOAD WAS CLAMPED AND REMOVED AND REPOSITIONED AT THE TIME OF THE HORIZONTAL FISSURE STAPLING, BUT THE PROBLEM REMAINED. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT WAS NOT ABLE TO WORK. IT WAS REPLACED BY ANOTHER FROM THE SAME BATCH AND THE TRIGGERING OCCURRED NORMALLY. FOLLOWING THE PROCEDURE AT THE TIME OF STAPLING THE PULMONARY ARTERY INTERLOBAR WITH THE ARTIC RELOAD, THE STAPLER MADE A STRONG POPPING AND LOCKED, NOT ALLOWING THE FIRING. THE STAPLER WAS REPLACED BY ANOTHER OF THE SAME BATCH WITHOUT INTERCURRENCES. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE MEDIASTINAL LYMPHADENECTOMY PROCEDURE, AT THE TIME OF CLAMPING THE PULMONARY ARTERY INTERLOBAR, THE CURVED RELOAD WAS CLAMPED AND REMOVED AND REPOSITIONED AT THE TIME OF THE HORIZONTAL FISSURE STAPLING, BUT THE PROBLEM REMAINED. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT WAS NOT ABLE TO WORK. IT WAS REPLACED BY ANOTHER FROM THE SAME BATCH AND THE TRIGGERING OCCURRED NORMALLY. FOLLOWING THE PROCEDURE AT THE TIME OF STAPLING THE PULMONARY ARTERY INTERLOBAR WITH THE ARTIC RELOAD, THE STAPLER MADE A STRONG POPPING AND LOCKED, NOT ALLOWING THE FIRING. THE STAPLER WAS REPLACED BY ANOTHER OF THE SAME BATCH WITHOUT INTERCURRENCES. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE MEDIASTINAL LYMPHADENECTOMY PROCEDURE, AT THE TIME OF CLAMPING THE PULMONARY ARTERY INTERLOBAR, THE CURVED RELOAD WAS CLAMPED AND REMOVED AND REPOSITIONED AT THE TIME OF THE HORIZONTAL FISSURE STAPLING, BUT THE PROBLEM REMAINED. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT WAS NOT ABLE TO WORK. IT WAS REPLACED BY ANOTHER FROM THE SAME BATCH AND THE TRIGGERING OCCURRED NORMALLY. FOLLOWING THE PROCEDURE AT THE TIME OF STAPLING THE PULMONARY ARTERY INTERLOBAR WITH THE ARTIC RELOAD, THE STAPLER MADE A STRONG POPPING AND LOCKED, NOT ALLOWING THE FIRING. THE STAPLER WAS REPLACED BY ANOTHER OF THE SAME BATCH WITHOUT INTERCURRENCES. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE MEDIASTINAL LYMPHADENECTOMY PROCEDURE, AT THE TIME OF CLAMPING THE PULMONARY ARTERY INTERLOBAR, THE CURVED RELOAD WAS CLAMPED AND REMOVED AND REPOSITIONED AT THE TIME OF THE HORIZONTAL FISSURE STAPLING, BUT THE PROBLEM REMAINED. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT WAS NOT ABLE TO WORK. IT WAS REPLACED BY ANOTHER FROM THE SAME BATCH AND THE TRIGGERING OCCURRED NORMALLY. FOLLOWING THE PROCEDURE AT THE TIME OF STAPLING THE PULMONARY ARTERY INTERLOBAR WITH THE ARTIC RELOAD, THE STAPLER MADE A STRONG POPPING AND LOCKED, NOT ALLOWING THE FIRING. THE STAPLER WAS REPLACED BY ANOTHER OF THE SAME BATCH WITHOUT INTERCURRENCES. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE MEDIASTINAL LYMPHADENECTOMY PROCEDURE, AT THE TIME OF CLAMPING THE PULMONARY ARTERY INTERLOBAR, THE CURVED RELOAD WAS CLAMPED AND REMOVED AND REPOSITIONED AT THE TIME OF THE HORIZONTAL FISSURE STAPLING, BUT THE PROBLEM REMAINED. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT WAS NOT ABLE TO WORK. IT WAS REPLACED BY ANOTHER FROM THE SAME BATCH AND THE TRIGGERING OCCURRED NORMALLY. FOLLOWING THE PROCEDURE AT THE TIME OF STAPLING THE PULMONARY ARTERY INTERLOBAR WITH THE ARTIC RELOAD, THE STAPLER MADE A STRONG POPPING AND LOCKED, NOT ALLOWING THE FIRING. THE STAPLER WAS REPLACED BY ANOTHER OF THE SAME BATCH WITHOUT INTER-OCCURRENCES. THERE WAS NO PATIENT INJURY.
15-10-2018 | Malfunction | ethicon endo-surgery | 10-01-2019 | ECHELON FLEX 60 LONG | Appropriate Term/Code Not Available |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE WAS RECEIVED WITH NO COMPLAINT INFORMATION. ATTEMPTS WERE MADE TO RECONCILE THE DEVICE WITH NO SUCCESS RESULTING IN THE CREATION OF A NEW COMPLAINT.
10-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, DURING STAPLING, THE DEVICE WAS NOT ABLE TO BE FIRED. THEY USED ANOTHER STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-11-2018 | Injury | covidien medical products | 10-01-2019 | PREMIUM PLUS CEEA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING COLORECTAL SURGERY, THE DEVICE DID NOT FIRED. THE KNIFE BLADE CUT BUT NO STAPLES WERE FIRED. WHEN THEY OPENED A SECOND STAPLER, IT FIRED 270 DEGREES AND CUT THE TISSUE. THEY FINISHED THE SURGERY WITH THE SUTURE MANUALLY. PATIENT STAYED AT HOSPITAL 10 DAYS MORE AND HE/SHE HAS ILEOSTOMY ALSO HAVE DEFECATION PROBLEM.
26-11-2018 | Injury | covidien medical products | 10-01-2019 | PREMIUM PLUS CEEA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLORECTAL SURGERY, THE DEVICE DID NOT FIRED. THE KNIFE BLADE CUT BUT NO STAPLES WERE FIRED. WHEN THEY OPENED A SECOND STAPLER, IT FIRED 270 DEGREES AND CUT THE TISSUE. THEY FINISHED THE SURGERY WITH THE SUTURE MANUALLY. PATIENT STAYED AT HOSPITAL 10 DAYS MORE AND HE/SHE HAS ILEOSTOMY ALSO HAVE DEFECATION PROBLEM.
26-11-2018 | Injury | covidien medical products | 10-01-2019 | PREMIUM PLUS CEEA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING COLORECTAL SURGERY, THE DEVICE DID NOT FIRED. THE KNIFE BLADE CUT BUT NO STAPLES WERE FIRED. WHEN THEY OPENED A SECOND STAPLER, IT FIRED 270 DEGREES AND CUT THE TISSUE. THEY FINISHED THE SURGERY WITH THE SUTURE MANUALLY. PATIENT STAYED AT HOSPITAL 10 DAYS MORE AND HE/SHE HAS ILEOSTOMY ALSO HAVE DEFECATION PROBLEM.
26-11-2018 | Injury | covidien medical products | 10-01-2019 | PREMIUM PLUS CEEA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING COLORECTAL SURGERY, THE DEVICE DID NOT FIRED. THE KNIFE BLADE CUT BUT NO STAPLES WERE FIRED. WHEN THEY OPENED A SECOND STAPLER, IT FIRED 180 DEGREES AND CUT THE TISSUE. THEY FINISHED THE SURGERY WITH THE SUTURE MANUALLY.
26-11-2018 | Injury | covidien medical products | 10-01-2019 | PREMIUM PLUS CEEA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLORECTAL SURGERY, THE DEVICE DID NOT FIRED. THE KNIFE BLADE CUT BUT NO STAPLES WERE FIRED. WHEN THEY OPENED A SECOND STAPLER, IT FIRED 180 DEGREES AND CUT THE TISSUE. THEY FINISHED THE SURGERY WITH THE SUTURE MANUALLY.
26-11-2018 | Injury | covidien medical products | 10-01-2019 | PREMIUM PLUS CEEA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING COLORECTAL SURGERY, THE DEVICE DID NOT FIRED. THE KNIFE BLADE CUT BUT NO STAPLES WERE FIRED. WHEN THEY OPENED A SECOND STAPLER, IT FIRED 180 DEGREES AND CUT THE TISSUE. THEY FINISHED THE SURGERY WITH THE SUTURE MANUALLY.
07-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LOW ANTERIOR RESECTION, THE ATTACHMENT OF DEVICES ARE CONNECTED PROPERLY BUT WHEN THE JAW WAS CLOSED, IT COULD NOT BE CLOSED COMPLETELY AT THE TIP PART. THEY TRIED FOR FOUR TIMES AND THE JAW CLOSED RIGIDLY AT THE 4TH TIME. THEY STILL USED IT IN THE PROCEDURE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INVOLVED.
07-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LOW ANTERIOR RESECTION, THE ATTACHMENT OF DEVICES ARE CONNECTED PROPERLY BUT WHEN THE JAW WAS CLOSED, IT COULD NOT BE CLOSED COMPLETELY AT THE TIP PART. THEY TRIED FOR FOUR TIMES AND THE JAW CLOSED RIGIDLY AT THE 4TH TIME. THEY STILL USED IT IN THE PROCEDURE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INVOLVED.
07-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | SIGNIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LOW ANTERIOR RESECTION, THE ATTACHMENT OF DEVICES ARE CONNECTED PROPERLY BUT WHEN THE JAW WAS CLOSED, IT COULD NOT BE CLOSED COMPLETELY AT THE TIP PART. THEY TRIED FOR FOUR TIMES AND THE JAW CLOSED RIGIDLY AT THE 4TH TIME. THEY STILL USED IT IN THE PROCEDURE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INVOLVED.
07-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LOW ANTERIOR RESECTION, THE ATTACHMENT OF DEVICES ARE CONNECTED PROPERLY BUT WHEN THE JAW WAS CLOSED, IT COULD NOT BE CLOSED COMPLETELY AT THE TIP PART. THEY TRIED FOR FOUR TIMES AND THE JAW CLOSED RIGIDLY AT THE 4TH TIME. THEY STILL USED IT IN THE PROCEDURE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INVOLVED.
07-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LOW ANTERIOR RESECTION, THE ATTACHMENT OF DEVICES ARE CONNECTED PROPERLY BUT WHEN THE JAW WAS CLOSED, IT COULD NOT BE CLOSED COMPLETELY AT THE TIP PART. THEY TRIED FOR FOUR TIMES AND THE JAW CLOSED RIGIDLY AT THE 4TH TIME. THEY STILL USED IT IN THE PROCEDURE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO PATIENT INVOLVED.
05-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SURGICAL RESECTION OF THE LEFT ADRENAL HYPERTROPHY ON A PATIENT WITH BRONCHIAL CARCINOMA, THERE WAS DIFFICULTY ON INSERTION OF THE STAPLER. IT DID NOT CLOSE. THE SURGEON USED ANOTHER RELOAD TO COMPLETE THE PROCEDURE WHICH CAUSED A SLIGHT DELAY IN CARE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SURGICAL RESECTION OF THE LEFT ADRENAL HYPERTROPHY ON A PATIENT WITH BRONCHIAL CARCINOMA, THERE WAS DIFFICULTY ON INSERTION OF THE STAPLER. IT DID NOT CLOSE. THE SURGEON USED ANOTHER RELOAD TO COMPLETE THE PROCEDURE WHICH CAUSED A SLIGHT DELAY IN CARE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 10-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLORECTAL PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 10-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLORECTAL PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
03-10-2018 | Death | us surgical puerto rico | 10-01-2019 | EEA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER THE PATIENT UNDERGONE ESOPHAGECTOMY WITH RECONSTRUCTION OF THE GASTRIC TUBE IN THE STOMACH, ON THE 5TH DAY POSTOPERATIVE DAY, THE PATIENT EVOLVED WITH A 50% OF THE ESOPHAGEAL GASTRIC ANASTOMOSIS IN THE POSTERIOR PART AND 5 CM IN THE GASTRIC TUBE. IT WAS ALSO NOTED THAT THERE WAS TISSUE DAMAGE. ON THE 47TH POST OPERATIVE DAY, THE PATIENT DIED.
03-10-2018 | Death | us surgical puerto rico | 10-01-2019 | EEA | Loss of or Failure to Bond | A040102 | Death | Event Description: ACCORDING TO THE REPORTER, AFTER THE PATIENT UNDERGONE ESOPHAGECTOMY WITH RECONSTRUCTION OF THE GASTRIC TUBE IN THE STOMACH, ON THE 5TH DAY POSTOPERATIVE DAY, THE PATIENT EVOLVED WITH A 50% OF THE ESOPHAGEAL GASTRIC ANASTOMOSIS IN THE POSTERIOR PART AND 5 CM IN THE GASTRIC TUBE. IT WAS ALSO NOTED THAT THERE WAS TISSUE DAMAGE. ON THE 47TH POST OPERATIVE DAY, THE PATIENT DIED.
03-10-2018 | Death | us surgical puerto rico | 10-01-2019 | EEA | Loss of or Failure to Bond | A040102 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER THE PATIENT UNDERGONE ESOPHAGECTOMY WITH RECONSTRUCTION OF THE GASTRIC TUBE IN THE STOMACH, ON THE 5TH DAY POSTOPERATIVE DAY, THE PATIENT EVOLVED WITH A 50% OF THE ESOPHAGEAL GASTRIC ANASTOMOSIS IN THE POSTERIOR PART AND 5 CM IN THE GASTRIC TUBE. IT WAS ALSO NOTED THAT THERE WAS TISSUE DAMAGE. ON THE 47TH POST OPERATIVE DAY, THE PATIENT DIED.
03-10-2018 | Death | us surgical puerto rico | 10-01-2019 | GIA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER THE PATIENT UNDERGONE ESOPHAGECTOMY WITH RECONSTRUCTION OF THE GASTRIC TUBE ON THE STOMACH, ON THE 5TH DAY POSTOPERATIVE DAY, THE PATIENT EVOLVED WITH A 50% OF THE ESOPHAGEAL GASTRIC ANASTOMOSIS IN THE POSTERIOR PART AND 5 CM IN THE GASTRIC TUBE. IT WAS ALSO NOTED THAT THERE WAS TISSUE DAMAGE. ON THE 47TH POST OPERATIVE DAY, THE PATIENT DIED.
03-10-2018 | Death | us surgical puerto rico | 10-01-2019 | GIA | Loss of or Failure to Bond | A040102 | Death | Event Description: ACCORDING TO THE REPORTER, AFTER THE PATIENT UNDERGONE ESOPHAGECTOMY WITH RECONSTRUCTION OF THE GASTRIC TUBE ON THE STOMACH, ON THE 5TH DAY POSTOPERATIVE DAY, THE PATIENT EVOLVED WITH A 50% OF THE ESOPHAGEAL GASTRIC ANASTOMOSIS IN THE POSTERIOR PART AND 5 CM IN THE GASTRIC TUBE. IT WAS ALSO NOTED THAT THERE WAS TISSUE DAMAGE. ON THE 47TH POST OPERATIVE DAY, THE PATIENT DIED.
03-10-2018 | Death | us surgical puerto rico | 10-01-2019 | GIA | Loss of or Failure to Bond | A040102 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER THE PATIENT UNDERGONE ESOPHAGECTOMY WITH RECONSTRUCTION OF THE GASTRIC TUBE ON THE STOMACH, ON THE 5TH DAY POSTOPERATIVE DAY, THE PATIENT EVOLVED WITH A 50% OF THE ESOPHAGEAL GASTRIC ANASTOMOSIS IN THE POSTERIOR PART AND 5 CM IN THE GASTRIC TUBE. IT WAS ALSO NOTED THAT THERE WAS TISSUE DAMAGE. ON THE 47TH POST OPERATIVE DAY, THE PATIENT DIED.
24-08-2018 | Malfunction | us surgical puerto rico | 11-01-2019 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THERE WAS AN ISSUE LOADING THE DEVICE AND SO THE DEVICE DID NOT FIRE. THEY USED ANOTHER HANDLE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | covidien lp llc north haven | 11-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, DURING STAPLING, THE DEVICE WAS NOT ABLE TO BE FIRED. THEY USED ANOTHER STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ENDO ILS 29MM, CURVED | Fluid/Blood Leak | A050401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, AFTER FIRING THE DEVICE, THE ANASTOMOTIC RING WAS INTACT, AN AIR LEAK TEST WAS PERFORMED AND A LEAK WAS DETECTED AT THE ANASTOMOTIC RING. THE DOCTOR OVER SEWED THE ANASTOMOSIS WITH SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THE SALES REP WAS PRESENT FOR THE PROCEDURE.
23-11-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE INSTRUMENT RELOAD THREE TIMES. THE SECOND MAGAZINE STAPLED NOT COMPLETELY AND IT HAD TO BE SEWN. NO HARM TO THE PATIENT.
10-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | FLEX 60 ARTICULATING | Material Twisted/Bent | A040609 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AT THE OPENING OF THE FORCEPS THE JOINT WAS TWISTED. ATTEMPT TO PLACE THE CARTRIDGE TO CLOSE AND MANIPULATE THE DEVICE BY THE SURGEON. WHEN CLOSING THE CLAMP, WAS BLOCKED; IMPOSSIBILITY TO REOPEN IT AND USE IT. NO PATIENT CONSEQUENCE.
19-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | 75MM SELECTABLE NEW TLC | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: SCREW FELL OFF. IT WAS REPORTED THAT THE SCREW WAS FOUND FALLED OFF WHEN OPENING THE STERILE PACKING DURING SURGERY. THE DEVICE WAS NOT USED AND THE SCREW DIDN'T FALL INTO THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE "PANCREATODUODENUM", AFTER FIRED, NOTED THE STAPLES MALFORMED AND BLEEDING. SUTURE WAS USED TO OVERSEW AND STOP BLEEDING WITHOUT BLOOD TRANSFUSION. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
01-01-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLING DIDN'T WORK OUT PROPERLY SO DOCTOR COULDN'T CONNECT STOMACH AND INTESTINES. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
17-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN TOTAL PROCTOCOLECTOMY, THE FIRING KNOB WAS BROKEN OFF AT THE FIRING ON THE COLON. NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
18-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE PHYSICIAN WAS PERFORMING A ROBOTIC PROSTATECTOMY, REQUESTED TO USE PVS35 FOR LIGATION AND TRANSECTION OF THE PEDICLE BUNDLE OUTSIDE OF PROSTATE. DISCUSSION PRIOR TO CASE WITH PHYSICIAN AND THE TERRITORY ACCOUNT LEADER REGARDING THE TISSUE THICKNESS AND QUESTION WHETHER STAPLES WOULD FORM CORRECTLY. ENCOURAGED TO USE PSEE45A. PHYSICIAN DETERMINED TO TRY SMALLER PROFILE PVS35 DURING PROCEDURE. FIRST FIRING ON RIGHT SIDE SHOWED INSUFFICIENT STAPLER LENGTH COVERAGE AND STAPLES NOT FORMED , CAUSING BLEEDING IN MID STAPLE LINE. SALES REP QUESTIONED PHYSICIAN AND PA FIRING STAPLER AS TO TISSUE THICKNESS AND DIFFICULTY CLOSING. PHYSICIAN INSISTED ON FIRING AGAIN ON LEFT SIDE. BETTER COVERAGE OF STAPLE LINE, BUT AGAIN STAPLES DID NOT FORM, RESULTING IN BLEEDING FROM TRANSECTION OF PEDICLE. PHYSICIAN FELT WRONG CHOICE IN STAPLER ADDED SURGICAL TIME AND INCREASE IN BLEEDING GAVE HIM MORE TO CORRECT DURING SURGERY. MALFORMED STAPLE LINE, CAUSING BLEEDING FROM VESSELS THAT WERE TRANSACTED AND NOT CLOSED WITH APPROPRIATE STAPLE FORMATION.
19-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT, PERFORMED RIGHT TRANSVERSE COLOSTOMY USING EES NEW LINEAR CUTTER 75MM NTLC75 AND 2PCS OF RELOAD SR75. 1ST FIRING NOTICED SOME STAPLES DIDN'T FORM PROPERLY. UPON 2ND FIRING, ALL STAPLES FAILED TO FORM NORMAL B SHAPE STAPLE (STAPLE LEGS WERE OPENED). DR USED "BLUE" SETTING. PATIENT DIDN'T UNDERGO RADIOTHERAPY, TISSUES WERE NORMAL. SURGERY WAS DELAYED 30 MINUTES. THE PROCEDURE WAS COMPLETED BY HAND SEWING USING SUTURE TO CLOSE THE MALFORMED STAPLE LINE. PROCEDURE WAS SUCCESSFULLY COMPLETED. PATIENT STATUS IN UNKNOWN, BUT NO PATIENT CONSEQUENCES WERE REPORTED.
19-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | 75MM SELECTABLE NEW TLC | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT, PERFORMED RIGHT TRANSVERSE COLOSTOMY USING EES NEW LINEAR CUTTER 75MM NTLC75 AND 2PCS OF RELOAD SR75. 1ST FIRING NOTICED SOME STAPLES DIDN'T FORM PROPERLY. UPON 2ND FIRING, ALL STAPLES FAILED TO FORM NORMAL B SHAPE STAPLE (STAPLE LEGS WERE OPENED). DR USED "BLUE" SETTING. PATIENT DIDN'T UNDERGO RADIOTHERAPY, TISSUES WERE NORMAL. SURGERY WAS DELAYED 30 MINUTES. THE PROCEDURE WAS COMPLETED BY HAND SEWING USING SUTURE TO CLOSE THE MALFORMED STAPLE LINE. PROCEDURE WAS SUCCESSFULLY COMPLETED. PATIENT STATUS IN UNKNOWN, BUT NO PATIENT CONSEQUENCES WERE REPORTED.
20-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 33MM, CURVED | Misfire | A050502 | Failure to Anastomose | Event Description: SEE USER FACILITY MEDWATCH.
20-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 33MM, CURVED | Misfire | A050502 | Tissue Damage | Event Description: SEE USER FACILITY MEDWATCH.
20-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 33MM, CURVED | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: SEE USER FACILITY MEDWATCH.
20-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 33MM, CURVED | Misfire | A050502 | No Code Available | Event Description: SEE USER FACILITY MEDWATCH.
01-01-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ILS 29MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE MISFIRED DURING USE. THE EXPLANATION OF EVENTS WERE: THE FIRING RANGE WAS WITHIN THE GREEN ZONE AND WHEN I FIRED THE DEVICE THE HANDLE COMPLETED THE PULL BACK BUT NO AUDIBLE FEEDBACK WAS GIVEN, THE STAPLES DID FIRE BUT I AM CONCERNED ABOUT THE FORMATION OF THE STAPLES, I THEN PERFORMED A JACUZZI TEST WHICH WAS A SUCCESS BUT I¿M CONCERNED FOR THE LONG TERM OF THE STAPLE LINE INTEGRITY. TWO DONUTS WERE SUCCESSFULLY RECEIVED FROM THE DEVICE.
01-11-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | GST BLACK RELOAD, 60MM, 6 ROW | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE THE DEVICE CUT BUT DID NOT FIRE STAPLES ON THE SPECIMEN SIDE. A NEW RELOAD WITH THE SAME DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON USED THE PVE35A AND VASECR35 IN LOBECTOMY SURGERY. WHEN HE APPLIED THE STAPLER INTO VASCULAR AND TRIED TO TRANSECT THE VASCULAR. HE FOUND THAT THE KNIFE CAN'T BE RETURNED. SO THEY USED THE MANUAL OVERRIDE SWITCH TO RETURN KNIFE. THE NURSE CHANGED A NEW PVE35A AND VASECR35 TO CONTINUE THE SURGERY.
01-01-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | PWRD ECH FLEX 45MM 34CM SHAFT | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON USED PSE45A IN LAPAROSCOPIC APPENDICECTOMY SURGERY. HE FOUND THAT THE KNIFE CAN'T BE RETURNED , SO THEY USED THE MANUAL OVERRIDE SWITCH TO SOLVE THE PROBLEM.
17-12-2018 | Injury | intuitive surgical | 11-01-2019 | ENDOWRIST | Failure to Sense | A070908 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THE SURGICAL STAFF ENCOUNTERED A STAPLER 45 RELOAD RECOGNITION ISSUE. THE INITIAL REPORTER CLAIMED THAT THE SURGICAL STAFF HAD TO MANUALLY SELECT EACH RELOAD. IN ADDITION, THE INITIAL REPORTER INDICATED THAT THE BLADE OF A STAPLER 45 RELOAD HAD NOT CUT CLEAN AND TISSUE WAS LEFT ON THE BLADE. ON POST-OPERATIVE DAY #1, THE PATIENT REPORTEDLY RETURNED TO THE SURGEON DUE TO BLEEDING IDENTIFIED ON ONE OF THE STAPLE LINES. NO OTHER CLINICAL INFORMATION WAS PROVIDED.
17-12-2018 | Injury | intuitive surgical | 11-01-2019 | ENDOWRIST | Failure to Cut | A050702 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THE SURGICAL STAFF ENCOUNTERED A STAPLER 45 RELOAD RECOGNITION ISSUE. THE INITIAL REPORTER CLAIMED THAT THE SURGICAL STAFF HAD TO MANUALLY SELECT EACH RELOAD. IN ADDITION, THE INITIAL REPORTER INDICATED THAT THE BLADE OF A STAPLER 45 RELOAD HAD NOT CUT CLEAN AND TISSUE WAS LEFT ON THE BLADE. ON POST-OPERATIVE DAY #1, THE PATIENT REPORTEDLY RETURNED TO THE SURGEON DUE TO BLEEDING IDENTIFIED ON ONE OF THE STAPLE LINES. NO OTHER CLINICAL INFORMATION WAS PROVIDED.
19-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | GST BLACK RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, THE BLACK LOAD ON THE FIRST FIRING ON THE SLEEVE (B)(4) AND THE STAPLES FORMED ON THE PATIENT SIDE, BUT THE SPECIMEN SIDE THE STAPLES WERE NOT FORMED PROPERLY AND SHOOTING ALL OVER THE PLACE. THEY WERE STICKING UP AND JUST ALL OVER THE PLACE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
04-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 21MM, CURVED | Insufficient Information | E2401 | Blood Loss | Event Description: IT WAS REPORTED: BLEEDING AFTER SURGERY. IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY, AT THE NIGHT AFTER OPERATION, NOTED BLOOD IN THE DRAINAGE TUBE OF STOMACH, WITH CONSERVATIVE TREATMENT TO STOP BLEEDING AND WITH BLOOD TRANSFUSION. AFTER ONE WEEK, NOTED THE PATIENT WAS WITH HEMORRHAGIC SHOCK, DID THE EMERGENCY SURGERY TO STOP BLEEDING, FOUND ONE STAPLE COVERED THE VENULE OF MUCOSA BUT STILL NOTED THE PULSATILE BLEEDER FROM VENULE, USED SUTURE TO STOP BLEEDING. THE PATIENT IS STABLE IN HOSPITAL NOW.
04-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 21MM, CURVED | Insufficient Information | E2401 | No Code Available | Event Description: IT WAS REPORTED: BLEEDING AFTER SURGERY. IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY, AT THE NIGHT AFTER OPERATION, NOTED BLOOD IN THE DRAINAGE TUBE OF STOMACH, WITH CONSERVATIVE TREATMENT TO STOP BLEEDING AND WITH BLOOD TRANSFUSION. AFTER ONE WEEK, NOTED THE PATIENT WAS WITH HEMORRHAGIC SHOCK, DID THE EMERGENCY SURGERY TO STOP BLEEDING, FOUND ONE STAPLE COVERED THE VENULE OF MUCOSA BUT STILL NOTED THE PULSATILE BLEEDER FROM VENULE, USED SUTURE TO STOP BLEEDING. THE PATIENT IS STABLE IN HOSPITAL NOW.
19-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | POWERED 60 ECHELON +, 280MM SHAFT | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON IS A THORACIC SURGEON AND AN EXPERIENCED ECHELON USER. HE SELECTED THE BLACK CARTRIDGE AND AS HE FIRED THE GUN FELT THAT THE NOISE MADE BY THE MOTOR DIDN¿T SOUND AS IT USUALLY DOES. THE BLADE DID NOT RETURN TO THE HOUSING UNIT AND HE HAD TO USE THE REVERSE BUTTON. HE WAS NOT HAPPY WITH THE STAPLE LINE AND REQUESTED THAT BE REPORTED AS A FAULTY PRODUCT. ANOTHER PCEE60A DEVICE WAS OPENED AND A BLACK RELOAD WAS SELECTED. HE FIRED THE DEVICE (NOTE: THE SURGEON AND STAFF ALL CONFIRMED THAT HE WAITED THE RECOMMENDED 15 SECONDS) THE BLADE MOVED FORWARD AND THEN BECAME JAMMED. HE REPORTED THAT THE REVERSE BUTTON DID NOT WORK ON THIS OCCASION AND THE DEVICE WAS COMPLETELY STUCK. HE THEN USED THE MANUAL OVERRIDE AND WAS ABLE TO REMOVE THE DEVICE FROM THE PATIENT. THEY CONTINUED USING AN ENDO GIA COVIDIEN PRODUCT AS THE SURGEON FEELS THERE IS AN ISSUE WITH THIS LOT/BATCH OF PCEE60A.
14-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | PROXIMATE*RL LINEAR STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE RADICAL GASTRECTOMY FOR GASTRIC CANCER SURGERY, AFTER FIRED, NOTED THE PART OF STAPLES MALFORMED WITH STRAIGHT LEG. SUTURE WAS USED TO SEW THE WOUND. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | POWERED 60 ECHELON +, 280MM SHAFT | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON IS A THORACIC SURGEON AND AN EXPERIENCED ECHELON USER. HE SELECTED THE BLACK CARTRIDGE AND AS HE FIRED THE GUN FELT THAT THE NOISE MADE BY THE MOTOR DIDN'T SOUND AS IT USUALLY DOES. THE BLADE DID NOT RETURN TO THE HOUSING UNIT AND HE HAD TO USE THE REVERSE BUTTON. HE WAS NOT HAPPY WITH THE STAPLE LINE AND REQUESTED THAT BE REPORTED AS A FAULTY PRODUCT. ANOTHER PCEE60A DEVICE WAS OPENED AND A BLACK RELOAD WAS SELECTED. HE FIRED THE DEVICE (NOTE-THE SURGEON AND STAFF ALL CONFIRMED THAT HE WAITED THE RECOMMENDED 15 SECONDS) THE BLADE MOVED FORWARD AND THEN BECAME JAMMED. HE REPORTED THAT THE REVERSE BUTTON DID NOT WORK ON THIS OCCASION AND THE DEVICE WAS COMPLETELY STUCK. HE THEN USED THE MANUAL OVERRIDE AND WAS ABLE TO REMOVE THE DEVICE FROM THE PATIENT. THEY CONTINUED USING AN ENDO GIA COVIDIEN PRODUCT AS THE SURGEON FEELS THERE IS AN ISSUE WITH THIS LOT/BATCH OF PCEE60A.
18-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF ESOPHAGEAL CANCER, WHEN DID THE ESOPHAGOGASTRIC ANASTOMOSIS, AFTER FIRED WITH AUDIBLE FEEDBACK, BUT NOTED STAPLE MALFORMED AND BLEEDING. CHANGED ANOTHER CDH25A TO STOP BLEEDING AND COMPLETE SURGERY. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
12-12-2018 | Malfunction | covidien lp llc north haven | 11-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO / LOBECTOMY PROCEDURE, AT THE FOURTH FIRING, THEY CHECKED THE OPENING AND CLOSING OF THE JAW AND APPROACHED IT TO THE TISSUE. WHEN THEY PRESSED THE FIRING BUTTON, THE FIRING STOPPED. THE INDICATOR SHOWED THE FIRING DISPLAY AS USUAL. AFTER BEING REPLACED WITH A NEW CARTRIDGE, IT WAS USED WITHOUT PROBLEM AND FIRED FOR 3 SHOTS.
06-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED: BLEEDING AFTER SURGERY. IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY+ CHILD RECONSTRUCTION OF DIGESTIVE TRACT, USED CDH29A TO DID THE SIDE-TO-SIDE GASTROJEJUNOSTOMY. ONE DAY AFTER OPERATION, NOTED BLOOD IN THE DRAINAGE TUBE OF STOMACH AND CONTINUOUS DECREASE OF HEMOGLOBIN, DID THE EMERGENCY SURGERY ON 7TH DEC, NOTED PULSATILE BLEEDER FROM POSTERIOR OF GASTROINTESTINAL ANASTOMOSIS AS VIDEO SHOWS, USED SUTURE TO STOP BLEEDING. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
06-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED: BLEEDING AFTER SURGERY. IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY+ CHILD RECONSTRUCTION OF DIGESTIVE TRACT, USED CDH29A TO DID THE SIDE-TO-SIDE GASTROJEJUNOSTOMY. ONE DAY AFTER OPERATION, NOTED BLOOD IN THE DRAINAGE TUBE OF STOMACH AND CONTINUOUS DECREASE OF HEMOGLOBIN, DID THE EMERGENCY SURGERY ON 7TH DEC, NOTED PULSATILE BLEEDER FROM POSTERIOR OF GASTROINTESTINAL ANASTOMOSIS AS VIDEO SHOWS, USED SUTURE TO STOP BLEEDING. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
01-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT THE PATIENT WAS BROUGHT IN WITH POST-OP BLOOD LOSS. PT WAS BROUGHT IN 2 DAYS POST OP WITH RECTAL BLOOD LOSS. SHE FAINTED TWO TIMES BEFORE FINDING OUT (AT HOME) THE BLOOD LOSS. PATIENT HAD A REDO PROCEDURE. BAND OUT AND R&Y BYPASS IN. LOTS OF ADHESIONS, NO FURTHER PARTICULARITIES. STAPLE LINES SEEMED ALL GOOD. SURGEON DOESN¿T KNOW WHERE IT COMES FROM.
01-12-2018 | Injury | ethicon endo-surgery | 11-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT THE PATIENT WAS BROUGHT IN WITH POST-OP BLOOD LOSS. PT WAS BROUGHT IN 2 DAYS POST OP WITH RECTAL BLOOD LOSS. SHE FAINTED TWO TIMES BEFORE FINDING OUT (AT HOME) THE BLOOD LOSS. PATIENT HAD A REDO PROCEDURE. BAND OUT AND R&Y BYPASS IN. LOTS OF ADHESIONS, NO FURTHER PARTICULARITIES. STAPLE LINES SEEMED ALL GOOD. SURGEON DOESN¿T KNOW WHERE IT COMES FROM.
29-11-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Infarction, Cerebral | Event Description: IT WAS REPORTED: BLEEDING AFTER SURGERY. IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY+ CHILD RECONSTRUCTION OF DIGESTIVE TRACT,AT THE NIGHT AFTER OPERATION, NOTED BLOOD IN THE DRAINAGE TUBE OF STOMACH AND CONTINUOUS DECREASE OF HEMOGLOBIN,DID THE EMERGENCY SURGERY ON (B)(6), NOTED THE PULSATILE BLEEDER AND MUCOSA WAS TORN, USED SUTURE TO STOP BLEEDING. AFTER SURGERY, THE PATIENT WAS WITH SOME SEVERE COMPLICATIONS: CEREBRAL INFARCTION, RIGHT LIMB HEMIPLEGIA, AND STRESS GASTROINTESTINAL BLEEDING. THE PATIENT IS STABLE IN HOSPITAL NOW.
29-11-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED: BLEEDING AFTER SURGERY. IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY+ CHILD RECONSTRUCTION OF DIGESTIVE TRACT,AT THE NIGHT AFTER OPERATION, NOTED BLOOD IN THE DRAINAGE TUBE OF STOMACH AND CONTINUOUS DECREASE OF HEMOGLOBIN,DID THE EMERGENCY SURGERY ON (B)(6), NOTED THE PULSATILE BLEEDER AND MUCOSA WAS TORN, USED SUTURE TO STOP BLEEDING. AFTER SURGERY, THE PATIENT WAS WITH SOME SEVERE COMPLICATIONS: CEREBRAL INFARCTION, RIGHT LIMB HEMIPLEGIA, AND STRESS GASTROINTESTINAL BLEEDING. THE PATIENT IS STABLE IN HOSPITAL NOW.
29-11-2018 | Injury | ethicon endo-surgery | 11-01-2019 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED: BLEEDING AFTER SURGERY. IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY+ CHILD RECONSTRUCTION OF DIGESTIVE TRACT,AT THE NIGHT AFTER OPERATION, NOTED BLOOD IN THE DRAINAGE TUBE OF STOMACH AND CONTINUOUS DECREASE OF HEMOGLOBIN,DID THE EMERGENCY SURGERY ON (B)(6), NOTED THE PULSATILE BLEEDER AND MUCOSA WAS TORN, USED SUTURE TO STOP BLEEDING. AFTER SURGERY, THE PATIENT WAS WITH SOME SEVERE COMPLICATIONS: CEREBRAL INFARCTION, RIGHT LIMB HEMIPLEGIA, AND STRESS GASTROINTESTINAL BLEEDING. THE PATIENT IS STABLE IN HOSPITAL NOW.
18-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE EXCISION OF DISTAL GASTRIC CANCER SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. SUTURE WAS USED TO SEW THE WOUND. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE WAS USED DURING LAP SLEEVE GASTRECTOMY. THE DEVICE WAS PLACED ON STOMACH NEAR PYLORUS WITH BLACK RELOAD AND SEAM GUARD. WHEN THEY ATTEMPTED TO FIRE THE DEVICE MADE AN UNUSUAL NOISE. WHEN THEY ATTEMPTED TO REMOVE DEVICE THE HANDLE WOULD OPEN BUT END EFFECTOR WOULD NOT. REVERSE BUTTON USED TWICE AND THE SECOND TIME THE BLADE REVERSED BUT NOT COMPLETELY. THE JAWS WOULD STILL NOT OPEN, ALTHOUGH HANDLE WOULD OPEN. A NEW BATTERY WAS USED WITHOUT SUCCESS. A THIRD DEVICE WAS OPENED AND USED TO COMPLETE THE CASE.
18-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE WAS USED DURING LAP SLEEVE GASTRECTOMY. THE DEVICE WAS PLACED ON STOMACH NEAR PYLORUS WITH BLACK RELOAD AND SEAM GUARD. WHEN THEY ATTEMPTED TO FIRE THE DEVICE MADE AN UNUSUAL NOISE. WHEN THEY ATTEMPTED TO REMOVE DEVICE THE HANDLE WOULD OPEN BUT END EFFECTOR WOULD NOT. REVERSE BUTTON USED TWICE AND THE SECOND TIME THE BLADE REVERSED BUT NOT COMPLETELY. THE JAWS WOULD STILL NOT OPEN, ALTHOUGH HANDLE WOULD OPEN. A NEW BATTERY WAS USED WITHOUT SUCCESS. A THIRD DEVICE WAS OPENED AND USED TO COMPLETE THE CASE.
18-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE WAS USED DURING LAP SLEEVE GASTRECTOMY. THE DEVICE WAS PLACED ON STOMACH NEAR PYLORUS WITH BLACK RELOAD AND SEAM GUARD. WHEN THEY ATTEMPTED TO FIRE THE DEVICE MADE AN UNUSUAL NOISE. WHEN THEY ATTEMPTED TO REMOVE DEVICE THE HANDLE WOULD OPEN BUT END EFFECTOR WOULD NOT. REVERSE BUTTON USED TWICE AND THE SECOND TIME THE BLADE REVERSED BUT NOT COMPLETELY. THE JAWS WOULD STILL NOT OPEN, ALTHOUGH HANDLE WOULD OPEN. A NEW BATTERY WAS USED WITHOUT SUCCESS. A THIRD DEVICE WAS OPENED AND USED TO COMPLETE THE CASE.
13-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE ROBOT THORACOSCOPIC LOBECTOMY, AFTER FIRED WITH ECR60B AND TWO ECR60DS, NOTED STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
13-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE ROBOT THORACOSCOPIC LOBECTOMY, AFTER FIRED WITH ECR60B AND TWO ECR60DS, NOTED STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
13-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE ROBOT THORACOSCOPIC LOBECTOMY, AFTER FIRED WITH ECR60B AND TWO ECR60DS, NOTED STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-06-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT FIRE AND MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE RADICAL GASTRECTOMY FOR GASTRIC CANCER SURGERY, COULD NOT FIRE WITH EC60A AND ECR60G(P58N0K). CHANGED ANOTHER ECR60G (P58N0K), STILL COULD NOT FIRE. THEN CHANGED WITH ECR60G(R58L3L), NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: HOLD FOR ESTHER 3/1/2023
19-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY, THE SURGEON FIRED THE PVS AND THE RELOAD ONLY DEPLOYED ON THE NON-PATIENT SIDE WHEN THEY WERE FIRING ACROSS THE PULMONARY ARTERY. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. DURING CASE THERE WAS A SUBSTANTIAL AMOUNT OF BLOOD LOSS. IT IS UNKNOWN IF THERE WERE ANY ADVERSE CONSEQUENCES TO THE PATIENT.
11-12-2018 | Malfunction | covidien lp llc north haven | 11-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A SIGMOIDECTOMY AT THE 1ST FIRING WHEN THE RELOAD WAS CONNECTED TO THE HANDLE, THE HANDLE COULD NOT BE SQUEEZED. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT HARM.
11-12-2018 | Malfunction | us surgical puerto rico | 11-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A SIGMOIDECTOMY AT THE 1ST FIRING WHEN THE RELOAD WAS CONNECTED TO THE HANDLE, THE HANDLE COULD NOT BE SQUEEZED. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT HARM.
14-11-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | POWERED 60 ECHELON +, 280MM SHAFT | Fitting Problem | A1208 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE THAT WHEN TECH WAS TRYING TO LOAD THE STAPLER AFTER FIRST FIRING, THE RELOAD WOULD NOT FIT INTO GUN. TECH SAID IT LOOKED LIKE SOMETHING WAS IN JAWS OF THE DEVICE AND WOULD NOT ALLOW RELOAD TO MOVE BACK.
23-10-2018 | Death | us surgical puerto rico | 11-01-2019 | GIA | Loss of or Failure to Bond | A040102 | Death | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVE A TRANSANAL TOTAL MESORECTAL EXCISION, THE SURGEON NOTICED THAT THERE WAS A HOLE IN THE STAPLE LINE AND ALSO NOTICED THAT THE PATIENT DEVELOPED CLOTTING PROBLEMS. THE HOLE WAS CLOSED BY SUTURING. AFTER LEAVING THE HOSPITAL, THE PATIENT DIED AT HOME. THE PATIENT'S DEATH WAS NOT RELATED TO THE DEVICE ISSUE.
23-10-2018 | Death | us surgical puerto rico | 11-01-2019 | GIA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVE A TRANSANAL TOTAL MESORECTAL EXCISION, THE SURGEON NOTICED THAT THERE WAS A HOLE IN THE STAPLE LINE AND ALSO NOTICED THAT THE PATIENT DEVELOPED CLOTTING PROBLEMS. THE HOLE WAS CLOSED BY SUTURING. AFTER LEAVING THE HOSPITAL, THE PATIENT DIED AT HOME. THE PATIENT'S DEATH WAS NOT RELATED TO THE DEVICE ISSUE.
20-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN PNEUMONECTOMY, BLEEDING OCCURRED WHEN THE JAWS WERE OPENED AFTER THE 3RD FIRING ON THE BLOOD VESSEL. THE STAPLES WERE UNFORMED. SOME STAPLES WERE LEFT INSIDE THE JAWS, AND THOSE WERE FORMED AS INTENDED. THERE WAS NO DIFFICULTY IN THE FIRING. THE AMOUNT OF THE BLEEDING WAS UNKNOWN. NO BLOOD TRANSFUSION WAS REQUIRED. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
20-12-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN PNEUMONECTOMY, BLEEDING OCCURRED WHEN THE JAWS WERE OPENED AFTER THE 3RD FIRING ON THE BLOOD VESSEL. THE STAPLES WERE UNFORMED. SOME STAPLES WERE LEFT INSIDE THE JAWS, AND THOSE WERE FORMED AS INTENDED. THERE WAS NO DIFFICULTY IN THE FIRING. THE AMOUNT OF THE BLEEDING WAS UNKNOWN. NO BLOOD TRANSFUSION WAS REQUIRED. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
16-05-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DID NOT FIRING. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 11-01-2019 | 75 MM NTLC SELECTABLE RELOAD | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE PLASTIC TIP AT THE END OF THE CARTRIDGE WAS ALREADY BROKEN. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
17-12-2018 | Injury | covidien lp llc north haven | 13-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLES DID NOT CLOSE CORRECTLY. THE STAPLE LINE WAS OVERSEWN TO RESOLVE THE ISSUE.
17-12-2018 | Injury | covidien lp llc north haven | 13-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLES DID NOT CLOSE CORRECTLY. THE STAPLE LINE WAS OVERSEWN TO RESOLVE THE ISSUE.
14-12-2018 | Malfunction | us surgical puerto rico | 13-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE.
14-12-2018 | Malfunction | covidien lp llc north haven | 13-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE.
17-12-2018 | Injury | us surgical puerto rico | 13-01-2019 | EEA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, WHILE THE DEVICE WAS BEING APPLIED TO THE CIRCULAR ANASTOMOSIS OF THE REMAINING PART OF THE COLON TO THE ANAL CANAL, THE STAPLES FELL INTO THE PATIENT'S CAVITY WITHOUT FIRING THE DEVICE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. A NEW DEVICE WAS USED TO COMPLETE THE CASE.
17-12-2018 | Injury | us surgical puerto rico | 13-01-2019 | EEA | Expulsion | A010401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, WHILE THE DEVICE WAS BEING APPLIED TO THE CIRCULAR ANASTOMOSIS OF THE REMAINING PART OF THE COLON TO THE ANAL CANAL, THE STAPLES FELL INTO THE PATIENT'S CAVITY WITHOUT FIRING THE DEVICE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. A NEW DEVICE WAS USED TO COMPLETE THE CASE.
17-12-2018 | Malfunction | us surgical puerto rico | 13-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACIC LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE BRONCHUS OF THE UPPER LOBE OF THE LUNG, THE HANDLE COULD NOT BE SQUEEZED USING THE 60MM RELOAD. A 45MM RELOAD WAS PUT ON THE HANDLE AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 13-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACIC LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE BRONCHUS OF THE UPPER LOBE OF THE LUNG, THE HANDLE COULD NOT BE SQUEEZED USING THE 60MM RELOAD. A 45MM RELOAD WAS PUT ON THE HANDLE AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 13-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, WHEN CHECKING THE CLOSING AND OPENING OF THE JAWS, IT WAS FOUND THAT THE JAWS COULD BE CLOSED BUT COULD NOT BE OPENED. A NEW RELOAD PACKAGE WAS OPENED, CONNECTED IT TO THE HANDLE AND THE DEVICE WAS USED WITH NO PROBLEM.THERE WAS NO PATIENT INVOLVEMENT.
14-12-2018 | Malfunction | us surgical puerto rico | 13-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, WHEN CHECKING THE CLOSING AND OPENING OF THE JAWS, IT WAS FOUND THAT THE JAWS COULD BE CLOSED BUT COULD NOT BE OPENED. A NEW RELOAD PACKAGE WAS OPENED, CONNECTED IT TO THE HANDLE AND THE DEVICE WAS USED WITH NO PROBLEM.THERE WAS NO PATIENT INVOLVEMENT.
17-12-2018 | Malfunction | us surgical puerto rico | 13-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORASCOPIC WEDGE SEGMENTAL RESECTION PROCEDURE, WHILE USING THE DEVICE FOR PNEUMOTHORAX ON BULLA RESECTION LINE, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON HOWEVER THE FIRING WAS NOT PERFORMED. THE RELOAD WAS REPLACED TO A NEW ONE, AND THE FIRING WAS DONE WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
16-12-2018 | Injury | covidien lp llc north haven | 13-01-2019 | ABSORBATACK | Loss of or Failure to Bond | A040102 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE TACKS OF THE DEVICE DID NOT STAY IN THE MESH AND WERE FALLING. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE DEVICE ISSUE. THE TACKS FELL INTO THE PATIENT CAVITY AND WERE NOT RETRIEVED. THEY SUTURED THE MESH TO RESOLVE THE ISSUE.
16-12-2018 | Injury | covidien lp llc north haven | 13-01-2019 | ABSORBATACK | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE TACKS OF THE DEVICE DID NOT STAY IN THE MESH AND WERE FALLING. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE DEVICE ISSUE. THE TACKS FELL INTO THE PATIENT CAVITY AND WERE NOT RETRIEVED. THEY SUTURED THE MESH TO RESOLVE THE ISSUE.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DOCTOR FIRED THE FIRST STAPLER AND WHEN HE OPENED IT, THE STAPLERS WERE DEPLOYED ONLY ON ONE SIDE OF THE STOMACH. IT WAS REPORTED THAT THE SPECIMEN SIDE OPENED BECAUSE THE STAPLER DID NOT FORM A B-SHAPE AT ALL. ANOTHER CARTRIDGE HAD TO BE USED TO COMPLETE THE CASE. IT WAS REPORTED THAT THERE WAS NO STAPLE FORMATION ON ALL THE LINE OF THE STAPLER (PROXIMAL, DISTAL AND CENTRAL) BUT ONLY ON THE SPECIMEN SIDE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DOCTOR FIRED THE FIRST STAPLER AND WHEN HE OPENED IT, THE STAPLERS WERE DEPLOYED ONLY ON ONE SIDE OF THE STOMACH. IT WAS REPORTED THAT THE SPECIMEN SIDE OPENED BECAUSE THE STAPLER DID NOT FORM A B-SHAPE AT ALL. ANOTHER CARTRIDGE HAD TO BE USED TO COMPLETE THE CASE. IT WAS REPORTED THAT THERE WAS NO STAPLE FORMATION ON ALL THE LINE OF THE STAPLER (PROXIMAL, DISTAL AND CENTRAL) BUT ONLY ON THE SPECIMEN SIDE. THERE WAS NO PATIENT INJURY.
16-12-2018 | Injury | us surgical puerto rico | 14-01-2019 | EEA | Difficult to Remove | A150207 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING INTESTINAL ANASTOMOSIS, THE ANVIL COULD NOT BE TILTED AFTER FIRING. IT WAS REPORTED ANASTOMOSIS WAS DONE MANUALLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2018 | Injury | us surgical puerto rico | 14-01-2019 | EEA | Difficult to Remove | A150207 | No Code Available | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING INTESTINAL ANASTOMOSIS, THE ANVIL COULD NOT BE TILTED AFTER FIRING. IT WAS REPORTED ANASTOMOSIS WAS DONE MANUALLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2018 | Injury | us surgical puerto rico | 14-01-2019 | EEA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING INTESTINAL ANASTOMOSIS, THE ANVIL COULD NOT BE TILTED AFTER FIRING. IT WAS REPORTED ANASTOMOSIS WAS DONE MANUALLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2018 | Injury | us surgical puerto rico | 14-01-2019 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING INTESTINAL ANASTOMOSIS, THE ANVIL COULD NOT BE TILTED AFTER FIRING. IT WAS REPORTED ANASTOMOSIS WAS DONE MANUALLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORASCOPIC WEDGE SEGMENTAL RESECTION PROCEDURE, WHILE USING THE DEVICE FOR PNEUMOTHORAX ON BULLA RESECTION LINE, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON HOWEVER THE FIRING WAS NOT PERFORMED. THE RELOAD WAS REPLACED TO A NEW ONE, AND THE FIRING WAS DONE WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | TRI-STAPLE 2.0 | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FOURTH FIRING ON THE STOMACH TISSUE IN THE FUNDUS AREA ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AND AN INCOMPLETE STAPLE LINE. IT WAS STATED THAT THE STAPLE LINE WAS NOT ABLE TO BE FULLY FIRED AND APPROXIMATELY THREE ROWS OF UNFORMED STAPLES WENT THROUGH THE TISSUE AT THE DISTAL. A NEW RELOAD WAS OPENED AND FIRED MEDIALLY TO THE MISFIRED STAPLES. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FOURTH FIRING ON THE STOMACH TISSUE IN THE FUNDUS AREA ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AND AN INCOMPLETE STAPLE LINE. IT WAS STATED THAT THE STAPLE LINE WAS NOT ABLE TO BE FULLY FIRED AND APPROXIMATELY THREE ROWS OF UNFORMED STAPLES WENT THROUGH THE TISSUE AT THE DISTAL. A NEW RELOAD WAS OPENED AND FIRED MEDIALLY TO THE MISFIRED STAPLES. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FOURTH FIRING ON THE STOMACH TISSUE IN THE FUNDUS AREA ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AND AN INCOMPLETE STAPLE LINE. IT WAS STATED THAT THE STAPLE LINE WAS NOT ABLE TO BE FULLY FIRED AND APPROXIMATELY THREE ROWS OF UNFORMED STAPLES WENT THROUGH THE TISSUE AT THE DISTAL. A NEW RELOAD WAS OPENED AND FIRED MEDIALLY TO THE MISFIRED STAPLES. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FOURTH FIRING ON THE STOMACH TISSUE IN THE FUNDUS AREA ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION AND AN INCOMPLETE STAPLE LINE. IT WAS STATED THAT THE STAPLE LINE WAS NOT ABLE TO BE FULLY FIRED AND APPROXIMATELY THREE ROWS OF UNFORMED STAPLES WENT THROUGH THE TISSUE AT THE DISTAL. A NEW RELOAD WAS OPENED AND FIRED MEDIALLY TO THE MISFIRED STAPLES. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ENDOSCOPIC LUMPECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE WAS UNABLE TO FIRE AND MADE A BUZZ SOUND. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ENDOSCOPIC LUMPECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE WAS UNABLE TO FIRE AND MADE A BUZZ SOUND. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ENDOSCOPIC LUMPECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE WAS UNABLE TO FIRE AND MADE A BUZZ SOUND. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC SURGERY, WHILE DOING A LEFT LOWER LOBE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE CANNOT BE SQUEEZED, AND THE SHOT CAN'T BE FIRED. A BUZZ SOUND WAS ALSO HEARD. THERE WAS NO PATIENT HARM.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC SURGERY, WHILE DOING A LEFT LOWER LOBE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE CANNOT BE SQUEEZED, AND THE SHOT CAN'T BE FIRED. A BUZZ SOUND WAS ALSO HEARD. THERE WAS NO PATIENT HARM.
17-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC SURGERY, WHILE DOING A LEFT LOWER LOBE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE CANNOT BE SQUEEZED, AND THE SHOT CAN'T BE FIRED. A BUZZ SOUND WAS ALSO HEARD. THERE WAS NO PATIENT HARM.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A THORACIC SURGERY ON A PATIENT THAT HAD RADIATION THERAPY, ON THE SIXTH FIRING OF THE STAPLER, UNKNOWN RELOAD USED, NO BUTTRESSING, THE STAPLER LOCKED OUT DURING FIRING. THE DOCTOR ATTEMPTED TO USE THE REVERSE BUTTON BUT THE REVERSE BUTTON WOULDN'T OPEN THE JAWS OF THE INSTRUMENT. THE DOCTOR TRIED TO USE MANUAL OVERRIDE BUT MANUAL OVERRIDE WOULDN'T WORK. THE DOCTOR CUT AROUND THE STAPLER AND USED SUTURE TO REPAIR THE CUT LINE. IT WAS NOT REPORTED HOW THE REMAINDER OF THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. WHEN THE DOCTOR REMOVED THE DEVICE FROM THE PATIENT, THE BLADE WAS FULLY ADVANCED, THE TISSUE WAS RADIATED BUT NOT TOO THICK.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A THORACIC SURGERY ON A PATIENT THAT HAD RADIATION THERAPY, ON THE SIXTH FIRING OF THE STAPLER, UNKNOWN RELOAD USED, NO BUTTRESSING, THE STAPLER LOCKED OUT DURING FIRING. THE DOCTOR ATTEMPTED TO USE THE REVERSE BUTTON BUT THE REVERSE BUTTON WOULDN'T OPEN THE JAWS OF THE INSTRUMENT. THE DOCTOR TRIED TO USE MANUAL OVERRIDE BUT MANUAL OVERRIDE WOULDN'T WORK. THE DOCTOR CUT AROUND THE STAPLER AND USED SUTURE TO REPAIR THE CUT LINE. IT WAS NOT REPORTED HOW THE REMAINDER OF THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. WHEN THE DOCTOR REMOVED THE DEVICE FROM THE PATIENT, THE BLADE WAS FULLY ADVANCED, THE TISSUE WAS RADIATED BUT NOT TOO THICK.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE FIRST FIRING WITH A GREEN RELOAD, GORE SEAM GUARD BUTTRESSING MATERIAL, THE STAPLER FIRED, THE RIGHT SIDE OF THE STAPLES DEPLOYED COMPLETELY BUT THE LEFT SIDE OF THE RELOAD DIDN'T DEPLOY ANY STAPLES BUT DID CUT THE STOMACH. BLEEDING OCCURRED, NO BLOOD PRODUCTS WERE NEEDED, BLEEDING WAS CONTROLLED WITH A SECOND LIKE GREEN RELOAD AND SECOND LIKE STAPLER, DEPLOYING A ROW OF STAPLES PARALLEL TO THE FIRST STAPLE LINE FIRED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
27-06-2017 | Malfunction | covidien lp llc north haven | 14-01-2019 | ENDO GIA | Separation Failure | A150301 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: DURING VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), THE DEVICE FAILED USING REINFORCED RELOAD. STAPLER LOCKED AND DID NOT CUT THROUGH. THE SURGEON HAD TO CUT OUT THE PERISTRIP STILL ATTACHED TO THE STAPLER. REPLACED BY NEW DEVICE TO COMPLETE THE CASE. NO PATIENT INJURY.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Self-Activation or Keying | A150104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS "ABNORMAL" WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS "ABNORMAL" WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Self-Activation or Keying | A150104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS "ABNORMAL" WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS "ABNORMAL" WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY 3. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY 4. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: 1. DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK 2. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY 3. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY 4. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Self-Activation or Keying | A150104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 25MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Self-Activation or Keying | A150104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY. ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Self-Activation or Keying | A150104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY, ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DIFFICULT TO FIRE. IT WAS REPORTED THAT OUR COMMERCIAL TEAM TRIED TO FIRE THE DEVICE ON THE PIG'S STOMACH AND SMALL INTESTINE, NOTED THESE ISSUES: DIFFICULT TO FIRE, NEED SQUEEZE DOWN THE FIRING TRIGGER WITH BIG FORCE, WITH AUDIBLE FEEDBACK, BUT THE SOUND IS ABNORMAL WEAK. DURING FIRING, THE ADJUSTING KNOB ROTATING AUTOMATICALLY, ROTATED THE ADJUSTING KNOB TO 1MM, BEFORE FIRING, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. DURING AIR FIRE, NOTED THE ADJUSTING KNOB ROTATING AUTOMATICALLY. THEY NOTED AFTER FIRING, THE FIRING TRIGGER COULD NOT RETURN. THERE WAS NO PATIENT INVOLVEMENT.
31-10-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | UNK-SIGADAPT | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX EN Y PROCEDURE, THE SURGICAL TECH LOADED THE STAPLE RELOAD ONTO THE HANDLE AND CYCLED THE DEVICE. THE RELOAD LOOKED VERY WOBBLY ON THE HANDLE AND THE SURGEON PLACED IT DOWN THE TROCAR AND ATTEMPTED TO PLACE THE JAWS AROUND TISSUE. THE JAWS WOULD NOT CLOSE ON THE TISSUE COMPLETELY. THE DEVICE WAS TAKEN OUT OF THE PATIENT, AND THE RELOAD WAS REMOVED THEN REATTACHED AND CYCLED AGAIN. THE RELOAD WAS STILL WOBBLY ON THE END OF THE HANDLE SO IT WAS TAKEN OFF AND REPLACED. THERE WAS NO PATIENT INJURY.
31-10-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | UNK-SIGADAPT | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX EN Y PROCEDURE, THE SURGICAL TECH LOADED THE STAPLE RELOAD ONTO THE HANDLE AND CYCLED THE DEVICE. THE RELOAD LOOKED VERY WOBBLY ON THE HANDLE AND THE SURGEON PLACED IT DOWN THE TROCAR AND ATTEMPTED TO PLACE THE JAWS AROUND TISSUE. THE JAWS WOULD NOT CLOSE ON THE TISSUE COMPLETELY. THE DEVICE WAS TAKEN OUT OF THE PATIENT, AND THE RELOAD WAS REMOVED THEN REATTACHED AND CYCLED AGAIN. THE RELOAD WAS STILL WOBBLY ON THE END OF THE HANDLE SO IT WAS TAKEN OFF AND REPLACED. THERE WAS NO PATIENT INJURY.
01-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | RELOAD- CURVED CUTTER STAPLER | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE WAS FIRED AND NO STAPLES DEPLOYED. UPON INSPECTION IT WAS NOTED THAT THE RELOAD DID NOT HAVE ANY STAPLES. A SECOND RELOAD WAS USED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCES REPORTED.
16-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THERE WAS PROBLEM WITH UNLOADING THE DEVICE AS THE JAWS WERE CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THERE WAS PROBLEM WITH UNLOADING THE DEVICE AS THE JAWS WERE CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | UNKNOWN ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THERE WAS PROBLEM WITH UNLOADING THE DEVICE AS THE JAWS WERE CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | UNKNOWN ENDO GI | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THERE WAS PROBLEM WITH UNLOADING THE DEVICE AS THE JAWS WERE CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | EEA | Loss of or Failure to Bond | A040102 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLOSTOMY CLOSURE, WHILE USING THE DEVICE IN COLORECTAL ANASTOMOSIS, THE DEVICE CUT BUT DID NOT STAPLED. THE STAPLES DID NOT HOLD THE FABRIC AT BOTH ENDS. THE EVENT LEAD TO OR EXTEND PATIENT HOSPITALIZATION FOR 20 MINUTES. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE RELOAD WAS CONNECTED TO THE HANDLE TO RESECT THE GASTRIC CORPUS DURING A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY, THE FIRING STOPPED WHEN ADVANCING FOR ABOUT 5MM. THE FIRING BUTTON WAS PRESSED AGAIN BUT THE KNIFE DID NOT ADVANCE. IT WAS STATED THAT WHEN ADVANCING WITH THE RELOAD IN ORDER TO MAKE A SPACE, THE DOCTOR EXPANDED THE FIRST STAPLE LINE. THE STAPLES WERE DEPLOYED WITH ONE SIDE OF THE STAPLE LINE TWISTED AND WAS POSSIBLE THAT JAWS WERE CLOSED IN A STATE THAT THE STOMACH ON THE BACK SIDE WAS TUCKED UP AND STACKED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE RELOAD WAS CONNECTED TO THE HANDLE TO RESECT THE GASTRIC CORPUS DURING A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY, THE FIRING STOPPED WHEN ADVANCING FOR ABOUT 5MM. THE FIRING BUTTON WAS PRESSED AGAIN BUT THE KNIFE DID NOT ADVANCE. IT WAS STATED THAT WHEN ADVANCING WITH THE RELOAD IN ORDER TO MAKE A SPACE, THE DOCTOR EXPANDED THE FIRST STAPLE LINE. THE STAPLES WERE DEPLOYED WITH ONE SIDE OF THE STAPLE LINE TWISTED AND WAS POSSIBLE THAT JAWS WERE CLOSED IN A STATE THAT THE STOMACH ON THE BACK SIDE WAS TUCKED UP AND STACKED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE RELOAD WAS CONNECTED TO THE HANDLE TO RESECT THE GASTRIC CORPUS DURING A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY, THE FIRING STOPPED WHEN ADVANCING FOR ABOUT 5MM. THE FIRING BUTTON WAS PRESSED AGAIN BUT THE KNIFE DID NOT ADVANCE. IT WAS STATED THAT WHEN ADVANCING WITH THE RELOAD IN ORDER TO MAKE A SPACE, THE DOCTOR EXPANDED THE FIRST STAPLE LINE. THE STAPLES WERE DEPLOYED WITH ONE SIDE OF THE STAPLE LINE TWISTED AND WAS POSSIBLE THAT JAWS WERE CLOSED IN A STATE THAT THE STOMACH ON THE BACK SIDE WAS TUCKED UP AND STACKED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE RELOAD WAS CONNECTED TO THE HANDLE TO RESECT THE GASTRIC CORPUS DURING A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY, THE FIRING STOPPED WHEN ADVANCING FOR ABOUT 5MM. THE FIRING BUTTON WAS PRESSED AGAIN BUT THE KNIFE DID NOT ADVANCE. IT WAS STATED THAT WHEN ADVANCING WITH THE RELOAD IN ORDER TO MAKE A SPACE, THE DOCTOR EXPANDED THE FIRST STAPLE LINE. THE STAPLES WERE DEPLOYED WITH ONE SIDE OF THE STAPLE LINE TWISTED AND WAS POSSIBLE THAT JAWS WERE CLOSED IN A STATE THAT THE STOMACH ON THE BACK SIDE WAS TUCKED UP AND STACKED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: SEE ATTACHED USER FACILITY MEDWATCH # (B)(4).
13-11-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ECHELON FLEX 45 STANDARD | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A PNEUMECTOMY PROCEDURE, THE STAPLER FAILED TO CUT UPPER RIGHT VEIN.
09-11-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RESECTION OF PULMONARY BULLAE, COULD NOT FIRE WITH EC60A,ECR60G. CHANGED ANOTHER ECR60G, COULD PERFORM GOOD FIRING. CUSTOMER SUSPECT THIS ISSUE IS NOT RELATED WITH GUN, SO ONLY COMPLIANT FOR ECR60G. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
05-12-2018 | Injury | covidien lp llc north haven | 14-01-2019 | ABSORBATACK | Loss of or Failure to Bond | A040102 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HERNIA PROCEDURE, THE TACKS OF THE DEVICE DON'T STAY AT THE MESH AND FALLING. TACK FELL INTO PATIENT CAVITY AND NOT RETRIEVED. THE PROCEDURE EXCEEDS MORE THAN 30 MINUTES. THE DOCTOR PUT SUTURES ON MESH.
05-12-2018 | Injury | covidien lp llc north haven | 14-01-2019 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HERNIA PROCEDURE, THE TACKS OF THE DEVICE DON'T STAY AT THE MESH AND FALLING. TACK FELL INTO PATIENT CAVITY AND NOT RETRIEVED. THE PROCEDURE EXCEEDS MORE THAN 30 MINUTES. THE DOCTOR PUT SUTURES ON MESH.
05-12-2018 | Injury | covidien lp llc north haven | 14-01-2019 | ABSORBATACK | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HERNIA PROCEDURE, THE TACKS OF THE DEVICE DON'T STAY AT THE MESH AND FALLING. TACK FELL INTO PATIENT CAVITY AND NOT RETRIEVED. THE PROCEDURE EXCEEDS MORE THAN 30 MINUTES. THE DOCTOR PUT SUTURES ON MESH.
05-12-2018 | Injury | covidien lp llc north haven | 14-01-2019 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HERNIA PROCEDURE, THE TACKS OF THE DEVICE DON'T STAY AT THE MESH AND FALLING. TACK FELL INTO PATIENT CAVITY AND NOT RETRIEVED. THE PROCEDURE EXCEEDS MORE THAN 30 MINUTES. THE DOCTOR PUT SUTURES ON MESH.
27-06-2017 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Separation Failure | A150301 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER: DURING VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), THE DEVICE FAILED USING REINFORCED RELOAD. STAPLER LOCKED AND DID NOT CUT THROUGH. THE SURGEON HAD TO CUT OUT THE PERISTRIP STILL ATTACHED TO THE STAPLER. REPLACED BY NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-10-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE THE DEVICE HAD AN UNKNOWN ISSUE.
29-10-2018 | Injury | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHILE ON SECOND FIRING DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THEY DID NOT NOTICE THAT THE TISSUE WAS NOT VERY THICK. THE STAPLER STOPPED FIRING ON ABOUT ONE THIRD OF THE WAY AND SHOWED THE RED WEDGE INDICATING THAT THE TISSUE WAS OUT OF RANGE FOR THE SELECTED TISSUE. THEY OPENED THE STAPLER TO FIND THAT IT HAD NOT FULLY FIRED AND THAT THERE WERE PARTIALLY FIRED STAPLES PRESENT. THEY EVENTUALLY REMOVED THE UNFORMED STAPLES AND TRIMMED THE SEAM GUARD IN PREPARATION FOR THE NEXT FIRING. THE NEXT FIRING STOPPED ALSO SHOWING THE WEDGE ONCE AGAIN AFTER IT WAS PARTIALLY FIRED. THE SEAM GUARD WAS LAIRED INCREASING THE THICKNESS THAT THE STAPLER HAD TO GET THROUGH. THE THIRD LOAD WAS FIRED, AND THE SAME PROBLEM OCCURRED. THE SURGEON HAD PUT IN AN EXTRA 5MM PORT AS THERE WAS BLEEDING AFTER THE FIRST INCIDENT FORM THE INFERIOR ASPECT OF THE REMNANT JUST BELOW WHERE THE MALFORMED STAPLES HAD BEEN. A CLIP APPLIER WAS USED TO CONTROL THE BLEEDING. THEY ALSO USED 15MM PORT FOR SUBSEQUENT FIRINGS. THEN, EVENTUALLY A MANUAL HANDLE WAS OPENED AND ONLY PARTIALLY FIRED AND THEN A SECOND ONE WAS OPENED AND ONLY PARTIALLY FIRED, AND BOTH HANDLES MADE LOUD CRACKING SOUNDS AS THEY WERE BEING FIRED. THERE WERE FOUR LOUD CRACKED WERE HEARD ON EACH STAPLER. THEY OPENED AND USED ANOTHER STAPLER TO COMPLETE THE CASE. THE SURGEON ESTIMATED ABOUT 150 ML OF BLOOD LOSS. THIS ALSO RESULTED TO EXTEND SURGICAL TIME OF 40 MINUTES DELAY.
29-10-2018 | Injury | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE ON SECOND FIRING DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THEY DID NOT NOTICE THAT THE TISSUE WAS NOT VERY THICK. THE STAPLER STOPPED FIRING ON ABOUT ONE THIRD OF THE WAY AND SHOWED THE RED WEDGE INDICATING THAT THE TISSUE WAS OUT OF RANGE FOR THE SELECTED TISSUE. THEY OPENED THE STAPLER TO FIND THAT IT HAD NOT FULLY FIRED AND THAT THERE WERE PARTIALLY FIRED STAPLES PRESENT. THEY EVENTUALLY REMOVED THE UNFORMED STAPLES AND TRIMMED THE SEAM GUARD IN PREPARATION FOR THE NEXT FIRING. THE NEXT FIRING STOPPED ALSO SHOWING THE WEDGE ONCE AGAIN AFTER IT WAS PARTIALLY FIRED. THE SEAM GUARD WAS LAIRED INCREASING THE THICKNESS THAT THE STAPLER HAD TO GET THROUGH. THE THIRD LOAD WAS FIRED, AND THE SAME PROBLEM OCCURRED. THE SURGEON HAD PUT IN AN EXTRA 5MM PORT AS THERE WAS BLEEDING AFTER THE FIRST INCIDENT FORM THE INFERIOR ASPECT OF THE REMNANT JUST BELOW WHERE THE MALFORMED STAPLES HAD BEEN. A CLIP APPLIER WAS USED TO CONTROL THE BLEEDING. THEY ALSO USED 15MM PORT FOR SUBSEQUENT FIRINGS. THEN, EVENTUALLY A MANUAL HANDLE WAS OPENED AND ONLY PARTIALLY FIRED AND THEN A SECOND ONE WAS OPENED AND ONLY PARTIALLY FIRED, AND BOTH HANDLES MADE LOUD CRACKING SOUNDS AS THEY WERE BEING FIRED. THERE WERE FOUR LOUD CRACKED WERE HEARD ON EACH STAPLER. THEY OPENED AND USED ANOTHER STAPLER TO COMPLETE THE CASE. THE SURGEON ESTIMATED ABOUT 150 ML OF BLOOD LOSS. THIS ALSO RESULTED TO EXTEND SURGICAL TIME OF 40 MINUTES DELAY.
29-10-2018 | Injury | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHILE ON SECOND FIRING DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THEY DID NOT NOTICE THAT THE TISSUE WAS NOT VERY THICK. THE STAPLER STOPPED FIRING ON ABOUT ONE THIRD OF THE WAY AND SHOWED THE RED WEDGE INDICATING THAT THE TISSUE WAS OUT OF RANGE FOR THE SELECTED TISSUE. THEY OPENED THE STAPLER TO FIND THAT IT HAD NOT FULLY FIRED AND THAT THERE WERE PARTIALLY FIRED STAPLES PRESENT. THEY EVENTUALLY REMOVED THE UNFORMED STAPLES AND TRIMMED THE SEAM GUARD IN PREPARATION FOR THE NEXT FIRING. THE NEXT FIRING STOPPED ALSO SHOWING THE WEDGE ONCE AGAIN AFTER IT WAS PARTIALLY FIRED. THE SEAM GUARD WAS LAIRED INCREASING THE THICKNESS THAT THE STAPLER HAD TO GET THROUGH. THE THIRD LOAD WAS FIRED, AND THE SAME PROBLEM OCCURRED. THE SURGEON HAD PUT IN AN EXTRA 5MM PORT AS THERE WAS BLEEDING AFTER THE FIRST INCIDENT FORM THE INFERIOR ASPECT OF THE REMNANT JUST BELOW WHERE THE MALFORMED STAPLES HAD BEEN. A CLIP APPLIER WAS USED TO CONTROL THE BLEEDING. THEY ALSO USED 15MM PORT FOR SUBSEQUENT FIRINGS. THEN, EVENTUALLY A MANUAL HANDLE WAS OPENED AND ONLY PARTIALLY FIRED AND THEN A SECOND ONE WAS OPENED AND ONLY PARTIALLY FIRED, AND BOTH HANDLES MADE LOUD CRACKING SOUNDS AS THEY WERE BEING FIRED. THERE WERE FOUR LOUD CRACKED WERE HEARD ON EACH STAPLER. THEY OPENED AND USED ANOTHER STAPLER TO COMPLETE THE CASE. THE SURGEON ESTIMATED ABOUT 150 ML OF BLOOD LOSS. THIS ALSO RESULTED TO EXTEND SURGICAL TIME OF 40 MINUTES DELAY.
29-10-2018 | Injury | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE ON SECOND FIRING DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THEY DID NOT NOTICE THAT THE TISSUE WAS NOT VERY THICK. THE STAPLER STOPPED FIRING ON ABOUT ONE THIRD OF THE WAY AND SHOWED THE RED WEDGE INDICATING THAT THE TISSUE WAS OUT OF RANGE FOR THE SELECTED TISSUE. THEY OPENED THE STAPLER TO FIND THAT IT HAD NOT FULLY FIRED AND THAT THERE WERE PARTIALLY FIRED STAPLES PRESENT. THEY EVENTUALLY REMOVED THE UNFORMED STAPLES AND TRIMMED THE SEAM GUARD IN PREPARATION FOR THE NEXT FIRING. THE NEXT FIRING STOPPED ALSO SHOWING THE WEDGE ONCE AGAIN AFTER IT WAS PARTIALLY FIRED. THE SEAM GUARD WAS LAIRED INCREASING THE THICKNESS THAT THE STAPLER HAD TO GET THROUGH. THE THIRD LOAD WAS FIRED, AND THE SAME PROBLEM OCCURRED. THE SURGEON HAD PUT IN AN EXTRA 5MM PORT AS THERE WAS BLEEDING AFTER THE FIRST INCIDENT FORM THE INFERIOR ASPECT OF THE REMNANT JUST BELOW WHERE THE MALFORMED STAPLES HAD BEEN. A CLIP APPLIER WAS USED TO CONTROL THE BLEEDING. THEY ALSO USED 15MM PORT FOR SUBSEQUENT FIRINGS. THEN, EVENTUALLY A MANUAL HANDLE WAS OPENED AND ONLY PARTIALLY FIRED AND THEN A SECOND ONE WAS OPENED AND ONLY PARTIALLY FIRED, AND BOTH HANDLES MADE LOUD CRACKING SOUNDS AS THEY WERE BEING FIRED. THERE WERE FOUR LOUD CRACKED WERE HEARD ON EACH STAPLER. THEY OPENED AND USED ANOTHER STAPLER TO COMPLETE THE CASE. THE SURGEON ESTIMATED ABOUT 150 ML OF BLOOD LOSS. THIS ALSO RESULTED TO EXTEND SURGICAL TIME OF 40 MINUTES DELAY.
14-12-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | UNKNOWN ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING THE RESECTION OF THE VESSELS; THE RELOAD COULD NOT BE FIRED (BLADE AND STAPLES DID NOT DEPLOY) AFTER PUSHING THE GREEN BUTTON. IT WAS NOTED THAT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER LOAD WAS THEN USED WITH SUCCESS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 14-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING THE RESECTION OF THE VESSELS; THE RELOAD COULD NOT BE FIRED (BLADE AND STAPLES DID NOT DEPLOY) AFTER PUSHING THE GREEN BUTTON. IT WAS NOTED THAT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER LOAD WAS THEN USED WITH SUCCESS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-10-2018 | Malfunction | us surgical puerto rico | 14-01-2019 | ENDO GIA ULTRA | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER BROKE OFF. THE DEVICE WAS CHANGED TO CONTINUE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2018 | Injury | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, THE STAPLER WAS USE FOR THE TERMINAL ANASTOMOSIS. DURING THE FIRING IT WAS HEARD A NOISE DIFFERENT FROM THE USUAL ONE (CUTTING WASHER). WHEN THE HEAD WAS OPENED, THE STAPLER HAD CUT BUT ONLY SOME CLIPS WERE APPLIED. SOME CLIPS WERE MALFORMED AND OTHER ABSENT. IT WAS PERFORMED A NEW RESECTION OF THE RECTUM WITH A NEW CDH OF THE SAME LOT NUMBER: THE SAME ISSUE OCCURRED. THE ANASTOMOSIS WAS COMPLETED MANUALLY WITH SUTURES. A COLOSTOMY WAS PERFORMED AS PROTECTION.
17-12-2018 | Injury | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, THE STAPLER WAS USE FOR THE TERMINAL ANASTOMOSIS. DURING THE FIRING IT WAS HEARD A NOISE DIFFERENT FROM THE USUAL ONE (CUTTING WASHER). WHEN THE HEAD WAS OPENED, THE STAPLER HAD CUT BUT ONLY SOME CLIPS WERE APPLIED. SOME CLIPS WERE MALFORMED AND OTHER ABSENT. IT WAS PERFORMED A NEW RESECTION OF THE RECTUM WITH A NEW CDH OF THE SAME LOT NUMBER: THE SAME ISSUE OCCURRED. THE ANASTOMOSIS WAS COMPLETED MANUALLY WITH SUTURES. A COLOSTOMY WAS PERFORMED AS PROTECTION.
17-12-2018 | Injury | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, THE STAPLER WAS USE FOR THE TERMINAL ANASTOMOSIS. DURING THE FIRING IT WAS HEARD A NOISE DIFFERENT FROM THE USUAL ONE (CUTTING WASHER). WHEN THE HEAD WAS OPENED, THE STAPLER HAD CUT BUT ONLY SOME CLIPS WERE APPLIED. SOME CLIPS WERE MALFORMED AND OTHER ABSENT. IT WAS PERFORMED A NEW RESECTION OF THE RECTUM WITH A NEW CDH OF THE SAME LOT NUMBER: THE SAME ISSUE OCCURRED. THE ANASTOMOSIS WAS COMPLETED MANUALLY WITH SUTURES. A COLOSTOMY WAS PERFORMED AS PROTECTION.
17-12-2018 | Injury | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Noise, Audible | A0508 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, THE STAPLER WAS USE FOR THE TERMINAL ANASTOMOSIS. DURING THE FIRING IT WAS HEARD A NOISE DIFFERENT FROM THE USUAL ONE (CUTTING WASHER). WHEN THE HEAD WAS OPENED, THE STAPLER HAD CUT BUT ONLY SOME CLIPS WERE APPLIED. SOME CLIPS WERE MALFORMED AND OTHER ABSENT. IT WAS PERFORMED A NEW RESECTION OF THE RECTUM WITH A NEW CDH OF THE SAME LOT NUMBER: THE SAME ISSUE OCCURRED. THE ANASTOMOSIS WAS COMPLETED MANUALLY WITH SUTURES. A COLOSTOMY WAS PERFORMED AS PROTECTION.
17-12-2018 | Injury | ethicon endo-surgery | 14-01-2019 | ILS 29MM, CURVED | Noise, Audible | A0508 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, THE STAPLER WAS "USE" FOR THE TERMINAL ANASTOMOSIS. DURING THE FIRING IT WAS HEARD A NOISE DIFFERENT FROM THE USUAL ONE (CUTTING WASHER). WHEN THE HEAD WAS OPENED, THE STAPLER HAD CUT BUT ONLY SOME CLIPS WERE APPLIED. SOME CLIPS WERE MALFORMED AND OTHER ABSENT. "IT WAS PERFORMED A NEW RESECTION OF THE RECTUM WITH A NEW CDH OF THE SAME LOT NUMBER"; THE SAME ISSUE OCCURRED. THE ANASTOMOSIS WAS COMPLETED MANUALLY WITH SUTURES. A COLOSTOMY WAS PERFORMED AS PROTECTION.
20-09-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLER WOULD NOT FIRE DURING PROCEDURE. RETRIEVED ANOTHER STAPLER TO USE. NO PATIENT CONSEQUENCE REPORTED.
30-11-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | PROXIMATE*LINEAR STAP RELOAD | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHILE THE LOBECTOMY SURGERY, XR30G DID NOT CLOSE THE BRONCHUS TISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ETS ARTIC LNR CUTR NO KNIFE | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE WOULD NOT FIRE. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME.
12-11-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | PWRD ECH FLEX 60MM 34CM SHAFT | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LEFT LATERAL HEPATECTOMY, WHEN SEVERING THE HEPATIC PEDICLE, COULD NOT FIRE FARTHER AFTER FIRED A HALF. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
23-11-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, AFTER THE FIRST SHOT THE MACHINE STAYED OPEN. IT WAS STRANGE FROM THE BEGINNING. PROBLEMS FROM THE SECOND SHOT BECAUSE YOU COULD NOT EASILY INTRODUCE THE ENDOCORTADORA BY THE TROCAR. COULD NOT BE USED THE SEAMGUARD DURING THE PROCEDURE AFTER THE SECOND SHOT. SEAMGUARDS COULD NOT BE USED AS THEY BROKE AND HAD DIFFICULTY INTRODUCING THE TROCAR IN THE FOLLOWING SHOTS. THERE WERE NO PATIENT CONSEQUENCE REPORTED.
18-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE THE DEVICE WAS FIRED USING A BLACK RELOAD. THE DEVICE CUT BUT DID NOT STAPLE. THE PROCEDURE WAS COMPLETED BY OVER SEWING THE STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THERE WAS ONLY PARTIAL FIRING, NOT FULL FIRING OF DEVICE. NO PATIENT CONSEQUENCE REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GALLBLADDER A PSE45A WAS LOADED WITH GST45B, BUT WOULD NOT RE-OPEN AFTER IT WAS CLOSED. IT WAS NEVER PUT INSIDE PATIENT OR USED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ECHELON FLEX 45 STANDARD | Unsealed Device Packaging | A020503 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DISTRIBUTOR INFORMED: THERE IS CASE OF PRODUCT OF BREACHED SEALING. THE SHIPPING BOXES WERE IN PERFECT CONDITION.
21-12-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ECHELON FLEX 45 STANDARD | Unsealed Device Packaging | A020503 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DISTRIBUTOR INFORMED: THERE IS CASE OF PRODUCT OF BREACHED SEALING. THE SHIPPING BOXES WERE IN PERFECT CONDITION.
01-01-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | GST BLUE RELOAD, 60MM, 6 ROW | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE GST60B RELOADS WAS LOADED INTO STAPLER. TECH NOTICED THAT IT APPEARED TO BE DAMAGED UPON INSERTION IN ABDOMINAL CAVITY. A NEW RELOAD WAS REPLACED BEFORE FIRING. THERE WAS NO DELAY AND SURGERY WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 14-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN CASE: FIRST RELOAD WOULD NOT FIRE; THIRD RELOAD QUIT HALF WAY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
03-11-2017 | Malfunction | us surgical puerto rico | 14-01-2019 | UNKNOWN GIA PRODUCT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SURGERY, THE STAPLER DID NOT FIRE. THE LOAD WAS CHANGED TO CONTINUE THE SURGERY. THERE WAS NO INJURY.
03-11-2017 | Malfunction | covidien lp llc north haven | 14-01-2019 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SURGERY, THE STAPLER DID NOT FIRE. THE LOAD WAS CHANGED TO CONTINUE THE SURGERY. THERE WAS NO INJURY.
03-11-2017 | Malfunction | covidien lp llc north haven | 14-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SURGERY, THE STAPLER DID NOT FIRE. THE LOAD WAS CHANGED TO CONTINUE THE SURGERY. THERE WAS NO INJURY.
25-03-2019 | Malfunction | covidien lp llc north haven | 14-01-2019 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS A DIFFICULTY WITH LOADING THE RELOAD INTO THE HANDLE. ANOTHER IDENTICAL PRODUCT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, IN A WEDGE SEGMENTAL RESECTION PROCEDURE, THE FIRST FIRING HAD NO PROBLEM. THE SURGEON SET A NEW RELOAD AND CONFIRMED THE GREEN BUTTON FLASHED, THEN TRIED TO FIRE BUT FAILED. THEY REPLACED IT WITH ANOTHER RELOAD, THE SURGEON SET IT AND CONFIRMED THE GREEN BUTTON FLASHED, THEN HE TRIED TO FIRE THE DEVICE, BUT THE FIRING WAS NOT PERFORMED, EITHER. THEY USED ANOTHER DEVICE AND THE PROCEDURE WAS COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, IN A WEDGE SEGMENTAL RESECTION PROCEDURE, THE FIRST FIRING HAD NO PROBLEM. THE SURGEON SET A NEW RELOAD AND CONFIRMED THE GREEN BUTTON FLASHED, THEN TRIED TO FIRE BUT FAILED. THEY REPLACED IT WITH ANOTHER RELOAD, THE SURGEON SET IT AND CONFIRMED THE GREEN BUTTON FLASHED, THEN HE TRIED TO FIRE THE DEVICE, BUT THE FIRING WAS NOT PERFORMED, EITHER. THEY USED ANOTHER DEVICE AND THE PROCEDURE WAS COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RECTOSIGMOIDECTOMY, DURING RECTAL TRANSECTION, THE STAPLER MADE A HUGE SOUND AND DID NOT FIRE THE RELOAD WHEN THE SURGEON SQUEEZED THE HANDLE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD AND STAPLER WERE OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RECTOSIGMOIDECTOMY, DURING RECTAL TRANSECTION, THE STAPLER MADE A HUGE SOUND AND DID NOT FIRE THE RELOAD WHEN THE SURGEON SQUEEZED THE HANDLE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD AND STAPLER WERE OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RECTOSIGMOIDECTOMY, DURING RECTAL TRANSECTION, THE STAPLER MADE A HUGE SOUND AND DID NOT FIRE THE RELOAD WHEN THE SURGEON SQUEEZED THE HANDLE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD AND STAPLER WERE OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | ETS*45 ENDO LIN CUT RELOAD WH | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE TR45W CARTRIDGE RELOAD WOULD NOT FIRE. NO OTHER INFORMATION IS AVAILABLE.
05-11-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Difficult or Delayed Positioning | A150203 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHEN THE DEVICE WAS USED ON THE SECOND FIRING, IT WAS DIFFICULT TO FIRE MAKING A LOUD CRACKING SOUND. THE NEXT SQUEEZED OF THE HANDLE ALSO MADE A LOUD CRACKING SOUND. THE LOAD WAS ABLE TO BE FULLY FIRED. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHEN THE DEVICE WAS USED ON THE SECOND FIRING, IT WAS DIFFICULT TO FIRE MAKING A LOUD CRACKING SOUND. THE NEXT SQUEEZED OF THE HANDLE ALSO MADE A LOUD CRACKING SOUND. THE LOAD WAS ABLE TO BE FULLY FIRED. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-11-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WE HAD A COMPLETE FAILURE OF STAPLING IN THE PENULTIMATE STAPLE WITH OPENING OF THE ENTIRE STAPLED AREA. THE SYSTEM WAS ECHELON POWER AND THE RELOAD USED FOR BODY AND BOTTOM TRANSITION REGION WAS BLUE AS PER MANUFACTURER'S RECOMMENDATION. WE NEED TO MAKE MANUAL RAFFIA FOR BEING A SLEEVE WITH NO SPACE FOR NEW STAPLERS.
18-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLES WERE UNFORMED WHEN THE DEVICE WAS USED ON THE GLISSON'S SHEATH. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
13-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC SURGERY, ON STAPLING THE LUNG, AFTER USE OF THE CHARGER, THE NURSE WANTED TO TAKE THE DEVICE OUT TO USE ANOTHER ONE AND IT WAS IMPOSSIBLE TO GET IT OUT. THE KNIFE OF THE CHARGER JAMMED IN THE CLAMP. THE KNIFE BLADE WAS DIFFICULT TO RETRACT. THEY CHANGED CLAMP AND THERE WAS NO PROBLEM AFTERWARDS. THEY TOOK ANOTHER HANDLE AND LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Mechanical Jam | A0506 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC SURGERY, ON STAPLING THE LUNG, AFTER USE OF THE CHARGER, THE NURSE WANTED TO TAKE THE DEVICE OUT TO USE ANOTHER ONE AND IT WAS IMPOSSIBLE TO GET IT OUT. THE KNIFE OF THE CHARGER JAMMED IN THE CLAMP. THE KNIFE BLADE WAS DIFFICULT TO RETRACT. THEY CHANGED CLAMP AND THERE WAS NO PROBLEM AFTERWARDS. THEY TOOK ANOTHER HANDLE AND LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Retraction Problem | A0510 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC SURGERY, ON STAPLING THE LUNG, AFTER USE OF THE CHARGER, THE NURSE WANTED TO TAKE THE DEVICE OUT TO USE ANOTHER ONE AND IT WAS IMPOSSIBLE TO GET IT OUT. THE KNIFE OF THE CHARGER JAMMED IN THE CLAMP. THE KNIFE BLADE WAS DIFFICULT TO RETRACT. THEY CHANGED CLAMP AND THERE WAS NO PROBLEM AFTERWARDS. THEY TOOK ANOTHER HANDLE AND LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC SURGERY, ON STAPLING THE LUNG, AFTER USE OF THE CHARGER, THE NURSE WANTED TO TAKE THE DEVICE OUT TO USE ANOTHER ONE AND IT WAS IMPOSSIBLE TO GET IT OUT. THE KNIFE OF THE CHARGER JAMMED IN THE CLAMP. THE KNIFE BLADE WAS DIFFICULT TO RETRACT. THEY CHANGED CLAMP AND THERE WAS NO PROBLEM AFTERWARDS. THEY TOOK ANOTHER HANDLE AND LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Mechanical Jam | A0506 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC SURGERY, ON STAPLING THE LUNG, AFTER USE OF THE CHARGER, THE NURSE WANTED TO TAKE THE DEVICE OUT TO USE ANOTHER ONE AND IT WAS IMPOSSIBLE TO GET IT OUT. THE KNIFE OF THE CHARGER JAMMED IN THE CLAMP. THE KNIFE BLADE WAS DIFFICULT TO RETRACT. THEY CHANGED CLAMP AND THERE WAS NO PROBLEM AFTERWARDS. THEY TOOK ANOTHER HANDLE AND LOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp | 15-01-2019 | ABSORBATACK | Material Twisted/Bent | A040609 | No Known Impact Or Consequence To Patient | Event Description: ABSTACK DEVICE BENT USED ON MESH DURING HERNIA REPAIR. NO ADVERSE EFFECT ON PATIENT. DEVICE WAS WORKING PROPERLY THEN BENT WHEN MOVING POSITION. NEW DEVICE OPENED AND "USE".
01-11-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, IN A KIDNEY/PANCREAS TRANSPLANT, DEVICE FAILED. THE STAPLE LOAD WOULD NOT OPEN AND WOULDN'T FIRE. THE ISSUE WAS NOTICED DURING BENCHING AND DID NOT REACH THE PATIENT. THERE WAS NO PATIENT INJURY.
01-11-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, IN A KIDNEY/PANCREAS TRANSPLANT, DEVICE FAILED. THE STAPLE LOAD WOULD NOT OPEN AND WOULDN'T FIRE. THE ISSUE WAS NOTICED DURING BENCHING AND DID NOT REACH THE PATIENT. THERE WAS NO PATIENT INJURY.
01-11-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, IN A KIDNEY/PANCREAS TRANSPLANT, DEVICE FAILED. THE STAPLE LOAD WOULD NOT OPEN AND WOULDN'T FIRE. THE ISSUE WAS NOTICED DURING BENCHING AND DID NOT REACH THE PATIENT. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTRECTOMY PROCEDURE, THE STAPLER FIRED TWO SHOTS AND THE THIRD WAS LOCKED. IT WAS NOTED THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. IT WAS REPORTED IT WAS NECESSARY TO USE ANOTHER STAPLER. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | PROXIMATE*PPH PROCEDURE SET | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLING WAS HEMI-CIRCUMFERENTIAL, EVEN THOUGH THE STANDARD PROCEDURE WAS RESPECTED AND THE BURSA WAS COMPLETELY TIGHT. SO, THE TISSUE DONUT WAS INCOMPLETE. THE SURGEON HAD TO RE-PERFORM A PROCEDURE. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE (SO, A SECOND STAPLING LINE WAS PERFORMED). PATIENT CONSEQUENCE: PROBABLY, MORE PAINFUL FOR THE PATIENT.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ROTICULATOR | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RECTOSIGMOIDECTOMY PROCEDURE, WHILE STAPLING IN THE MEDIUM STRAIGHT AND IRRADIATED STRAIGHT FABRIC, THE STAPLER FAILED TO COMPLETE THE CLOSURE AND WITH THE RESISTANCE IT BROKE. THE DEVICE WAS DIFFICULT OR UNABLE TO CLOSE AND THE CLOSING LEVER BROKE. NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ROTICULATOR | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RECTOSIGMOIDECTOMY PROCEDURE, WHILE STAPLING IN THE MEDIUM STRAIGHT AND IRRADIATED STRAIGHT FABRIC, THE STAPLER FAILED TO COMPLETE THE CLOSURE AND WITH THE RESISTANCE IT BROKE. THE DEVICE WAS DIFFICULT OR UNABLE TO CLOSE AND THE CLOSING LEVER BROKE. NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | ILS 29MM, STRAIGHT | Inappropriate Audible Prompt/Feedback | A090101 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT CHIEF USED THE SDH29A IN LAR SURGERY. HE DID NOT "AWARE" AUDIO FEEDBACK WHEN SDH29A WAS FIRING. EVEN HE HAD CHECKED THE ANASTOMOSIS AREA IS PERFECT, HE ASKED ME REPORT THE PRODUCT COMPLAINT. THERE WAS NO DELAY AND THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
18-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE COULD NOT BE ADVANCED EVEN AFTER MULTIPLE TRIES. THE CASE WAS COMPLETED BY A NEW DEVICE AND A NEW CARTRIDGE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE COULD NOT BE ADVANCED EVEN AFTER MULTIPLE TRIES. THE CASE WAS COMPLETED BY A NEW DEVICE AND A NEW CARTRIDGE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE COULD NOT BE ADVANCED EVEN AFTER MULTIPLE TRIES. THE CASE WAS COMPLETED BY A NEW DEVICE AND A NEW CARTRIDGE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE COULD NOT BE ADVANCED EVEN AFTER MULTIPLE TRIES. THE CASE WAS COMPLETED BY A NEW DEVICE AND A NEW CARTRIDGE. THERE WAS NO PATIENT INJURY.
05-11-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN THE DEVICE WAS CLOSE AROUND THE STOMACH TISSUE, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED AND WAS TRIED TO FIRE, THE HANDLE LOCKED IN POSITION PRIOR TO FIRING AND WAS UNABLE TO ADVANCE TO ACTIVATE THE STAPLER. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE SHOWED FAILURE FROM THE SECOND FIRING AFTER THE TRIGGER WAS TRIGGERED, AFTER WHICH THE STAPLER AND THE LOAD WERE LOCKED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND DID NOT FIRE. NEW HANDLE AND RELOAD WERE OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE SHOWED FAILURE FROM THE SECOND FIRING AFTER THE TRIGGER WAS TRIGGERED, AFTER WHICH THE STAPLER AND THE LOAD WERE LOCKED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND DID NOT FIRE. NEW HANDLE AND RELOAD WERE OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE SHOWED FAILURE FROM THE SECOND FIRING AFTER THE TRIGGER WAS TRIGGERED, AFTER WHICH, THE STAPLER AND THE LOAD WERE LOCKED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE AND DID NOT FIRE. NEW HANDLE AND RELOAD WERE OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE SHOWED FAILURE FROM THE SECOND FIRING AFTER THE TRIGGER WAS TRIGGERED, AFTER WHICH, THE STAPLER AND THE LOAD WERE LOCKED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE AND DID NOT FIRE. NEW HANDLE AND RELOAD WERE OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVE A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS A COMPLICATION OF A LEAK. REOPERATION WAS DONE TO CORRECT THE ISSUE.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVE A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS A COMPLICATION OF A LEAK. REOPERATION WAS DONE TO CORRECT THE ISSUE.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | TRI-STAPLE 2.0 | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVE A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS A COMPLICATION OF A LEAK. REOPERATION WAS DONE TO CORRECT THE ISSUE.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVE A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS A COMPLICATION OF A LEAK. REOPERATION WAS DONE TO CORRECT THE ISSUE.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 15 DAYS AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS LEAK TOWARDS TOP OF THE STOMACH. PATIENT WAS BROUGHT BACK TO THE HOSPITAL TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 15 DAYS AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS LEAK TOWARDS TOP OF THE STOMACH. PATIENT WAS BROUGHT BACK TO THE HOSPITAL TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 15 DAYS AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS LEAK TOWARDS TOP OF THE STOMACH. PATIENT WAS BROUGHT BACK TO THE HOSPITAL TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 15 DAYS AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS LEAK TOWARDS TOP OF THE STOMACH. PATIENT WAS BROUGHT BACK TO THE HOSPITAL TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION.
12-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A RECTOSIGMOIDECTOMY PROCEDURE, STAPLER STAPLED BUT DID NOT CUT.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | TRI-STAPLE 2.0 | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 15 DAYS AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS LEAK TOWARDS TOP OF THE STOMACH. PATIENT WAS BROUGHT BACK TO THE HOSPITAL TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 15 DAYS AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS LEAK TOWARDS TOP OF THE STOMACH. PATIENT WAS BROUGHT BACK TO THE HOSPITAL TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 15 DAYS AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS LEAK TOWARDS TOP OF THE STOMACH. PATIENT WAS BROUGHT BACK TO THE HOSPITAL TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 15 DAYS AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS LEAK TOWARDS TOP OF THE STOMACH. PATIENT WAS BROUGHT BACK TO THE HOSPITAL TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE PATIENT HAD A LEAK NEEDED TO BE TENDED TO BUT IT WAS STATED THAT THERE WAS NO RE-OPERATION NEEDED. THIS RESULTED TO AN EXTENDED HOSPITALIZATION BUT IT WAS UNKNOWN FOR HOW LONG.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE PATIENT HAD A LEAK NEEDED TO BE TENDED TO BUT IT WAS STATED THAT THERE WAS NO RE-OPERATION NEEDED. THIS RESULTED TO AN EXTENDED HOSPITALIZATION BUT IT WAS UNKNOWN FOR HOW LONG.
26-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING AN OPEN COLECTOMY, THE FIRING KNOB WAS BROKEN OFF ON THE WAY OF THE 4TH FIRING OF FEEA. NO PIECES WERE LEFT INSIDE THE PATIENT. NO ADDITIONAL TREATMENT WAS REQUIRED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
26-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | 75MM SELECTABLE NEW TLC | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING AN OPEN COLECTOMY, THE FIRING KNOB WAS BROKEN OFF ON THE WAY OF THE 4TH FIRING OF FEEA. NO PIECES WERE LEFT INSIDE THE PATIENT. NO ADDITIONAL TREATMENT WAS REQUIRED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | TRI-STAPLE 2.0 | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE PATIENT HAD A LEAK NEEDED TO BE TENDED TO BUT IT WAS STATED THAT THERE WAS NO RE-OPERATION NEEDED. THIS RESULTED TO AN EXTENDED HOSPITALIZATION BUT IT WAS UNKNOWN FOR HOW LONG.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE PATIENT HAD A LEAK NEEDED TO BE TENDED TO BUT IT WAS STATED THAT THERE WAS NO RE-OPERATION NEEDED. THIS RESULTED TO AN EXTENDED HOSPITALIZATION BUT IT WAS UNKNOWN FOR HOW LONG.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTRECTOMY PROCEDURE, THE STAPLER FIRED TWO SHOTS AND THE THIRD WAS LOCKED. IT WAS NOTED THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. IT WAS REPORTED IT WAS NECESSARY TO USE ANOTHER STAPLER. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | POWERED PLUS STANDARD 45 | Leak/Splash | A0504 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP APPENDECTOMY, THE STAPLES DIDN'T OCCLUDE THE BLOOD VESSEL. VESSELS NEAR THE APPENDAGE, THEY WERE OOZING. UNKNOWN FOR SURE HOW THE CASE WAS COMPLETED, THOUGHT ABOUT OVER STAPLING. THE VESSEL THAT FEEDS THE APPENDIX WHEN HE WENT TO CLIP IT AND SEAL IT WITH THE STAPLER, IT WAS OOZING AND IT HAD SOME BLOOD SPURTING. IT DID CUT THE TISSUE. THE APPEARANCE OF THE DEPLOYED STAPLES LOOK NORMAL, DIDN'T DEPLOY ALL OF THEM. THE STAPLES THAT MADE IT ON THE TISSUE WERE NOT MALFORMED. THEY LOOK LIKE THEY HAVE THE RIGHT FORM TO THEM. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | POWERED PLUS STANDARD 45 | Leak/Splash | A0504 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP APPENDECTOMY, THE STAPLES DIDN'T OCCLUDE THE BLOOD VESSEL. VESSELS NEAR THE APPENDAGE, THEY WERE OOZING. UNKNOWN FOR SURE HOW THE CASE WAS COMPLETED, THOUGHT ABOUT OVER STAPLING. THE VESSEL THAT FEEDS THE APPENDIX WHEN HE WENT TO CLIP IT AND SEAL IT WITH THE STAPLER, IT WAS OOZING AND IT HAD SOME BLOOD SPURTING. IT DID CUT THE TISSUE. THE APPEARANCE OF THE DEPLOYED STAPLES LOOK NORMAL, DIDN'T DEPLOY ALL OF THEM. THE STAPLES THAT MADE IT ON THE TISSUE WERE NOT MALFORMED. THEY LOOK LIKE THEY HAVE THE RIGHT FORM TO THEM. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER 20 POST-OPERATIVE-DAY OF SLEEVE GASTRECTOMY, PATIENT HAD TO BE BROUGHT BACK IN DUE TO POST OPERATIVE LEAK. ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER 20 POST-OPERATIVE-DAY OF SLEEVE GASTRECTOMY, PATIENT HAD TO BE BROUGHT BACK IN DUE TO POST OPERATIVE LEAK. ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
01-01-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | POWERED PLUS STANDARD 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP APPENDECTOMY, THE STAPLES DIDN'T OCCLUDE THE BLOOD VESSEL. VESSELS NEAR THE APPENDAGE, THEY WERE OOZING. UNKNOWN FOR SURE HOW THE CASE WAS COMPLETED, THOUGHT ABOUT OVER STAPLING. THE VESSEL THAT FEEDS THE APPENDIX WHEN HE WENT TO CLIP IT AND SEAL IT WITH THE STAPLER, IT WAS OOZING AND IT HAD SOME BLOOD SPURTING. IT DID CUT THE TISSUE. THE APPEARANCE OF THE DEPLOYED STAPLES LOOK NORMAL, DIDN'T DEPLOY ALL OF THEM. THE STAPLES THAT MADE IT ON THE TISSUE WERE NOT MALFORMED. THEY LOOK LIKE THEY HAVE THE RIGHT FORM TO THEM. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER 20 POST-OPERATIVE-DAY OF SLEEVE GASTRECTOMY, PATIENT HAD TO BE BROUGHT BACK IN DUE TO POST OPERATIVE LEAK. ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
15-01-2019 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER 20 POST-OPERATIVE-DAY OF SLEEVE GASTRECTOMY, PATIENT HAD TO BE BROUGHT BACK IN DUE TO POST OPERATIVE LEAK. ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
01-04-2017 | Injury | davol inc sub c.r bard | 15-01-2019 | CAPSURE FIXATION SYSTEM | Defective Device | A0203 | Pain | Event Description: THE FOLLOWING WAS REPORTED BY THE PATIENT'S ATTORNEY: (B)(6) 2017: PATIENT HAD A LAPAROSCOPIC RECTOPEXY VENTRAL MESH INSERTION WITH A NON BARD/DAVOL ETHICON PROLEEN MESH AND A CAPSURE FIXATION DEVICE. SPECIFIC DEVICE ISSUE AND/OR PATIENT INJURIES WERE NOT IDENTIFIED. ATTORNEY REQUESTS CONFIRMATION THAT WE ARE THE MANUFACTURERS OF THIS FIXATION PRODUCT AND CONFIRMATION OF DATE THAT THESE PRODUCTS WERE FIRST MANUFACTURED AND THE DATE THEY FIRST WENT INTO CIRCULATION. ADDENDUM: AS REPORTED FROM THE PATIENT¿S ATTORNEY: (B)(6) 2017, THE PATIENT UNDERWENT A LAPAROSCOPIC RECTOPEXY WITH IMPLANT OF A NON-BARD/DAVOL MESH DUE TO A DIAGNOSIS OF RECTAL PROLAPSE. PER THE HOSPITAL DISCHARGE NOTE, "THE MESH WAS DEPLOYED IN THE ANTERIOR WALL OF THE RECTUM AND FIXED TO THE SACRAL PROMONTORY." THE MESH WAS FIXED INTO PLACE WITH THE BARD/DAVOL CAPSURE DEVICE. (B)(6) 2017, DURING POSTOPERATIVE VISIT, THE PATIENT COMPLAINED OF BACK PAIN, FREQUENT LEAKAGE [BOWEL] AND PAIN HIGH UP IN THE ANAL CANAL. (B)(6) 2017, THE PATIENT UNDERWENT AN MRI WHICH SHOWED SOME ARTIFACT LYING ADJACENT TO THE SACRAL PROMONTORY, "LIKELY DUE TO THE METALLIC TACKS USED DURING SURGERY." AS REPORTED, ORTHOPAEDIC INVESTIGATIONS RULED OUT ANY ORTHOPAEDIC CAUSE OF THE PATIENT'S SYMPTOMS. (B)(6) 2017, DURING FOLLOW-UP, THE SURGEON OFFERED TO PERFORM ADDITIONAL SURGERY BUT THE PATIENT WAS APPREHENSIVE DUE TO THE NATURE OF HER SYMPTOMS FOLLOWING THE INITIAL SURGERY. (B)(6) 2018, DURING A SECOND FOLLOW-UP, THE SURGEON OFFERED TO "REMOVE THE TACKS" AS THE SURGEON THOUGHT THE TACKS WERE THE "MAIN CAUSE OF THE PATIENT'S PROBLEMS." (B)(6) 2018, FOLLOWING MULTIPLE INVESTIGATIONS FROM THE ATTENDING CLINICIANS, THE SOURCE OF THE PAIN WAS THOUGHT TO BE RELATED TO THE FIXATION DEVICE [CAPSURE] INSERTED DURING THE OPERATION DATED ON (B)(6) 2017. (B)(6) 2019, THE PATIENT SOUGHT TREATMENT WITH ANOTHER PHYSICIAN WHO REPORTED: "COULD PALPATE RECTOPEXY MESH [NON-BARD/DAVOL] & STITCHES OVER POSTERIOR VAGINAL WALL THROUGH VAGINAL EPITHELIUM." (B)(6) 2019, THE PATIENT UNDERWENT AN EXPLORATORY PROCEDURE AND THE OPERATIVE NOTES STATED: "NO EVIDENCE OF EROSIONS. AT LAPAROSCOPY MESH IN CORRECT POSITION, SLIGHTLY HIGH AT THE LEVEL OF THE VAGINAL SEPTUM, PERITONEALIZED." AS REPORTED, IN (B)(6) 2019, THE PATIENT UNDERWENT ADDITIONAL SURGERY AND WAS ALSO OFFERED FURTHER SURGERY. THE MOST RECENT APPOINTMENTS HAVE BEEN POSTPONED DUE TO THE ONGOING PANDEMIC [COVID-19]. ATTORNEY ALLEGES THE FIXATION PRODUCTS [CAPSURE] WERE DEFECTIVE. IT IS ALLEGED THAT HE PATIENT EXPERIENCED CHRONIC PAIN, DYSPAREUNIA, ALTERED BOWEL MOVEMENTS, ALTERED SENSATION AND PAIN IN THE LOWER SPINE IN THE REGION THE TACKS WERE PLACED. IT IS ALSO ALLEGED THAT THE PATIENT'S PHYSICAL AND PSYCHIATRIC DIFFICULTIES ARE CONTINUING WHICH AFFECT THE PATIENT'S QUALITY OF LIFE EVERY DAY.
01-04-2017 | Injury | davol inc sub c.r bard | 15-01-2019 | CAPSURE FIXATION SYSTEM | Defective Device | A0203 | Injury | Event Description: THE FOLLOWING WAS REPORTED BY THE PATIENT'S ATTORNEY: (B)(6) 2017: PATIENT HAD A LAPAROSCOPIC RECTOPEXY VENTRAL MESH INSERTION WITH A NON BARD/DAVOL ETHICON PROLEEN MESH AND A CAPSURE FIXATION DEVICE. SPECIFIC DEVICE ISSUE AND/OR PATIENT INJURIES WERE NOT IDENTIFIED. ATTORNEY REQUESTS CONFIRMATION THAT WE ARE THE MANUFACTURERS OF THIS FIXATION PRODUCT AND CONFIRMATION OF DATE THAT THESE PRODUCTS WERE FIRST MANUFACTURED AND THE DATE THEY FIRST WENT INTO CIRCULATION. ADDENDUM: AS REPORTED FROM THE PATIENT¿S ATTORNEY: (B)(6) 2017, THE PATIENT UNDERWENT A LAPAROSCOPIC RECTOPEXY WITH IMPLANT OF A NON-BARD/DAVOL MESH DUE TO A DIAGNOSIS OF RECTAL PROLAPSE. PER THE HOSPITAL DISCHARGE NOTE, "THE MESH WAS DEPLOYED IN THE ANTERIOR WALL OF THE RECTUM AND FIXED TO THE SACRAL PROMONTORY." THE MESH WAS FIXED INTO PLACE WITH THE BARD/DAVOL CAPSURE DEVICE. (B)(6) 2017, DURING POSTOPERATIVE VISIT, THE PATIENT COMPLAINED OF BACK PAIN, FREQUENT LEAKAGE [BOWEL] AND PAIN HIGH UP IN THE ANAL CANAL. (B)(6) 2017, THE PATIENT UNDERWENT AN MRI WHICH SHOWED SOME ARTIFACT LYING ADJACENT TO THE SACRAL PROMONTORY, "LIKELY DUE TO THE METALLIC TACKS USED DURING SURGERY." AS REPORTED, ORTHOPAEDIC INVESTIGATIONS RULED OUT ANY ORTHOPAEDIC CAUSE OF THE PATIENT'S SYMPTOMS. (B)(6) 2017, DURING FOLLOW-UP, THE SURGEON OFFERED TO PERFORM ADDITIONAL SURGERY BUT THE PATIENT WAS APPREHENSIVE DUE TO THE NATURE OF HER SYMPTOMS FOLLOWING THE INITIAL SURGERY. (B)(6) 2018, DURING A SECOND FOLLOW-UP, THE SURGEON OFFERED TO "REMOVE THE TACKS" AS THE SURGEON THOUGHT THE TACKS WERE THE "MAIN CAUSE OF THE PATIENT'S PROBLEMS." (B)(6) 2018, FOLLOWING MULTIPLE INVESTIGATIONS FROM THE ATTENDING CLINICIANS, THE SOURCE OF THE PAIN WAS THOUGHT TO BE RELATED TO THE FIXATION DEVICE [CAPSURE] INSERTED DURING THE OPERATION DATED ON (B)(6) 2017. (B)(6) 2019, THE PATIENT SOUGHT TREATMENT WITH ANOTHER PHYSICIAN WHO REPORTED: "COULD PALPATE RECTOPEXY MESH [NON-BARD/DAVOL] & STITCHES OVER POSTERIOR VAGINAL WALL THROUGH VAGINAL EPITHELIUM." (B)(6) 2019, THE PATIENT UNDERWENT AN EXPLORATORY PROCEDURE AND THE OPERATIVE NOTES STATED: "NO EVIDENCE OF EROSIONS. AT LAPAROSCOPY MESH IN CORRECT POSITION, SLIGHTLY HIGH AT THE LEVEL OF THE VAGINAL SEPTUM, PERITONEALIZED." AS REPORTED, IN (B)(6) 2019, THE PATIENT UNDERWENT ADDITIONAL SURGERY AND WAS ALSO OFFERED FURTHER SURGERY. THE MOST RECENT APPOINTMENTS HAVE BEEN POSTPONED DUE TO THE ONGOING PANDEMIC [COVID-19]. ATTORNEY ALLEGES THE FIXATION PRODUCTS [CAPSURE] WERE DEFECTIVE. IT IS ALLEGED THAT HE PATIENT EXPERIENCED CHRONIC PAIN, DYSPAREUNIA, ALTERED BOWEL MOVEMENTS, ALTERED SENSATION AND PAIN IN THE LOWER SPINE IN THE REGION THE TACKS WERE PLACED. IT IS ALSO ALLEGED THAT THE PATIENT'S PHYSICAL AND PSYCHIATRIC DIFFICULTIES ARE CONTINUING WHICH AFFECT THE PATIENT'S QUALITY OF LIFE EVERY DAY.
01-04-2017 | Injury | davol inc sub c.r bard | 15-01-2019 | CAPSURE FIXATION SYSTEM | Defective Device | A0203 | Disability | Event Description: THE FOLLOWING WAS REPORTED BY THE PATIENT'S ATTORNEY: (B)(6) 2017: PATIENT HAD A LAPAROSCOPIC RECTOPEXY VENTRAL MESH INSERTION WITH A NON BARD/DAVOL ETHICON PROLEEN MESH AND A CAPSURE FIXATION DEVICE. SPECIFIC DEVICE ISSUE AND/OR PATIENT INJURIES WERE NOT IDENTIFIED. ATTORNEY REQUESTS CONFIRMATION THAT WE ARE THE MANUFACTURERS OF THIS FIXATION PRODUCT AND CONFIRMATION OF DATE THAT THESE PRODUCTS WERE FIRST MANUFACTURED AND THE DATE THEY FIRST WENT INTO CIRCULATION. ADDENDUM: AS REPORTED FROM THE PATIENT¿S ATTORNEY: (B)(6) 2017, THE PATIENT UNDERWENT A LAPAROSCOPIC RECTOPEXY WITH IMPLANT OF A NON-BARD/DAVOL MESH DUE TO A DIAGNOSIS OF RECTAL PROLAPSE. PER THE HOSPITAL DISCHARGE NOTE, "THE MESH WAS DEPLOYED IN THE ANTERIOR WALL OF THE RECTUM AND FIXED TO THE SACRAL PROMONTORY." THE MESH WAS FIXED INTO PLACE WITH THE BARD/DAVOL CAPSURE DEVICE. (B)(6) 2017, DURING POSTOPERATIVE VISIT, THE PATIENT COMPLAINED OF BACK PAIN, FREQUENT LEAKAGE [BOWEL] AND PAIN HIGH UP IN THE ANAL CANAL. (B)(6) 2017, THE PATIENT UNDERWENT AN MRI WHICH SHOWED SOME ARTIFACT LYING ADJACENT TO THE SACRAL PROMONTORY, "LIKELY DUE TO THE METALLIC TACKS USED DURING SURGERY." AS REPORTED, ORTHOPAEDIC INVESTIGATIONS RULED OUT ANY ORTHOPAEDIC CAUSE OF THE PATIENT'S SYMPTOMS. (B)(6) 2017, DURING FOLLOW-UP, THE SURGEON OFFERED TO PERFORM ADDITIONAL SURGERY BUT THE PATIENT WAS APPREHENSIVE DUE TO THE NATURE OF HER SYMPTOMS FOLLOWING THE INITIAL SURGERY. (B)(6) 2018, DURING A SECOND FOLLOW-UP, THE SURGEON OFFERED TO "REMOVE THE TACKS" AS THE SURGEON THOUGHT THE TACKS WERE THE "MAIN CAUSE OF THE PATIENT'S PROBLEMS." (B)(6) 2018, FOLLOWING MULTIPLE INVESTIGATIONS FROM THE ATTENDING CLINICIANS, THE SOURCE OF THE PAIN WAS THOUGHT TO BE RELATED TO THE FIXATION DEVICE [CAPSURE] INSERTED DURING THE OPERATION DATED ON (B)(6) 2017. (B)(6) 2019, THE PATIENT SOUGHT TREATMENT WITH ANOTHER PHYSICIAN WHO REPORTED: "COULD PALPATE RECTOPEXY MESH [NON-BARD/DAVOL] & STITCHES OVER POSTERIOR VAGINAL WALL THROUGH VAGINAL EPITHELIUM." (B)(6) 2019, THE PATIENT UNDERWENT AN EXPLORATORY PROCEDURE AND THE OPERATIVE NOTES STATED: "NO EVIDENCE OF EROSIONS. AT LAPAROSCOPY MESH IN CORRECT POSITION, SLIGHTLY HIGH AT THE LEVEL OF THE VAGINAL SEPTUM, PERITONEALIZED." AS REPORTED, IN (B)(6) 2019, THE PATIENT UNDERWENT ADDITIONAL SURGERY AND WAS ALSO OFFERED FURTHER SURGERY. THE MOST RECENT APPOINTMENTS HAVE BEEN POSTPONED DUE TO THE ONGOING PANDEMIC [COVID-19]. ATTORNEY ALLEGES THE FIXATION PRODUCTS [CAPSURE] WERE DEFECTIVE. IT IS ALLEGED THAT HE PATIENT EXPERIENCED CHRONIC PAIN, DYSPAREUNIA, ALTERED BOWEL MOVEMENTS, ALTERED SENSATION AND PAIN IN THE LOWER SPINE IN THE REGION THE TACKS WERE PLACED. IT IS ALSO ALLEGED THAT THE PATIENT'S PHYSICAL AND PSYCHIATRIC DIFFICULTIES ARE CONTINUING WHICH AFFECT THE PATIENT'S QUALITY OF LIFE EVERY DAY.
01-04-2017 | Injury | davol inc sub c.r bard | 15-01-2019 | CAPSURE FIXATION SYSTEM | Insufficient Information | E2401 | Pain | Event Description: THE FOLLOWING WAS REPORTED BY THE PATIENT'S ATTORNEY: (B)(6) 2017: PATIENT HAD A LAPAROSCOPIC RECTOPEXY VENTRAL MESH INSERTION WITH A NON BARD/DAVOL ETHICON PROLEEN MESH AND A CAPSURE FIXATION DEVICE. SPECIFIC DEVICE ISSUE AND/OR PATIENT INJURIES WERE NOT IDENTIFIED. ATTORNEY REQUESTS CONFIRMATION THAT WE ARE THE MANUFACTURERS OF THIS FIXATION PRODUCT AND CONFIRMATION OF DATE THAT THESE PRODUCTS WERE FIRST MANUFACTURED AND THE DATE THEY FIRST WENT INTO CIRCULATION. ADDENDUM: AS REPORTED FROM THE PATIENT¿S ATTORNEY: (B)(6) 2017, THE PATIENT UNDERWENT A LAPAROSCOPIC RECTOPEXY WITH IMPLANT OF A NON-BARD/DAVOL MESH DUE TO A DIAGNOSIS OF RECTAL PROLAPSE. PER THE HOSPITAL DISCHARGE NOTE, "THE MESH WAS DEPLOYED IN THE ANTERIOR WALL OF THE RECTUM AND FIXED TO THE SACRAL PROMONTORY." THE MESH WAS FIXED INTO PLACE WITH THE BARD/DAVOL CAPSURE DEVICE. (B)(6) 2017, DURING POSTOPERATIVE VISIT, THE PATIENT COMPLAINED OF BACK PAIN, FREQUENT LEAKAGE [BOWEL] AND PAIN HIGH UP IN THE ANAL CANAL. (B)(6) 2017, THE PATIENT UNDERWENT AN MRI WHICH SHOWED SOME ARTIFACT LYING ADJACENT TO THE SACRAL PROMONTORY, "LIKELY DUE TO THE METALLIC TACKS USED DURING SURGERY." AS REPORTED, ORTHOPAEDIC INVESTIGATIONS RULED OUT ANY ORTHOPAEDIC CAUSE OF THE PATIENT'S SYMPTOMS. (B)(6) 2017, DURING FOLLOW-UP, THE SURGEON OFFERED TO PERFORM ADDITIONAL SURGERY BUT THE PATIENT WAS APPREHENSIVE DUE TO THE NATURE OF HER SYMPTOMS FOLLOWING THE INITIAL SURGERY. (B)(6) 2018, DURING A SECOND FOLLOW-UP, THE SURGEON OFFERED TO "REMOVE THE TACKS" AS THE SURGEON THOUGHT THE TACKS WERE THE "MAIN CAUSE OF THE PATIENT'S PROBLEMS." (B)(6) 2018, FOLLOWING MULTIPLE INVESTIGATIONS FROM THE ATTENDING CLINICIANS, THE SOURCE OF THE PAIN WAS THOUGHT TO BE RELATED TO THE FIXATION DEVICE [CAPSURE] INSERTED DURING THE OPERATION DATED ON (B)(6) 2017. (B)(6) 2019, THE PATIENT SOUGHT TREATMENT WITH ANOTHER PHYSICIAN WHO REPORTED: "COULD PALPATE RECTOPEXY MESH [NON-BARD/DAVOL] & STITCHES OVER POSTERIOR VAGINAL WALL THROUGH VAGINAL EPITHELIUM." (B)(6) 2019, THE PATIENT UNDERWENT AN EXPLORATORY PROCEDURE AND THE OPERATIVE NOTES STATED: "NO EVIDENCE OF EROSIONS. AT LAPAROSCOPY MESH IN CORRECT POSITION, SLIGHTLY HIGH AT THE LEVEL OF THE VAGINAL SEPTUM, PERITONEALIZED." AS REPORTED, IN (B)(6) 2019, THE PATIENT UNDERWENT ADDITIONAL SURGERY AND WAS ALSO OFFERED FURTHER SURGERY. THE MOST RECENT APPOINTMENTS HAVE BEEN POSTPONED DUE TO THE ONGOING PANDEMIC [COVID-19]. ATTORNEY ALLEGES THE FIXATION PRODUCTS [CAPSURE] WERE DEFECTIVE. IT IS ALLEGED THAT HE PATIENT EXPERIENCED CHRONIC PAIN, DYSPAREUNIA, ALTERED BOWEL MOVEMENTS, ALTERED SENSATION AND PAIN IN THE LOWER SPINE IN THE REGION THE TACKS WERE PLACED. IT IS ALSO ALLEGED THAT THE PATIENT'S PHYSICAL AND PSYCHIATRIC DIFFICULTIES ARE CONTINUING WHICH AFFECT THE PATIENT'S QUALITY OF LIFE EVERY DAY.
01-04-2017 | Injury | davol inc sub c.r bard | 15-01-2019 | CAPSURE FIXATION SYSTEM | Insufficient Information | E2401 | Injury | Event Description: THE FOLLOWING WAS REPORTED BY THE PATIENT'S ATTORNEY: (B)(6) 2017: PATIENT HAD A LAPAROSCOPIC RECTOPEXY VENTRAL MESH INSERTION WITH A NON BARD/DAVOL ETHICON PROLEEN MESH AND A CAPSURE FIXATION DEVICE. SPECIFIC DEVICE ISSUE AND/OR PATIENT INJURIES WERE NOT IDENTIFIED. ATTORNEY REQUESTS CONFIRMATION THAT WE ARE THE MANUFACTURERS OF THIS FIXATION PRODUCT AND CONFIRMATION OF DATE THAT THESE PRODUCTS WERE FIRST MANUFACTURED AND THE DATE THEY FIRST WENT INTO CIRCULATION. ADDENDUM: AS REPORTED FROM THE PATIENT¿S ATTORNEY: (B)(6) 2017, THE PATIENT UNDERWENT A LAPAROSCOPIC RECTOPEXY WITH IMPLANT OF A NON-BARD/DAVOL MESH DUE TO A DIAGNOSIS OF RECTAL PROLAPSE. PER THE HOSPITAL DISCHARGE NOTE, "THE MESH WAS DEPLOYED IN THE ANTERIOR WALL OF THE RECTUM AND FIXED TO THE SACRAL PROMONTORY." THE MESH WAS FIXED INTO PLACE WITH THE BARD/DAVOL CAPSURE DEVICE. (B)(6) 2017, DURING POSTOPERATIVE VISIT, THE PATIENT COMPLAINED OF BACK PAIN, FREQUENT LEAKAGE [BOWEL] AND PAIN HIGH UP IN THE ANAL CANAL. (B)(6) 2017, THE PATIENT UNDERWENT AN MRI WHICH SHOWED SOME ARTIFACT LYING ADJACENT TO THE SACRAL PROMONTORY, "LIKELY DUE TO THE METALLIC TACKS USED DURING SURGERY." AS REPORTED, ORTHOPAEDIC INVESTIGATIONS RULED OUT ANY ORTHOPAEDIC CAUSE OF THE PATIENT'S SYMPTOMS. (B)(6) 2017, DURING FOLLOW-UP, THE SURGEON OFFERED TO PERFORM ADDITIONAL SURGERY BUT THE PATIENT WAS APPREHENSIVE DUE TO THE NATURE OF HER SYMPTOMS FOLLOWING THE INITIAL SURGERY. (B)(6) 2018, DURING A SECOND FOLLOW-UP, THE SURGEON OFFERED TO "REMOVE THE TACKS" AS THE SURGEON THOUGHT THE TACKS WERE THE "MAIN CAUSE OF THE PATIENT'S PROBLEMS." (B)(6) 2018, FOLLOWING MULTIPLE INVESTIGATIONS FROM THE ATTENDING CLINICIANS, THE SOURCE OF THE PAIN WAS THOUGHT TO BE RELATED TO THE FIXATION DEVICE [CAPSURE] INSERTED DURING THE OPERATION DATED ON (B)(6) 2017. (B)(6) 2019, THE PATIENT SOUGHT TREATMENT WITH ANOTHER PHYSICIAN WHO REPORTED: "COULD PALPATE RECTOPEXY MESH [NON-BARD/DAVOL] & STITCHES OVER POSTERIOR VAGINAL WALL THROUGH VAGINAL EPITHELIUM." (B)(6) 2019, THE PATIENT UNDERWENT AN EXPLORATORY PROCEDURE AND THE OPERATIVE NOTES STATED: "NO EVIDENCE OF EROSIONS. AT LAPAROSCOPY MESH IN CORRECT POSITION, SLIGHTLY HIGH AT THE LEVEL OF THE VAGINAL SEPTUM, PERITONEALIZED." AS REPORTED, IN (B)(6) 2019, THE PATIENT UNDERWENT ADDITIONAL SURGERY AND WAS ALSO OFFERED FURTHER SURGERY. THE MOST RECENT APPOINTMENTS HAVE BEEN POSTPONED DUE TO THE ONGOING PANDEMIC [COVID-19]. ATTORNEY ALLEGES THE FIXATION PRODUCTS [CAPSURE] WERE DEFECTIVE. IT IS ALLEGED THAT HE PATIENT EXPERIENCED CHRONIC PAIN, DYSPAREUNIA, ALTERED BOWEL MOVEMENTS, ALTERED SENSATION AND PAIN IN THE LOWER SPINE IN THE REGION THE TACKS WERE PLACED. IT IS ALSO ALLEGED THAT THE PATIENT'S PHYSICAL AND PSYCHIATRIC DIFFICULTIES ARE CONTINUING WHICH AFFECT THE PATIENT'S QUALITY OF LIFE EVERY DAY.
01-04-2017 | Injury | davol inc sub c.r bard | 15-01-2019 | CAPSURE FIXATION SYSTEM | Insufficient Information | E2401 | Disability | Event Description: THE FOLLOWING WAS REPORTED BY THE PATIENT'S ATTORNEY: (B)(6) 2017: PATIENT HAD A LAPAROSCOPIC RECTOPEXY VENTRAL MESH INSERTION WITH A NON BARD/DAVOL ETHICON PROLEEN MESH AND A CAPSURE FIXATION DEVICE. SPECIFIC DEVICE ISSUE AND/OR PATIENT INJURIES WERE NOT IDENTIFIED. ATTORNEY REQUESTS CONFIRMATION THAT WE ARE THE MANUFACTURERS OF THIS FIXATION PRODUCT AND CONFIRMATION OF DATE THAT THESE PRODUCTS WERE FIRST MANUFACTURED AND THE DATE THEY FIRST WENT INTO CIRCULATION. ADDENDUM: AS REPORTED FROM THE PATIENT¿S ATTORNEY: (B)(6) 2017, THE PATIENT UNDERWENT A LAPAROSCOPIC RECTOPEXY WITH IMPLANT OF A NON-BARD/DAVOL MESH DUE TO A DIAGNOSIS OF RECTAL PROLAPSE. PER THE HOSPITAL DISCHARGE NOTE, "THE MESH WAS DEPLOYED IN THE ANTERIOR WALL OF THE RECTUM AND FIXED TO THE SACRAL PROMONTORY." THE MESH WAS FIXED INTO PLACE WITH THE BARD/DAVOL CAPSURE DEVICE. (B)(6) 2017, DURING POSTOPERATIVE VISIT, THE PATIENT COMPLAINED OF BACK PAIN, FREQUENT LEAKAGE [BOWEL] AND PAIN HIGH UP IN THE ANAL CANAL. (B)(6) 2017, THE PATIENT UNDERWENT AN MRI WHICH SHOWED SOME ARTIFACT LYING ADJACENT TO THE SACRAL PROMONTORY, "LIKELY DUE TO THE METALLIC TACKS USED DURING SURGERY." AS REPORTED, ORTHOPAEDIC INVESTIGATIONS RULED OUT ANY ORTHOPAEDIC CAUSE OF THE PATIENT'S SYMPTOMS. (B)(6) 2017, DURING FOLLOW-UP, THE SURGEON OFFERED TO PERFORM ADDITIONAL SURGERY BUT THE PATIENT WAS APPREHENSIVE DUE TO THE NATURE OF HER SYMPTOMS FOLLOWING THE INITIAL SURGERY. (B)(6) 2018, DURING A SECOND FOLLOW-UP, THE SURGEON OFFERED TO "REMOVE THE TACKS" AS THE SURGEON THOUGHT THE TACKS WERE THE "MAIN CAUSE OF THE PATIENT'S PROBLEMS." (B)(6) 2018, FOLLOWING MULTIPLE INVESTIGATIONS FROM THE ATTENDING CLINICIANS, THE SOURCE OF THE PAIN WAS THOUGHT TO BE RELATED TO THE FIXATION DEVICE [CAPSURE] INSERTED DURING THE OPERATION DATED ON (B)(6) 2017. (B)(6) 2019, THE PATIENT SOUGHT TREATMENT WITH ANOTHER PHYSICIAN WHO REPORTED: "COULD PALPATE RECTOPEXY MESH [NON-BARD/DAVOL] & STITCHES OVER POSTERIOR VAGINAL WALL THROUGH VAGINAL EPITHELIUM." (B)(6) 2019, THE PATIENT UNDERWENT AN EXPLORATORY PROCEDURE AND THE OPERATIVE NOTES STATED: "NO EVIDENCE OF EROSIONS. AT LAPAROSCOPY MESH IN CORRECT POSITION, SLIGHTLY HIGH AT THE LEVEL OF THE VAGINAL SEPTUM, PERITONEALIZED." AS REPORTED, IN (B)(6) 2019, THE PATIENT UNDERWENT ADDITIONAL SURGERY AND WAS ALSO OFFERED FURTHER SURGERY. THE MOST RECENT APPOINTMENTS HAVE BEEN POSTPONED DUE TO THE ONGOING PANDEMIC [COVID-19]. ATTORNEY ALLEGES THE FIXATION PRODUCTS [CAPSURE] WERE DEFECTIVE. IT IS ALLEGED THAT HE PATIENT EXPERIENCED CHRONIC PAIN, DYSPAREUNIA, ALTERED BOWEL MOVEMENTS, ALTERED SENSATION AND PAIN IN THE LOWER SPINE IN THE REGION THE TACKS WERE PLACED. IT IS ALSO ALLEGED THAT THE PATIENT'S PHYSICAL AND PSYCHIATRIC DIFFICULTIES ARE CONTINUING WHICH AFFECT THE PATIENT'S QUALITY OF LIFE EVERY DAY.
01-04-2017 | Injury | davol inc sub c.r bard | 15-01-2019 | CAPSURE FIXATION SYSTEM | Patient Device Interaction Problem | A01 | Pain | Event Description: THE FOLLOWING WAS REPORTED BY THE PATIENT'S ATTORNEY: (B)(6) 2017: PATIENT HAD A LAPAROSCOPIC RECTOPEXY VENTRAL MESH INSERTION WITH A NON BARD/DAVOL ETHICON PROLEEN MESH AND A CAPSURE FIXATION DEVICE. SPECIFIC DEVICE ISSUE AND/OR PATIENT INJURIES WERE NOT IDENTIFIED. ATTORNEY REQUESTS CONFIRMATION THAT WE ARE THE MANUFACTURERS OF THIS FIXATION PRODUCT AND CONFIRMATION OF DATE THAT THESE PRODUCTS WERE FIRST MANUFACTURED AND THE DATE THEY FIRST WENT INTO CIRCULATION. ADDENDUM: AS REPORTED FROM THE PATIENT¿S ATTORNEY: (B)(6) 2017, THE PATIENT UNDERWENT A LAPAROSCOPIC RECTOPEXY WITH IMPLANT OF A NON-BARD/DAVOL MESH DUE TO A DIAGNOSIS OF RECTAL PROLAPSE. PER THE HOSPITAL DISCHARGE NOTE, "THE MESH WAS DEPLOYED IN THE ANTERIOR WALL OF THE RECTUM AND FIXED TO THE SACRAL PROMONTORY." THE MESH WAS FIXED INTO PLACE WITH THE BARD/DAVOL CAPSURE DEVICE. (B)(6) 2017, DURING POSTOPERATIVE VISIT, THE PATIENT COMPLAINED OF BACK PAIN, FREQUENT LEAKAGE [BOWEL] AND PAIN HIGH UP IN THE ANAL CANAL. (B)(6) 2017, THE PATIENT UNDERWENT AN MRI WHICH SHOWED SOME ARTIFACT LYING ADJACENT TO THE SACRAL PROMONTORY, "LIKELY DUE TO THE METALLIC TACKS USED DURING SURGERY." AS REPORTED, ORTHOPAEDIC INVESTIGATIONS RULED OUT ANY ORTHOPAEDIC CAUSE OF THE PATIENT'S SYMPTOMS. (B)(6) 2017, DURING FOLLOW-UP, THE SURGEON OFFERED TO PERFORM ADDITIONAL SURGERY BUT THE PATIENT WAS APPREHENSIVE DUE TO THE NATURE OF HER SYMPTOMS FOLLOWING THE INITIAL SURGERY. (B)(6) 2018, DURING A SECOND FOLLOW-UP, THE SURGEON OFFERED TO "REMOVE THE TACKS" AS THE SURGEON THOUGHT THE TACKS WERE THE "MAIN CAUSE OF THE PATIENT'S PROBLEMS." (B)(6) 2018, FOLLOWING MULTIPLE INVESTIGATIONS FROM THE ATTENDING CLINICIANS, THE SOURCE OF THE PAIN WAS THOUGHT TO BE RELATED TO THE FIXATION DEVICE [CAPSURE] INSERTED DURING THE OPERATION DATED ON (B)(6) 2017. (B)(6) 2019, THE PATIENT SOUGHT TREATMENT WITH ANOTHER PHYSICIAN WHO REPORTED: "COULD PALPATE RECTOPEXY MESH [NON-BARD/DAVOL] & STITCHES OVER POSTERIOR VAGINAL WALL THROUGH VAGINAL EPITHELIUM." (B)(6) 2019, THE PATIENT UNDERWENT AN EXPLORATORY PROCEDURE AND THE OPERATIVE NOTES STATED: "NO EVIDENCE OF EROSIONS. AT LAPAROSCOPY MESH IN CORRECT POSITION, SLIGHTLY HIGH AT THE LEVEL OF THE VAGINAL SEPTUM, PERITONEALIZED." AS REPORTED, IN (B)(6) 2019, THE PATIENT UNDERWENT ADDITIONAL SURGERY AND WAS ALSO OFFERED FURTHER SURGERY. THE MOST RECENT APPOINTMENTS HAVE BEEN POSTPONED DUE TO THE ONGOING PANDEMIC [COVID-19]. ATTORNEY ALLEGES THE FIXATION PRODUCTS [CAPSURE] WERE DEFECTIVE. IT IS ALLEGED THAT HE PATIENT EXPERIENCED CHRONIC PAIN, DYSPAREUNIA, ALTERED BOWEL MOVEMENTS, ALTERED SENSATION AND PAIN IN THE LOWER SPINE IN THE REGION THE TACKS WERE PLACED. IT IS ALSO ALLEGED THAT THE PATIENT'S PHYSICAL AND PSYCHIATRIC DIFFICULTIES ARE CONTINUING WHICH AFFECT THE PATIENT'S QUALITY OF LIFE EVERY DAY.
01-04-2017 | Injury | davol inc sub c.r bard | 15-01-2019 | CAPSURE FIXATION SYSTEM | Patient Device Interaction Problem | A01 | Injury | Event Description: THE FOLLOWING WAS REPORTED BY THE PATIENT'S ATTORNEY: (B)(6) 2017: PATIENT HAD A LAPAROSCOPIC RECTOPEXY VENTRAL MESH INSERTION WITH A NON BARD/DAVOL ETHICON PROLEEN MESH AND A CAPSURE FIXATION DEVICE. SPECIFIC DEVICE ISSUE AND/OR PATIENT INJURIES WERE NOT IDENTIFIED. ATTORNEY REQUESTS CONFIRMATION THAT WE ARE THE MANUFACTURERS OF THIS FIXATION PRODUCT AND CONFIRMATION OF DATE THAT THESE PRODUCTS WERE FIRST MANUFACTURED AND THE DATE THEY FIRST WENT INTO CIRCULATION. ADDENDUM: AS REPORTED FROM THE PATIENT¿S ATTORNEY: (B)(6) 2017, THE PATIENT UNDERWENT A LAPAROSCOPIC RECTOPEXY WITH IMPLANT OF A NON-BARD/DAVOL MESH DUE TO A DIAGNOSIS OF RECTAL PROLAPSE. PER THE HOSPITAL DISCHARGE NOTE, "THE MESH WAS DEPLOYED IN THE ANTERIOR WALL OF THE RECTUM AND FIXED TO THE SACRAL PROMONTORY." THE MESH WAS FIXED INTO PLACE WITH THE BARD/DAVOL CAPSURE DEVICE. (B)(6) 2017, DURING POSTOPERATIVE VISIT, THE PATIENT COMPLAINED OF BACK PAIN, FREQUENT LEAKAGE [BOWEL] AND PAIN HIGH UP IN THE ANAL CANAL. (B)(6) 2017, THE PATIENT UNDERWENT AN MRI WHICH SHOWED SOME ARTIFACT LYING ADJACENT TO THE SACRAL PROMONTORY, "LIKELY DUE TO THE METALLIC TACKS USED DURING SURGERY." AS REPORTED, ORTHOPAEDIC INVESTIGATIONS RULED OUT ANY ORTHOPAEDIC CAUSE OF THE PATIENT'S SYMPTOMS. (B)(6) 2017, DURING FOLLOW-UP, THE SURGEON OFFERED TO PERFORM ADDITIONAL SURGERY BUT THE PATIENT WAS APPREHENSIVE DUE TO THE NATURE OF HER SYMPTOMS FOLLOWING THE INITIAL SURGERY. (B)(6) 2018, DURING A SECOND FOLLOW-UP, THE SURGEON OFFERED TO "REMOVE THE TACKS" AS THE SURGEON THOUGHT THE TACKS WERE THE "MAIN CAUSE OF THE PATIENT'S PROBLEMS." (B)(6) 2018, FOLLOWING MULTIPLE INVESTIGATIONS FROM THE ATTENDING CLINICIANS, THE SOURCE OF THE PAIN WAS THOUGHT TO BE RELATED TO THE FIXATION DEVICE [CAPSURE] INSERTED DURING THE OPERATION DATED ON (B)(6) 2017. (B)(6) 2019, THE PATIENT SOUGHT TREATMENT WITH ANOTHER PHYSICIAN WHO REPORTED: "COULD PALPATE RECTOPEXY MESH [NON-BARD/DAVOL] & STITCHES OVER POSTERIOR VAGINAL WALL THROUGH VAGINAL EPITHELIUM." (B)(6) 2019, THE PATIENT UNDERWENT AN EXPLORATORY PROCEDURE AND THE OPERATIVE NOTES STATED: "NO EVIDENCE OF EROSIONS. AT LAPAROSCOPY MESH IN CORRECT POSITION, SLIGHTLY HIGH AT THE LEVEL OF THE VAGINAL SEPTUM, PERITONEALIZED." AS REPORTED, IN (B)(6) 2019, THE PATIENT UNDERWENT ADDITIONAL SURGERY AND WAS ALSO OFFERED FURTHER SURGERY. THE MOST RECENT APPOINTMENTS HAVE BEEN POSTPONED DUE TO THE ONGOING PANDEMIC [COVID-19]. ATTORNEY ALLEGES THE FIXATION PRODUCTS [CAPSURE] WERE DEFECTIVE. IT IS ALLEGED THAT HE PATIENT EXPERIENCED CHRONIC PAIN, DYSPAREUNIA, ALTERED BOWEL MOVEMENTS, ALTERED SENSATION AND PAIN IN THE LOWER SPINE IN THE REGION THE TACKS WERE PLACED. IT IS ALSO ALLEGED THAT THE PATIENT'S PHYSICAL AND PSYCHIATRIC DIFFICULTIES ARE CONTINUING WHICH AFFECT THE PATIENT'S QUALITY OF LIFE EVERY DAY.
01-04-2017 | Injury | davol inc sub c.r bard | 15-01-2019 | CAPSURE FIXATION SYSTEM | Patient Device Interaction Problem | A01 | Disability | Event Description: THE FOLLOWING WAS REPORTED BY THE PATIENT'S ATTORNEY: (B)(6) 2017: PATIENT HAD A LAPAROSCOPIC RECTOPEXY VENTRAL MESH INSERTION WITH A NON BARD/DAVOL ETHICON PROLEEN MESH AND A CAPSURE FIXATION DEVICE. SPECIFIC DEVICE ISSUE AND/OR PATIENT INJURIES WERE NOT IDENTIFIED. ATTORNEY REQUESTS CONFIRMATION THAT WE ARE THE MANUFACTURERS OF THIS FIXATION PRODUCT AND CONFIRMATION OF DATE THAT THESE PRODUCTS WERE FIRST MANUFACTURED AND THE DATE THEY FIRST WENT INTO CIRCULATION. ADDENDUM: AS REPORTED FROM THE PATIENT¿S ATTORNEY: (B)(6) 2017, THE PATIENT UNDERWENT A LAPAROSCOPIC RECTOPEXY WITH IMPLANT OF A NON-BARD/DAVOL MESH DUE TO A DIAGNOSIS OF RECTAL PROLAPSE. PER THE HOSPITAL DISCHARGE NOTE, "THE MESH WAS DEPLOYED IN THE ANTERIOR WALL OF THE RECTUM AND FIXED TO THE SACRAL PROMONTORY." THE MESH WAS FIXED INTO PLACE WITH THE BARD/DAVOL CAPSURE DEVICE. (B)(6) 2017, DURING POSTOPERATIVE VISIT, THE PATIENT COMPLAINED OF BACK PAIN, FREQUENT LEAKAGE [BOWEL] AND PAIN HIGH UP IN THE ANAL CANAL. (B)(6) 2017, THE PATIENT UNDERWENT AN MRI WHICH SHOWED SOME ARTIFACT LYING ADJACENT TO THE SACRAL PROMONTORY, "LIKELY DUE TO THE METALLIC TACKS USED DURING SURGERY." AS REPORTED, ORTHOPAEDIC INVESTIGATIONS RULED OUT ANY ORTHOPAEDIC CAUSE OF THE PATIENT'S SYMPTOMS. (B)(6) 2017, DURING FOLLOW-UP, THE SURGEON OFFERED TO PERFORM ADDITIONAL SURGERY BUT THE PATIENT WAS APPREHENSIVE DUE TO THE NATURE OF HER SYMPTOMS FOLLOWING THE INITIAL SURGERY. (B)(6) 2018, DURING A SECOND FOLLOW-UP, THE SURGEON OFFERED TO "REMOVE THE TACKS" AS THE SURGEON THOUGHT THE TACKS WERE THE "MAIN CAUSE OF THE PATIENT'S PROBLEMS." (B)(6) 2018, FOLLOWING MULTIPLE INVESTIGATIONS FROM THE ATTENDING CLINICIANS, THE SOURCE OF THE PAIN WAS THOUGHT TO BE RELATED TO THE FIXATION DEVICE [CAPSURE] INSERTED DURING THE OPERATION DATED ON (B)(6) 2017. (B)(6) 2019, THE PATIENT SOUGHT TREATMENT WITH ANOTHER PHYSICIAN WHO REPORTED: "COULD PALPATE RECTOPEXY MESH [NON-BARD/DAVOL] & STITCHES OVER POSTERIOR VAGINAL WALL THROUGH VAGINAL EPITHELIUM." (B)(6) 2019, THE PATIENT UNDERWENT AN EXPLORATORY PROCEDURE AND THE OPERATIVE NOTES STATED: "NO EVIDENCE OF EROSIONS. AT LAPAROSCOPY MESH IN CORRECT POSITION, SLIGHTLY HIGH AT THE LEVEL OF THE VAGINAL SEPTUM, PERITONEALIZED." AS REPORTED, IN (B)(6) 2019, THE PATIENT UNDERWENT ADDITIONAL SURGERY AND WAS ALSO OFFERED FURTHER SURGERY. THE MOST RECENT APPOINTMENTS HAVE BEEN POSTPONED DUE TO THE ONGOING PANDEMIC [COVID-19]. ATTORNEY ALLEGES THE FIXATION PRODUCTS [CAPSURE] WERE DEFECTIVE. IT IS ALLEGED THAT HE PATIENT EXPERIENCED CHRONIC PAIN, DYSPAREUNIA, ALTERED BOWEL MOVEMENTS, ALTERED SENSATION AND PAIN IN THE LOWER SPINE IN THE REGION THE TACKS WERE PLACED. IT IS ALSO ALLEGED THAT THE PATIENT'S PHYSICAL AND PSYCHIATRIC DIFFICULTIES ARE CONTINUING WHICH AFFECT THE PATIENT'S QUALITY OF LIFE EVERY DAY.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, FOLLOWING THE FIRST FIRING, BLACK FRAGMENTS OF DIFFERENT SIZES FELL OFF FROM THE DEVICE. NO FRAGMENTS FELL INTO THE PATIENT'S CAVITY. AFTER THE FIRING, THE BLACK FRAGMENTS WERE FOUND AND COLLECTED, AND THEN THE PROCEDURE WAS CONTINUED. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT FIRE AND WOULD NOT OPEN JAW. IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC RIGHT LOWER LOBECTOMY, COULD NOT FIRE AND COULD NOT OPEN THE JAW. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT FIRE AND WOULD NOT OPEN JAW. IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC RIGHT LOWER LOBECTOMY, COULD NOT FIRE AND COULD NOT OPEN THE JAW. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | IDRIVE ULTRA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, FOLLOWING THE FIRST FIRING, BLACK FRAGMENTS OF DIFFERENT SIZES FELL OFF FROM THE DEVICE. NO FRAGMENTS FELL INTO THE PATIENT'S CAVITY. AFTER THE FIRING, THE BLACK FRAGMENTS WERE FOUND AND COLLECTED, AND THEN THE PROCEDURE WAS CONTINUED. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, FOLLOWING THE FIRST FIRING, BLACK FRAGMENTS OF DIFFERENT SIZES FELL OFF FROM THE DEVICE. NO FRAGMENTS FELL INTO THE PATIENT'S CAVITY. AFTER THE FIRING, THE BLACK FRAGMENTS WERE FOUND AND COLLECTED, AND THEN THE PROCEDURE WAS CONTINUED. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A MIDDLE UMBILICAL LAPAROTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, THE DEVICE DID NOT FIRE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOWER ANTERIOR RESECTION, THE PUSHER THUMB PIECE OF THE DEVICE COULD NOT BE MOVED, THEREFORE THE DEVICE DID NOT FIRE. A SECOND STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
05-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN THORACIC PROCEDURE THAT THE DEVICE MIS-FIRED. UNKNOWN AS TO HOW EXACTLY THE DEVICE MIS-FIRED. UNKNOWN AS TO HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
17-04-2019 | Malfunction | covidien lp llc north haven | 15-01-2019 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INCISIONAL HERNIA REPAIR, THREE TACKERS HAD A DIFFICULTY WITH LOADING THE RELOAD. THEY OPENED ANOTHER TACKER UNTIL A TACKER WORKED. THERE WAS NO PATIENT INJURY.
17-04-2019 | Malfunction | covidien lp llc north haven | 15-01-2019 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INCISIONAL HERNIA REPAIR, THREE TACKERS HAD A DIFFICULTY WITH LOADING THE RELOAD. THEY OPENED ANOTHER TACKER UNTIL A TACKER WORKED. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | ILS 25MM, CURVED | Appropriate Term/Code Not Available |  | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE JNJ (B)(6) COMPLIANT SAMPLE HANDLING CENTER RECEIVED THE SAMPLE WITH NO COMPLAINT INFORMATION. ATTEMPTS WERE MADE TO RECONCILE THE DEVICE WITH NO SUCCESS RESULTING IN THE CREATION OF A NEW COMPLAINT. THE WASHER WAS NOT CUT OFF.
26-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT FIRE. IT WAS REPORTED THAT DURING THE ENDOSCOPIC RADICAL ESOPHAGECTOMY FOR ESOPHAGEAL CANCER, WHEN DID ESOPHAGOGASTRIC ANASTOMOSIS, CHOOSE 2MM STAPLE HEIGHT, TRIED TO FIRE BUT WITHOUT AUDIBLE FEEDBACK, THE WASHER WAS NOT CUT OFF, THE CUT LINE IS COMPLETE, BUT THE STAPLES MALFORMED WITH IRREGULAR SHAPE. THE TISSUE WAS BLEEDING, USED SUTURE TO OVERSEW AND STOP BLEEDING. THE SURGERY DELAYED ABOUT 20 MINS. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
26-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | ILS 25MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT FIRE. IT WAS REPORTED THAT DURING THE ENDOSCOPIC RADICAL ESOPHAGECTOMY FOR ESOPHAGEAL CANCER, WHEN DID ESOPHAGOGASTRIC ANASTOMOSIS, CHOOSE 2MM STAPLE HEIGHT, TRIED TO FIRE BUT WITHOUT AUDIBLE FEEDBACK, THE WASHER WAS NOT CUT OFF, THE CUT LINE IS COMPLETE, BUT THE STAPLES MALFORMED WITH IRREGULAR SHAPE. THE TISSUE WAS BLEEDING, USED SUTURE TO OVERSEW AND STOP BLEEDING. THE SURGERY DELAYED ABOUT 20 MINS. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
25-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, A ROUND METAL PART OF THE CARTRIDGE FORK WAS BROKEN OFF. THE BROKEN PART WAS A CONNECTING PART WHICH THE ANVIL HALF AND THE CARTRIDGE HALF WERE JOINED TOGETHER. THE BROKEN PART WAS RETRIEVED, AND NO PIECES WERE LEFT INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
05-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | GST GOLD RELOAD, 60MM, 6 ROW | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, IT WAS FOUND THAT THE CARTRIDGE HAD BEEN DAMAGED BEFORE USE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
25-12-2018 | Malfunction | ethicon endo-surgery | 15-01-2019 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT FIRE. IT WAS REPORTED THAT DURING THE LAPAROSCOPIC RADICAL RESECTION OF LOW RECTAL CANCER, COULD NOT FIRE, EVEN THE SURGEON HAD SQUEEZED DOWN THE FIRING TRIGGER WITH BIG FORCE, BUT WITHOUT AUDIBLE FEEDBACK, CHECKED THE WASHER, IT WAS NOT CUT OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE FAILED TO FIRE. IT WAS REPORTED THAT THE RELOAD WAS TESTED BY THE SCRUB NURSE PRIOR TO USE, AND IT WOULD OPEN AND CLOSE BUT DID NOT FIRE. A DIFFERENT RELOAD WAS USED TO COMPLETE THE PROCEDURE.THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | UNKNOWN ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE FAILED TO FIRE. IT WAS REPORTED THAT THE RELOAD WAS TESTED BY THE SCRUB NURSE PRIOR TO USE, AND IT WOULD OPEN AND CLOSE BUT DID NOT FIRE. A DIFFERENT RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | EEA | Gas/Air Leak | A050402 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC SIGMOID RESECTION, THE SURGEON PERFORMED LEAK TEST AND NOTICED MICRO BUBBLES DURING FIRST PUMP. TWO MORE LEAK TESTS WERE PERFORMED AND NO BUBBLES WERE OBSERVED. A DRAIN WAS PLACED TO OBSERVE FLUID MANAGEMENT POST OPERATIVELY. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE ARTERY DURING A VIDEO-ASSISTED THORACOSCOPIC UPPER LEFT LOBECTOMY, THE STAPLER WAS NOT ABLE TO FIRE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE ARTERY DURING A VIDEO-ASSISTED THORACOSCOPIC UPPER LEFT LOBECTOMY, THE STAPLER WAS NOT ABLE TO FIRE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE INSTRUMENT DID NOT FIRE. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | UNEGIATRI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE INSTRUMENT DID NOT FIRE. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-10-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR COLON RESECTION, THE SURGEON USED A 60 MM LOAD, BUT THE STAPLER ONLY STAPLED 45MM INSTEAD OF 60MM. AN ADDITIONAL RELOAD WAS USED TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
01-10-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | UNKNOWN ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR COLON RESECTION, THE SURGEON USED A 60 MM LOAD, BUT THE STAPLER ONLY STAPLED 45MM INSTEAD OF 60MM. AN ADDITIONAL RELOAD WAS USED TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, WHILE STAPLING THE GASTRIC TISSUE, THE RING HANDLE COULD NOT BE SQUEEZED ONCE THE GREEN BUTTON WAS PRESSED IN PREPARATION TO FIRE. A NEW LOAD WAS OPENED AND FIRED WITH THE SAME HANDLE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, WHILE STAPLING THE GASTRIC TISSUE, THE RING HANDLE COULD NOT BE SQUEEZED ONCE THE GREEN BUTTON WAS PRESSED IN PREPARATION TO FIRE. A NEW LOAD WAS OPENED AND FIRED WITH THE SAME HANDLE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, WHILE STAPLING THE GASTRIC TISSUE, THE RING HANDLE COULD NOT BE SQUEEZED ONCE THE GREEN BUTTON WAS PRESSED IN PREPARATION TO FIRE. A NEW LOAD WAS OPENED AND FIRED WITH THE SAME HANDLE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, WHILE STAPLING THE GASTRIC TISSUE, THE RING HANDLE COULD NOT BE SQUEEZED ONCE THE GREEN BUTTON WAS PRESSED IN PREPARATION TO FIRE. A NEW LOAD WAS OPENED AND FIRED WITH THE SAME HANDLE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
17-12-2018 | Injury | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING IN A LAPAROSCOPIC SIGMOIDECTOMY WITH DOUBLE-STAPLING TECHNIQUE (DST) ANASTOMOSIS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED IN THE COLON FOR COLECTOMY, THE FIRING STOPPED IMMEDIATELY AND THE SCREEN AT THAT MOMENT DISPLAYED THE RESISTANCE LIMIT MARK SCREEN. IN THE COLON, THE STAPLE(S) WHICH WAS NOT FORMED REMAINED IN THE U SHAPE AND THE RESECTION WAS NOT COMPLETED. THE ANAL SIDE WAS TREATED AND THE FIRING WAS PERFORMED WITH A NEW RELOAD AND THE RESECTION WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. WHEN THE DEVICE WAS CHECKED, THEY FOUND THE PLASTIC PART OF CARTRIDGE WAS MISALIGNED.
17-12-2018 | Injury | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING IN A LAPAROSCOPIC SIGMOIDECTOMY WITH DOUBLE-STAPLING TECHNIQUE (DST) ANASTOMOSIS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED IN THE COLON FOR COLECTOMY, THE FIRING STOPPED IMMEDIATELY AND THE SCREEN AT THAT MOMENT DISPLAYED THE RESISTANCE LIMIT MARK SCREEN. IN THE COLON, THE STAPLE(S) WHICH WAS NOT FORMED REMAINED IN THE U SHAPE AND THE RESECTION WAS NOT COMPLETED. THE ANAL SIDE WAS TREATED AND THE FIRING WAS PERFORMED WITH A NEW RELOAD AND THE RESECTION WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. WHEN THE DEVICE WAS CHECKED, THEY FOUND THE PLASTIC PART OF CARTRIDGE WAS MISALIGNED.
17-12-2018 | Injury | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING IN A LAPAROSCOPIC SIGMOIDECTOMY WITH DOUBLE-STAPLING TECHNIQUE (DST) ANASTOMOSIS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED IN THE COLON FOR COLECTOMY, THE FIRING STOPPED IMMEDIATELY AND THE SCREEN AT THAT MOMENT DISPLAYED THE RESISTANCE LIMIT MARK SCREEN. IN THE COLON, THE STAPLE(S) WHICH WAS NOT FORMED REMAINED IN THE U SHAPE AND THE RESECTION WAS NOT COMPLETED. THE ANAL SIDE WAS TREATED AND THE FIRING WAS PERFORMED WITH A NEW RELOAD AND THE RESECTION WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. WHEN THE DEVICE WAS CHECKED, THEY FOUND THE PLASTIC PART OF CARTRIDGE WAS MISALIGNED.
17-12-2018 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING IN A LAPAROSCOPIC SIGMOIDECTOMY WITH DOUBLE-STAPLING TECHNIQUE (DST) ANASTOMOSIS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED IN THE COLON FOR COLECTOMY, THE FIRING STOPPED IMMEDIATELY AND THE SCREEN AT THAT MOMENT DISPLAYED THE RESISTANCE LIMIT MARK SCREEN. IN THE COLON, THE STAPLE(S) WHICH WAS NOT FORMED REMAINED IN THE U SHAPE AND THE RESECTION WAS NOT COMPLETED. THE ANAL SIDE WAS TREATED AND THE FIRING WAS PERFORMED WITH A NEW RELOAD AND THE RESECTION WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. WHEN THE DEVICE WAS CHECKED, THEY FOUND THE PLASTIC PART OF CARTRIDGE WAS MISALIGNED.
17-12-2018 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Device Displays Incorrect Message | A090201 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING IN A LAPAROSCOPIC SIGMOIDECTOMY WITH DOUBLE-STAPLING TECHNIQUE (DST) ANASTOMOSIS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED IN THE COLON FOR COLECTOMY, THE FIRING STOPPED IMMEDIATELY AND THE SCREEN AT THAT MOMENT DISPLAYED THE RESISTANCE LIMIT MARK SCREEN. IN THE COLON, THE STAPLE(S) WHICH WAS NOT FORMED REMAINED IN THE U SHAPE AND THE RESECTION WAS NOT COMPLETED. THE ANAL SIDE WAS TREATED AND THE FIRING WAS PERFORMED WITH A NEW RELOAD AND THE RESECTION WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. WHEN THE DEVICE WAS CHECKED, THEY FOUND THE PLASTIC PART OF CARTRIDGE WAS MISALIGNED.
17-12-2018 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING IN A LAPAROSCOPIC SIGMOIDECTOMY WITH DOUBLE-STAPLING TECHNIQUE (DST) ANASTOMOSIS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED IN THE COLON FOR COLECTOMY, THE FIRING STOPPED IMMEDIATELY AND THE SCREEN AT THAT MOMENT DISPLAYED THE RESISTANCE LIMIT MARK SCREEN. IN THE COLON, THE STAPLE(S) WHICH WAS NOT FORMED REMAINED IN THE U SHAPE AND THE RESECTION WAS NOT COMPLETED. THE ANAL SIDE WAS TREATED AND THE FIRING WAS PERFORMED WITH A NEW RELOAD AND THE RESECTION WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. WHEN THE DEVICE WAS CHECKED, THEY FOUND THE PLASTIC PART OF CARTRIDGE WAS MISALIGNED.
17-12-2018 | Injury | covidien lp llc north haven | 15-01-2019 | SIGNIA | Device Displays Incorrect Message | A090201 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOR THE FIRST FIRING IN A LAPAROSCOPIC SIGMOIDECTOMY WITH DOUBLE-STAPLING TECHNIQUE (DST) ANASTOMOSIS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED IN THE COLON FOR COLECTOMY, THE FIRING STOPPED IMMEDIATELY AND THE SCREEN AT THAT MOMENT DISPLAYED THE RESISTANCE LIMIT MARK SCREEN. IN THE COLON, THE STAPLE(S) WHICH WAS NOT FORMED REMAINED IN THE U SHAPE AND THE RESECTION WAS NOT COMPLETED. THE ANAL SIDE WAS TREATED AND THE FIRING WAS PERFORMED WITH A NEW RELOAD AND THE RESECTION WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. WHEN THE DEVICE WAS CHECKED, THEY FOUND THE PLASTIC PART OF CARTRIDGE WAS MISALIGNED.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR STOMACH RESECTION, THE DEVICE WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION. THE SURGEON STATED THAT THE LAST GROUP OF STAPLES LOOKED MANGLED AND CLUMPED. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED FOR SMALL BOWEL RESECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. THEY USED A NEW RELOAD WITH THE SAME HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-09-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC POUCH CREATION OF A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE HANDLE MADE AN AUDIBLE 'CRACK' AFTER THE SECOND FIRING. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THE SAME RELOAD WAS USED WITH ANOTHER HANDLE AND WORKED FINE. THERE WAS NO PATIENT INJURY.
14-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO ROU-EN Y BYPASS, STAPLE LOAD WOULD NOT FIRE. THE HANDLE WAS UNABLE TO MOVE ONCE LOADED. THEY OPEN A NEW RELOAD TO COMPLETE THE CASE. IT WAS NOT USED FOR PATIENT.
14-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO ROU-EN Y BYPASS, STAPLE LOAD WOULD NOT FIRE. THE HANDLE WAS UNABLE TO MOVE ONCE LOADED. THEY OPEN A NEW RELOAD TO COMPLETE THE CASE. IT WAS NOT USED FOR PATIENT.
18-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, DURING EXCISION OF INTERLOBULAR FISSURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THEY CHANGED THE HANDLE AND STAPLE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, DURING EXCISION OF INTERLOBULAR FISSURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THEY CHANGED THE HANDLE AND STAPLE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | SIGNIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A COLORECTAL PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A COLORECTAL PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A COLORECTAL PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEY USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS PROCEDURE, WHEN THE SURGEON WAS MAKING THE POUCH, THE SURGEON WAS BALE SQUEEZE THE HANDLE AND PRESS THE GREEN BUTTON BUT THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 15-01-2019 | UNKNOWN ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS PROCEDURE, WHEN THE SURGEON WAS MAKING THE POUCH, THE SURGEON WAS BALE SQUEEZE THE HANDLE AND PRESS THE GREEN BUTTON BUT THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Injury | us surgical puerto rico | 15-01-2019 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOWER ANTERIOR RESECTION OF THE COLON, THE SAFETY LEVER WAS NOT RELEASING AND BROKE-OFF. THE DOCTOR THEN SQUEEZED THE HANDLE. THERE WAS NO RESISTANCE WHEN FIRING AND RELEASING THE STAPLER. THE KNIFE BLADE AND STAPLES DEPLOYED, HOWEVER THE STAPLES DID NOT FULLY FORM. THEY IMMEDIATELY NOTICED AIR. THEY HAD TO SUTURE THE ANASTOMOSIS BECAUSE THE STAPLE LINE FELL APART. THE CASE WAS PROLONGED AROUND 4 HOURS. THE EVENT LED TO EXTENDED HOSPITALIZATION.
18-12-2018 | Injury | us surgical puerto rico | 15-01-2019 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOWER ANTERIOR RESECTION OF THE COLON, THE SAFETY LEVER WAS NOT RELEASING AND BROKE-OFF. THE DOCTOR THEN SQUEEZED THE HANDLE. THERE WAS NO RESISTANCE WHEN FIRING AND RELEASING THE STAPLER. THE KNIFE BLADE AND STAPLES DEPLOYED, HOWEVER THE STAPLES DID NOT FULLY FORM. THEY IMMEDIATELY NOTICED AIR. THEY HAD TO SUTURE THE ANASTOMOSIS BECAUSE THE STAPLE LINE FELL APART. THE CASE WAS PROLONGED AROUND 4 HOURS. THE EVENT LED TO EXTENDED HOSPITALIZATION.
18-12-2018 | Injury | us surgical puerto rico | 15-01-2019 | EEA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOWER ANTERIOR RESECTION OF THE COLON, THE SAFETY LEVER WAS NOT RELEASING AND BROKE-OFF. THE DOCTOR THEN SQUEEZED THE HANDLE. THERE WAS NO RESISTANCE WHEN FIRING AND RELEASING THE STAPLER. THE KNIFE BLADE AND STAPLES DEPLOYED, HOWEVER THE STAPLES DID NOT FULLY FORM. THEY IMMEDIATELY NOTICED AIR. THEY HAD TO SUTURE THE ANASTOMOSIS BECAUSE THE STAPLE LINE FELL APART. THE CASE WAS PROLONGED AROUND 4 HOURS. THE EVENT LED TO EXTENDED HOSPITALIZATION.
18-12-2018 | Injury | us surgical puerto rico | 15-01-2019 | EEA | Gas/Air Leak | A050402 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOWER ANTERIOR RESECTION OF THE COLON, THE SAFETY LEVER WAS NOT RELEASING AND BROKE-OFF. THE DOCTOR THEN SQUEEZED THE HANDLE. THERE WAS NO RESISTANCE WHEN FIRING AND RELEASING THE STAPLER. THE KNIFE BLADE AND STAPLES DEPLOYED, HOWEVER THE STAPLES DID NOT FULLY FORM. THEY IMMEDIATELY NOTICED AIR. THEY HAD TO SUTURE THE ANASTOMOSIS BECAUSE THE STAPLE LINE FELL APART. THE CASE WAS PROLONGED AROUND 4 HOURS. THE EVENT LED TO EXTENDED HOSPITALIZATION.
18-12-2018 | Injury | us surgical puerto rico | 15-01-2019 | EEA | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOWER ANTERIOR RESECTION OF THE COLON, THE SAFETY LEVER WAS NOT RELEASING AND BROKE-OFF. THE DOCTOR THEN SQUEEZED THE HANDLE. THERE WAS NO RESISTANCE WHEN FIRING AND RELEASING THE STAPLER. THE KNIFE BLADE AND STAPLES DEPLOYED, HOWEVER THE STAPLES DID NOT FULLY FORM. THEY IMMEDIATELY NOTICED AIR. THEY HAD TO SUTURE THE ANASTOMOSIS BECAUSE THE STAPLE LINE FELL APART. THE CASE WAS PROLONGED AROUND 4 HOURS. THE EVENT LED TO EXTENDED HOSPITALIZATION.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | ENDO GIA | Protective Measures Problem | A16 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE THE BLADE OF THE CARTRIDGE COMES OUT WHEN THE VESSEL IS STILL. THERE WAS NO PATIENT INVOLVED.
18-12-2018 | Malfunction | covidien lp llc north haven | 15-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, IN A WEDGE SEGMENTAL RESECTION PROCEDURE, THE FIRST FIRING HAD NO PROBLEM. THE SURGEON SETS A NEW RELOAD AND CONFIRMED THE GREEN BUTTON FLASHED, THEN TRIED TO FIRE BUT FAILED. THEY REPLACED IT WITH ANOTHER RELOAD, THE SURGEON SET IT AND CONFIRMED THE GREEN BUTTON FLASHED, THEN HE TRIED TO FIRE THE DEVICE, BUT THE FIRING WAS NOT PERFORMED, EITHER. THEY USED ANOTHER DEVICE AND THE PROCEDURE WAS COMPLETED WITH NO PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE EVENT OCCURRED DURING THE PROCEDURE, BUT THE PRODUCT WAS NOT USED FOR PATIENT.
02-01-2019 | Injury | unk | 15-01-2019 | TITANIUM SURGICAL TACKS | Migration or Expulsion of Device | A0104 | Vessel Or Plaque, Device Embedded In | Event Description: TITANIUM SURGICAL TACKS IMPLANTED IN HERNIA MESH SURGERY (AT THE WAIST) IN 2013 HAVE MIGRATED THROUGHOUT THE ABDOMINAL AREA, HIP, GROIN AND EMBEDDED IN INTESTINES, RIB AREAS, ETC. ON (B)(6) 2019, SURGERY FOR CHRONIC PAIN REVEALED A CLUSTER OF TACKS AT LOWER LEFT RIB. APPROX 8 CTS OVER LAST 5 YEARS DID NOT SHOW TITANIUM TACKS. NEXT SURGERY SCHEDULED FOR (B)(6) 2019 TO REMOVE AS MANY TACKS AS POSSIBLE WITHOUT JEOPARDIZING PT'S LIFE.
02-01-2019 | Injury | unk | 15-01-2019 | TITANIUM SURGICAL TACKS | Migration or Expulsion of Device | A0104 | Pain | Event Description: TITANIUM SURGICAL TACKS IMPLANTED IN HERNIA MESH SURGERY (AT THE WAIST) IN 2013 HAVE MIGRATED THROUGHOUT THE ABDOMINAL AREA, HIP, GROIN AND EMBEDDED IN INTESTINES, RIB AREAS, ETC. ON (B)(6) 2019, SURGERY FOR CHRONIC PAIN REVEALED A CLUSTER OF TACKS AT LOWER LEFT RIB. APPROX 8 CTS OVER LAST 5 YEARS DID NOT SHOW TITANIUM TACKS. NEXT SURGERY SCHEDULED FOR (B)(6) 2019 TO REMOVE AS MANY TACKS AS POSSIBLE WITHOUT JEOPARDIZING PT'S LIFE.
03-01-2019 | Malfunction | justright surgical | 16-01-2019 | JUSTRIGHT 5MM STAPLER | Mechanical Problem | A05 | Tissue Damage | Event Description: STAPLER LOADED BY SCRUB, TESTED PRIOR TO ENTRY INTO PORT BY SURGEON, PLACED INTO PORT, FIRED, JAWS OPENED TO STAPLER ENOUGH FOR TISSUE TO FALL FROM STAPLER, #5 BUTTON IN THE BACK OF THE STAPLER DID NOT OPEN FULLY, SURGEON PULLED STAPLER BACK RECOILED AND A HOLE OCCURRED IN THE DIAPHRAGM. THE OPENING WAS SUTURED INTRAOPERATIVELY. THE PATIENT REMAINED HEMODYNAMICALLY STABLE THROUGHOUT.
12-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC WEDGE SEGMENTAL RESECTION, THE LUNG PARENCHYMA WAS APPROACHED FOR TUMOR RESECTION. THE DEVICE PERFORMED THE FIRST FIRING THEN THE JAWS WERE ATTEMPTED TO BE OPENED AND THEY WERE NOT ABLE TO OPEN SMOOTHLY. WHEN THE DEVICE WAS CHECKED, THERE WERE STAPLES THAT WERE UNABLE TO BE CLOSED PROPERLY AND WERE NOT B-SHAPED. IT THE JAWS WERE HARD TO BE OPENED DUE TO NON B-SHAPED STAPLES. WHEN THE RELOAD WAS COLLECTED AND CHECKED, THE STAPLES REMAINED AT 10MM FROM THE DISTAL END OF THE CARTRIDGE. THE TISSUE WAS NOT SO THICK. A NEW RELOAD WAS USED AND THE TUMOR WAS EXTRACTED. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC WEDGE SEGMENTAL RESECTION, THE LUNG PARENCHYMA WAS APPROACHED FOR TUMOR RESECTION. THE DEVICE PERFORMED THE FIRST FIRING THEN THE JAWS WERE ATTEMPTED TO BE OPENED AND THEY WERE NOT ABLE TO OPEN SMOOTHLY. WHEN THE DEVICE WAS CHECKED, THERE WERE STAPLES THAT WERE UNABLE TO BE CLOSED PROPERLY AND WERE NOT B-SHAPED. IT THE JAWS WERE HARD TO BE OPENED DUE TO NON B-SHAPED STAPLES. WHEN THE RELOAD WAS COLLECTED AND CHECKED, THE STAPLES REMAINED AT 10MM FROM THE DISTAL END OF THE CARTRIDGE. THE TISSUE WAS NOT SO THICK. A NEW RELOAD WAS USED AND THE TUMOR WAS EXTRACTED. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC WEDGE SEGMENTAL RESECTION, THE LUNG PARENCHYMA WAS APPROACHED FOR TUMOR RESECTION. THE DEVICE PERFORMED THE FIRST FIRING THEN THE JAWS WERE ATTEMPTED TO BE OPENED AND THEY WERE NOT ABLE TO OPEN SMOOTHLY. WHEN THE DEVICE WAS CHECKED, THERE WERE STAPLES THAT WERE UNABLE TO BE CLOSED PROPERLY AND WERE NOT B-SHAPED. IT THE JAWS WERE HARD TO BE OPENED DUE TO NON B-SHAPED STAPLES. WHEN THE RELOAD WAS COLLECTED AND CHECKED, THE STAPLES REMAINED AT 10MM FROM THE DISTAL END OF THE CARTRIDGE. THE TISSUE WAS NOT SO THICK. A NEW RELOAD WAS USED AND THE TUMOR WAS EXTRACTED. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC WEDGE SEGMENTAL RESECTION, THE LUNG PARENCHYMA WAS APPROACHED FOR TUMOR RESECTION. THE DEVICE PERFORMED THE FIRST FIRING THEN THE JAWS WERE ATTEMPTED TO BE OPENED AND THEY WERE NOT ABLE TO OPEN SMOOTHLY. WHEN THE DEVICE WAS CHECKED, THERE WERE STAPLES THAT WERE UNABLE TO BE CLOSED PROPERLY AND WERE NOT B-SHAPED. IT THE JAWS WERE HARD TO BE OPENED DUE TO NON B-SHAPED STAPLES. WHEN THE RELOAD WAS COLLECTED AND CHECKED, THE STAPLES REMAINED AT 10MM FROM THE DISTAL END OF THE CARTRIDGE. THE TISSUE WAS NOT SO THICK. A NEW RELOAD WAS USED AND THE TUMOR WAS EXTRACTED. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC WEDGE SEGMENTAL RESECTION, THE LUNG PARENCHYMA WAS APPROACHED FOR TUMOR RESECTION. THE DEVICE PERFORMED THE FIRST FIRING THEN THE JAWS WERE ATTEMPTED TO BE OPENED AND THEY WERE NOT ABLE TO OPEN SMOOTHLY. WHEN THE DEVICE WAS CHECKED, THERE WERE STAPLES THAT WERE UNABLE TO BE CLOSED PROPERLY AND WERE NOT B-SHAPED. IT THE JAWS WERE HARD TO BE OPENED DUE TO NON B-SHAPED STAPLES. WHEN THE RELOAD WAS COLLECTED AND CHECKED, THE STAPLES REMAINED AT 10MM FROM THE DISTAL END OF THE CARTRIDGE. THE TISSUE WAS NOT SO THICK. A NEW RELOAD WAS USED AND THE TUMOR WAS EXTRACTED. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC WEDGE SEGMENTAL RESECTION, THE LUNG PARENCHYMA WAS APPROACHED FOR TUMOR RESECTION. THE DEVICE PERFORMED THE FIRST FIRING THEN THE JAWS WERE ATTEMPTED TO BE OPENED AND THEY WERE NOT ABLE TO OPEN SMOOTHLY. WHEN THE DEVICE WAS CHECKED, THERE WERE STAPLES THAT WERE UNABLE TO BE CLOSED PROPERLY AND WERE NOT B-SHAPED. IT THE JAWS WERE HARD TO BE OPENED DUE TO NON B-SHAPED STAPLES. WHEN THE RELOAD WAS COLLECTED AND CHECKED, THE STAPLES REMAINED AT 10MM FROM THE DISTAL END OF THE CARTRIDGE. THE TISSUE WAS NOT SO THICK. A NEW RELOAD WAS USED AND THE TUMOR WAS EXTRACTED. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | UNKNOWN ENDO GI | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC WEDGE SEGMENTAL RESECTION, THE LUNG PARENCHYMA WAS APPROACHED FOR TUMOR RESECTION. THE DEVICE PERFORMED THE FIRST FIRING THEN THE JAWS WERE ATTEMPTED TO BE OPENED AND THEY WERE NOT ABLE TO OPEN SMOOTHLY. WHEN THE DEVICE WAS CHECKED, THERE WERE STAPLES THAT WERE UNABLE TO BE CLOSED PROPERLY AND WERE NOT B-SHAPED. IT THE JAWS WERE HARD TO BE OPENED DUE TO NON B-SHAPED STAPLES. WHEN THE RELOAD WAS COLLECTED AND CHECKED, THE STAPLES REMAINED AT 10MM FROM THE DISTAL END OF THE CARTRIDGE. THE TISSUE WAS NOT SO THICK. A NEW RELOAD WAS USED AND THE TUMOR WAS EXTRACTED. THERE WAS NO PATIENT INJURY.
12-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | UNKNOWN ENDO GI | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC WEDGE SEGMENTAL RESECTION, THE LUNG PARENCHYMA WAS APPROACHED FOR TUMOR RESECTION. THE DEVICE PERFORMED THE FIRST FIRING THEN THE JAWS WERE ATTEMPTED TO BE OPENED AND THEY WERE NOT ABLE TO OPEN SMOOTHLY. WHEN THE DEVICE WAS CHECKED, THERE WERE STAPLES THAT WERE UNABLE TO BE CLOSED PROPERLY AND WERE NOT B-SHAPED. IT THE JAWS WERE HARD TO BE OPENED DUE TO NON B-SHAPED STAPLES. WHEN THE RELOAD WAS COLLECTED AND CHECKED, THE STAPLES REMAINED AT 10MM FROM THE DISTAL END OF THE CARTRIDGE. THE TISSUE WAS NOT SO THICK. A NEW RELOAD WAS USED AND THE TUMOR WAS EXTRACTED. THERE WAS NO PATIENT INJURY.
03-11-2018 | Malfunction | ethicon endo-surgery | 16-01-2019 | ECHELON FLEX 45 STANDARD | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP GASTROPLASTY PROCEDURE,THE ECHELON 45 FLEX CAUGHT DURING THE STAPLING HAVING TO BE ACTIVATED THE EMERGENCY BUTTON TO BE ABLE TO OPEN, NOT PERFORMING THE STAPLING, LOSING THE LOAD.
18-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE SURGEON FIRED THE RELOAD WITH THE UNIVERSAL HANDLE, BUT IT STOPPED AT THE VERY BEGINNING. THE HANDLE AND RELOAD WERE BOTH CHANGED TO COMPLETE THE PROCEDURE.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | TRI-STAPLE 2.0 | Component Missing | A020602 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE STAPLING OF THE STOMACH, THE LOADING UNIT CAME OUT OF PACKAGING WITH NO STAPLES IN THE CARTRIDGE. THE UNIT WAS LOADED AND GRASPED THE STOMACH BUT WOULD NOT FIRE. UPON REMOVAL FROM THE STOMACH, THE TECHNICIAN NOTICED THAT THE CARTRIDGE DID NOT HAVE STAPLES IN THE UNIT. IT WAS NOTED THAT THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD WAS OPENED IN ORDER TO COMPLETE THE CASE. IT WAS NOTED THAT THE SAME ADAPTER AND HANDLE WERE USED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE STAPLING OF THE STOMACH, THE LOADING UNIT CAME OUT OF PACKAGING WITH NO STAPLES IN THE CARTRIDGE. THE UNIT WAS LOADED AND GRASPED THE STOMACH BUT WOULD NOT FIRE. UPON REMOVAL FROM THE STOMACH, THE TECHNICIAN NOTICED THAT THE CARTRIDGE DID NOT HAVE STAPLES IN THE UNIT. IT WAS NOTED THAT THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD WAS OPENED IN ORDER TO COMPLETE THE CASE. IT WAS NOTED THAT THE SAME ADAPTER AND HANDLE WERE USED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE STAPLING OF THE STOMACH, THE LOADING UNIT CAME OUT OF PACKAGING WITH NO STAPLES IN THE CARTRIDGE. THE UNIT WAS LOADED AND GRASPED THE STOMACH BUT WOULD NOT FIRE. UPON REMOVAL FROM THE STOMACH, THE TECHNICIAN NOTICED THAT THE CARTRIDGE DID NOT HAVE STAPLES IN THE UNIT. IT WAS NOTED THAT THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD WAS OPENED IN ORDER TO COMPLETE THE CASE. IT WAS NOTED THAT THE SAME ADAPTER AND HANDLE WERE USED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE STAPLING OF THE STOMACH, THE LOADING UNIT CAME OUT OF PACKAGING WITH NO STAPLES IN THE CARTRIDGE. THE UNIT WAS LOADED AND GRASPED THE STOMACH BUT WOULD NOT FIRE. UPON REMOVAL FROM THE STOMACH, THE TECHNICIAN NOTICED THAT THE CARTRIDGE DID NOT HAVE STAPLES IN THE UNIT. IT WAS NOTED THAT THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD WAS OPENED IN ORDER TO COMPLETE THE CASE. IT WAS NOTED THAT THE SAME ADAPTER AND HANDLE WERE USED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 16-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE RELOAD CAN NOT BE FIRE OFF. NO OTHER INFORMATION IS AVAILABLE. NO PATIENT CONSEQUENCE REPORTED.
12-09-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE TROCATTER DID NOT FIRE. THE DEVICE WAS CHANGED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
12-09-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE TROCATTER DID NOT FIRE. THE DEVICE WAS CHANGED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | ethicon endo-surgery | 16-01-2019 | PWRD ECH FLEX 60MM 28CM SHAFT | Overheating of Device | A1005 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE TOTAL LAPAROSCOPIC RADICAL GASTRECTOMY SURGERY, COULD NOT FIRE FARTHER AFTER FIRED 2/3 WHEN SEVERING GASTRIC BODY TISSUE. TRIED REVERSING BLADE AND BATTERY PROCESS BUT DID NOT WORK, AND THE BATTERY IS GETTING HOT. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2018 | Malfunction | ethicon endo-surgery | 16-01-2019 | PWRD ECH FLEX 60MM 28CM SHAFT | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE TOTAL LAPAROSCOPIC RADICAL GASTRECTOMY SURGERY, COULD NOT FIRE FARTHER AFTER FIRED 2/3 WHEN SEVERING GASTRIC BODY TISSUE. TRIED REVERSING BLADE AND BATTERY PROCESS BUT DID NOT WORK, AND THE BATTERY IS GETTING HOT. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-11-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLER WAS ABLE TO FIRE BUT THE RELOAD WAS STUCK AND COULD NOT BE OPENED. IT WAS STATED THAT THE SURGEON USED THE EGG FORCEPS TO REMOVED THE STAPLES. THEY CHANGED DEVICE TO COMPLETE THE CASE. IT WAS ALSO STATED THAT THE PATIENT RECEIVED CHEMOTHERAPY OR RADIATION THERAPY. THERE WAS NO PATIENT INJURY.
20-11-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLER WAS ABLE TO FIRE BUT THE RELOAD WAS STUCK AND COULD NOT BE OPENED. IT WAS STATED THAT THE SURGEON USED THE EGG FORCEPS TO REMOVED THE STAPLES. THEY CHANGED DEVICE TO COMPLETE THE CASE. IT WAS ALSO STATED THAT THE PATIENT RECEIVED CHEMOTHERAPY OR RADIATION THERAPY. THERE WAS NO PATIENT INJURY.
17-10-2018 | Malfunction | ethicon endo-surgery | 16-01-2019 | ECHELON FLEX 60 LONG | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE SLEEVE GASTRIC SURGERY, COULD NOT FIRE WITH LONG60A AND ECR60D, THE CHANGED ANOTHER ECR60D, COULD NOT OPEN THE JAW AFTER FIRED. OPENED THE JAW MANUALLY WITH BIG FORCE. THEN CHANGED WITH ANOTHER NEW LONG60A AND ECR60D, AFTER FIRED, NOTED THE STAPLES MALFORMED. SUTURE WAS USED TO SEW THE WOUND. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2018 | Malfunction | ethicon endo-surgery | 16-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A DISTAL PANC, THE STAPLER FIRING SEQUENCE STOPPED ON THE FIRST FIRE ON EXTREMELY THICK TISSUE USING A GREEN RELOAD. FIRING SEQUENCE STOPPED, TRIED REVERSE, TRIED MANUAL RATCHET, OPENED UP HANDLE BUT JAWS WEREN¿T OPENING. SURGEON WAS ASKED TO SEE IF KNIFE BLADE IS IN HOME POSITION, HE SAID LOOKED 10MM FROM HOME POSITION. MANAGED TO FINALLY GET OFF TISSUE. IT WAS ALSO REPORTED THAT THE PATIENT ENDED UP DOING VERY WELL, HE LOOKS GOOD. IT WAS NOT A DEVICE FAILURE, HIS PANCREAS WAS SO HARD AND CALCIFIED FROM HIS CHRONIC PANCREATITIS THAT IT BROKE STAPLER.
01-01-2018 | Malfunction | ethicon endo-surgery | 16-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A DISTAL PANC, THE STAPLER FIRING SEQUENCE STOPPED ON THE FIRST FIRE ON EXTREMELY THICK TISSUE USING A GREEN RELOAD. FIRING SEQUENCE STOPPED, TRIED REVERSE, TRIED MANUAL RATCHET, OPENED UP HANDLE BUT JAWS WEREN¿T OPENING. SURGEON WAS ASKED TO SEE IF KNIFE BLADE IS IN HOME POSITION, HE SAID LOOKED 10MM FROM HOME POSITION. MANAGED TO FINALLY GET OFF TISSUE. IT WAS ALSO REPORTED THAT THE PATIENT ENDED UP DOING VERY WELL, HE LOOKS GOOD. IT WAS NOT A DEVICE FAILURE, HIS PANCREAS WAS SO HARD AND CALCIFIED FROM HIS CHRONIC PANCREATITIS THAT IT BROKE STAPLER.
13-12-2018 | Malfunction | covidien lp | 16-01-2019 | ENDO GIA ULTRA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: THE DOCTOR WAS USING AN ENDOGIA STAPLER AND IT MALFUNCTIONED AFTER FIRING. A SECOND ENDOGIA STAPLER WAS INTRODUCED TO THE FIELD AND BROKE UPON UNLOCKING THE STAPLER. ANOTHER DEVICE WAS INTRODUCED TO THE FIELD AND THE PROCEDURE WAS COMPLETED AS PLANNED WITH NO HARM TO THE PATIENT. MANUFACTURER RESPONSE FOR STAPLE, IMPLANTABLE, ENDO GIA ULTRA (PER SITE REPORTER). DEVICE WILL BE RETURNED FOR FAILURE ANALYSIS. ANY FURTHER CORRESPONDENCE WILL BE SHARED. MANUFACTURER RESPONSE FOR STAPLE, IMPLANTABLE, ENDO GIA ULTRA (PER SITE REPORTER). DEVICE WILL BE RETURNED FOR FAILURE ANALYSIS. ANY FURTHER CORRESPONDENCE WILL BE SHARED.
13-12-2018 | Malfunction | covidien lp | 16-01-2019 | ENDO GIA ULTRA | Defective Component | A0202 | No Consequences Or Impact To Patient | Event Description: THE DOCTOR WAS USING AN ENDOGIA STAPLER AND IT MALFUNCTIONED AFTER FIRING. A SECOND ENDOGIA STAPLER WAS INTRODUCED TO THE FIELD AND BROKE UPON UNLOCKING THE STAPLER. ANOTHER DEVICE WAS INTRODUCED TO THE FIELD AND THE PROCEDURE WAS COMPLETED AS PLANNED WITH NO HARM TO THE PATIENT. MANUFACTURER RESPONSE FOR STAPLE, IMPLANTABLE, ENDO GIA ULTRA (PER SITE REPORTER). DEVICE WILL BE RETURNED FOR FAILURE ANALYSIS. ANY FURTHER CORRESPONDENCE WILL BE SHARED. MANUFACTURER RESPONSE FOR STAPLE, IMPLANTABLE, ENDO GIA ULTRA (PER SITE REPORTER). DEVICE WILL BE RETURNED FOR FAILURE ANALYSIS. ANY FURTHER CORRESPONDENCE WILL BE SHARED.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | PREMIUM PLUS CEEA | Material Twisted/Bent | A040609 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLON CANCER RADICAL SURGERY, THE DEVICE BROKE AS THE ANVIL ROLL OVER THEREFORE WAS NOT USED AND WAS BENT. DEVICE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE PERFORMING RESECTION OF LUNG, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE. A NEW RELOAD WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE PERFORMING RESECTION OF LUNG, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE. A NEW RELOAD WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2018 | Injury | ethicon endo-surgery | 16-01-2019 | WHITE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RIGHT LOWER LOBE RESECTION, WHEN SEVERING THE PULMONARY VEIN ADJACENT TO THE PERICARDIUM, AFTER FIRED, NOTED SOME MIDDLE STAPLES MISSING. ENLARGED THE SMALL INCISION AND USED SUTURE TO OVERSEW AND STOP BLEEDING. THE PATIENT LOST ABOUT 1000ML BLOOD, THE SURGERY DELAYED ABOUT 40 MINS. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS EXTENDED, BUT COULD NOT CONFIRM THE SPECIFIC TIME.
15-12-2018 | Injury | ethicon endo-surgery | 16-01-2019 | WHITE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RIGHT LOWER LOBE RESECTION, WHEN SEVERING THE PULMONARY VEIN ADJACENT TO THE PERICARDIUM, AFTER FIRED, NOTED SOME MIDDLE STAPLES MISSING. ENLARGED THE SMALL INCISION AND USED SUTURE TO OVERSEW AND STOP BLEEDING. THE PATIENT LOST ABOUT 1000ML BLOOD, THE SURGERY DELAYED ABOUT 40 MINS. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW. THE PATIENT¿S LENGTH OF HOSPITAL STAY IS EXTENDED, BUT COULD NOT CONFIRM THE SPECIFIC TIME.
17-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORASCOPIC WEDGE SEGMENTAL RESECTION, THE RELOAD WAS CONNECTED TO THE HANDLE, WITH THE YELLOW TAB STILL ATTACHED TO THE RELOAD. THE OPENING AND CLOSING THE JAWS WAS CHECKED WITH NO PROBLEM HOWEVER ON THE FIRST FIRING, THE SURGEON CLAMPED THE PULMONARY TISSUE, BUT THE CLAMPING WAS LOOSE AND THE TISSUE WAS SLIPPED OUT. THEY RE-CONNECTED THE RELOAD BUT THE STATUS WAS THE SAME. A NEW DEVICE WAS USED AND IT WORKED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORASCOPIC WEDGE SEGMENTAL RESECTION, THE RELOAD WAS CONNECTED TO THE HANDLE, WITH THE YELLOW TAB STILL ATTACHED TO THE RELOAD. THE OPENING AND CLOSING THE JAWS WAS CHECKED WITH NO PROBLEM HOWEVER ON THE FIRST FIRING, THE SURGEON CLAMPED THE PULMONARY TISSUE, BUT THE CLAMPING WAS LOOSE AND THE TISSUE WAS SLIPPED OUT .THEY RE-CONNECTED THE RELOAD BUT THE STATUS WAS THE SAME. A NEW DEVICE WAS USED AND IT WORKED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | UNKNOWN ENDO GI | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS, ON JEJUNOSTOMY ANASTOMOSIS, AFTER FIRING THE 45MM BLUE CARTRIDGE, THE B FORMATION DID NOT OCCUR IN THE STAPLES AT THE LAST PART OF THE CARTRIDGE AND THE STAPLE LINE WAS INCOMPLETE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS, ON JEJUNOSTOMY ANASTOMOSIS, AFTER FIRING THE 45MM BLUE CARTRIDGE, THE B FORMATION DID NOT OCCUR IN THE STAPLES AT THE LAST PART OF THE CARTRIDGE AND THE STAPLE LINE WAS INCOMPLETE. THERE WAS NO PATIENT INJURY.
12-12-2018 | Injury | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER A RIGHT COLECTOMY BEING PERFORMED, PATIENT ENDED UP WITH LEAK AT ANASTOMOSIS; SURGEON BROUGHT PATIENT BACK AND RE-DID ANASTOMOSIS.
12-12-2018 | Injury | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER A RIGHT COLECTOMY BEING PERFORMED, PATIENT ENDED UP WITH LEAK AT ANASTOMOSIS; SURGEON BROUGHT PATIENT BACK AND RE-DID ANASTOMOSIS.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Read Input Signal | A1301 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, DURING THE LAST FIRING ON FUNDUS, WHILE LOADING THE RELOAD INTO ADAPTER THE RELOAD WAS AUTO ARTICULATING ITSELF. HANDLE WAS NOT ABLE TO DETECT RELOAD.. THERE WAS CALIBRATION PROBLEM BETWEEN RELOAD AND ADAPTER. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Unintended Movement | A0512 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, DURING THE LAST FIRING ON FUNDUS, WHILE LOADING THE RELOAD INTO ADAPTER THE RELOAD WAS AUTO ARTICULATING ITSELF. HANDLE WAS NOT ABLE TO DETECT RELOAD.. THERE WAS CALIBRATION PROBLEM BETWEEN RELOAD AND ADAPTER. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | IDRIVE ULTRA | Failure to Read Input Signal | A1301 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, DURING THE LAST FIRING ON FUNDUS, WHILE LOADING THE RELOAD INTO ADAPTER THE RELOAD WAS AUTO ARTICULATING ITSELF. HANDLE WAS NOT ABLE TO DETECT RELOAD. THERE WAS CALIBRATION PROBLEM BETWEEN RELOAD AND ADAPTER. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | IDRIVE ULTRA | Unintended Movement | A0512 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, DURING THE LAST FIRING ON FUNDUS, WHILE LOADING THE RELOAD INTO ADAPTER THE RELOAD WAS AUTO ARTICULATING ITSELF. HANDLE WAS NOT ABLE TO DETECT RELOAD. THERE WAS CALIBRATION PROBLEM BETWEEN RELOAD AND ADAPTER. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | SIGNIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN APPENDECTOMY PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THE RELOAD WAS REPLACED. TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN APPENDECTOMY PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THE RELOAD WAS REPLACED. TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE SURGEON DID THE FIRST FIRE OF THE POUCH DURING A LAPAROSCOPIC GASTROPLASTY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT HE STAPLER DID NOT FIRE. THEY OPENED A NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE SURGEON DID THE FIRST FIRE OF THE POUCH DURING A LAPAROSCOPIC GASTROPLASTY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HE STAPLER DID NOT FIRE. THEY OPENED A NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Device Slipped | A051204 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO LOBECTOMY PROCEDURE, ON THE SECOND FIRING OF THE DEVICE, THE GREEN BUTTON WAS PRESSED, BUT THE DEVICE WAS UNABLE TO BE FIRED, AND WHEN THE SURGEON TRIED TO SQUEEZED THE HANDLE A CRACK SOUND WAS HEARD. THEY USED ANOTHER RELOAD, HOWEVER, THE HANDLE RAN IDLE AND BECAME UNUSABLE, SO THE HANDLE WAS ALSO EXCHANGED.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO LOBECTOMY PROCEDURE, ON THE SECOND FIRING OF THE DEVICE, THE GREEN BUTTON WAS PRESSED, BUT THE DEVICE WAS UNABLE TO BE FIRED, AND WHEN THE SURGEON TRIED TO SQUEEZED THE HANDLE A CRACK SOUND WAS HEARD. THEY USED ANOTHER RELOAD, HOWEVER, THE HANDLE RAN IDLE AND BECAME UNUSABLE, SO THE HANDLE WAS ALSO EXCHANGED.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO LOBECTOMY PROCEDURE, ON THE SECOND FIRING OF THE DEVICE, THE GREEN BUTTON WAS PRESSED, BUT THE DEVICE WAS UNABLE TO BE FIRED, AND WHEN THE SURGEON TRIED TO SQUEEZED THE HANDLE A CRACK SOUND WAS HEARD. THEY USED ANOTHER RELOAD, HOWEVER, THE HANDLE RAN IDLE AND BECAME UNUSABLE, SO THE HANDLE WAS ALSO EXCHANGED.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO LOBECTOMY PROCEDURE, ON THE SECOND FIRING OF THE DEVICE, THE GREEN BUTTON WAS PRESSED BUT THE DEVICE WAS UNABLE TO BE FIRED, AND WHEN THE SURGEON TRIED TO SQUEEZED THE HANDLE A CRACK SOUND WAS HEARD. THEY USED ANOTHER RELOAD, HOWEVER THE HANDLE RAN IDLE AND BECAME UNUSABLE, SO THE HANDLE WAS ALSO EXCHANGED.
17-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOIDECTOMY, WHEN RESECTING SIGMOID COLON, AFTER CLAMPING THE DEVICE INTO THE TISSUE, AND WHEN THE SURGEON ATTEMPTED TO FIRE THE DEVICE, THE BLADE DID NOT ADVANCE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOIDECTOMY, WHEN RESECTING SIGMOID COLON, AFTER CLAMPING THE DEVICE INTO THE TISSUE, AND WHEN THE SURGEON ATTEMPTED TO FIRE THE DEVICE, THE BLADE DID NOT ADVANCE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-11-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC LOBECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-11-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC LOBECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2018 | Malfunction | ethicon endo-surgery | 16-01-2019 | ILS 29MM, CURVED | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DURING STAPLING, THE DEVICE HAD TO BE PASSED THROUGH THE RECTUM. AT THE TIME OF ENGAGEMENT, THE DEVICE DIDN'T WORK DESPITE ACTIVATION BY THE SURGEON. THE DEVICE TORE THE PATIENT'S RECTUM. THE SURGEON HAD TO MAKE AGAIN THE TIMES BEFORE THE USE OF THE NEW CLAMP FROM A DIFFERENT SUPPLIER. THE NEW CLAMP WAS LARGER THAN NECESSARY. EXTENSION OF THE PROCEDURE.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WERE ISSUES WITH 3 RELOADS. FOR THE FIRST REPORTED RELOAD, THE SURGEON WAS ABLE TO PUSH THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THE OTHER RELOAD HAD AN INCOMPLETE STAPLE LINE PROXIMALLY. THE LAST RELOAD HAD POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE IN THE CENTER. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE STOMACH DURING A LAPAROSCOPIC SLEEVE PROCEDURE, THE STAPLER BLOCKED AND THE GREEN BUTTON COULD NOT BE PRESSED AND DOES NOT ALLOW THE RELOAD TO ADVANCE. THE RELOAD WAS CHANGED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE STOMACH DURING A LAPAROSCOPIC SLEEVE PROCEDURE, THE STAPLER BLOCKED AND THE GREEN BUTTON COULD NOT BE PRESSED AND DOES NOT ALLOW THE RELOAD TO ADVANCE. THE RELOAD WAS CHANGED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | UNKNOWN ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE STOMACH DURING A LAPAROSCOPIC SLEEVE PROCEDURE, THE STAPLER BLOCKED AND THE GREEN BUTTON COULD NOT BE PRESSED AND DOES NOT ALLOW THE RELOAD TO ADVANCE. THE RELOAD WAS CHANGED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | UNKNOWN ENDO GI | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE STOMACH DURING A LAPAROSCOPIC SLEEVE PROCEDURE, THE STAPLER BLOCKED AND THE GREEN BUTTON COULD NOT BE PRESSED AND DOES NOT ALLOW THE RELOAD TO ADVANCE. THE RELOAD WAS CHANGED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-03-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC, COLON CANCER RESECTION PROCEDURE, THE DEVICE GREEN BUTTON WAS PRESSED, HOWEVER THE HANDLE WAS NOT SQUEEZE, THUS THE DEVICE DID NOT FIRE. REPLACEMENT OF SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
05-03-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC, COLON CANCER RESECTION PROCEDURE, THE DEVICE GREEN BUTTON WAS PRESSED, HOWEVER THE HANDLE WAS NOT SQUEEZE, THUS THE DEVICE DID NOT FIRE. REPLACEMENT OF SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
05-03-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC, COLON CANCER RESECTION PROCEDURE, THE DEVICE GREEN BUTTON WAS PRESSED, HOWEVER THE HANDLE WAS NOT SQUEEZE, THUS THE DEVICE DID NOT FIRE. REPLACEMENT OF SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
05-03-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC, COLON CANCER RESECTION PROCEDURE, THE DEVICE GREEN BUTTON WAS PRESSED, HOWEVER THE HANDLE WAS NOT SQUEEZE, THUS THE DEVICE DID NOT FIRE. REPLACEMENT OF SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
13-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ESOPHAGECTOMY PROCEDURE, FOR THE FIRST FIRING, THE RELOAD WAS LOADED AND APPROACHED TO THE TISSUE, EVEN THOUGH THE GREEN BUTTON WAS PRESSED, THE DEVICE WAS UNABLE TO BE FIRED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO INJURY.
13-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ESOPHAGECTOMY PROCEDURE, FOR THE FIRST FIRING, THE RELOAD WAS LOADED AND APPROACHED TO THE TISSUE, EVEN THOUGH THE GREEN BUTTON WAS PRESSED, THE DEVICE WAS UNABLE TO BE FIRED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT. NO INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTRECTOMY AND WHILE STAPLING THE VESSEL, THE STAPLER DID NOT FIRE. THE DOCTOR REMOVED THE BROKEN ONE AND REPLACED BY A NEW STAPLER AND RELOAD TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTRECTOMY AND WHILE STAPLING THE VESSEL, THE STAPLER DID NOT FIRE. THE DOCTOR REMOVED THE BROKEN ONE AND REPLACED BY A NEW STAPLER AND RELOAD TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT HEMICOLECTOMY, THE STAPLER WAS STICKING OUT OF THE DISTAL AREA. IT WAS REPORTED A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | SIGNIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE TRANSECTING THE STOMACH, THE JAWS OF THE DEVICE LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGICAL TEAM REMOVED THE ADAPTER FROM THE SHELL/HANDLE COMPONENT, WAITED, THEN RECONNECTED AND THE DEVICE RESET, OPENING THE JAWS. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | SIGNIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE TRANSECTING THE STOMACH, THE JAWS OF THE DEVICE LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGICAL TEAM REMOVED THE ADAPTER FROM THE SHELL/HANDLE COMPONENT, WAITED, THEN RECONNECTED AND THE DEVICE RESET, OPENING THE JAWS. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN TRANSECTING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT STAPLER IMMEDIATELY STOPPED WHEN THE STAPLER STARTED TO FIRE. THIS RESULTED TO AN INCOMPLETE STAPLE LINE PROXIMALLY. THEY OPENED JAWS AND GOT ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN TRANSECTING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT STAPLER IMMEDIATELY STOPPED WHEN THE STAPLER STARTED TO FIRE. THIS RESULTED TO AN INCOMPLETE STAPLE LINE PROXIMALLY. THEY OPENED JAWS AND GOT ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN TRANSECTING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT STAPLER IMMEDIATELY STOPPED WHEN THE STAPLER STARTED TO FIRE. THIS RESULTED TO AN INCOMPLETE STAPLE LINE PROXIMALLY. THEY OPENED JAWS AND GOT ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN TRANSECTING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT STAPLER IMMEDIATELY STOPPED WHEN THE STAPLER STARTED TO FIRE. THIS RESULTED TO AN INCOMPLETE STAPLE LINE PROXIMALLY. THEY OPENED JAWS AND GOT ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
17-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | EEA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOWER ANTERIOR RESECTION PROCEDURE, WHILE PERFORMING ANASTAMOSIS OF ILEUM, THE FIRST STAPLER USED DID NOT MAKE A COMPLETE FIRE AND AIR LEAK TEST REVEALED LEAKING. SURGEON REPORTED NOT FEELING/HEARING THE NORMAL CRUNCH THAT THE STAPLER MAKES WHEN FIRING. THE SECOND STAPLER WOULD NOT RELEASE FROM THE TISSUE. IN ORDER TO AVOID FURTHER PATIENT INJURY, THE ANVIL WAS COMPLETELY RELEASED. PORTION OF ILEUM WAS DAMAGED WITH MISFIRED STAPLES. TWO STAPLERS HAD TO BE USED WHICH ADDED TIME, AND UNPLANNED ILEOSTOMY HAD TO BE CREATED TO COMPLETE THE CASE. EXTENDED HOSPITALIZATION WAS ALSO NOTED.
17-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | EEA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOWER ANTERIOR RESECTION PROCEDURE, WHILE PERFORMING ANASTAMOSIS OF ILEUM, THE FIRST STAPLER USED DID NOT MAKE A COMPLETE FIRE AND AIR LEAK TEST REVEALED LEAKING. SURGEON REPORTED NOT FEELING/HEARING THE NORMAL CRUNCH THAT THE STAPLER MAKES WHEN FIRING. THE SECOND STAPLER WOULD NOT RELEASE FROM THE TISSUE. IN ORDER TO AVOID FURTHER PATIENT INJURY, THE ANVIL WAS COMPLETELY RELEASED. PORTION OF ILEUM WAS DAMAGED WITH MISFIRED STAPLES. TWO STAPLERS HAD TO BE USED WHICH ADDED TIME, AND UNPLANNED ILEOSTOMY HAD TO BE CREATED TO COMPLETE THE CASE. EXTENDED HOSPITALIZATION WAS ALSO NOTED.
17-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | EEA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOWER ANTERIOR RESECTION PROCEDURE, WHILE PERFORMING ANASTAMOSIS OF ILEUM, THE FIRST STAPLER USED DID NOT MAKE A COMPLETE FIRE AND AIR LEAK TEST REVEALED LEAKING. SURGEON REPORTED NOT FEELING/HEARING THE NORMAL CRUNCH THAT THE STAPLER MAKES WHEN FIRING. THE SECOND STAPLER WOULD NOT RELEASE FROM THE TISSUE. IN ORDER TO AVOID FURTHER PATIENT INJURY, THE ANVIL WAS COMPLETELY RELEASED. PORTION OF ILEUM WAS DAMAGED WITH MISFIRED STAPLES. TWO STAPLERS HAD TO BE USED WHICH ADDED TIME, AND UNPLANNED ILEOSTOMY HAD TO BE CREATED TO COMPLETE THE CASE. EXTENDED HOSPITALIZATION WAS ALSO NOTED.
17-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | EEA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOWER ANTERIOR RESECTION PROCEDURE, WHILE PERFORMING ANASTAMOSIS OF ILEUM, THE FIRST STAPLER USED DID NOT MAKE A COMPLETE FIRE AND AIR LEAK TEST REVEALED LEAKING. SURGEON REPORTED NOT FEELING/HEARING THE NORMAL CRUNCH THAT THE STAPLER MAKES WHEN FIRING. THE SECOND STAPLER WOULD NOT RELEASE FROM THE TISSUE. IN ORDER TO AVOID FURTHER PATIENT INJURY, THE ANVIL WAS COMPLETELY RELEASED. PORTION OF ILEUM WAS DAMAGED WITH MISFIRED STAPLES. TWO STAPLERS HAD TO BE USED WHICH ADDED TIME, AND UNPLANNED ILEOSTOMY HAD TO BE CREATED TO COMPLETE THE CASE. EXTENDED HOSPITALIZATION WAS ALSO NOTED.
17-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | EEA | Gas/Air Leak | A050402 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOWER ANTERIOR RESECTION PROCEDURE, WHILE PERFORMING ANASTAMOSIS OF ILEUM, THE FIRST STAPLER USED DID NOT MAKE A COMPLETE FIRE AND AIR LEAK TEST REVEALED LEAKING. SURGEON REPORTED NOT FEELING/HEARING THE NORMAL CRUNCH THAT THE STAPLER MAKES WHEN FIRING. THE SECOND STAPLER WOULD NOT RELEASE FROM THE TISSUE. IN ORDER TO AVOID FURTHER PATIENT INJURY, THE ANVIL WAS COMPLETELY RELEASED. PORTION OF ILEUM WAS DAMAGED WITH MISFIRED STAPLES. TWO STAPLERS HAD TO BE USED WHICH ADDED TIME, AND UNPLANNED ILEOSTOMY HAD TO BE CREATED TO COMPLETE THE CASE. EXTENDED HOSPITALIZATION WAS ALSO NOTED.
17-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | EEA | Gas/Air Leak | A050402 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOWER ANTERIOR RESECTION PROCEDURE, WHILE PERFORMING ANASTAMOSIS OF ILEUM, THE FIRST STAPLER USED DID NOT MAKE A COMPLETE FIRE AND AIR LEAK TEST REVEALED LEAKING. SURGEON REPORTED NOT FEELING/HEARING THE NORMAL CRUNCH THAT THE STAPLER MAKES WHEN FIRING. THE SECOND STAPLER WOULD NOT RELEASE FROM THE TISSUE. IN ORDER TO AVOID FURTHER PATIENT INJURY, THE ANVIL WAS COMPLETELY RELEASED. PORTION OF ILEUM WAS DAMAGED WITH MISFIRED STAPLES. TWO STAPLERS HAD TO BE USED WHICH ADDED TIME, AND UNPLANNED ILEOSTOMY HAD TO BE CREATED TO COMPLETE THE CASE. EXTENDED HOSPITALIZATION WAS ALSO NOTED.
17-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | EEA | Noise, Audible | A0508 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOWER ANTERIOR RESECTION PROCEDURE, WHILE PERFORMING ANASTAMOSIS OF ILEUM, THE FIRST STAPLER USED DID NOT MAKE A COMPLETE FIRE AND AIR LEAK TEST REVEALED LEAKING. SURGEON REPORTED NOT FEELING/HEARING THE NORMAL CRUNCH THAT THE STAPLER MAKES WHEN FIRING. THE SECOND STAPLER WOULD NOT RELEASE FROM THE TISSUE. IN ORDER TO AVOID FURTHER PATIENT INJURY, THE ANVIL WAS COMPLETELY RELEASED. PORTION OF ILEUM WAS DAMAGED WITH MISFIRED STAPLES. TWO STAPLERS HAD TO BE USED WHICH ADDED TIME, AND UNPLANNED ILEOSTOMY HAD TO BE CREATED TO COMPLETE THE CASE. EXTENDED HOSPITALIZATION WAS ALSO NOTED.
17-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | EEA | Noise, Audible | A0508 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOWER ANTERIOR RESECTION PROCEDURE, WHILE PERFORMING ANASTAMOSIS OF ILEUM, THE FIRST STAPLER USED DID NOT MAKE A COMPLETE FIRE AND AIR LEAK TEST REVEALED LEAKING. SURGEON REPORTED NOT FEELING/HEARING THE NORMAL CRUNCH THAT THE STAPLER MAKES WHEN FIRING. THE SECOND STAPLER WOULD NOT RELEASE FROM THE TISSUE. IN ORDER TO AVOID FURTHER PATIENT INJURY, THE ANVIL WAS COMPLETELY RELEASED. PORTION OF ILEUM WAS DAMAGED WITH MISFIRED STAPLES. TWO STAPLERS HAD TO BE USED WHICH ADDED TIME, AND UNPLANNED ILEOSTOMY HAD TO BE CREATED TO COMPLETE THE CASE. EXTENDED HOSPITALIZATION WAS ALSO NOTED.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | TA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER SMALL BOWEL RESECTION, PATIENT WAS READMITTED A COUPLE DAYS LATER. UPON INSPECTION OF THE ORIGINAL STAPLE LINE IT WAS CLEAR THE STAPLE LINE CAME APART AND THAT FECES WERE LEAKING THROUGH THE STAPLE LINE. THE SURGEON RESECTED AND THE ANASTOMOSES AND PERFORMED AN ADDITIONAL SMALL BOWEL ANASTOMOSES USING NEW RELOADS.
13-12-2018 | Injury | us surgical puerto rico | 16-01-2019 | TA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER SMALL BOWEL RESECTION, PATIENT WAS READMITTED A COUPLE DAYS LATER. UPON INSPECTION OF THE ORIGINAL STAPLE LINE IT WAS CLEAR THE STAPLE LINE CAME APART AND THAT FECES WERE LEAKING THROUGH THE STAPLE LINE. THE SURGEON RESECTED AND THE ANASTOMOSES AND PERFORMED AN ADDITIONAL SMALL BOWEL ANASTOMOSES USING NEW RELOADS.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WERE ISSUES WITH 3 RELOADS. FOR THE FIRST REPORTED RELOAD, THE SURGEON WAS ABLE TO PUSH THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THE OTHER RELOAD HAD AN INCOMPLETE STAPLE LINE PROXIMALLY. THE LAST RELOAD HAD A POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE IN THE CENTER. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WERE ISSUES WITH 3 RELOADS. FOR THE FIRST REPORTED RELOAD, THE SURGEON WAS ABLE TO PUSH THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THE OTHER RELOAD HAD AN INCOMPLETE STAPLE LINE PROXIMALLY. THE LAST RELOAD HAD A POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE IN THE CENTER. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WERE ISSUES WITH 3 RELOADS. FOR THE FIRST REPORTED RELOAD, THE SURGEON WAS ABLE TO PUSH THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THE OTHER RELOAD HAD AN INCOMPLETE STAPLE LINE PROXIMALLY. THE LAST RELOAD HAD A POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE IN THE CENTER. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | TA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THERE WAS FAULTY IN FIRING THE DEVICE.
19-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | TA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THERE WAS FAULTY IN FIRING THE DEVICE.
25-05-2018 | Injury | us surgical puerto rico | 16-01-2019 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY,POST-OPERATIVE A LAPAROSCOPIC VENTRAL HERNIA REPAIR, PATIENTS WITH VENTRAL HERNIAS BETWEEN 2 AND 7 CM WERE RANDOMIZED INTO ONE OF THREE MESH FIXATION GROUPS: PERMANENT TACKS, ABSORBABLE TACKS AND ABSORBABLE SYNTHETIC GLUE. TWENTY FIVE PATIENT'S WERE ASSIGNED TO EACH GROUP. FOR THE PRIMARY ENDPOINT, UNSPECIFIED PAIN ON THE SECOND POSTOPERATIVE DAY WAS OBSERVED. AFTER 1 MONTH, INFECTION PRESENTING AS REDNESS OR DISCHARGE OF PUS WAS OBSERVED AT THE PORT SITE WAS OBSERVED ON FIVE PATIENTS, NINE PATIENTS WERE CLINICALLY DIAGNOSED WITH SEROMA AND TWO PATIENTS WERE DIAGNOSED WITH RECURRENT HERNIA AFTER 6 AND 12 MONTHS.
25-05-2018 | Injury | us surgical puerto rico | 16-01-2019 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY,POST-OPERATIVE A LAPAROSCOPIC VENTRAL HERNIA REPAIR, PATIENTS WITH VENTRAL HERNIAS BETWEEN 2 AND 7 CM WERE RANDOMIZED INTO ONE OF THREE MESH FIXATION GROUPS: PERMANENT TACKS, ABSORBABLE TACKS AND ABSORBABLE SYNTHETIC GLUE. TWENTY FIVE PATIENT'S WERE ASSIGNED TO EACH GROUP. FOR THE PRIMARY ENDPOINT, UNSPECIFIED PAIN ON THE SECOND POSTOPERATIVE DAY WAS OBSERVED. AFTER 1 MONTH, INFECTION PRESENTING AS REDNESS OR DISCHARGE OF PUS WAS OBSERVED AT THE PORT SITE WAS OBSERVED ON FIVE PATIENTS, NINE PATIENTS WERE CLINICALLY DIAGNOSED WITH SEROMA AND TWO PATIENTS WERE DIAGNOSED WITH RECURRENT HERNIA AFTER 6 AND 12 MONTHS.
25-05-2018 | Injury | us surgical puerto rico | 16-01-2019 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Seroma | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY,POST-OPERATIVE A LAPAROSCOPIC VENTRAL HERNIA REPAIR, PATIENTS WITH VENTRAL HERNIAS BETWEEN 2 AND 7 CM WERE RANDOMIZED INTO ONE OF THREE MESH FIXATION GROUPS: PERMANENT TACKS, ABSORBABLE TACKS AND ABSORBABLE SYNTHETIC GLUE. TWENTY FIVE PATIENT'S WERE ASSIGNED TO EACH GROUP. FOR THE PRIMARY ENDPOINT, UNSPECIFIED PAIN ON THE SECOND POSTOPERATIVE DAY WAS OBSERVED. AFTER 1 MONTH, INFECTION PRESENTING AS REDNESS OR DISCHARGE OF PUS WAS OBSERVED AT THE PORT SITE WAS OBSERVED ON FIVE PATIENTS, NINE PATIENTS WERE CLINICALLY DIAGNOSED WITH SEROMA AND TWO PATIENTS WERE DIAGNOSED WITH RECURRENT HERNIA AFTER 6 AND 12 MONTHS.
25-05-2018 | Injury | us surgical puerto rico | 16-01-2019 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Hernia | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY,POST-OPERATIVE A LAPAROSCOPIC VENTRAL HERNIA REPAIR, PATIENTS WITH VENTRAL HERNIAS BETWEEN 2 AND 7 CM WERE RANDOMIZED INTO ONE OF THREE MESH FIXATION GROUPS: PERMANENT TACKS, ABSORBABLE TACKS AND ABSORBABLE SYNTHETIC GLUE. TWENTY FIVE PATIENT'S WERE ASSIGNED TO EACH GROUP. FOR THE PRIMARY ENDPOINT, UNSPECIFIED PAIN ON THE SECOND POSTOPERATIVE DAY WAS OBSERVED. AFTER 1 MONTH, INFECTION PRESENTING AS REDNESS OR DISCHARGE OF PUS WAS OBSERVED AT THE PORT SITE WAS OBSERVED ON FIVE PATIENTS, NINE PATIENTS WERE CLINICALLY DIAGNOSED WITH SEROMA AND TWO PATIENTS WERE DIAGNOSED WITH RECURRENT HERNIA AFTER 6 AND 12 MONTHS.
08-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN EMERGENCY PNEUMONECTOMY PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN EMERGENCY PNUEMONECTOMY PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Injury | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROJEJUNOSTOMY PROCEDURE, THE STAPLES DID NOT FORM THE PROPER "B" SHAPE. THE POINT WAS SUTURED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
18-12-2018 | Injury | covidien lp llc north haven | 16-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROJEJUNOSTOMY PROCEDURE, THE STAPLES DID NOT FORM THE PROPER "B" SHAPE. THE POINT WAS SUTURED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
18-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | UNKNOWN GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT HEMICOLECTOMY, THE STAPLER WAS STICKING OUT OF THE DISTAL AREA. IT WAS REPORTED A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN TRANSECTING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT STAPLER IMMEDIATELY STOPPED WHEN THE STAPLER STARTED TO FIRE. THIS RESULTED IN AN INCOMPLETE STAPLE LINE PROXIMALLY. THEY OPENED JAWS AND GOT ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | covidien lp llc north haven | 16-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN TRANSECTING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT STAPLER IMMEDIATELY STOPPED WHEN THE STAPLER STARTED TO FIRE. THIS RESULTED IN AN INCOMPLETE STAPLE LINE PROXIMALLY. THEY OPENED JAWS AND GOT ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE SURGEON FIRED THE RELOAD WITH THE UNIVERSAL HANDLE, BUT IT STOPPED AT THE VERY BEGINNING. THE HANDLE AND RELOAD WERE BOTH CHANGED TO COMPLETE THE PROCEDURE.
18-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMICOLECTOMY, ON THE 4TH FIRING, THE SURGEON TOKE THE DEVICE OUT OF THE CAVITY, AND USED IT FOR CLOSING THE COMMON HOLE OF FUNCTIONAL END-TO-END ANASTOMOSIS (FEEA). THE SURGEON GRASPED THE TISSUE, PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE ATTEMPTING TO FIRE, BUT THE HANDLE WAS TOO STIFF AND IT WAS NOT ABLE TO BE SQUEEZED. WHEN THE SURGEON SQUEEZED THE HANDLE AGAIN WITH FORCE, A CRACK SOUND WAS HEARD. THEY REPLACED THE HANDLE AND THE CARTRIDGE, THEN IT WAS ABLE TO BE FIRED WITHOUT PROBLEM. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMICOLECTOMY, ON THE 4TH FIRING, THE SURGEON TOKE THE DEVICE OUT OF THE CAVITY, AND USED IT FOR CLOSING THE COMMON HOLE OF FUNCTIONAL END-TO-END ANASTOMOSIS (FEEA). THE SURGEON GRASPED THE TISSUE, PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE ATTEMPTING TO FIRE, BUT THE HANDLE WAS TOO STIFF AND IT WAS NOT ABLE TO BE SQUEEZED. WHEN THE SURGEON SQUEEZED THE HANDLE AGAIN WITH FORCE, A CRACK SOUND WAS HEARD. THEY REPLACED THE HANDLE AND THE CARTRIDGE, THEN IT WAS ABLE TO BE FIRED WITHOUT PROBLEM. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 16-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMICOLECTOMY, AT THE 4TH FIRING, THE SURGEON TOKE THE DEVICE OUT OF THE CAVITY, AND USED IT FOR CLOSING THE COMMON HOLE OF FUNCTIONAL END-TO-END ANASTOMOSIS (FEEA). THE SURGEON GRASPED THE TISSUE, PRESSED GREEN BUTTON AND SQUEEZED THE HANDLE ATTEMPTING TO FIRE, BUT THE HANDLE WAS TOO STIFF AND IT WAS NOT ABLE TO BE SQUEEZED. WHEN THE SURGEON SQUEEZED THE HANDLE AGAIN WITH FORCE, A CRACK SOUND WAS HEARD. THEY REPLACED THE HANDLE AND THE CARTRIDGE, THEN IT WAS ABLE TO BE FIRED WITHOUT PROBLEM. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
16-01-2019 | Malfunction | intuitive surgical | 17-01-2019 | ENDOWRIST | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: PATIENT PRESENTED TO (B)(6) FOR LEFT PARTIAL NEPHRECTOMY. DUE TO THE SIZE OF THE MASS THE SURGICAL TEAM DECIDED TO DO A FULL NEPHRECTOMY. THE 45 MM STRAIGHT DA VINCI XI STAPLER WAS USED. AFTER THE STAPLER WAS PLACED ON THE RENAL ARTERY THE STAPLER ONLY FIRED TO 24 MM AND STOPPED. IT WOULD NOT COMPLETE THE FIRING. AN ENDO GIA WAS QUICKLY OBTAINED AND USED TO COMPLETE THE DISSECTION WITHOUT BLEEDING. THE DA VINCI XI WRENCH WAS USED TO REMOVE THE STAPLER MANUALLY. AFTER CLOSE INSPECTION OF THE UNIT IT WAS NOTED THAT ONE SMALL WHITE PLASTIC PEG ON THE INSIDE OF THE STAPLER WAS BROKEN. THE STAPLE LOAD AND THE STAPLER WERE GIVEN TO THE INVENTORY SPECIALIST.
20-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA ULTRA | Device Slipped | A051204 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, WHILE ON SMALL INTESTINE, THE DEVICE DID NOT FIRE. IT WAS NOT POSSIBLE TO PRESS THE GREEN BUTTON AND WHEN THE SURGEON SQUEEZED THE HANDLE IT WAS LOOSE. THE RELOAD AND HANDLE WERE CHANGED TO COMPLETE THEPROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, WHILE ON SMALL INTESTINE, THE DEVICE DID NOT FIRE. IT WAS NOT POSSIBLE TO PRESS THE GREEN BUTTON AND WHEN THE SURGEON SQUEEZED THE HANDLE IT WAS LOOSE. THE RELOAD AND HANDLE WERE CHANGED TO COMPLETE THEPROCEDURE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | ENDOPATH | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: THE STAPLER FAILED TO STAPLE AND JAMMED. PER HOSPITAL, THE MANUFACTURER PROVIDED A RETURN GOODS AUTHORIZATION NUMBER AND PRODUCT RETURN SHIPPING CONTAINER.
17-12-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | ENDOPATH | Mechanical Jam | A0506 | No Known Impact Or Consequence To Patient | Event Description: THE STAPLER FAILED TO STAPLE AND JAMMED. PER HOSPITAL, THE MANUFACTURER PROVIDED A RETURN GOODS AUTHORIZATION NUMBER AND PRODUCT RETURN SHIPPING CONTAINER.
20-12-2018 | Injury | covidien lp llc north haven | 17-01-2019 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, WHILE ON STOMACH - POUCH, THE DOCTOR REPORTED THAT EVEN WITH THE LOAD-BEARING FILM THE CLAMP LINE BLED. THE DOCTOR NEEDED TO OVERSEW THE CLAMP LINE TO RESOLVE THE ISSUE.
11-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | UNKNOWN ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, THE DEVICE WOULD NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, THE DEVICE WOULD NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-01-2019 | Injury | davis & geck caribe | 17-01-2019 | UNKNOWN ENDO ST | Detachment of Device or Device Component | A0501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NISSEN, THE DEVICE WAS DIFFICULT TO TOGGLE AND THE NEEDLE BROKE OFF. WHILE SUTURING, THE NEEDLE DISLODGED AND FELL INTO THE PATIENT. THE DOCTOR RESECTED A PORTION OF THE STOMACH BUT COULD NOT FIND THE NEEDLE AND DECIDED NOT TO RESECT FURTHER. THE DOCTOR ALSO STATED THAT HE DID NOT BELIEVE THIS WOULD BE A PROBLEM FOR THE PATIENT AND THAT THE PATIENT HAS NOT REPORTED ANY ISSUES. AN UNANTICIPATED TISSUE LOSS WAS NOTED.
17-01-2019 | Injury | davis & geck caribe | 17-01-2019 | UNKNOWN ENDO ST | Detachment of Device or Device Component | A0501 | Device Embedded In Tissue or Plaque | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NISSEN, THE DEVICE WAS DIFFICULT TO TOGGLE AND THE NEEDLE BROKE OFF. WHILE SUTURING, THE NEEDLE DISLODGED AND FELL INTO THE PATIENT. THE DOCTOR RESECTED A PORTION OF THE STOMACH BUT COULD NOT FIND THE NEEDLE AND DECIDED NOT TO RESECT FURTHER. THE DOCTOR ALSO STATED THAT HE DID NOT BELIEVE THIS WOULD BE A PROBLEM FOR THE PATIENT AND THAT THE PATIENT HAS NOT REPORTED ANY ISSUES. AN UNANTICIPATED TISSUE LOSS WAS NOTED.
17-01-2019 | Injury | davis & geck caribe | 17-01-2019 | UNKNOWN ENDO ST | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NISSEN, THE DEVICE WAS DIFFICULT TO TOGGLE AND THE NEEDLE BROKE OFF. WHILE SUTURING, THE NEEDLE DISLODGED AND FELL INTO THE PATIENT. THE DOCTOR RESECTED A PORTION OF THE STOMACH BUT COULD NOT FIND THE NEEDLE AND DECIDED NOT TO RESECT FURTHER. THE DOCTOR ALSO STATED THAT HE DID NOT BELIEVE THIS WOULD BE A PROBLEM FOR THE PATIENT AND THAT THE PATIENT HAS NOT REPORTED ANY ISSUES. AN UNANTICIPATED TISSUE LOSS WAS NOTED.
17-01-2019 | Injury | davis & geck caribe | 17-01-2019 | UNKNOWN ENDO ST | Human-Device Interface Problem | A22 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NISSEN, THE DEVICE WAS DIFFICULT TO TOGGLE AND THE NEEDLE BROKE OFF. WHILE SUTURING, THE NEEDLE DISLODGED AND FELL INTO THE PATIENT. THE DOCTOR RESECTED A PORTION OF THE STOMACH BUT COULD NOT FIND THE NEEDLE AND DECIDED NOT TO RESECT FURTHER. THE DOCTOR ALSO STATED THAT HE DID NOT BELIEVE THIS WOULD BE A PROBLEM FOR THE PATIENT AND THAT THE PATIENT HAS NOT REPORTED ANY ISSUES. AN UNANTICIPATED TISSUE LOSS WAS NOTED.
17-01-2019 | Injury | davis & geck caribe | 17-01-2019 | UNKNOWN ENDO ST | Human-Device Interface Problem | A22 | Device Embedded In Tissue or Plaque | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NISSEN, THE DEVICE WAS DIFFICULT TO TOGGLE AND THE NEEDLE BROKE OFF. WHILE SUTURING, THE NEEDLE DISLODGED AND FELL INTO THE PATIENT. THE DOCTOR RESECTED A PORTION OF THE STOMACH BUT COULD NOT FIND THE NEEDLE AND DECIDED NOT TO RESECT FURTHER. THE DOCTOR ALSO STATED THAT HE DID NOT BELIEVE THIS WOULD BE A PROBLEM FOR THE PATIENT AND THAT THE PATIENT HAS NOT REPORTED ANY ISSUES. AN UNANTICIPATED TISSUE LOSS WAS NOTED.
17-01-2019 | Injury | davis & geck caribe | 17-01-2019 | UNKNOWN ENDO ST | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NISSEN, THE DEVICE WAS DIFFICULT TO TOGGLE AND THE NEEDLE BROKE OFF. WHILE SUTURING, THE NEEDLE DISLODGED AND FELL INTO THE PATIENT. THE DOCTOR RESECTED A PORTION OF THE STOMACH BUT COULD NOT FIND THE NEEDLE AND DECIDED NOT TO RESECT FURTHER. THE DOCTOR ALSO STATED THAT HE DID NOT BELIEVE THIS WOULD BE A PROBLEM FOR THE PATIENT AND THAT THE PATIENT HAS NOT REPORTED ANY ISSUES. AN UNANTICIPATED TISSUE LOSS WAS NOTED.
17-01-2019 | Injury | us surgical puerto rico | 17-01-2019 | EEA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC COLECTOMY, THE ANASTOMOSIS COMPLETELY OPENED. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THE PROBLEM. A COLOSTOMY WAS PERFORMED AS A SURGICAL INTERVENTION. A NEW AND ADDITIONAL OPERATION WAS DONE TO RESOLVE THE ISSUE. THE PATIENT WAS STILL HOSPITALIZED. THERE WAS TEMPORARY PATIENT INJURY.
17-01-2019 | Injury | us surgical puerto rico | 17-01-2019 | EEA | Loss of or Failure to Bond | A040102 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC COLECTOMY, THE ANASTOMOSIS COMPLETELY OPENED. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THE PROBLEM. A COLOSTOMY WAS PERFORMED AS A SURGICAL INTERVENTION. A NEW AND ADDITIONAL OPERATION WAS DONE TO RESOLVE THE ISSUE. THE PATIENT WAS STILL HOSPITALIZED. THERE WAS TEMPORARY PATIENT INJURY.
17-01-2019 | Injury | us surgical puerto rico | 17-01-2019 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC COLECTOMY, THE ANASTOMOSIS COMPLETELY OPENED. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THE PROBLEM. A COLOSTOMY WAS PERFORMED AS A SURGICAL INTERVENTION. A NEW AND ADDITIONAL OPERATION WAS DONE TO RESOLVE THE ISSUE. THE PATIENT WAS STILL HOSPITALIZED. THERE WAS TEMPORARY PATIENT INJURY.
06-12-2018 | Injury | us surgical puerto rico | 17-01-2019 | EEA | Loss of or Failure to Bond | A040102 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC COLECTOMY, DURING ANASTOMOSIS, THE ANASTOMOSED SITE COMPLETELY OPENED. THERE WAS BLOOD LOSS OF MORE THAN 500CC WHICH REQUIRED A BLOOD TRANSFUSION. THE PRODUCT PROBLEM RESULTED TO COLOSTOMY. "K-OVAIO" WAS DONE TO PREVENT PERMANENT IMPAIRMENT. THE PATIENT'S STAY IN THE HOSPITAL HAS BEEN EXTENDED. CURRENTLY, THE PATIENT IS STILL AT THE HOSPITAL. IT WAS STATED THAT THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
06-12-2018 | Injury | us surgical puerto rico | 17-01-2019 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC COLECTOMY, DURING ANASTOMOSIS, THE ANASTOMOSED SITE COMPLETELY OPENED. THERE WAS BLOOD LOSS OF MORE THAN 500CC WHICH REQUIRED A BLOOD TRANSFUSION. THE PRODUCT PROBLEM RESULTED TO COLOSTOMY. "K-OVAIO" WAS DONE TO PREVENT PERMANENT IMPAIRMENT. THE PATIENT'S STAY IN THE HOSPITAL HAS BEEN EXTENDED. CURRENTLY, THE PATIENT IS STILL AT THE HOSPITAL. IT WAS STATED THAT THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
18-12-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LOBECTOMY, DURING THE TRANSECTION OF THE PULMONARY ARTERY, AT THE FIRST FIRING, THE KNIFE STOPPED ADVANCING. THE DOCTOR GOT THE ANVIL TO THE BACK SIDE AND CLOSED THE JAWS. THE DEVICE WAS FIRED AFTER CONFIRMING THAT THE GREEN BUTTON WAS FLASHING. AT THE BEGINNING OF THE FIRING, A TONE OF FIRING WAS HEARD, BUT THE KNIFE STOPPED ADVANCING IMMEDIATELY AND WOULD NOT MOVE EVEN THOUGH THE DOCTOR PRESSED THE FIRE BUTTON AGAIN. THE SURGICAL TIME WAS EXTENDED FOR NOT MORE THAN 30 MINUTES. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LOBECTOMY, DURING THE TRANSECTION OF THE PULMONARY ARTERY, AT THE FIRST FIRING, THE KNIFE STOPPED ADVANCING. THE DOCTOR GOT THE ANVIL TO THE BACK SIDE AND CLOSED THE JAWS. THE DEVICE WAS FIRED AFTER CONFIRMING THAT THE GREEN BUTTON WAS FLASHING. AT THE BEGINNING OF THE FIRING, A TONE OF FIRING WAS HEARD, BUT THE KNIFE STOPPED ADVANCING IMMEDIATELY AND WOULD NOT MOVE EVEN THOUGH THE DOCTOR PRESSED THE FIRE BUTTON AGAIN. THE SURGICAL TIME WAS EXTENDED FOR NOT MORE THAN 30 MINUTES. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LOBECTOMY, DURING THE TRANSECTION OF THE PULMONARY ARTERY, AT THE FIRST FIRING, THE KNIFE STOPPED ADVANCING. THE DOCTOR GOT THE ANVIL TO THE BACK SIDE AND CLOSED THE JAWS. THE DEVICE WAS FIRED AFTER CONFIRMING THAT THE GREEN BUTTON WAS FLASHING. AT THE BEGINNING OF THE FIRING, A TONE OF FIRING WAS HEARD, BUT THE KNIFE STOPPED ADVANCING IMMEDIATELY AND WOULD NOT MOVE EVEN THOUGH THE DOCTOR PRESSED THE FIRE BUTTON AGAIN. THE SURGICAL TIME WAS EXTENDED FOR NOT MORE THAN 30 MINUTES. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/LOBECTOMY PROCEDURE, BLEEDING OCCURRED FREQUENTLY. THE BLEEDING WAS MORE THAN OOZING BLEEDING, LESS THAN 200 CC, IN A SHORT PERIOD. THE ISSUE WAS SOLVED BY ASTRICTION. TENSION WAS NOT APPLIED ON THE BLOOD VESSEL. THE STATUS OF THE PATIENT WAS REPORTED AS NO PROBLEM.
19-12-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, BEFORE ENTERING THE ABDOMINAL CAVITY DURING CYCLE OF THE FIRST RELOAD, THE STAPLER COULD NOT BE OPENED AND HAD A VISIBLE STAPLES. THE LOAD WAS EXCHANGED FOR A NEW ONE, IT WAS POSSIBLE TO CLOSE AND OPEN THE JAWS, HOWEVER, IT WAS IMPOSSIBLE TO MAKE ARTICULATION. STAPLER AND LOADS WERE PUT AWAY AND USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THE PACKAGING WAS NOT DAMAGED, THE EQUIPMENT WAS NEW. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, BEFORE ENTERING THE ABDOMINAL CAVITY DURING CYCLE OF THE FIRST RELOAD, THE STAPLER COULD NOT BE OPENED AND HAD A VISIBLE STAPLES. THE LOAD WAS EXCHANGED FOR A NEW ONE, IT WAS POSSIBLE TO CLOSE AND OPEN THE JAWS, HOWEVER, IT WAS IMPOSSIBLE TO MAKE ARTICULATION. STAPLER AND LOADS WERE PUT AWAY AND USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THE PACKAGING WAS NOT DAMAGED, THE EQUIPMENT WAS NEW. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA | Positioning Problem | A1502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, BEFORE ENTERING THE ABDOMINAL CAVITY DURING CYCLE OF THE FIRST RELOAD, THE STAPLER COULD NOT BE OPENED AND HAD A VISIBLE STAPLES. THE LOAD WAS EXCHANGED FOR A NEW ONE, IT WAS POSSIBLE TO CLOSE AND OPEN THE JAWS, HOWEVER, IT WAS IMPOSSIBLE TO MAKE ARTICULATION. STAPLER AND LOADS WERE PUT AWAY AND USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THE PACKAGING WAS NOT DAMAGED, THE EQUIPMENT WAS NEW. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, THE CLAMP WITH THE REFILLS DOES NOT CUT OR STAPLE. THEY CHANGED RELOAD WITHOUT SUCCESS AND AFTER CHANGING PLIERS EVERYTHING WENT BACK TO WORKING. THERE WAS NO PATIENT INJURY.
14-11-2018 | Injury | ethicon endo-surgery | 17-01-2019 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | Not Applicable | Event Description: IT WAS REPORTED THAT WHEN PERFORMING AN ANASTOMOSIS FOR A PATIENT TREATED REGARDING FOR A COLORECTAL CANCER, THE DEVICE GOT JAMMED WITH STAPLING ISSUE. THE ISSUE HAS BEEN NOTICED DURING USING. THERE HAVE BEEN LEAKS OF FECAL SAMPLES WHICH LEADING TO PERITONEAL LAVAGE. THEY HAVE CHANGED THE DEVICE AND THE CARTRIDGE TO FINISH THE PROCEDURE.
14-11-2018 | Injury | ethicon endo-surgery | 17-01-2019 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | No Code Available | Event Description: IT WAS REPORTED THAT WHEN PERFORMING AN ANASTOMOSIS FOR A PATIENT TREATED REGARDING FOR A COLORECTAL CANCER, THE DEVICE GOT JAMMED WITH STAPLING ISSUE. THE ISSUE HAS BEEN NOTICED DURING USING. THERE HAVE BEEN LEAKS OF FECAL SAMPLES WHICH LEADING TO PERITONEAL LAVAGE. THEY HAVE CHANGED THE DEVICE AND THE CARTRIDGE TO FINISH THE PROCEDURE.
18-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, THE CLAMP WITH THE REFILLS DOES NOT CUT OR STAPLE. THEY CHANGED RELOAD WITHOUT SUCCESS AND AFTER CHANGING PLIERS EVERYTHING WENT BACK TO WORKING. THERE WAS NO PATIENT INJURY.
17-12-2018 | Injury | us surgical puerto rico | 17-01-2019 | EEA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, A FEW HOURS FOLLOWING A RECTUM RESECTION PROCEDURE, THE ANASTOMOSIS LEAKED. A COLOSTOMY WAS MADE A DAY LATER TO RESOLVE THE ISSUE. PATIENT WAS PLACED IN INTENSIVE CARE.
17-12-2018 | Injury | us surgical puerto rico | 17-01-2019 | EEA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, A FEW HOURS FOLLOWING A RECTUM RESECTION PROCEDURE, THE ANASTOMOSIS LEAKED. A COLOSTOMY WAS MADE A DAY LATER TO RESOLVE THE ISSUE. PATIENT WAS PLACED IN INTENSIVE CARE.
27-12-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT FIRE. IT WAS REPORTED THAT DURING THE LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, COULD NOT FIRE. THE TISSUE WAS CUT OFF, BUT THE STAPLES MALFORMED AND WASHER WAS NOT CUT OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
27-12-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT FIRE. IT WAS REPORTED THAT DURING THE LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, COULD NOT FIRE. THE TISSUE WAS CUT OFF, BUT THE STAPLES MALFORMED AND WASHER WAS NOT CUT OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
27-12-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT FIRE. IT WAS REPORTED THAT DURING THE LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, COULD NOT FIRE. THE TISSUE WAS CUT OFF, BUT THE STAPLES MALFORMED AND WASHER WAS NOT CUT OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
27-12-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT CUT WASHER. DURING THE RADICAL RESECTION OF LOW RECTAL CANCER, AFTER FIRED (CUSTOMER CONFIRMED HAS SQUEEZED DOWN THE FIRING TRIGGER, PTOP - PLASTIC TO PLASTIC), BUT WITHOUT AUDIBLE FEEDBACK, 60% STAPLES MALFORMED AND WASHER WAS NOT CUT OFF. REMOVED ALL THE MALFORMED STAPLES, AND USED SUTURE TO OVERSEW. THE SURGERY DELAYED 2 HOURS. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
27-12-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: WOULD NOT CUT WASHER. DURING THE RADICAL RESECTION OF LOW RECTAL CANCER, AFTER FIRED (CUSTOMER CONFIRMED HAS SQUEEZED DOWN THE FIRING TRIGGER, PTOP - PLASTIC TO PLASTIC), BUT WITHOUT AUDIBLE FEEDBACK, 60% STAPLES MALFORMED AND WASHER WAS NOT CUT OFF. REMOVED ALL THE MALFORMED STAPLES, AND USED SUTURE TO OVERSEW. THE SURGERY DELAYED 2 HOURS. THE PATIENT IS STABLE AND IN THE HOSPITAL NOW.
20-12-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, WHILE THE DEVICE WAS BEING APPLIED IN THE COLON TUMOR, THE DEVICE WAS NOT ABLE TO FIRE. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-10-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | RELOAD- CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, STAPLES DIDN¿T FIRE FULLY/PROPERLY WHEN THE CONTOUR WAS APPLIED ON THE RECTUM. WHICH IN TURN LEAD TO USE HAND SUTURING TO CLOSE THE SEMI-FIRED STAPLERS AT THE RECTAL STUMP. PROCEDURE PROLONGED 1.5 HOURS,
20-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, WHILE THE DEVICE WAS BEING APPLIED IN THE COLON TUMOR, THE DEVICE WAS NOT ABLE TO FIRE. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, WHEN THE CARTRIDGE WAS CONNECTED PRIOR A LOBECTOMY PROCEDURE, THE DEVICE WAS STIFF AND THE JAWS DID NOT CLOSE WHEN OPENING AND CLOSING WAS CHECKED. A NEW HANDLE WAS OPENED WITH THE SAME CARTRIDGE AND THE PROCEDURE WAS CONTINUED. THERE WAS NO PATIENT INVOLVEMENT.
20-12-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | UNKNOWN EGIA SU | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, WHEN THE CARTRIDGE WAS CONNECTED PRIOR A LOBECTOMY PROCEDURE, THE DEVICE WAS STIFF AND THE JAWS DID NOT CLOSE WHEN OPENING AND CLOSING WAS CHECKED. A NEW HANDLE WAS OPENED WITH THE SAME CARTRIDGE AND THE PROCEDURE WAS CONTINUED. THERE WAS NO PATIENT INVOLVEMENT.
31-07-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. THEY REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-07-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. THEY REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-07-2018 | Malfunction | covidien lp llc north haven | 17-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. THEY REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-07-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. THEY REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-11-2018 | Injury | us surgical puerto rico | 17-01-2019 | EEA | Loss of or Failure to Bond | A040102 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, WHILE PERFORMING AN END-TO-END ANASTOMOSIS, THE DEVICE IS DIFFICULT OR WAS UNABLE TO CLOSE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE DEVICE PROBLEM. THERE WAS ALSO AN INCISION EXTENSION OF 2 INCH DUE TO THE PRODUCT PROBLEM. AN UNANTICIPATED TISSUE LOSS OF AROUND 2 INCHES OF RECTUM OCCURRED DUE TO THE DEVICE PROBLEM.
29-11-2018 | Injury | us surgical puerto rico | 17-01-2019 | EEA | Positioning Problem | A1502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, WHILE PERFORMING AN END-TO-END ANASTOMOSIS, THE DEVICE IS DIFFICULT OR WAS UNABLE TO CLOSE. THEY USED ANOTHER DEVICE TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE DEVICE PROBLEM. THERE WAS ALSO AN INCISION EXTENSION OF 2 INCH DUE TO THE PRODUCT PROBLEM. AN UNANTICIPATED TISSUE LOSS OF AROUND 2 INCHES OF RECTUM OCCURRED DUE TO THE DEVICE PROBLEM.
12-10-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, GOING ACROSS MESENTERY, STAPLER WOULD CLOSE BUT NOT FIRE WITH VASCULAR LOAD. 10 MIN DELAY GETTING NEW STAPLERS AND RELOADS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
18-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE RESECTION, THE ARTICULATION BUTTON WAS BROKEN DOWN. IT WAS REPORTED THE SURGEON USED A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
12-10-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, GOING ACROSS MESENTERY, STAPLER WOULD CLOSE BUT NOT FIRE WITH VASCULAR LOAD. 10 MIN DELAY GETTING NEW STAPLERS AND RELOADS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
14-12-2018 | Malfunction | us surgical puerto rico | 17-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE.
19-11-2018 | Malfunction | ethicon endo-surgery | 17-01-2019 | GST WHITE RELOAD, 45MM, 6 ROW | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS FOUND THAT THE PROXIMAL END OF THE CARTRIDGE PAN HAD BEEN COME OFF FROM THE CARTRIDGE BEFORE USE. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
17-01-2019 | Injury | davol inc sub c.r bard | 17-01-2019 | SORBAFIX ENHANCED | Patient Device Interaction Problem | A01 | Itching Sensation | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE PATIENT UNDERWENT A LAPAROSCOPIC REPAIR OF A VENTRAL HERNIA WITH IMPLANT A BARD VENTRALIGHT ST MESH. THE MESH WAS PRE-SUTURED FOR PLACEMENT PRIOR TO MECHANICAL FIXATION PERFORMED WITH A BARD SORBAFIX DEVICE. AS REPORTED SIX DAYS POSTOPERATIVELY, THE PATIENT DEVELOPED A RASH ON HER ABDOMEN. THE PATIENT PRESENTED FOR TWO FOLLOW UP VISITS WITH THE SURGEON, WHO DID NOT OFFER A DIAGNOSIS OR TREATMENT FOR THE RASH. THE PATIENT PRESENTED TO A DERMATOLOGIST TO ASSESS THE RASH. AS REPORTED, THE RASH SPREAD FROM HEAD TO TOE AND THE PATIENT HAS SEVERE ITCHING WITH THE RASH OUTBREAKS. THE PATIENT HAS UNDERGONE FOUR ROUNDS OF TREATMENT WITH ORAL PREDNISONE AND FIVE CORTISONE SHOTS. THE RASH CONTINUES TO RE-EMERGE ONCE THE STEROIDS WEAR OFF. THE PATIENT HAS UNDERGONE MULTIPLE ALLERGY TESTS AND HAS ONLY BEEN FOUND ALLERGIC TO COBALT BLUE AND SYNTHETIC COBALT BLUE. THE PATIENT UNDERWENT HER ANNUAL CT SCAN FOR ONGOING CANCER PREVENTION (HISTORY OF CANCER X2) AND AT THIS TIME REPORTS THERE IS FLUID BEHIND THE MESH. AS REPORT THE PATIENT'S MD ADVISED THAT THIS WILL RESOLVE OVER TIME AND NO TREATMENT WILL BE PROVIDED AT THIS TIME. THE PATIENT REMAINS CANCER FREE AND HER LAST CHEMOTHERAPY TREATMENT ENDED 1 YR. PRIOR TO THE HERNIA REPAIR.
17-01-2019 | Injury | davol inc sub c.r bard | 17-01-2019 | SORBAFIX ENHANCED | Patient Device Interaction Problem | A01 | Rash | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE PATIENT UNDERWENT A LAPAROSCOPIC REPAIR OF A VENTRAL HERNIA WITH IMPLANT A BARD VENTRALIGHT ST MESH. THE MESH WAS PRE-SUTURED FOR PLACEMENT PRIOR TO MECHANICAL FIXATION PERFORMED WITH A BARD SORBAFIX DEVICE. AS REPORTED SIX DAYS POSTOPERATIVELY, THE PATIENT DEVELOPED A RASH ON HER ABDOMEN. THE PATIENT PRESENTED FOR TWO FOLLOW UP VISITS WITH THE SURGEON, WHO DID NOT OFFER A DIAGNOSIS OR TREATMENT FOR THE RASH. THE PATIENT PRESENTED TO A DERMATOLOGIST TO ASSESS THE RASH. AS REPORTED, THE RASH SPREAD FROM HEAD TO TOE AND THE PATIENT HAS SEVERE ITCHING WITH THE RASH OUTBREAKS. THE PATIENT HAS UNDERGONE FOUR ROUNDS OF TREATMENT WITH ORAL PREDNISONE AND FIVE CORTISONE SHOTS. THE RASH CONTINUES TO RE-EMERGE ONCE THE STEROIDS WEAR OFF. THE PATIENT HAS UNDERGONE MULTIPLE ALLERGY TESTS AND HAS ONLY BEEN FOUND ALLERGIC TO COBALT BLUE AND SYNTHETIC COBALT BLUE. THE PATIENT UNDERWENT HER ANNUAL CT SCAN FOR ONGOING CANCER PREVENTION (HISTORY OF CANCER X2) AND AT THIS TIME REPORTS THERE IS FLUID BEHIND THE MESH. AS REPORT THE PATIENT'S MD ADVISED THAT THIS WILL RESOLVE OVER TIME AND NO TREATMENT WILL BE PROVIDED AT THIS TIME. THE PATIENT REMAINS CANCER FREE AND HER LAST CHEMOTHERAPY TREATMENT ENDED 1 YR. PRIOR TO THE HERNIA REPAIR.
17-01-2019 | Injury | davol inc sub c.r bard | 17-01-2019 | SORBAFIX ENHANCED | Patient Device Interaction Problem | A01 | Reaction | Event Description: IT WAS REPORTED THAT ON (B)(6) 2018 THE PATIENT UNDERWENT A LAPAROSCOPIC REPAIR OF A VENTRAL HERNIA WITH IMPLANT A BARD VENTRALIGHT ST MESH. THE MESH WAS PRE-SUTURED FOR PLACEMENT PRIOR TO MECHANICAL FIXATION PERFORMED WITH A BARD SORBAFIX DEVICE. AS REPORTED SIX DAYS POSTOPERATIVELY, THE PATIENT DEVELOPED A RASH ON HER ABDOMEN. THE PATIENT PRESENTED FOR TWO FOLLOW UP VISITS WITH THE SURGEON, WHO DID NOT OFFER A DIAGNOSIS OR TREATMENT FOR THE RASH. THE PATIENT PRESENTED TO A DERMATOLOGIST TO ASSESS THE RASH. AS REPORTED, THE RASH SPREAD FROM HEAD TO TOE AND THE PATIENT HAS SEVERE ITCHING WITH THE RASH OUTBREAKS. THE PATIENT HAS UNDERGONE FOUR ROUNDS OF TREATMENT WITH ORAL PREDNISONE AND FIVE CORTISONE SHOTS. THE RASH CONTINUES TO RE-EMERGE ONCE THE STEROIDS WEAR OFF. THE PATIENT HAS UNDERGONE MULTIPLE ALLERGY TESTS AND HAS ONLY BEEN FOUND ALLERGIC TO COBALT BLUE AND SYNTHETIC COBALT BLUE. THE PATIENT UNDERWENT HER ANNUAL CT SCAN FOR ONGOING CANCER PREVENTION (HISTORY OF CANCER X2) AND AT THIS TIME REPORTS THERE IS FLUID BEHIND THE MESH. AS REPORT THE PATIENT'S MD ADVISED THAT THIS WILL RESOLVE OVER TIME AND NO TREATMENT WILL BE PROVIDED AT THIS TIME. THE PATIENT REMAINS CANCER FREE AND HER LAST CHEMOTHERAPY TREATMENT ENDED 1 YR. PRIOR TO THE HERNIA REPAIR.
14-12-2018 | Malfunction | ethicon endo-surgery | 18-01-2019 | ECHELON FLEX | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: DURING THE PROCEDURE AN ETHICON POWERED PLUS LONG ARTICULATING ENDOSCOPIC LINEAR CUTTER WAS BEING USED. AFTER MULTIPLE SUCCESSFUL FIRINGS, A CARTRIDGE WAS LOADED AND MISFIRED. IT WAS REMOVED, ANOTHER CARTRIDGE WAS PLACED, AND IT MISFIRED. THE CUTTER WAS REMOVED FROM THE FIELD AND A NEW CUTTER WAS GIVEN TO THE SCRUB TECH SO THE CASE COULD BE COMPLETED. NO PATIENT HARM.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE WILL NOT PULLED WHEN FIRED. THERE WAS NO PATIENT INJURY.
18-01-2019 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Unintended System Motion | A051206 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY ON THE STOMACH, A REINFORCEMENT MATERIAL WAS USED IN CONJUNCTION WITH THE STAPLING DEVICE, THE STAPLER DID NOT FIRE AND THE RELOAD KEPT STRAIGHTENING BY ITSELF. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-01-2019 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY ON THE STOMACH, A REINFORCEMENT MATERIAL WAS USED IN CONJUNCTION WITH THE STAPLING DEVICE, THE STAPLER DID NOT FIRE AND THE RELOAD KEPT STRAIGHTENING BY ITSELF. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-01-2019 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Unintended System Motion | A051206 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY ON THE STOMACH, A REINFORCEMENT MATERIAL WAS USED IN CONJUNCTION WITH THE STAPLING DEVICE, THE STAPLER DID NOT FIRE AND THE RELOAD KEPT STRAIGHTENING BY ITSELF. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-01-2019 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY ON THE STOMACH, A REINFORCEMENT MATERIAL WAS USED IN CONJUNCTION WITH THE STAPLING DEVICE, THE STAPLER DID NOT FIRE AND THE RELOAD KEPT STRAIGHTENING BY ITSELF. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, THE STAPLING STOPPED AT 1/5 POINT.
21-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, THE STAPLING STOPPED AT 1/5 POINT.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE WILL NOT PULLED WHEN FIRED. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON COULD NOT ARTICULATE THE RELOAD WITH THE POWERED STAPLER. THEY FIGURED OUT THAT THE RELOAD HAD SOMETHING WRONG AND ORDERED THE NURSE TO CHANGE IT WITH NEW ONE. WHEN THE OPERATING ROOM NURSE TRIED TO TAKE THE RELOAD OUT, THE CONNECTION PART BROKE OUT. THEY PREPARED A NEW RELOAD AND THE SURGEON USED IT.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Positioning Problem | A1502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON COULD NOT ARTICULATE THE RELOAD WITH THE POWERED STAPLER. THEY FIGURED OUT THAT THE RELOAD HAD SOMETHING WRONG AND ORDERED THE NURSE TO CHANGE IT WITH NEW ONE. WHEN THE OPERATING ROOM NURSE TRIED TO TAKE THE RELOAD OUT, THE CONNECTION PART BROKE OUT. THEY PREPARED A NEW RELOAD AND THE SURGEON USED IT.
21-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE WILL NOT PULLED WHEN FIRED. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY ON THE STOMACH, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. A REPLACEMENT WAS GIVEN TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY ON THE STOMACH, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. A REPLACEMENT WAS GIVEN TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY ON THE STOMACH, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. A REPLACEMENT WAS GIVEN TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRIOR A LAPAROSCOPIC INGUINAL HERNIA PROCEDURE, THERE WERE DIFFICULTIES WITH LOADING THE RELOAD ONTO THE HANDLE, DIFFICULTY IN PRESSING THE "PUSH TO RELOAD" BUTTON, AND DIFFICULTY WITH NAVIGATING THE LOCK AND UNLOCK BUTTON. AFTER THE LOADING ISSUES WERE EXPERIENCED, THE RELOAD WAS NOT USED BY THE PATIENT. THERE WAS NO PATIENT INVOLVEMENT.
28-05-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MAJOR GASTRIC RESECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. REPLACEMENT OF SURGICAL INSTRUMENTS WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-05-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MAJOR GASTRIC RESECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. REPLACEMENT OF SURGICAL INSTRUMENTS WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-05-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MAJOR GASTRIC RESECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. REPLACEMENT OF SURGICAL INSTRUMENTS WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-05-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MAJOR GASTRIC RESECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. REPLACEMENT OF SURGICAL INSTRUMENTS WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-05-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MAJOR GASTRIC RESECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. REPLACEMENT OF SURGICAL INSTRUMENTS WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-11-2018 | Injury | us surgical puerto rico | 18-01-2019 | EEA | Loss of or Failure to Bond | A040102 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY AFTER A LAPAROSCOPIC COLOSTOMY PROCEDURE, THERE WAS A COMPLETE OPENING OF THE ANASTOMOSIS. MEDICAL OR SURGICAL INTERVENTION WAS NEEDED TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION. THE EVENT LEADS TO OR EXTEND PATIENT HOSPITALIZATION. THERE WAS TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT OF THE PROBLEM. THERE WERE PERITONITIS AND THE CREATION OF TEMPORARY/DEFINITIVE COLOSTOMY TO BE DECIDED. THE INCISION WAS EXTENDED MORE THAN 1 INCH DUE TO THE PRODUCT PROBLEM. THE PATIENT WAS STILL HOSPITALIZED AT THE TIME OF THE REPORT.
30-11-2018 | Injury | us surgical puerto rico | 18-01-2019 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY AFTER A LAPAROSCOPIC COLOSTOMY PROCEDURE, THERE WAS A COMPLETE OPENING OF THE ANASTOMOSIS. MEDICAL OR SURGICAL INTERVENTION WAS NEEDED TO PREVENT A PERMANENT IMPAIRMENT OF A FUNCTION. THE EVENT LEADS TO OR EXTEND PATIENT HOSPITALIZATION. THERE WAS TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT OF THE PROBLEM. THERE WERE PERITONITIS AND THE CREATION OF TEMPORARY/DEFINITIVE COLOSTOMY TO BE DECIDED. THE INCISION WAS EXTENDED MORE THAN 1 INCH DUE TO THE PRODUCT PROBLEM. THE PATIENT WAS STILL HOSPITALIZED AT THE TIME OF THE REPORT.
26-11-2018 | Malfunction | ethicon endo-surgery | 18-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY THE STAPLER WOULDN'T FIRE. A SECOND LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
02-07-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGEAL CANCER SURGERY, THE SURGEON WAS NOT ABLE TO PRESSED THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED AND THE STAPLER DID NOT FIRE. IT WAS NOTED THEY REPLACED THE CARTRIDGE WITH A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
02-07-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGEAL CANCER SURGERY, THE SURGEON WAS NOT ABLE TO PRESSED THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED AND THE STAPLER DID NOT FIRE. IT WAS NOTED THEY REPLACED THE CARTRIDGE WITH A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IT WAS UNABLE TO FIRE DEVICE. THE RELOAD WAS NOT FIRED IN APPEARANCE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
19-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC PROSTATECTOMY PROCEDURE, DURING STAPLING THE DORSAL VENOUS, THERE WAS A POOR STAPLE LINE FORMATION AND THE STAPLE LINE IS INCOMPLETE ON THE PROXIMAL PART. THE DOCTOR ALSO MENTIONED THE DEVICE WAS VERY HARD TO OPEN AFTER BEING FIRED WHEREIN THE EXCESS FORCE WAS NEEDED. A NEW UNIT WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC PROSTATECTOMY PROCEDURE, DURING STAPLING THE DORSAL VENOUS, THERE WAS A POOR STAPLE LINE FORMATION AND THE STAPLE LINE IS INCOMPLETE ON THE PROXIMAL PART. THE DOCTOR ALSO MENTIONED THE DEVICE WAS VERY HARD TO OPEN AFTER BEING FIRED WHEREIN THE EXCESS FORCE WAS NEEDED. A NEW UNIT WAS USED TO COMPLETE THE PROCEDURE. THERE IS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. IT WAS REPORTED THAT IT WAS IMPOSSIBLE TO CLOSE THE INSTRUMENT.
18-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. IT WAS REPORTED THAT IT WAS IMPOSSIBLE TO CLOSE THE INSTRUMENT.
21-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNK-SIGADAPT | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, THE STAPLING STOPPED AT 1/5 POINT.
19-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | MULTIFIRE ENDO TA | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE JAWS OF THE DEVICE COULD NOT BE OPENED. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | GIA | Detachment of Device or Device Component | A0501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN OPENING THE PRODUCT, THE LITTLE BLACK PART OF THE DEVICE WAS BROKEN AND WAS DISENGAGED. THERE WAS NO PATIENT INVOLVEMENT.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | IDRIVE ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IT WAS UNABLE TO FIRE DEVICE. THE RELOAD WAS NOT FIRED IN APPEARANCE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
11-08-2017 | Malfunction | ethicon endo-surgery | 18-01-2019 | ILS 33MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLING FAILURE. THERE WAS IMMEDIATE OPENING OF THE ANASTOMOSIS SHORTLY AFTER FIRING.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IT WAS UNABLE TO FIRE DEVICE. THE RELOAD WAS NOT FIRED IN APPEARANCE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
14-06-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGEAL CANCER PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY REPLACED THE OLD CARTRIDGE WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-06-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGEAL CANCER PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY REPLACED THE OLD CARTRIDGE WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-06-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGEAL CANCER PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY REPLACED THE OLD CARTRIDGE WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-09-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LUNG RESECTION FOR REMOVAL OF NODULE, THE HANDLE WAS MAKING A CRACKING SOUND DURING THE FIRING SEQUENCE IN THE THICK TISSUE, THE SCRUB TECH WAS UNABLE TO UNLOAD THE STAPLER AFTER FIRING. ANOTHER HANDLE WAS OPENED AND USED AND THAT HANDLE CRACKED AS WELL. IN BOTH FIRINGS, THE STAPLER FIRED COMPLETELY. UPON VISUAL INSPECTION, HE COULD SEE THE BLACK RETRACTION KNOBS WERE NOT FULLY RETRACTED ON THE HANDLE THAT COULD NOT BE UNLOADED. HE COMPLETED THE RETRACTION AND THE RELOAD DISENGAGED FROM THE STAPLER PROPERLY. THERE WAS NO PATIENT HARM.
10-09-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LUNG RESECTION FOR REMOVAL OF NODULE, THE HANDLE WAS MAKING A CRACKING SOUND DURING THE FIRING SEQUENCE IN THE THICK TISSUE, THE SCRUB TECH WAS UNABLE TO UNLOAD THE STAPLER AFTER FIRING. ANOTHER HANDLE WAS OPENED AND USED AND THAT HANDLE CRACKED AS WELL. IN BOTH FIRINGS, THE STAPLER FIRED COMPLETELY. UPON VISUAL INSPECTION, HE COULD SEE THE BLACK RETRACTION KNOBS WERE NOT FULLY RETRACTED ON THE HANDLE THAT COULD NOT BE UNLOADED. HE COMPLETED THE RETRACTION AND THE RELOAD DISENGAGED FROM THE STAPLER PROPERLY. THERE WAS NO PATIENT HARM.
31-05-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN ENDO GIA SULU | Component or Accessory Incompatibility | A1701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THERE WAS A PROBLEM IN LOADING THE DEVICE AND IT COULD NOT BE COMPLETELY CLOSED. REPLACEMENT OF THE SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
31-05-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN ENDO GIA SULU | Component or Accessory Incompatibility | A1701 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THERE WAS A PROBLEM IN LOADING THE DEVICE AND IT COULD NOT BE COMPLETELY CLOSED. REPLACEMENT OF THE SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
31-05-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THERE WAS A PROBLEM IN LOADING THE DEVICE AND IT COULD NOT BE COMPLETELY CLOSED. REPLACEMENT OF THE SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
31-05-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THERE WAS A PROBLEM IN LOADING THE DEVICE AND IT COULD NOT BE COMPLETELY CLOSED. REPLACEMENT OF THE SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
31-05-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THERE WAS A PROBLEM IN LOADING THE DEVICE AND IT COULD NOT BE COMPLETELY CLOSED. REPLACEMENT OF THE SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
31-05-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THERE WAS A PROBLEM IN LOADING THE DEVICE AND IT COULD NOT BE COMPLETELY CLOSED. REPLACEMENT OF THE SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
31-05-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THERE WAS A PROBLEM IN LOADING THE DEVICE AND IT COULD NOT BE COMPLETELY CLOSED. REPLACEMENT OF THE SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
31-05-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THERE WAS A PROBLEM IN LOADING THE DEVICE AND IT COULD NOT BE COMPLETELY CLOSED. REPLACEMENT OF THE SURGICAL INSTRUMENT WAS DONE TO COMPLETE THE CASE. THERE IS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 18-01-2019 | 75MM SELECTABLE NEW TLC | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, BLADE DOESN'T MOVE FORWARD. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
11-09-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Insufficient Information | E2401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTRECTOMY PROCEDURE, AT THE SECOND FIRING, THE RELOAD WAS CONNECTED TO THE HANDLE AND CLAMPED THE TISSUE, BUT THE HANDLE'S GREEN INDICATOR DID NOT ILLUMINATE AND FIRING COULD NOT BE PERFORMED. THE RELOAD WAS REPLACED WITH ANOTHER ONE, AND THE PROCEDURE FINISHED WITHOUT A PROBLEM. THERE WAS NO PATIENT INJURY.
21-12-2018 | Injury | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AT THE SECOND FIRING, THE STAPLING STOPPED MIDWAY THROUGH THE FIRING. IT WAS STATED THAT THE DEVICE WAS REMOVED FROM THE TISSUE BY FORCE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN.
21-12-2018 | Injury | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AT THE SECOND FIRING, THE STAPLING STOPPED MIDWAY THROUGH THE FIRING. IT WAS STATED THAT THE DEVICE WAS REMOVED FROM THE TISSUE BY FORCE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN.
10-09-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LUNG RESECTION FOR REMOVAL OF NODULE, THE HANDLE WAS MAKING A CRACKING SOUND DURING THE FIRING SEQUENCE IN THE THICK TISSUE, THE SCRUB TECH WAS UNABLE TO UNLOAD THE STAPLER AFTER FIRING. ANOTHER HANDLE WAS OPENED AND USED AND THAT HANDLE CRACKED AS WELL. IN BOTH FIRINGS, THE STAPLER FIRED COMPLETELY. UPON VISUAL INSPECTION, HE COULD SEE THE BLACK RETRACTION KNOBS WERE NOT FULLY RETRACTED ON THE HANDLE THAT COULD NOT BE UNLOADED. HE COMPLETED THE RETRACTION AND THE RELOAD DISENGAGED FROM THE STAPLER PROPERLY. THERE WAS NO PATIENT HARM.
10-09-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LUNG RESECTION FOR REMOVAL OF NODULE, THE HANDLE WAS MAKING A CRACKING SOUND DURING THE FIRING SEQUENCE IN THE THICK TISSUE, THE SCRUB TECH WAS UNABLE TO UNLOAD THE STAPLER AFTER FIRING. ANOTHER HANDLE WAS OPENED AND USED AND THAT HANDLE CRACKED AS WELL. IN BOTH FIRINGS, THE STAPLER FIRED COMPLETELY. UPON VISUAL INSPECTION, HE COULD SEE THE BLACK RETRACTION KNOBS WERE NOT FULLY RETRACTED ON THE HANDLE THAT COULD NOT BE UNLOADED. HE COMPLETED THE RETRACTION AND THE RELOAD DISENGAGED FROM THE STAPLER PROPERLY. THERE WAS NO PATIENT HARM.
11-09-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Difficult or Delayed Positioning | A150203 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THEY TRANSECTED THE BRONCHUS WITH THE STAPLE IN A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY WITH LOBECTOMY DX, THEY PRESSED THE GREEN BUTTON AND THE FIRST STAPLER¿S JAWS CLOSED AND THEY WERE UNABLE TO OPEN JAWS TO REMOVE IT FROM THE TISSUE. AFTER 40 SECONDS OF ATTEMPTING TO OPEN THE JAWS, IT REOPENED BUT THE STAPLER DID NOT FIRE THE STAPLE. THEY OPENED A SECOND STAPLER, THAT THEY ALSO HAD DIFFICULTY IN OPERATING, BUT MANAGED TO FINISH THE INTERVENTION. THERE WAS NO PATIENT INJURY.
02-08-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGEAL CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS NOT ABLE TO FIRE. THEY REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-08-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER INSTALLING THE CARTRIDGE DURING A LAPAROSCOPIC RADICAL ESOPHAGEAL CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS NOT ABLE TO FIRE. THEY REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Injury | covidien lp llc north haven | 18-01-2019 | SIGNIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AT THE SECOND FIRING, THE STAPLING STOPPED MIDWAY THROUGH THE FIRING. IT WAS STATED THAT THE DEVICE WAS REMOVED FROM THE TISSUE BY FORCE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN.
21-12-2018 | Injury | covidien lp llc north haven | 18-01-2019 | SIGNIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AT THE SECOND FIRING, THE STAPLING STOPPED MIDWAY THROUGH THE FIRING. IT WAS STATED THAT THE DEVICE WAS REMOVED FROM THE TISSUE BY FORCE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
21-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
21-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
10-08-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLON CANCER RESECTION, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLER FORMATION. THEY REPLACED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Injury | covidien lp llc north haven | 18-01-2019 | SIGNIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AT THE SECOND FIRING, THE STAPLING STOPPED MIDWAY THROUGH THE FIRING. IT WAS STATED THAT THE DEVICE WAS REMOVED FROM THE TISSUE BY FORCE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN.
21-12-2018 | Injury | covidien lp llc north haven | 18-01-2019 | SIGNIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AT THE SECOND FIRING, THE STAPLING STOPPED MIDWAY THROUGH THE FIRING. IT WAS STATED THAT THE DEVICE WAS REMOVED FROM THE TISSUE BY FORCE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. ADDITIONAL TISSUE RESECTION WAS REQUIRED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN.
06-09-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN THORACIC PROCEDURE, THE DEVICE HANDLES MADE A CRUNCHING NOISE WHILE BEING FIRED. THE FIRING DID NOT STOP, HOWEVER, ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER LOADING THE CARTRIDGE DURING SIGMOIDECTOMY PROCEDURE, THE SURGEON PUSHED THE GREEN BUTTON AND TRIED TO FIRE BUT COULD NOT. THEY CHECKED THE PRODUCT AND NOTED PRE-FIRING AT THE PROXIMAL END OF THE CARTRIDGE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
19-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER LOADING THE CARTRIDGE DURING SIGMOIDECTOMY PROCEDURE, THE SURGEON PUSHED THE GREEN BUTTON AND TRIED TO FIRE BUT COULD NOT. THEY CHECKED THE PRODUCT AND NOTED PRE-FIRING AT THE PROXIMAL END OF THE CARTRIDGE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
18-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, BLEEDING OF LESS THAN 200CC OCCURRED FREQUENTLY. THE ISSUE WAS RESOLVED BY ASTRICTION. THERE WAS NO PATIENT INJURY.
21-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WERE TWO RELOADS THAT WERE NOT FIRED AS THEY WERE BLOCKED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | UNKNOWN ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WERE TWO RELOADS THAT WERE NOT FIRED AS THEY WERE BLOCKED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
06-09-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN THORACIC PROCEDURE, THE DEVICE HANDLES MADE A CRUNCHING NOISE WHILE BEING FIRED. THE FIRING DID NOT STOP, HOWEVER, ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | IDRIVE ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, DURING PULMONARY ARTERY LIGATION, THE SURGEON COULD NOT FIRE THE DEVICE. THE GREEN BUTTON COULD BE PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, DURING PULMONARY ARTERY LIGATION, THE POWERED STAPLER COULD NOT BE FIRED. THE SURGEON THEN CHANGED TO A MANUAL STAPLER, USING THE RELOAD THAT WAS INITIALLY USED WITH THE FIRST DEVICE, HOWEVER THE SECOND STAPLER STILL WOULD NOT FIRE. THE GREEN BUTTON COULD BE PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED. THE SURGEON TRIED TO REPLACE THE RELOAD WITH A NEW ONE BUT THE RESULT IS STILL THE SAME. FINALLY, A NEW MANUAL HANDLE AND RELOAD WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER LOADING THE CARTRIDGE DURING SIGMOIDECTOMY PROCEDURE, THE SURGEON PUSHED THE GREEN BUTTON AND TRIED TO FIRE BUT COULD NOT. THEY CHECKED THE PRODUCT AND NOTED PRE-FIRING AT THE PROXIMAL END OF THE CARTRIDGE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
19-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER LOADING THE CARTRIDGE DURING SIGMOIDECTOMY PROCEDURE, THE SURGEON PUSHED THE GREEN BUTTON AND TRIED TO FIRE BUT COULD NOT. THEY CHECKED THE PRODUCT AND NOTED PRE-FIRING AT THE PROXIMAL END OF THE CARTRIDGE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, AFTER THE CARTRIDGE WAS CONNECTED WITH THE HANDLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON WAS ABLE TO SQUEEZE THE HANDLE HOWEVER, IT COULD NOT FIRE. THEN THE CARTRIDGE WAS CONNECTED WITH A DEMO HANDLE TO CHECK THE SITUATION AND WAS FOUND TO BE UNABLE TO FIRE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT INJURY.
01-11-2018 | Injury | us surgical puerto rico | 18-01-2019 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Internal Organ Perforation | Event Description: ACCORDING TO LITERATURE STUDY, JOURNAL OF ENDOUROLOGY VOLUME 32, NUMBER 11, NOVEMBER 2018, BY SOKOL REXHEPI, MD, ENTELA REXHEPI, MD, MARTIN STUMM, MD, PETER MALLMANN, MD, PHD, AND SEBASTIAN LUDWIG, MD. THE LITERATURE ARTICLE REVIEWS SURGICAL TECHNIQUE FOR TREATMENT OF APICAL PROLAPSE WITH URINARY INCONTINENCE, AFTER UNDERGOING LAPAROSCOPIC BILATERAL SACROCOLPOPEXY, IT WAS REPORTED IN THE ARTICLE THAT OUT OF 120, TOTAL OF 6 PATIENTS HAD THE FOLLOWING COMPLICATIONS AND WERE REPORTED: BLADDER INJURY (1), BOWEL PERFORATION (1), APICAL PROLAPSE (4). PATIENTS HAD TO BE REOPERATED TO RESOLVE THE ISSUE.
01-11-2018 | Injury | us surgical puerto rico | 18-01-2019 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Damage | Event Description: ACCORDING TO LITERATURE STUDY, JOURNAL OF ENDOUROLOGY VOLUME 32, NUMBER 11, NOVEMBER 2018, BY SOKOL REXHEPI, MD, ENTELA REXHEPI, MD, MARTIN STUMM, MD, PETER MALLMANN, MD, PHD, AND SEBASTIAN LUDWIG, MD. THE LITERATURE ARTICLE REVIEWS SURGICAL TECHNIQUE FOR TREATMENT OF APICAL PROLAPSE WITH URINARY INCONTINENCE, AFTER UNDERGOING LAPAROSCOPIC BILATERAL SACROCOLPOPEXY, IT WAS REPORTED IN THE ARTICLE THAT OUT OF 120, TOTAL OF 6 PATIENTS HAD THE FOLLOWING COMPLICATIONS AND WERE REPORTED: BLADDER INJURY (1), BOWEL PERFORATION (1), APICAL PROLAPSE (4). PATIENTS HAD TO BE REOPERATED TO RESOLVE THE ISSUE.
01-11-2018 | Injury | us surgical puerto rico | 18-01-2019 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Prolapse | Event Description: ACCORDING TO LITERATURE STUDY, JOURNAL OF ENDOUROLOGY VOLUME 32, NUMBER 11, NOVEMBER 2018, BY SOKOL REXHEPI, MD, ENTELA REXHEPI, MD, MARTIN STUMM, MD, PETER MALLMANN, MD, PHD, AND SEBASTIAN LUDWIG, MD. THE LITERATURE ARTICLE REVIEWS SURGICAL TECHNIQUE FOR TREATMENT OF APICAL PROLAPSE WITH URINARY INCONTINENCE, AFTER UNDERGOING LAPAROSCOPIC BILATERAL SACROCOLPOPEXY, IT WAS REPORTED IN THE ARTICLE THAT OUT OF 120, TOTAL OF 6 PATIENTS HAD THE FOLLOWING COMPLICATIONS AND WERE REPORTED: BLADDER INJURY (1), BOWEL PERFORATION (1), APICAL PROLAPSE (4). PATIENTS HAD TO BE REOPERATED TO RESOLVE THE ISSUE.
01-11-2018 | Injury | us surgical puerto rico | 18-01-2019 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO LITERATURE STUDY, JOURNAL OF ENDOUROLOGY VOLUME 32, NUMBER 11, NOVEMBER 2018, BY SOKOL REXHEPI, MD, ENTELA REXHEPI, MD, MARTIN STUMM, MD, PETER MALLMANN, MD, PHD, AND SEBASTIAN LUDWIG, MD. THE LITERATURE ARTICLE REVIEWS SURGICAL TECHNIQUE FOR TREATMENT OF APICAL PROLAPSE WITH URINARY INCONTINENCE, AFTER UNDERGOING LAPAROSCOPIC BILATERAL SACROCOLPOPEXY, IT WAS REPORTED IN THE ARTICLE THAT OUT OF 120, TOTAL OF 6 PATIENTS HAD THE FOLLOWING COMPLICATIONS AND WERE REPORTED: BLADDER INJURY (1), BOWEL PERFORATION (1), APICAL PROLAPSE (4). PATIENTS HAD TO BE REOPERATED TO RESOLVE THE ISSUE.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN IDRIVE HANDLE | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, DURING PULMONARY ARTERY LIGATION, THE SURGEON COULD NOT FIRE THE DEVICE. THE GREEN BUTTON COULD BE PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: DUPLICATE OF 702928167.
17-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, AFTER THE GREEN BUTTON WAS PRESSED, THE HANDLE WAS LOCKED AND THE FIRING COULD NOT BE PERFORMED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, AFTER THE GREEN BUTTON WAS PRESSED, THE HANDLE WAS LOCKED AND THE FIRING COULD NOT BE PERFORMED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN EGIA ADAPTER | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, DURING PULMONARY ARTERY LIGATION, THE SURGEON COULD NOT FIRE THE DEVICE. THE GREEN BUTTON COULD BE PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-10-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN BYPASS PROCEDURE, THE DEVICE FIRED 2 RELOADS; HOWEVER, DURING THE FINAL OF SECOND FIRE, THE DEVICE DID NOT FIRE AND IT WAS NECESSARY REPLACE THE DEVICE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, AFTER THE CARTRIDGE WAS CONNECTED WITH THE HANDLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON WAS ABLE TO SQUEEZE THE HANDLE HOWEVER, IT COULD NOT FIRE. THEN THE CARTRIDGE WAS CONNECTED WITH A DEMO HANDLE TO CHECK THE SITUATION AND WAS FOUND TO BE UNABLE TO FIRE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT INJURY.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | GIA | Failure to Align | A050701 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, WHEN PLACED ON THE FIELD AND BEFORE USING IT ON THE PATIENT, THE HANDLE FELL APART WHEN THE CERTIFIED SURGICAL TECHNOLOGIST WAS PREPARING TO HAND IT TO THE SURGEON. IT WAS THEN SET TO THE SIDE AND ANOTHER DEVICE WAS OPENED AND USED.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SINGLE HOLE LOBECTOMY PROCEDURE, WHILE CUTTING LUNG TISSUE, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS REPORTED THAT THE STAPLE BURST. THE RELOAD AND HANDLE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA ULTRA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SINGLE HOLE LOBECTOMY PROCEDURE, WHILE CUTTING LUNG TISSUE, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS REPORTED THAT THE STAPLE BURST. THE RELOAD AND HANDLE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SINGLE HOLE LOBECTOMY PROCEDURE, WHILE CUTTING LUNG TISSUE, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS REPORTED THAT THE STAPLE BURST. THE RELOAD AND HANDLE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA | Expulsion | A010401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SINGLE HOLE LOBECTOMY PROCEDURE, WHILE CUTTING LUNG TISSUE, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS REPORTED THAT THE STAPLE BURST. THE RELOAD AND HANDLE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2018 | Malfunction | ethicon endo-surgery | 18-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DAMAGED FIRING KNOB. IT WAS REPORTED THAT DURING THE RIGHT HEMICOLECTOMY, DURING THE 2ND FIRING, THE FIRING KNOB WAS BROKEN OFF. ALL THE PIECES WERE RETRIEVED FROM THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | UNKNOWN TRI STAPLE PRODUCT | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, DURING DISTAL RECTUM TRANSECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, DURING DISTAL RECTUM TRANSECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Human-Device Interface Problem | A22 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, DURING DISTAL RECTUM TRANSECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO INJURY.
21-12-2018 | Malfunction | us surgical puerto rico | 18-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2018 | Malfunction | ethicon endo-surgery | 18-01-2019 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF RECTAL CANCER, AFTER FIRED, NOTED THE STAPLES MALFORMED WITH STRAIGHT LEG. THE WASHER WAS CUT OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-12-2018 | Malfunction | ethicon endo-surgery | 18-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: DAMAGED FIRING KNOB. IT WAS REPORTED THAT DURING THE RIGHT HEMICOLECTOMY, DURING THE 2ND FIRING, THE FIRING KNOB WAS BROKEN OFF. ALL THE PIECES WOULD BE RETURNED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2019 | Malfunction | ethicon endo-surgery | 18-01-2019 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER'S FIRING KNOB HAD BROKEN DURING THE FIRING SEQUENCE.
22-11-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Difficult or Delayed Positioning | A150203 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE RESECTION IN A LAPAROSCOPIC UPPER LEFT LOBECTOMY PROCEDURE, THE SURGEON TRIED TO TRANSECT THE TISSUE, BUT SQUEEZING OF THE HANDLE WAS HARD. THE FIRING STOPPED IN THE MIDDLE AND THE DEVICE WAS NOT ABLE TO FIRE AFTER THAT. IN THE MIDDLE, THE HANDLE BROKE WITH ¿POP¿ SOUND FEEDBACK. ANOTHER DEVICE WAS USED BUT THE SAME ISSUE HAPPENED. THERE WAS AN INCOMPLETE STAPLE LINE IN PROXIMAL AREA WHICH WAS FIXED WITH DOUBLE STAPLING. THE ¿SAW TOOTH¿ OF THE STAPLER BODY ALSO BROKE OFF. THEY USED ADDITIONAL HANDLE AND RELOAD AND FIRED VERY SLOWLY AND CAREFULLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-11-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE RESECTION IN A LAPAROSCOPIC UPPER LEFT LOBECTOMY PROCEDURE, THE SURGEON TRIED TO TRANSECT THE TISSUE, BUT SQUEEZING OF THE HANDLE WAS HARD. THE FIRING STOPPED IN THE MIDDLE AND THE DEVICE WAS NOT ABLE TO FIRE AFTER THAT. IN THE MIDDLE, THE HANDLE BROKE WITH ¿POP¿ SOUND FEEDBACK. ANOTHER DEVICE WAS USED BUT THE SAME ISSUE HAPPENED. THERE WAS AN INCOMPLETE STAPLE LINE IN PROXIMAL AREA WHICH WAS FIXED WITH DOUBLE STAPLING. THE ¿SAW TOOTH¿ OF THE STAPLER BODY ALSO BROKE OFF. THEY USED ADDITIONAL HANDLE AND RELOAD AND FIRED VERY SLOWLY AND CAREFULLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-11-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE RESECTION IN A LAPAROSCOPIC UPPER LEFT LOBECTOMY PROCEDURE, THE SURGEON TRIED TO TRANSECT THE TISSUE, BUT SQUEEZING OF THE HANDLE WAS HARD. THE FIRING STOPPED IN THE MIDDLE AND THE DEVICE WAS NOT ABLE TO FIRE AFTER THAT. IN THE MIDDLE, THE HANDLE BROKE WITH ¿POP¿ SOUND FEEDBACK. ANOTHER DEVICE WAS USED BUT THE SAME ISSUE HAPPENED. THERE WAS AN INCOMPLETE STAPLE LINE IN PROXIMAL AREA WHICH WAS FIXED WITH DOUBLE STAPLING. THE ¿SAW TOOTH¿ OF THE STAPLER BODY ALSO BROKE OFF. THEY USED ADDITIONAL HANDLE AND RELOAD AND FIRED VERY SLOWLY AND CAREFULLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-10-2018 | Malfunction | covidien lp llc north haven | 18-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN BYPASS PROCEDURE, THE DEVICE FIRED 2 RELOADS; HOWEVER, DURING THE FINAL OF SECOND FIRE, THE DEVICE DID NOT FIRE AND IT WAS NECESSARY TO REPLACE THE DEVICE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
22-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS NOT ABLE TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | EEA | Difficult to Remove | A150207 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS NOT MOVING AND WAS STUCK INSIDE.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE DISCONNECTED VASCULAR ARTERY, WHEN THE CARTRIDGE WAS LOADED, IT COULD NOT BE FIRED. THE DEVICE REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | UNKNOWN ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE DISCONNECTED VASCULAR ARTERY, WHEN THE CARTRIDGE WAS LOADED, IT COULD NOT BE FIRED. THE DEVICE REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | ABSORBATACK | Difficult to Remove | A150207 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA OPERATION, WHILE THE DEVICE WAS BEING APPLIED FOR MESH FIXATION, THE DEVICE WAS BENDING AND STICKING IN THE TROCAR. THE DEVICE WAS TAKEN OUT OF WITH TROCAR AND ANOTHER DEVICE WAS PLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | ABSORBATACK | Material Twisted/Bent | A040609 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA OPERATION, WHILE THE DEVICE WAS BEING APPLIED FOR MESH FIXATION, THE DEVICE WAS BENDING AND STICKING IN THE TROCAR. THE DEVICE WAS TAKEN OUT OF WITH TROCAR AND ANOTHER DEVICE WAS PLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE DISCONNECTED AIR TUBE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THEY USED NEW NAILS TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | UNKNOWN ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC LOBECTOMY PROCEDURE, WHILE THE DEVICE WAS BEING APPLIED TO THE DISCONNECTED AIR TUBE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THEY USED NEW NAILS TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY WEIGHT LOSS SURGERY, WHEN THEY DISCONNECTED THE GASTRIC ANTRUM. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY REPLACED WITH A NEW STAPLER AND CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY WEIGHT LOSS SURGERY, WHEN THEY DISCONNECTED THE GASTRIC ANTRUM. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY REPLACED WITH A NEW STAPLER AND CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY WEIGHT LOSS SURGERY, WHEN THEY DISCONNECTED THE GASTRIC ANTRUM. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY REPLACED WITH A NEW STAPLER AND CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY WEIGHT LOSS SURGERY, WHEN THEY DISCONNECTED THE GASTRIC ANTRUM. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY REPLACED WITH A NEW STAPLER AND CARTRIDGE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | TRI-STAPLE 2.0 | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STAPLING AND DIVIDING THE STOMACH, THE POWERED STAPLER FIRED PARTIALLY ON TISSUE AND WOULD NOT ADVANCE AFTER ABOUT HALFWAY THROUGH THE FIRING. THE SURGEON TRIED FIRING THE RELOADS WITH MULTIPLE POWERED STAPLER AND A MANUAL HANDLE, BUT THE RELOAD WAS ONLY ABLE TO PARTIALLY FIRE EACH TIME. THE SURGEON WAS ABLE TO COMPLETE THE SLEEVE WHEN HE SWITCHED TO ANOTHER POWERED STAPLER AND FIRED MEDIAL TO THE PRIOR PARTIAL FIRINGS. THE STAPLER DID SHOW AN EXCESSIVE FORCE ICON HALFWAY THROUGH THE FIRST FIRING AT THE BASE OF THE STOMACH, WHICH WAS VERY THICK, BUT AFTER A WAIT, THE FIRING WAS ABLE TO FINISH. IT WAS NOTED THAT THE DISTAL STAPLE LINE WAS INCOMPLETE. THE SURGEON COMPLETED THE SLEEVE WITH 8 RELOADS AND WAS READING ZONE 3 AND FIRING SLOWLY THE WHOLE WAY UP. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | UNKNOWN IDRIVE HANDLE | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STAPLING AND DIVIDING THE STOMACH, THE POWERED STAPLER FIRED PARTIALLY ON TISSUE AND WOULD NOT ADVANCE AFTER ABOUT HALFWAY THROUGH THE FIRING. THE SURGEON TRIED FIRING THE RELOADS WITH MULTIPLE POWERED STAPLER AND A MANUAL HANDLE, BUT THE RELOAD WAS ONLY ABLE TO PARTIALLY FIRE EACH TIME. THE SURGEON WAS ABLE TO COMPLETE THE SLEEVE WHEN HE SWITCHED TO ANOTHER POWERED STAPLER AND FIRED MEDIAL TO THE PRIOR PARTIAL FIRINGS. THE STAPLER DID SHOW AN EXCESSIVE FORCE ICON HALFWAY THROUGH THE FIRST FIRING AT THE BASE OF THE STOMACH, WHICH WAS VERY THICK, BUT AFTER A WAIT, THE FIRING WAS ABLE TO FINISH. IT WAS NOTED THAT THE DISTAL STAPLE LINE WAS INCOMPLETE. THE SURGEON COMPLETED THE SLEEVE WITH 8 RELOADS AND WAS READING ZONE 3 AND FIRING SLOWLY THE WHOLE WAY UP. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | UNKNOWN IDRIVE HANDLE | Difficult to Advance | A150205 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STAPLING AND DIVIDING THE STOMACH, THE POWERED STAPLER FIRED PARTIALLY ON TISSUE AND WOULD NOT ADVANCE AFTER ABOUT HALFWAY THROUGH THE FIRING. THE SURGEON TRIED FIRING THE RELOADS WITH MULTIPLE POWERED STAPLER AND A MANUAL HANDLE, BUT THE RELOAD WAS ONLY ABLE TO PARTIALLY FIRE EACH TIME. THE SURGEON WAS ABLE TO COMPLETE THE SLEEVE WHEN HE SWITCHED TO ANOTHER POWERED STAPLER AND FIRED MEDIAL TO THE PRIOR PARTIAL FIRINGS. THE STAPLER DID SHOW AN EXCESSIVE FORCE ICON HALFWAY THROUGH THE FIRST FIRING AT THE BASE OF THE STOMACH, WHICH WAS VERY THICK, BUT AFTER A WAIT, THE FIRING WAS ABLE TO FINISH. IT WAS NOTED THAT THE DISTAL STAPLE LINE WAS INCOMPLETE. THE SURGEON COMPLETED THE SLEEVE WITH 8 RELOADS AND WAS READING ZONE 3 AND FIRING SLOWLY THE WHOLE WAY UP. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | UNKNOWN ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STAPLING AND DIVIDING THE STOMACH, THE POWERED STAPLER FIRED PARTIALLY ON TISSUE AND WOULD NOT ADVANCE AFTER ABOUT HALFWAY THROUGH THE FIRING. THE SURGEON TRIED FIRING THE RELOADS WITH MULTIPLE POWERED STAPLER AND A MANUAL HANDLE, BUT THE RELOAD WAS ONLY ABLE TO PARTIALLY FIRE EACH TIME. THE SURGEON WAS ABLE TO COMPLETE THE SLEEVE WHEN HE SWITCHED TO ANOTHER POWERED STAPLER AND FIRED MEDIAL TO THE PRIOR PARTIAL FIRINGS. THE STAPLER DID SHOW AN EXCESSIVE FORCE ICON HALFWAY THROUGH THE FIRST FIRING AT THE BASE OF THE STOMACH, WHICH WAS VERY THICK, BUT AFTER A WAIT, THE FIRING WAS ABLE TO FINISH. IT WAS NOTED THAT THE DISTAL STAPLE LINE WAS INCOMPLETE. THE SURGEON COMPLETED THE SLEEVE WITH 8 RELOADS AND WAS READING ZONE 3 AND FIRING SLOWLY THE WHOLE WAY UP. THERE WAS NO PATIENT INJURY.
23-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT HEPATIC INCISION, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT HEPATIC INCISION, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRO JEJUNOSTOMY PROCEDURE, AFTER PUSHING THE GREEN BUTTON SURGEON CAN ABLE SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTRECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTRECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | SIGNIA | Device Displays Incorrect Message | A090201 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STAPLING AND DIVIDING THE STOMACH, THE POWERED STAPLER FIRED PARTIALLY ON TISSUE AND WOULD NOT ADVANCE AFTER ABOUT HALFWAY THROUGH THE FIRING. THE SURGEON TRIED FIRING THE RELOADS WITH MULTIPLE POWERED STAPLER AND A MANUAL HANDLE, BUT THE RELOAD WAS ONLY ABLE TO PARTIALLY FIRE EACH TIME. THE SURGEON WAS ABLE TO COMPLETE THE SLEEVE WHEN HE SWITCHED TO ANOTHER POWERED STAPLER AND FIRED MEDIAL TO THE PRIOR PARTIAL FIRINGS. THE STAPLER DID SHOW AN EXCESSIVE FORCE ICON HALFWAY THROUGH THE FIRST FIRING AT THE BASE OF THE STOMACH, WHICH WAS VERY THICK, BUT AFTER A WAIT, THE FIRING WAS ABLE TO FINISH. IT WAS NOTED THAT THE DISTAL STAPLE LINE WAS INCOMPLETE. THE SURGEON COMPLETED THE SLEEVE WITH 8 RELOADS AND WAS READING ZONE 3 AND FIRING SLOWLY THE WHOLE WAY UP. THERE WAS NO PATIENT INJURY.
14-11-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY, WHILE ON A BROKEN BLOOD VESSEL, THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY. THE HANDLE ALSO MAKES A STRANGE NOISE. TO RESOLVE THE ISSUE, THEY REPLACED THE DEVICE INTO A NEW ONE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | EEA ORVIL | Difficult to Remove | A150207 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER IN A LAPAROSCOPIC TOTAL GASTRECTOMY, DURING THE ESOPHAGUS JEJUNAL ANASTOMOSIS. AFTER FIRING, THE DOCTOR REMOVED THE DEVICE FROM THE JEJUNUM, BUT THE ANVIL DID NOT TILT AND THE DEVICE COULD NOT BE REMOVED. AT THE TIME OF FIRING, THE DOCTOR FIRMLY SQUEEZED AND ROTATED THE RETURN KNOB WITHOUT PROBLEM TWO TIMES NECESSARILY FOR TILTING THE ANVIL. THE DOCTOR PUSHED THE DEVICE TO THE ORAL SIDE IN ORDER TO TILT THE ANVIL, BUT THE ANVIL DID NOT TILT, AND THE DOCTOR REPEATED THIS SAME OPERATION CONTINUOUSLY FOR ABOUT 15 MINUTES, AND FINALLY MANAGED TO REMOVE THE DEVICE FROM THE ANASTOMOSIS PART. AT THAT TIME IT WAS DIFFICULT TO CHECK THE STATUS OF THE ANVIL, AND WHEN THE DEVICE WAS REMOVED OUT OF THE JEJUNUM, THE ANVIL WAS INDEED NOT TILTED. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE DEVICE WAS REMOVED FROM THE TISSUE BY FORCE BUT NO TISSUE DAMAGE WAS CAUSED.
25-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AT THE FIRST FIRING, THE INSTRUMENT DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AT THE FIRST FIRING, THE INSTRUMENT DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Injury | us surgical puerto rico | 20-01-2019 | ILA | Failure to Cut | A050702 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A DURING FUNCTIONAL END-TO-END ANASTOMOSIS IN TRANSVERSE COLECTOMY. THE DEFECT OCCURRED WHEN CLOSING THE COMMON HOLE AFTER SIDE-BY-SIDE ANASTOMOSIS. AT THE FIRST FIRING, WHILE GRASPING THE TISSUE AND PROCEEDING THE FIRING, IT BECAME HARD FROM THE HALFWAY AND THE KNIFE BAR BECAME BENT ALTHOUGH THE FIRING ADVANCED TO THE END. WHEN OBSERVING THE TISSUE AFTER RELEASING THE DEVICE, IT WAS FOUND THAT THE STAPLES WERE APPLIED, BUT THE CUTTING WAS ONLY HALF WAY DONE. THE TISSUE THAT WAS NOT SEPARATED COMPLETELY WAS CUT WITH AN ELECTROSURGICAL KNIFE AND REINFORCED WITH THE NEEDLE AND THREAD, AND THE OPERATION WAS COMPLETED.
21-12-2018 | Injury | us surgical puerto rico | 20-01-2019 | ILA | Material Twisted/Bent | A040609 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A DURING FUNCTIONAL END-TO-END ANASTOMOSIS IN TRANSVERSE COLECTOMY. THE DEFECT OCCURRED WHEN CLOSING THE COMMON HOLE AFTER SIDE-BY-SIDE ANASTOMOSIS. AT THE FIRST FIRING, WHILE GRASPING THE TISSUE AND PROCEEDING THE FIRING, IT BECAME HARD FROM THE HALFWAY AND THE KNIFE BAR BECAME BENT ALTHOUGH THE FIRING ADVANCED TO THE END. WHEN OBSERVING THE TISSUE AFTER RELEASING THE DEVICE, IT WAS FOUND THAT THE STAPLES WERE APPLIED, BUT THE CUTTING WAS ONLY HALF WAY DONE. THE TISSUE THAT WAS NOT SEPARATED COMPLETELY WAS CUT WITH AN ELECTROSURGICAL KNIFE AND REINFORCED WITH THE NEEDLE AND THREAD, AND THE OPERATION WAS COMPLETED.
21-12-2018 | Injury | covidien lp llc north haven | 20-01-2019 | ILA | Failure to Cut | A050702 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A DURING FUNCTIONAL END-TO-END ANASTOMOSIS IN TRANSVERSE COLECTOMY. THE DEFECT OCCURRED WHEN CLOSING THE COMMON HOLE AFTER SIDE-BY-SIDE ANASTOMOSIS. AT THE FIRST FIRING, WHILE GRASPING THE TISSUE AND PROCEEDING THE FIRING, IT BECAME HARD FROM THE HALFWAY AND THE KNIFE BAR BECAME BENT ALTHOUGH THE FIRING ADVANCED TO THE END. WHEN OBSERVING THE TISSUE AFTER RELEASING THE DEVICE, IT WAS FOUND THAT THE STAPLES WERE APPLIED, BUT THE CUTTING WAS ONLY HALF WAY DONE. THE TISSUE THAT WAS NOT SEPARATED COMPLETELY WAS CUT WITH AN ELECTROSURGICAL KNIFE AND REINFORCED WITH THE NEEDLE AND THREAD, AND THE OPERATION WAS COMPLETED.
21-12-2018 | Injury | covidien lp llc north haven | 20-01-2019 | ILA | Material Twisted/Bent | A040609 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A DURING FUNCTIONAL END-TO-END ANASTOMOSIS IN TRANSVERSE COLECTOMY. THE DEFECT OCCURRED WHEN CLOSING THE COMMON HOLE AFTER SIDE-BY-SIDE ANASTOMOSIS. AT THE FIRST FIRING, WHILE GRASPING THE TISSUE AND PROCEEDING THE FIRING, IT BECAME HARD FROM THE HALFWAY AND THE KNIFE BAR BECAME BENT ALTHOUGH THE FIRING ADVANCED TO THE END. WHEN OBSERVING THE TISSUE AFTER RELEASING THE DEVICE, IT WAS FOUND THAT THE STAPLES WERE APPLIED, BUT THE CUTTING WAS ONLY HALF WAY DONE. THE TISSUE THAT WAS NOT SEPARATED COMPLETELY WAS CUT WITH AN ELECTROSURGICAL KNIFE AND REINFORCED WITH THE NEEDLE AND THREAD, AND THE OPERATION WAS COMPLETED.
25-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HEPATECTOMY, WHEN THE FIRST HEPATIC GATE WAS DISCONNECTED, THE DEVICE DID NOT FIRE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | us surgical puerto rico | 20-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HEPATECTOMY, WHEN THE FIRST HEPATIC GATE WAS DISCONNECTED, THE DEVICE DID NOT FIRE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | covidien lp llc north haven | 20-01-2019 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HEPATECTOMY, WHEN THE FIRST HEPATIC GATE WAS DISCONNECTED, THE DEVICE DID NOT FIRE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRO JEJUNOSTOMY PROCEDURE, AFTER PUSHING THE GREEN BUTTON SURGEON CAN ABLE SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE ENDOSCOPIC-ASSISTED RADICAL RESECTION OF RECTAL CANCER, AFTER FIRED, NOTED THE STAPLES MALFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
16-11-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | 75MM SELECTABLE NEW TLC | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN COLECTOMY, THE FIRING KNOB DID NOT RETURN TO THE ORIGINAL POSITION DURING USE, AND THE ANVIL HALF AND THE CARTRIDGE HALF COULD NOT BE JOINED TOGETHER. THE DEVICE WAS USED FOR FUNCTIONAL END-TO-END ANASTOMOSIS (FEEA) ON THE COLON. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
21-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | ENDO GIA | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE SURGEON FOUND FRAGMENT IN PATIENT'S THORACIC CAVITY. IT WAS COLLECTED BY THE USE OF FORCEPS. THE SURGERY WAS CONTINUED AND WAS COMPLETED WITHOUT PROBLEM. THE YELLOW FRAGMENT WAS CONSIDERED TO BE A PART OF THE YELLOW TAB OF THE DEVICE. THE SALES REP CHECKED WITH THE SCRUB NURSE. IT WAS SAID THAT THE PROCEDURE WAS OPERATED IN THE FOLLOWING ORDER: AFTER CONNECTING THE CARTRIDGE WITH HANDLE, PULLED THE YELLOW TAB AND THEN CHECKED OPENING/CLOSING OF THE JAWS. THE STATUS OF THE PATIENT WAS REPORTED AS NO PROBLEM.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS, AT THE FIRST FIRING, THE SURGEON GRASPED THE SMALL BOWEL WITH THE PRODUCT AND PRESSED THE GREEN BUTTON, THEN TRIED TO SQUEEZE THE HANDLE ATTEMPTING TO FIRE, BUT THE HANDLE WAS NOT ABLE TO BE SQUEEZED. A NEW HANDLE WAS USED AND CONNECTED TO THE SAME RELOAD THEN TRIED TO FIRE, BUT THE HANDLE STILL WAS NOT ABLE TO BE SQUEEZED. A NEW HANDLE AND A NEW RELOAD WERE TAKEN OUT, AND THE SURGERY WAS CONTINUED. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS, AT THE FIRST FIRING, THE SURGEON GRASPED THE SMALL BOWEL WITH THE PRODUCT AND PRESSED THE GREEN BUTTON, THEN TRIED TO SQUEEZE THE HANDLE ATTEMPTING TO FIRE, BUT THE HANDLE WAS NOT ABLE TO BE SQUEEZED. A NEW HANDLE WAS USED AND CONNECTED TO THE SAME RELOAD THEN TRIED TO FIRE, BUT THE HANDLE STILL WAS NOT ABLE TO BE SQUEEZED. A NEW HANDLE AND A NEW RELOAD WERE TAKEN OUT, AND THE SURGERY WAS CONTINUED. THERE WAS NO PATIENT INJURY.
28-12-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, AT THE TIME OF FIRING THE CDH29A, THE BLADE DID NOT WORK, ONLY STAPLING THE FABRIC LEAVING IT UNCUT. THE SURGEON TRIED TO EXTRACT THE CIRCULAR STAPLER, HOWEVER IT WAS VERY DIFFICULT, SO HE DECIDED TO KEEP THE CDH29A AND CUT THE SURROUNDING TISSUE TO RELEASE IT AND PERFORM ANOTHER ANASTOMOSIS. THIS IS WHY A CS40G PLUS AN SR75 WERE USED TO ACHIEVE THE TERMINO-TERMINAL ANATOMY AGAIN.
25-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THEY THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN THE SURGEON FIRED THE STAPLE TO TARGET TISSUE HE FELT SOMETHING STUCK ON STAPLE'S JAW. SURGEON MOVE IT BACK THE KNIFE AND OPEN THE JAW. SURGEON AND SCRUB NURSE SAID THEY FOUND NON-SHAPED STAPLES AND STAPLE BODY'S KNIFE WAS LOOKED IT WAS NOT ENOUGH SHARP TO CUT FOR TISSUE.
01-01-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN THE SURGEON FIRED THE STAPLE TO TARGET TISSUE HE FELT SOMETHING STUCK ON STAPLE'S JAW. SURGEON MOVE IT BACK THE KNIFE AND OPEN THE JAW. SURGEON AND SCRUB NURSE SAID THEY FOUND NON-SHAPED STAPLES AND STAPLE BODY'S KNIFE WAS LOOKED IT WAS NOT ENOUGH SHARP TO CUT FOR TISSUE.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL CANCER RESECTION, WHEN THE RECTUM WAS DISCONNECTED, THERE WAS A LOUD NOISE WHEN THE DEVICE WAS FIRST FIRED. IT COULD NOT BE FIRED WHEN THE SECOND NAIL WAS LOADED. THEY REPLACED THE DEVICE WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL CANCER RESECTION, WHEN THE RECTUM WAS DISCONNECTED, THERE WAS A LOUD NOISE WHEN THE DEVICE WAS FIRST FIRED. IT COULD NOT BE FIRED WHEN THE SECOND NAIL WAS LOADED. THEY REPLACED THE DEVICE WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | PWRD ECH FLEX 60MM 34CM SHAFT | Material Deformation | A0406 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RADICAL RESECTION OF RECTAL CANCER, COULD NOT INSERT INTO THE TROCAR. CHECKED THE DEVICE, NOTED SHAFT WAS DEFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
26-12-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE THAT THE TECH HAD A HARD TIME LOADING RELOAD, BUT FINALLY DID IT. NOT BECAUSE OF LACK OF FAMILIARITY, SHE HAS LOADED THE STAPLERS BEFORE. USED SEAM GUARD. THE STAPLES ON THE PORTION OF THE STOMACH LEFT INSIDE THE PATIENT WERE FIRED PROPERLY, BUT THE STAPLES ON THE REMNANTS OF THE STOMACH COMING OUT OF THE PATIENT WERE NOT FIRED. REMNANT OF THE STOMACH LEFT OPEN. KNIFE DID TRANSECT. NOT SURE WHAT PATIENT OUTCOME HAS BEEN, WAITING TO SEE HOW THEY RECOVER. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED.
25-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL CANCER RESECTION, WHEN THE RECTUM WAS DISCONNECTED, THERE WAS A LOUD NOISE WHEN THE DEVICE WAS FIRST FIRED. IT COULD NOT BE FIRED WHEN THE SECOND NAIL WAS LOADED. THEY REPLACED THE DEVICE WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE SECOND RELOAD DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLES STARTED TO AND JAMMED IN THE MIDDLE. THE JAWS LOCKED ON THE STOMACH. IT WAS STATED THAT THEY HAD TO USE FORCE TO RETRACT THE KNIFE. ADDITIONALLY, IT WAS REPORTED THAT PRIOR TO THE INCIDENT, THERE WAS A LOUD BREAKING SOUND COMING FROM THE STAPLER AND AFTER THAT, THEY COULD NOT COMPLETE THE FIRING OR RETRACT THE KNIFE. THE SURGERY WAS EXTENDED TO 20 MINUTES TO SOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE SECOND RELOAD DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLES STARTED TO AND JAMMED IN THE MIDDLE. THE JAWS LOCKED ON THE STOMACH. IT WAS STATED THAT THEY HAD TO USE FORCE TO RETRACT THE KNIFE. ADDITIONALLY, IT WAS REPORTED THAT PRIOR TO THE INCIDENT, THERE WAS A LOUD BREAKING SOUND COMING FROM THE STAPLER AND AFTER THAT, THEY COULD NOT COMPLETE THE FIRING OR RETRACT THE KNIFE. THE SURGERY WAS EXTENDED TO 20 MINUTES TO SOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE SECOND RELOAD DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLES STARTED TO AND JAMMED IN THE MIDDLE. THE JAWS LOCKED ON THE STOMACH. IT WAS STATED THAT THEY HAD TO USE FORCE TO RETRACT THE KNIFE. ADDITIONALLY, IT WAS REPORTED THAT PRIOR TO THE INCIDENT, THERE WAS A LOUD BREAKING SOUND COMING FROM THE STAPLER AND AFTER THAT, THEY COULD NOT COMPLETE THE FIRING OR RETRACT THE KNIFE. THE SURGERY WAS EXTENDED TO 20 MINUTES TO SOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | ENDO GIA | Entrapment of Device | A150208 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE SECOND RELOAD DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLES STARTED TO AND JAMMED IN THE MIDDLE. THE JAWS LOCKED ON THE STOMACH. IT WAS STATED THAT THEY HAD TO USE FORCE TO RETRACT THE KNIFE. ADDITIONALLY, IT WAS REPORTED THAT PRIOR TO THE INCIDENT, THERE WAS A LOUD BREAKING SOUND COMING FROM THE STAPLER AND AFTER THAT, THEY COULD NOT COMPLETE THE FIRING OR RETRACT THE KNIFE. THE SURGERY WAS EXTENDED TO 20 MINUTES TO SOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE SECOND RELOAD DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLES STARTED TO AND JAMMED IN THE MIDDLE. THE JAWS LOCKED ON THE STOMACH. IT WAS STATED THAT THEY HAD TO USE FORCE TO RETRACT THE KNIFE. ADDITIONALLY, IT WAS REPORTED THAT PRIOR TO THE INCIDENT, THERE WAS A LOUD BREAKING SOUND COMING FROM THE STAPLER AND AFTER THAT, THEY COULD NOT COMPLETE THE FIRING OR RETRACT THE KNIFE. THE SURGERY WAS EXTENDED TO 20 MINUTES TO SOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN SLEEVE GASTRECTOMY, THE ENDOCUTTER, PLEE60A, AND HE HAD A GREEN RELOAD GST60G, ON THE LAST FIRING IT STAPLED ON THE REMNANT SIDE OF THE STOMACH AND LEFT THE PATIENT SIDE OF THE STOMACH WIDE OPEN. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE, THEN HAD TO STAPLE IN THE STOMACH TO WHERE IT ALMOST MEETS THE ESOPHAGUS TO GET IT STAPLED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | EEA | Difficult to Remove | A150207 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER IN A LAPAROSCOPIC TOTAL GASTRECTOMY, DURING THE ESOPHAGUS JEJUNAL ANASTOMOSIS. AFTER FIRING, THE DOCTOR REMOVED THE DEVICE FROM THE JEJUNUM, BUT THE ANVIL DID NOT TILT AND THE DEVICE COULD NOT BE REMOVED. AT THE TIME OF FIRING, THE DOCTOR FIRMLY SQUEEZED AND ROTATED THE RETURN KNOB WITHOUT PROBLEM TWO TIMES NECESSARILY FOR TILTING THE ANVIL. THE DOCTOR PUSHED THE DEVICE TO THE ORAL SIDE IN ORDER TO TILT THE ANVIL, BUT THE ANVIL DID NOT TILT, AND THE DOCTOR REPEATED THIS SAME OPERATION CONTINUOUSLY FOR ABOUT 15 MINUTES, AND FINALLY MANAGED TO REMOVE THE DEVICE FROM THE ANASTOMOSIS PART. AT THAT TIME IT WAS DIFFICULT TO CHECK THE STATUS OF THE ANVIL, AND WHEN THE DEVICE WAS REMOVED OUT OF THE JEJUNUM, THE ANVIL WAS INDEED NOT TILTED. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THE DEVICE WAS REMOVED FROM THE TISSUE BY FORCE BUT NO TISSUE DAMAGE WAS CAUSED.
01-01-2018 | Injury | ethicon endo-surgery | 21-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED A LAP APPENDECTOMY ON (B)(6) 2018. AFTER FIRING THE ATS45 ENDOCUTTER, THE JAWS WOULD NOT OPEN WHEN THE RELEASE BUTTON WAS PRESSED. THE OR CALLED THE LOCAL REPRESENTATIVE FOR TROUBLESHOOTING SUGGESTIONS. WHILE ON THE PHONE WITH THE REPRESENTATIVE, IT WAS LEARNED BY THE REPRESENTATIVE THAT THE RELEASE BUTTON WAS PRESSED, FULLY DOWN, AND THE JAWS WOULD STILL NOT OPEN. THE REPRESENTATIVE WAS TOLD THE STAPLER MADE A "NOISE" WHEN THE BUTTON WAS PRESSED. THE REPRESENTATIVE SUGGESTED PRYING THE FIRING HANDLE AWAY FROM THE CLOSING HANDLE. ONCE THAT WAS DONE, THE JAWS STILL DID NOT OPEN AND THE REP SUGGESTED DOING THE SAME WITH THE CLOSING HANDLE. THE JAWS THEN OPENED AND THE STAPLER WAS PUT ASIDE. A STAFF MEMBER IN THE ROOM TOLD THE REPRESENTATIVE THAT THE STAPLES "DIDN'T WORK" AND THAT THEY WERE GOING TO HAVE TO OPEN.
01-01-2018 | Injury | ethicon endo-surgery | 21-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED A LAP APPENDECTOMY ON (B)(6) 2018. AFTER FIRING THE ATS45 ENDOCUTTER, THE JAWS WOULD NOT OPEN WHEN THE RELEASE BUTTON WAS PRESSED. THE OR CALLED THE LOCAL REPRESENTATIVE FOR TROUBLESHOOTING SUGGESTIONS. WHILE ON THE PHONE WITH THE REPRESENTATIVE, IT WAS LEARNED BY THE REPRESENTATIVE THAT THE RELEASE BUTTON WAS PRESSED, FULLY DOWN, AND THE JAWS WOULD STILL NOT OPEN. THE REPRESENTATIVE WAS TOLD THE STAPLER MADE A "NOISE" WHEN THE BUTTON WAS PRESSED. THE REPRESENTATIVE SUGGESTED PRYING THE FIRING HANDLE AWAY FROM THE CLOSING HANDLE. ONCE THAT WAS DONE, THE JAWS STILL DID NOT OPEN AND THE REP SUGGESTED DOING THE SAME WITH THE CLOSING HANDLE. THE JAWS THEN OPENED AND THE STAPLER WAS PUT ASIDE. A STAFF MEMBER IN THE ROOM TOLD THE REPRESENTATIVE THAT THE STAPLES "DIDN'T WORK" AND THAT THEY WERE GOING TO HAVE TO OPEN.
01-01-2018 | Injury | ethicon endo-surgery | 21-01-2019 | ETS FLEX ARTICNG LNR CUTR 45MM | Noise, Audible | A0508 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED A LAP APPENDECTOMY ON (B)(6) 2018. AFTER FIRING THE ATS45 ENDOCUTTER, THE JAWS WOULD NOT OPEN WHEN THE RELEASE BUTTON WAS PRESSED. THE OR CALLED THE LOCAL REPRESENTATIVE FOR TROUBLESHOOTING SUGGESTIONS. WHILE ON THE PHONE WITH THE REPRESENTATIVE, IT WAS LEARNED BY THE REPRESENTATIVE THAT THE RELEASE BUTTON WAS PRESSED, FULLY DOWN, AND THE JAWS WOULD STILL NOT OPEN. THE REPRESENTATIVE WAS TOLD THE STAPLER MADE A "NOISE" WHEN THE BUTTON WAS PRESSED. THE REPRESENTATIVE SUGGESTED PRYING THE FIRING HANDLE AWAY FROM THE CLOSING HANDLE. ONCE THAT WAS DONE, THE JAWS STILL DID NOT OPEN AND THE REP SUGGESTED DOING THE SAME WITH THE CLOSING HANDLE. THE JAWS THEN OPENED AND THE STAPLER WAS PUT ASIDE. A STAFF MEMBER IN THE ROOM TOLD THE REPRESENTATIVE THAT THE STAPLES "DIDN'T WORK" AND THAT THEY WERE GOING TO HAVE TO OPEN.
24-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO ASSISTED THORACOSCOPIC SURGERY FOR SEGMENTECTOMY, THERE WAS WEIRD SOUND HEARD AND THE DEVICE DID NOT WORK. IN TRANSECTION OF PARENCHYMA, THERE WAS POOR STAPLE SHAPE NOTED IN THE RELOAD. ANOTHER RELOAD USED FOR BRONCHUS TRANSECTION WAS NOT BALE TO FIRE. NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO ASSISTED THORACOSCOPIC SURGERY FOR SEGMENTECTOMY, THERE WAS WEIRD SOUND HEARD AND THE DEVICE DID NOT WORK. IN TRANSECTION OF PARENCHYMA, THERE WAS POOR STAPLE SHAPE NOTED IN THE RELOAD. ANOTHER RELOAD USED FOR BRONCHUS TRANSECTION WAS NOT BALE TO FIRE. NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO ASSISTED THORACOSCOPIC SURGERY FOR SEGMENTECTOMY, THERE WAS WEIRD SOUND HEARD AND THE DEVICE DID NOT WORK. IN TRANSECTION OF PARENCHYMA, THERE WAS POOR STAPLE SHAPE NOTED IN THE RELOAD. ANOTHER RELOAD USED FOR BRONCHUS TRANSECTION WAS NOT BALE TO FIRE. NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO ASSISTED THORACOSCOPIC SURGERY FOR SEGMENTECTOMY, THERE WAS WEIRD SOUND HEARD AND THE DEVICE DID NOT WORK. IN TRANSECTION OF PARENCHYMA, THERE WAS POOR STAPLE SHAPE NOTED IN THE RELOAD. ANOTHER RELOAD USED FOR BRONCHUS TRANSECTION WAS NOT BALE TO FIRE. NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: HOLD FOR ESTHER (2/17/2023)
24-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | GIA | Detachment of Device or Device Component | A0501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVE AN OPEN GASTRECTOMY PROCEDURE, THE BLACK KNOB OF THE DEVICE FELL OUT. A NEW DEVICE WAS USED TO COMPLETE THE CASE. NO PATIENT INVOLVEMENT.
26-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RIGHT UPPER LOBECTOMY, DURING THE ANASTOMOTIC RESECTION OF THE RIGHT UPPER LOBE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THE RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | UNKNOWN ENDO GI | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RIGHT UPPER LOBECTOMY, DURING THE ANASTOMOTIC RESECTION OF THE RIGHT UPPER LOBE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THE RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | TA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN THORACIC ESOPHAGECTOMY, WHILE USING THE PRODUCT TO CLOSE THE TISSUE, IT WAS FOUND THAT THE STAPLER COULD BE CLAMPED BUT COULDN'T BE FIRED. THE SURGEON PERFORMED FULL, COMPLETE SQUEEZE OF THE HANDLE BUT IT DID NOT RETURN TO OPEN POSITION FOLLOWING THE FIRST SQUEEZE AND THE HANDLE REMAIN CLOSED POSITION FOLLOWING THE SECOND SQUEEZE OF THE HANDLE. AFTER REPEATED ATTEMPTS, THE DOCTOR USED ANOTHER PRODUCT OF THE SAME TYPE, AND THE OPERATION ENDED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY RESECTION, INTERPULMONARY RUPTURE, THE DEVICE COULDN'T BE FIRED AFTER IT SOUNDED LIKE "PIPA". THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY RESECTION, INTERPULMONARY RUPTURE, THE DEVICE COULDN'T BE FIRED AFTER IT SOUNDED LIKE "PIPA". THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY RESECTION, INTERPULMONARY RUPTURE, THE DEVICE COULDN'T BE FIRED AFTER IT SOUNDED LIKE "PIPA". THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RIGHT HEMICOLECTOMY, WHILE CLOSING THE COMMON OPENING OF THE INTESTINE, THE DEVICE DID NOT FIRE. THE BLACK PUSHER THUMB PIECE CANNOT BE MOVED AT ALL. REPLACED THE OLD DEVICE WITH A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RIGHT HEMICOLECTOMY, WHILE CLOSING THE COMMON OPENING OF THE INTESTINE, THE DEVICE DID NOT FIRE. THE BLACK PUSHER THUMB PIECE CANNOT BE MOVED AT ALL. REPLACED THE OLD DEVICE WITH A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
29-10-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER LOCKED AND DID NOT FIRE. THERE WAS NO PATIENT INJURY.
29-10-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER LOCKED AND DID NOT FIRE. THERE WAS NO PATIENT INJURY.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Detachment of Device or Device Component | A0501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Detachment of Device or Device Component | A0501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Injury | us surgical puerto rico | 21-01-2019 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE WHEN CUTTING AND SUTURING THE LUNGS, THE DEVICE COULD NOT BE FIRED. THE JAWS COULD NOT BE OPENED AND WAS LOCKED ON TISSUE AS WELL. THE DEVICE WAS DISCONNECTED TO THE TISSUE WITH A KNIFE. IT WAS ALSO STATED THAT THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THE ISSUE RESULTED TO PERMANENT TISSUE DAMAGE. MORE TISSUES OF THE LUNG WAS CUT AND TAKEN OUT. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
26-12-2018 | Malfunction | ethicon endo-surgery | 21-01-2019 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE SECOND FIRING OF THE STAPLER, WITH A BLACK RELOAD, PERISTRIPS BUTTRESSING MATERIAL, THE JAWS OF THE INSTRUMENT WERE CLOSED, TO INTRODUCE THE DEVICE INTO THE PATIENT AND THE JAWS WOULDN'T OPEN, BEING LOCKED ON THE BUTTRESSING MATERIAL. THE DEVICE WAS REMOVED FROM THE PATIENT AND A SECOND LIKE STAPLER, RELOAD AND BUTTRESSING MATERIAL WAS USED TO CONTINUE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
26-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | ENDO GIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, UPON THE CREATION OF JEJUNOSTOMY, THE STAPLE LINE DID NOT APPEAR TO FORM PROPERLY. ANOTHER DEVICE WAS USED TO RESECT THE PART OF THE STAPLE LINE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, UPON THE CREATION OF JEJUNOSTOMY, THE STAPLE LINE DID NOT APPEAR TO FORM PROPERLY. ANOTHER DEVICE WAS USED TO RESECT THE PART OF THE STAPLE LINE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | SIGNIA | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, UPON THE CREATION OF JEJUNOSTOMY, THE STAPLE LINE DID NOT APPEAR TO FORM PROPERLY. ANOTHER DEVICE WAS USED TO RESECT THE PART OF THE STAPLE LINE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN PERFORMING THE FIRST RESECTION OF THE STOMACH POUCH IN A LAPAROSCOPIC GASTRIC BYPASS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS STATED THAT THE TRIGGER WAS LOCKED. A NEW DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | UNKNOWN EGIA SU | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN PERFORMING THE FIRST RESECTION OF THE STOMACH POUCH IN A LAPAROSCOPIC GASTRIC BYPASS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS STATED THAT THE TRIGGER WAS LOCKED. A NEW DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | ENDO GIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE FIRST FIRING IN A LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY, THE ASSEMBLY WAS PERFORMED AS USUAL AND COMPLETED UNTIL THE OPENING AND CLOSING OF THE JAWS WAS CHECKED. DURING THE USE OF THE DEVICE, WHEN THE TISSUE WAS GRASPED, AND FIRING WAS ATTEMPTED TO BE PERFORMED, THE DEVICE STOPPED WORKING SHORTLY AFTER THE KNIFE WAS ATTEMPTED TO ADVANCE BY LESS FEW MILLIMETERS. THE DOCTOR DID NOT RECOGNIZE THAT THE KNIFE HAD ADVANCED, AND THEY SET A NEW POWERED STAPLER AGAIN, ATTACHED THE CARTRIDGE, BUT AN ERROR OCCURRED. IT WAS STATED THE TISSUE WAS NOT PARTICULARLY THICK OR HARD. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE FIRST FIRING IN A LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY, THE ASSEMBLY WAS PERFORMED AS USUAL AND COMPLETED UNTIL THE OPENING AND CLOSING OF THE JAWS WAS CHECKED. DURING THE USE OF THE DEVICE, WHEN THE TISSUE WAS GRASPED, AND FIRING WAS ATTEMPTED TO BE PERFORMED, THE DEVICE STOPPED WORKING SHORTLY AFTER THE KNIFE WAS ATTEMPTED TO ADVANCE BY LESS FEW MILLIMETERS. THE DOCTOR DID NOT RECOGNIZE THAT THE KNIFE HAD ADVANCED, AND THEY SET A NEW POWERED STAPLER AGAIN, ATTACHED THE CARTRIDGE, BUT AN ERROR OCCURRED. IT WAS STATED THE TISSUE WAS NOT PARTICULARLY THICK OR HARD. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | SIGNIA | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE FIRST FIRING IN A LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY, THE ASSEMBLY WAS PERFORMED AS USUAL AND COMPLETED UNTIL THE OPENING AND CLOSING OF THE JAWS WAS CHECKED. DURING THE USE OF THE DEVICE, WHEN THE TISSUE WAS GRASPED, AND FIRING WAS ATTEMPTED TO BE PERFORMED, THE DEVICE STOPPED WORKING SHORTLY AFTER THE KNIFE WAS ATTEMPTED TO ADVANCE BY LESS FEW MILLIMETERS. THE DOCTOR DID NOT RECOGNIZE THAT THE KNIFE HAD ADVANCED, AND THEY SET A NEW POWERED STAPLER AGAIN, ATTACHED THE CARTRIDGE, BUT AN ERROR OCCURRED. IT WAS STATED THE TISSUE WAS NOT PARTICULARLY THICK OR HARD. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/ LOBECTOMY, THE RELOAD WAS USED FOR RESECTION, WHEN FIRING WAS BEING PERFORMED AT THE FOURTH FIRING, THE HANDLE WAS LOCKED AND BECAME UNABLE TO BE SQUEEZED. THERE WERE NO COMPLICATIONS SUCH AS BLEEDING DUE TO THAT FIRING HAS STOPPED BEFORE THE VESSEL RESECTION. A NEW PRODUCT WAS USED INSTEAD, AND THE OPERATION WAS CONTINUED AND COMPLETED SAFELY. THERE WAS NO PATIENT INJURY.
25-12-2018 | Malfunction | covidien lp llc north haven | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/ LOBECTOMY, THE RELOAD WAS USED FOR RESECTION, WHEN FIRING WAS BEING PERFORMED AT THE FOURTH FIRING, THE HANDLE WAS LOCKED AND BECAME UNABLE TO BE SQUEEZED. THERE WERE NO COMPLICATIONS SUCH AS BLEEDING DUE TO THAT FIRING HAS STOPPED BEFORE THE VESSEL RESECTION. A NEW PRODUCT WAS USED INSTEAD, AND THE OPERATION WAS CONTINUED AND COMPLETED SAFELY. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RECTAL CANCER PROCEDURE, WHILE USING THE DEVICE ON A BROKEN RECTUM, THE DEVICE DID NOT FIRE. THE SURGEON ABLE TO PRESSED THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. RE-CHANGING THE DEVICE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2018 | Malfunction | us surgical puerto rico | 21-01-2019 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RECTAL CANCER PROCEDURE, WHILE USING THE DEVICE ON A BROKEN RECTUM, THE DEVICE DID NOT FIRE. THE SURGEON ABLE TO PRESSED THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. RE-CHANGING THE DEVICE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | ILS 29MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN THE COURSE OF AN ANTERIOR RESECTION PROCEDURE USING CDH29A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE DID NOT PROCEED AND STAPLES ALSO WEREN'T FIRED. AS AN EMERGENCY MEASURE, THE SURGEON DETACHED IT FROM THE TISSUE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES FOR THE PATIENT SO FAR.
10-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DOCTOR WAS PERFORMING THE LEAK TEST, AFTER USING THE CIRCULAR STAPLER, AND NOTICED A MALFORMED STAPLE IN THE ANASTOMOSES. THE LEAK TEST PASSED SUCCESSFULLY AND THE DOCTOR LEFT THE MALFORMED STAPLE INTACT. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
10-12-2019 | Injury | ethicon endo-surgery | 02-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN ANTERIOR RESECTION ANASTOMOSIS THE ANTERIOR PART OF ANASTOMOSIS IS NOT COMPLETELY STAPLED. NOT MUCH BLEEDING WAS OBSERVED HOWEVER DURING CHECKING THE DONUT WAS OBSERVED TO BE INCOMPLETELY FORMED AT THE DISTAL END SO THE SURGEON PROCEEDED TO SUTURE THE TISSUE. PATIENT IS CURRENTLY ON ANTIBIOTIC AND UNDER HOSPITAL MONITORING.
10-12-2019 | Injury | ethicon endo-surgery | 02-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN ANTERIOR RESECTION ANASTOMOSIS THE ANTERIOR PART OF ANASTOMOSIS IS NOT COMPLETELY STAPLED. NOT MUCH BLEEDING WAS OBSERVED HOWEVER DURING CHECKING THE DONUT WAS OBSERVED TO BE INCOMPLETELY FORMED AT THE DISTAL END SO THE SURGEON PROCEEDED TO SUTURE THE TISSUE. PATIENT IS CURRENTLY ON ANTIBIOTIC AND UNDER HOSPITAL MONITORING.
10-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | PROXIMATE*PPH PROCEDURE SET | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON FOUND THE STAPLER HARD TO FIRE, AND DIDN'T CUT TISSUE. TRIED AGAIN THE SECOND TIME (TRIED TO FIRE STAPLER AGAIN BY PRESSING HARDER) AND STAPLER FINALLY FIRED, BUT THEY WISH TO KNOW WHAT¿S THE ISSUE WITH THE PRODUCT. NO ADVERSE EVENT TO PATIENT REPORTED. FORCE TO FIRE WAS REPORTED TO BE HIGHER THAN USUAL. SURGERY WAS DELAYED 15 MINUTES. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
10-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | PROXIMATE*PPH PROCEDURE SET | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON FOUND THE STAPLER HARD TO FIRE, AND DIDN'T CUT TISSUE. TRIED AGAIN THE SECOND TIME (TRIED TO FIRE STAPLER AGAIN BY PRESSING HARDER) AND STAPLER FINALLY FIRED, BUT THEY WISH TO KNOW WHAT¿S THE ISSUE WITH THE PRODUCT. NO ADVERSE EVENT TO PATIENT REPORTED. FORCE TO FIRE WAS REPORTED TO BE HIGHER THAN USUAL. SURGERY WAS DELAYED 15 MINUTES. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
10-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING LUNG TISSUE RESECTION IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THREE HANDLES AND FIVE RELOADS WERE OPENED IN TOTAL, UNFORTUNATELY WITHOUT SUCCESS. A DIFFERENT STAPLER WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING LUNG TISSUE RESECTION IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THREE HANDLES AND FIVE RELOADS WERE OPENED IN TOTAL, UNFORTUNATELY WITHOUT SUCCESS. A DIFFERENT STAPLER WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING LUNG TISSUE RESECTION IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THREE HANDLES AND FIVE RELOADS WERE OPENED IN TOTAL, UNFORTUNATELY WITHOUT SUCCESS. A DIFFERENT STAPLER WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING LUNG TISSUE RESECTION IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THREE HANDLES AND FIVE RELOADS WERE OPENED IN TOTAL, UNFORTUNATELY WITHOUT SUCCESS. A DIFFERENT STAPLER WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
13-11-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | PROXIMATE*75MM LINEAR CUTTER | Unintended Ejection | A050302 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, ONCE THE ORANGE TAB IS REMOVED IF THE GUN IS HELD UPSIDE DOWN STAPLES FALL OUT OF THE GUN. STAPLE IN JAR THAT FELL OUT. OCCURRED OUTSIDE OF PATIENT, NO PATIENT HARM.
17-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | EEA | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION, WHEN THE ANVIL WAS CONNECTED, AND THE HANDLE WAS INSERTED INTO THE ANUS OF THE ABDOMINAL CAVITY, ALTHOUGH THE ANVIL WAS INSERTED UNTIL THE ORANGE BAND BECAME UNABLE TO BE SEEN, IT DID NOT STOP SOLIDLY. THE SURGEON TRIED SEVERAL TIMES BUT WHEN THOUGHT THAT IT WAS CONNECTED SUCCESSFULLY AND STARTED TO WIND THE WING NUT, IT CAME OFF IMMEDIATELY. ANOTHER DEVICE WAS USED TO COMPLETE CASE. SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
07-11-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Break | A0401 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, ONE OF THE SMALL PARTS OF THE YELLOW SHIPPING WEDGE BROKE AND FELL INTO THE PATIENT'S CAVITY. THE SURGEON NOTICED IT AND WAS ABLE TO REMOVE IT.
07-11-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, ONE OF THE SMALL PARTS OF THE YELLOW SHIPPING WEDGE BROKE AND FELL INTO THE PATIENT'S CAVITY. THE SURGEON NOTICED IT AND WAS ABLE TO REMOVE IT.
13-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE SURGEON WAS ABLE TO FIRE THE THREE LOADING UNITS WITHOUT A PROBLEM. AT THE 4TH FIRING, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE LOADING UNIT COULD NOT BE FIRED. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-11-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | PROXIMATE*PPH PROCEDURE SET | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DOCTOR WAS PERFORMING A HEMORRHOIDECTOMY AND THE BLADE DID NOT ENGAGE DURING FIRING. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED, BUT SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, UPON OPENING THE ABDOMEN TO RESECT THE DUODENUM DURING A DISTAL GASTRECTOMY, THE OPENING AND CLOSING WAS CHECKED. THE GREEN BUTTON WAS THEN PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED. THE FIRING COULD NOT BE PERFORMED. ANOTHER RELOAD WAS USED THE SAME EVENT HAPPENED. A NEW RELOAD WAS USED AND FIRING WAS PERFORMED.
16-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC, ON RELOAD OF THE STAPLER, THERE WAS A PROBLEM LOADING THE DEVICE. AFTER EIGHT RELOADS USED, THE STAPLER WAS BLOCKED AND WAS IMPOSSIBLE TO USE THE DEVICE AGAIN. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY USED A NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC, ON RELOAD OF THE STAPLER, THERE WAS A PROBLEM LOADING THE DEVICE. AFTER EIGHT RELOADS USED, THE STAPLER WAS BLOCKED AND WAS IMPOSSIBLE TO USE THE DEVICE AGAIN. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY USED A NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, THE POWERED CIRCULAR STAPLER WAS UNABLE TO FIRE. THE INDICATOR WAS WITHIN RANGE AND THE SAFETY STILL ENGAGED. THE ANVIL WAS PROPERLY ATTACHED BUT THE DEVICE WOULD NOT FIRE. THE SCREEN WAS ILLUMINATED. A SECOND LIKE DEVICE WAS ATTAINED FROM ANOTHER FACILITY. THERE WAS A DELAY OF FORTY-FIVE MINUTES. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
04-11-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO A VASCULAR TISSUE DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, IT WAS REPORTED THAT THE STAPLER DID FIRE WITH THE CURVED TIP 45 TAN BUT THE STAPLES DID NOT HOLD AND STAPLES DID NOT CLOSE TIGHTLY ENOUGH, AS A RESULT THE STAPLE LINE BLED. THE SURGEON HAD TO PLACE CLIPS ALONG THE STAPLE LINE TO CONTROL BLEEDING.
04-11-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO A VASCULAR TISSUE DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, IT WAS REPORTED THAT THE STAPLER DID FIRE WITH THE CURVED TIP 45 TAN BUT THE STAPLES DID NOT HOLD AND STAPLES DID NOT CLOSE TIGHTLY ENOUGH, AS A RESULT THE STAPLE LINE BLED. THE SURGEON HAD TO PLACE CLIPS ALONG THE STAPLE LINE TO CONTROL BLEEDING.
04-11-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO A VASCULAR TISSUE DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, IT WAS REPORTED THAT THE STAPLER DID FIRE WITH THE CURVED TIP 45 TAN BUT THE STAPLES DID NOT HOLD AND STAPLES DID NOT CLOSE TIGHTLY ENOUGH, AS A RESULT THE STAPLE LINE BLED. THE SURGEON HAD TO PLACE CLIPS ALONG THE STAPLE LINE TO CONTROL BLEEDING.
17-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, UPON BLOOD VESSEL TREATMENT, THE REPORTED RELOAD WAS CONNECTED TO THE HANDLE AND THE SURGEON TRIED TO FIRE THE DEVICE, BUT SQUEEZING HANDLE WAS NOT SMOOTH AND IT WAS STIFFER THAN USUAL. NEW RELOAD WAS TAKEN AND CONTINUED THE PROCEDURE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. ALSO, AN EXCESSIVE PRESSURE WAS REQUIRED TO UNLOAD THE RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THIS RESULTED TO EXTENDED SURGICAL TIME OF MORE 30 MINUTES. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. ALSO, AN EXCESSIVE PRESSURE WAS REQUIRED TO UNLOAD THE RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THIS RESULTED TO EXTENDED SURGICAL TIME OF MORE 30 MINUTES. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. ALSO, AN EXCESSIVE PRESSURE WAS REQUIRED TO UNLOAD THE RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THIS RESULTED TO EXTENDED SURGICAL TIME OF MORE 30 MINUTES. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. ALSO, AN EXCESSIVE PRESSURE WAS REQUIRED TO UNLOAD THE RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THIS RESULTED TO EXTENDED SURGICAL TIME OF MORE 30 MINUTES. THERE WAS NO PATIENT INJURY.
16-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ESOPHAGOJEJUNAL IN GASTRECTOMY OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT HAD INCOMPLETE STAPLE LINE. THE STAPLE LINE WAS FIXED BY OVERSEWING.
11-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | PROXIMATE*PPH PROCEDURE SET | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LONGO PROCEDURE, THE DEVICE BLOCKED INTRAOPERATIVELY AND DID NOT TRIGGER. AN ADDITIONAL DEVICE WAS OPENED AND THE OPERATION ENDED WITH NO INFLUENCE ON SURGERY AND PATIENT SAFETY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, AT THE TIME OF INTERLOBAR FORMATION, THE LEAK TEST WAS ALSO COMPLETED, THERE WAS NO PROBLEM AT THAT TIME, SO WOUND CLOSING WAS PERFORMED. AFTER THAT, WHEN THE PATIENT STATUS WAS CHECKED AGAIN, ONE OUTSIDE STAPLE LINE OF RELOAD WAS NOTED TO BE LEAKED THEN REOPERATION WAS PERFORMED. THE STAPLE LINE WAS SUTURED BY MANUAL SUTURING AND THE WOUND WAS CLOSED AGAIN. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN.
17-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | EEA | Loss of or Failure to Bond | A040102 | Sepsis | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF BOWEL IN A LEFT HEMICOLECTOMY OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLE LINE HAD LEAK. THE ANASTOMOSIS WAS REMOVED AND STOMA BAG WAS PLACED. THE PATIENT HAD SEPSIS AND NEED TO HAVE ANOTHER SURGERY. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
10-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF BOWEL IN A LEFT HEMICOLECTOMY OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLE LINE HAD LEAK. THE ANASTOMOSIS WAS REMOVED AND STOMA BAG WAS PLACED. THE PATIENT HAD SEPSIS AND NEED TO HAVE ANOTHER SURGERY. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
10-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | EEA | Fluid/Blood Leak | A050401 | Sepsis | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF BOWEL IN A LEFT HEMICOLECTOMY OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLE LINE HAD LEAK. THE ANASTOMOSIS WAS REMOVED AND STOMA BAG WAS PLACED. THE PATIENT HAD SEPSIS AND NEED TO HAVE ANOTHER SURGERY. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
10-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | EEA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF BOWEL IN A LEFT HEMICOLECTOMY OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLE LINE HAD LEAK. THE ANASTOMOSIS WAS REMOVED AND STOMA BAG WAS PLACED. THE PATIENT HAD SEPSIS AND NEED TO HAVE ANOTHER SURGERY. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
11-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Activation, Positioning or Separation Problem | A15 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THERE WAS DIFFICULTY IN KEEPING THE ANVIL OF A CDH25P ATTACHED WHILST CLOSING THE GUN. MR N PLACED 3 STAY SUTURES AROUND THE OESOPHAGUS AND PULLED THEM OUT USING AN ENDOCLOSE AND CLIPPED THE SUTURE OUTSIDE THE CHEST AND THEN A PURSE STRING (WHIP STITCH) WAS MADE AROUND. A BOUGIE WAS USED TO DILATE THE OESOPHAGUS AND THEN THE ANVIL WAS INSERTED INTO THE OESOPHAGUS USING A 10AG ANVIL GRASPER. THE ANVIL WAS GRASPED AT THE GRASPING NOTCH AND ALSO AROUND THE PURSE STRING TYING NOTCH. THE LOCKING SPRING WAS AVOIDED AT ALL TIMES. THE GUN WAS PUT IN THE STOMACH AND THE SPIKE EXTENDED, A BOX STITCH WAS RUN AROUND THE BASE OF THE SPIKE (THIS IS THEIR USUAL PRACTICE TO PREVENT THE SEROSA OF THE STOMACH TEARING) THE PURSE STRING WAS ATTEMPTED TO BE TIGHTENED BUT WOULD NOT RUN. MR N TRIED TO GET THE PURSE STRING TO DRAW BUT IT WOULD NOT. THE ANVIL WAS REMOVED FROM THE OESOPHAGUS AND LEFT IN THE THORACIC CAVITY. MR I CAME IN BUT HE COULD NOT RECTIFY THE SITUATION. SO MR I DECIDED TO PLACE ANOTHER PURSE STRING. THE STAY SUTURES WERE REMOVED. THE ANVIL WAS KEPT IN THE OESOPHAGUS. A NEW STAY SUTURE WAS PLACED A FEW MMS FURTHER UP ON THE OESOPHAGUS AND WHILST HOLDING THE GRASPING NOTCH OF THE ANVIL, DIATHERMY WAS USED TO REMOVE THE TISSUE AT THE END OF THE OESOPHAGUS COMPLETE WITH THE PURSE STRING. THEN ANOTHER STAY SUTURE WAS PLACED AND CONTINUED WITH THE DIATHERMY UNTIL THERE WERE 4 STAY SUTURES AND THE PURSE STRING AND TISSUE WERE FREE AND REMOVED. THIS TOOK AT LEAST AN HOUR. MR I PUT IN ANOTHER PURSE STRING (WHIP STITCH) AND PLACED THE ANVIL IN USING JOHAN GRASPERS HOLDING AT THE GRASPING NOTCH. THE PURSE STRING DREW AS IT SHOULD. IT WAS ENSURED THAT THE PURSE STRING WAS UP ON THE PURSE STRING SUTURE TIE AREA AND NOT ON THE LOCKING SPRING. THE HEAD OF THE GUN WAS PUT THROUGH THE STOMACH AND THE SPIKE WOUND FULLY OUT. A NEW BOX STITCH WAS RUN AROUND THE BASE OF THE SPIKE. THE GUN WAS PUT IN THE CHEST AND THEN THEY ATTEMPTED TO ATTACH THE HEAD OF THE GUN. IT WAS A LITTLE TRICKY AND MR I WAS MINDFUL THAT ADJUSTING KNOB WAS CLOSING OF ITS OWN VOLITION AS HE WAS TRYING TO ATTACH THE ANVIL, SO HE HELD THE KNOB TO PREVENT THIS. THE ORANGE BAND WAS HIDDEN BY THE ANVIL AND THE THOUGHT TO HAVE CLICK ATTACHED. THE GUN WAS STARTED TO BE WOUND DOWN BUT AS IT CLOSED THE ANVIL CLICKED OFF THE SPIKE. MR I THEN ATTEMPTED TO PUT THE ANVIL ON THE SPIKE A SECOND TIME, THIS WAS NOT EASY. THE ORANGE BAND WAS COVERED. THIS TIME THE HEAD WAS DEFINITELY FELT AND HEARD TO CLICK ON. THE GUN WAS WOUND DOWN BUT ONCE AGAIN THE ANVIL DETACHED. THE SURGEON WAS EXCEPTIONALLY CONCERNED WHY THIS WOULD HAPPEN. MR I TRIED A THIRD TIME, THE ORANGE BAND WAS COVERED, THE HEAD WAS HEARD AND FELT TO CLICK ON JUST AS IT HAD THE SECOND TIME. THE GUN WAS WOUND DOWN (WHILST THIS WAS HAPPENING A THORACOTOMY TRAY WAS BEING OPENED JUST IN CASE). I AM VERY RELIEVED TO REPORT THAT THE HEAD REMAINED ATTACHED AND THE GUN WOUND DOWN AS IT SHOULD. THE INDICATOR WAS IN THE GREEN (JUST OVERHALF WAY) 15 SECONDS OBSERVED, A FINAL LITTLE TIGHTEN AND THEN THE GUN WAS FIRED. THE FIRING PROCESS WAS SMOOTH AND WITHOUT INCIDENT, AND THE GREEN TICK CONFIRMED THE FIRING SEQUENCE COMPLETE. 4 HALF TURNS COUNTERCLOCKWISE ON THE ADJUSTING KNOB AND THE GUN WAS REMOVED. THE DONUTS WERE SEEN TO BE COMPLETE. THE PROBLEMS OF GETTING THE HEAD OF THE GUN TO ATTACH AND THEN FIRE EXTENDED THE OPERATION BY ABOUT AN HOUR. I DID WONDER IF SOME TISSUE HAD MANAGED TO GO INTO THE ANVIL STALK WHILST IT WAS LEFT IN THE THORACIC CAVITY, IT SEEMED A VERY LOW POSSIBILITY. I DID LOOK IN AFTERWARDS BUT COULD SEE NO EVIDENCE OF ANY TISSUE. THROUGHOUT THE CASE I KEPT AN EYE ON THE LOCKING SPRING AND THAT THE PURSE STRING WAS TIED UP AT THE TOP AND AWAY FROM THE LOCKING SPRING. MR N HAS USED ECHELON CIRCULAR POWERED TWICE BEFORE, ONCE FOR AN OPEN OESOPHAGECTOMY AND ONCE FOR A THORACOSCOPIC OESOPHAGECTOMY. THIS WAS THE FIRST TIME MR I HAS USED THE DEVICE.
11-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THERE WAS DIFFICULTY IN KEEPING THE ANVIL OF A CDH25P ATTACHED WHILST CLOSING THE GUN. MR N PLACED 3 STAY SUTURES AROUND THE OESOPHAGUS AND PULLED THEM OUT USING AN ENDOCLOSE AND CLIPPED THE SUTURE OUTSIDE THE CHEST AND THEN A PURSE STRING (WHIP STITCH) WAS MADE AROUND. A BOUGIE WAS USED TO DILATE THE OESOPHAGUS AND THEN THE ANVIL WAS INSERTED INTO THE OESOPHAGUS USING A 10AG ANVIL GRASPER. THE ANVIL WAS GRASPED AT THE GRASPING NOTCH AND ALSO AROUND THE PURSE STRING TYING NOTCH. THE LOCKING SPRING WAS AVOIDED AT ALL TIMES. THE GUN WAS PUT IN THE STOMACH AND THE SPIKE EXTENDED, A BOX STITCH WAS RUN AROUND THE BASE OF THE SPIKE (THIS IS THEIR USUAL PRACTICE TO PREVENT THE SEROSA OF THE STOMACH TEARING) THE PURSE STRING WAS ATTEMPTED TO BE TIGHTENED BUT WOULD NOT RUN. MR N TRIED TO GET THE PURSE STRING TO DRAW BUT IT WOULD NOT. THE ANVIL WAS REMOVED FROM THE OESOPHAGUS AND LEFT IN THE THORACIC CAVITY. MR I CAME IN BUT HE COULD NOT RECTIFY THE SITUATION. SO MR I DECIDED TO PLACE ANOTHER PURSE STRING. THE STAY SUTURES WERE REMOVED. THE ANVIL WAS KEPT IN THE OESOPHAGUS. A NEW STAY SUTURE WAS PLACED A FEW MMS FURTHER UP ON THE OESOPHAGUS AND WHILST HOLDING THE GRASPING NOTCH OF THE ANVIL, DIATHERMY WAS USED TO REMOVE THE TISSUE AT THE END OF THE OESOPHAGUS COMPLETE WITH THE PURSE STRING. THEN ANOTHER STAY SUTURE WAS PLACED AND CONTINUED WITH THE DIATHERMY UNTIL THERE WERE 4 STAY SUTURES AND THE PURSE STRING AND TISSUE WERE FREE AND REMOVED. THIS TOOK AT LEAST AN HOUR. MR I PUT IN ANOTHER PURSE STRING (WHIP STITCH) AND PLACED THE ANVIL IN USING JOHAN GRASPERS HOLDING AT THE GRASPING NOTCH. THE PURSE STRING DREW AS IT SHOULD. IT WAS ENSURED THAT THE PURSE STRING WAS UP ON THE PURSE STRING SUTURE TIE AREA AND NOT ON THE LOCKING SPRING. THE HEAD OF THE GUN WAS PUT THROUGH THE STOMACH AND THE SPIKE WOUND FULLY OUT. A NEW BOX STITCH WAS RUN AROUND THE BASE OF THE SPIKE. THE GUN WAS PUT IN THE CHEST AND THEN THEY ATTEMPTED TO ATTACH THE HEAD OF THE GUN. IT WAS A LITTLE TRICKY AND MR I WAS MINDFUL THAT ADJUSTING KNOB WAS CLOSING OF ITS OWN VOLITION AS HE WAS TRYING TO ATTACH THE ANVIL, SO HE HELD THE KNOB TO PREVENT THIS. THE ORANGE BAND WAS HIDDEN BY THE ANVIL AND THE THOUGHT TO HAVE CLICK ATTACHED. THE GUN WAS STARTED TO BE WOUND DOWN BUT AS IT CLOSED THE ANVIL CLICKED OFF THE SPIKE. MR I THEN ATTEMPTED TO PUT THE ANVIL ON THE SPIKE A SECOND TIME, THIS WAS NOT EASY. THE ORANGE BAND WAS COVERED. THIS TIME THE HEAD WAS DEFINITELY FELT AND HEARD TO CLICK ON. THE GUN WAS WOUND DOWN BUT ONCE AGAIN THE ANVIL DETACHED. THE SURGEON WAS EXCEPTIONALLY CONCERNED WHY THIS WOULD HAPPEN. MR I TRIED A THIRD TIME, THE ORANGE BAND WAS COVERED, THE HEAD WAS HEARD AND FELT TO CLICK ON JUST AS IT HAD THE SECOND TIME. THE GUN WAS WOUND DOWN (WHILST THIS WAS HAPPENING A THORACOTOMY TRAY WAS BEING OPENED JUST IN CASE). I AM VERY RELIEVED TO REPORT THAT THE HEAD REMAINED ATTACHED AND THE GUN WOUND DOWN AS IT SHOULD. THE INDICATOR WAS IN THE GREEN (JUST OVERHALF WAY) 15 SECONDS OBSERVED, A FINAL LITTLE TIGHTEN AND THEN THE GUN WAS FIRED. THE FIRING PROCESS WAS SMOOTH AND WITHOUT INCIDENT, AND THE GREEN TICK CONFIRMED THE FIRING SEQUENCE COMPLETE. 4 HALF TURNS COUNTERCLOCKWISE ON THE ADJUSTING KNOB AND THE GUN WAS REMOVED. THE DONUTS WERE SEEN TO BE COMPLETE. THE PROBLEMS OF GETTING THE HEAD OF THE GUN TO ATTACH AND THEN FIRE EXTENDED THE OPERATION BY ABOUT AN HOUR. I DID WONDER IF SOME TISSUE HAD MANAGED TO GO INTO THE ANVIL STALK WHILST IT WAS LEFT IN THE THORACIC CAVITY, IT SEEMED A VERY LOW POSSIBILITY. I DID LOOK IN AFTERWARDS BUT COULD SEE NO EVIDENCE OF ANY TISSUE. THROUGHOUT THE CASE I KEPT AN EYE ON THE LOCKING SPRING AND THAT THE PURSE STRING WAS TIED UP AT THE TOP AND AWAY FROM THE LOCKING SPRING. MR N HAS USED ECHELON CIRCULAR POWERED TWICE BEFORE, ONCE FOR AN OPEN OESOPHAGECTOMY AND ONCE FOR A THORACOSCOPIC OESOPHAGECTOMY. THIS WAS THE FIRST TIME MR I HAS USED THE DEVICE.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE WAS ABLE TO FIRED; HOWEVER ONLY THE CENTRAL PART OF THE STAPLE LINE WAS STAPLED AND THE FIRST 1.5CM OF ANASTOMOSIS AND LAST 1.5CM WERE OPENED. TO RESOLVE THE ISSUE, THE SURGEON HAD TO SUTURE THE TISSUE RESULTING TO TISSUE DAMAGE. THE SURGICAL TIME ALSO WAS EXTENDED FOR MORE THAN 30 MINUTES DUE TO THE DEVICE PROBLEM.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE WAS ABLE TO FIRED; HOWEVER ONLY THE CENTRAL PART OF THE STAPLE LINE WAS STAPLED AND THE FIRST 1.5CM OF ANASTOMOSIS AND LAST 1.5CM WERE OPENED. TO RESOLVE THE ISSUE, THE SURGEON HAD TO SUTURE THE TISSUE RESULTING TO TISSUE DAMAGE. THE SURGICAL TIME ALSO WAS EXTENDED FOR MORE THAN 30 MINUTES DUE TO THE DEVICE PROBLEM.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE, WHEN THE STAPLER WAS APPLIED TO THE DISTAL END OF THE RECTUM, THE JAWS OF THE TWO BLUE CARTRIDGES AND ONE PURPLE CARTRIDGE WERE UNABLE TO CLOSE AND THERE WAS A DIFFICULTY IN UNLOADING THE CARTRIDGE. THEY HAD TO APPLY A HUGE AMOUNT OF PRESSURE TO REMOVE IT. THE DEVICE WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE SURGICAL TIME WAS EXTENDED BY THIRTY MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THEY REPLACED WITH ANOTHER PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, BEFORE CLAMPING THE TISSUE DURING WEDGE/SEGMENTAL RESECTION, AFTER THE DEVICE WAS SET UP AND CHECKED THE OPENING AND CLOSING, THE NURSE CLOSED THE CARTRIDGE AGAIN SO THAT IT WAS ABLE TO BE EASILY INSERTED FROM THE TROCAR. THE STAPLER WAS INSERTED FROM THE TROCAR AND TRIED TO BE OPENED, BUT THE CARTRIDGE DID NOT OPEN. THE PROCEDURE WAS EXTENDED BY LESS THAN 30 MINUTES. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | SIGNIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, BEFORE CLAMPING THE TISSUE DURING WEDGE/SEGMENTAL RESECTION, AFTER THE DEVICE WAS SET UP AND CHECKED THE OPENING AND CLOSING, THE NURSE CLOSED THE CARTRIDGE AGAIN SO THAT IT WAS ABLE TO BE EASILY INSERTED FROM THE TROCAR. THE STAPLER WAS INSERTED FROM THE TROCAR AND TRIED TO BE OPENED, BUT THE CARTRIDGE DID NOT OPEN. THE PROCEDURE WAS EXTENDED BY LESS THAN 30 MINUTES. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | SIGNIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, BEFORE CLAMPING THE TISSUE DURING WEDGE/SEGMENTAL RESECTION, AFTER THE DEVICE WAS SET UP AND CHECKED THE OPENING AND CLOSING, THE NURSE CLOSED THE CARTRIDGE AGAIN SO THAT IT WAS ABLE TO BE EASILY INSERTED FROM THE TROCAR. THE STAPLER WAS INSERTED FROM THE TROCAR AND TRIED TO BE OPENED, BUT THE CARTRIDGE DID NOT OPEN. THE PROCEDURE WAS EXTENDED BY LESS THAN 30 MINUTES. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-11-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | TA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR BOWEL RESECTION OPEN PROCEDURE, THE SURGEON TRIED TO SQUEEZE THE HANDLE BUT IT WOULD NOT MOVE AND THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED BUT SAME THING HAPPENED. ANOTHER PRE-LOADED STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. ALSO, AN EXCESSIVE PRESSURE WAS REQUIRED TO UNLOAD THE RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THIS RESULTED TO EXTENDED SURGICAL TIME OF MORE 30 MINUTES. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. ALSO, AN EXCESSIVE PRESSURE WAS REQUIRED TO UNLOAD THE RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THIS RESULTED TO EXTENDED SURGICAL TIME OF MORE 30 MINUTES. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. ALSO, AN EXCESSIVE PRESSURE WAS REQUIRED TO UNLOAD THE RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THIS RESULTED TO EXTENDED SURGICAL TIME OF MORE 30 MINUTES. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. ALSO, AN EXCESSIVE PRESSURE WAS REQUIRED TO UNLOAD THE RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THIS RESULTED TO EXTENDED SURGICAL TIME OF MORE 30 MINUTES. THERE WAS NO PATIENT INJURY.
27-11-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | TA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR BOWEL RESECTION OPEN PROCEDURE, THE SURGEON TRIED TO SQUEEZE THE HANDLE BUT IT WOULD NOT MOVE AND THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED BUT SAME THING HAPPENED. ANOTHER PRE-LOADED STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | PREMIUM PLUS CEEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ANTERIOR RECTAL WALL IN A LAPAROSCOPIC PROCEDURE, THE COMPONENT WAS DISENGAGED INTO PATIENT'S CAVITY AND IT WAS STATED THAT A HANDLE WAS TURNED TO EXTERIORIZE THE PEAK AND IT WAS RATHER THE STAPLES WHICH CROSSED THE ANTERIOR WALL OF THE RECTUM. THE DEVICE WAS THEN REMOVED VIA ABDOMEN AND STAPLES RECTALLY. ANASTOMOSIS WAS PERFORMED WITH ANOTHER MACHINE AND SUTURED A LAYER OVER THE LINE ANTERIOR CLIP. A PROTECTIVE ILEOSTOMY WAS ALREADY PLANNED FOR THIS PATIENT. IT WAS IN ILEUS BUT THERE WAS NO SIGN OF ANASTOMOTIC LEAK.
11-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | PREMIUM PLUS CEEA | Failure to Form Staple | A050704 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ANTERIOR RECTAL WALL IN A LAPAROSCOPIC PROCEDURE, THE COMPONENT WAS DISENGAGED INTO PATIENT'S CAVITY AND IT WAS STATED THAT A HANDLE WAS TURNED TO EXTERIORIZE THE PEAK AND IT WAS RATHER THE STAPLES WHICH CROSSED THE ANTERIOR WALL OF THE RECTUM. THE DEVICE WAS THEN REMOVED VIA ABDOMEN AND STAPLES RECTALLY. ANASTOMOSIS WAS PERFORMED WITH ANOTHER MACHINE AND SUTURED A LAYER OVER THE LINE ANTERIOR CLIP. A PROTECTIVE ILEOSTOMY WAS ALREADY PLANNED FOR THIS PATIENT. IT WAS IN ILEUS BUT THERE WAS NO SIGN OF ANASTOMOTIC LEAK.
11-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | PREMIUM PLUS CEEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ANTERIOR RECTAL WALL IN A LAPAROSCOPIC PROCEDURE, THE COMPONENT WAS DISENGAGED INTO PATIENT'S CAVITY AND IT WAS STATED THAT A HANDLE WAS TURNED TO EXTERIORIZE THE PEAK AND IT WAS RATHER THE STAPLES WHICH CROSSED THE ANTERIOR WALL OF THE RECTUM. THE DEVICE WAS THEN REMOVED VIA ABDOMEN AND STAPLES RECTALLY. ANASTOMOSIS WAS PERFORMED WITH ANOTHER MACHINE AND SUTURED A LAYER OVER THE LINE ANTERIOR CLIP. A PROTECTIVE ILEOSTOMY WAS ALREADY PLANNED FOR THIS PATIENT. IT WAS IN ILEUS BUT THERE WAS NO SIGN OF ANASTOMOTIC LEAK.
11-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | PREMIUM PLUS CEEA | Detachment of Device or Device Component | A0501 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ANTERIOR RECTAL WALL IN A LAPAROSCOPIC PROCEDURE, THE COMPONENT WAS DISENGAGED INTO PATIENT'S CAVITY AND IT WAS STATED THAT A HANDLE WAS TURNED TO EXTERIORIZE THE PEAK AND IT WAS RATHER THE STAPLES WHICH CROSSED THE ANTERIOR WALL OF THE RECTUM. THE DEVICE WAS THEN REMOVED VIA ABDOMEN AND STAPLES RECTALLY. ANASTOMOSIS WAS PERFORMED WITH ANOTHER MACHINE AND SUTURED A LAYER OVER THE LINE ANTERIOR CLIP. A PROTECTIVE ILEOSTOMY WAS ALREADY PLANNED FOR THIS PATIENT. IT WAS IN ILEUS BUT THERE WAS NO SIGN OF ANASTOMOTIC LEAK.
11-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | PREMIUM PLUS CEEA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ANTERIOR RECTAL WALL IN A LAPAROSCOPIC PROCEDURE, THE COMPONENT WAS DISENGAGED INTO PATIENT'S CAVITY AND IT WAS STATED THAT A HANDLE WAS TURNED TO EXTERIORIZE THE PEAK AND IT WAS RATHER THE STAPLES WHICH CROSSED THE ANTERIOR WALL OF THE RECTUM. THE DEVICE WAS THEN REMOVED VIA ABDOMEN AND STAPLES RECTALLY. ANASTOMOSIS WAS PERFORMED WITH ANOTHER MACHINE AND SUTURED A LAYER OVER THE LINE ANTERIOR CLIP. A PROTECTIVE ILEOSTOMY WAS ALREADY PLANNED FOR THIS PATIENT. IT WAS IN ILEUS BUT THERE WAS NO SIGN OF ANASTOMOTIC LEAK.
11-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | PREMIUM PLUS CEEA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ANTERIOR RECTAL WALL IN A LAPAROSCOPIC PROCEDURE, THE COMPONENT WAS DISENGAGED INTO PATIENT'S CAVITY AND IT WAS STATED THAT A HANDLE WAS TURNED TO EXTERIORIZE THE PEAK AND IT WAS RATHER THE STAPLES WHICH CROSSED THE ANTERIOR WALL OF THE RECTUM. THE DEVICE WAS THEN REMOVED VIA ABDOMEN AND STAPLES RECTALLY. ANASTOMOSIS WAS PERFORMED WITH ANOTHER MACHINE AND SUTURED A LAYER OVER THE LINE ANTERIOR CLIP. A PROTECTIVE ILEOSTOMY WAS ALREADY PLANNED FOR THIS PATIENT. IT WAS IN ILEUS BUT THERE WAS NO SIGN OF ANASTOMOTIC LEAK.
11-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | PREMIUM PLUS CEEA | Device Contamination with Chemical or Other Material | A180104 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ANTERIOR RECTAL WALL IN A LAPAROSCOPIC PROCEDURE, THE COMPONENT WAS DISENGAGED INTO PATIENT'S CAVITY AND IT WAS STATED THAT A HANDLE WAS TURNED TO EXTERIORIZE THE PEAK AND IT WAS RATHER THE STAPLES WHICH CROSSED THE ANTERIOR WALL OF THE RECTUM. THE DEVICE WAS THEN REMOVED VIA ABDOMEN AND STAPLES RECTALLY. ANASTOMOSIS WAS PERFORMED WITH ANOTHER MACHINE AND SUTURED A LAYER OVER THE LINE ANTERIOR CLIP. A PROTECTIVE ILEOSTOMY WAS ALREADY PLANNED FOR THIS PATIENT. IT WAS IN ILEUS BUT THERE WAS NO SIGN OF ANASTOMOTIC LEAK.
11-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | PREMIUM PLUS CEEA | Device Contamination with Chemical or Other Material | A180104 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ANTERIOR RECTAL WALL IN A LAPAROSCOPIC PROCEDURE, THE COMPONENT WAS DISENGAGED INTO PATIENT'S CAVITY AND IT WAS STATED THAT A HANDLE WAS TURNED TO EXTERIORIZE THE PEAK AND IT WAS RATHER THE STAPLES WHICH CROSSED THE ANTERIOR WALL OF THE RECTUM. THE DEVICE WAS THEN REMOVED VIA ABDOMEN AND STAPLES RECTALLY. ANASTOMOSIS WAS PERFORMED WITH ANOTHER MACHINE AND SUTURED A LAYER OVER THE LINE ANTERIOR CLIP. A PROTECTIVE ILEOSTOMY WAS ALREADY PLANNED FOR THIS PATIENT. IT WAS IN ILEUS BUT THERE WAS NO SIGN OF ANASTOMOTIC LEAK.
11-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | PREMIUM PLUS CEEA | Device Contamination with Chemical or Other Material | A180104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ANTERIOR RECTAL WALL IN A LAPAROSCOPIC PROCEDURE, THE COMPONENT WAS DISENGAGED INTO PATIENT'S CAVITY AND IT WAS STATED THAT A HANDLE WAS TURNED TO EXTERIORIZE THE PEAK AND IT WAS RATHER THE STAPLES WHICH CROSSED THE ANTERIOR WALL OF THE RECTUM. THE DEVICE WAS THEN REMOVED VIA ABDOMEN AND STAPLES RECTALLY. ANASTOMOSIS WAS PERFORMED WITH ANOTHER MACHINE AND SUTURED A LAYER OVER THE LINE ANTERIOR CLIP. A PROTECTIVE ILEOSTOMY WAS ALREADY PLANNED FOR THIS PATIENT. IT WAS IN ILEUS BUT THERE WAS NO SIGN OF ANASTOMOTIC LEAK.
16-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN THE SURGEON WAS FIRING OVER STOMACH DURING TRANSECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE RELOAD WAS CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN OPEN PROCEDURE, THE STAPLER WAS FIRED 2 TIMES WITH 2 RELOADS. THE FIRST TIME THE STAPLER WAS OK. THE SECOND TIME THE CLIPS DID NOT BEND CORRECTLY (THE CLIPS LEGS ARE NOT BENT). THERE WAS NO HARM DONE TO THE PATIENT. I HAVE AT THIS TIME NOT ANY MORE INFORMATION.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | SIGNIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING THE STOMACH DURING A LAPAROSCOPIC GASTRIC SLEEVE, THE RELOAD WOULD NOT FIRE. THEY TESTED THE DEVICE AND FIRED ONE STAPLE LOAD WITH THE ADAPTER AND THE LOAD FIRED COMPLETELY BUT SLOWER THAN NORMAL. ANOTHER ATTEMPT WAS MADE TO FIRE ANOTHER LOAD AND IT WOULD NOT FIRE. IN ADDITION, THE DEVICE ACTED AS IF IT WOULD FIRE BUT SLOWER THAN NORMAL. SAME THING HAPPENED WHEN THE SURGEON FIRED THE DEVICE DURING PROCEDURE EXCEPT ABOUT 10MM OF STAPLES FIRED BEFORE THE BLADE STOPPED. IT WAS ALSO STATED THAT THERE WAS STAPLE LINE INCOMPLETE PROXIMALLY. THE LINE WAS RE-INITIATED WITH A NEW STAPLER, ADAPTER AND RELOAD. THERE WAS NO PATIENT INJURY.
18-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | EEA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOIDECTOMY LAPAROSCOPIC PROCEDURE, THE DEVICE CUT BUT DID NOT DEPLOY STAPLES. THE SURGICAL TIME WAS EXTENDED 30 MINUTES AND MORE DUE TO PRODUCT PROBLEM WHICH LEADS TO LONGER EXPOSURE TO ANESTHESIA. A SMALL PORTION OF INTESTINE WAS LOSS. MANUAL ANASTOMOSIS WAS PERFORMED TO COMPLETE THE CASE.
18-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | EEA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOIDECTOMY LAPAROSCOPIC PROCEDURE, THE DEVICE CUT BUT DID NOT DEPLOY STAPLES. THE SURGICAL TIME WAS EXTENDED 30 MINUTES AND MORE DUE TO PRODUCT PROBLEM WHICH LEADS TO LONGER EXPOSURE TO ANESTHESIA. A SMALL PORTION OF INTESTINE WAS LOSS. MANUAL ANASTOMOSIS WAS PERFORMED TO COMPLETE THE CASE.
09-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO ANASTOMOSIS LUNG TISSUE DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, WHEN THE SURGEON USED THE FIRST RELOAD AN INCOMPLETE TISSUE ANASTOMOSIS AND STAPLE LINE LEAKAGE WERE OBSERVED. THEN THE SURGEON USED THE SECOND RELOAD AT THE LEAKAGE LOCATION BUT THE STAPLE LINE WAS INCOMPLETE. THE SURGEON OPENED AND USED A NEW HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Gas/Air Leak | A050402 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO ANASTOMOSIS LUNG TISSUE DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, WHEN THE SURGEON USED THE FIRST RELOAD AN INCOMPLETE TISSUE ANASTOMOSIS AND STAPLE LINE LEAKAGE WERE OBSERVED. THEN THE SURGEON USED THE SECOND RELOAD AT THE LEAKAGE LOCATION BUT THE STAPLE LINE WAS INCOMPLETE. THE SURGEON OPENED AND USED A NEW HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO ANASTOMOSIS LUNG TISSUE DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, WHEN THE SURGEON USED THE FIRST RELOAD AN INCOMPLETE TISSUE ANASTOMOSIS AND STAPLE LINE LEAKAGE WERE OBSERVED. THEN THE SURGEON USED THE SECOND RELOAD AT THE LEAKAGE LOCATION BUT THE STAPLE LINE WAS INCOMPLETE. THE SURGEON OPENED AND USED NEW HANDLE AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | EEA | Material Twisted/Bent | A040609 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, 6 DAYS FOLLOWING A LAPAROSCOPIC MYOMECTOMY, IT WAS STATED THAT THE DEVICE WORKED NORMALLY WITH ITS NORMAL FEATURES, AS OPENING THE GREEN SAFETY BUTTON SENSING DURING CLIPPING AND THE GIVING CLIPPING. IT WAS ALSO REPORTED THAT THE ANVIL WAS BENT. THE SURGERY WAS PERFORMED ON 9TH OF DECEMBER AND ON 17TH OF SAME MONTH. ON THE 6TH DUST MADE A FISTULA 70% OF ANASTOMOSIS CAUSING THE PATIENT MORE SERIOUS CONDITION. THERE WAS A TISSUE DAMAGE, ADDITIONAL OPERATION, EXTENDED HOSPITALIZATION AND THE PATIENT WAS CONFINED TO INTENSIVE CARE UNIT AS A RESULT OF THE EVENT.
09-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | EEA | Material Twisted/Bent | A040609 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 6 DAYS FOLLOWING A LAPAROSCOPIC MYOMECTOMY, IT WAS STATED THAT THE DEVICE WORKED NORMALLY WITH ITS NORMAL FEATURES, AS OPENING THE GREEN SAFETY BUTTON SENSING DURING CLIPPING AND THE GIVING CLIPPING. IT WAS ALSO REPORTED THAT THE ANVIL WAS BENT. THE SURGERY WAS PERFORMED ON 9TH OF DECEMBER AND ON 17TH OF SAME MONTH. ON THE 6TH DUST MADE A FISTULA 70% OF ANASTOMOSIS CAUSING THE PATIENT MORE SERIOUS CONDITION. THERE WAS A TISSUE DAMAGE, ADDITIONAL OPERATION, EXTENDED HOSPITALIZATION AND THE PATIENT WAS CONFINED TO INTENSIVE CARE UNIT AS A RESULT OF THE EVENT.
16-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY PROCEDURE, DURING TRANSECTION OF STOMACH, SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON, PROCEED FIRING BY SQUEEZING HANDLE, HOWEVER CANNOT PROCEED FURTHER FIRING AS THE HANDLE CANNOT BE SQUEEZED AT THE MIDDLE. SURGEON CHANGED TO ANOTHER NEW DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. NO PATIENT INJURY.
16-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY PROCEDURE, DURING TRANSECTION OF STOMACH, SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON, PROCEED FIRING BY SQUEEZING HANDLE, HOWEVER CANNOT PROCEED FURTHER FIRING AS THE HANDLE CANNOT BE SQUEEZED AT THE MIDDLE. SURGEON CHANGED TO ANOTHER NEW DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC TRANSIT BIPARTITION, THE SURGEON BEGAN TO STAPLE THE STOMACH STARTING FROM THE ANTRUM. SECONDLY CAME TO ¿ANGULARUSA¿, THE EMBEZZLEMENTS IN THE LAST PART OF THE RELOAD DID NOT RECEIVE A B FORMATION. THE SURGEON CUT THE TISSUE BUT DID NOT STAPLING. THE SURGEON COMPLETED THE SURGERY WITH OTHER PRODUCTS. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE ABLE TO BE FIRED, BUT WHEN THE SURGEON TRYING TO RETRACT THE KNIFE THE DEVICE LOCKED ON TISSUE. TO RESOLVE THE ISSUE THE SURGEON THEN HAD TO FORCEFULLY PULL THE RETURN KNOB. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE ABLE TO BE FIRED, BUT WHEN THE SURGEON TRYING TO RETRACT THE KNIFE THE DEVICE LOCKED ON TISSUE. TO RESOLVE THE ISSUE THE SURGEON THEN HAD TO FORCEFULLY PULL THE RETURN KNOB. THERE WAS NO PATIENT INJURY.
02-01-2020 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Loss of or Failure to Bond | A040102 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, 4 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC GIST REMOVAL, THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE DISTALLY. IT WAS STATED BY THE SURGEON THAT THE PATIENT HAD STAPLE LINE PERFORATION. A SECOND PROCEDURE WAS NEEDED TO OVERSEW THE STAPLE LINE LAPAROSCOPICALLY DUE TO PERITONITIS. THE PATIENT EXPERIENCED HAD EXTENDED HOSPITALIZATION AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. IT WAS ALSO NOTED THAT THE SURGICAL PROCEDURE WAS EXTENDED FOR 30 MINUTES.
02-01-2020 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Loss of or Failure to Bond | A040102 | Peritonitis | Event Description: ACCORDING TO THE REPORTER, 4 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC GIST REMOVAL, THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE DISTALLY. IT WAS STATED BY THE SURGEON THAT THE PATIENT HAD STAPLE LINE PERFORATION. A SECOND PROCEDURE WAS NEEDED TO OVERSEW THE STAPLE LINE LAPAROSCOPICALLY DUE TO PERITONITIS. THE PATIENT EXPERIENCED HAD EXTENDED HOSPITALIZATION AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. IT WAS ALSO NOTED THAT THE SURGICAL PROCEDURE WAS EXTENDED FOR 30 MINUTES.
02-01-2020 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 4 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC GIST REMOVAL, THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE DISTALLY. IT WAS STATED BY THE SURGEON THAT THE PATIENT HAD STAPLE LINE PERFORATION. A SECOND PROCEDURE WAS NEEDED TO OVERSEW THE STAPLE LINE LAPAROSCOPICALLY DUE TO PERITONITIS. THE PATIENT EXPERIENCED HAD EXTENDED HOSPITALIZATION AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. IT WAS ALSO NOTED THAT THE SURGICAL PROCEDURE WAS EXTENDED FOR 30 MINUTES.
02-01-2020 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, 4 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC GIST REMOVAL, THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE DISTALLY. IT WAS STATED BY THE SURGEON THAT THE PATIENT HAD STAPLE LINE PERFORATION. A SECOND PROCEDURE WAS NEEDED TO OVERSEW THE STAPLE LINE LAPAROSCOPICALLY DUE TO PERITONITIS. THE PATIENT EXPERIENCED HAD EXTENDED HOSPITALIZATION AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. IT WAS ALSO NOTED THAT THE SURGICAL PROCEDURE WAS EXTENDED FOR 30 MINUTES.
02-01-2020 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Misfire | A050502 | Peritonitis | Event Description: ACCORDING TO THE REPORTER, 4 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC GIST REMOVAL, THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE DISTALLY. IT WAS STATED BY THE SURGEON THAT THE PATIENT HAD STAPLE LINE PERFORATION. A SECOND PROCEDURE WAS NEEDED TO OVERSEW THE STAPLE LINE LAPAROSCOPICALLY DUE TO PERITONITIS. THE PATIENT EXPERIENCED HAD EXTENDED HOSPITALIZATION AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. IT WAS ALSO NOTED THAT THE SURGICAL PROCEDURE WAS EXTENDED FOR 30 MINUTES.
02-01-2020 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 4 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC GIST REMOVAL, THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE DISTALLY. IT WAS STATED BY THE SURGEON THAT THE PATIENT HAD STAPLE LINE PERFORATION. A SECOND PROCEDURE WAS NEEDED TO OVERSEW THE STAPLE LINE LAPAROSCOPICALLY DUE TO PERITONITIS. THE PATIENT EXPERIENCED HAD EXTENDED HOSPITALIZATION AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. IT WAS ALSO NOTED THAT THE SURGICAL PROCEDURE WAS EXTENDED FOR 30 MINUTES.
02-01-2020 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, 4 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC GIST REMOVAL, THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE DISTALLY. IT WAS STATED BY THE SURGEON THAT THE PATIENT HAD STAPLE LINE PERFORATION. A SECOND PROCEDURE WAS NEEDED TO OVERSEW THE STAPLE LINE LAPAROSCOPICALLY DUE TO PERITONITIS. THE PATIENT EXPERIENCED HAD EXTENDED HOSPITALIZATION AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. IT WAS ALSO NOTED THAT THE SURGICAL PROCEDURE WAS EXTENDED FOR 30 MINUTES.
02-01-2020 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Peritonitis | Event Description: ACCORDING TO THE REPORTER, 4 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC GIST REMOVAL, THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE DISTALLY. IT WAS STATED BY THE SURGEON THAT THE PATIENT HAD STAPLE LINE PERFORATION. A SECOND PROCEDURE WAS NEEDED TO OVERSEW THE STAPLE LINE LAPAROSCOPICALLY DUE TO PERITONITIS. THE PATIENT EXPERIENCED HAD EXTENDED HOSPITALIZATION AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. IT WAS ALSO NOTED THAT THE SURGICAL PROCEDURE WAS EXTENDED FOR 30 MINUTES.
02-01-2020 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 4 DAYS POST-OPERATIVELY ON A LAPAROSCOPIC GASTRIC GIST REMOVAL, THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE DISTALLY. IT WAS STATED BY THE SURGEON THAT THE PATIENT HAD STAPLE LINE PERFORATION. A SECOND PROCEDURE WAS NEEDED TO OVERSEW THE STAPLE LINE LAPAROSCOPICALLY DUE TO PERITONITIS. THE PATIENT EXPERIENCED HAD EXTENDED HOSPITALIZATION AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. IT WAS ALSO NOTED THAT THE SURGICAL PROCEDURE WAS EXTENDED FOR 30 MINUTES.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING THE STOMACH DURING A LAPAROSCOPIC GASTRIC SLEEVE, THE RELOAD WOULD NOT FIRE. THEY TESTED THE DEVICE AND FIRED ONE STAPLE LOAD WITH THE ADAPTER AND THE LOAD FIRED COMPLETELY BUT SLOWER THAN NORMAL. ANOTHER ATTEMPT WAS MADE TO FIRE ANOTHER LOAD AND IT WOULD NOT FIRE. IN ADDITION, THE DEVICE ACTED AS IF IT WOULD FIRE BUT SLOWER THAN NORMAL. SAME THING HAPPENED WHEN THE SURGEON FIRED THE DEVICE DURING PROCEDURE EXCEPT ABOUT 10MM OF STAPLES FIRED BEFORE THE BLADE STOPPED. IT WAS ALSO STATED THAT THERE WAS STAPLE LINE INCOMPLETE PROXIMALLY. THE LINE WAS RE-INITIATED WITH A NEW STAPLER, ADAPTER AND RELOAD. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING THE STOMACH DURING A LAPAROSCOPIC GASTRIC SLEEVE, THE RELOAD WOULD NOT FIRE. THEY TESTED THE DEVICE AND FIRED ONE STAPLE LOAD WITH THE ADAPTER AND THE LOAD FIRED COMPLETELY BUT SLOWER THAN NORMAL. ANOTHER ATTEMPT WAS MADE TO FIRE ANOTHER LOAD AND IT WOULD NOT FIRE. IN ADDITION, THE DEVICE ACTED AS IF IT WOULD FIRE BUT SLOWER THAN NORMAL. SAME THING HAPPENED WHEN THE SURGEON FIRED THE DEVICE DURING PROCEDURE EXCEPT ABOUT 10MM OF STAPLES FIRED BEFORE THE BLADE STOPPED. IT WAS ALSO STATED THAT THERE WAS STAPLE LINE INCOMPLETE PROXIMALLY. THE LINE WAS RE-INITIATED WITH A NEW STAPLER, ADAPTER AND RELOAD. THERE WAS NO PATIENT INJURY.
11-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE PRODUCT WAS BEING USED AT THE BEGINNING OF A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO WAS PRESS THE GREEN BUTTON BUT THE DEVICE COULD NOT BE FIRED. ANOTHER HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
29-07-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | EEA | Failure to Align | A050701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF INTESTINE IN COLORECTAL LAPAROSCOPIC PROCEDURE, THE STAPLER GOT SKEWED AFTER THE BASE WAS CONNECTED WITH THE STAPLING ROD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE PRODUCT WAS BEING USED AT THE BEGINNING OF A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO WAS PRESS THE GREEN BUTTON BUT THE DEVICE COULD NOT BE FIRED. ANOTHER HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Pulmonary Emphysema | Event Description: ACCORDING TO THE REPORTER, THREE DAYS AFTER A LAPAROSCOPIC SEGMENTAL RESECTION AND AFTER A SESSION OF KINESIOLOGY WITH HYPERPRESSION, THE PATIENT SUFFERED SUBCUTANEOUS EMPHYSEMA. REINTERVENTION WAS DONE AND A DRAINAGE WAS INSTALLED. THE EVENT LEAD TO OR EXTEND PATIENT HOSPITALIZATION.
15-12-2019 | Injury | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, THREE DAYS AFTER A LAPAROSCOPIC SEGMENTAL RESECTION AND AFTER A SESSION OF KINESIOLOGY WITH HYPERPRESSION, THE PATIENT SUFFERED SUBCUTANEOUS EMPHYSEMA. REINTERVENTION WAS DONE AND A DRAINAGE WAS INSTALLED. THE EVENT LEAD TO OR EXTEND PATIENT HOSPITALIZATION.
13-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED. HENCE, THE DEVICE DID NOT FIRE.
19-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN DEALING WITH THE MIDDLE HEPATIC VEIN DURING A LAPAROSCOPIC PARTIAL HEPATECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE GUT DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS ALSO STATED THAT THE RELOAD BROKE OFF. ANOTHER RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE GUT DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS ALSO STATED THAT THE RELOAD BROKE OFF. ANOTHER RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
13-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED. HENCE, THE DEVICE DID NOT FIRE.
16-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, UPON OPENING THE ABDOMEN TO RESECT THE DUODENUM DURING A DISTAL GASTRECTOMY, THE OPENING AND CLOSING WAS CHECKED. THE GREEN BUTTON WAS THEN PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED. THE FIRING COULD NOT BE PERFORMED. ANOTHER RELOAD WAS USED THE SAME EVENT HAPPENED. A NEW RELOAD WAS USED AND FIRING WAS PERFORMED.
10-12-2019 | Injury | us surgical puerto rico | 02-01-2020 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, A POST-OPERATIVE BLEEDING OCCURRED.
19-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Injury | us surgical puerto rico | 02-01-2020 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, A POST-OPERATIVE BLEEDING OCCURRED.
06-12-2019 | Injury | ethicon endo-surgery | 02-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT THE DOCTOR PERFORMED A LAPAROSCOPIC GASTRIC BYPASS ON A PATIENT ON (B)(6) 2019, USING A PLEE60A STAPLER AND GST60W RELOAD, NO BUTTRESSING MATERIAL. THE PATIENT RETURNED ON (B)(6) 2019 WITH A CLOT AT THE JEJUNOSTOMY. THE PATIENT WAS TAKEN BACK TO SURGERY, A SCAN SUCTION WAS PERFORMED. THE PATIENT IS OKAY.
06-12-2019 | Injury | ethicon endo-surgery | 02-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT THE DOCTOR PERFORMED A LAPAROSCOPIC GASTRIC BYPASS ON A PATIENT ON (B)(6) 2019, USING A PLEE60A STAPLER AND GST60W RELOAD, NO BUTTRESSING MATERIAL. THE PATIENT RETURNED ON (B)(6) 2019 WITH A CLOT AT THE JEJUNOSTOMY. THE PATIENT WAS TAKEN BACK TO SURGERY, A SCAN SUCTION WAS PERFORMED. THE PATIENT IS OKAY.
29-10-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | POWERED PLUS COMPACT 45 | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON WAS USING POWER ECHELON FOR HIS LOBECTOMY AND THE DEVICE STOPPED WORKING AFTER 11 ACTIVATIONS. THEY OPENED ANOTHER DEVICE FOR THE LAST ACTIVATION. THERE WERE NO PATIENT CONSEQUENCES REPORTED, AND IT'S UNKNOWN WHETHER THE SURGERY WAS DELAYED OR SUCCESSFULLY COMPLETED. IT'S UNKNOWN WHETHER ANY OTHER MEDICAL INTERVENTION WAS REQUIRED.
18-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Mechanics Altered | A0507 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A BILROTH I PROCEDURE, THE STAPLER DID CUT THE TISSUE BUT THE CARTRIDGE GOT STUCK AFTER FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-11-2019 | Malfunction | davol inc sub c.r bard | 02-01-2020 | OPTIFIX 30 CT STRAIGHT 39CM | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS INITIALLY REPORTED THAT THE OPTIFIX FIXATION DEVICE FIRED CORRECTLY ONCE DURING A VENTRAL HERNIA REPAIR, THEN MISFIRED 4 TIMES. THE DEVICE WAS REMOVED FROM THE PORT, AND THE CLINICIAN REMOVED THE TACK, AND IT THEN PROCEEDED TO FIRE CORRECTLY. THERE WAS NO REPORTED PATIENT INJURY. DURING THE DEVICE EVALUATION IT WAS NOTED THAT ONE OF THE BARREL INSERT AT THE DISTAL TIP HAD DETACHED ALONG WITH ONE OF FOUR RETENTION TABS. THE MISSING PIECES WERE NOT INCLUDED WITH THE DEVICE. IN FOLLOW UP, IT WAS CONFIRMED THAT THE OPTIFIX DEVICE HAD BEEN REMOVED FROM THE PATIENT IN AN ATTEMPT TO CLEAR THE DEVICE. THE DAMAGE TO THE DISTAL TIP AND DETACHMENT OF COMPONENTS PRESENTED AFTER BEING REMOVED FROM THE BODY. THE BARREL INSERT AND TAB WERE NOT BROKEN IN THE BODY. CONTACT REPORTED THAT BROKEN FASTENERS WERE REMOVED FROM THE ABDOMINAL CAVITY DURING THE PROCEDURE. IT WAS NOT PREVIOUSLY REPORTED THAT FASTENERS WERE BROKEN IN USE. THE CASE WAS COMPLETED AND THERE WAS NO PATIENT INJURY AS A RESULT OF THE PROBLEM EXPERIENCED.
04-11-2019 | Malfunction | davol inc sub c.r bard | 02-01-2020 | OPTIFIX 30 CT STRAIGHT 39CM | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS INITIALLY REPORTED THAT THE OPTIFIX FIXATION DEVICE FIRED CORRECTLY ONCE DURING A VENTRAL HERNIA REPAIR, THEN MISFIRED 4 TIMES. THE DEVICE WAS REMOVED FROM THE PORT, AND THE CLINICIAN REMOVED THE TACK, AND IT THEN PROCEEDED TO FIRE CORRECTLY. THERE WAS NO REPORTED PATIENT INJURY. DURING THE DEVICE EVALUATION IT WAS NOTED THAT ONE OF THE BARREL INSERT AT THE DISTAL TIP HAD DETACHED ALONG WITH ONE OF FOUR RETENTION TABS. THE MISSING PIECES WERE NOT INCLUDED WITH THE DEVICE. IN FOLLOW UP, IT WAS CONFIRMED THAT THE OPTIFIX DEVICE HAD BEEN REMOVED FROM THE PATIENT IN AN ATTEMPT TO CLEAR THE DEVICE. THE DAMAGE TO THE DISTAL TIP AND DETACHMENT OF COMPONENTS PRESENTED AFTER BEING REMOVED FROM THE BODY. THE BARREL INSERT AND TAB WERE NOT BROKEN IN THE BODY. CONTACT REPORTED THAT BROKEN FASTENERS WERE REMOVED FROM THE ABDOMINAL CAVITY DURING THE PROCEDURE. IT WAS NOT PREVIOUSLY REPORTED THAT FASTENERS WERE BROKEN IN USE. THE CASE WAS COMPLETED AND THERE WAS NO PATIENT INJURY AS A RESULT OF THE PROBLEM EXPERIENCED.
04-11-2019 | Malfunction | davol inc sub c.r bard | 02-01-2020 | OPTIFIX 30 CT STRAIGHT 39CM | Material Twisted/Bent | A040609 | No Consequences Or Impact To Patient | Event Description: IT WAS INITIALLY REPORTED THAT THE OPTIFIX FIXATION DEVICE FIRED CORRECTLY ONCE DURING A VENTRAL HERNIA REPAIR, THEN MISFIRED 4 TIMES. THE DEVICE WAS REMOVED FROM THE PORT, AND THE CLINICIAN REMOVED THE TACK, AND IT THEN PROCEEDED TO FIRE CORRECTLY. THERE WAS NO REPORTED PATIENT INJURY. DURING THE DEVICE EVALUATION IT WAS NOTED THAT ONE OF THE BARREL INSERT AT THE DISTAL TIP HAD DETACHED ALONG WITH ONE OF FOUR RETENTION TABS. THE MISSING PIECES WERE NOT INCLUDED WITH THE DEVICE. IN FOLLOW UP, IT WAS CONFIRMED THAT THE OPTIFIX DEVICE HAD BEEN REMOVED FROM THE PATIENT IN AN ATTEMPT TO CLEAR THE DEVICE. THE DAMAGE TO THE DISTAL TIP AND DETACHMENT OF COMPONENTS PRESENTED AFTER BEING REMOVED FROM THE BODY. THE BARREL INSERT AND TAB WERE NOT BROKEN IN THE BODY. CONTACT REPORTED THAT BROKEN FASTENERS WERE REMOVED FROM THE ABDOMINAL CAVITY DURING THE PROCEDURE. IT WAS NOT PREVIOUSLY REPORTED THAT FASTENERS WERE BROKEN IN USE. THE CASE WAS COMPLETED AND THERE WAS NO PATIENT INJURY AS A RESULT OF THE PROBLEM EXPERIENCED.
01-01-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE CUSTOMER STATED THAT AFTER FIRING THE STAPLER, THEY NOTICED THAT ONE SIDE OF THE TISSUE DID NOT GET STAPLED. ONLY ONE SIDE OF THE RELOAD FIRED STAPLES. THE DOCTOR OVER SEWED THE TISSUE. THERE WAS A DELAY OF TWENTY TO THIRTY MINUTES. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
19-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | TA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMI-HEPATECTOMY, AT THE LIGATION AND DIVISION OF THE BRANCH OF THE HEPATIC VEIN, AFTER CLAMPING THE TARGETED VESSEL, THE OPERATOR TRIED TO SQUEEZE THE HANDLE TO FIRE THE STAPLES. HOWEVER, THE DEVICE DID NOT WORK. THE HANDLE DID NOT MOVE AT ALL. THE DEVICE DID NOT FIRE. SO THEY PUSHED THE RETRACTION BLACK BUTTON AND TOOK THE JAW OUT OF THE VESSEL AND ORDERED TO REPLACE THE ABNORMAL STAPLER WITH THE A NEW ONE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, SURGEON WAS NOT ABLE TO FIRE THE DEVICE. THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. NO PATIENT INJURY.
12-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLE DID NOT FIRE EVEN AFTER MULTIPLE RE-TRIES AND RE-ADJUSTMENT. THEY USED ANOTHER RELOAD TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-06-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, WHEN CUTTING THE LUNG PARENCHYMA, THE RELOAD WAS ABLE TO CLIP THE TISSUE, THE GREEN BUTTON WAS ABLE TO BE PRESSED, BUT IT COULD NOT BE FIRED NORMALLY. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED WITH THE SAME BRAND. THERE WAS NO PATIENT INJURY.
16-04-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN USING THE HANDLE TO STAPLE THE PATIENT, THE JAWS COULD NOT BE CLOSED NORMALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
16-04-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN USING THE HANDLE TO STAPLE THE PATIENT, THE JAWS COULD NOT BE CLOSED NORMALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
01-01-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | POWERED PLUS LONG 45 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE REP, DURING A LAPAROSCOPIC/ ROBOTIC CYSTECTOMY, THE STAPLER WAS CLOSED WITH DIFFICULTY ON THE BLADDER PEDICLE. DURING FIRING, THE STAPLER FIRED PART WAY AND THEN STOPPED. THE FA USED THE REVERSE BUTTON TO BRING THE KNIFE BLADE BACK AND OPEN THE STAPLER. UPON REMOVAL, THE STAFF NOTED THAT THE KNIFE BLADE BENT THE ANVIL. THE SURGEON IS AWARE THAT THE TISSUE WAS TOO THICK FOR THE RELOAD HE CHOSE TO USE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT CONSEQUENCES REPORTED.
01-01-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | POWERED PLUS LONG 45 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE REP, DURING A LAPAROSCOPIC/ ROBOTIC CYSTECTOMY, THE STAPLER WAS CLOSED WITH DIFFICULTY ON THE BLADDER PEDICLE. DURING FIRING, THE STAPLER FIRED PART WAY AND THEN STOPPED. THE FA USED THE REVERSE BUTTON TO BRING THE KNIFE BLADE BACK AND OPEN THE STAPLER. UPON REMOVAL, THE STAFF NOTED THAT THE KNIFE BLADE BENT THE ANVIL. THE SURGEON IS AWARE THAT THE TISSUE WAS TOO THICK FOR THE RELOAD HE CHOSE TO USE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT CONSEQUENCES REPORTED.
16-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING ANTRUM DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER COULD NOT BE FIRED AND THERE WAS AN UNLOADING ISSUE. THE SURGEON OPENED A NEW STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING ANTRUM DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER COULD NOT BE FIRED AND THERE WAS AN UNLOADING ISSUE. THE SURGEON OPENED A NEW STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2019 | Malfunction | ethicon endo-surgery | 02-01-2020 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A SLEEVE GASTRECTOMY, AT THE APEX OF THE THIRD TO LAST FIRING THEY NOTICED BLUE PLASTIC FRAGMENTS IN A SMALL PIECE OF TISSUE. THERE WAS NO DISRUPTION TO THE STAPLE LINE. THEY WERE RETRIEVED WITH EASE AND THE CASE WAS COMPLETED WITH THE DEVICE. THERE WERE NO PATIENT CONSEQUENCES. FOUR RELOADS WERE USED IN CASE AND WILL BE RETURNED.
16-12-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
17-06-2019 | Malfunction | covidien lp llc north haven | 02-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DETACHING THE INTESTINAL CANAL ON A LAPAROSCOPIC COLORECTAL SURGERY, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING LUNG TISSUE RESECTION IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THREE HANDLES AND FIVE RELOADS WERE OPENED IN TOTAL, UNFORTUNATELY WITHOUT SUCCESS. A DIFFERENT STAPLER WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
17-06-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DETACHING THE INTESTINAL CANAL ON A LAPAROSCOPIC COLORECTAL SURGERY, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | us surgical puerto rico | 02-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING LUNG TISSUE RESECTION IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THREE HANDLES AND FIVE RELOADS WERE OPENED IN TOTAL, UNFORTUNATELY WITHOUT SUCCESS. A DIFFERENT STAPLER WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | ECHELON FLEX ENDOPATH | Difficult to Open or Close | A051104 | Tissue Damage | Event Description: MD USING AN ARTICULATING ENDOSCOPIC LINEAR CUTTER. PER STAFF REPORT, THE STAPLER WAS CLAMPED, THE STAPLES FIRED, AND THE CUTTER RUN. THEN THE STAPLER WOULD NOT OPEN, CAUSING THE TISSUE TO BE CAUGHT IN THE STAPLER. AFTER MULTIPLE ATTEMPTS TO OPEN THE STAPLER, IT OPENED AND RELEASED THE TISSUE. FOLLOWING THIS, THE MD WAS NOT CONFIDENT THE STAPLES WERE ADEQUATELY PLACED, AND A DECISION WAS MADE TO OVER-SEW THE STAPLE LINE.
10-12-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | ECHELON FLEX ENDOPATH | Mechanical Jam | A0506 | Tissue Damage | Event Description: MD USING AN ARTICULATING ENDOSCOPIC LINEAR CUTTER. PER STAFF REPORT, THE STAPLER WAS CLAMPED, THE STAPLES FIRED, AND THE CUTTER RUN. THEN THE STAPLER WOULD NOT OPEN, CAUSING THE TISSUE TO BE CAUGHT IN THE STAPLER. AFTER MULTIPLE ATTEMPTS TO OPEN THE STAPLER, IT OPENED AND RELEASED THE TISSUE. FOLLOWING THIS, THE MD WAS NOT CONFIDENT THE STAPLES WERE ADEQUATELY PLACED, AND A DECISION WAS MADE TO OVER-SEW THE STAPLE LINE.
29-11-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | ECHELON FLEX | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: STAPLER MISFIRED AND REQUIRED USE OF MANUAL RELEASE BUTTON.
19-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | EEA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SPACE BETWEEN THE CARTRIDGE AND ANVIL CLOSED, THE GREEN BAR VISIBLE IN THE INDICATOR WINDOWS BUT THE DEVICE DID NOT FIRE.
17-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY LAPAROSCOPIC PROCEDURE, THE DEVICE FIRED BUT STOPPED ONE CM FROM DISTAL END. ALSO, THERE WERE UNFORMED STAPLES AT DISTAL END. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY LAPAROSCOPIC PROCEDURE, THE DEVICE FIRED BUT STOPPED ONE CM FROM DISTAL END. ALSO, THERE WERE UNFORMED STAPLES AT DISTAL END. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-11-2019 | Injury | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSE. THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND THE PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
17-11-2019 | Injury | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | Radiation Exposure, Unintended | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSE. THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND THE PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
17-11-2019 | Injury | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSE. THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND THE PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
17-11-2019 | Injury | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSE. THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND THE PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
17-11-2019 | Injury | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Radiation Exposure, Unintended | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSE. THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND THE PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
17-11-2019 | Injury | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSE. THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND THE PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
17-11-2019 | Injury | us surgical puerto rico | 03-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSE. THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND THE PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
17-11-2019 | Injury | us surgical puerto rico | 03-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Radiation Exposure, Unintended | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSE. THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND THE PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
17-11-2019 | Injury | us surgical puerto rico | 03-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSE. THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND THE PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
10-07-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE DEVICE DID NOT FORM PROPERLY WHEN CUTTING THE TISSUE. THE STAPLE DID NOT FORM "B" SHAPE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE DEVICE FIRED, BUT THERE WAS POOR STAPLE FORMATION. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC LOBECTOMY, THE CARTRIDGE WAS LOADED TO THE HANDLE AS USUAL, AND THE OPENING/CLOSING OF JAWS WAS CHECKED. WHEN TRYING TO FIRE, THE DEVICE FAILED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | ILS 29MM, CURVED | Use of Device Problem | A23 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE SURGERY OF LAPAROSCOPIC RADICAL RESECTION OF ASCENDING COLON CANCER, COULD NOT SQUEEZE DOWN THE FIRING TRIGGER EVEN WITH BIG FORCE, USED NTLC75 TO COMPLETE THE SURGERY. AFTER SURGERY, THE SURGEON TRIED 3 TIMES WITH BIG FORCE. THE LAST TIME, COULD FIRED AND WITH AUDIBLE FEEDBACK. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-07-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING WHEN DETACHING THE LUNG TISSUE DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THERE WAS RESISTANCE ON SQUEEZING THE HANDLE AND MADE A BEEP SOUND. HENCE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD AND HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-07-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING WHEN DETACHING THE LUNG TISSUE DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THERE WAS RESISTANCE ON SQUEEZING THE HANDLE AND MADE A BEEP SOUND. HENCE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD AND HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE RELOAD WOULD NOT DEPLOY STAPLES. ANOTHER RELOAD WAS USED. THERE WAS NO DELAY TO THE SURGERY.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE RELOAD WOULD NOT DEPLOY STAPLES. ANOTHER RELOAD WAS USED. THERE WAS NO DELAY TO THE SURGERY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO RECTUM DURING AN OPEN LOWER ANTERIOR RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | PROXIMATE*PPH PROCEDURE SET | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, OPENING AND INSERTING THE STAPLER INTO THE ANAL CANAL WAS POSSIBLE AS USUAL. AFTERWARDS THEY POSITIONED THE ANVIL BEYOND THE PURSE-STRING SUTURE AND CLOSED THE STAPLER. THE RED SAFETY HAS BEEN PULLED BACK TO START THE FIRING PROCESS, BUT THE AUDIBLE CLICK HAS NOT BEEN HEARD. AFTER A ANOTHER TRY THEY OPENED AND REMOVED THE STAPLER COMPLETELY. THEY OBSERVED A FEW STAPLES WHICH DID NOT CLOSE PROPERLY. ADDITIONALLY THEY SAW THAT THE STAPLER DID NOT CUT AT ALL, DUE TO LATER MANIPULATIONS OF THE STAPLER DURING THE REMOVAL THE TISSUE RIPPED. FURTHER RESECTION OF MUCOSA, RE-ADAPTION DONE BY HAND, THEREFORE SUTURING HAD TO BE DONE MORE DISTALLY THAN NORMAL WHICH MEANS MORE POSTOPERATIVE PAIN AND URGE.
25-07-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING WHEN DETACHING THE LUNG TISSUE DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THERE WAS RESISTANCE ON SQUEEZING THE HANDLE AND MADE A BEEP SOUND. HENCE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD AND HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | TA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACIC LOBECTOMY PROCEDURE, AFTER OPENING THE PACKAGE, THE GRIP COULD NOT BE CLOSED AND SUBSEQUENT OPERATION COULD NOT BE CARRIED OUT. AFTER REPLACING ONE, THE SAME SITUATION STILL OCCURRED, AND THE THIRD ONE COULD BE USED NORMALLY. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | TA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACIC LOBECTOMY PROCEDURE, AFTER OPENING THE PACKAGE, THE GRIP COULD NOT BE CLOSED AND SUBSEQUENT OPERATION COULD NOT BE CARRIED OUT. AFTER REPLACING ONE, THE SAME SITUATION STILL OCCURRED, AND THE THIRD ONE COULD BE USED NORMALLY. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, SIX RELOADS DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
21-11-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
21-11-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
13-10-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SEGMENTECTOMY LAPAROSCOPIC PROCEDURE, THE RELOAD COULD NOT BE FIRED. IT WAS ALSO STATED THAT THE RELOAD COULD NOT BE CLOSED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. IT WAS STATED THAT TWO DEVICES HAD THE SAME MALFUNCTION. THERE WAS NO PATIENT INJURY.
13-10-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SEGMENTECTOMY LAPAROSCOPIC PROCEDURE, THE RELOAD COULD NOT BE FIRED. IT WAS ALSO STATED THAT THE RELOAD COULD NOT BE CLOSED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. IT WAS STATED THAT TWO DEVICES HAD THE SAME MALFUNCTION. THERE WAS NO PATIENT INJURY.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | SIGNIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS APPLIED TO LUNG TISSUE (PULMONARY PARENCHYMA) DURING A LAPAROSCOPIC WEDGE RESECTION CONVERTED TO OPEN THORACOTOMY, THE STAPLER WAS ABLE TO BE FIRED BUT THE KNIFE BLADE DID NOT RETRACT AND THE JAWS STAYED CLOSED WITH TISSUE INSIDE. A CRACKING NOISE WAS HEARD DURING THE PROCEDURE. IN ADDITION, THE RELOAD WAS BROKEN AND CURVED ON THE CONNECTION TO THE ADAPTER. OUTSIDE THE PATIENT CAVITY, AN ATTEMPT WAS MADE TO REMOVE THE RELOAD BY FORCE CAUSING IT TO BREAK. SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES AND THE INCISION WAS ALSO EXTENDED MORE THAN 1 INCH. ANOTHER DEVICE WAS USED TO CLOSE THE LUNG TISSUE.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | SIGNIA | Retraction Problem | A0510 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS APPLIED TO LUNG TISSUE (PULMONARY PARENCHYMA) DURING A LAPAROSCOPIC WEDGE RESECTION CONVERTED TO OPEN THORACOTOMY, THE STAPLER WAS ABLE TO BE FIRED BUT THE KNIFE BLADE DID NOT RETRACT AND THE JAWS STAYED CLOSED WITH TISSUE INSIDE. A CRACKING NOISE WAS HEARD DURING THE PROCEDURE. IN ADDITION, THE RELOAD WAS BROKEN AND CURVED ON THE CONNECTION TO THE ADAPTER. OUTSIDE THE PATIENT CAVITY, AN ATTEMPT WAS MADE TO REMOVE THE RELOAD BY FORCE CAUSING IT TO BREAK. SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES AND THE INCISION WAS ALSO EXTENDED MORE THAN 1 INCH. ANOTHER DEVICE WAS USED TO CLOSE THE LUNG TISSUE.
08-11-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY, STAPLE INSTALLATION PROBLEM DURING SURGERY. THE SR75 COULD NOT INSTALL ONTO THE NTLC75 SMOOTHLY. ALSO, IT WAS FOUND THAT THE PLASTIC OF SR75 WAS FALL OFF. FINALLY, THE SURGEON OPENED ANOTHER NEW NTLC75 AND SR75 FOR THE SURGERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
18-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC DISTAL GASTRECTOMY WHEN APPROACHING DUODENUM FOR THE FIRST FIRING AS SOON AS THE TISSUE WAS CLAMPED AND THE GREEN BUTTON WAS PRESSED, THE ADAPTER CAME OFF. A NEW SHELL WAS ATTACHED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Break | A0401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS APPLIED TO LUNG TISSUE (PULMONARY PARENCHYMA) DURING A LAPAROSCOPIC WEDGE RESECTION CONVERTED TO OPEN THORACOTOMY, THE STAPLER WAS ABLE TO BE FIRED BUT THE KNIFE BLADE DID NOT RETRACT AND THE JAWS STAYED CLOSED WITH TISSUE INSIDE. A CRACKING NOISE WAS HEARD DURING THE PROCEDURE. IN ADDITION, THE RELOAD WAS BROKEN AND CURVED ON THE CONNECTION TO THE ADAPTER. OUTSIDE THE PATIENT CAVITY, AN ATTEMPT WAS MADE TO REMOVE THE RELOAD BY FORCE CAUSING IT TO BREAK. SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES AND THE INCISION WAS ALSO EXTENDED MORE THAN 1 INCH. ANOTHER DEVICE WAS USED TO CLOSE THE LUNG TISSUE.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS APPLIED TO LUNG TISSUE (PULMONARY PARENCHYMA) DURING A LAPAROSCOPIC WEDGE RESECTION CONVERTED TO OPEN THORACOTOMY, THE STAPLER WAS ABLE TO BE FIRED BUT THE KNIFE BLADE DID NOT RETRACT AND THE JAWS STAYED CLOSED WITH TISSUE INSIDE. A CRACKING NOISE WAS HEARD DURING THE PROCEDURE. IN ADDITION, THE RELOAD WAS BROKEN AND CURVED ON THE CONNECTION TO THE ADAPTER. OUTSIDE THE PATIENT CAVITY, AN ATTEMPT WAS MADE TO REMOVE THE RELOAD BY FORCE CAUSING IT TO BREAK. SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES AND THE INCISION WAS ALSO EXTENDED MORE THAN 1 INCH. ANOTHER DEVICE WAS USED TO CLOSE THE LUNG TISSUE.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Retraction Problem | A0510 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS APPLIED TO LUNG TISSUE (PULMONARY PARENCHYMA) DURING A LAPAROSCOPIC WEDGE RESECTION CONVERTED TO OPEN THORACOTOMY, THE STAPLER WAS ABLE TO BE FIRED BUT THE KNIFE BLADE DID NOT RETRACT AND THE JAWS STAYED CLOSED WITH TISSUE INSIDE. A CRACKING NOISE WAS HEARD DURING THE PROCEDURE. IN ADDITION, THE RELOAD WAS BROKEN AND CURVED ON THE CONNECTION TO THE ADAPTER. OUTSIDE THE PATIENT CAVITY, AN ATTEMPT WAS MADE TO REMOVE THE RELOAD BY FORCE CAUSING IT TO BREAK. SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES AND THE INCISION WAS ALSO EXTENDED MORE THAN 1 INCH. ANOTHER DEVICE WAS USED TO CLOSE THE LUNG TISSUE.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Noise, Audible | A0508 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS APPLIED TO LUNG TISSUE (PULMONARY PARENCHYMA) DURING A LAPAROSCOPIC WEDGE RESECTION CONVERTED TO OPEN THORACOTOMY, THE STAPLER WAS ABLE TO BE FIRED BUT THE KNIFE BLADE DID NOT RETRACT AND THE JAWS STAYED CLOSED WITH TISSUE INSIDE. A CRACKING NOISE WAS HEARD DURING THE PROCEDURE. IN ADDITION, THE RELOAD WAS BROKEN AND CURVED ON THE CONNECTION TO THE ADAPTER. OUTSIDE THE PATIENT CAVITY, AN ATTEMPT WAS MADE TO REMOVE THE RELOAD BY FORCE CAUSING IT TO BREAK. SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES AND THE INCISION WAS ALSO EXTENDED MORE THAN 1 INCH. ANOTHER DEVICE WAS USED TO CLOSE THE LUNG TISSUE.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, SIX RELOADS DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
18-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY PROCEDURE, THE TISSUE WAS CUT BUT THE CARTRIDGE GOT STUCK AFTER FIRING. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A BARIATRIC LAPAROSCOPIC PROCEDURE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOIDECTOMY, THE PRODUCT WAS NOT USED ON THE PATIENT, AT THE FIRST FIRING, A CARTRIDGE WAS ATTACHED TO THE HANDLE, AND WHEN AN ATTEMPT WAS MADE TO FIRE THE DEVICE, FIRING COULD NOT BE PERFORMED. WHEN THE CARTRIDGE WAS REATTACHED TO ANOTHER NEW HANDLE TO PERFORM A FIRING, THE CARTRIDGE WAS ABLE TO BE USED. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, SIX RELOADS DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, SIX RELOADS DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, SIX RELOADS DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
11-12-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | ILS 33MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, ECS33A STAPLER WAS ATTEMPTED TO BE FIRED TO COMPLETE AN END TO END ANASTOMOSIS. THE STAPLER CUT BUT STAPLES DID NOT DEPLOY. THE SURGEON NOTED THAT HE DID NOT HEAR THE CRUNCHING SOUND. ¿SURGEON IS UNSURE IF IT WAS DEVICE OR USER ERROR¿. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
19-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOLAPAROTOMY, AT THE SECOND FIRING AT THE TIME OF INTERLOBAR RESECTION, AFTER THE SURGEON PRESSED THE GREEN BUTTON AND TRIED TO SQUEEZE THE HANDLE OF THE STAPLER, IT WAS STIFF AND UNABLE TO FIRE, THE HANDLE COULD NOT BE SQUEEZED. RELEASE WAS PERFORMED WITH NO PROBLEM, AND AFTER THAT A NEW HANDLE AND CARTRIDGE WERE USED AND THE OPERATION WAS COMPLETED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOLAPAROTOMY, AT THE SECOND FIRING AT THE TIME OF INTERLOBAR RESECTION, AFTER THE SURGEON PRESSED THE GREEN BUTTON AND TRIED TO SQUEEZE THE HANDLE OF THE STAPLER, IT WAS STIFF AND UNABLE TO FIRE, THE HANDLE COULD NOT BE SQUEEZED. RELEASE WAS PERFORMED WITH NO PROBLEM, AND AFTER THAT A NEW HANDLE AND CARTRIDGE WERE USED AND THE OPERATION WAS COMPLETED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOLAPAROTOMY, AT THE SECOND FIRING AT THE TIME OF INTERLOBAR RESECTION, AFTER THE SURGEON PRESSED THE GREEN BUTTON AND TRIED TO SQUEEZE THE HANDLE OF THE STAPLER, IT WAS STIFF AND UNABLE TO FIRE, THE HANDLE COULD NOT BE SQUEEZED. RELEASE WAS PERFORMED WITH NO PROBLEM, AND AFTER THAT A NEW HANDLE AND CARTRIDGE WERE USED AND THE OPERATION WAS COMPLETED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, SIX RELOADS DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
13-12-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN A RECTOSIGMOIDECTOMY SURGERY, THE DOCTOR USED THE CDH AND EVEN AFTER USING THE CORRECT TECHNIQUE OF USING THE DEVICE AS HE ALWAYS USED THE PRODUCT HE NOTICED THAT SOME STAPLES WERE NOT TIGHTLY CLOSED SO THE DOCTOR SUTURED THE STAPLES TO STRENGTHEN THE ANASTOMOSIS.
20-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | EEA | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING AN ANTERIOR RESECTION WHEN CREATING THE ANASTOMOSES OF THE COLON, AFTER THE STAPLER WAS FIRED, THE SAFETY CATCH WAS RE-ENGAGED AND THE HANDLE WAS TURNED 2 FULL TURNS AND AN AUDIBLE CLICK HEARD, HOWEVER THE ANVIL DIDN¿T TILT. THE DOCTOR REQUESTED THE ASSISTANT KEEP TURNING THE BLACK KNOB HOWEVER THE ANVIL STILL DIDN¿T TILT SO THE DOCTOR HAD TO MANUALLY TILT THE ANVIL SO THE STAPLER COULD BE REMOVED. NO PATIENT HARM.
06-03-2020 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN COLORECTAL /THORACIC SURGERY DURING THE FIRST FIRING, AN UNKNOWN ISSUE OCCURRED WITH THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN A RECTOSIGMOIDECTOMY SURGERY, THE DOCTOR USED THE CDH AND EVEN AFTER USING THE CORRECT TECHNIQUE OF USING THE DEVICE AS HE ALWAYS USED THE PRODUCT HE NOTICED THAT SOME STAPLES WERE NOT TIGHTLY CLOSED SO THE DOCTOR SUTURED THE STAPLES TO STRENGTHEN THE ANASTOMOSIS.
18-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY PROCEDURE, THE TISSUE WAS CUT BUT THE CARTRIDGE GOT STUCK AFTER FIRING. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE HENCE THE STAPLER DID NOT FIRE.
20-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE HENCE THE STAPLER DID NOT FIRE.
19-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE HENCE THE DEVICE DID NOT FIRE.
15-05-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE HAD SOMETHING WRONG, AND THE RELOAD COULD NOT BE OPENED. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | davis & geck caribe | 03-01-2020 | UNKNOWN ENDO STITCH SULU | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, THE DEVICE COULD NOT BE LOADED. WHEN IT WAS TRIED TO BE LOADED FOR THE SECOND TIME, A PIECE OF THE DEVICE BROKE.
20-12-2019 | Malfunction | davis & geck caribe | 03-01-2020 | UNKNOWN ENDO STITCH SULU | Break | A0401 | No Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE COULD NOT BE LOADED. WHEN IT WAS TRIED TO BE LOADED FOR THE SECOND TIME, A PIECE OF THE DEVICE BROKE.
20-12-2019 | Malfunction | davis & geck caribe | 03-01-2020 | UNKNOWN ENDO STITCH SULU | Component or Accessory Incompatibility | A1701 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, THE DEVICE COULD NOT BE LOADED. WHEN IT WAS TRIED TO BE LOADED FOR THE SECOND TIME, A PIECE OF THE DEVICE BROKE.
20-12-2019 | Malfunction | davis & geck caribe | 03-01-2020 | UNKNOWN ENDO STITCH SULU | Component or Accessory Incompatibility | A1701 | No Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE COULD NOT BE LOADED. WHEN IT WAS TRIED TO BE LOADED FOR THE SECOND TIME, A PIECE OF THE DEVICE BROKE.
20-12-2019 | Malfunction | davis & geck caribe | 03-01-2020 | UNKNOWN ENDO STITCH SULU | Human-Device Interface Problem | A22 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, THE DEVICE COULD NOT BE LOADED. WHEN IT WAS TRIED TO BE LOADED FOR THE SECOND TIME, A PIECE OF THE DEVICE BROKE.
20-12-2019 | Malfunction | davis & geck caribe | 03-01-2020 | UNKNOWN ENDO STITCH SULU | Human-Device Interface Problem | A22 | No Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE COULD NOT BE LOADED. WHEN IT WAS TRIED TO BE LOADED FOR THE SECOND TIME, A PIECE OF THE DEVICE BROKE.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE HENCE THE DEVICE DID NOT FIRE.
25-11-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN UNKNOWN PROCEDURE, THE DEVICE COULD NOT BE FIRED AT THE FIRST FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | TA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SMALL BOWEL RESECTION, WHEN APPROXIMATING TISSUE AFTER AN ANASTOMOSIS IN THE BOWEL, THE STAPLES DID NOT DEPLOY. THE SURGEON CUT THE TISSUE THINKING THE STAPLES DEPLOYED. THE SURGEON RE-STAPLED THE ANASTOMOSIS TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | TA | Mechanics Altered | A0507 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SMALL BOWEL RESECTION, WHEN APPROXIMATING TISSUE AFTER AN ANASTOMOSIS IN THE BOWEL, THE STAPLES DID NOT DEPLOY. THE SURGEON CUT THE TISSUE THINKING THE STAPLES DEPLOYED. THE SURGEON RE-STAPLED THE ANASTOMOSIS TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | TA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SMALL BOWEL RESECTION, WHEN APPROXIMATING TISSUE AFTER AN ANASTOMOSIS IN THE BOWEL, THE STAPLES DID NOT DEPLOY. THE SURGEON CUT THE TISSUE THINKING THE STAPLES DEPLOYED. THE SURGEON RE-STAPLED THE ANASTOMOSIS TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-10-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | TA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION WHEN CUTTING THE RECTUM BOTH OF THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT HARM.
18-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | TA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACIC LOBECTOMY PROCEDURE, THE DEVICE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | EEA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE FOUR STAPLERS WERE NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE.
20-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | EEA | Detachment of Device or Device Component | A0501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, A DEVICE COMPONENT DISENGAGED.
19-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY PROCEDURE, WHEN THE DEVICE WAS FIRED, THE STAPLE FAILED TO FORM B SHAPE, THAT RESULT TO INCOMPLETE STAPLE LINE AT THE PROXIMAL PART WHICH CAUSE LEAK. TO RESOLVE THE ISSUE, THE SURGEON HAD TO SUTURE THE TISSUE.
19-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY PROCEDURE, WHEN THE DEVICE WAS FIRED, THE STAPLE FAILED TO FORM B SHAPE, THAT RESULT TO INCOMPLETE STAPLE LINE AT THE PROXIMAL PART WHICH CAUSE LEAK. TO RESOLVE THE ISSUE, THE SURGEON HAD TO SUTURE THE TISSUE.
16-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING JEJUNUM TO JEJUNUM DURING A LAPAROSCOPIC GASTRIC BYPASS, THE STAPLER WAS ABLE TO FIRE BUT SEVERAL STAPLES DID NOT FORM AT THE DISTAL TIP OF THE STAPLE LINE AND APPEARED TO HAVE MISSED THE ANVIL AS THE STAPLES WERE STICKING OUT OF THE TISSUE. THERE WAS A NOTICEABLE TISSUE BRUISING. IT WAS ALSO STATED THAT STAPLE LINE WAS INCOMPLETE DISTALLY AND SURGEON OVERSEWN THE STAPLE LINE TO FIX IT.
16-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING JEJUNUM TO JEJUNUM DURING A LAPAROSCOPIC GASTRIC BYPASS, THE STAPLER WAS ABLE TO FIRE BUT SEVERAL STAPLES DID NOT FORM AT THE DISTAL TIP OF THE STAPLE LINE AND APPEARED TO HAVE MISSED THE ANVIL AS THE STAPLES WERE STICKING OUT OF THE TISSUE. THERE WAS A NOTICEABLE TISSUE BRUISING. IT WAS ALSO STATED THAT STAPLE LINE WAS INCOMPLETE DISTALLY AND SURGEON OVERSEWN THE STAPLE LINE TO FIX IT.
13-09-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | 75MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLER GOT STUCK HALFWAY WHILE FIRING. STAPLE LINE BROKEN. THERE WERE NO PATIENT CONSEQUENCES.
20-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE FIRST FEW PINS POPPED OUT AND THE RELOAD COULD NOT CONTINUE FIRING. IT WAS REPORTED THAT THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL PART. ANOTHER RELOAD WAS THEN USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, USING BOTH RELOADS, THE STAPLE LINE WAS INCOMPLETE DISTALLY. WHEN THE SURGEON PRESSED THE GREEN BUTTON, FEW FIRST PINS POPPED OUT AND THE STAPLE COULD NOT CONTINUE FIRING. REPLACEMENT WITH ANOTHER RELOADS WAS DONE TO COMPLETE THE CASE.
03-01-2020 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE, AND LOCKED ON TISSUE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE, AND LOCKED ON TISSUE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Malfunction | covidien lp llc north haven | 03-01-2020 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE, AND LOCKED ON TISSUE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Malfunction | covidien lp llc north haven | 03-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE, AND LOCKED ON TISSUE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING JEJUNUM TO JEJUNUM DURING A LAPAROSCOPIC GASTRIC BYPASS, THE STAPLER WAS ABLE TO FIRE BUT SEVERAL STAPLES DID NOT FORM AT THE DISTAL TIP OF THE STAPLE LINE AND APPEARED TO HAVE MISSED THE ANVIL AS THE STAPLES WERE STICKING OUT OF THE TISSUE. THERE WAS A NOTICEABLE TISSUE BRUISING. IT WAS ALSO STATED THAT STAPLE LINE WAS INCOMPLETE DISTALLY AND SURGEON OVERSEWN THE STAPLE LINE TO FIX IT.
16-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING JEJUNUM TO JEJUNUM DURING A LAPAROSCOPIC GASTRIC BYPASS, THE STAPLER WAS ABLE TO FIRE BUT SEVERAL STAPLES DID NOT FORM AT THE DISTAL TIP OF THE STAPLE LINE AND APPEARED TO HAVE MISSED THE ANVIL AS THE STAPLES WERE STICKING OUT OF THE TISSUE. THERE WAS A NOTICEABLE TISSUE BRUISING. IT WAS ALSO STATED THAT STAPLE LINE WAS INCOMPLETE DISTALLY AND SURGEON OVERSEWN THE STAPLE LINE TO FIX IT.
21-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT DEVICE MISFIRED AND THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE. THE CASE WAS CONVERTED INTO AN OPEN PROCEDURE AND USED AN OPEN STAPLER TO COMPLETE THE CASE. THE FIRING KNOB OF THE OPEN STAPLER BROKE AND THE STAPLER PARTIALLY FIRED. THE STAPLE LINE WAS RESECTED AND RE-STAPLED. IT WAS ALSO NOTED THAT THERE WAS AN ISSUE WITH THE SEAL OF THE PACKAGING OF THE OPEN STAPLER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
19-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE FOUR STAPLERS WERE NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE.
21-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL END. FIRST FEW PINS POPPED OUT BUT THE REST COULDN'T BE FIRED.
21-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL END. FIRST FEW PINS POPPED OUT BUT THE REST COULDN'T BE FIRED.
21-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | UNKNOWN GIA PRODUCT | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT DEVICE MISFIRED AND THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE. THE CASE WAS CONVERTED INTO AN OPEN PROCEDURE AND USED AN OPEN STAPLER TO COMPLETE THE CASE. THE FIRING KNOB OF THE OPEN STAPLER BROKE AND THE STAPLER PARTIALLY FIRED. THE STAPLE LINE WAS RESECTED AND RE-STAPLED. IT WAS ALSO NOTED THAT THERE WAS AN ISSUE WITH THE SEAL OF THE PACKAGING OF THE OPEN STAPLER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
21-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | RELIAMAX | Unsealed Device Packaging | A020503 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT DEVICE MISFIRED AND THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE. THE CASE WAS CONVERTED INTO AN OPEN PROCEDURE AND USED AN OPEN STAPLER TO COMPLETE THE CASE. THE FIRING KNOB OF THE OPEN STAPLER BROKE AND THE STAPLER PARTIALLY FIRED. THE STAPLE LINE WAS RESECTED AND RE-STAPLED. IT WAS ALSO NOTED THAT THERE WAS AN ISSUE WITH THE SEAL OF THE PACKAGING OF THE OPEN STAPLER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
21-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | RELIAMAX | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT DEVICE MISFIRED AND THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE. THE CASE WAS CONVERTED INTO AN OPEN PROCEDURE AND USED AN OPEN STAPLER TO COMPLETE THE CASE. THE FIRING KNOB OF THE OPEN STAPLER BROKE AND THE STAPLER PARTIALLY FIRED. THE STAPLE LINE WAS RESECTED AND RE-STAPLED. IT WAS ALSO NOTED THAT THERE WAS AN ISSUE WITH THE SEAL OF THE PACKAGING OF THE OPEN STAPLER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
21-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | RELIAMAX | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT DEVICE MISFIRED AND THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE. THE CASE WAS CONVERTED INTO AN OPEN PROCEDURE AND USED AN OPEN STAPLER TO COMPLETE THE CASE. THE FIRING KNOB OF THE OPEN STAPLER BROKE AND THE STAPLER PARTIALLY FIRED. THE STAPLE LINE WAS RESECTED AND RE-STAPLED. IT WAS ALSO NOTED THAT THERE WAS AN ISSUE WITH THE SEAL OF THE PACKAGING OF THE OPEN STAPLER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
21-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT DEVICE MISFIRED AND THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE. THE CASE WAS CONVERTED INTO AN OPEN PROCEDURE AND USED AN OPEN STAPLER TO COMPLETE THE CASE. THE FIRING KNOB OF THE OPEN STAPLER BROKE AND THE STAPLER PARTIALLY FIRED. THE STAPLE LINE WAS RESECTED AND RE-STAPLED. IT WAS ALSO NOTED THAT THERE WAS AN ISSUE WITH THE SEAL OF THE PACKAGING OF THE OPEN STAPLER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
21-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, USING BOTH RELOADS, THE STAPLE LINE WAS INCOMPLETE DISTALLY. WHEN THE SURGEON PRESSED THE GREEN BUTTON, FEW FIRST PINS POPPED OUT AND THE STAPLE COULD NOT CONTINUE FIRING. REPLACEMENT WITH ANOTHER RELOADS WAS DONE TO COMPLETE THE CASE.
20-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AT THE DISCONNECTED BOWEL DURING AN OPEN COLECTOMY, THE STAPLES HAD POOR FORMATION, AND NO B-SHAPE WAS FORMED. SUTURE WAS USED FOR STITCHING IN ORDER TO COMPLETE THE CASE.
21-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, THERE WAS POOR STAPLE FORMATION AND THE STAPLER COULD NOT FINISH STAPLE LINE. PART OF STAPLER WAS STUCK IN RELOAD. THE STAPLER FIRED BUT STOPPED AND STAPLE LINE WAS INCOMPLETE CENTRALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, THERE WAS POOR STAPLE FORMATION AND THE STAPLER COULD NOT FINISH STAPLE LINE. PART OF STAPLER WAS STUCK IN RELOAD. THE STAPLER FIRED BUT STOPPED AND STAPLE LINE WAS INCOMPLETE CENTRALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY PROCEDURE, THE TISSUE WAS CUT BUT THE CARTRIDGE GOT STUCK AFTER FIRING. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY PROCEDURE, THE TISSUE WAS CUT BUT THE CARTRIDGE GOT STUCK AFTER FIRING. THERE WAS NO PATIENT INJURY.
21-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL END. FIRST FEW PINS POPPED OUT BUT THE REST COULDN'T BE FIRED.
21-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL END. FIRST FEW PINS POPPED OUT BUT THE REST COULDN'T BE FIRED.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE. NO PATIENT HARM.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE. NO PATIENT HARM.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE FOR ALL RELOADS.
17-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY PROCEDURE, THE TISSUE WAS CUT BUT THE CARTRIDGE GOT STUCK AFTER FIRING. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY PROCEDURE, THE TISSUE WAS CUT BUT THE CARTRIDGE GOT STUCK AFTER FIRING. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE FOR ALL RELOADS.
17-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, THE SURGEON TRIED FIRING INTO THE PULMONARY VESSEL HOWEVER, AFTER SQUEEZING THE HANDLE ONCE AND PARTIALLY FIRING, THE HANDLE JAMMED AND THE BLADE WILL NOT MOVE ANY FURTHER. THE SURGEON APPLIED TWO CLAMPS ON EITHER SIDE BEFORE REMOVING THE STAPLER. ONCE REMOVED THE SURGEON SAW THAT THERE WAS BRUISING ON THE VESSEL. A NEW RELOAD WAS USED. IT WAS ALSO REPORTED THAT THERE WAS POOR STAPLE FORMATION. THE SURGEON WAS ABLE TO USED THE HANDLE TO FIRE MULTIPLE TIMES AFTERWARDS. THERE WAS NO PATIENT HARM.
17-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, THE SURGEON TRIED FIRING INTO THE PULMONARY VESSEL HOWEVER, AFTER SQUEEZING THE HANDLE ONCE AND PARTIALLY FIRING, THE HANDLE JAMMED AND THE BLADE WILL NOT MOVE ANY FURTHER. THE SURGEON APPLIED TWO CLAMPS ON EITHER SIDE BEFORE REMOVING THE STAPLER. ONCE REMOVED THE SURGEON SAW THAT THERE WAS BRUISING ON THE VESSEL. A NEW RELOAD WAS USED. IT WAS ALSO REPORTED THAT THERE WAS POOR STAPLE FORMATION. THE SURGEON WAS ABLE TO USED THE HANDLE TO FIRE MULTIPLE TIMES AFTERWARDS. THERE WAS NO PATIENT HARM.
22-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE UPON STAPLING FOURTH FITTING, THE STAPLE LOCKED ON THE TISSUE AND WAS NOT ABLE TO OPEN. DOCTOR USED SUTURE AND ENDOCLINCH TO REWARD THE KNIFE MANUALLY. REPLACEMENT OF THE HANDLE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, ON THE GREATER CURVATURE, THE STAPLER WAS ABLE TO FIRE, THE STAPLE LINE WAS INCOMPLETE PROXIMALLY ON TWO RELOADS. THEY SUTURED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE. THEY REPLACED THE MISFIRED ONE WITH ANOTHER RELOAD TO COMPLETE THE CASE.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE. NO PATIENT HARM.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE FOR ALL RELOADS.
16-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE RELOAD FIRED BUT STOPPED AND WAS UNABLE TO CONTINUE TO FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-05-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE HAD SOMETHING WRONG, AND THE RELOAD COULD NOT BE OPENED. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-06-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN ESOPHAGEAL CANCER OPERATION, UPON USING THE DEVICE TO CLOSE VESSELS, THE TISSUE COULD NOT BE CLIPPED AND IT COULD NOT BE FIRED NORMALLY. THE SURGEON REPLACED WITH THE SAME TYPE OF RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE FOR ALL RELOADS.
21-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, ON THE SECOND PART OF THE STOMACH, STAPLE LINE WAS INCOMPLETE IN PROXIMAL. SUTURING WAS DONE TO FIX THE STAPLE LINE. SURGEON REPLACED THE MISFIRED ONE WITH ANOTHER RELOAD.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE FOR ALL RELOADS.
22-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IN THE STOMACH, THE KNIFE DID NOT CUT TILL THE END AND THE REINFORCED MATERIAL AND STITCH WERE STUCK TO THE END OF THE RELOAD. THE SURGEONS HAVE TO CAREFULLY REMOVE THE TISSUE BY GRASPER AND CUT THE STUCK STITCH BY A LAPAROSCOPIC SCISSOR.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE FOR ALL RELOADS.
22-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | UNKNOWN STAPLING DEVICE | Failure to Cut | A050702 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IN THE STOMACH, THE KNIFE DID NOT CUT TILL THE END AND THE REINFORCED MATERIAL AND STITCH WERE STUCK TO THE END OF THE RELOAD. THE SURGEONS HAVE TO CAREFULLY REMOVE THE TISSUE BY GRASPER AND CUT THE STUCK STITCH BY A LAPAROSCOPIC SCISSOR.
31-10-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | GST GREEN RELOAD, 60MM, 6 ROW | Device Fell | A051203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, RELOAD WAS OPENED, NURSE WANTED TO INSERT THE RELOAD TO THE POWERED ECHELON GUN, WHEN A RED PLASTIC PIECE FELL OF THE RELOAD. AND LATER THE RELOAD COULDN¿T BE PLACED INSIDE THE POWERED GUN. THERE WERE NO PATIENT CONSEQUENCES.
20-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE FOR ALL RELOADS.
22-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE MISFIRED. TO RESOLVE THE ISSUE, THE SURGEON USED ANOTHER DEVICE RELOAD. THERE WAS NO PATIENT INJURY.
22-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE MISFIRED. TO RESOLVE THE ISSUE, THE SURGEON USED ANOTHER DEVICE RELOAD. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, DURING THE SIXTH FIRING, THE DEVICE APPROACHED TO THE TISSUE AND THE CARTRIDGE WAS CLOSED. AFTER THAT, THE DOCTOR ATTEMPTED TO PRESS THE GREEN BUTTON, BUT IT DID NOT FLASH. THE OPERATION WAS REPEATED SEVERAL TIMES, BUT THE GREEN BUTTON DID NOT TURN ON AFTER ALL, SO THE CARTRIDGE WAS RE-ATTACHED. AFTER ATTACHING THE SAME CARTRIDGE, THE GREEN BUTTON TURNED ON WITHOUT ANY PROBLEMS AND THE FIRING WAS ABLE TO BE PERFORMED, BUT THE FIRING SUDDENLY STOPPED WHEN ABOUT HALF FIRING WAS COMPLETED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Insufficient Information | E2401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, DURING THE SIXTH FIRING, THE DEVICE APPROACHED TO THE TISSUE AND THE CARTRIDGE WAS CLOSED. AFTER THAT, THE DOCTOR ATTEMPTED TO PRESS THE GREEN BUTTON, BUT IT DID NOT FLASH. THE OPERATION WAS REPEATED SEVERAL TIMES, BUT THE GREEN BUTTON DID NOT TURN ON AFTER ALL, SO THE CARTRIDGE WAS RE-ATTACHED. AFTER ATTACHING THE SAME CARTRIDGE, THE GREEN BUTTON TURNED ON WITHOUT ANY PROBLEMS AND THE FIRING WAS ABLE TO BE PERFORMED, BUT THE FIRING SUDDENLY STOPPED WHEN ABOUT HALF FIRING WAS COMPLETED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, DURING THE SIXTH FIRING, THE DEVICE APPROACHED TO THE TISSUE AND THE CARTRIDGE WAS CLOSED. AFTER THAT, THE DOCTOR ATTEMPTED TO PRESS THE GREEN BUTTON, BUT IT DID NOT FLASH. THE OPERATION WAS REPEATED SEVERAL TIMES, BUT THE GREEN BUTTON DID NOT TURN ON AFTER ALL, SO THE CARTRIDGE WAS RE-ATTACHED. AFTER ATTACHING THE SAME CARTRIDGE, THE GREEN BUTTON TURNED ON WITHOUT ANY PROBLEMS AND THE FIRING WAS ABLE TO BE PERFORMED, BUT THE FIRING SUDDENLY STOPPED WHEN ABOUT HALF FIRING WAS COMPLETED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
22-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE ON GREATER CURVATURE, THE DEVICE FIRED HOWEVER THERE WAS AN INCOMPLETE STAPLE LINE ON THE PROXIMAL. SUTURING AND REPLACING OF THE MISFIRED DEVICE WERE DONE TO COMPLETE THE PROCEDURE. AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
23-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LATERAL HEPATIC LOBECTOMY PROCEDURE, WHILE USING THE DEVICE ON THE LIVER PARENCHYMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THREE RELOADS DID NOT FIRE.
17-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING THE STOMACH DURING A LAPAROSCOPIC GASTRIC SLEEVE, THE RELOAD WOULD NOT FIRE. THE DEVICE WAS TESTED AND FIRED ONE STAPLE LOAD WITH THE ADAPTER, THE LOAD FIRED COMPLETELY BUT SLOWER THAN THE NORMAL. ANOTHER ATTEMPT WAS MADE TO FIRE ANOTHER LOAD AND IT WOULD NOT FIRE. IN ADDITION, THE DEVICE ACTED AS IF IT WOULD FIRE BUT SLOWER THAN NORMAL. SAME THING HAPPENED WHEN THE SURGEON FIRED THE DEVICE DURING PROCEDURE EXCEPT ABOUT 10MM OF STAPLES FIRED BEFORE THE BLADE STOPPED. IT WAS ALSO STATED THAT THERE WAS STAPLE LINE INCOMPLETE PROXIMALLY. THE LINE WAS RE-INITIATED WITH A NEW STAPLER, ADAPTER AND RELOAD. THERE WAS NO PATIENT INJURY.
30-10-2019 | Malfunction | ethicon endo-surgery | 03-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AFTER A FIRST USE OF THE STAPLER WITHOUT ANY PROBLEM, IT WAS UNABLE TO REINSERT A NEW RELOAD, THE INNER FINS DID NOT RETRACT. THE INSTALLATION OF THE CARTRIDGE RELOAD INTO THE JAW OF THE DEVICE WAS IMPOSSIBLE. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES.
23-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LATERAL HEPATIC LOBECTOMY PROCEDURE, WHILE USING THE DEVICE ON THE LIVER PARENCHYMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THREE RELOADS DID NOT FIRE.
18-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, DURING THE SIXTH FIRING, THE DEVICE APPROACHED TO THE TISSUE AND THE CARTRIDGE WAS CLOSED. AFTER THAT, THE DOCTOR ATTEMPTED TO PRESS THE GREEN BUTTON, BUT IT DID NOT FLASH. THE OPERATION WAS REPEATED SEVERAL TIMES, BUT THE GREEN BUTTON DID NOT TURN ON AFTER ALL, SO THE CARTRIDGE WAS RE-ATTACHED. AFTER ATTACHING THE SAME CARTRIDGE, THE GREEN BUTTON TURNED ON WITHOUT ANY PROBLEMS AND THE FIRING WAS ABLE TO BE PERFORMED, BUT THE FIRING SUDDENLY STOPPED WHEN ABOUT HALF FIRING WAS COMPLETED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THIRD FIRING ON A SUB TOTAL GASTRECTOMY, THERE WERE TWO DEVICES THAT COULD NOT BE FIRED WITH THE S AME HANDLE. THERE WAS NO ERROR INDICATION ON THE POWERED HANDLE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Entrapment of Device | A150208 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
19-12-2019 | Injury | celestica de monterrey s.a de c.v | 03-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER FIRED AND THE TISSUE WAS CUT, BUT THE STAPLES DID NOT FORM CORRECTLY. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED USING A HAND INSTRUMENT AND BY PULLING THE STAPLER BACK. THE STAPLE PINS FELL INTO THE CAVITY OF THE PATIENT. THE SURGICAL TIME WAS EXTENDED BY ALMOST 1 HOUR TO CONTROL THE BLEEDING AFTER WHICH THE SURGERY WAS CONVERTED TO BYPASS. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THE DEVICE ISSUE.
23-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED LOBECTOMY PROCEDURE, THE DEVICE WAS ABLE TO BE FIRED NINE (9) TIMES WITHOUT ISSUE, HOWEVER, ON THE 10TH FIRING THE RELOAD CANNOT BE FIRED COMPLETELY BECAUSE THE DEVICE JAMMED AND THE KNIFE STOPPED MOVING AND COULD NOT CONTINUE. THE SURGEON USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-02-2020 | Malfunction | davol inc sub c.r bard | 03-01-2020 | SORBAFIX ABSORBABLE FIXATION SYSTEM | Incorrect, Inadequate or Imprecise Result or Readings | A0908 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SORBAFIX FIXATION DEVICE HAD TO BE CHANGED DURING A LAPAROSCOPIC SURGERY BECAUSE THERE WAS DIFFICULTY IN CLAMPING/MAKING PRESSURE. A SECURESTRAP WAS REQUESTED AND THE SURGERY WAS UNEVENTFULLY COMPLETED. THERE WAS NO INJURY OR IMPACT TO THE PATIENT REPORTED. BASED ON THE INFORMATION PROVIDED, THE SORBAFIX DID NOT ADEQUATELY SEAT. ADDENDUM: THE SURGEON WAS IMPLANTING FASTENERS IN OR NEAR THE COOPER'S LIGAMENT BEFORE THE PROBLEM OCCURRED. ALTHOUGH THE EXACT TIMING OF THE PROBLEM COULD NOT BE RECALLED, IT WAS BELIEVE TO HAVE OCCURRED AFTER THE THIRD FASTENER. RIGHT AFTER THE LAST APPLICATION, BEFORE THE STAPLER LOCKED UP AND WAS UNABLE TO BE USED, THE FASTENERS DID MANAGE TO OBTAIN PROPER FASTENING. THE CASE WAS PROLONGED FOR APPROXIMATELY 10 MINUTES. IT WAS UNKNOWN IF THE TEMPERATURE INDICATOR HAD BEEN ACTIVATED.
14-02-2020 | Malfunction | davol inc sub c.r bard | 03-01-2020 | SORBAFIX ABSORBABLE FIXATION SYSTEM | Inaccurate Synchronization | A071202 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SORBAFIX FIXATION DEVICE HAD TO BE CHANGED DURING A LAPAROSCOPIC SURGERY BECAUSE THERE WAS DIFFICULTY IN CLAMPING/MAKING PRESSURE. A SECURESTRAP WAS REQUESTED AND THE SURGERY WAS UNEVENTFULLY COMPLETED. THERE WAS NO INJURY OR IMPACT TO THE PATIENT REPORTED. BASED ON THE INFORMATION PROVIDED, THE SORBAFIX DID NOT ADEQUATELY SEAT. ADDENDUM: THE SURGEON WAS IMPLANTING FASTENERS IN OR NEAR THE COOPER'S LIGAMENT BEFORE THE PROBLEM OCCURRED. ALTHOUGH THE EXACT TIMING OF THE PROBLEM COULD NOT BE RECALLED, IT WAS BELIEVE TO HAVE OCCURRED AFTER THE THIRD FASTENER. RIGHT AFTER THE LAST APPLICATION, BEFORE THE STAPLER LOCKED UP AND WAS UNABLE TO BE USED, THE FASTENERS DID MANAGE TO OBTAIN PROPER FASTENING. THE CASE WAS PROLONGED FOR APPROXIMATELY 10 MINUTES. IT WAS UNKNOWN IF THE TEMPERATURE INDICATOR HAD BEEN ACTIVATED.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHEN THE DUODENUM WAS RESECTED IN A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, THE STAPLES DID NOT FORM WELL WHICH RESULTED IN STAPLE BURST AND BLEEDING. SOME OF THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. THERE WAS ALSO AN INCOMPLETE STAPLE LINE WHICH WAS FIXED USING SUTURES. A NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THE DEVICE PROBLEM.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DUODENUM WAS RESECTED IN A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, THE STAPLES DID NOT FORM WELL WHICH RESULTED IN STAPLE BURST AND BLEEDING. SOME OF THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. THERE WAS ALSO AN INCOMPLETE STAPLE LINE WHICH WAS FIXED USING SUTURES. A NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THE DEVICE PROBLEM.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHEN THE DUODENUM WAS RESECTED IN A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, THE STAPLES DID NOT FORM WELL WHICH RESULTED IN STAPLE BURST AND BLEEDING. SOME OF THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. THERE WAS ALSO AN INCOMPLETE STAPLE LINE WHICH WAS FIXED USING SUTURES. A NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THE DEVICE PROBLEM.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DUODENUM WAS RESECTED IN A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, THE STAPLES DID NOT FORM WELL WHICH RESULTED IN STAPLE BURST AND BLEEDING. SOME OF THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. THERE WAS ALSO AN INCOMPLETE STAPLE LINE WHICH WAS FIXED USING SUTURES. A NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THE DEVICE PROBLEM.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHEN THE DUODENUM WAS RESECTED IN A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, THE STAPLES DID NOT FORM WELL WHICH RESULTED IN STAPLE BURST AND BLEEDING. SOME OF THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. THERE WAS ALSO AN INCOMPLETE STAPLE LINE WHICH WAS FIXED USING SUTURES. A NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THE DEVICE PROBLEM.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN THE DUODENUM WAS RESECTED IN A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, THE STAPLES DID NOT FORM WELL WHICH RESULTED IN STAPLE BURST AND BLEEDING. SOME OF THE STAPLES FELL INTO THE CAVITY OF THE PATIENT. THERE WAS ALSO AN INCOMPLETE STAPLE LINE WHICH WAS FIXED USING SUTURES. A NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THE DEVICE PROBLEM.
23-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED LOBECTOMY PROCEDURE, THE DEVICE WAS ABLE TO BE FIRED NINE (9) TIMES WITHOUT ISSUE, HOWEVER, ON THE 10TH FIRING THE RELOAD CANNOT BE FIRED COMPLETELY BECAUSE THE DEVICE JAMMED AND THE KNIFE STOPPED MOVING AND COULD NOT CONTINUE. THE SURGEON USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND POOR STAPLE FORMATION. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE.
23-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THIRD FIRING ON A SUB TOTAL GASTRECTOMY, THERE WERE TWO DEVICES THAT COULD NOT BE FIRED WITH THE S AME HANDLE. THERE WAS NO ERROR INDICATION ON THE POWERED HANDLE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-07-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, WHEN DISCONNECTING THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. ANOTHER RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. NO PATIENT HARM.
23-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENECTOMY PROCEDURE, WHILE APPLYING THE DEVICE TO THE DETACHED TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Overheating of Device | A1005 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SUB TOTAL GASTRECTOMY PROCEDURE, AFTER THE SECOND FIRING THE DEVICE HANDLE DISPLAY DISAPPEARED AND THE DEVICE BECAME INOPERATIVE. TO RESOLVE THE ISSUE, THE SURGEON USED MANUAL STAPLER. IT WAS NOTED THAT THE DEVICE HEATED UP TO 50-60 DEGREE AFTER THE DISPLAY DISAPPEARED. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Complete Loss of Power | A070801 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SUB TOTAL GASTRECTOMY PROCEDURE, AFTER THE SECOND FIRING THE DEVICE HANDLE DISPLAY DISAPPEARED AND THE DEVICE BECAME INOPERATIVE. TO RESOLVE THE ISSUE, THE SURGEON USED MANUAL STAPLER. IT WAS NOTED THAT THE DEVICE HEATED UP TO 50-60 DEGREE AFTER THE DISPLAY DISAPPEARED. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENECTOMY PROCEDURE, WHILE APPLYING THE DEVICE TO THE DETACHED TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMI-HEPATECTOMY, WHEN THE SURGEON TRIED TO TRANSECT THE LIVER PARENCHYMA WITH THE FIRST RELOAD, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. SAME ISSUE OCCURRED ON ANOTHER TWO RELOADS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMI-HEPATECTOMY, WHEN THE SURGEON TRIED TO TRANSECT THE LIVER PARENCHYMA WITH THE FIRST RELOAD, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. SAME ISSUE OCCURRED ON ANOTHER TWO RELOADS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, WHILE CUTTING THE RECTUM SURGEON HEARD A BIG SOUND AND HANDLE BROKE. THE SURGEON ALSO NOTICED THAT THE RELOAD DIDN'T FORM A PERFECT B SHAPE. SURGEON REPLACED THE HANDLE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Device Displays Incorrect Message | A090201 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, AT THE SECOND FIRING, THE GREEN BUTTON FLASHED, THEN THE SURGEON PUSHED THE CENTER CONTROL BUTTON DOWN, BUT THE DEVICE STOPPED REACTING IMMEDIATELY. ON THE PREVIOUS SETTING, THE COMMUNICATION ERROR OF THE ADAPTER OCCURRED. PROBABLY ,THE REMOVAL OF THE CARTRIDGE HAD BEEN PERFORMED IMPROPERLY AND THERE WAS A POSSIBILITY IT HAD CAUSED LOADING FAILURE AND PRE-FIRING WITH THE REPORTED CARTRIDGE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Component or Accessory Incompatibility | A1701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, AT THE SECOND FIRING, THE GREEN BUTTON FLASHED, THEN THE SURGEON PUSHED THE CENTER CONTROL BUTTON DOWN, BUT THE DEVICE STOPPED REACTING IMMEDIATELY. ON THE PREVIOUS SETTING, THE COMMUNICATION ERROR OF THE ADAPTER OCCURRED. PROBABLY ,THE REMOVAL OF THE CARTRIDGE HAD BEEN PERFORMED IMPROPERLY AND THERE WAS A POSSIBILITY IT HAD CAUSED LOADING FAILURE AND PRE-FIRING WITH THE REPORTED CARTRIDGE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, AT THE SECOND FIRING, THE GREEN BUTTON FLASHED, THEN THE SURGEON PUSHED THE CENTER CONTROL BUTTON DOWN, BUT THE DEVICE STOPPED REACTING IMMEDIATELY. ON THE PREVIOUS SETTING, THE COMMUNICATION ERROR OF THE ADAPTER OCCURRED. PROBABLY ,THE REMOVAL OF THE CARTRIDGE HAD BEEN PERFORMED IMPROPERLY AND THERE WAS A POSSIBILITY IT HAD CAUSED LOADING FAILURE AND PRE-FIRING WITH THE REPORTED CARTRIDGE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Complete Loss of Power | A070801 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, AT THE SECOND FIRING, THE GREEN BUTTON FLASHED, THEN THE SURGEON PUSHED THE CENTER CONTROL BUTTON DOWN, BUT THE DEVICE STOPPED REACTING IMMEDIATELY. ON THE PREVIOUS SETTING, THE COMMUNICATION ERROR OF THE ADAPTER OCCURRED. PROBABLY ,THE REMOVAL OF THE CARTRIDGE HAD BEEN PERFORMED IMPROPERLY AND THERE WAS A POSSIBILITY IT HAD CAUSED LOADING FAILURE AND PRE-FIRING WITH THE REPORTED CARTRIDGE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
25-06-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZED THE HANDLE. HENCE, THE STAPLER DID NOT FIRE DURING DETACHING THE LUNG TISSUE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | SIGNIA | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, DURING SET UP, THE ADAPTER WAS REPORTED TO BE BROKEN. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT HEMICOLECTOMY, DURING EXTRACORPOREAL TRANSECTION OF THE BOWEL, THE SURGEON STATED THAT THERE WAS NO PROBLEM WITH THE HANDLE BUT THE SECOND RELOAD USED COULD NOT BE FIRED. SURGEON REPLACED THE RELOAD WITH A NEW ONE. NO PATIENT INJURY.
11-12-2019 | Injury | us surgical puerto rico | 03-01-2020 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON PROCEDURE, WHEN RESTORING CONTINUITY ON A COLON, THE DEVICE WAS UNABLE TO SNAP THE ANVIL ONTO THE CLAMP. THE FINS AT THE ANVIL HAVE MOVED AWAY READY TO BREAK. SURGEON NEEDED TO PERFORM SURGICAL INTERVENTION. SURGEON NEED TO CONVERT THE PROCEDURE INTO AN OPEN PROCEDURE. PATIENT HAD EXTENDED HOSPITALIZATION. EXTENDED INCISION OF MORE THAN 1 INCH WAS ALSO NOTED.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, WHILE THE RELOAD WAS APPLIED ONTO THE DISTAL END OF THE RECTUM, THE DEVICE WAS NOT ABLE TO BE FIRED. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. SURGEON ALSO EXPERIENCED DIFFICULTY IN UNLOADING THE DEVICE. OPERATOR APPLIED EXCESSIVE PRESSURE TO UNLOAD THE RELOAD. SURGEON USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE DUE TO DEVICE PROBLEM.
18-12-2019 | Injury | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, WHILE THE RELOAD WAS APPLIED ONTO THE DISTAL END OF THE RECTUM, THE DEVICE WAS NOT ABLE TO BE FIRED. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. SURGEON ALSO EXPERIENCED DIFFICULTY IN UNLOADING THE DEVICE. OPERATOR APPLIED EXCESSIVE PRESSURE TO UNLOAD THE RELOAD. SURGEON USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE DUE TO DEVICE PROBLEM.
25-07-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LUNG CANCER SURGERY, WHEN DISCONNECTING THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. ANOTHER RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT HARM.
17-12-2019 | Malfunction | covidien lp llc north haven | 03-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT GOT STUCK AFTER FIRING. THE DEVICE WAS REMOVED WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
17-12-2019 | Malfunction | us surgical puerto rico | 03-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT GOT STUCK AFTER FIRING. THE DEVICE WAS REMOVED WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
09-12-2019 | Injury | intuitive surgical | 06-01-2020 | ENDOWRIST | Misfire | A050502 | Injury | Event Description: IT WAS INITIALLY REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, WHEN THE STAPLER 30 INSTRUMENT WAS FIRED WITH A GREEN STAPLER 30 RELOAD, THERE WAS A BLADE EXPOSED MESSAGE. THEN, A BACK-UP STAPLER INSTRUMENT AND RELOAD WERE TRIED; HOWEVER, THE SAME BLADE EXPOSED ERROR MESSAGE WAS GENERATED. THEN THE SURGEON "OPTED" TO PLACE SUTURES IN ORDER TO "SEAL" THE AREA. INTUITIVE SURGICAL INC. (ISI) MADE MULTIPLE FOLLOW-UP ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION. HOWEVER, NO FURTHER DETAILS HAVE BEEN RECEIVED AS OF THE DATE OF THIS REPORT.
18-10-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE STOMACH DURING A ROBOTIC LAPAROSCOPIC GASTROPLASTY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS A STAPLE FORMATION ISSUE. IT WAS ALSO STATED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE DOCTOR NEEDED TO FORCE THE DEVICE TO LOOSEN THE TISSUE. IT WAS CONFIRMED THAT TWO RELOADS HAD THE SAME MALFUNCTION. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSONIC GASTRECTOMY PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ABDOMINAL WALL HERNIA PROCEDURE, THE PRODUCT WAS ABLE TO FIRE BUT THE PENETRATION WAS BAD. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, WHILE USING THE DEVICE ON THE STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. A NEW CUTTING CLOSURE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHEN STAPLING, THE FIRST HANDLE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE SECOND HANDLES HAD AN UNKNOWN ISSUE. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHEN STAPLING, THE FIRST HANDLE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE SECOND HANDLES HAD AN UNKNOWN ISSUE. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING THE FISSURES IN A LAPAROSCOPIC PULMONARY LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GRE EN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. THE DEVICE REPLACED WITH A NEW CUTTING CLOSURE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORT, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING THE FISSURES IN A LAPAROSCOPIC PULMONARY LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. THE DEVICE REPLACED WITH A NEW CUTTING CLOSURE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORT, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, WHILE USING THE DEVICE ON THE STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. A NEW CUTTING CLOSURE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY LAPAROSCOPIC PROCEDURE, THE STAPLES DISENGAGED INTO THE CAVITY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING FIRST FIRE AT THE ANTRUM OF THE STOMACH ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS ABLE TO FIRE BUT THE RELOAD LOCKED UP AND DIDN'T CLOSE WELL. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | EEA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLORECTAL ANASTOMOSIS ON A LAPAROSCOPIC PROCTECTOMY, THE DEVICE DID NOT CUT AND STAPLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LEFT HEMIHEPATECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HANDLE COULD NOT BE SQUEEZED. THE DEVICE WAS UNABLE TO FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | EEA | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE ESOPHAGUS AND STOMACH WERE ANATOMIZED DURING AN OPEN ESOPHAGOGASTROSTOMY, THE ANVIL DID NOT TILT AFTER FIRING, AND ONLY PART OF THE STAPLE WAS OUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | EEA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE ESOPHAGUS AND STOMACH WERE ANATOMIZED DURING AN OPEN ESOPHAGOGASTROSTOMY, THE ANVIL DID NOT TILT AFTER FIRING, AND ONLY PART OF THE STAPLE WAS OUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY WHILE WORKING ON THE FUNDUS, THE REPORTED DEVICES MISFIRED. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE DEVICE PROBLEM. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
21-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY WHILE WORKING ON THE FUNDUS, THE REPORTED DEVICES MISFIRED. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE DEVICE PROBLEM. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
25-10-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A DISTAL GASTRECTOMY, WHEN RESECTING THE JEJUNUM, THE HANDLE WAS SQUEEZED AND THE FIRING WAS PERFORMED, HOWEVER BLACK RETURN KNOB ONLY ADVANCED 30MM AND STOPPED. THE HANDLE WAS TOO STIFF TO BE SQUEEZED FURTHER. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ESOPHAGEAL TRANSECTION, THE RELOAD WAS PLACED AT THE TIME OF FIRING BUT IT DID NOT PERFORM FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
21-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY WHILE WORKING ON THE FUNDUS, THE REPORTED DEVICES MISFIRED. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE DEVICE PROBLEM. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
21-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY WHILE WORKING ON THE FUNDUS, THE REPORTED DEVICES MISFIRED. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE DEVICE PROBLEM. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
27-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY, A DEVICE COMPONENT DISENGAGED. THE DEVICE WAS CHANGED BY A NEW ONE TO COMPLETE THE CASE. THE EVENT HAPPENED DURING A PROCEDURE BUT THE DEVICE WAS NOT USED IN PATIENT.
26-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE LUNG LOBE IN A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE LUNG LOBE IN A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-08-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG CANCER SURGERY, THE SURGEON COULD NOT PRESSED THE GREEN BUTTON AND COULD NOT SQUEEZED THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ESOPHAGEAL TRANSECTION, THE RELOAD WAS PLACED AT THE TIME OF FIRING BUT IT DID NOT PERFORM FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | EEA | Failure to Align | A050701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A HIGHER ANTERIOR RESECTION WHEN THE ANVIL AND THE STAPLER WERE TIGHTENED, AND GREEN MARK WAS INDICATED ONCE, AND SAFETY WAS FUNCTIONAL. AFTER THAT, THE INTESTINAL TRACT WAS FOUND TO BE TWISTED, THE PRODUCT WAS ROTATED SLIGHTLY IN THE OPEN DIRECTION AND THE TWIST ISSUE WAS SOLVED. WHEN THE DOCTOR TRIED TO TIGHTEN THE PRODUCT AGAIN, THE ANVIL WAS FOUND TO BE TILTED AND IT WAS IN THE CONDITION THAT UNABLE TO TIGHTEN THE PRODUCT PROPERLY. THE TILT ISSUE WAS SOLVED BY PRESSING WITH FORCEPS, ATTACHMENT WAS PERFORMED, AND FIRING WAS PERFORMED. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT HARM.
01-08-2019 | Malfunction | covidien medical products | 06-01-2020 | PREMIUM PLUS CEEA | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROUND SIGMOID SURGERY, THE HEAD WAS DETACHED AFTER TRIGGERING THE ANASTOMOSIS AND INITIALLY REMAINED IN THE COLON. THE HEAD WAS REMOVED USING GRAIN FORCEPS WITHOUT DAMAGING THE COLON. THERE WAS NO PATIENT INJURY.
24-12-2019 | Injury | us surgical puerto rico | 06-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A THORASCOPIC WEDGE RESECTION WHEN STAPLING THRU THE LUNG TISSUE, THERE WERE NO STAPLES IN THE PROXIMAL END. THE STAPLE LINE WAS INCOMPLETE. THE DOCTOR PERFORMED MANUAL SUTURING TO THE PART THAT HAS BEEN CUT BUT WAS NOT STAPLED. THERE WAS EXTENDED HOSPITAL STAY DUE TO THE ISSUE.
08-08-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC COLORECTAL PROCEDURE, WHILE DETACHING THE INTESTINAL CANAL, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/ LOBECTOMY, THE RELOAD WAS ATTACHED TO THE HANDLE AND THE FIRST FIRING WAS COMPLETED WITHOUT PROBLEM. AT THE SECOND FIRING, THERE WAS NO PROBLEM WITH THE JAWS OPEN/CLOSE AND APPROACHING, BUT AFTER THE GREEN BUTTON WAS PRESSED, THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. WHEN A NEW CARTRIDGE WAS OPENED TO DEAL WITH THE ISSUE, FIRING WAS ABLE TO BE COMPLETED WITHOUT PROBLEM. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
25-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS USED IN LOWER LOBE DURING A THORACO WEDGE SEGMENTAL RESECTION, IT WAS STATED THAT A CRACKING SOUND WAS HEARD SEVERAL TIMES. THE SURGEON WAS ABLE TO SQUEEZED THE HANDLE HOWEVER THE DEVICE DID NOT FIRE. THERE WAS NO PROBLEM WITH THE STAPLES. A NEW HANDLE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
25-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS USED IN LOWER LOBE DURING A THORACO WEDGE SEGMENTAL RESECTION, IT WAS STATED THAT A CRACKING SOUND WAS HEARD SEVERAL TIMES. THE SURGEON WAS ABLE TO SQUEEZED THE HANDLE HOWEVER THE DEVICE DID NOT FIRE. THERE WAS NO PROBLEM WITH THE STAPLES. A NEW HANDLE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
08-08-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LUNG CANCER SURGERY, UPON CUTTING AND CLOSING PULMONARY BLOOD VESSELS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE DEVICE WAS REPLACED WITH THE SAME SURGICAL INSTRUMENT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOADING OF HERNIA TACKER ON A LAPAROSCOPIC HERNIA REPAIR, THE DEVICE WAS DIFFICULT TO LOAD AND UNABLE TO REMOVE THE YELLOW PIECE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | RELIATACK | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOADING OF HERNIA TACKER ON A LAPAROSCOPIC HERNIA REPAIR, THE DEVICE WAS DIFFICULT TO LOAD AND UNABLE TO REMOVE THE YELLOW PIECE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | RELIATACK | Component or Accessory Incompatibility | A1701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LOADING OF HERNIA TACKER ON A LAPAROSCOPIC HERNIA REPAIR, THE DEVICE WAS DIFFICULT TO LOAD AND UNABLE TO REMOVE THE YELLOW PIECE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACO/LOBECTOMY, UPON THE BLOOD VESSEL RESECTION, THE GREEN BUTTON OF THE DEVICE WAS PRESSED BUT THE KNIFE STOPPED DURING FIRING. THE KNIFE HAD NOT REACHED THE BLOOD VESSEL. AS THE KNIFE WAS UNABLE TO BE ADVANCED AS IT WAS, THE KNIFE WAS RETURNED AND THE JAWS WERE OPENED AND RELEASED. THE DEVICE WAS REMOVED FROM THE INNER CHEST CAVITY ONCE, AND THE RELOAD WAS REPLACED. THE SAME HANDLE WAS USED AND IT WAS ABLE TO BE FIRED WITHOUT PROBLEMS. WHEN THE RELOAD WITH ISSUE WAS CHECKED, A BLACK PART WAS NOTED TO BE CAUGHT(BURIED) IN THE GROOVE WHERE THE KNIFE RAN PREVENTING THE KNIFE BEING UNABLE TO BE ADVANCED. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
25-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | PREMIUM PLUS CEEA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL GASTRECTOMY JEJUNOSTOMY, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | SIGNIA | Failure to Charge | A070603 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC SIGMOIDECTOMY, INDICATOR OF THE CHARGER WAS FLASHING RED. THOUGH THE POWER HANDLE WAS REMOVED, IT WAS IN A STATUS OF BEING VERY HOT AND NOT STARTING UP. AGAIN, EVEN THOUGH IT WAS SET TO THE CHARGER, IT FLASHED RED. THERE WAS A TRACE THAT THE PLASTIC NEAR THE CHARGER'S CONNECTION PLUG WITH THE POWER HANDLE SEEMED BEING MELTED DUE TO THE HEAT. THERE WAS NO PATIENT INVOLVED.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | SIGNIA | Melted | A040610 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC SIGMOIDECTOMY, INDICATOR OF THE CHARGER WAS FLASHING RED. THOUGH THE POWER HANDLE WAS REMOVED, IT WAS IN A STATUS OF BEING VERY HOT AND NOT STARTING UP. AGAIN, EVEN THOUGH IT WAS SET TO THE CHARGER, IT FLASHED RED. THERE WAS A TRACE THAT THE PLASTIC NEAR THE CHARGER'S CONNECTION PLUG WITH THE POWER HANDLE SEEMED BEING MELTED DUE TO THE HEAT. THERE WAS NO PATIENT INVOLVED.
08-08-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC COLORECTAL PROCEDURE, WHILE DETACHING THE INTESTINAL CANAL, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
24-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
08-08-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LUNG CANCER SURGERY, UPON CUTTING AND CLOSING PULMONARY BLOOD VESSELS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE DEVICE WAS REPLACED WITH THE SAME SURGICAL INSTRUMENT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE RELOAD WAS ATTACHED TO THE HANDLE. AT THE FIRST FIRING WHEN TRYING TO RESECT THE RECTUM, THE FIRING WAS UNABLE TO BE PERFORMED. IT WAS REPLACED WITH ANOTHER CARTRIDGE, AND THERE WAS NO PROBLEM IN USE. THERE WAS NO PATIENT INJURY.
17-11-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED LOBECTOMY PROCEDURE, AFTER FIRING THE SURGEON TRIED TO LOAD A NEW RELOAD INTO THE HANDLE, HOWEVER IT COULD NOT BE LOADED. THE SURGEON THEN USED ANOTHER HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST FIRING ON A LAPAROSCOPIC SIGMOIDECTOMY, THE RELOAD WAS MADE COMPLETELY CLOSED BUT THE GAP OF THE RELOAD WAS FOUND TO BE OPENED WIDE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
25-12-2019 | Malfunction | covidien medical products | 06-01-2020 | PREMIUM PLUS CEEA | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION PROCEDURE, THE DEVICE DID NOT CUT THE TISSUE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
24-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE FAILED TO FIRE ANY RELOAD. REPLACEMENT HANDLE WAS USED TO CONTINUE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN DETACHING THE ESOPHAGUS DURING AN ESOPHAGEAL CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. ANOTHER RELOAD AND HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE STAPLER AND RELOAD WERE BEING APPLIED TO THE DORSAL VASCULAR COMPLEX DURING A LAPAROSCOPIC ROBOTIC PROSTATECTOMY, THE JAWS OF THE STAPLER WERE UNABLE TO CLOSE ON THE VENOUS COMPLEX. ANOTHER STAPLER AND RELOAD WERE OPENED. THE DEVICE WAS ADJUSTED CLOSED ON THE TISSUE AND IT COULD STAPLE WITHOUT ISSUE.
24-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN DETACHING THE ESOPHAGUS DURING AN ESOPHAGEAL CANCER SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. ANOTHER RELOAD AND HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF LIVER CANCER, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HANDLE COULD NOT BE SQUEEZED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE STAPLER AND RELOAD WERE BEING APPLIED TO THE DORSAL VASCULAR COMPLEX DURING A LAPAROSCOPIC ROBOTIC PROSTATECTOMY, THE JAWS OF THE STAPLER WERE UNABLE TO CLOSE ON THE VENOUS COMPLEX. ANOTHER STAPLER AND RELOAD WERE OPENED. THE DEVICE WAS ADJUSTED CLOSED ON THE TISSUE AND IT COULD STAPLE WITHOUT ISSUE.
26-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW RECTAL RESECTION, ON DISCONNECTION, THE STAPLER WAS NOT ABLE TO FIRE, THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE WAS NOT SQUEEZED. THEY CHANGED THE RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
26-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
26-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
26-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
26-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
26-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE.
12-11-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | 75MM SELECTABLE NEW TLC | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: DAMAGED DEVICE. IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY SURGERY, WHEN 2ND FIRING, NOTED ONE METAL PIECE WAS DETACHED FROM THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
26-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN DOING COLON TRANSECTION ON A LAPAROSCOPIC COLECTOMY, THE STAPLER WAS ABLE TO FIRE BUT STAPLES DID NOT FORM A PERFECT ¿B¿ SHAPE CENTRALLY. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | ECHELON FLEX 45 STANDARD | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE VATS LEFT UPPER LOBECTOMY SURGERY, COULD NOT FIRE WHEN SEVERING PULMONARY FISSURE TISSUE. CHANGED ANOTHER ECR60G, COULD NOT FIRE, AND CHANGED ANOTHER ECR60G, THE CLOSURE TRIGGER WAS BROKEN. CHANGED ANOTHER NEW DEVICE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-11-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE VATS LEFT UPPER LOBECTOMY SURGERY, COULD NOT FIRE WHEN SEVERING PULMONARY FISSURE TISSUE. CHANGED ANOTHER ECR60G, COULD NOT FIRE, AND CHANGED ANOTHER ECR60G, THE CLOSURE TRIGGER WAS BROKEN. CHANGED ANOTHER NEW DEVICE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE THERE WAS AN ISSUE WITH AN ATS45 WHERE THE DEVICE PARTIALLY FIRED ON HUGE SPLENIC ARTERY NEITHER SEALING IT OR FULLY DIVIDING IT. LUCKILY WE MANAGED TO CONTROL THE ARTERY BUT THIS WAS A HAIR AWAY FROM AN UNCONTROLLABLE HEMORRHAGE.
12-12-2019 | Injury | ethicon endo-surgery | 06-01-2020 | ENDO ILS 29MM, CURVED | Difficult to Open or Close | A051104 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A BOWEL RESECTION PROCEDURE, "AFTER FIRING, THE CIRCULAR STAPLER WAS DIFFICULT TO OPEN, SURGEON OVERSEWED AND PATIENT WAS GIVEN A STOMA". SURGERY WAS DELAYED 180 MINUTES.
18-12-2019 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, WHILE THE RELOAD WAS APPLIED ONTO THE DISTAL END OF THE RECTUM, THE DEVICE WAS NOT ABLE TO BE FIRED. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. SURGEON ALSO EXPERIENCED DIFFICULTY IN UNLOADING THE DEVICE. OPERATOR APPLIED EXCESSIVE PRESSURE TO UNLOAD THE RELOAD. SURGEON USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE DUE TO DEVICE PROBLEM.
18-12-2019 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, WHILE THE RELOAD WAS APPLIED ONTO THE DISTAL END OF THE RECTUM, THE DEVICE WAS NOT ABLE TO BE FIRED. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. SURGEON ALSO EXPERIENCED DIFFICULTY IN UNLOADING THE DEVICE. OPERATOR APPLIED EXCESSIVE PRESSURE TO UNLOAD THE RELOAD. SURGEON USED ANOTHER DEVICE TO RESOLVE THE ISSUE. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE DUE TO DEVICE PROBLEM.
30-11-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF RECTAL CANCER SURGERY, WHEN SEVERING THE RECTUM TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED. DONUTS WERE COMPLETE AND WASHER WAS CUT OFF. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IT WAS APPLIED ON THE ANTRUM SIDE, THE RELOAD WAS NOT GETTING UNLOADED FROM THE ADAPTER, THE STAPLER WAS NOT ABLE TO FIRE. THE JAWS OF THE DEVICE LOCKED ON THE TISSUE AND WAS REMOVED BY PULLING BACK WITH THE HELP OF A GRASPER. THE POWERED STAPLER WAS REPLACED WITH A MANUAL STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IT WAS APPLIED ON THE ANTRUM SIDE, THE RELOAD WAS NOT GETTING UNLOADED FROM THE ADAPTER, THE STAPLER WAS NOT ABLE TO FIRE. THE JAWS OF THE DEVICE LOCKED ON THE TISSUE AND WAS REMOVED BY PULLING BACK WITH THE HELP OF A GRASPER. THE POWERED STAPLER WAS REPLACED WITH A MANUAL STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | UNKNOWN ENDO GIA SULU | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IT WAS APPLIED ON THE ANTRUM SIDE, THE RELOAD WAS NOT GETTING UNLOADED FROM THE ADAPTER, THE STAPLER WAS NOT ABLE TO FIRE. THE JAWS OF THE DEVICE LOCKED ON THE TISSUE AND WAS REMOVED BY PULLING BACK WITH THE HELP OF A GRASPER. THE POWERED STAPLER WAS REPLACED WITH A MANUAL STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IT WAS APPLIED ON THE ANTRUM SIDE, THE RELOAD WAS NOT GETTING UNLOADED FROM THE ADAPTER, THE STAPLER WAS NOT ABLE TO FIRE. THE JAWS OF THE DEVICE LOCKED ON THE TISSUE AND WAS REMOVED BY PULLING BACK WITH THE HELP OF A GRASPER. THE POWERED STAPLER WAS REPLACED WITH A MANUAL STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IT WAS APPLIED ON THE ANTRUM SIDE, THE RELOAD WAS NOT GETTING UNLOADED FROM THE ADAPTER, THE STAPLER WAS NOT ABLE TO FIRE. THE JAWS OF THE DEVICE LOCKED ON THE TISSUE AND WAS REMOVED BY PULLING BACK WITH THE HELP OF A GRASPER. THE POWERED STAPLER WAS REPLACED WITH A MANUAL STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | FLEX 60 ARTICULATING | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC COLON SURGERY, THE DEVICE WOULD ONLY PARTIALLY FIRE WHEN SEVERING COLON TISSUE. KNIFE COULD NOT BE RETURNED, USED ANOTHER DEVICE TO CUT THE TISSUE AND REMOVED THE ARTICULATED DEVICE WITH TROCAR. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | FLEX 60 ARTICULATING | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC COLON SURGERY, THE DEVICE WOULD ONLY PARTIALLY FIRE WHEN SEVERING COLON TISSUE. KNIFE COULD NOT BE RETURNED, USED ANOTHER DEVICE TO CUT THE TISSUE AND REMOVED THE ARTICULATED DEVICE WITH TROCAR. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER WAS FIRED TWO TIMES WITH 2 RELOADS. THE SECOND TIME THE CLIPS DID NOT CLOSE PROPERLY. I HAVE NO MORE INFORMATION AT THIS TIME.
07-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THERE WAS A GREEN RELOAD FAILURE. FOUR STAPLE ROWS FAILED, ONLY TWO LINES FOLLOWED BY THE CUT LINE COULD BE FIRED AND THE STAPLES DID NOT FORM CORRECTLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. IT WAS REPORTED THAT THE PROCEDURE WAS SUCCESSFULLY COMPLETED, BUT IT'S UNKNOWN IF ANY OTHER MEDICAL INTERVENTION WAS REQUIRED.
04-11-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBECTOMY SURGERY, THE DEVICE WOULD NOT FIRE WHEN SEVERING PULMONARY LOBE TISSUE ON THE SIXTH FIRING. THE KNIFE REVERSE BUTTON DID NOT WORK, USED MANUAL OVERRIDE LEVER TO RETURN THE BLADE. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-11-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBECTOMY SURGERY, THE DEVICE WOULD NOT FIRE WHEN SEVERING PULMONARY LOBE TISSUE ON THE SIXTH FIRING. THE KNIFE REVERSE BUTTON DID NOT WORK, USED MANUAL OVERRIDE LEVER TO RETURN THE BLADE. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC NEOBLADDER SURGERY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS UNABLE TO SQUEEZE THE HANDLE, HENCE THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
03-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | PROXIMATE*PPH PROCEDURE SET | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A PPH PROCEDURE, AFTER STAPLING AND REMOVING THE DEVICE, THERE WERE BLEEDING AT 3,6,12 O'CLOCK POSITIONS. THE DOCTOR CHECKED AND FOUND THAT THE TISSUE WAS CUT ALREADY BUT NOT BE STITCHED AND THE PINS STILL INSIDE THE DEVICE. THE DOCTOR PERFORMS SUTURE TOGETHER WITH HANGING HEMORRHOIDS TECHNIQUE FOR PATIENT. THIS PROLONGED SURGERY BY ANESTHETIC FOR AN ADDITIONAL 20 MINUTES , INCREASING THE ANESTHETIC AND HEMORRHOIDS TO HANGING STITCHES IN THE UNPAIRED POSITION. UPDATE ON THE PATIENT'S HEALTH: WAS FINE AND WAS DISCHARGED ON 07-DEC.
06-01-2020 | Malfunction | us surgical puerto rico | 06-01-2020 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | No Information | Event Description: ACCORDING TO THE REPORTER, THERE WAS NO COMPLAINT RECEIVED FOR THE DEVICE.
23-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC NEOBLADDER SURGERY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS UNABLE TO SQUEEZE THE HANDLE, HENCE THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Adhesion(s) | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Bacterial Infection | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Calcium Deposits/Calcification | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Cellulitis | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Diarrhea | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Purulent Discharge | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Dyspnea | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Erythema | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Fatigue | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Hyperbilirubinemia | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Inflammation | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Pulmonary Edema | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Scar Tissue | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Septic Shock | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Seroma | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Tachycardia | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Vomiting | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Chills | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Hypernatremia | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Impaired Healing | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Respiratory Failure | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Ascites | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Distention | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Fibrosis | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Thrombosis/Thrombus | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
26-12-2017 | Injury | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: LEGAL - ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION AND AN INCIDENTAL APPENDECTOMY. A SMALL SEROSAL TEAR CLOSE TO THE POINT OF RESECTION WHICH WAS OVERSEWN AND IMBRICATED WITH A SINGLE -3-0 SILK STITCH WAS NOTED. NO OTHER COMPLICATIONS WERE REPORTED DURING THE PROCEDURE. THE PATIENT WAS ADMITTED TO ICU AFTER DEVELOPING PULMONARY ISSUES AND WAS TREATED WITH ANTIBIOTICS. A CT SCAN WAS DONE ON THE PATIENT AND REVEALED SCATTERED FREE AIR PRESENT THROUGHOUT THE PATIENTS UPPER ABDOMEN AND APPARENT WALL THICKENING OF MULTIPLE SMALL BOWEL LOOP WITHIN THE LOWER LEFT QUADRANT OF THE ABDOMEN. FURTHER COMPLICATIONS IDENTIFIED A SMALL BOWEL ANASTOMOTIC LEAK PRESENT IN THE PATIENTS MIDDLE UPPER ABDOMEN. THE ANASTOMOSIS IS VISUALIZED TO REVEAL A LEAK COMING FROM THE UPPER PORTION OF THE TA STAPLE LINE. THIS DEVELOPED INTRA-ABDOMINAL ABSCESS AND REQUIRED FURTHER SURGICAL INTENTIONS FOR THE PATIENT.
31-10-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | ECHELON 60 ENDOPATH STAPLER | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC GASTRECTOMY SURGERY, THE DEVICE WOULD NOT FIRE WHEN SEVERING DUODENUM TISSUE ON THE 1ST FIRING. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN ENTERECTOMY, DURING THE DETACHMENT OF THE SMALL INTESTINE, THE DEVICE DID NOT FIRE. THE BLACK PUSHER THUMB PIECE COUNT NOT BE MOVED AT ALL. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE FIFTH FIRING ON SLEEVE GASTRECTOMY, THE OUTSIDE ROW OF STAPLES HAD A COMPLETELY MALFORMED STAPLE PROTRUDING OUT OF STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED AND NO OTHER MEDICAL INTERVENTION WAS REQUIRED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
19-07-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Insufficient Information | E2401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DOCTOR HAD AN UNKNOWN ISSUE WITH THE DEVICE. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED OR IF THERE WERE ANY CONSEQUENCES TO THE PATIENT.
09-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN EXPLORATORY LAPAROTOMY, THE STAPLER JAMMED ON THE FOURTH FIRING. THE STAPLER WOULD NOT FIRE AFTER THE FIFTH RELOAD WAS INSERTED AND LOCKED OUT ON TISSUE. THE CUSTOMER TRIED TO USE MANUAL OVERRIDE AND THE STAPLER WOULD NOT OPEN. A SECOND LIKE STAPLER AND RELOAD WAS USED TO STAPLER AROUND THE STUCK TISSUE. THERE WAS A TEN-MINUTE DELAY IN THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THIS IS ALL THE INFORMATION KNOWN/AVAILABLE REGARDING THIS EVENT.
09-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN EXPLORATORY LAPAROTOMY, THE STAPLER JAMMED ON THE FOURTH FIRING. THE STAPLER WOULD NOT FIRE AFTER THE FIFTH RELOAD WAS INSERTED AND LOCKED OUT ON TISSUE. THE CUSTOMER TRIED TO USE MANUAL OVERRIDE AND THE STAPLER WOULD NOT OPEN. A SECOND LIKE STAPLER AND RELOAD WAS USED TO STAPLER AROUND THE STUCK TISSUE. THERE WAS A TEN-MINUTE DELAY IN THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THIS IS ALL THE INFORMATION KNOWN/AVAILABLE REGARDING THIS EVENT.
09-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A VATS SEGMENTAL RESECTION OF RIGHT LUNG, THE KNIFE STOPPED WHILE FIRING. THE DEVICE WAS USED ON THE LUNG PARENCHYMA, AND THE TARGET TISSUE WAS THICK. THE KNIFE DID NOT RETURN TO HOME POSITION, AND THE KNIFE REVERSE SWITCH AND THE MANUAL OVERRIDE LEVER WERE USED. THE SURGEON CONVERTED THE PROCEDURE SINCE THE JAWS REMAINED CLAMPING THE TISSUE, AND ADDITIONAL SUTURE WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE REP THAT THE DEVICES DID NOT WORK. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
28-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLON CANCER PROCEDURE, THE STAPLING WAS NOT DONE PERFECTLY, SO THE BLEEDING OCCURRED. THEY EXCHANGED THE PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-09-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RESECTION, POOR STAPLE FORMATION WAS NOTED EVEN THOUGH THE FIRING CYCLE WAS COMPLETED. THE STAPLE LINE WAS ALSO INCOMPLETE CENTRALLY. UNFORMED STAPLES FROM TWO RELOADS FELL INTO THE PATIENT'S CAVITY AND WAS MANUALLY PICKED, RINSED AND SUCKED. THIS CAUSED LEAKAGE IN THE LUNG AND FORCES THE SURGEON TO RESECT TISSUE AGAIN, WHICH LED TO TISSUE LOSS AND TISSUE DAMAGE. FIRST THE HANDLE WAS REPLACED, THEN ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
30-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT HEMICOLECTOMY PROCEDURE, WHILE DOING A TRANSVERSE ASCENDING COLON, THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON CHANGED THE DEVICE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT HEMICOLECTOMY PROCEDURE, WHILE DOING A TRANSVERSE ASCENDING COLON, THE DEVICE WAS NOT ABLE TO FIRE. SURGEON CHANGED THE DEVICE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY PROCEDURE, DURING PULMONARY ARTERY DISCONNECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS UNABLE TO SQUEEZE THE HANDLE, HENCE THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | UNKNOWN ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY PROCEDURE, DURING PULMONARY ARTERY DISCONNECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS UNABLE TO SQUEEZE THE HANDLE, HENCE THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY PROCEDURE, DURING PULMONARY ARTERY DISCONNECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS UNABLE TO SQUEEZE THE HANDLE, HENCE THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A WHIPPLE PROCEDURE WHEN FIRING IN THE STOMACH, THE STAPLES OF THE RELOAD WERE MALFORMED AND THE FIRING WAS UNABLE TO BE COMPLETED, THE STAPLE LINE WAS INCOMPLETE IN THE CENTER PART. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 06-01-2020 | UNKNOWN EGIA SULU | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATIC SEGMENTECTOMY PROCEDURE, SOME OF THE STAPLES WERE MALFORMED. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, THE SURGEON COULD NOT PRESSED THE GREEN BUTTON BUT WAS ABLE TO SQUEEZED THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES . MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES . MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 06-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES . MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | ENDOSCOPIC LINEAR CUTTER LAPAROSCOPIC STAPLER | Break | A0401 | No Known Impact Or Consequence To Patient | Event Description: DURING ROUTINE LAPAROSCOPIC APPENDECTOMY THE LAPAROSCOPIC STAPLER MISFIRED WHILE USING A BLUE STAPLE LOAD. THE EQUIPMENT IN AN ETHICON ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER, 45 MM STAPLE LINE. NO TISSUE WAS DAMAGED FROM THE MISFIRE. ANOTHER STAPLER AND STAPLE LOAD WAS INTRODUCED TO THE FIELD AND USED SUCCESSFULLY. FDA MEDWATCH COMPLETED. FDA SAFETY REPORT ID # (B)(4).
24-12-2019 | Malfunction | ethicon endo-surgery | 06-01-2020 | ENDOSCOPIC LINEAR CUTTER LAPAROSCOPIC STAPLER | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: DURING ROUTINE LAPAROSCOPIC APPENDECTOMY THE LAPAROSCOPIC STAPLER MISFIRED WHILE USING A BLUE STAPLE LOAD. THE EQUIPMENT IN AN ETHICON ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER, 45 MM STAPLE LINE. NO TISSUE WAS DAMAGED FROM THE MISFIRE. ANOTHER STAPLER AND STAPLE LOAD WAS INTRODUCED TO THE FIELD AND USED SUCCESSFULLY. FDA MEDWATCH COMPLETED. FDA SAFETY REPORT ID # (B)(4).
27-11-2019 | Malfunction | ethicon endo-surgery | 07-01-2020 | 55MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLES DID NOT CLOSE PROPERLY: SHAPE "B" MALFORMED. THE PROCEDURE WAS EXTENDED TO 15 MINUTES. THE STAPLING AND THE CUT WAS RE-PERFORMED WITH ANOTHER DEVICE.
29-10-2019 | Malfunction | ethicon endo-surgery | 07-01-2020 | ILS 29MM, CURVED | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, WHEN THE SURGEON PULLED A SAFETY BAR, THE SAFETY BAR WAS BROKEN. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHEN THE FIVE STAPLER HANDLES WERE USED AND EACH WHEN CLOSED DOWN ON THE TISSUE WOULD NOT CLOSE FULLY AND MAKE A CRACKING NOISE. IN ADDITION, THE HANDLES WOULD NOT RELEASE TO OPEN THE JAWS AND LOCKED ON TISSUE. THIS OCCURRED ON FOUR REPORTED HANDLES. MORE STAPLER HANDLES WERE PULLED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Failure to Fire | A050501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Difficult to Open or Close | A051104 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
23-11-2019 | Injury | covidien lp llc north haven | 07-01-2020 | EEA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING OPEN HEMORRHOID PROCEDURE, UPON FORMING THE STAPLES, THE DEVICE LACK OF 1/4 STAPLE CYCLE LEAD TO SEVERE BLEEDING AND SURGEON HAD TO STOP BLEEDING BY SUTURING AND EMBOLIZATION WAS DONE DUE TO HEMORRHOIDS. THERE WAS TISSUE DAMAGE AND INCISION WAS EXTENDED FOR MORE THAN 1 INCH AS A RESULT OF THIS PROBLEM. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE PATIENT HAD TO STAY 5 DAYS MORE IN THE HOSPITAL BECAUSE OF SURGICAL PERFORMANCE WAS NOT GOOD AND PATIENT EXPERIENCED DIFFICULTY WITH BOWEL MOVEMENT.
23-11-2019 | Injury | covidien lp llc north haven | 07-01-2020 | EEA | Failure to Form Staple | A050704 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING OPEN HEMORRHOID PROCEDURE, UPON FORMING THE STAPLES, THE DEVICE LACK OF 1/4 STAPLE CYCLE LEAD TO SEVERE BLEEDING AND SURGEON HAD TO STOP BLEEDING BY SUTURING AND EMBOLIZATION WAS DONE DUE TO HEMORRHOIDS. THERE WAS TISSUE DAMAGE AND INCISION WAS EXTENDED FOR MORE THAN 1 INCH AS A RESULT OF THIS PROBLEM. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE PATIENT HAD TO STAY 5 DAYS MORE IN THE HOSPITAL BECAUSE OF SURGICAL PERFORMANCE WAS NOT GOOD AND PATIENT EXPERIENCED DIFFICULTY WITH BOWEL MOVEMENT.
23-11-2019 | Injury | covidien lp llc north haven | 07-01-2020 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING OPEN HEMORRHOID PROCEDURE, UPON FORMING THE STAPLES, THE DEVICE LACK OF 1/4 STAPLE CYCLE LEAD TO SEVERE BLEEDING AND SURGEON HAD TO STOP BLEEDING BY SUTURING AND EMBOLIZATION WAS DONE DUE TO HEMORRHOIDS. THERE WAS TISSUE DAMAGE AND INCISION WAS EXTENDED FOR MORE THAN 1 INCH AS A RESULT OF THIS PROBLEM. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE PATIENT HAD TO STAY 5 DAYS MORE IN THE HOSPITAL BECAUSE OF SURGICAL PERFORMANCE WAS NOT GOOD AND PATIENT EXPERIENCED DIFFICULTY WITH BOWEL MOVEMENT.
27-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC LOBECTOMY WHEN IT WAS FIRED THE THIRD TIME, THE HANDLE COULD NOT NORMALLY FIRE BOTH RELOADS. THE STAPLES DIDN'T COME OUT WHEN FIRING. HANDLE AND RELOADS WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Failure to Fire | A050501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Difficult to Open or Close | A051104 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | SIGNIA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
29-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC LOBECTOMY WHEN IT WAS FIRED THE THIRD TIME, THE HANDLE COULD NOT NORMALLY FIRE BOTH RELOADS. THE STAPLES DIDN'T COME OUT WHEN FIRING. HANDLE AND RELOADS WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
18-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A LOW ANTERIOR RESECTION, THEY TRIED TO PERFORM STAPLING ON THE RECTUM WITH THE COMBINATION OF POWERED STAPLER AND RELOAD BUT COULD NOT FIRE, AND FURTHERMORE THE JAWS BECAME UNABLE TO OPEN. AT THAT SITUATION, THEY REMOVED THE ADAPTER FROM THE HANDLE AND USED A MANUAL RETRACTION TOOL, BUT THE ADAPTER TOOL COULD NOT TURN, SO ADDITIONAL TISSUE RESECTION WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MIN. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER ON RESECTION OF THE RECTUM, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY REPLACED WITH A PRODUCT OF ANOTHER BRAND TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC SURGERY RIGHT LOWER LUNG RESECTION, WHEN ON SEVERED THE BLOOD VESSELS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE TWO STAPLERS MISFIRED AND THE STAPLES FELL INTO THE PATIENT'S CAVITY AND WERE RETRIEVED BY THE SURGEON. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER ON RESECTION OF THE RECTUM, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY REPLACED WITH A PRODUCT OF ANOTHER BRAND TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS UNABLE TO SQUEEZE THE HANDLE, HENCE THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RIGHT COLON CANCER PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LEFT LOWER LOBE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE AND THE DEVICE WAS NOT ABLE TO FIRE. SURGEON USED COMPETITOR'S DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
09-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, A HEMOPERITONEUM WAS OBSERVED. THE SURGEON HAD TO RE-OPERATE THE PATIENT TO RESOLVE THE ISSUE.
18-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A BILROTH I PROCEDURE, THE STAPLER DID CUT THE TISSUE BUT THE CARTRIDGE GOT STUCK AFTER FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Premature End-of-Life Indicator | A160602 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Unraveled Material | A040508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Device Displays Incorrect Message | A090201 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Positioning Problem | A1502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Injury | us surgical puerto rico | 07-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AT THE FIRST FIRE ON THE ANTRUM, THE DEVICE WAS ABLE TO FIRE AND HAD AN UNSPECIFIED ISSUE. SUTURING AND FIRING ANOTHER CARTRIDGE WAS PERFORMED TO RESOLVE THE ISSUE.
26-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE PANCREAS WAS BEING CUT IN HALF USING BLUE RELOAD IN AN OPEN PANCREATECTOMY PROCEDURE, THE DEVICE WAS UNABLE TO FIRE. THE RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LEFT LOWER LOBE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE AND THE DEVICE WAS NOT ABLE TO FIRE. SURGEON USED COMPETITOR'S DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Positioning Problem | A1502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
12-12-2019 | Malfunction | ethicon endo-surgery | 07-01-2020 | ILS 29MM, CURVED | Activation Problem | A1501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, AS PER THE DOCTOR AFTER FIRING THE STAPLER, THE STAPLER DID NOT FULLY DEPLOY. AROUND 25% OF THE COLON WAS ONLY ANASTOMOSED. SURGERY WAS DELAYED 45 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED AND NO PATIENT CONSEQUENCES WERE REPORTED.
01-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | UNKNOWN ENDO GIA SULU | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE TWO STAPLERS MISFIRED AND THE STAPLES FELL INTO THE PATIENT'S CAVITY AND WERE RETRIEVED BY THE SURGEON. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING REMOVAL OF DUODENUM ON A LAPAROSCOPIC PANCREATICODUODENECTOMY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HANDLE COULD NOT BE SQUEEZED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, THERE WAS POOR STAPLE FORMATION AND THAT THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS OVER-SEWN TO FIX THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
31-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, THERE WAS POOR STAPLE FORMATION AND THAT THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS OVER-SEWN TO FIX THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING REMOVAL OF DUODENUM ON A LAPAROSCOPIC PANCREATICODUODENECTOMY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HANDLE COULD NOT BE SQUEEZED. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Injury | us surgical puerto rico | 07-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, THERE WAS POOR STAPLE FORMATION AND THAT THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS OVER-SEWN TO FIX THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Positioning Problem | A1502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE JAWS WOULD NOT OPEN AND WOULD NOT ARTICULATE RIGHT AFTER INITIAL FIRING. A NEW SHELL AND HANDLE WAS USED BUT IT FELL ONTO THE FLOOR. SO A NEW SHELL AND HANDLE WERE USED TOGETHER WITH A THIRD ADAPTER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
26-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY LOBECTOMY PROCEDURE, WHILE APPLYING THE DEVICE ON THE PULMONARY PARENCHYMA, THE STAPLES COULD NOT FORM A B SHAPE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING THE COLON DURING A LAPAROSCOPIC LEFT COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE; HENCE THE STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
26-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE.
26-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE.
29-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HEPATECTOMY PROCEDURE, WHEN SURGEON TRIED TO FIRE THE DEVICE FOR THE THIRD TIME, THE HANDLE COULD NOT FIRE THE RELOAD. THE SURGEON THEN USED ANOTHER DEVICE HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
29-11-2019 | Injury | us surgical puerto rico | 07-01-2020 | EEA | Difficult to Remove | A150207 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHEN CLOSING THE ANVIL DURING AN OPEN ARTIFICIAL ANAL ANATOMY SURGERY, THE DEVICE COULD NOT CLOSE THE ANVIL TO FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PATIENT HAD TISSUE DAMAGE, SEVERE BLEEDING AND SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE AS A RESULT OF THE EVENT.
29-11-2019 | Injury | us surgical puerto rico | 07-01-2020 | EEA | Difficult to Remove | A150207 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, WHEN CLOSING THE ANVIL DURING AN OPEN ARTIFICIAL ANAL ANATOMY SURGERY, THE DEVICE COULD NOT CLOSE THE ANVIL TO FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PATIENT HAD TISSUE DAMAGE, SEVERE BLEEDING AND SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE AS A RESULT OF THE EVENT.
29-11-2019 | Injury | us surgical puerto rico | 07-01-2020 | EEA | Difficult to Remove | A150207 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN CLOSING THE ANVIL DURING AN OPEN ARTIFICIAL ANAL ANATOMY SURGERY, THE DEVICE COULD NOT CLOSE THE ANVIL TO FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PATIENT HAD TISSUE DAMAGE, SEVERE BLEEDING AND SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE AS A RESULT OF THE EVENT.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING THE COLON DURING A LAPAROSCOPIC LEFT COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE; HENCE THE STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE STAPLER WAS ABLE TO FIRE, THERE WAS POOR STAPLE FORMATION, THE B SHAPE WAS INCOMPLETE. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THEY SUTURED TO FIX THE STAPLE LINE AND COMPLETE THE CASE.
29-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC LOBECTOMY, WHEN IT WAS FIRED THE THIRD TIME, THE HANDLE COULD NOT NORMALLY FIRE BOTH RELOADS. THE STAPLES DIDN'T COME OUT WHEN FIRING. HANDLE AND RELOADS WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE.
26-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RIGHT COLON CANCER PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Component or Accessory Incompatibility | A1701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Component or Accessory Incompatibility | A1701 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, AND THE JAWS OF THE RELOAD COULD NOT BE OPENED.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
30-12-2019 | Injury | us surgical puerto rico | 07-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A THORASCOPIC LOBECTOMY, WHEN PERFORMING A LUNG LOBE RESECTION, THE STAPLES WERE ONLY PARTIALLY FORMED AND THE STAPLE LINE IS INCOMPLETE IN THE REMOTE PART. THE SURGEON USED SUTURE REINFORCEMENT TO STITCH THE TISSUE THAT WAS NOT FULLY STAPLED.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY ON STAPLING VESSEL, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY ON STAPLING VESSEL, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY ON STAPLING VESSEL, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE. THEY CHANGED THE DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | EEA | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A ROBOTIC LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER FIRING THE DEVICE AND CREATING ANASTOMOSES OF THE COLON AND RECTUM, THERE WAS DIFFICULTY IN REMOVING THE DEVICE FROM THE CAVITY. THE DEVICE WAS PUSHED FORWARD FOR ABOUT 1 MILLIMETER, ROTATED A QUARTER TO THE RIGHT AND TO THE LEFT, AND REPEATED. HOWEVER, THE DEVICE STILL COULD NOT BE REMOVED. THE SURGEON HAD TO USE A GRASPER TO HOLD ONTO THE DISTAL COLON TO REMOVE THE DEVICE EASILY. THERE WAS NO PATIENT INJURY.
30-12-2019 | Injury | us surgical puerto rico | 07-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A THORASCOPIC LOBECTOMY, WHEN PERFORMING A LUNG LOBE RESECTION, THE STAPLES WERE ONLY PARTIALLY FORMED AND THE STAPLE LINE IS INCOMPLETE IN THE REMOTE PART. THE SURGEON USED SUTURE REINFORCEMENT TO STITCH THE TISSUE THAT WAS NOT FULLY STAPLED.
30-12-2019 | Injury | us surgical puerto rico | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER DURING A THORASCOPIC LOBECTOMY, WHEN PERFORMING A LUNG LOBE RESECTION, THE STAPLES WERE ONLY PARTIALLY FORMED AND THE STAPLE LINE IS INCOMPLETE IN THE REMOTE PART. THE SURGEON USED SUTURE REINFORCEMENT TO STITCH THE TISSUE THAT WAS NOT FULLY STAPLED.
25-12-2019 | Malfunction | celestica de monterrey s.a de c.v | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, WHEN THE FIRST 45-2.5MM RELOAD WAS BEING APPLIED TO THE BLOOD VESSEL, THE HANDLE OF THE DEVICE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. AFTER CHANGING THE RELOAD TO ANOTHER 45-2.5MM RELOAD, THE FIRING WAS SUCCESSFUL. WHILE TREATING THE INTERSTITIAL TISSUE, A 60MM RELOAD WAS USED BUT THE SAME PROBLEM AS THE FIRST RELOAD WAS ENCOUNTERED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO INJURY.
28-11-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON THE 6TH APPLICATION OF STAPLE ON THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS ABLE TO FIRE BUT THERE WAS A POOR STAPLE FORMATION AND STAPLE LINE WAS INCOMPLETE CENTRALLY. ANOTHER RELOAD WAS USED TO FIX THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
11-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Misfire | A050502 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY PROCEDURE, THE SURGEON FOUND THAT THERE IS NO CORRECT STAPLE FORMATION, AND HAD AN INCOMPLETE STAPLE LINE DISTALLY THAT LED TO BLOOD LOSS. THE SURGEON THEN HAD TO MANUALLY SUTURE THE TISSUE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
11-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY PROCEDURE, THE SURGEON FOUND THAT THERE IS NO CORRECT STAPLE FORMATION, AND HAD AN INCOMPLETE STAPLE LINE DISTALLY THAT LED TO BLOOD LOSS. THE SURGEON THEN HAD TO MANUALLY SUTURE THE TISSUE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
11-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY PROCEDURE, THE SURGEON FOUND THAT THERE IS NO CORRECT STAPLE FORMATION, AND HAD AN INCOMPLETE STAPLE LINE DISTALLY THAT LED TO BLOOD LOSS. THE SURGEON THEN HAD TO MANUALLY SUTURE THE TISSUE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
11-12-2019 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY PROCEDURE, THE SURGEON FOUND THAT THERE IS NO CORRECT STAPLE FORMATION, AND HAD AN INCOMPLETE STAPLE LINE DISTALLY THAT LED TO BLOOD LOSS. THE SURGEON THEN HAD TO MANUALLY SUTURE THE TISSUE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
25-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, WHEN THE FIRST 45-2.5MM RELOAD WAS BEING APPLIED TO THE BLOOD VESSEL, THE HANDLE OF THE DEVICE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. AFTER CHANGING THE RELOAD TO ANOTHER 45-2.5MM RELOAD, THE FIRING WAS SUCCESSFUL. WHILE TREATING THE INTERSTITIAL TISSUE, A 60MM RELOAD WAS USED BUT THE SAME PROBLEM AS THE FIRST RELOAD WAS ENCOUNTERED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO INJURY.
04-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THE STAPLER LOCKED AT FIRING AND WAS NOT ABLE TO FIRE. THEY REPLACED WITH A NEW LOAD AND CONTINUED TO LOCK. THEY REPLACED WITH A NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THE STAPLER LOCKED AT FIRING AND WAS NOT ABLE TO FIRE. THEY REPLACED WITH A NEW LOAD AND CONTINUED TO LOCK. THEY REPLACED WITH A NEW STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE DEVICE WAS ABLE TO BE FIRED, HOWEVER ONLY 3 NAILS FORMED PROPERLY. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESECTION OF SMALL INTESTINE TUMOR PROCEDURE WHILE DISCONNECTING THE INTESTINE, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. IT WAS ALSO REPORTED THAT THE DEVICE WAS NOT FULLY FIRED AND WAS OPENED AFTER BEING PULLED BACK. ANOTHER RELOAD WAS USED BUT THE SAME ISSUE PERSISTED. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
22-02-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | EEA | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY PROCEDURE, THE DEVICE'S OGIVA AND PIN WERE BROKEN. IT WAS ALSO REPORTED THAT THE OGIVA DID NOT REMAIN COUPLING PROPERLY IN THE CONNECTOR, EVEN AFTER REMOVING THE ORANGE LOCK, PREVENTING APPROACH TO TISSUE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
22-02-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | EEA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY PROCEDURE, THE DEVICE'S OGIVA AND PIN WERE BROKEN. IT WAS ALSO REPORTED THAT THE OGIVA DID NOT REMAIN COUPLING PROPERLY IN THE CONNECTOR, EVEN AFTER REMOVING THE ORANGE LOCK, PREVENTING APPROACH TO TISSUE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Malfunction | us surgical puerto rico | 07-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESECTION OF SMALL INTESTINE TUMOR PROCEDURE WHILE DISCONNECTING THE INTESTINE, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. IT WAS ALSO REPORTED THAT THE DEVICE WAS NOT FULLY FIRED AND WAS OPENED AFTER BEING PULLED BACK. ANOTHER RELOAD WAS USED BUT THE SAME ISSUE PERSISTED. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY PROCEDURE, THE RELOAD GOT STUCK AT THE START OF FIRING PROCESS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Positioning Problem | A1502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO LIGATION AND DIVISION OF PULMONARY VEIN DURING A VIDEO ASSISTED THO RADIOSCOPIC LOBECTOMY PROCEDURE, IT WAS STATED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE FIRST RELOAD WAS INSERTED WITH THE POWER HANDLE THROUGH THE WINDOW PORT OF PATIENT AND THE PULMONARY VEIN WAS CLAMPED WITH RELOADS COMPLETELY AND PUSHED THE GREEN SAFETY BUTTON. WHEN THE TOGGLE BUTTON WAS PUSHED DOWN TO FIRE THE RELOAD, IT WAS OBSERVED THAT THE BLADE GOT STOPPED AT THE BEGINNING PHASE OF THE FIRING PROCESS. THE TOGGLE BUTTON WAS TRIED TO PUSH DOWN ONE MORE TIME HOWEVER THE BLADE DID NOT MOVE AND PUSHED IT UP TO RETRACT THE BLADE. ANOTHER RELOAD WAS THEN USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS AN ERROR SOUND HEARD FROM THE HANDLE, AS ARTICULATING THE RELOAD IN LEFT SIDE. THE SCRUB NURSE DISCONNECTED THE NEW RELOAD FROM THE HANDLE AND REPLACED THE POWER HANDLE SET WITH ANOTHER ONE AND THE RELOAD WAS SUCCESSFULLY FIRED. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | SIGNIA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO LIGATION AND DIVISION OF PULMONARY VEIN DURING A VIDEO ASSISTED THO RADIOSCOPIC LOBECTOMY PROCEDURE, IT WAS STATED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE FIRST RELOAD WAS INSERTED WITH THE POWER HANDLE THROUGH THE WINDOW PORT OF PATIENT AND THE PULMONARY VEIN WAS CLAMPED WITH RELOADS COMPLETELY AND PUSHED THE GREEN SAFETY BUTTON. WHEN THE TOGGLE BUTTON WAS PUSHED DOWN TO FIRE THE RELOAD, IT WAS OBSERVED THAT THE BLADE GOT STOPPED AT THE BEGINNING PHASE OF THE FIRING PROCESS. THE TOGGLE BUTTON WAS TRIED TO PUSH DOWN ONE MORE TIME HOWEVER THE BLADE DID NOT MOVE AND PUSHED IT UP TO RETRACT THE BLADE. ANOTHER RELOAD WAS THEN USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS AN ERROR SOUND HEARD FROM THE HANDLE, AS ARTICULATING THE RELOAD IN LEFT SIDE. THE SCRUB NURSE DISCONNECTED THE NEW RELOAD FROM THE HANDLE AND REPLACED THE POWER HANDLE SET WITH ANOTHER ONE AND THE RELOAD WAS SUCCESSFULLY FIRED. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCESS OF INTESTINAL DISSECTION ON A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE A STAPLE LINE AND DID NOT CUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCESS OF INTESTINAL DISSECTION ON A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE A STAPLE LINE AND DID NOT CUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Injury | covidien medical products | 07-01-2020 | PREMIUM PLUS CEEA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLORECTAL SURGERY, THE STAPLE WAS ABLE TO CUT BUT DID NOT DEPLOYED STAPLES. SURGICAL INCISION WAS EXTENDED MORE THAN ONCE INCH AND THE PROCEDURE WAS EXTENDED FOR ONE HOUR AND 30 MINUTES DUE TO THE PRODUCT PROBLEM.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO LIGATION AND DIVISION OF PULMONARY VEIN DURING A VIDEO ASSISTED THO RACOSCOPIC LOBECTOMY PROCEDURE, IT WAS STATED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE FIRST RELOAD WAS INSERTED WITH THE POWER HANDLE THROUGH THE WINDOW PORT OF PATIENT AND THE PULMONARY VEIN WAS CLAMPED WITH RELOADS COMPLETELY AND PUSHED THE GREEN SAFETY BUTTON. WHEN THE TOGGLE BUTTON WAS PUSHED DOWN TO FIRE THE RELOAD, IT WAS OBSERVED THAT THE BLADE GOT STOPPED AT THE BEGINNING PHASE OF THE FIRING PROCESS. THE TOGGLE BUTTON WAS TRIED TO PUSH DOWN ONE MORE TIME HOWEVER THE BLADE DID NOT MOVE AND PUSHED IT UP TO RETRACT THE BLADE. ANOTHER RELOAD WAS THEN USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS AN ERROR SOUND HEARD FROM THE HANDLE, AS ARTICULATING THE RELOAD IN LEFT SIDE. THE SCRUB NURSE DISCONNECTED THE NEW RELOAD FROM THE HANDLE AND REPLACED THE POWER HANDLE SET WITH ANOTHER ONE AND THE RELOAD WAS SUCCESSFULLY FIRED. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO LIGATION AND DIVISION OF PULMONARY VEIN DURING A VIDEO ASSISTED THO RACOSCOPIC LOBECTOMY PROCEDURE, IT WAS STATED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE FIRST RELOAD WAS INSERTED WITH THE POWER HANDLE THROUGH THE WINDOW PORT OF PATIENT AND THE PULMONARY VEIN WAS CLAMPED WITH RELOADS COMPLETELY AND PUSHED THE GREEN SAFETY BUTTON. WHEN THE TOGGLE BUTTON WAS PUSHED DOWN TO FIRE THE RELOAD, IT WAS OBSERVED THAT THE BLADE GOT STOPPED AT THE BEGINNING PHASE OF THE FIRING PROCESS. THE TOGGLE BUTTON WAS TRIED TO PUSH DOWN ONE MORE TIME HOWEVER THE BLADE DID NOT MOVE AND PUSHED IT UP TO RETRACT THE BLADE. ANOTHER RELOAD WAS THEN USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS AN ERROR SOUND HEARD FROM THE HANDLE, AS ARTICULATING THE RELOAD IN LEFT SIDE. THE SCRUB NURSE DISCONNECTED THE NEW RELOAD FROM THE HANDLE AND REPLACED THE POWER HANDLE SET WITH ANOTHER ONE AND THE RELOAD WAS SUCCESSFULLY FIRED. THERE WAS NO PATIENT INJURY.
01-01-2020 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING AT DEODUNEM DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE HOWEVER THERE WAS A POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL PART. IT WAS STATED THAT THE STAPLES WAS NOT SEATED PROPERLY ON TISSUE, IT WAS MILKING OUT. THE DEVICE WAS ALSO LOCKED ON TISSUE, THEN THE SURGEON RETRACTED A KNOB TO UNLOAD THE RELOAD. THE STAPLE LINE WAS FIXED BY SUTURING AND ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
01-01-2020 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING AT DEODUNEM DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE HOWEVER THERE WAS A POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL PART. IT WAS STATED THAT THE STAPLES WAS NOT SEATED PROPERLY ON TISSUE, IT WAS MILKING OUT. THE DEVICE WAS ALSO LOCKED ON TISSUE, THEN THE SURGEON RETRACTED A KNOB TO UNLOAD THE RELOAD. THE STAPLE LINE WAS FIXED BY SUTURING AND ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
01-01-2020 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING AT DEODUNEM DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE HOWEVER THERE WAS A POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL PART. IT WAS STATED THAT THE STAPLES WAS NOT SEATED PROPERLY ON TISSUE, IT WAS MILKING OUT. THE DEVICE WAS ALSO LOCKED ON TISSUE, THEN THE SURGEON RETRACTED A KNOB TO UNLOAD THE RELOAD. THE STAPLE LINE WAS FIXED BY SUTURING AND ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
01-01-2020 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING AT DEODUNEM DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE HOWEVER THERE WAS A POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL PART. IT WAS STATED THAT THE STAPLES WAS NOT SEATED PROPERLY ON TISSUE, IT WAS MILKING OUT. THE DEVICE WAS ALSO LOCKED ON TISSUE, THEN THE SURGEON RETRACTED A KNOB TO UNLOAD THE RELOAD. THE STAPLE LINE WAS FIXED BY SUTURING AND ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Unintended Movement | A0512 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, AT THE TRANSECTION OF THE STOMACH BODY, THE OPERATOR TRIED TO FIRE THE STAPLES OF THE RELOADS ON THE TARGETED TISSUE, AFTER SUCCESSFULLY CLAMPING IT AND PUSHING THE GREEN SAFETY BUTTON, AS THEY PUSHED DOWN THE TOGGLE BUTTON TO FIRE THE RELOADS, THEY NOTICED THAT THE RELOADS TWITCHED THREE TIMES, THE RELOADS UNINTENTIONALLY MOVED IN FAST SWING WAY, IN THE LEFT AND RIGHT WAY FOR A SHORT TIME; AT THE BEGINNING PHASE OF FIRING PROCESS, AT THE COMPLETE PHASE OF FIRING PROCESS, AND AT THE BEGINNING PHASE OF THE RETRACTION PROCESS. THEY REPLACED THE ABNORMAL RELOADS WITH A NEW ONE. THERE WAS NO PATIENT INJURY.
12-12-2019 | Injury | ethicon endo-surgery | 07-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: IT WAS REPORTED THAT DURING REVIEW OF A JOURNAL ARTICLE, TITLE : A COMPARATIVE RETROSPECTIVE STUDY OF ROBOTIC SLEEVE GASTRECTOMY VS ROBOTIC GASTRIC BYPASS. AUTHOR : RADOMIR KOSANOVIC, REY JESÚS ROMERO, CHARAN DONKOR ARMANDO SARASUA, JORGE RAFAEL RABAZA, ANTHONY MICHAEL GONZALEZ. CITATION: INT J MED ROBOTICS COMPUT ASSIST SURG 2015; 11: 275¿283. / DOI: 10.1002/RCS.1621. THE PURPOSE OF THIS COMPARATIVE RETROSPECTIVE STUDY FROM A SINGLE INSTITUTION WAS TO MATCH ROBOTIC GASTRIC BYPASS (RGB) VS ROBOTIC SLEEVE GASTRECTOMY (RSG) AND COMPARE THEM TO THOSE OBSERVED WITH THE LAPAROSCOPIC PLATFORM. FROM SEP 2009 TO MAY 2012, A TOTAL OF 299 PATIENTS UNDERWENT EITHER ROBOTIC GASTRIC BYPASS (RGB) (N=165) OR ROBOTIC SLEEVE GASTRECTOMY (RSG) (N=134). IN BOTH GROUPS, A MIDLINE SUPRA-UMBILICAL OR UMBILICAL VERTICAL INCISION WAS PERFORMED AND PNEUMOPERITONEUM (12¿15MMHG) WAS CREATED USING A VERESS NEEDLE OR AN OPTICAL TROCAR (ENDOPATH XCEL ETHICON). DURING RGB, THE LAPAROSCOPICALLY STITCH-MARKED BILIOPANCREATIC LIMB WAS TRANSECTED WITH A LINEAR STAPLER (ECHELON ENDOPATH, ETHICON ENDOSURGERY) LOADED WITH A WHITE CARTRIDGE (45MM) THROUGH THE ASSISTANT PORT, LEAVING THE STITCH ON THE ALIMENTARY LIMB. THE INTERNAL PART OF THE JEJUNO¿JEJUNAL ANASTOMOSIS BETWEEN THE ALIMENTARY AND BILIOPANCREATIC LIMBS WAS CREATED USING A WHITE CARTRIDGE, THEN THE ENTEROTOMY WAS CLOSED WITH WHITE RELOADS AND REINFORCED WITH ONE LAYER OF RUNNING SUTURE (VYCRIL 2-0). THE LESSER OMENTUM WAS OPENED AND THE RETROGASTRIC SPACE ENTERED FOR TRANSECTION ACROSS THE STOMACH, 3¿4 CM DISTALLY FROM THE GASTROESOPHAGEAL JUNCTION, USING ONE HORIZONTAL APPLICATION WITH A 45MM STAPLER (BLUE). CLOSURE OF THE ENTEROTOMY IN PERFORMED WITH TWO LAYERS OF 2-0 VICRYL SUTURES IN A RUNNING FASHION. REINFORCEMENT TO THE STAPLE LINE WAS DONE EITHER BY BIOABSORBABLE MATERIAL (PERI-STRIPS DRY) OR FIBRIN GLUE APPLICATION (EVICEL FIBRIN SEALANT, ETHICON). FOR RSG, THE GREATER CURVATURE OF THE STOMACH WAS DIVIDED, IN MOST OF THE CASES WITH THE USE OF A HARMONIC SCALPEL (ETHICON ENDO-SURGERY) OR, MORE RECENTLY, WITH USE OF AN ENDOWRIST ONE VESSEL SEALER. RESECTION OF THE STOMACH WAS STARTED 4¿6 CM FROM THE PYLORUS, WITH SEQUENTIAL TRANSECTIONS USING A LINEAR STAPLER (ECHELON ENDOPATH, ETHICON ENDO-SURGERY) THROUGH THE ASSISTANT PORT. FOR THE FIRST TWO OR THREE APPLICATIONS IN THE ANTRAL AREA, GREEN CARTRIDGES (45MM) WERE UTILIZED, THEN BLUE (45MM) FOR THE THREE OR FOUR REMAINING TRANSECTIONS TO REACH THE ANGLE OF HIS. THE ANAESTHESIOLOGIST WAS POSITIONED OVER THE LEFT SIDE OF THE PATIENT¿S HEAD AND FROM THIS POSITION HE MANAGED A 38-FR BOUGIE (GASTRIC CALIBRATION TUBE, ETHICON) IN THE STOMACH. THE STAPLE LINE WAS REINFORCED WITH OVERSEWING PLUS FIBRIN GLUE (EVICEL FIBRIN SEALANT, ETHICON), OVERSEWN ALONE IN, AND BIOABSORBABLE-BUTTRESSED MATERIAL (PERI-STRIPS DRY). POSTOPERATIVE COMPLICATIONS UNDER RGB GROUP INCLUDED BLEEDING (N=2) ONE REQUIRED BLOOD TRANSFUSION, AND A HAEMATOMA REQUIRED PERCUTANEOUS DRAINAGE; AND INTRAABDOMINAL ABSCESS (N=1) TREATED WITH ANTIBIOTICS AND OBSERVATION; SMALL BOWEL PERFORATION (N=1) NEEDED REOPERATION; TREATED WITH BOWEL RESECTION AND SIDE-TO-SIDE ANASTOMOSIS AND DISCHARGED 12 DAYS LATER; INCARCERATED HERNIAS IN TROCAR SITES (N=1) AND INCISIONAL HERNIA (N=1). POSTOPERATIVE COMPLICATIONS UNDER RSG GROUP INCLUDED SUPERIOR MESENTERIC VEIN THROMBOSIS (N=1) TREATED WITH ANTICOAGULATION; BLEEDING FROM STAPLE LINE (N=1) REQUIRED REOPERATION; AND STOMACH TORSION (N=1) INITIALLY REQUIRED DILATATIONS, THEN EVENTUALLY DIAGNOSTIC LAPAROSCOPY, LYSIS OF ADHESIONS AND GASTROPEXY. RSGS HAD SHORTER SURGICAL TIMES AND A LOWER INCIDENCE OF LONGTERM COMPLICATIONS WHEN COMPARED WITH RGBS...
12-12-2019 | Injury | ethicon endo-surgery | 07-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Thrombosis | Event Description: IT WAS REPORTED THAT DURING REVIEW OF A JOURNAL ARTICLE, TITLE : A COMPARATIVE RETROSPECTIVE STUDY OF ROBOTIC SLEEVE GASTRECTOMY VS ROBOTIC GASTRIC BYPASS. AUTHOR : RADOMIR KOSANOVIC, REY JESÚS ROMERO, CHARAN DONKOR ARMANDO SARASUA, JORGE RAFAEL RABAZA, ANTHONY MICHAEL GONZALEZ. CITATION: INT J MED ROBOTICS COMPUT ASSIST SURG 2015; 11: 275¿283. / DOI: 10.1002/RCS.1621. THE PURPOSE OF THIS COMPARATIVE RETROSPECTIVE STUDY FROM A SINGLE INSTITUTION WAS TO MATCH ROBOTIC GASTRIC BYPASS (RGB) VS ROBOTIC SLEEVE GASTRECTOMY (RSG) AND COMPARE THEM TO THOSE OBSERVED WITH THE LAPAROSCOPIC PLATFORM. FROM SEP 2009 TO MAY 2012, A TOTAL OF 299 PATIENTS UNDERWENT EITHER ROBOTIC GASTRIC BYPASS (RGB) (N=165) OR ROBOTIC SLEEVE GASTRECTOMY (RSG) (N=134). IN BOTH GROUPS, A MIDLINE SUPRA-UMBILICAL OR UMBILICAL VERTICAL INCISION WAS PERFORMED AND PNEUMOPERITONEUM (12¿15MMHG) WAS CREATED USING A VERESS NEEDLE OR AN OPTICAL TROCAR (ENDOPATH XCEL ETHICON). DURING RGB, THE LAPAROSCOPICALLY STITCH-MARKED BILIOPANCREATIC LIMB WAS TRANSECTED WITH A LINEAR STAPLER (ECHELON ENDOPATH, ETHICON ENDOSURGERY) LOADED WITH A WHITE CARTRIDGE (45MM) THROUGH THE ASSISTANT PORT, LEAVING THE STITCH ON THE ALIMENTARY LIMB. THE INTERNAL PART OF THE JEJUNO¿JEJUNAL ANASTOMOSIS BETWEEN THE ALIMENTARY AND BILIOPANCREATIC LIMBS WAS CREATED USING A WHITE CARTRIDGE, THEN THE ENTEROTOMY WAS CLOSED WITH WHITE RELOADS AND REINFORCED WITH ONE LAYER OF RUNNING SUTURE (VYCRIL 2-0). THE LESSER OMENTUM WAS OPENED AND THE RETROGASTRIC SPACE ENTERED FOR TRANSECTION ACROSS THE STOMACH, 3¿4 CM DISTALLY FROM THE GASTROESOPHAGEAL JUNCTION, USING ONE HORIZONTAL APPLICATION WITH A 45MM STAPLER (BLUE). CLOSURE OF THE ENTEROTOMY IN PERFORMED WITH TWO LAYERS OF 2-0 VICRYL SUTURES IN A RUNNING FASHION. REINFORCEMENT TO THE STAPLE LINE WAS DONE EITHER BY BIOABSORBABLE MATERIAL (PERI-STRIPS DRY) OR FIBRIN GLUE APPLICATION (EVICEL FIBRIN SEALANT, ETHICON). FOR RSG, THE GREATER CURVATURE OF THE STOMACH WAS DIVIDED, IN MOST OF THE CASES WITH THE USE OF A HARMONIC SCALPEL (ETHICON ENDO-SURGERY) OR, MORE RECENTLY, WITH USE OF AN ENDOWRIST ONE VESSEL SEALER. RESECTION OF THE STOMACH WAS STARTED 4¿6 CM FROM THE PYLORUS, WITH SEQUENTIAL TRANSECTIONS USING A LINEAR STAPLER (ECHELON ENDOPATH, ETHICON ENDO-SURGERY) THROUGH THE ASSISTANT PORT. FOR THE FIRST TWO OR THREE APPLICATIONS IN THE ANTRAL AREA, GREEN CARTRIDGES (45MM) WERE UTILIZED, THEN BLUE (45MM) FOR THE THREE OR FOUR REMAINING TRANSECTIONS TO REACH THE ANGLE OF HIS. THE ANAESTHESIOLOGIST WAS POSITIONED OVER THE LEFT SIDE OF THE PATIENT¿S HEAD AND FROM THIS POSITION HE MANAGED A 38-FR BOUGIE (GASTRIC CALIBRATION TUBE, ETHICON) IN THE STOMACH. THE STAPLE LINE WAS REINFORCED WITH OVERSEWING PLUS FIBRIN GLUE (EVICEL FIBRIN SEALANT, ETHICON), OVERSEWN ALONE IN, AND BIOABSORBABLE-BUTTRESSED MATERIAL (PERI-STRIPS DRY). POSTOPERATIVE COMPLICATIONS UNDER RGB GROUP INCLUDED BLEEDING (N=2) ONE REQUIRED BLOOD TRANSFUSION, AND A HAEMATOMA REQUIRED PERCUTANEOUS DRAINAGE; AND INTRAABDOMINAL ABSCESS (N=1) TREATED WITH ANTIBIOTICS AND OBSERVATION; SMALL BOWEL PERFORATION (N=1) NEEDED REOPERATION; TREATED WITH BOWEL RESECTION AND SIDE-TO-SIDE ANASTOMOSIS AND DISCHARGED 12 DAYS LATER; INCARCERATED HERNIAS IN TROCAR SITES (N=1) AND INCISIONAL HERNIA (N=1). POSTOPERATIVE COMPLICATIONS UNDER RSG GROUP INCLUDED SUPERIOR MESENTERIC VEIN THROMBOSIS (N=1) TREATED WITH ANTICOAGULATION; BLEEDING FROM STAPLE LINE (N=1) REQUIRED REOPERATION; AND STOMACH TORSION (N=1) INITIALLY REQUIRED DILATATIONS, THEN EVENTUALLY DIAGNOSTIC LAPAROSCOPY, LYSIS OF ADHESIONS AND GASTROPEXY. RSGS HAD SHORTER SURGICAL TIMES AND A LOWER INCIDENCE OF LONGTERM COMPLICATIONS WHEN COMPARED WITH RGBS...
12-12-2019 | Injury | ethicon endo-surgery | 07-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING REVIEW OF A JOURNAL ARTICLE, TITLE : A COMPARATIVE RETROSPECTIVE STUDY OF ROBOTIC SLEEVE GASTRECTOMY VS ROBOTIC GASTRIC BYPASS. AUTHOR : RADOMIR KOSANOVIC, REY JESÚS ROMERO, CHARAN DONKOR ARMANDO SARASUA, JORGE RAFAEL RABAZA, ANTHONY MICHAEL GONZALEZ. CITATION: INT J MED ROBOTICS COMPUT ASSIST SURG 2015; 11: 275¿283. / DOI: 10.1002/RCS.1621. THE PURPOSE OF THIS COMPARATIVE RETROSPECTIVE STUDY FROM A SINGLE INSTITUTION WAS TO MATCH ROBOTIC GASTRIC BYPASS (RGB) VS ROBOTIC SLEEVE GASTRECTOMY (RSG) AND COMPARE THEM TO THOSE OBSERVED WITH THE LAPAROSCOPIC PLATFORM. FROM SEP 2009 TO MAY 2012, A TOTAL OF 299 PATIENTS UNDERWENT EITHER ROBOTIC GASTRIC BYPASS (RGB) (N=165) OR ROBOTIC SLEEVE GASTRECTOMY (RSG) (N=134). IN BOTH GROUPS, A MIDLINE SUPRA-UMBILICAL OR UMBILICAL VERTICAL INCISION WAS PERFORMED AND PNEUMOPERITONEUM (12¿15MMHG) WAS CREATED USING A VERESS NEEDLE OR AN OPTICAL TROCAR (ENDOPATH XCEL ETHICON). DURING RGB, THE LAPAROSCOPICALLY STITCH-MARKED BILIOPANCREATIC LIMB WAS TRANSECTED WITH A LINEAR STAPLER (ECHELON ENDOPATH, ETHICON ENDOSURGERY) LOADED WITH A WHITE CARTRIDGE (45MM) THROUGH THE ASSISTANT PORT, LEAVING THE STITCH ON THE ALIMENTARY LIMB. THE INTERNAL PART OF THE JEJUNO¿JEJUNAL ANASTOMOSIS BETWEEN THE ALIMENTARY AND BILIOPANCREATIC LIMBS WAS CREATED USING A WHITE CARTRIDGE, THEN THE ENTEROTOMY WAS CLOSED WITH WHITE RELOADS AND REINFORCED WITH ONE LAYER OF RUNNING SUTURE (VYCRIL 2-0). THE LESSER OMENTUM WAS OPENED AND THE RETROGASTRIC SPACE ENTERED FOR TRANSECTION ACROSS THE STOMACH, 3¿4 CM DISTALLY FROM THE GASTROESOPHAGEAL JUNCTION, USING ONE HORIZONTAL APPLICATION WITH A 45MM STAPLER (BLUE). CLOSURE OF THE ENTEROTOMY IN PERFORMED WITH TWO LAYERS OF 2-0 VICRYL SUTURES IN A RUNNING FASHION. REINFORCEMENT TO THE STAPLE LINE WAS DONE EITHER BY BIOABSORBABLE MATERIAL (PERI-STRIPS DRY) OR FIBRIN GLUE APPLICATION (EVICEL FIBRIN SEALANT, ETHICON). FOR RSG, THE GREATER CURVATURE OF THE STOMACH WAS DIVIDED, IN MOST OF THE CASES WITH THE USE OF A HARMONIC SCALPEL (ETHICON ENDO-SURGERY) OR, MORE RECENTLY, WITH USE OF AN ENDOWRIST ONE VESSEL SEALER. RESECTION OF THE STOMACH WAS STARTED 4¿6 CM FROM THE PYLORUS, WITH SEQUENTIAL TRANSECTIONS USING A LINEAR STAPLER (ECHELON ENDOPATH, ETHICON ENDO-SURGERY) THROUGH THE ASSISTANT PORT. FOR THE FIRST TWO OR THREE APPLICATIONS IN THE ANTRAL AREA, GREEN CARTRIDGES (45MM) WERE UTILIZED, THEN BLUE (45MM) FOR THE THREE OR FOUR REMAINING TRANSECTIONS TO REACH THE ANGLE OF HIS. THE ANAESTHESIOLOGIST WAS POSITIONED OVER THE LEFT SIDE OF THE PATIENT¿S HEAD AND FROM THIS POSITION HE MANAGED A 38-FR BOUGIE (GASTRIC CALIBRATION TUBE, ETHICON) IN THE STOMACH. THE STAPLE LINE WAS REINFORCED WITH OVERSEWING PLUS FIBRIN GLUE (EVICEL FIBRIN SEALANT, ETHICON), OVERSEWN ALONE IN, AND BIOABSORBABLE-BUTTRESSED MATERIAL (PERI-STRIPS DRY). POSTOPERATIVE COMPLICATIONS UNDER RGB GROUP INCLUDED BLEEDING (N=2) ONE REQUIRED BLOOD TRANSFUSION, AND A HAEMATOMA REQUIRED PERCUTANEOUS DRAINAGE; AND INTRAABDOMINAL ABSCESS (N=1) TREATED WITH ANTIBIOTICS AND OBSERVATION; SMALL BOWEL PERFORATION (N=1) NEEDED REOPERATION; TREATED WITH BOWEL RESECTION AND SIDE-TO-SIDE ANASTOMOSIS AND DISCHARGED 12 DAYS LATER; INCARCERATED HERNIAS IN TROCAR SITES (N=1) AND INCISIONAL HERNIA (N=1). POSTOPERATIVE COMPLICATIONS UNDER RSG GROUP INCLUDED SUPERIOR MESENTERIC VEIN THROMBOSIS (N=1) TREATED WITH ANTICOAGULATION; BLEEDING FROM STAPLE LINE (N=1) REQUIRED REOPERATION; AND STOMACH TORSION (N=1) INITIALLY REQUIRED DILATATIONS, THEN EVENTUALLY DIAGNOSTIC LAPAROSCOPY, LYSIS OF ADHESIONS AND GASTROPEXY. RSGS HAD SHORTER SURGICAL TIMES AND A LOWER INCIDENCE OF LONGTERM COMPLICATIONS WHEN COMPARED WITH RGBS...
12-12-2019 | Injury | ethicon endo-surgery | 07-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING REVIEW OF A JOURNAL ARTICLE, TITLE : A COMPARATIVE RETROSPECTIVE STUDY OF ROBOTIC SLEEVE GASTRECTOMY VS ROBOTIC GASTRIC BYPASS. AUTHOR : RADOMIR KOSANOVIC, REY JESÚS ROMERO, CHARAN DONKOR ARMANDO SARASUA, JORGE RAFAEL RABAZA, ANTHONY MICHAEL GONZALEZ. CITATION: INT J MED ROBOTICS COMPUT ASSIST SURG 2015; 11: 275¿283. / DOI: 10.1002/RCS.1621. THE PURPOSE OF THIS COMPARATIVE RETROSPECTIVE STUDY FROM A SINGLE INSTITUTION WAS TO MATCH ROBOTIC GASTRIC BYPASS (RGB) VS ROBOTIC SLEEVE GASTRECTOMY (RSG) AND COMPARE THEM TO THOSE OBSERVED WITH THE LAPAROSCOPIC PLATFORM. FROM SEP 2009 TO MAY 2012, A TOTAL OF 299 PATIENTS UNDERWENT EITHER ROBOTIC GASTRIC BYPASS (RGB) (N=165) OR ROBOTIC SLEEVE GASTRECTOMY (RSG) (N=134). IN BOTH GROUPS, A MIDLINE SUPRA-UMBILICAL OR UMBILICAL VERTICAL INCISION WAS PERFORMED AND PNEUMOPERITONEUM (12¿15MMHG) WAS CREATED USING A VERESS NEEDLE OR AN OPTICAL TROCAR (ENDOPATH XCEL ETHICON). DURING RGB, THE LAPAROSCOPICALLY STITCH-MARKED BILIOPANCREATIC LIMB WAS TRANSECTED WITH A LINEAR STAPLER (ECHELON ENDOPATH, ETHICON ENDOSURGERY) LOADED WITH A WHITE CARTRIDGE (45MM) THROUGH THE ASSISTANT PORT, LEAVING THE STITCH ON THE ALIMENTARY LIMB. THE INTERNAL PART OF THE JEJUNO¿JEJUNAL ANASTOMOSIS BETWEEN THE ALIMENTARY AND BILIOPANCREATIC LIMBS WAS CREATED USING A WHITE CARTRIDGE, THEN THE ENTEROTOMY WAS CLOSED WITH WHITE RELOADS AND REINFORCED WITH ONE LAYER OF RUNNING SUTURE (VYCRIL 2-0). THE LESSER OMENTUM WAS OPENED AND THE RETROGASTRIC SPACE ENTERED FOR TRANSECTION ACROSS THE STOMACH, 3¿4 CM DISTALLY FROM THE GASTROESOPHAGEAL JUNCTION, USING ONE HORIZONTAL APPLICATION WITH A 45MM STAPLER (BLUE). CLOSURE OF THE ENTEROTOMY IN PERFORMED WITH TWO LAYERS OF 2-0 VICRYL SUTURES IN A RUNNING FASHION. REINFORCEMENT TO THE STAPLE LINE WAS DONE EITHER BY BIOABSORBABLE MATERIAL (PERI-STRIPS DRY) OR FIBRIN GLUE APPLICATION (EVICEL FIBRIN SEALANT, ETHICON). FOR RSG, THE GREATER CURVATURE OF THE STOMACH WAS DIVIDED, IN MOST OF THE CASES WITH THE USE OF A HARMONIC SCALPEL (ETHICON ENDO-SURGERY) OR, MORE RECENTLY, WITH USE OF AN ENDOWRIST ONE VESSEL SEALER. RESECTION OF THE STOMACH WAS STARTED 4¿6 CM FROM THE PYLORUS, WITH SEQUENTIAL TRANSECTIONS USING A LINEAR STAPLER (ECHELON ENDOPATH, ETHICON ENDO-SURGERY) THROUGH THE ASSISTANT PORT. FOR THE FIRST TWO OR THREE APPLICATIONS IN THE ANTRAL AREA, GREEN CARTRIDGES (45MM) WERE UTILIZED, THEN BLUE (45MM) FOR THE THREE OR FOUR REMAINING TRANSECTIONS TO REACH THE ANGLE OF HIS. THE ANAESTHESIOLOGIST WAS POSITIONED OVER THE LEFT SIDE OF THE PATIENT¿S HEAD AND FROM THIS POSITION HE MANAGED A 38-FR BOUGIE (GASTRIC CALIBRATION TUBE, ETHICON) IN THE STOMACH. THE STAPLE LINE WAS REINFORCED WITH OVERSEWING PLUS FIBRIN GLUE (EVICEL FIBRIN SEALANT, ETHICON), OVERSEWN ALONE IN, AND BIOABSORBABLE-BUTTRESSED MATERIAL (PERI-STRIPS DRY). POSTOPERATIVE COMPLICATIONS UNDER RGB GROUP INCLUDED BLEEDING (N=2) ONE REQUIRED BLOOD TRANSFUSION, AND A HAEMATOMA REQUIRED PERCUTANEOUS DRAINAGE; AND INTRAABDOMINAL ABSCESS (N=1) TREATED WITH ANTIBIOTICS AND OBSERVATION; SMALL BOWEL PERFORATION (N=1) NEEDED REOPERATION; TREATED WITH BOWEL RESECTION AND SIDE-TO-SIDE ANASTOMOSIS AND DISCHARGED 12 DAYS LATER; INCARCERATED HERNIAS IN TROCAR SITES (N=1) AND INCISIONAL HERNIA (N=1). POSTOPERATIVE COMPLICATIONS UNDER RSG GROUP INCLUDED SUPERIOR MESENTERIC VEIN THROMBOSIS (N=1) TREATED WITH ANTICOAGULATION; BLEEDING FROM STAPLE LINE (N=1) REQUIRED REOPERATION; AND STOMACH TORSION (N=1) INITIALLY REQUIRED DILATATIONS, THEN EVENTUALLY DIAGNOSTIC LAPAROSCOPY, LYSIS OF ADHESIONS AND GASTROPEXY. RSGS HAD SHORTER SURGICAL TIMES AND A LOWER INCIDENCE OF LONGTERM COMPLICATIONS WHEN COMPARED WITH RGBS...
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, AT THE TRANSECTION OF PULMONARY BRONCHUS, THE OPERATOR TRIED TO FIRE THE STAPLES ON THE TARGETED TISSUE WITH HANDLE. AFTER CLAMPING IT AND PUSHING THE SAFETY BUTTON SUCCESSFULLY, THE HANDLE COULD NOT BE SQUEEZED AND THE BLADE OF THE RELOAD DID NOT MOVE AT ALL. THE ABNORMAL RELOAD WERE REPLACED WITH A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC MEDIASTINAL TUMOR RESECTION PROCEDURE, THE DEVICE DID NOT FIRE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | us surgical puerto rico | 07-01-2020 | UNKNOWN EGIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC MEDIASTINAL TUMOR RESECTION PROCEDURE, THE DEVICE DID NOT FIRE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC RADICAL RESECTION OF COLORECTAL CANCER PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, HENCE THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE RELOAD ONLY FIRED FOR THE FIRST PART AND AUTOMATICALLY RECALLED THE BLADE SINCE THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. A DIFFERENT RELOAD WAS USED AND THE SAME ISSUE OCCURRED. IT WAS ALSO STATED THAT FOR THE BOTH RELOADS THERE WAS A POOR STAPLE FORMATION. A DIFFERENT BRAND WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE RELOAD ONLY FIRED FOR THE FIRST PART AND AUTOMATICALLY RECALLED THE BLADE SINCE THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. A DIFFERENT RELOAD WAS USED AND THE SAME ISSUE OCCURRED. IT WAS ALSO STATED THAT FOR THE BOTH RELOADS THERE WAS A POOR STAPLE FORMATION. A DIFFERENT BRAND WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, WHILE WORKING ON THE LUNG PARENCHYMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE THEREFORE, THE DEVICE WAS NOT ABLE TO FIRE. SURGEON USED ANOTHER RELOAD TO RESOLVE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE RELOAD ONLY FIRED FOR THE FIRST PART AND AUTOMATICALLY RECALLED THE BLADE SINCE THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. A DIFFERENT RELOAD WAS USED AND THE SAME ISSUE OCCURRED. IT WAS ALSO STATED THAT FOR THE BOTH RELOADS THERE WAS A POOR STAPLE FORMATION. A DIFFERENT BRAND WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE RELOAD ONLY FIRED FOR THE FIRST PART AND AUTOMATICALLY RECALLED THE BLADE SINCE THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. A DIFFERENT RELOAD WAS USED AND THE SAME ISSUE OCCURRED. IT WAS ALSO STATED THAT FOR THE BOTH RELOADS THERE WAS A POOR STAPLE FORMATION. A DIFFERENT BRAND WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING DISSECTION OF LUNG PARENCHYMA ON A THORACOSCOPIC LOBECTOMY, THE STAPLER WAS ABLE TO FIRE BUT THE STAPLES DID NOT FORM B SHAPE. A NEW HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | Sepsis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENTS DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENTS DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Pulmonary Embolism | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Abdominal Pain | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Abscess | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Adhesion(s) | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Anemia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Bacterial Infection | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Diarrhea | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Purulent Discharge | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Dyspnea | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Erythema | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Fatigue | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Fever | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Fistula | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | High Blood Pressure/ Hypertension | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Low Blood Pressure/ Hypotension | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Infection | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Inflammation | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Nausea | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Necrosis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Internal Organ Perforation | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Pain | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Pleural Effusion | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Pneumonia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Renal Failure | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Sepsis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Septic Shock | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Seroma | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tachycardia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Vomiting | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Chills | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Hernia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Hypernatremia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Peritonitis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Anxiety | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Discomfort | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Depression | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Fungal Infection | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Respiratory Acidosis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Respiratory Failure | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Fluid Discharge | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Fibrosis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Thrombosis/Thrombus | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Cramp(s) /Muscle Spasm(s) | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Appropriate Term / Code Not Available | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Pulmonary Embolism | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Adhesion(s) | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Anemia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Bacterial Infection | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Diarrhea | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Purulent Discharge | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Dyspnea | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Erythema | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Fatigue | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | High Blood Pressure/ Hypertension | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Low Blood Pressure/ Hypotension | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Inflammation | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Internal Organ Perforation | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Pneumonia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Renal Failure | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Septic Shock | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Seroma | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tachycardia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Vomiting | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Chills | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hernia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hypernatremia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Peritonitis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Anxiety | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Discomfort | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Depression | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Fungal Infection | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Respiratory Acidosis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Respiratory Failure | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Fibrosis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Thrombosis/Thrombus | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Cramp(s) /Muscle Spasm(s) | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Pulmonary Embolism | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Abdominal Pain | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Abscess | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Adhesion(s) | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Anemia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Bacterial Infection | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Diarrhea | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Purulent Discharge | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Dyspnea | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Erythema | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Fatigue | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Fever | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Fistula | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | High Blood Pressure/ Hypertension | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Low Blood Pressure/ Hypotension | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Unspecified Infection | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Inflammation | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Nausea | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Necrosis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Internal Organ Perforation | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Pain | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Pleural Effusion | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Pneumonia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Renal Failure | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Sepsis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Septic Shock | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Seroma | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Tachycardia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Vomiting | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Chills | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Hernia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Hypernatremia | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Peritonitis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Anxiety | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Discomfort | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Depression | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Fungal Infection | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Respiratory Acidosis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Respiratory Failure | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Fluid Discharge | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Fibrosis | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Thrombosis/Thrombus | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Cramp(s) /Muscle Spasm(s) | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
21-12-2017 | Injury | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Patient Device Interaction Problem | A01 | Appropriate Term / Code Not Available | Event Description: LEGAL: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2017, THE PATIENT UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY. DURING THE PROCEDURE THE DEVICES WERE USED. THE PATIENT WAS ADMITTED TO THE ER AFTER EXPERIENCING SEVERE ABDOMINAL PAIN. THE PATIENT'S DOCTOR DISCOVERED THE PATIENT WAS SEPTIC AND IT WAS CAUSED BY A FAILED STAPLE LINE DUE TO INCOMPLETE STAPLE FORMATION. IT WAS DISCOVERED A STAPLE HAD "POPPED" BECAUSE IT HAD BEEN IMPROPERLY FORMED BY THE STAPLER. THE PATIENT WAS HOSPITALIZED FOR OVER A MONTH AND EXPERIENCED FURTHER MEDICAL COMPLICATIONS AFTERWARDS.
19-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | EEA | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, THE EXCESS TISSUE COULD NOT BE COMPLETELY CUT OFF BY THE STAPLER, WHICH MADE IT DIFFICULT TO REMOVE THE STAPLER AFTER ANASTOMOSIS. THE TISSUE WAS CUT WITH SCISSORS. THE EVENT RESULTED TO AN INCREASE IN BLEEDING, AFFECTED THE QUALITY OF ANASTOMOSIS AND RISK OF ANASTOMOTIC FISTULA AFTER PROCEDURE. NO PATIENT COMPLICATION WAS REPORTED AS THE RESULT OF THE EVENT.
12-12-2019 | Injury | ethicon | 07-01-2020 | SECURESTRAP UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: TITLE: SEROMA FOLLOWING TRANSABDOMINAL PREPERITONEAL PATCH PLASTY (TAPP): INCIDENCE, RISK FACTORS, AND PREVENTIVE MEASURES THE PURPOSE OF THIS PROSPECTIVE STUDY WAS TO DISCUSS INFLUENCING FACTORS FOR SEROMA FORMATION IN MALE PATIENTS AFTER TAPP REPAIR. FROM SEP 01, 2009 TO SEP 01, 2015, 20,000 MALE PATIENTS WITH PRIMARY UNILATERAL INGUINAL HERNIA UNDERWENT TAPP REPAIR. THE PATIENTS WERE DIVIDED INTO THREE GROUPS: WITHOUT MESH FIXATION (N=8799, MEDIAN AGE 55.0 ± 15.6 YEARS, MEAN BMI 25.9 ± 3.3 KG/M2); USING TACKS FOR FIXATION (N=6387, MEDIAN AGE 58.8 ± 14.7 YEARS, MEAN BMI 26.0 ± 3.4 KG/M2); AND PATIENTS WITH GLUE FIXATION (N=4818, MEDIAN AGE 56.4 ± 15.0 YEARS, MEAN BMI 25.8 ± 3.4 KG/M2). DURING THE PROCEDURE, THE PATIENTS USED OTHER MESHES, PROLENE MESH (N=465) OR ULTRAPRO MESH (N=5021). FOR FIXATION, THOSE WHO UTILIZED TACKS INCLUDED OTHER TACKS, EMS STAPLER (N=1299) OR SECURESTRAP (N=328). AND THOSE WHO USED GLUE INCLUDED OTHER GLUES OR EVICEL (N=1607). POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING (N=?), SEROMA (N=?), INFECTION (N=?), AND WOUND HEALING DISORDERS (N=?). IN SUMMARY, IT HAS BEEN DEMONSTRATED THAT THE SEROMA RATE IN MALE PATIENTS WITH TAPP REPAIR OF PRIMARY UNILATERAL INGUINAL HERNIA IS NEGATIVELY INFLUENCED BY MESH FIXATION WITH TACKS OR GLUE. NON-MESH FIXATION WAS ASSOCIATED WITH THE LOWEST SEROMA RATE.
12-12-2019 | Injury | ethicon | 07-01-2020 | SECURESTRAP UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Seroma | Event Description: TITLE: SEROMA FOLLOWING TRANSABDOMINAL PREPERITONEAL PATCH PLASTY (TAPP): INCIDENCE, RISK FACTORS, AND PREVENTIVE MEASURES THE PURPOSE OF THIS PROSPECTIVE STUDY WAS TO DISCUSS INFLUENCING FACTORS FOR SEROMA FORMATION IN MALE PATIENTS AFTER TAPP REPAIR. FROM SEP 01, 2009 TO SEP 01, 2015, 20,000 MALE PATIENTS WITH PRIMARY UNILATERAL INGUINAL HERNIA UNDERWENT TAPP REPAIR. THE PATIENTS WERE DIVIDED INTO THREE GROUPS: WITHOUT MESH FIXATION (N=8799, MEDIAN AGE 55.0 ± 15.6 YEARS, MEAN BMI 25.9 ± 3.3 KG/M2); USING TACKS FOR FIXATION (N=6387, MEDIAN AGE 58.8 ± 14.7 YEARS, MEAN BMI 26.0 ± 3.4 KG/M2); AND PATIENTS WITH GLUE FIXATION (N=4818, MEDIAN AGE 56.4 ± 15.0 YEARS, MEAN BMI 25.8 ± 3.4 KG/M2). DURING THE PROCEDURE, THE PATIENTS USED OTHER MESHES, PROLENE MESH (N=465) OR ULTRAPRO MESH (N=5021). FOR FIXATION, THOSE WHO UTILIZED TACKS INCLUDED OTHER TACKS, EMS STAPLER (N=1299) OR SECURESTRAP (N=328). AND THOSE WHO USED GLUE INCLUDED OTHER GLUES OR EVICEL (N=1607). POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING (N=?), SEROMA (N=?), INFECTION (N=?), AND WOUND HEALING DISORDERS (N=?). IN SUMMARY, IT HAS BEEN DEMONSTRATED THAT THE SEROMA RATE IN MALE PATIENTS WITH TAPP REPAIR OF PRIMARY UNILATERAL INGUINAL HERNIA IS NEGATIVELY INFLUENCED BY MESH FIXATION WITH TACKS OR GLUE. NON-MESH FIXATION WAS ASSOCIATED WITH THE LOWEST SEROMA RATE.
12-12-2019 | Injury | ethicon | 07-01-2020 | SECURESTRAP UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Impaired Healing | Event Description: TITLE: SEROMA FOLLOWING TRANSABDOMINAL PREPERITONEAL PATCH PLASTY (TAPP): INCIDENCE, RISK FACTORS, AND PREVENTIVE MEASURES THE PURPOSE OF THIS PROSPECTIVE STUDY WAS TO DISCUSS INFLUENCING FACTORS FOR SEROMA FORMATION IN MALE PATIENTS AFTER TAPP REPAIR. FROM SEP 01, 2009 TO SEP 01, 2015, 20,000 MALE PATIENTS WITH PRIMARY UNILATERAL INGUINAL HERNIA UNDERWENT TAPP REPAIR. THE PATIENTS WERE DIVIDED INTO THREE GROUPS: WITHOUT MESH FIXATION (N=8799, MEDIAN AGE 55.0 ± 15.6 YEARS, MEAN BMI 25.9 ± 3.3 KG/M2); USING TACKS FOR FIXATION (N=6387, MEDIAN AGE 58.8 ± 14.7 YEARS, MEAN BMI 26.0 ± 3.4 KG/M2); AND PATIENTS WITH GLUE FIXATION (N=4818, MEDIAN AGE 56.4 ± 15.0 YEARS, MEAN BMI 25.8 ± 3.4 KG/M2). DURING THE PROCEDURE, THE PATIENTS USED OTHER MESHES, PROLENE MESH (N=465) OR ULTRAPRO MESH (N=5021). FOR FIXATION, THOSE WHO UTILIZED TACKS INCLUDED OTHER TACKS, EMS STAPLER (N=1299) OR SECURESTRAP (N=328). AND THOSE WHO USED GLUE INCLUDED OTHER GLUES OR EVICEL (N=1607). POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING (N=?), SEROMA (N=?), INFECTION (N=?), AND WOUND HEALING DISORDERS (N=?). IN SUMMARY, IT HAS BEEN DEMONSTRATED THAT THE SEROMA RATE IN MALE PATIENTS WITH TAPP REPAIR OF PRIMARY UNILATERAL INGUINAL HERNIA IS NEGATIVELY INFLUENCED BY MESH FIXATION WITH TACKS OR GLUE. NON-MESH FIXATION WAS ASSOCIATED WITH THE LOWEST SEROMA RATE.
12-12-2019 | Injury | ethicon | 07-01-2020 | SECURESTRAP UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: TITLE: SEROMA FOLLOWING TRANSABDOMINAL PREPERITONEAL PATCH PLASTY (TAPP): INCIDENCE, RISK FACTORS, AND PREVENTIVE MEASURES THE PURPOSE OF THIS PROSPECTIVE STUDY WAS TO DISCUSS INFLUENCING FACTORS FOR SEROMA FORMATION IN MALE PATIENTS AFTER TAPP REPAIR. FROM SEP 01, 2009 TO SEP 01, 2015, 20,000 MALE PATIENTS WITH PRIMARY UNILATERAL INGUINAL HERNIA UNDERWENT TAPP REPAIR. THE PATIENTS WERE DIVIDED INTO THREE GROUPS: WITHOUT MESH FIXATION (N=8799, MEDIAN AGE 55.0 ± 15.6 YEARS, MEAN BMI 25.9 ± 3.3 KG/M2); USING TACKS FOR FIXATION (N=6387, MEDIAN AGE 58.8 ± 14.7 YEARS, MEAN BMI 26.0 ± 3.4 KG/M2); AND PATIENTS WITH GLUE FIXATION (N=4818, MEDIAN AGE 56.4 ± 15.0 YEARS, MEAN BMI 25.8 ± 3.4 KG/M2). DURING THE PROCEDURE, THE PATIENTS USED OTHER MESHES, PROLENE MESH (N=465) OR ULTRAPRO MESH (N=5021). FOR FIXATION, THOSE WHO UTILIZED TACKS INCLUDED OTHER TACKS, EMS STAPLER (N=1299) OR SECURESTRAP (N=328). AND THOSE WHO USED GLUE INCLUDED OTHER GLUES OR EVICEL (N=1607). POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING (N=?), SEROMA (N=?), INFECTION (N=?), AND WOUND HEALING DISORDERS (N=?). IN SUMMARY, IT HAS BEEN DEMONSTRATED THAT THE SEROMA RATE IN MALE PATIENTS WITH TAPP REPAIR OF PRIMARY UNILATERAL INGUINAL HERNIA IS NEGATIVELY INFLUENCED BY MESH FIXATION WITH TACKS OR GLUE. NON-MESH FIXATION WAS ASSOCIATED WITH THE LOWEST SEROMA RATE.
26-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Separation Failure | A150301 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO LOBECTOMY PROCEDURE, WHEN USING THE REPORTED REINFORCEMENT CARTRIDGE ON INTERLOBAR TISSUE, THE SHEET COULD NOT BE CUT PROPERLY AND IT GOT STACKED AT THE TIP. IT WAS ALSO REPORTED THAT THERE WAS A LITTLE PART OF FIRING THAT IS INCOMPLETE SINCE THE FIRST ROW OF STAPLES AND WHERE THE KNIFE DID NOT CUT. THERE WAS NO PATIENT INJURY.
26-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO LOBECTOMY PROCEDURE, WHEN USING THE REPORTED REINFORCEMENT CARTRIDGE ON INTERLOBAR TISSUE, THE SHEET COULD NOT BE CUT PROPERLY AND IT GOT STACKED AT THE TIP. IT WAS ALSO REPORTED THAT THERE WAS A LITTLE PART OF FIRING THAT IS INCOMPLETE SINCE THE FIRST ROW OF STAPLES AND WHERE THE KNIFE DID NOT CUT. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF THE INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN THE DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF THE INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN THE DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Positioning Problem | A1502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF THE INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN THE DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF THE INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN THE DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF THE INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN THE DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Positioning Problem | A1502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF THE INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN THE DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF THE INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN THE DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 07-01-2020 | ENDO GIA | Positioning Problem | A1502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS SURGERY, THE TWO HANDLES' HEAD OF THE INSTRUMENT SHAFT BROKE OFF. WITH THE FIRST RELOAD, THE LOADING UNIT CONNECTOR AND SHAFT BROKE OFF. IT WAS NOT EASY TO APPROACH WHEN THE DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. THE NEXT THREE RELOADS WERE ALSO NOT EASY TO APPROACH WHEN DOCTOR WAS TRYING TO CUT THE ESOPHAGUS TISSUE. IT WAS ALSO MENTIONED THAT THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT UPPER LOBE LOBECTOMY, THE STAPLER STOPPED WORKING AFTER FOUR USES. IT WAS STATED THAT THE HANDLE WOULD NO LONGER CLOSE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | UNKNOWN EGIA SULU | Difficult to Open or Close | A051104 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT UPPER LOBE LOBECTOMY, THE STAPLER STOPPED WORKING AFTER FOUR USES. IT WAS STATED THAT THE HANDLE WOULD NO LONGER CLOSE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-05-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLER WOULD NOT FIRE PROPERLY.
02-05-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLER WOULD NOT FIRE PROPERLY.
01-12-2019 | Malfunction | us surgical puerto rico | 07-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, THE HANDLE WOULD NOT FIRE AFTER THE FIRST FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-12-2019 | Malfunction | covidien lp llc north haven | 07-01-2020 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, THE HANDLE WOULD NOT FIRE AFTER THE FIRST FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC PROCEDURE DURING STAPLING, THE RELOADS LOCKED AND THE DEVICE COULD NOT BE ARMED FOR FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE RELOAD GOT STUCK AND DID NOT ADVANCE. THE SURGEON TRIED TO SQUEEZE THE HANDLE TWICE, BUT COULD NOT SQUEEZE IT ANYMORE. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | us surgical puerto rico | 08-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC PROCEDURE DURING STAPLING, THE RELOADS LOCKED AND THE DEVICE COULD NOT BE ARMED FOR FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE RELOAD GOT STUCK AND DID NOT ADVANCE. THE SURGEON TRIED TO SQUEEZE THE HANDLE TWICE, BUT COULD NOT SQUEEZE IT ANYMORE. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2019 | Malfunction | davis & geck caribe | 08-01-2020 | UNKNOWN ENDO STITCH SULU | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS STICKING. THE SURGEON CONTINUED TO USE THE DEVICE; HOWEVER WHILE MAKING THE LAST STITCH, THE DEVICE GOT STUCK AND THE NEEDLE BROKE IN HALF. THERE WAS NO PATIENT INJURY.
11-12-2019 | Malfunction | davis & geck caribe | 08-01-2020 | UNKNOWN ENDO STITCH SULU | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS STICKING. THE SURGEON CONTINUED TO USE THE DEVICE; HOWEVER WHILE MAKING THE LAST STITCH, THE DEVICE GOT STUCK AND THE NEEDLE BROKE IN HALF. THERE WAS NO PATIENT INJURY.
12-12-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | ILS 33MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE UNKNOWN SURGERY, WHEN SEVERING GASTROINTESTINAL TISSUE ON THE FIRST FIRING, AFTER FIRED, NOTED THE STAPLES MALFORMED. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TRANSECTION OF COLON ON A LAPAROSCOPIC RIGHT COLECTOMY, THE STAPLER WOULD NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2020 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING FIRING OF TISSUE NEXT TO FUNDUS ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND COULD SQUEEZE THE HANDLE BUT THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON STATED THAT THE KNIFE DID NOT CUT ALL THE WAY THROUGH THE END. IT WAS ALSO REPORTED THAT THE SURGEON MANIPULATED THE STUCK TISSUE ON THE REINFORCED MATERIAL TO GET IT OUT. AS SEEN ON THE ATTACHMENT SENT, IT WAS SEEN THAT THERE WAS ROUGH EDGES ON HOW THE TISSUE WAS CUT. THERE WAS NO PATIENT INJURY.
22-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON STATED THAT THE KNIFE DID NOT CUT ALL THE WAY THROUGH THE END. IT WAS ALSO REPORTED THAT THE SURGEON MANIPULATED THE STUCK TISSUE ON THE REINFORCED MATERIAL TO GET IT OUT. AS SEEN ON THE ATTACHMENT SENT, IT WAS SEEN THAT THERE WAS ROUGH EDGES ON HOW THE TISSUE WAS CUT. THERE WAS NO PATIENT INJURY.
22-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Material Deformation | A0406 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON STATED THAT THE KNIFE DID NOT CUT ALL THE WAY THROUGH THE END. IT WAS ALSO REPORTED THAT THE SURGEON MANIPULATED THE STUCK TISSUE ON THE REINFORCED MATERIAL TO GET IT OUT. AS SEEN ON THE ATTACHMENT SENT, IT WAS SEEN THAT THERE WAS ROUGH EDGES ON HOW THE TISSUE WAS CUT. THERE WAS NO PATIENT INJURY.
01-01-2018 | Malfunction | ethicon endo-surgery | 08-01-2020 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AT THE END OF A SIGMOID COLON RESECTION, THE DOCTOR MENTIONED A CIRCULAR STAPLER FIRING IN THE SPRING OF 2018, WHERE THERE WAS NO WASHER PRESENT IN THE CIRCULAR DEVICE IN AN RNY PROCEDURE. TISSUES HAD STAPLES BUT DID FORM AN ANASTOMOSIS. THE DOCTOR REFIRED A SECOND LIKE DEVICE AND IT WORKED FINE. PATIENT IS DOING GREAT. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | SIGNIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON STATED THAT THE KNIFE DID NOT CUT ALL THE WAY THROUGH THE END. IT WAS ALSO REPORTED THAT THE SURGEON MANIPULATED THE STUCK TISSUE ON THE REINFORCED MATERIAL TO GET IT OUT. AS SEEN ON THE ATTACHMENT SENT, IT WAS SEEN THAT THERE WAS ROUGH EDGES ON HOW THE TISSUE WAS CUT. THERE WAS NO PATIENT INJURY.
01-01-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | CONTOUR CURVED CUTTER STAPLER | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, AFTER FIRING THE INSTRUMENT, THERE WERE ONLY SEEN STAPLES AT 1/3 OF THE ROW. THE KNIFE WAS ALSO A BIT DAMAGE OR TORN, AS SEEN ON THE INSTRUMENT. IT DID CUT. SURGERY WAS DELAYED 15 MINUTES. NEEDED TO HANDSEW THE STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
22-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON STATED THAT THE KNIFE DID NOT CUT ALL THE WAY THROUGH THE END. IT WAS ALSO REPORTED THAT THE SURGEON MANIPULATED THE STUCK TISSUE ON THE REINFORCED MATERIAL TO GET IT OUT. AS SEEN ON THE ATTACHMENT SENT, IT WAS SEEN THAT THERE WAS ROUGH EDGES ON HOW THE TISSUE WAS CUT. THERE WAS NO PATIENT INJURY
03-10-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | GST GREEN RELOAD, 60MM, 6 ROW | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THERE WAS A MISFIRING IN TWO DIFFERENT CASES (FOUR MISSING STAPLES FROM THE STAPLE LINE IN EACH CASE). THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
03-10-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | GST GREEN RELOAD, 60MM, 6 ROW | Component Missing | A020602 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THERE WAS A MISFIRING IN TWO DIFFERENT CASES (FOUR MISSING STAPLES FROM THE STAPLE LINE IN EACH CASE). NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
10-12-2019 | Injury | ethicon endo-surgery | 08-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Misfire | A050502 | Cardiac Arrest | Event Description: IT WAS REPORTED THAT DURING A HEPATOBILIARY, RIGHT HEPATECTOMY DONOR SURGERY, PROCEDURE, MISFIRING OF PVE35A ON HEPATIC VEIN, LEAD TO SEVER BLEEDING FOR 10 MINS, AND THE PATIENT GOT CARDIAC ARRESTED. PATIENT CURRENTLY IN ICU.
10-12-2019 | Injury | ethicon endo-surgery | 08-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Misfire | A050502 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A HEPATOBILIARY, RIGHT HEPATECTOMY DONOR SURGERY, PROCEDURE, MISFIRING OF PVE35A ON HEPATIC VEIN, LEAD TO SEVER BLEEDING FOR 10 MINS, AND THE PATIENT GOT CARDIAC ARRESTED. PATIENT CURRENTLY IN ICU.
10-12-2019 | Injury | ethicon endo-surgery | 08-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Misfire | A050502 | Not Applicable | Event Description: IT WAS REPORTED THAT DURING A HEPATOBILIARY, RIGHT HEPATECTOMY DONOR SURGERY, PROCEDURE, MISFIRING OF PVE35A ON HEPATIC VEIN, LEAD TO SEVER BLEEDING FOR 10 MINS, AND THE PATIENT GOT CARDIAC ARRESTED. PATIENT CURRENTLY IN ICU.
10-12-2019 | Injury | ethicon endo-surgery | 08-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Misfire | A050502 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A HEPATOBILIARY, RIGHT HEPATECTOMY DONOR SURGERY, PROCEDURE, MISFIRING OF PVE35A ON HEPATIC VEIN, LEAD TO SEVER BLEEDING FOR 10 MINS, AND THE PATIENT GOT CARDIAC ARRESTED. PATIENT CURRENTLY IN ICU.
11-10-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | ECHELON FLEX | Defective Device | A0203 | No Known Impact Or Consequence To Patient | Event Description: STAPLER FAILED TO FUNCTION PROPERLY.
09-10-2019 | nan | ethicon endo-surgery | 08-01-2020 | ECHELON FLEX | Defective Device | A0203 | No Known Impact Or Consequence To Patient | Event Description: STAPLER FAILED TO FUNCTION PROPERLY.
09-10-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | ENDOPATH | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: TWO STAPLERS ATS45 MISFIRED DURING SURGERY.
30-10-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, THE STAPLER BLOCKED WHILE STAPLING. SURGEON HAD TO CUT MANUALLY. NO PATIENT CONSEQUENCE MENTIONED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
27-11-2019 | Injury | ethicon endo-surgery | 08-01-2020 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT IT WAS THE THIRD PATIENT OPERATED FOR A GASTRIC BYPASS OF THE DAY. NO UNUSUAL BLEEDING OBSERVED DURING THE INITIAL PROCEDURE. THE PATIENT WAS REOPERATED ON THE SAME DAY AS THE INITIAL PROCEDURE. THE ISSUE WAS DETECTED THE AFTERNOON AFTER THE PROCEDURE. THE SURGEON OBSERVED A LEAKAGE DURING THE TEST ON THE STAPLE LINE: THE SUTURE WAS REINFORCED WITH A STITCH. NO PATIENT CONSEQUENCE AFTER THE REOPERATION. THE PATIENT LEFT THE HOSPITAL ON (B)(6) 2019.
27-11-2019 | Injury | ethicon endo-surgery | 08-01-2020 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT IT WAS THE THIRD PATIENT OPERATED FOR A GASTRIC BYPASS OF THE DAY. NO UNUSUAL BLEEDING OBSERVED DURING THE INITIAL PROCEDURE. THE PATIENT WAS REOPERATED ON THE SAME DAY AS THE INITIAL PROCEDURE. THE ISSUE WAS DETECTED THE AFTERNOON AFTER THE PROCEDURE. THE SURGEON OBSERVED A LEAKAGE DURING THE TEST ON THE STAPLE LINE: THE SUTURE WAS REINFORCED WITH A STITCH. NO PATIENT CONSEQUENCE AFTER THE REOPERATION. THE PATIENT LEFT THE HOSPITAL ON (B)(6) 2019.
10-12-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Device Fell | A051203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A THORACIC SURGERY PROCEDURE, THE CUSTOMER CLAIMS THAT A SPRING FELL OUT OF THE STAPLER. THIS OCCURRED PRIOR TO HANDING THE STAPLER BACK TO THE DOCTOR. THE SPRING FELL ONTO THE BACK TABLE. THE STAPLER PERFORMED AND STAPLED WELL BEFORE THE SPRING FELL OUT. A SECOND LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THIS IS ALL THE INFORMATION KNOWN/AVAILABLE REGARDING THIS EVENT.
11-12-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A BARIATRIC PROCEDURE, THE STAPLER ONLY WORKED ON THE RIGHT SIDE OF THE CUT LINE. THE LEFT SIDE DIDN'T STAPLE. A SIMILAR DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2020 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA | Protective Measures Problem | A16 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PRIOR TO THORACOSCOPIC PULMONARY SEGMENTECTOMY, THE BLADE AUTOMATICALLY PUSHED FORWARD PRIOR REMOVING THE YELLOW PROTECTIVE WEDGE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | PROXIMATE*PPH PROCEDURE SET | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY PROCEDURE, THE STAPLES FROM THE PPH03 DIDN'T FIRE PROPERLY, LEADING TO HEAVY BLEEDING. SURGERY WAS DELAYED 35 MINUTES. MONOPOLAR AND HAEMOSTATIC AGENT USED TO MANAGE THE BLEEDING. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
05-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Material Rupture | A0412 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS A POOR STAPLE FORMATION, THE KNIFE DID NOT CUT ALL THE WAY TO THE END. IT WAS ALSO REPORTED THAT THE REINFORCED MATERIAL WAS TORN AND NOT COMPLETELY CUT. THE SURGEON CLIP THE ANASTOMOSIS TO PREVENT BLEEDING AND ANOTHER RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS A POOR STAPLE FORMATION, THE KNIFE DID NOT CUT ALL THE WAY TO THE END. IT WAS ALSO REPORTED THAT THE REINFORCED MATERIAL WAS TORN AND NOT COMPLETELY CUT. THE SURGEON CLIP THE ANASTOMOSIS TO PREVENT BLEEDING AND ANOTHER RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS A POOR STAPLE FORMATION, THE KNIFE DID NOT CUT ALL THE WAY TO THE END. IT WAS ALSO REPORTED THAT THE REINFORCED MATERIAL WAS TORN AND NOT COMPLETELY CUT. THE SURGEON CLIP THE ANASTOMOSIS TO PREVENT BLEEDING AND ANOTHER RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THERE WAS A POOR STAPLE FORMATION, THE KNIFE DID NOT CUT ALL THE WAY TO THE END. IT WAS ALSO REPORTED THAT THE REINFORCED MATERIAL WAS TORN AND NOT COMPLETELY CUT. THE SURGEON CLIP THE ANASTOMOSIS TO PREVENT BLEEDING AND ANOTHER RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING THE LUNG WEDGE DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE HOWEVER THE HANDLE EXPLODED WHILE FIRING, AND AFTER LOADING A NEW STAPLES, THE HANDLE COULD NOT BE FIRED AGAIN. A SUTURE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING THE LUNG WEDGE DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE HOWEVER THE HANDLE EXPLODED WHILE FIRING, AND AFTER LOADING A NEW STAPLES, THE HANDLE COULD NOT BE FIRED AGAIN. A SUTURE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING THE LUNG WEDGE DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE HOWEVER THE HANDLE EXPLODED WHILE FIRING, AND AFTER LOADING A NEW STAPLES, THE HANDLE COULD NOT BE FIRED AGAIN. A SUTURE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN RESECTING THE LUNG WEDGE DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE HOWEVER THE HANDLE EXPLODED WHILE FIRING, AND AFTER LOADING A NEW STAPLES, THE HANDLE COULD NOT BE FIRED AGAIN. A SUTURE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | PROXIMATE*PPH PROCEDURE SET | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY PROCEDURE, THE STAPLES FROM THE PPH03 DIDN'T FIRE PROPERLY, LEADING TO HEAVY BLEEDING. SURGERY WAS DELAYED 35 MINUTES. MONOPOLAR AND HAEMOSTATIC AGENT USED TO MANAGE THE BLEEDING. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
10-12-2019 | Injury | ethicon endo-surgery | 08-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT POST-OP TO AN UNKNOWN PROCEDURE THE PATIENT EXPERIENCED BLOOD LOSS. IT IS UNKNOWN HOW LONG AFTER THE PROCEDURE THIS OCCURRED, AND PHYSICIAN IS UNSURE IF BLEEDING OCCURRED AT STAPLE LINE. PATIENT WAS GIVEN TWO PINTS OF BLOOD AND IS STABLE. NO SURGERY WAS NEEDED.
10-12-2019 | Injury | ethicon endo-surgery | 08-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT POST-OP TO AN UNKNOWN PROCEDURE THE PATIENT EXPERIENCED BLOOD LOSS. IT IS UNKNOWN HOW LONG AFTER THE PROCEDURE THIS OCCURRED, AND PHYSICIAN IS UNSURE IF BLEEDING OCCURRED AT STAPLE LINE. PATIENT WAS GIVEN TWO PINTS OF BLOOD AND IS STABLE. NO SURGERY WAS NEEDED.
11-12-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE GOT JAMMED IN A CLOSING POSITION INSIDE THE SMALL INTESTINE. IMPOSSIBILITY TO UNLATCH THE DEVICE. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE. PROCEDURE: GASTRO-INTESTINAL
06-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | SIGNIA | Difficult or Delayed Activation | A150102 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC WEDGE RESECTION, THE LEFT CONTROL BUTTON WAS TOO SENSITIVE AND VAGUE TO USE IT. WHEN SURGEON JUST TRIED TO OPEN THE JAW, THE LOADING UNIT ARTICULATED TO LEFT SIDE. THE SURGEON TRIED SAME THING SEVERAL TIMES BUT SAME THING HAPPENED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | SIGNIA | Unintended Movement | A0512 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC WEDGE RESECTION, THE LEFT CONTROL BUTTON WAS TOO SENSITIVE AND VAGUE TO USE IT. WHEN SURGEON JUST TRIED TO OPEN THE JAW, THE LOADING UNIT ARTICULATED TO LEFT SIDE. THE SURGEON TRIED SAME THING SEVERAL TIMES BUT SAME THING HAPPENED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | UNKNOWN ENDO GIA SULU | Unintended Movement | A0512 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC WEDGE RESECTION, THE LEFT CONTROL BUTTON WAS TOO SENSITIVE AND VAGUE TO USE IT. WHEN SURGEON JUST TRIED TO OPEN THE JAW, THE LOADING UNIT ARTICULATED TO LEFT SIDE. THE SURGEON TRIED SAME THING SEVERAL TIMES BUT SAME THING HAPPENED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Injury | us surgical puerto rico | 08-01-2020 | GIA | Loss of or Failure to Bond | A040102 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A WHIPPLE OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE AND CUT THE TISSUE, BUT THE STAPLE LINE FELL APART AND ALL STAPLES WERE PARTIALLY FORMED. THE UNFORMED STAPLES FELL INTO THE CAVITY OF THE PATIENT. THE STAPLE LINE WAS FIXED BY OVER SEWING.
03-01-2020 | Injury | us surgical puerto rico | 08-01-2020 | GIA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A WHIPPLE OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE AND CUT THE TISSUE, BUT THE STAPLE LINE FELL APART AND ALL STAPLES WERE PARTIALLY FORMED. THE UNFORMED STAPLES FELL INTO THE CAVITY OF THE PATIENT. THE STAPLE LINE WAS FIXED BY OVER SEWING.
03-01-2020 | Injury | us surgical puerto rico | 08-01-2020 | GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A WHIPPLE OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE AND CUT THE TISSUE, BUT THE STAPLE LINE FELL APART AND ALL STAPLES WERE PARTIALLY FORMED. THE UNFORMED STAPLES FELL INTO THE CAVITY OF THE PATIENT. THE STAPLE LINE WAS FIXED BY OVER SEWING.
03-01-2020 | Injury | us surgical puerto rico | 08-01-2020 | GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A WHIPPLE OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE AND CUT THE TISSUE, BUT THE STAPLE LINE FELL APART AND ALL STAPLES WERE PARTIALLY FORMED. THE UNFORMED STAPLES FELL INTO THE CAVITY OF THE PATIENT. THE STAPLE LINE WAS FIXED BY OVER SEWING.
03-01-2020 | Injury | us surgical puerto rico | 08-01-2020 | GIA | Detachment of Device or Device Component | A0501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A WHIPPLE OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE AND CUT THE TISSUE, BUT THE STAPLE LINE FELL APART AND ALL STAPLES WERE PARTIALLY FORMED. THE UNFORMED STAPLES FELL INTO THE CAVITY OF THE PATIENT. THE STAPLE LINE WAS FIXED BY OVER SEWING.
03-01-2020 | Injury | us surgical puerto rico | 08-01-2020 | GIA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A WHIPPLE OPEN PROCEDURE, THE DEVICE WAS ABLE TO FIRE AND CUT THE TISSUE, BUT THE STAPLE LINE FELL APART AND ALL STAPLES WERE PARTIALLY FORMED. THE UNFORMED STAPLES FELL INTO THE CAVITY OF THE PATIENT. THE STAPLE LINE WAS FIXED BY OVER SEWING.
31-12-2019 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO PULMONARY ARTERY OF THE LUNG DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE DEVICE COULD NOT BE CLOSED PROPERLY AT THE END AND THE STAPLER DID NOT FIRE.
31-12-2019 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO PULMONARY ARTERY OF THE LUNG DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE DEVICE COULD NOT BE CLOSED PROPERLY AT THE END AND THE STAPLER DID NOT FIRE.
31-12-2019 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO PULMONARY ARTERY OF THE LUNG DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE DEVICE COULD NOT BE CLOSED PROPERLY AT THE END AND THE STAPLER DID NOT FIRE.
31-12-2019 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Information | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO PULMONARY ARTERY OF THE LUNG DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE DEVICE COULD NOT BE CLOSED PROPERLY AT THE END AND THE STAPLER DID NOT FIRE.
31-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO PULMONARY ARTERY OF THE LUNG DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE DEVICE COULD NOT BE CLOSED PROPERLY AT THE END AND THE STAPLER DID NOT FIRE.
31-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO PULMONARY ARTERY OF THE LUNG DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE DEVICE COULD NOT BE CLOSED PROPERLY AT THE END AND THE STAPLER DID NOT FIRE.
31-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO PULMONARY ARTERY OF THE LUNG DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE DEVICE COULD NOT BE CLOSED PROPERLY AT THE END AND THE STAPLER DID NOT FIRE.
31-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Information | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO PULMONARY ARTERY OF THE LUNG DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE DEVICE COULD NOT BE CLOSED PROPERLY AT THE END AND THE STAPLER DID NOT FIRE.
16-09-2019 | Injury | covidien lp llc north haven | 08-01-2020 | UNKNOWN EGIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE STUDY, AFTER USING A LINEAR STAPLING METHOD ON LAPAROSCOPIC TOTAL GASTRECTOMY (LTG) FOR GASTRIC CANCER IN 29 PATIENT, ANASTOMOTIC COMPLICATIONS WERE REPORTED INCLUDING STRICTURE, LEAKAGE AND BLEEDING. ADDITIONAL EXAMINATIONS, INCLUDING BLOOD TESTS AND COMPUTED TOMOGRAPHY, WERE PERFORMED.
16-09-2019 | Injury | covidien lp llc north haven | 08-01-2020 | UNKNOWN EGIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Radiation Exposure, Unintended | Event Description: ACCORDING TO THE LITERATURE STUDY, AFTER USING A LINEAR STAPLING METHOD ON LAPAROSCOPIC TOTAL GASTRECTOMY (LTG) FOR GASTRIC CANCER IN 29 PATIENT, ANASTOMOTIC COMPLICATIONS WERE REPORTED INCLUDING STRICTURE, LEAKAGE AND BLEEDING. ADDITIONAL EXAMINATIONS, INCLUDING BLOOD TESTS AND COMPUTED TOMOGRAPHY, WERE PERFORMED.
16-09-2019 | Injury | covidien lp llc north haven | 08-01-2020 | UNKNOWN EGIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE LITERATURE STUDY, AFTER USING A LINEAR STAPLING METHOD ON LAPAROSCOPIC TOTAL GASTRECTOMY (LTG) FOR GASTRIC CANCER IN 29 PATIENT, ANASTOMOTIC COMPLICATIONS WERE REPORTED INCLUDING STRICTURE, LEAKAGE AND BLEEDING. ADDITIONAL EXAMINATIONS, INCLUDING BLOOD TESTS AND COMPUTED TOMOGRAPHY, WERE PERFORMED.
16-09-2019 | Injury | covidien lp llc north haven | 08-01-2020 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE STUDY, AFTER USING CIRCULAR STAPLING METHOD ON LAPAROSCOPIC TOTAL GASTRECTOMY (LTG) FOR GASTRIC CANCER IN 77 PATIENTS, ANASTOMOTIC COMPLICATIONS WERE REPORTED INCLUDING STRICTURE, LEAKAGE AND BLEEDING. ADDITIONAL EXAMINATIONS, INCLUDING BLOOD TESTS AND COMPUTED TOMOGRAPHY, WERE PERFORMED.
16-09-2019 | Injury | covidien lp llc north haven | 08-01-2020 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Radiation Exposure, Unintended | Event Description: ACCORDING TO THE LITERATURE STUDY, AFTER USING CIRCULAR STAPLING METHOD ON LAPAROSCOPIC TOTAL GASTRECTOMY (LTG) FOR GASTRIC CANCER IN 77 PATIENTS, ANASTOMOTIC COMPLICATIONS WERE REPORTED INCLUDING STRICTURE, LEAKAGE AND BLEEDING. ADDITIONAL EXAMINATIONS, INCLUDING BLOOD TESTS AND COMPUTED TOMOGRAPHY, WERE PERFORMED.
16-09-2019 | Injury | covidien lp llc north haven | 08-01-2020 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE LITERATURE STUDY, AFTER USING CIRCULAR STAPLING METHOD ON LAPAROSCOPIC TOTAL GASTRECTOMY (LTG) FOR GASTRIC CANCER IN 77 PATIENTS, ANASTOMOTIC COMPLICATIONS WERE REPORTED INCLUDING STRICTURE, LEAKAGE AND BLEEDING. ADDITIONAL EXAMINATIONS, INCLUDING BLOOD TESTS AND COMPUTED TOMOGRAPHY, WERE PERFORMED.
19-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY PROCEDURE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON AND ABLE TO SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT WORK. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 08-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY PROCEDURE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON AND ABLE TO SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT WORK. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | us surgical puerto rico | 08-01-2020 | UNKNOWN ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY PROCEDURE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON AND ABLE TO SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT WORK. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-07-2019 | Injury | davis & geck caribe | 08-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM (B)(6) 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | davis & geck caribe | 08-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Abscess | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM (B)(6) 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | davis & geck caribe | 08-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Fistula | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM (B)(6) 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | davis & geck caribe | 08-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Injury | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM (B)(6) 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | davis & geck caribe | 08-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Obstruction/Occlusion | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM (B)(6) 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | davis & geck caribe | 08-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM (B)(6) 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | us surgical puerto rico | 08-01-2020 | UNKNOWN ENDO GIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | us surgical puerto rico | 08-01-2020 | UNKNOWN ENDO GIA | Fluid/Blood Leak | A050401 | Abscess | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | us surgical puerto rico | 08-01-2020 | UNKNOWN ENDO GIA | Fluid/Blood Leak | A050401 | Fistula | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | us surgical puerto rico | 08-01-2020 | UNKNOWN ENDO GIA | Fluid/Blood Leak | A050401 | Injury | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | us surgical puerto rico | 08-01-2020 | UNKNOWN ENDO GIA | Fluid/Blood Leak | A050401 | Obstruction/Occlusion | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | us surgical puerto rico | 08-01-2020 | UNKNOWN ENDO GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
02-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE STAPLES WOULD NOT FORM PROPERLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-10-2019 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE SEPARATION OF LUNG PARENCHYMA ON A VIDEO ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, THE DEVICE WAS ABLE TO FIRE STAPLES, HOWEVER POOR STAPLE FORMATION WAS OBSERVED AS IT WAS UNABLE TO FORMED B SHAPE. BOTH HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2020 | Malfunction | covidien lp llc north haven | 08-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, WHEN THE DEVICE WAS BEING FIRED THE FIRST TIME ACROSS THE STOMACH, THE RELOAD WOULD NOT ADVANCE TO STAPLE OR CUT. THE STAPLER WAS NOT ABLE TO FIRE, THE GREEN BUTTON WAS PRESSED AND THE HANDLE DID NOT SQUEEZE. THEY OPENED A NEW STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2020 | Malfunction | us surgical puerto rico | 08-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, WHEN THE DEVICE WAS BEING FIRED THE FIRST TIME ACROSS THE STOMACH, THE RELOAD WOULD NOT ADVANCE TO STAPLE OR CUT. THE STAPLER WAS NOT ABLE TO FIRE, THE GREEN BUTTON WAS PRESSED AND THE HANDLE DID NOT SQUEEZE. THEY OPENED A NEW STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | ethicon endo-surgery | 08-01-2020 | ECHELON STPLR VASC PWR 35X320 P | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: ROBOTIC STAPLER MISFIRED WHILE ATTEMPTING TO STAPLE THE BRONCH. FDA SAFETY REPORT ID# (B)(4).
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A PULMONARY LOBECTOMY, WHEN THE 7TH SHOT WAS FIRED THE LOAD WAS TO BE FIRED THE LOAD LOCKED AND COULDN'T BE FIRED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE ASSEMBLY TO RESOLVE THE ISSUE. THERE WAS NO PATIENT HARM.
13-11-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, THE DEVICE HAD A LOADING PROBLEM. WHEN FIRING THE SIXTH STAPLE, THE DEVICE POPPED AND DID NOT WORK ANYMORE. IT WAS ALSO REPORTED THAT THE DEVICE LOOKED BROKEN. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-11-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, THE DEVICE HAD A LOADING PROBLEM. WHEN FIRING THE SIXTH STAPLE, THE DEVICE POPPED AND DID NOT WORK ANYMORE. IT WAS ALSO REPORTED THAT THE DEVICE LOOKED BROKEN. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-11-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, THE DEVICE HAD A LOADING PROBLEM. WHEN FIRING THE SIXTH STAPLE, THE DEVICE POPPED AND DID NOT WORK ANYMORE. IT WAS ALSO REPORTED THAT THE DEVICE LOOKED BROKEN. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN STAPLING DEVICE | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PULMONARY SEGMENTECTOMY WHEN TRYING TO STAPLE THE TISSUE, THE LOAD OF THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT HARM.
18-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PULMONARY SEGMENTECTOMY WHEN TRYING TO STAPLE THE TISSUE, THE LOAD OF THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT HARM.
12-12-2019 | Injury | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHEN ON LOBAR ARTERY, INNERMOST STAPLES CLOSEST TO THE INC ISION LINE HAVE NOT BEEN FULLY FORMED INTO THE B FORMATION. STAPLES HANG LOOSELY IN THE VESSEL. TO SOLVE THE PROBLEM, A METAL CLIP WAS APPLIED. THE JAWS OF THE DEVICE WAS ALSO LOCKED ON TISSUE. THE PRODUCT STUCK DURING FIRING AND COULD NOT BE DONE UNTIL THE END. A LTHOUGH THE FIRING COULD NOT BE COMPLETED, THE RELOAD COULD BE WITHDRAWN WHEN IT WAS OPEN. THIS SITUATION LED TO A PROLONGED PERIOD OF 10-15 MINUTES AND TENSE MOMENTS.
12-12-2019 | Injury | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHEN ON LOBAR ARTERY, INNERMOST STAPLES CLOSEST TO THE INC ISION LINE HAVE NOT BEEN FULLY FORMED INTO THE B FORMATION. STAPLES HANG LOOSELY IN THE VESSEL. TO SOLVE THE PROBLEM, A METAL CLIP WAS APPLIED. THE JAWS OF THE DEVICE WAS ALSO LOCKED ON TISSUE. THE PRODUCT STUCK DURING FIRING AND COULD NOT BE DONE UNTIL THE END. A LTHOUGH THE FIRING COULD NOT BE COMPLETED, THE RELOAD COULD BE WITHDRAWN WHEN IT WAS OPEN. THIS SITUATION LED TO A PROLONGED PERIOD OF 10-15 MINUTES AND TENSE MOMENTS.
25-07-2019 | Injury | covidien lp llc north haven | 09-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | covidien lp llc north haven | 09-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | Abscess | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | covidien lp llc north haven | 09-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | Fistula | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | covidien lp llc north haven | 09-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | Injury | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | covidien lp llc north haven | 09-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | Obstruction/Occlusion | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
25-07-2019 | Injury | covidien lp llc north haven | 09-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED FROM APRIL 11 TO MAY 2018, 166 PATIENTS UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH A ROUX-EN-Y RECONSTRUCTION WHILE 36 PATIENTS HAD LAPAROSCOPY-ASSISTED PROXIMAL GASTRECTOMY WITH JEJUNAL INTERPOSITION. ESOPHAGEAL INJURY DURING THE ANASTOMOTIC PROCEDURES OCCURRED IN 2 PATIENTS DURING LATG, WHICH REQUIRED OPERATION BE CONVERTED TO OPEN SURGERY BECAUSE THE ESOPHAGEAL WALL WAS WITHDRAWN INTO THE MEDIASTINUM, AND THE HAND-SEWN PURSE-STRING SUTURING WAS IMPOSSIBLE LAPAROSCOPICALLY. ON THE OTHER HAND, 63 PATIENTS NEEDED ADDITIONAL SUTURE REINFORCEMENT AT THE ESOPHAGOJEJUNOSTOMY SITE. ANASTOMOTIC LEAKAGE AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 4 PATIENTS. AMONG THEM, 1 PATIENT DEVELOPED A GRADE IIIA PANCREATIC FISTULA FOLLOWED BY A SECONDARY ANASTOMOTIC LEAKAGE, AND 1 PATIENT REQUIRED ADDITIONAL SURGICAL PROCEDURES. ANASTOMOTIC STRICTURES AT THE ESOPHAGOJEJUNOSTOMY SITE OCCURRED IN 12 PATIENTS. ALL OF THEM WERE SUCCESSFULLY TREATED BY ENDOSCOPIC BALLOON DILATATION. OTHER POSTOPERATIVE COMPLICATIONS ARE BOWEL OBSTRUCTION, PANCREATIC FISTULA, AND INTRA-ABDOMINAL ABSCESS. REOPERATIONS WERE REQUIRED IN 11 PATIENTS. THE CAUSES OF REOPERATION WERE PANCREATIC FISTULA (1 PATIENT), MEDIASTINITIS DUE TO ANASTOMOTIC LEAKAGE (1 PATIENT), AND BOWEL OBSTRUCTION (9 PATIENTS).
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Break | A0401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Break | A0401 | Disability | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Break | A0401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Misfire | A050502 | Disability | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Failure to Form Staple | A050704 | Disability | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Detachment of Device or Device Component | A0501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Detachment of Device or Device Component | A0501 | Disability | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Misassembly by Users | A230501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Misassembly by Users | A230501 | Disability | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | EEA | Misassembly by Users | A230501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, WHILE FIRING, THE HANDLE BROKE, AND THE ANVIL GOT DISENGAGED INTO THE CAVITY. THIS RESULTED TO POOR STAPLE FORMATION AND THE DONUT WAS INCOMPLETE - ONLY HALF WAS FORMED. IT WAS NOTED THAT AN IMPROPERLY MATCHED INSTRUMENT AND ANVIL COMBINATION WERE USED FOR THE PROCEDURE, AND THERE WAS EXCESSIVE TISSUE INCORPORATED INTO THE BARREL OF THE STAPLER. THE INCISION WAS EXTENDED MORE THAN ONE INCH. A COLOSTOMY WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
27-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | GIA | Misfire | A050502 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE PARTIALLY FIRED. THE STAPLER DID NOT STAPLE COMPLETELY THE TISSUE.
19-10-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PROCEDURE, RELOAD HAS FAILED AND WAS NOT TRIGGERING. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE AND THE SURGERY WENT SMOOTHLY. NO PATIENT HARM.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, DURING STOMACH TRANSECTION, THE RELOAD FIRED ONLY HALF AND RETRACTED AUTOMATICALLY. SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
05-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY PROCEDURE, THE DEVICE WERE ABLE TO BE FIRED, HOWEVER THE RELOAD COMPONENT WHICH IS THE "PIN" FELL INTO THE PATIENT CAVITY. TO RESOLVE THE ISSUE THE SURGEON RETRIEVED THE COMPONENT USING GRASPER. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | TA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE WORKING ON THE INTESTINE, THE DEVICE WAS NOT ABLE TO FIRE. SURGEON DID THE ANASTOMOSE BY HAND WITH SUTURE. NO PATIENT INJURY.
12-12-2019 | Malfunction | davol inc sub c.r bard | 09-01-2020 | SORBAFIX 30 COUNT | Incorrect, Inadequate or Imprecise Result or Readings | A0908 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHILE FIXATING A VENTRALIGHT ST WITH THE CANNULA PRESSED AGAINST THE MESH WITH ADEQUATE BACK PRESSURE, NONE OF THE SORBAFIX TACKS WERE GRABBING ON TO THE FASCIA AND THEY FAILED TO TACK. THE TACKS WERE FIXATED THROUGH THE MESH, HOWEVER WERE VERY EASILY PULLED AWAY FROM THE FASCIA AS THEY DIDN'T PENETRATE THE ABDOMINAL WALL. THE TACKS WERE ALSO JAMMING AND THE SORBAFIX HAD TO BE PULLED OUT OF THE TROCAR TWICE IN AN ATTEMPT TO FIRE AND TRY TO ¿CLEAR¿ THE CANNULA OF THE JAMMED TACKS. A DIFFERENT SORBAFIX WAS USED TO FIXATE THE MESH TO THE ABDOMINAL WALL. THE FASTENERS WERE SECURELY FIXATED TO MESH AND WERE NOT REQUIRED TO BE REMOVED FROM ABDOMEN. AS REPORTED, THERE WAS NO PATIENT IMPACT OR INJURY.
12-12-2019 | Malfunction | davol inc sub c.r bard | 09-01-2020 | SORBAFIX 30 COUNT | Inaccurate Synchronization | A071202 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHILE FIXATING A VENTRALIGHT ST WITH THE CANNULA PRESSED AGAINST THE MESH WITH ADEQUATE BACK PRESSURE, NONE OF THE SORBAFIX TACKS WERE GRABBING ON TO THE FASCIA AND THEY FAILED TO TACK. THE TACKS WERE FIXATED THROUGH THE MESH, HOWEVER WERE VERY EASILY PULLED AWAY FROM THE FASCIA AS THEY DIDN'T PENETRATE THE ABDOMINAL WALL. THE TACKS WERE ALSO JAMMING AND THE SORBAFIX HAD TO BE PULLED OUT OF THE TROCAR TWICE IN AN ATTEMPT TO FIRE AND TRY TO ¿CLEAR¿ THE CANNULA OF THE JAMMED TACKS. A DIFFERENT SORBAFIX WAS USED TO FIXATE THE MESH TO THE ABDOMINAL WALL. THE FASTENERS WERE SECURELY FIXATED TO MESH AND WERE NOT REQUIRED TO BE REMOVED FROM ABDOMEN. AS REPORTED, THERE WAS NO PATIENT IMPACT OR INJURY.
06-01-2020 | Injury | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
06-01-2020 | Injury | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
01-01-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A RECTOSIGMOIDECTOMY SURGERY, THE DOCTOR USED THE CDH AND EVEN AFTER USING THE CORRECT TECHNIQUE OF USING THE DEVICE AS HE ALWAYS USED THE PRODUCT HE NOTICED THAT SOME STAPLES WERE NOT TIGHTLY CLOSED SO THE DOCTOR SUTURED THE STAPLES TO STRENGTHEN THE ANASTOMOSIS; SUTURE WITH SUTURE OVER THE REINFORCEMENT CLIPS WAS NECESSARY. SURGERY WAS DELAYED 30 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED AND THERE WERE NO PATIENT CONSEQUENCES REPORTED.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | SIGNIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE GASTRECTOMY SURGERY, THE ADAPTER WOULD NOT TRIGGER LOAD APPLIED. ANOTHER ADAPTER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING TOTALLY LAPAROSCOPIC TOTAL GASTRECTOMY (TLTG), STAPLE LINE WAS COMPLETED BUT CUTTING WAS NOT PERFORMED. SURGICAL SHEARS HAD TO BE USED TO CUT AND ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
07-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER PRIOR TO USE, THE RELOAD WAS BROKEN. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY AND PULMONARY WEDGE PROCEDURE, THERE WAS A POOR STAPLE FORMATION WHEREIN THE STAPLES WERE NOT SHAPED OR BADLY SHAPED. A NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER ON A LAPAROSCOPIC SLEEVE GASTRECTOMY WHILE THE DEVICE WAS BEING APPLIED TO THE FUNDUS, THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. ANOTHER RELOAD WAS USED TO FIX THE STAPLE LINE AND TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER ON A LAPAROSCOPIC SLEEVE GASTRECTOMY WHILE THE DEVICE WAS BEING APPLIED TO THE FUNDUS, THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. ANOTHER RELOAD WAS USED TO FIX THE STAPLE LINE AND TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER AT FIFTH FIRING DURING A LUNG PARENCHYMA RESECTION ON A THORACOSCOPIC VIDEO ASSISTED SURGERY, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE; HOWEVER, THE ANCHORING SUTURE ON THE CARTRIDGE SIDE WAS DISENGAGED. AS A RESULT, THE LUNG PARENCHYMA AND THE CARTRIDGE COULD NOT BE SEPARATED BUT THE SHEET CAME OFF WITHOUT ANY PROBLEMS. A FORCEPS AND SCISSORS WERE USED TO CUT THE TISSUE BY FORCE TO BE REMOVED FROM THE DEVICE. NO TISSUE DAMAGE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER AT FIFTH FIRING DURING A LUNG PARENCHYMA RESECTION ON A THORACOSCOPIC VIDEO ASSISTED SURGERY, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE; HOWEVER, THE ANCHORING SUTURE ON THE CARTRIDGE SIDE WAS DISENGAGED. AS A RESULT, THE LUNG PARENCHYMA AND THE CARTRIDGE COULD NOT BE SEPARATED BUT THE SHEET CAME OFF WITHOUT ANY PROBLEMS. A FORCEPS AND SCISSORS WERE USED TO CUT THE TISSUE BY FORCE TO BE REMOVED FROM THE DEVICE. NO TISSUE DAMAGE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | SIGNIA | Self-Activation or Keying | A150104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE HANDLE AND ADAPTER TRIGGERED AUTOMATICALLY WHILE IN USE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE HANDLE AND ADAPTER TRIGGERED AUTOMATICALLY WHILE IN USE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, ON THE FIRST FIRING, THE STAPLER MADE A NOISE AND THE JAWS WENT LOOSE AND BROKE OFF. CHANGED THE RELOAD BUT NO SUCCESS. CHANGED THE RELOAD AND THE HANDLE THEN EVERYTHING WORKED NORMALLY AND THE PROCEDURE WAS COMPLETE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, ON THE FIRST FIRING, THE STAPLER MADE A NOISE AND THE JAWS WENT LOOSE AND BROKE OFF. CHANGED THE RELOAD BUT NO SUCCESS. CHANGED THE RELOAD AND THE HANDLE THEN EVERYTHING WORKED NORMALLY AND THE PROCEDURE WAS COMPLETE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN ENDO GIA SULU | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, ON THE FIRST FIRING, THE STAPLER MADE A NOISE AND THE JAWS WENT LOOSE AND BROKE OFF. CHANGED THE RELOAD BUT NO SUCCESS. CHANGED THE RELOAD AND THE HANDLE THEN EVERYTHING WORKED NORMALLY AND THE PROCEDURE WAS COMPLETE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, ON THE FIRST FIRING, THE STAPLER MADE A NOISE AND THE JAWS WENT LOOSE AND BROKE OFF. CHANGED THE RELOAD BUT NO SUCCESS. CHANGED THE RELOAD AND THE HANDLE THEN EVERYTHING WORKED NORMALLY AND THE PROCEDURE WAS COMPLETE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER AT FIFTH FIRING DURING A LUNG PARENCHYMA RESECTION ON A THORACOSCOPIC VIDEO ASSISTED SURGERY, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE; HOWEVER, THE ANCHORING SUTURE ON THE CARTRIDGE SIDE WAS DISENGAGED. AS A RESULT, THE LUNG PARENCHYMA AND THE CARTRIDGE COULD NOT BE SEPARATED BUT THE SHEET CAME OFF WITHOUT ANY PROBLEMS. A FORCEPS AND SCISSORS WERE USED TO CUT THE TISSUE BY FORCE TO BE REMOVED FROM THE DEVICE. NO TISSUE DAMAGE. THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MIN. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL RESECTION ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS NOT ABLE TO FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL RESECTION ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS NOT ABLE TO FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL RESECTION ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS NOT ABLE TO FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL RESECTION ON A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE WAS NOT ABLE TO FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-11-2019 | Injury | ethicon endo-surgery | 09-01-2020 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: BLEEDING POST-OP. IT WAS REPORTED THAT THE PATIENT INDICATION IS RECTUM CANCER. 9 DAYS AFTER THE RECTUM SURGERY, NOTED THE BLEEDING.
16-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | PROXIMATE*RL LINEAR STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A TUMOR DEBULKING PROCEDURE, WERE STAPLING ON THE COLON, CLOSED ON THICK TISSUE AND WHEN THE PHYSICIAN WENT TO SQUEEZE THE HANDLE THE DEVICE 'POPPED' OPEN AND THE STAPLES DIDN'T FORM PROPERLY. HE TRIED IT AGAIN AND COMPLETED THE CASE WITH THE DEVICE. THERE WERE NO PATIENT CONSEQUENCES.
06-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, BYPASS OF GASTROENTEROSTOMY AT THE TIME OF SMALL BOWEL RESECTION DURING RY RECONSTRUCTION, THE HANDLE WAS ABLE TO BE SQUEEZED A LITTLE, BUT IT BECAME UNABLE TO BE SQUEEZED. SINCE NO STAPLES WERE PLACED TO THE TISSUE, A NEW HANDLE AND CARTRIDGE WERE USED AND THERE WAS NO PROBLEM. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, BYPASS OF GASTROENTEROSTOMY AT THE TIME OF SMALL BOWEL RESECTION DURING RY RECONSTRUCTION, THE HANDLE WAS ABLE TO BE SQUEEZED A LITTLE, BUT IT BECAME UNABLE TO BE SQUEEZED. SINCE NO STAPLES WERE PLACED TO THE TISSUE, A NEW HANDLE AND CARTRIDGE WERE USED AND THERE WAS NO PROBLEM. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ALSO ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ALSO ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ILS 21MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE SMALL INTESTINAL STOMA CLOSURE SURGERY, AFTER FIRED WITHOUT AUDIBLE FEEDBACK, DIFFICULT TO REMOVE THE DEVICE, USED SCALPEL TO CUT THE TISSUE, REMOVED THE DEVICE FROM PATIENT AND NOTED SOME STAPLES MALFORMED. USED SUTURE TO OVERSEW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
05-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | GIA | Misfire | A050502 | No Information | Event Description: ACCORDING TO THE REPORTER, THE SURGEON SQUEEZED THE HANDLE; HOWEVER, THE DEVICE GOT STUCK AND STOPPED MOVING DURING FIRING. THE SAME EVENT HAPPENED ON ANOTHER RELOAD.
05-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | GIA | Misfire | A050502 | No Information | Event Description: ACCORDING TO THE REPORTER, THE SURGEON SQUEEZED THE HANDLE; HOWEVER, THE DEVICE GOT STUCK AND STOPPED MOVING DURING FIRING. THE SAME EVENT HAPPENED ON ANOTHER RELOAD.
31-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS UNABLE TO SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. IT WAS ALSO REPORTED THAT BOTH RELOADS DID NOT FIRE SO A THIRD RELOAD WAS USED TOGETHER WITH A NEW HANDLE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
12-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE ENDOSCOPIC RECTAL SURGERY, WHEN SEVERING RECTUM TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED, AND SOME STAPLES WERE NOT IN THE TISSUE BUT ON THE TISSUE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, WHILE DETACHING THE LOBAR TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, WHILE DETACHING THE LOBAR TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ILS 21MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE SMALL INTESTINAL STOMA CLOSURE SURGERY, AFTER FIRED WITHOUT AUDIBLE FEEDBACK, DIFFICULT TO REMOVE THE DEVICE, USED SCALPEL TO CUT THE TISSUE, REMOVED THE DEVICE FROM PATIENT AND NOTED SOME STAPLES MALFORMED. USED SUTURE TO OVERSEW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
10-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL PROCEDURE, WHEN THEY WENT TO GO FIRE THE STAPLER. THE STAPLER ONLY CUT IT DID NOT STAPLE. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN FIRING THE INSTRUMENT, THE INSTRUMENT DID NOT STAPLE DURING THE ENTIRE ROW. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-11-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ILS 29MM, CURVED | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: CUTTING ISSUE. IT WAS REPORTED THAT THE PATIENT'S INDICATION IS RECTUM CANCER. DURING THE LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER SURGERY,WHEN SEVERING THE RECTUM AND SIGMOID,AFTER FIRED, COULD NOT REMOVED THE DEVICE, AFTER CHECKING, NOTED THE TISSUE WAS NOT CUT OFF, CHANGED ANOTHER NEW DEVICE TO RE-ANASTOMOSE AND USED SUTURE TO OVERSEW.
30-11-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: INCOMPLETE ANASTOMOSIS. IT WAS REPORTED THAT DURING THE LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER SURGERY,WHEN SEVERING THE RECTUM AND SIGMOID, AFTER FIRED, NOTED THE ANASTOMOSIS WAS INCOMPLETE, USED SUTURE TO OVERSEW.
11-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A SLEEVE GASTRECTOMY, THE STAPLES FROM TWO RELOADS WERE MALFORMED WHEN FIRING ON THE STOMACH. THE STAPLE LINE WAS OVER-SEWED WITH UNKNOWN SUTURE. THERE WAS A DELAY IN THE PROCEDURE OF FIVE MINUTES. FRAGMENTS WERE INDICATED AS BEING GENERATED AND WERE REMOVED SUCCESSFULLY FROM THE PATIENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THIS IS ALL THE INFORMATION KNOWN/AVAILABLE REGARDING THIS EVENT.
12-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHEN ON LOBAR ARTERY, INNERMOST STAPLES CLOSEST TO THE INCISION LINE HAVE NOT BEEN FULLY FORMED INTO THE B FORMATION. STAPLES HANG LOOSELY IN THE VESSEL. TO SOLVE THE PROBLEM, A METAL CLIP WAS APPLIED. THE JAWS OF THE DEVICE WAS ALSO LOCKED ON TISSUE. THE PRODUCT STUCK DURING FIRING AND COULD NOT BE DONE UNTIL THE END. ALTHOUGH THE FIRING COULD NOT BE COMPLETED, THE RELOAD COULD BE WITHDRAWN WHEN IT WAS OPEN. THIS SITUATION LED TO A PROLONGED PERIOD OF 10-15 MINUTES AND TENSE MOMENTS.
22-11-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE DEVICE WAS FIRED AND THE HANDLE MALFUNCTIONED. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE WITH NO PATIENT CONSEQUENCES REPORTED.
17-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE FAILED TO FIRE STAPLES.
17-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE FAILED TO FIRE STAPLES.
06-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIGESTIVE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ENDO ILS 29MM, CURVED | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL ANASTOMOSIS, THE STAPLING WAS DEFECTIVE. THE ANASTOMOSIS HAS BEEN PERFORMED AGAIN WITH ANOTHER DEVICE.
06-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIGESTIVE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LIVER RESECTION, THE DEVICE DID NOT FIRE. ANOTHER NEW DEVICE AND RELOAD WAS OPENED AND WORKED PERFECTLY. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LIVER RESECTION, THE DEVICE DID NOT FIRE. ANOTHER NEW DEVICE AND RELOAD WAS OPENED AND WORKED PERFECTLY. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | SIGNIA | Overheating of Device | A1005 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A WEDGE SEGMENTAL RESECTION PROCEDURE, AT FIRST FIRING, AFTER THE HANDLE WAS INSERTED TO THE SHELL, BLACK OUT OF THE DISPLAY WAS SEEN AND THE HANDLE COULD NOT BE USED. THE HANDLE WAS REMOVED FROM THE CLEAN FIELD AND WAS INSERTED INTO CHARGER, THEN RE-START UP WAS ATTEMPTED, BUT THE STATUS DID NOT CHARGE. THE CHARGER TURNED GREEN WHEN THE POWER HANDLE WAS INSERTED. IT WAS ALSO REPORTED THAT THE SURGEON FELT THE HANDLE TO BE HEATED, NOT HOT BUT WARM LASTING FOR SEVERAL HOURS. IT WAS NOTED THAT THE HEAT SOURCE WAS AROUND THE INSIDE OF THE CORNER WHERE WAS CURVED INTO L-SHAPE OF THE HANDLE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED ON THE PATIENT. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 09-01-2020 | SIGNIA | Complete Loss of Power | A070801 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A WEDGE SEGMENTAL RESECTION PROCEDURE, AT FIRST FIRING, AFTER THE HANDLE WAS INSERTED TO THE SHELL, BLACK OUT OF THE DISPLAY WAS SEEN AND THE HANDLE COULD NOT BE USED. THE HANDLE WAS REMOVED FROM THE CLEAN FIELD AND WAS INSERTED INTO CHARGER, THEN RE-START UP WAS ATTEMPTED, BUT THE STATUS DID NOT CHARGE. THE CHARGER TURNED GREEN WHEN THE POWER HANDLE WAS INSERTED. IT WAS ALSO REPORTED THAT THE SURGEON FELT THE HANDLE TO BE HEATED, NOT HOT BUT WARM LASTING FOR SEVERAL HOURS. IT WAS NOTED THAT THE HEAT SOURCE WAS AROUND THE INSIDE OF THE CORNER WHERE WAS CURVED INTO L-SHAPE OF THE HANDLE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED ON THE PATIENT. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY BARIATRIC BYPASS SURGERY PROCEDURE, AT THE TIME OF THE ENTERO-ANASTOMOSIS THE BEIGE LOAD LOCKED AND DID NOT FIRE. IT WAS ASSEMBLED AND DISASSEMBLED SEVERAL TIMES AND YET IT REMAINED LOCKED. AFTER THE RELOAD WAS REPLACED THE PROCEDURE WAS COMPLETED WITHOUT ISSUE. THERE WAS NO PATIENT INJURY.
08-11-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | GREEN RELOAD FOR ECHELON 45 | Device Fell | A051203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC LEFT LOWER LOBECTOMY SURGERY, OPENED THE PACKING AND NOTED SOME COMPONENTS FELL OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
18-11-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON FIRED ON TARGET THIN TISSUE AND LOOKED ON FOR CHECKING WHETHER IT WELL CUT OR NOT. AFTER REMOVING THE STAPLER BODY, SURGEON SAW TISSUE WAS NOT CUT CLEARLY BUT IT DISTORTED WITH NO REASON. LATER, SURGEON CUT MORE TISSUE AND STAPLED AGAIN WITH ANOTHER NEW STAPLER AND SURGERY WAS DONE WITH NO NEGATIVE ISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
11-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE POWERED ECHELON WHILE FIRING IN A BYPASS PROCEDURE THAT STAPLER DRAGS FROM THE TISSUE (PULLED BACKWARDS WHILE STAPLING) AND CAUSES MALFORMATION OF THE STAPLES AND IT MAY CAUSE LEAKS.
06-01-2020 | Injury | celestica de monterrey s.a de c.v | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
06-01-2020 | Injury | celestica de monterrey s.a de c.v | 09-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
06-01-2020 | Injury | celestica de monterrey s.a de c.v | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
06-01-2020 | Injury | celestica de monterrey s.a de c.v | 09-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE OVER THE RENAL ARTERY OR VEIN DURING A ROBOTIC LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE, AND THE JAWS WOULD NOT CLOSE. HENCE, THE DEVICE DID NOT FIRE. THE SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO HAVING TO OPEN MULTIPLE HANDLES AND RELOADS. THE PROCEDURE WAS CONVERTED INTO AN OPEN PROCEDURE TO COMPLETE THE CASE.
10-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | PWRD ECH FLEX 45MM 34CM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RIGHT HEMIHEPATECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN STAPLING DEVICE | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IT WAS APPLIED ON THE ANTRUM SIDE, THE RELOAD WAS NOT GETTING UNLOADED FROM THE ADAPTER, THE STAPLER WAS NOT ABLE TO FIRE. THE JAWS OF THE DEVICE LOCKED ON THE TISSUE AND WAS REMOVED BY PULLING BACK WITH THE HELP OF A GRASPER. THE POWERED STAPLER WAS REPLACED WITH A MANUAL STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN STAPLING DEVICE | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IT WAS APPLIED ON THE ANTRUM SIDE, THE RELOAD WAS NOT GETTING UNLOADED FROM THE ADAPTER, THE STAPLER WAS NOT ABLE TO FIRE. THE JAWS OF THE DEVICE LOCKED ON THE TISSUE AND WAS REMOVED BY PULLING BACK WITH THE HELP OF A GRASPER. THE POWERED STAPLER WAS REPLACED WITH A MANUAL STAPLER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER LOCKED OUT ONTO TISSUE AND WOULD NOT OPEN ON RELEASE. THE USER CUT THE STAPLER OFF OF THE TISSUE TO REMOVE IT. A SECOND LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS A FIVE MINUTE DELAY IN THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
13-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER LOCKED OUT ONTO TISSUE AND WOULD NOT OPEN ON RELEASE. THE USER CUT THE STAPLER OFF OF THE TISSUE TO REMOVE IT. A SECOND LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS A FIVE MINUTE DELAY IN THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
03-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LUNG VOLUME REDUCTION SURGERY, COULD NOT FIRE WHEN SEVERING EMPHYSEMA TISSUE. THIS OCCURRED ON THE THIRD FIRING. COULD NOT OPEN THE JAW SO USED ANOTHER MEDICAL DEVICE TO PRY OPEN THE JAWS OF THE DEVICE AND THE JAWS DID EVENTUALLY OPEN (SO DID NOT NEED TO CUT OFF THE DEVICE). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LUNG VOLUME REDUCTION SURGERY, COULD NOT FIRE WHEN SEVERING EMPHYSEMA TISSUE. THIS OCCURRED ON THE THIRD FIRING. COULD NOT OPEN THE JAW SO USED ANOTHER MEDICAL DEVICE TO PRY OPEN THE JAWS OF THE DEVICE AND THE JAWS DID EVENTUALLY OPEN (SO DID NOT NEED TO CUT OFF THE DEVICE). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE RELOAD WAS NOT FIRING. ANOTHER RELOAD WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A NEPHRECTOMY PROCEDURE, THE DEVICE WITH WHITE RELOAD WAS USED. THE REP WAS NOTIFIED THAT STAPLER TRANSECTED, BUT ONLY PARTIALLY STAPLED. SOME STAPLES WERE FORMED AT THE DISTAL TIP, BUT IT WAS REPORTED THEY DID NOT ALL FORM PROXIMALLY. INTRAOPERATIVE BLEEDING WAS EXPERIENCED, BUT WAS EFFECTIVELY MANAGED. A NEW STAPLER AND RELOAD WERE USED TO SUCCESSFULLY COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | us surgical puerto rico | 09-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN EGIA SULU | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN EGIA SULU | Entrapment of Device | A150208 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN EGIA SULU | Failure to Fire | A050501 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN EGIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
30-12-2019 | Malfunction | covidien lp llc north haven | 09-01-2020 | UNKNOWN EGIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRING OF THE RELOADS IN A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE AND THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE INSTRUMENT WAS REMOVED USING THE BLACK RETURN KNOB. ANOTHER STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AS A RESULT OF THIS PROBLEM.
16-12-2019 | Malfunction | ethicon endo-surgery | 09-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, THE DEVICE COULD ONLY BE FIRED PARTLY AND THE STAPLES WERE NOT FORMED IN B-SHAPE. SURGERY WAS COMPLETED WITH THE ENDO GIA FROM MEDTRONIC. THERE WAS NO HARM TO THE PATIENT REPORTED.
11-11-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING THE VESSEL ON THE 2ND FIRING, AFTER FIRED, NOTED OOZING. STOPPED OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Device Slipped | A051204 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN MAKING THE LAST SHOT DURING A BARIATRIC PROCEDURE WITH A POWERED ECHELON FLEX STAPLER, THE SURGEON FELT THAT THE CARTRIDGE, AT THE TIME OF STAPLING, SLIPPED, PRESENTING A LINE THAT WAS NOT UNIFORM AND PRESENTED LOOSE STAPLES. THIS DID NOT PRESENT ADEQUATE HEMOSTASIS AND BLED IN THE STAPLE LINE. AT THAT SAME TIME, ANOTHER CARTRIDGE WAS DELIVERED TO THE DOCTOR TO MAKE ANOTHER SHOT IN THE APPROPRIATE AREA. THE PATIENT HAS NOT PRESENTED COMPLICATIONS UNTIL THIS MOMENT AFTER ALMOST 17 HOURS THAT HAVE ELAPSED SINCE THE SURGERY.
12-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Device Slipped | A051204 | Blood Loss | Event Description: IT WAS REPORTED THAT WHEN MAKING THE LAST SHOT DURING A BARIATRIC PROCEDURE WITH A POWERED ECHELON FLEX STAPLER, THE SURGEON FELT THAT THE CARTRIDGE, AT THE TIME OF STAPLING, SLIPPED, PRESENTING A LINE THAT WAS NOT UNIFORM AND PRESENTED LOOSE STAPLES. THIS DID NOT PRESENT ADEQUATE HEMOSTASIS AND BLED IN THE STAPLE LINE. AT THAT SAME TIME, ANOTHER CARTRIDGE WAS DELIVERED TO THE DOCTOR TO MAKE ANOTHER SHOT IN THE APPROPRIATE AREA. THE PATIENT HAS NOT PRESENTED COMPLICATIONS UNTIL THIS MOMENT AFTER ALMOST 17 HOURS THAT HAVE ELAPSED SINCE THE SURGERY.
12-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Device Slipped | A051204 | No Code Available | Event Description: IT WAS REPORTED THAT WHEN MAKING THE LAST SHOT DURING A BARIATRIC PROCEDURE WITH A POWERED ECHELON FLEX STAPLER, THE SURGEON FELT THAT THE CARTRIDGE, AT THE TIME OF STAPLING, SLIPPED, PRESENTING A LINE THAT WAS NOT UNIFORM AND PRESENTED LOOSE STAPLES. THIS DID NOT PRESENT ADEQUATE HEMOSTASIS AND BLED IN THE STAPLE LINE. AT THAT SAME TIME, ANOTHER CARTRIDGE WAS DELIVERED TO THE DOCTOR TO MAKE ANOTHER SHOT IN THE APPROPRIATE AREA. THE PATIENT HAS NOT PRESENTED COMPLICATIONS UNTIL THIS MOMENT AFTER ALMOST 17 HOURS THAT HAVE ELAPSED SINCE THE SURGERY.
12-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Align | A050701 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN MAKING THE LAST SHOT DURING A BARIATRIC PROCEDURE WITH A POWERED ECHELON FLEX STAPLER, THE SURGEON FELT THAT THE CARTRIDGE, AT THE TIME OF STAPLING, SLIPPED, PRESENTING A LINE THAT WAS NOT UNIFORM AND PRESENTED LOOSE STAPLES. THIS DID NOT PRESENT ADEQUATE HEMOSTASIS AND BLED IN THE STAPLE LINE. AT THAT SAME TIME, ANOTHER CARTRIDGE WAS DELIVERED TO THE DOCTOR TO MAKE ANOTHER SHOT IN THE APPROPRIATE AREA. THE PATIENT HAS NOT PRESENTED COMPLICATIONS UNTIL THIS MOMENT AFTER ALMOST 17 HOURS THAT HAVE ELAPSED SINCE THE SURGERY.
12-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Align | A050701 | Blood Loss | Event Description: IT WAS REPORTED THAT WHEN MAKING THE LAST SHOT DURING A BARIATRIC PROCEDURE WITH A POWERED ECHELON FLEX STAPLER, THE SURGEON FELT THAT THE CARTRIDGE, AT THE TIME OF STAPLING, SLIPPED, PRESENTING A LINE THAT WAS NOT UNIFORM AND PRESENTED LOOSE STAPLES. THIS DID NOT PRESENT ADEQUATE HEMOSTASIS AND BLED IN THE STAPLE LINE. AT THAT SAME TIME, ANOTHER CARTRIDGE WAS DELIVERED TO THE DOCTOR TO MAKE ANOTHER SHOT IN THE APPROPRIATE AREA. THE PATIENT HAS NOT PRESENTED COMPLICATIONS UNTIL THIS MOMENT AFTER ALMOST 17 HOURS THAT HAVE ELAPSED SINCE THE SURGERY.
12-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Align | A050701 | No Code Available | Event Description: IT WAS REPORTED THAT WHEN MAKING THE LAST SHOT DURING A BARIATRIC PROCEDURE WITH A POWERED ECHELON FLEX STAPLER, THE SURGEON FELT THAT THE CARTRIDGE, AT THE TIME OF STAPLING, SLIPPED, PRESENTING A LINE THAT WAS NOT UNIFORM AND PRESENTED LOOSE STAPLES. THIS DID NOT PRESENT ADEQUATE HEMOSTASIS AND BLED IN THE STAPLE LINE. AT THAT SAME TIME, ANOTHER CARTRIDGE WAS DELIVERED TO THE DOCTOR TO MAKE ANOTHER SHOT IN THE APPROPRIATE AREA. THE PATIENT HAS NOT PRESENTED COMPLICATIONS UNTIL THIS MOMENT AFTER ALMOST 17 HOURS THAT HAVE ELAPSED SINCE THE SURGERY.
12-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN MAKING THE LAST SHOT DURING A BARIATRIC PROCEDURE WITH A POWERED ECHELON FLEX STAPLER, THE SURGEON FELT THAT THE CARTRIDGE, AT THE TIME OF STAPLING, SLIPPED, PRESENTING A LINE THAT WAS NOT UNIFORM AND PRESENTED LOOSE STAPLES. THIS DID NOT PRESENT ADEQUATE HEMOSTASIS AND BLED IN THE STAPLE LINE. AT THAT SAME TIME, ANOTHER CARTRIDGE WAS DELIVERED TO THE DOCTOR TO MAKE ANOTHER SHOT IN THE APPROPRIATE AREA. THE PATIENT HAS NOT PRESENTED COMPLICATIONS UNTIL THIS MOMENT AFTER ALMOST 17 HOURS THAT HAVE ELAPSED SINCE THE SURGERY.
12-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | Blood Loss | Event Description: IT WAS REPORTED THAT WHEN MAKING THE LAST SHOT DURING A BARIATRIC PROCEDURE WITH A POWERED ECHELON FLEX STAPLER, THE SURGEON FELT THAT THE CARTRIDGE, AT THE TIME OF STAPLING, SLIPPED, PRESENTING A LINE THAT WAS NOT UNIFORM AND PRESENTED LOOSE STAPLES. THIS DID NOT PRESENT ADEQUATE HEMOSTASIS AND BLED IN THE STAPLE LINE. AT THAT SAME TIME, ANOTHER CARTRIDGE WAS DELIVERED TO THE DOCTOR TO MAKE ANOTHER SHOT IN THE APPROPRIATE AREA. THE PATIENT HAS NOT PRESENTED COMPLICATIONS UNTIL THIS MOMENT AFTER ALMOST 17 HOURS THAT HAVE ELAPSED SINCE THE SURGERY.
12-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT WHEN MAKING THE LAST SHOT DURING A BARIATRIC PROCEDURE WITH A POWERED ECHELON FLEX STAPLER, THE SURGEON FELT THAT THE CARTRIDGE, AT THE TIME OF STAPLING, SLIPPED, PRESENTING A LINE THAT WAS NOT UNIFORM AND PRESENTED LOOSE STAPLES. THIS DID NOT PRESENT ADEQUATE HEMOSTASIS AND BLED IN THE STAPLE LINE. AT THAT SAME TIME, ANOTHER CARTRIDGE WAS DELIVERED TO THE DOCTOR TO MAKE ANOTHER SHOT IN THE APPROPRIATE AREA. THE PATIENT HAS NOT PRESENTED COMPLICATIONS UNTIL THIS MOMENT AFTER ALMOST 17 HOURS THAT HAVE ELAPSED SINCE THE SURGERY.
11-12-2019 | Malfunction | covidien lp | 10-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: ENDO GIA STAPLER STOPPED WORKING, DID NOT FIRE LOAD.
20-11-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | ETS*45 ENDO LIN CUT RELOAD WH | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY, A DEVICE LOADING THE TR45W WAS LOCKED OUT AT THE FIRST FIRING. THE DEVICE WAS USED ON THE BLOOD VESSEL. THE SAME DEVICE LOADING ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-11-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | GST GOLD RELOAD, 60MM, 6 ROW | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN COLECTOMY, THE STAPLE DRIVERS FELL OFF THE CARTRIDGE BEFORE USE. NO PIECES FELL INTO THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
20-11-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | GST GREEN RELOAD, 60MM, 6 ROW | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN THE DEVICE WAS OPENED, IT WAS NOTED THAT THE METAL PIECE THAT SITS UNDERNEATH THE RELOAD WAS NOT ATTACHED. A SIMILAR DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
23-11-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | ENDOPATH | Break | A0401 | No Consequences Or Impact To Patient | Event Description: THE ETHICON ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER WAS USED FOR A LAPAROSCOPIC APPENDECTOMY. A BLUE LOAD WAS PLACED ON THE STAPLER END AND THE STAPLER WAS INSERTED INTO THE PATIENT THROUGH A TROCAR. UPON SQUEEZING THE HANDLE OF THE STAPLER, A PIECE OF THE HANDLE BROKE OFF, RENDERING THE STAPLER UNUSABLE. THE DOCTOR WAS ABLE TO EXTRACT THE STAPLER, WITH THE BLUE LOAD STILL INTACT, FROM THE PATIENT WITHOUT INCIDENT. AN AUTOMATIC STAPLER WAS THEN REQUESTED AND USED THROUGHOUT THE REMAINDER OF THE CASE.
14-12-2019 | Injury | ethicon endo-surgery | 10-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DDRA) FOR PATIENTS UNDERGOING UTILIZATION AND COMPARATIVE ANALYSIS OF ENDOPATH® ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: 34 PATIENTS HAD BLEEDING PERI-OPERATIVE COMPLICATION AFTER CHOLECYSTECTOMY.
14-12-2019 | Injury | ethicon endo-surgery | 10-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UTILIZATION AND COMPARATIVE ANALYSIS OF ENDOPATH® ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: 53 PATIENTS HAD BLEEDING PERI-OPERATIVE COMPLICATION AFTER THORACIC SURGERY.
14-12-2019 | Injury | ethicon endo-surgery | 10-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UTILIZATION AND COMPARATIVE ANALYSIS OF ENDOPATH® ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: 403 PATIENTS HAD BLEEDING PERI-OPERATIVE COMPLICATION AFTER UROLOGICAL SURGERY.
14-12-2019 | Injury | ethicon endo-surgery | 10-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UTILIZATION AND COMPARATIVE ANALYSIS OF ENDOPATH® ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: 30 PATIENTS HAD SEPSIS PERI-OPERATIVE COMPLICATION AFTER CHOLECYSTECTOMY.
14-12-2019 | Injury | ethicon endo-surgery | 10-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UTILIZATION AND COMPARATIVE ANALYSIS OF ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: 22 PATIENTS HAD SURGICAL SITE INFECTION PERI-OPERATIVE COMPLICATION AFTER UROLOGICAL SURGERY.
16-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE, THE STAPLER (T40L1X) CUT "THE BLOW" (THE BOWEL) BUT THE "STAPLE NOT COMPLETED SEW UP." THE SURGERY WAS DELAYED BY 40-60 MINUTES AND THEY USED SUTURE TO REPAIR THE BOWEL. THE SURGERY WAS SUCCESSFULLY COMPLETED AND THERE WAS NO PATIENT CONSEQUENCE.
14-12-2019 | Injury | ethicon endo-surgery | 10-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Gas/Air Leak | A050402 | No Code Available | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UTILIZATION AND COMPARATIVE ANALYSIS OF ENDOPATH® ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: 193 PATIENTS HAD AN AIR LEAK PERI-OPERATIVE COMPLICATION AFTER THORACIC SURGERY.
14-12-2019 | Injury | ethicon endo-surgery | 10-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UTILIZATION AND COMPARATIVE ANALYSIS OF ENDOPATH® ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: 65 PATIENTS HAD PNEUMONIA PERI-OPERATIVE COMPLICATION AFTER THORACIC SURGERY.
14-12-2019 | Injury | ethicon endo-surgery | 10-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UTILIZATION AND COMPARATIVE ANALYSIS OF ENDOPATH® ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: 27 PATIENTS HAD PLEURAL EFFUSION PERI-OPERATIVE COMPLICATION AFTER THORACIC SURGERY.
14-12-2019 | Injury | ethicon endo-surgery | 10-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Urinary Tract Infection | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UTILIZATION AND COMPARATIVE ANALYSIS OF ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: 25 PATIENTS HAD A URINARY TRACT INFECTION PERI-OPERATIVE COMPLICATION AFTER THORACIC SURGERY.
13-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE BLUE RELOAD HAD A THIRD ROW OF STAPLES NOT FORM AND CAUSED SOME EXTRA BLEEDING, WHICH WAS CONTROLLED WITH CLIPS. SURGERY WAS NOT DELAYED DUE TO THE REPORTED EVENT AND WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
13-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE BLUE RELOAD HAD A THIRD ROW OF STAPLES NOT FORM AND CAUSED SOME EXTRA BLEEDING, WHICH WAS CONTROLLED WITH CLIPS. SURGERY WAS NOT DELAYED DUE TO THE REPORTED EVENT AND WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
13-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | BLUE RELOAD FOR ECHELON 60 | Activation, Positioning or Separation Problem | A15 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE BLUE RELOAD HAD A THIRD ROW OF STAPLES NOT FORM AND CAUSED SOME EXTRA BLEEDING, WHICH WAS CONTROLLED WITH CLIPS. SURGERY WAS NOT DELAYED DUE TO THE REPORTED EVENT AND WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
10-12-2019 | Injury | intuitive surgical | 10-01-2020 | ENDOWRIST | Insufficient Information | E2401 | Blood Loss | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED HEMICOLECTOMY SURGICAL PROCEDURE, THE PATIENT EXPERIENCED POST-OPERATIVE BLEEDING. AS A RESULT, THERE WAS A BLOOD TRANSFUSION RENDERED TO THE PATIENT. ON (B)(4) 2020, INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: IT WAS REPORTED TO HIM BY THE SURGEON, THAT THE PATIENT ALLEGEDLY EXPERIENCED POST-OP BLEEDING, AND HE BELIEVES THE PATIENT MIGHT HAVE HAD A BLOOD TRANSFUSION. IT WAS CONFIRMED THERE WAS NO SURGICAL INTERVENTION, AND THE SURGEON DID NOT ALLEGE ANY MALFUNCTION OF A DA VINCI SYSTEM, INSTRUMENT, OR ACCESSORY OCCURRED.
14-11-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Device Fell | A051203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION, WHEN THEY OPENED THE DEVICE THE BATTERY FELL ONTO THE FLOOR. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
04-10-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAP COLON PROCEDURE, IN AN ATTEMPT TO FULLY EXTEND THE TROCAR ON THE CDH29P, THE SURGEON TURNED THE KNOB COUNTER-CLOCKWISE TO IT¿S STOP AND ATTEMPTED TO TURN IT FURTHER, BUT NOT WITH EXCESSIVE FORCE. ONCE THE ANVIL WAS CONNECTED, THE KNOB ALLOWED ABOUT 1/4 TO 1/2 ROTATION CLOCKWISE AND STOPPED. A LARGE AMOUNT OF FORCE WAS REQUIRED TO OVERCOME THE STOP, AND THEN THE DEVICE CLOSED NORMALLY. IT SEEMED AS THOUGH THE GEARING OR THREADING INSIDE HAD JUMPED A TOOTH OR OTHERWISE BOUND. THE SURGEON OPTED TO CONTINUE WITH THE FIRING, AND ALL ELSE APPEARED TO WORK APPROPRIATELY. GOOD DONUTS, NEGATIVE LEAK TEST. THE SURGERY WAS NOT DELAYED DUE TO THE REPORTED EVENT. NO FRAGMENTS WERE GENERATED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. NO OTHER MEDICAL INTERVENTION WAS REQUIRED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
20-11-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: SURGEON DID ROBOT SURGERY PROCEDURE FOR GASTRECTOMY. SHE FIRED FOR CUT JEJUNUM SITE TISSUE WITH THIS PRODUCT. AFTER FIRED SHE SAW TARGET TISSUE WERE NOT CUT & SEAL CLEARLY BUT, TISSUE WERE NOT CUT AND STAPLES WERE NOT DEPLOYED ON TISSUE BUT OPENED ITS STAPLE LEGS. SURGEON USED ONE ANOTHER BODY AND CARTRIDGE FOR NEW FIRE. THERE WERE NO PATIENT CONSEQUENCES.
20-11-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: SURGEON DID ROBOT SURGERY PROCEDURE FOR GASTRECTOMY. SHE FIRED FOR CUT JEJUNUM SITE TISSUE WITH THIS PRODUCT. AFTER FIRED SHE SAW TARGET TISSUE WERE NOT CUT & SEAL CLEARLY BUT, TISSUE WERE NOT CUT AND STAPLES WERE NOT DEPLOYED ON TISSUE BUT OPENED ITS STAPLE LEGS. SURGEON USED ONE ANOTHER BODY AND CARTRIDGE FOR NEW FIRE. THERE WERE NO PATIENT CONSEQUENCES.
18-11-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE PSEE45A JAMMED DURING THE PROCEDURE. NO FURTHER INFORMATION WAS PROVIDED. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-10-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | FLEX 60 ARTICULATING | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE SURGEON USED AN EC60A IN LAPAROSCOPIC LAR, THE DEVICE ARTICULATION SITE WAS BROKEN, AND THE SPRING COIL EXPOSED. SO, THEY CHANGE THE NEW EC60A. PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
16-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THERE WAS NO FEEDBACK OF BREAKING THE WASHER AT THE FIRING, AND THE STAPLES WERE NOT FORMED PROPERLY. THE DEVICE WAS USED ON THE COLON. THE DEVICE COULD NOT BE REMOVED FROM THE PATENT, SO THAT THE ADJUSTING KNOB WAS FULLY ROTATED TO OPEN DIRECTION TO REMOVE THE DEVICE. THEN, IT WAS FOUND THAT THE TARGET TISSUE WAS NOT ANASTOMOSED AND IT WAS NOT STAPLED. THEREFORE, THE PLACED ANVIL WAS ALSO REMOVED FROM THE PATIENT. THE SOUND WAS HEARD THAT THE ANVIL AND THE HOUSING WERE ATTACHED PROPERLY BEFORE THE FIRING. THE GAP SETTING SCALE MARKED 1.5MM AT THE FIRING. THE TARGET TISSUE WAS FATTY AND THICK. THE SURGEON WHO FIRED THE DEVICE JUST RETURNED FROM HIS HAND INJURY AND ITS TREATMENT, SO THAT THERE WAS A POSSIBILITY THAT HIS GRIP STRENGTH WAS WEAK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. ADDITIONAL INFORMATION WAS PROVIDED THAT THE SURGEON REMOVED THE ANVIL AND ANASTOMOSED WITH THE REPLACEMENT WITHOUT CHANGING THE PROCEDURE. THERE WAS NO BLEEDING AND LEAKAGE FOR THE PATIENT. THE PATIENT IS STABLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
16-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | ILS 25MM, CURVED | Activation Problem | A1501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN THE COURSE OF A ESOPHAGOJEJUNOSTOMY PROCEDURE USING CDH25A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE DID NOT PROCEED AND STAPLES ALSO WEREN'T FIRED. AS AN EMERGENCY MEASURE, THE SURGEON DETACHED IT FROM THE TISSUE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES FOR THE PATIENT SO FAR.
17-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | ILS 25MM, CURVED | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: IN THE COURSE OF A ANTERIOR RESECTION PROCEDURE USING CDH25A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE DID NOT PROCEED AND STAPLES ALSO WEREN'T FIRED. AS AN EMERGENCY MEASURE, THE SURGEON DETACHED IT FROM THE TISSUE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES FOR THE PATIENT SO FAR.
17-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IN THE COURSE OF A ANTERIOR RESECTION PROCEDURE USING CDH25A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE DID NOT PROCEED AND STAPLES ALSO WEREN'T FIRED. AS AN EMERGENCY MEASURE, THE SURGEON DETACHED IT FROM THE TISSUE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES FOR THE PATIENT SO FAR.
13-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | GREEN RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC PNEUMONECTOMY SURGERY, WHEN SEVERING LUNG TISSUE , AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, WHEN THE DEVICE WAS FIRED, THE STAPLES DID NOT FORM PROPERLY. IT IS UNKNOWN HOW PROCEDURE WAS COMPLETED AND PATIENT CONSEQUENCE WAS NOT REPORTED.
17-12-2019 | Malfunction | ethicon endo-surgery | 10-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN HEPATECTOMY, AT THE FIRST FIRING, THE KNIFE MOVED FORWARD ABOUT 18MM AND RETURNED TO HOME POSITION ON THE WAY. AT THE 2ND FIRING, THE KNIFE MOVED FORWARD ABOUT 20MM AND RETURNED TO THE HOME POSITION ON THE WAY. THE DEVICE WAS USED ON THE GLISSON CAPSULE OF THE LEFT SEGMENT. THE SURGEON DID NOT ATTEMPT THE RELEASING DEVICE PROCESS SUCH USING THE KNIFE REVERSE SWITCH WHEN THE EVENT OCCURRED. NO UNUSUAL SOUND WAS HEARD DURING USE. THE SURGEON FIRED THE DEVICE DIVIDING SEVERAL STROKES. ANOTHER DEVICE LOADING ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
06-01-2020 | Malfunction | covidien lp llc north haven | 10-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO DETACH THE FISSURES DURING LAPAROSCOPIC PULMONARY LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. A NEW RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-11-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, BLEEDING OCCURRED FROM THE CUT LINE AFTER THE DEVICE WAS USED ON THE LEFT ATRIAL APPENDAGE, AND THE ARTIFICIAL HEART LUNG APPARATUS WAS STOPPED RESTORE THE PRESSURE FROM LEFT ATRIUM. THE SURGEON TRIED TO FIRE THE SAME DEVICE AGAIN, BUT THE CARDIAC RUPTURE OCCURRED. THEN, THE ARTIFICIAL HEART LUNG APPARATUS WAS REACTIVATED TO TREAT THE CARDIAC RUPTURE, AND ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE AMOUNT OF THE BLEEDING WAS UNKNOWN. NO BLOOD TRANSFUSION WAS REQUIRED. THE PATIENT IS IN THE HOSPITAL, AND THE PROGRESS OF THE PATIENT IS AS USUAL. THE SURGEON AND A NURSE HAD EXPERIENCE IN USING ECHELON DEVICES.
01-01-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DUE TO THE RELOADING UNIT OF THE LONG60A CODED STAPLER WAS CURVED, RELOAD COULDN'T LOCATE ON THE UNIT. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
03-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER, WITH UNKNOWN RELOAD, NO BUTTRESSING MATERIAL, WOULDN'T FIRE ON THE FIRST FIRING. ANOTHER LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
03-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER, WITH UNKNOWN RELOAD, NO BUTTRESSING MATERIAL, WOULDN'T FIRE ON THE FIRST FIRING. ANOTHER LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
17-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | UNKNOWN CONTOUR CURVED CUTTER STAPLER | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AT THE MOMENT THAT THE DEVICE WAS USED, THE DEVICE STAPLED BUT DIDN'T CUT CAUSING THE NEED FOR IMMEDIATE EXCHANGE OF MATERIAL. MATERIAL WAS CHANGED SOON AFTER FAILURE DETECTED. THERE WERE NO PATIENT CONSEQUENCES.
01-01-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE MALFORMED STAPLES FROM GST60B LOAD WERE RECOVERED FROM RELOAD. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
27-11-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | 55 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: STAPLE RETENTION ISSUE & WOULD NOT FIRE. IT WAS REPORTED THAT DURING THE RIGHT HEMICOLECTOMY SURGERY, AFTER INSTALLED THE CARTRIDGE, REMOVED THE RETAINING CAP, NOTED THE STAPLES PROTRUDING. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE WHEN SEVERING THE RECTUM TISSUE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | ILS 29MM, CURVED | Activation, Positioning or Separation Problem | A15 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN HE FIRED THE DEVICE, HE SQUEEZED PLASTIC TO PLASTIC BUT DIDN¿T GET A CRUNCH TO INDICATE THE WASHER WAS BROKEN. THANKFULLY THEY HAD ENOUGH BOWEL LEFT TO TAKE IT DOWN AND TRY AGAIN. THIS ADDED ABOUT AN HOUR ONTO THE PROCEDURE. PATIENT HAS NOW BEEN DISCHARGED AND IS HOME.
14-11-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | CONTOUR CURVED CUTTER STAPLER | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A RECTUM AMPUTATION, THE RELEASE BUTTON DID NOT WORK. NO KNOWN COMPLICATIONS OF THE PATIENT AT THE END OF SURGERY. THE INTESTINE WAS SUTURED WITH A CLASSIC TECHNIQUE. IT WAS PERFORMED BY GYNECOLOGY SURGEONS. SURGERY WAS DONE (B)(6) 2019.
17-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | PROXIMATE*RL TITAN HEAVY WIRE | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, THE STAPLES WERE NOT DEPLOYED AT THE FIRING ON THE BRONCHIAL TUBE. IT WAS FOUND THAT THE STAPLES WERE SCATTERED ON THE MAYO TABLE. THERE WAS NO BLEEDING AND LEAKAGE FOR THE PATIENT. THE SURGEON SUTURED THE STAPLE LINE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THE PATIENT IS STABLE. NO FURTHER INFORMATION IS AVAILABLE.
17-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY, THE DOCTOR WAS COMPRESSING THE TISSUE TO COMPLETE THE ANASTOMOSIS WITH THE POWERED CIRCULAR STAPLER AND THE ANVIL TIP DETACHED FROM THE TROCAR. THE DOCTOR CONTINUED TO RETRACT THE TROCAR IN THE DEVICE, NOT KNOWING IT WASN'T ATTACHED. AS A RESULT, THE DOCTOR HAD TO REATTACH THE ANVIL HEAD TO THE TROCAR, COMPRESS THE TISSUE A SECOND TIME, TO PERFORM THE ANASTOMOSIS. THE DOCTOR WAS ABLE TO PERFORM THE STAPLING PROCEDURE. THE ANASTOMOSIS PASSED THE LEAK TEST. DOCTOR INDICATED THIS DIDN'T HAPPEN WITH THE MANUAL STAPLER.
18-11-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THE METAL PLATE BENEATH THE RELOAD CARTRIDGE DISLODGED POST FIRE. THIS WAS IN THE CHANGING FIELD PRIOR TO LOADING OF NEW RELOAD. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED AND THERE WERE NO PATIENT CONSEQUENCES REPORTED.
17-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGEAL ANASTOMOSIS, THE DEVICE DID NOT WORK. THE STAPLES WERE NOT IN PLACE. USE OF ANOTHER DEVICE. THERE WERE NO PATIENT CONSEQUENCES.
23-09-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: AT THE TIME OF PERFORMING THE COLORECTAL ANASTOMOSIS, THE STAPLER IS INSERTED TO FIT IT WITH THE ANVIL AND GRADUALLY CLOSED TO TRIGGER IT. AT THE MOMENT THE STAPLER IS REMOVED FROM THE RECTUM, THE COLON IS PERFORATED SINCE IT ONLY CUTS BUT DOES NOT STAPLE. THE STAPLES ARE COMPLETELY OPEN. THE STAPLER IS REMOVED, AND THE SURGERY IS FINISHED MANUALLY. CAUSING THE PATIENT TO BE EXPOSED TO MORE HOURS OF ANESTHESIA.
09-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | ILS 29MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: DIFFICULT TO FIRE. IT WAS REPORTED THAT DURING THE SURGERY OF RIGHT HEMICOLECTOMY, WHEN SEVERING THE ILEUM AND COLON, FELT IT WAS DIFFICULT TO SQUEEZE DOWN THE FIRING TRIGGER, AT LAST, FIRING FINISHED WITH BIG FORCE, THE WASHER WAS CUT OFF, THE DONUTS WERE COMPLETE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
08-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING CONFIGURATION OF ILEOSIGMOID ANASTOMOSIS, THE CLOSURE OF THE MECHANICAL SUTURE (STAPLE LINE) FAILED. A NEW DEVICE IS USED. THE SURGERY WAS NOT DELAYED. PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
15-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | PROXIMATE*75MM LINEAR CUTTER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED, PATIENT WHO ENTERS FOR SCHEDULED EXPLORATORY LAPAROTOMY COLOSTOMY CLOSURE, THE ETHICON ENDO-SURGERY LINEAR SUTURE GUN WAS PASSED, AT THE TIME THE REMODELING PROCEDURE WAS FINISHED, THE DEVICE WAS REMOVED AND THERE WAS A BAD CUT OF THE SAME AND POOR FIT OF THE SUTURES. THE HOSPITAL DIDN'T REPORT PATIENT CONSEQUENCES SO THERE WASN'T ANY NEGATIVE OUTCOME. AFTER THE DEVICE FAILURE, THEY USED ANOTHER DEVICE TO CONCLUDE THE PROCEDURE.
18-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | PROXIMATE*RL LINEAR STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, THE DOCTOR REPORTED THAT AFTER FIRING THE STAPLER, SOME OF THE STAPLES WERE UNFORMED, FAILED TO MAKE A B SHAPE. A SECOND LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
09-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | POWERED PLUS STANDARD 45 | Difficult to Open or Remove Packaging Material | A020501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STERILE PACKAGING WAS DAMAGED UPON OPENING. DEVICE WAS NOT USED. NEW DEVICE WAS OPENED.
14-11-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE LUNG NODULE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | PROXIMATE*PPH PROCEDURE SET | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, PPH03 WAS CUT BUT THERE WERE NOT LOADED ANY CLIPS. FORTUNATELY, THERE WAS ANOTHER STAPLER TO USE IN THE HOSPITAL. SO, THE SURGERY WAS DONE SUCCESSFULLY BUT HOSPITAL WAS COMPLAINED TO US.
12-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE DISTAL GASTRIC CANCER OPERATION SURGERY, WHEN SEVERING DUODENAL TISSUE, AFTER FIRED, NOTED THE DISTAL STAPLES MALFORMED WITH IRREGULAR SHAPE, AND NOTED OOZING. USED SUTURE TO OVERSEW AND STOP OOZING. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | GREEN RELOAD FOR ECHELON 60 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RECTAL CANCER SURGERY, AFTER THE DEVICE WAS FIRED, IT WAS INSERTED INTO TROCAR, BUT COULD NOT OPEN THE JAW. TOOK OUT THE GUN, STILL COULD NOT OPEN JAW, AND THE KNIFE REVERSE BUTTON WAS USED TO OPEN THE JAW. THE SURFACE OF RELOADS WERE NORMAL, THEN CHANGE TO A NEW ECR60G. AFTER IT WAS CHANGED TO A NEW ECR60G, IT WAS NOTED THAT THE STAPLES WERE MALFORMED. CHANGED TO A NEW ONE AGAIN TO COMPLETE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT IN A SLEEVE SURGERY WHERE THE GREEN RELOADS STAPLE LINE WAS APPLIED, ONE STAPLE COME OUT OF THE LINE. NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE.
19-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON INFORMED US THAT THE STAPLES WERE MALFORMED. SHE ALSO EXPERIENCED ANOTHER TACTILE FEELING IN THE HANDLE COMPARED TO THE USUAL FEELING. THE STAPLE LINE HAD BEEN CHECKED AND WERE NEEDED ADDITIONALLY STITCHING WAS PERFORMED, HENCE NO PATIENT CONSEQUENCE.
13-01-2020 | Death | us surgical puerto rico | 13-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Death | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN JANUARY 2015 AND MARCH 2016, THIS WAS A RANDOMIZED-CONTROLLED STUDY FOR THE EFFECT OF VITAMIN E OINTMENT ON INCISIONAL SURGICAL SITE INFECTION IN 101 PATIENTS TREATED WITH ELECTIVE COLORECTAL LAPAROSCOPY. THE SURGICAL PROCEDURES INCLUDED LEFT COLECTOMY ON 33 PATIENTS, RIGHT COLECTOMY ON 28 PATIENTS, TRANSVERSE COLECTOMY ON 3 PATIENTS, LOW ANTERIOR RESECTION ON 33 PATIENTS AND TOTAL COLECTOMY ON 4 PATIENTS. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE SURGICAL TECHNIQUES BETWEEN GROUPS. THE ANASTOMOSES WERE STAPLED EITHER WITH A CIRCULAR STAPLER IN THE LEFT COLON AND RECTUM OF 66 PATIENTS OR A LINEAR STAPLER IN THE TRANSVERSE AND ILEOCOLIC ANASTOMOSES ON 35 PATIENTS. THERE WERE 3 CASES OF ANASTOMOTIC LEAKAGE WHICH ALL APPEARED IN COLORECTAL ANASTOMOSES, HENCE ALL DUE TO THE CIRCULAR STAPLER. THE STUDY MORTALITY RATE WAS 1%, WITH ONE PATIENT DEATH 30 DAYS AFTER OPERATION. THE CAUSE OF PERI-OPERATIVE DEATH WAS AN ANASTOMOTIC LEAK AND SEPSIS.
16-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY, BUT THEN TURNED INTO A BILOBECTOMY; USED THE PVE35A WITH A VASCULAR RELOAD. ON THE PA, DEVICE FIRED AND CUT BUT DID NOT STAPLE PROPERLY AND HAD TO BE RECTIFIED WITH A CLIP. NO RESISTANCE FELT OR AND STRANGE NOISES. IT WAS ALMOST LIKE SOME OF THE STAPLES WERE NOT FORMED PROPERLY. VESSEL WAS NICE AND FLAT AND NO TENSION WHEN FIRING.
19-12-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | PWRD ECH FLEX 45MM 34CM SHAFT | Material Protrusion/Extrusion | A0411 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT UPON STARTING TO FIRE THE STAPLER, THE RELOAD WAS EXTRUDED OUT THE DISTAL END OF STAPLER. THE PROCEDURE WAS DELAYED FOR 10 MINUTES, BUT A NEW STAPLER AND RELOAD WAS USED TO SUCCESSFULLY COMPLETE THE PROCEDURE. NO FRAGMENTS WERE GENERATED AND NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL MEDICAL INTERVENTION WAS REQUIRED.
01-01-2019 | Injury | ethicon endo-surgery | 13-01-2020 | POWERED PLUS STANDARD 45 | Sharp Edges | A051102 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A VATS PROCEDURE, SURGEON WAS ACCESSING LUNG TISSUE WITH PSEE45A DEVICE WITH GST RELOAD. THE SURGEON ADVANCED THE STAPLER, WHICH CREATED A TEAR IN THE PULMONARY ARTERY. THE SURGEON COMMENTED THAT THE DISTAL END OF THE RELOAD WAS SHARP. THE PROCEDURE CONVERTED TO AN OPEN PROCEDURE. PATIENT IS OK.
01-01-2019 | Injury | ethicon endo-surgery | 13-01-2020 | POWERED PLUS STANDARD 45 | Sharp Edges | A051102 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A VATS PROCEDURE, SURGEON WAS ACCESSING LUNG TISSUE WITH PSEE45A DEVICE WITH GST RELOAD. THE SURGEON ADVANCED THE STAPLER, WHICH CREATED A TEAR IN THE PULMONARY ARTERY. THE SURGEON COMMENTED THAT THE DISTAL END OF THE RELOAD WAS SHARP. THE PROCEDURE CONVERTED TO AN OPEN PROCEDURE. PATIENT IS OK.
09-10-2019 | Malfunction | ethicon endo-surgery | 13-01-2020 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, OT SCRUB NURSE TRIED TO LOAD THE GST60D RELOAD IN PLEE60A, WHEN FOUND CARTRIDGE BASE METAL SHEET IS OUT OF ITS POSITION- KEPT ASIDE AND LOADED A NEW ONE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
18-12-2019 | Injury | ethicon endo-surgery | 13-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Firing Problem | A0505 | No Code Available | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A HYBRID-VATS LEFT UPPER LOBECTOMY, THE MOTOR SOUND WAS HEARD FOR A MOMENT AND THE DEVICE STOPPED AT THE 2ND FIRING ON THE PULMONARY ARTERY. THE DEVICE DID NOT MOVE FURTHER THOUGH THE SURGEON PULLED THE FIRING TRIGGER. JUST IN CASE BLEEDING OCCURRED, THE PROCEDURE WAS CONVERTED FROM A SEMI-OPEN PROCEDURE TO AN OPEN PROCEDURE . THEREFORE, THE RANGE OF THE DISSECTION WAS EXPANDED (ITS INCISION SIZE WAS LONGER 1.5 TIMES THAN ORIGINALLY). AFTER THAT, THE KNIFE REVERSE SWITCH WAS USED TO RETURN THE KNIFE TO HOME POSITION. IT WAS FOUND THAT THE TARGET TISSUE WAS NOT CUT, AND NO STAPLES WERE DEPLOYED. NO BLEEDING OCCURRED. ANOTHER PVE35A LOADING ANOTHER VASECR35 WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ONE PVE35A AND ONE VASECR35 WILL BE RETURNING. NO FURTHER INFORMATION IS AVAILABLE. AFFILIATE REFERENCE NUMBER: (B)(4). PLEASE TAKE PHOTOS FOR (B)(6) CUSTOMER.
18-12-2019 | Injury | ethicon endo-surgery | 13-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Activation Problem | A1501 | No Code Available | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A HYBRID-VATS LEFT UPPER LOBECTOMY, THE MOTOR SOUND WAS HEARD FOR A MOMENT AND THE DEVICE STOPPED AT THE 2ND FIRING ON THE PULMONARY ARTERY. THE DEVICE DID NOT MOVE FURTHER THOUGH THE SURGEON PULLED THE FIRING TRIGGER. JUST IN CASE BLEEDING OCCURRED, THE PROCEDURE WAS CONVERTED FROM A SEMI-OPEN PROCEDURE TO AN OPEN PROCEDURE . THEREFORE, THE RANGE OF THE DISSECTION WAS EXPANDED (ITS INCISION SIZE WAS LONGER 1.5 TIMES THAN ORIGINALLY). AFTER THAT, THE KNIFE REVERSE SWITCH WAS USED TO RETURN THE KNIFE TO HOME POSITION. IT WAS FOUND THAT THE TARGET TISSUE WAS NOT CUT, AND NO STAPLES WERE DEPLOYED. NO BLEEDING OCCURRED. ANOTHER PVE35A LOADING ANOTHER VASECR35 WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ONE PVE35A AND ONE VASECR35 WILL BE RETURNING. NO FURTHER INFORMATION IS AVAILABLE. AFFILIATE REFERENCE NUMBER: (B)(4). PLEASE TAKE PHOTOS FOR (B)(6) CUSTOMER.
24-12-2019 | Malfunction | us surgical puerto rico | 13-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A CHOLECYSTECTOMY PROCEDURE, WHILE EXTRACTION WAS BEING DONE TO THE GALLBLADDER, IT WAS NOTED THAT THE GALLBLADDER WAS FIBROTIC, AND THE DEVICE WAS NOT ABLE TO FIRE. IT WAS NOTED THAT THE DEVICE ALSO LOCKED ON TISSUE, BUT EVENTUALLY WAS REMOVED BY MANUAL EXTRACTION. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | us surgical puerto rico | 13-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A CHOLECYSTECTOMY PROCEDURE, WHILE EXTRACTION WAS BEING DONE TO THE GALLBLADDER, IT WAS NOTED THAT THE GALLBLADDER WAS FIBROTIC, AND THE DEVICE WAS NOT ABLE TO FIRE. IT WAS NOTED THAT THE DEVICE ALSO LOCKED ON TISSUE, BUT EVENTUALLY WAS REMOVED BY MANUAL EXTRACTION. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 13-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A CHOLECYSTECTOMY PROCEDURE, WHILE EXTRACTION WAS BEING DONE TO THE GALLBLADDER, IT WAS NOTED THAT THE GALLBLADDER WAS FIBROTIC, AND THE DEVICE WAS NOT ABLE TO FIRE. IT WAS NOTED THAT THE DEVICE ALSO LOCKED ON TISSUE, BUT EVENTUALLY WAS REMOVED BY MANUAL EXTRACTION. THERE WAS NO PATIENT INJURY.
24-12-2019 | Malfunction | covidien lp llc north haven | 13-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A CHOLECYSTECTOMY PROCEDURE, WHILE EXTRACTION WAS BEING DONE TO THE GALLBLADDER, IT WAS NOTED THAT THE GALLBLADDER WAS FIBROTIC, AND THE DEVICE WAS NOT ABLE TO FIRE. IT WAS NOTED THAT THE DEVICE ALSO LOCKED ON TISSUE, BUT EVENTUALLY WAS REMOVED BY MANUAL EXTRACTION. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | us surgical puerto rico | 14-01-2020 | ENDO GIA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ENDOSCOPIC RECTAL SURGERY, AFTER FIRING, IT WAS DIFFICULT TO REMOVE THE DEVICE FROM THE TROCAR. AFTER REMOVING THE DEVICE, IT WAS FOUND THAT THE ARTICULATION JOINT WAS DEFORMED WHICH CAUSED THE DIFFICULTY IN REMOVING THE DEVICE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | us surgical puerto rico | 14-01-2020 | ENDO GIA | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN ENDOSCOPIC RECTAL SURGERY, AFTER FIRING, IT WAS DIFFICULT TO REMOVE THE DEVICE FROM THE TROCAR. AFTER REMOVING THE DEVICE, IT WAS FOUND THAT THE ARTICULATION JOINT WAS DEFORMED WHICH CAUSED THE DIFFICULTY IN REMOVING THE DEVICE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | covidien lp llc north haven | 14-01-2020 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE RELOAD WAS INSTALLED, THE JAW COULD NOT CLOSE. THERE WAS NO PATIENT INVOLVEMENT.
31-12-2019 | Malfunction | covidien lp llc north haven | 14-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE RELOAD WAS INSTALLED, THE JAW COULD NOT CLOSE. THERE WAS NO PATIENT INVOLVEMENT.
31-12-2019 | Malfunction | us surgical puerto rico | 14-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE RELOAD WAS INSTALLED, THE JAW COULD NOT CLOSE. THERE WAS NO PATIENT INVOLVEMENT.
16-12-2019 | Malfunction | us surgical puerto rico | 14-01-2020 | TA | Mechanics Altered | A0507 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY PROCEDURE, WHILE CLOSING THE ENTEROTOMY, THE STAPLES DID NOT DEPLOY. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | us surgical puerto rico | 14-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN INSTALLING THE RELOAD, THE JAWS WOULD NOT CLOSE. THERE WAS NO PATIENT INVOLVEMENT.
31-12-2019 | Malfunction | covidien lp llc north haven | 14-01-2020 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN INSTALLING THE RELOAD, THE JAWS WOULD NOT CLOSE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2019 | Injury | ethicon endo-surgery | 14-01-2020 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT THE PATIENT'S INDICATION IS COLITIC POLYP, RECTAL CANCER, DURING THE LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER SURGERY, NO ANY ABNORMAL SITUATION WAS OBSERVED. THE OPERATION PROCESS WAS SMOOTH, THE ANASTOMOTIC STOMA WAS WELL EXAMINED DURING THE OPERATION, AND THERE WAS NO BLEEDING UNDER COLONOSCOPY. THE PATIENT HAD HEMATOCHEZIA 4 HOURS AFTER OPERATION, WHICH WAS TEMPORARILY STABILIZED BY CONSERVATIVE TREATMENT WITH DRUGS, AND THE BLEEDING WAS REDUCED, AND THE BLOOD COLOR WAS DARK RED ON THE THIRD DAY. THE PATIENT IS STABLE AND IN HOSPITAL NOW.
16-12-2019 | Injury | ethicon endo-surgery | 14-01-2020 | ILS 25MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT THE PATIENT'S INDICATION IS MALIGNANT NEOPLASM OF ASCENDING COLON, DURING THE LAPAROSCOPIC RADICAL RESECTION OF RIGHT COLON CANCER SURGERY, NO ANY ABNORMAL SITUATION WAS OBSERVED. THE OPERATION PROCESS WAS SMOOTH, THE ANASTOMOTIC STOMA WAS WELL EXAMINED DURING THE OPERATION, AND THERE WAS NO BLEEDING UNDER COLONOSCOPY. THE PATIENT HAD HEMATOCHEZIA 4 HOURS AFTER OPERATION, WHICH WAS TEMPORARILY STABILIZED BY CONSERVATIVE TREATMENT WITH DRUGS, AND THE BLEEDING WAS REDUCED, AND THE BLOOD COLOR WAS DARK RED ON THE THIRD DAY. THE PATIENT IS STABLE AND IN HOSPITAL NOW.
19-12-2019 | Malfunction | davol inc sub c.r bard | 14-01-2020 | PERMAFIX FIXATION SYSTEM 15 COUNT | Incorrect, Inadequate or Imprecise Result or Readings | A0908 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHILE FIXATING A BARD/DAVOL SOFT MESH DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE FASTENERS IN THE PERMAFIX FIXATION DEVICE WERE COMING OUT AND WERE WOULD NOT FIXATE THE TISSUE / HOLD THE MESH. THIS WAS TRIED WITH 5-6 FASTENERS AND OCCURRED REPEATEDLY. A NEW FIXATION DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY AS A RESULT OF THE PROBLEM EXPERIENCED.
19-12-2019 | Malfunction | davol inc sub c.r bard | 14-01-2020 | PERMAFIX FIXATION SYSTEM 15 COUNT | Insufficient Information | E2401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHILE FIXATING A BARD/DAVOL SOFT MESH DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE FASTENERS IN THE PERMAFIX FIXATION DEVICE WERE COMING OUT AND WERE WOULD NOT FIXATE THE TISSUE / HOLD THE MESH. THIS WAS TRIED WITH 5-6 FASTENERS AND OCCURRED REPEATEDLY. A NEW FIXATION DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY AS A RESULT OF THE PROBLEM EXPERIENCED.
14-01-2020 | Malfunction | us surgical puerto rico | 14-01-2020 | GIA | Misfire | A050502 | No Information | Event Description: ACCORDING TO THE REPORTER, DURING THE SECOND FIRING ON AN OPEN PROCEDURE, THE DEVICE FIRED ABOUT 1/3 OF THE WAY, IT FELT JAMMED AND WOULD NO LONGER ADVANCE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
19-12-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLE LINE DID NOT FORM A B SHAPE. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
11-11-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | PROXIMATE*RL TITAN HEAVY WIRE | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LUNG CANCER SURGERY, WHEN SEVERING THE BRONCHUS, IT WAS NOTED THAT THE STAPLES WERE MALFORMED.
31-12-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | ECHELON FLEX | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: DOCTOR WAS USING ECHELON FLEX 35 MM WHITE VASCULAR STAPLER. STAPLER WOULD NOT FIRE.
06-12-2019 | Malfunction | justright surgical | 15-01-2020 | JUSTRIGHT5MM RELOAD | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: DOCTOR WAS DOING A RIGHT UPPER LOBE WEDGE RESECTION. DOCTOR SUCCESSFULLY USED THE JUST RIGHT 5MM STAPLER TWICE. THE ON THE THIRD LOAD DOCTOR NOTED THAT THE STAPLER DID NOT SEEM TO OPEN AS WIDE AS THE PREVIOUS TIMES AND HE ASKED THE REP ABOUT IT. REP SAID THAT IT MIGHT OPEN WIDER WHEN TISSUE WAS PLACED IN THE JAWS. DOCTOR PROCEEDED AND FIRED THE STAPLE LOAD. AFTER THE JAWS OF THE LOAD WERE OPENED DOCTOR NOTED THAT THE LOAD ONLY PARTIALLY FIRED AND THAT THERE WAS A NOTED INDENT ON THE LUNG TISSUE THAT DID NOT HAVE STAPLES. LOAD WAS SEQUESTERED AND A NEW LOAD WAS USED. NO PATIENT HARM. DOCTOR WAS RECORDING THE WHOLE PROCEDURE WITH THE NSTEAM SYSTEM. MANUFACTURER RESPONSE FOR JUST RIGHT 5MM RELOAD, (BRAND NOT PROVIDED) (PER SITE REPORTER). WAITING ON THE RETURN OF THE DEVICE.
06-12-2019 | Malfunction | justright surgical | 15-01-2020 | JUSTRIGHT5MM RELOAD | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: DOCTOR WAS DOING A RIGHT UPPER LOBE WEDGE RESECTION. DOCTOR SUCCESSFULLY USED THE JUST RIGHT 5MM STAPLER TWICE. THE ON THE THIRD LOAD DOCTOR NOTED THAT THE STAPLER DID NOT SEEM TO OPEN AS WIDE AS THE PREVIOUS TIMES AND HE ASKED THE REP ABOUT IT. REP SAID THAT IT MIGHT OPEN WIDER WHEN TISSUE WAS PLACED IN THE JAWS. DOCTOR PROCEEDED AND FIRED THE STAPLE LOAD. AFTER THE JAWS OF THE LOAD WERE OPENED DOCTOR NOTED THAT THE LOAD ONLY PARTIALLY FIRED AND THAT THERE WAS A NOTED INDENT ON THE LUNG TISSUE THAT DID NOT HAVE STAPLES. LOAD WAS SEQUESTERED AND A NEW LOAD WAS USED. NO PATIENT HARM. DOCTOR WAS RECORDING THE WHOLE PROCEDURE WITH THE NSTEAM SYSTEM. MANUFACTURER RESPONSE FOR JUST RIGHT 5MM RELOAD, (BRAND NOT PROVIDED) (PER SITE REPORTER). WAITING ON THE RETURN OF THE DEVICE.
01-10-2019 | Malfunction | covidien lp llc north haven | 15-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO GASTROPLASTY PROCEDURE, WHILE ATTEMPTING TO FIRE THE DEVICE, THE DEVICE LOCKED AND DID NOT WORK. IT WAS ALSO NOTED THAT THERE WAS A DIFFICULTY WITH UNLOADING THE DEVICE. A NEW RELOAD WAS USED BUT THE PROBLEM PERSISTED. THE RELOAD WAS REPLACED AGAIN AND IT WORKED NORMALLY. THERE WAS NO PATIENT INJURY.
01-10-2019 | Malfunction | covidien lp llc north haven | 15-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO GASTROPLASTY PROCEDURE, WHILE ATTEMPTING TO FIRE THE DEVICE, THE DEVICE LOCKED AND DID NOT WORK. IT WAS ALSO NOTED THAT THERE WAS A DIFFICULTY WITH UNLOADING THE DEVICE. A NEW RELOAD WAS USED BUT THE PROBLEM PERSISTED. THE RELOAD WAS REPLACED AGAIN AND IT WORKED NORMALLY. THERE WAS NO PATIENT INJURY.
06-12-2019 | Malfunction | covidien lp llc north haven | 15-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY PROCEDURE, THE RELOAD COULD NOT BE CLOSED CORRECTLY. THERE WAS NO PATIENT INJURY.
06-12-2019 | Malfunction | us surgical puerto rico | 15-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY PROCEDURE, THE RELOAD COULD NOT BE CLOSED CORRECTLY. THERE WAS NO PATIENT INJURY.
06-12-2019 | Malfunction | us surgical puerto rico | 15-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY PROCEDURE, THE RELOAD COULD NOT BE CLOSED CORRECTLY. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 15-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVE A THORACOSCOPIC LEFT SUPERIOR TIP PULMONARY BULLAE RESECTION, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, THE STAPLER COULD NOT BE FIRED. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
04-12-2019 | Malfunction | covidien lp llc north haven | 15-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING CLIPPING AT THE BEGINNING SHOT IN A LAPAROSCOPIC VIDEO GASTROPLASTY, THE BEIGE RELOAD FAILED AND DID NOT FIRE. ANOTHER BEIGE RELOAD WAS USED AS REPLACEMENT FOR THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2019 | Malfunction | covidien lp llc north haven | 15-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING CLIPPING AT THE BEGINNING SHOT IN A LAPAROSCOPIC VIDEO GASTROPLASTY, THE BEIGE RELOAD FAILED AND DID NOT FIRE. ANOTHER BEIGE RELOAD WAS USED AS REPLACEMENT FOR THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Malfunction | us surgical puerto rico | 15-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LOBECTOMY, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON HENCE THE DEVICE WAS UNABLE TO FIRE TO CLAMP THE BLOOD VESSELS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Malfunction | us surgical puerto rico | 15-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LOBECTOMY, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON HENCE THE DEVICE WAS UNABLE TO FIRE TO CLAMP THE BLOOD VESSELS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2019 | Injury | us surgical puerto rico | 15-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY PROCEDURE, THE SURGEON FOUND THAT THERE IS NO CORRECT STAPLE FORMATION, AND HAD AN INCOMPLETE STAPLE LINE DISTALLY, THAT LED TO BLOOD LOSS. THE SURGEON THEN HAD TO MANUALLY SUTURED THE TISSUE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
11-12-2019 | Injury | us surgical puerto rico | 15-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY PROCEDURE, THE SURGEON FOUND THAT THERE IS NO CORRECT STAPLE FORMATION, AND HAD AN INCOMPLETE STAPLE LINE DISTALLY, THAT LED TO BLOOD LOSS. THE SURGEON THEN HAD TO MANUALLY SUTURED THE TISSUE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
29-08-2019 | Malfunction | covidien lp llc north haven | 15-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN THORACIC PROCEDURE, THE SURGEON SEEMED TO BE ABLE TO BE SQUEEZED THE HANDLE BUT THE LOAD DID NOT FIRE. THE SURGEON THEN USED ANOTHER DEVICE HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLER BLOCKED DURING TURN OFF. NO FURTHER INFORMATION. NO CONSEQUENCES FOR THE PATIENT. PROCEDURE: UNKNOWN.
19-11-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: DURING SURGERY LOW ANTERIOR RECTUM RESECTION IN SURGEON COULD NOT OPEN CONTOUR AFTER FIRING. FINALLY, AFTER A WHILE AND BY USING OF HIGH STRENGTH FROM THE SURGEON SIDE IT WAS RELEASED. THE STAPLER LINE WAS INCOMPLETE AND MANY OF THE CLIPS REMAINED IN THE HUB BUT WERE UNFORMED. IT WAS PERFORMED BY EXPERIENCED SURGEON WHO USES CONTOUR REGULARLY AND BY SALES REP THERE WAS NOT TOO MUCH TISSUE. THE PATIENT IS FINE, AND NO COMPLICATIONS HAVE OCCURRED.
20-12-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT A STAPLE SEAM WITH THE CDH29A APPARATUS DID NOT FORM DURING SURGERY AT THE STAGE OF IMPOSING A CIRCULAR ANASTAMOSIS. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
19-12-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | ILS 25MM, CURVED | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, LOCKED TRIGGER. THERE WERE NO PATIENT CONSEQUENCES.
18-12-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC SIGMOID RESECTION, THE STAPLER WAS CLOSED DOWN WITHIN THE RANGE, THE BACK LIGHT WAS LIT, AND WAS FIRED WITHOUT ISSUE BY THE ATTENDING PHYSICIAN. THE STAPLER WAS OPENED WITH TWO COMPLETE REVOLUTIONS AND FISHTAILED OUT. AS THE PHYSICIAN WAS TESTING THE ANASTOMOSIS WITH AN ASEPTO SYRINGE AND AIR, THE ANASTOMOSIS FELL APART. THE CUT LINE WAS COMPLETE. THEY EXCISED THE TISSUE AND INSPECTED FOR MALFORMED STAPLES. THEY DID NOT FIND ANY STAPLES IN THE TISSUE. A COMPETITIVE DEVICE WAS OPENED AND USED SUCCESSFULLY. THE PATIENT RECEIVED A DIVERTING COLOSTOMY, BUT THAT WAS PREPLANNED AND NOT A RESULT OF THE STAPLED ANASTOMOSIS. THE PROCEDURE WAS DELAYED BY NINETY MINUTES. ANASTOMOSIS WAS EXCISED AND ANOTHER ANASTOMOSIS PERFORMED.
17-12-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE( MADE IN CHINA) TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-11-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | 55MM SELECTABLE NEW TLC | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DURING CONNECTING THE TWO PIECES OF THE NTLC A PLASTIC PIECE WAS BROKEN. A NEW DEVICE WAS OPENED. THERE WERE NO PATIENT CONSEQUENCES.
20-12-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | ILS 29MM, CURVED | Break | A0401 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE ANASTOMOTIC SITE STAPLES MALFORMED AND TISSUE COMPRESSION SCALE IS DAMAGE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE ANASTOMOTIC SITE STAPLES MALFORMED AND TISSUE COMPRESSION SCALE IS DAMAGE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2019 | Injury | intuitive surgical | 15-01-2020 | ENDOWRIST | Failure to Form Staple | A050704 | Injury | Event Description: REFER TO H10/H11 FOR ADDITIONAL INFORMATION.
21-11-2019 | Malfunction | ethicon endo-surgery | 15-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING AN APPENDECTOMY THE STAPLER WAS FIRED SUCCESSFULLY, WHEN LOAD WAS REMOVED THE LOAD SPLIT IN HALF. THE PAN OF THE SLED REMAINED WITHIN THE STAPLER. THE STAPLER WAS NOT ABLE TO BE RELOADED. A SIMILAR DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES.
02-01-2020 | Malfunction | us surgical puerto rico | 15-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE LIGATION AND DIVISION OF PULMONARY VEIN IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, THE OPERATOR TRIED TO CLAMP THE JAWS BETWEEN THE TARGETED VESSELS, AND THE HANDLE KNOB COULD BE SQUEEZED BUT THE JAWS OF THE RELOAD APPROXIMATED SUCCESSFULLY BUT WAS NOT LOCKED. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 15-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE LIGATION AND DIVISION OF PULMONARY VEIN IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC LOBECTOMY PROCEDURE, THE OPERATOR TRIED TO CLAMP THE JAWS BETWEEN THE TARGETED VESSELS, AND THE HANDLE KNOB COULD BE SQUEEZED BUT THE JAWS OF THE RELOAD APPROXIMATED SUCCESSFULLY BUT WAS NOT LOCKED. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | covidien lp llc north haven | 15-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, AFTER CLAMPING THE BRANCH OF THE HEPATIC VEIN AND PUSHING THE GREEN SAFETY BUTTON SUCCESSFULLY, THE HANDLE COULD NOT BE FULLY SQUEEZED AND THE BLADE OF THE RELOADS GOT STUCK AT THE BEGINNING PHASE OF FIRING. THE SURGEON RETRACTED THE RETRACTION KNOB FULLY AND OPENED THE JAW OUT OF THE TARGET. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-11-2019 | Injury | covidien lp llc north haven | 15-01-2020 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSED. IT WAS ALSO REPORTED THAT THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
17-11-2019 | Injury | covidien lp llc north haven | 15-01-2020 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | Radiation Exposure, Unintended | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING THE FIRST FIRING OF THE STAPLER ON THE STOMACH, DISTAL STAPLES AT THE END OF THE RELOAD DID NOT CLOSED. IT WAS ALSO REPORTED THAT THE PLASTIC TABS WERE NOT PUSHED AND THE DISTAL SUTURE DID NOT RELEASED. THE REINFORCEMENT STRIP HAD TO BE REMOVED WITH SHEARS. AND PROCEDURE WAS COMPLETED USING A POWER HANDLE. POST-OPERATIVELY THE PATIENT HAD ABDOMINAL PAIN AND THE SURGEON CONDUCTED A COMPUTED TOMOGRAPHY SCAN AND IT WAS CONFIRMED THAT THERE WAS NO LEAK.
12-12-2019 | Malfunction | covidien lp llc north haven | 15-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE START OF STAPLE FIRING DURING ANASTOMOSIS IN A LAPAROSCOPIC GASTROPLASTY PROCEDURE, THE LOAD LOCKED AND DID NOT FIRE. THE PRODUCT WAS CHANGED WITH ANOTHER LOT NUMBER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | GST BLACK RELOAD, 60MM, 6 ROW | Device Slipped | A051204 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT PRIOR TO THE START OF AN UNKNOWN PROCEDURE, THE METAL CASING ON THE BOTTOM OF THE RELOAD WAS BENT AND LOOSE FROM THE CARTRIDGE, RIGHT OUT OF THE PACKAGE. AS A RESULT, THE SLED WAS ALSO PUSHED BACK IN THE RELOAD, MAKING THE RELOAD UNUSABLE. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | GST BLACK RELOAD, 60MM, 6 ROW | Material Twisted/Bent | A040609 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT PRIOR TO THE START OF AN UNKNOWN PROCEDURE, THE METAL CASING ON THE BOTTOM OF THE RELOAD WAS BENT AND LOOSE FROM THE CARTRIDGE, RIGHT OUT OF THE PACKAGE. AS A RESULT, THE SLED WAS ALSO PUSHED BACK IN THE RELOAD, MAKING THE RELOAD UNUSABLE. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
18-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, SOME OF THE STAPLES HAVE NOT BEEN CLOSED. HEALTH PROFESSIONAL INFORMED ME THAT THERE ARE NO ISSUES OBSERVED WITH THE PATIENT. THERE WAS NO NEED FOR EXTENDED HOSPITAL STAY OR ADDITIONAL SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-11-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE PANCREATODUODENECTOMY SURGERY, THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A HALF WHEN SEVERING GASTRIC TISSUE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | GST WHITE RELOAD, 45MM, 6 ROW | Packaging Problem | A0205 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LUNG CANCER PROCEDURE, BEFORE THE DEVICE WAS USED ON THE PATIENT, IT WAS NOTED THAT THE PACKAGING WAS DAMAGED (THE SEALED SITE WAS OPEN). ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-12-2019 | Malfunction | us surgical puerto rico | 16-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE START OF STAPLE FIRING DURING ANASTOMOSIS IN A LAPAROSCOPIC GASTROPLASTY PROCEDURE, THE LOAD LOCKED AND DID NOT FIRE. THE PRODUCT WAS CHANGED WITH ANOTHER LOT NUMBER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-08-2019 | Malfunction | covidien lp llc north haven | 16-01-2020 | SIGNIA | Device Misassembled During Manufacturing /Shipping | A0206 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE SCREEN WAS NOT ILLUMINATING BRIGHT ENOUGH. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
18-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 29MM, STRAIGHT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE RIGHT COLON SURGERY, AFTER FIRED, THERE WAS NO VOICE FEEDBACK WHEN FIRED TO THE END. IT WAS SEEN THAT THE STAPLE ONLY HALF COMPLETE AND THE OTHER HALF WERE INCOMPLETE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
19-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLES WERE MALFORMED WHEN THE DEVICE WAS USED ON THE RIGHT MIDDLE LOBAR BRONCHUS. THE TARGET TISSUE WAS NOT STAPLED, SO AN ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. ADDITIONAL INFORMATION WAS PROVIDED THAT THE STAPLES WERE C SHAPED DUE TO SHORT STALING LEGS AGAINST THE TISSUE. THE BRONCHI WAS NOT SUTURED, SO THE SURGEON SUTURED THE LINE BY HIMSELF UNDER VATS AND COMPLETED THE PROCEDURE. THERE WAS NO INFORMATION ABOUT LEAKAGE AND BLEEDING. THE PATIENT IS STABLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
20-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AM OPEN TOTAL GASTRECTOMY, THE TROCAR WAS MOVED INTO THE HOUSING EASILY WHEN IT CONTACTED TO THE ANVIL WHILE THE SURGEON TRIED TO ATTACH THE TROCAR TO THE ANVIL. THE DEVICE INTENDED TO BE USED BETWEEN THE ESOPHAGUS AND THE JEJUNUM. SINCE THE ISSUE WAS NOT RESOLVED, THE SURGEON FIXED THE ADJUSTING KNOB BY HIS HAND TO FIRE THE DEVICE. AFTER THE FIRING, IT WAS FOUND THAT THE MUCOUS MEMBRANE WAS DAMAGED, AND THE DEPLOYED STAPLES WERE NOT FORMED PROPERLY. THE SURGEON USED THE REPLACEMENT AND ANASTOMOSED WITHOUT CHANGING THE PROCEDURE EVEN THOUGH THE MUCOSAL DAMAGE AROUND THE ANASTOMOSIS WAS CONFIRMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THE PATIENT IS STABLE. NO FURTHER INFORMATION IS AVAILABLE.
12-11-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ENDOPATH | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: SURGEONS USING STAPLER DEVICE FOR LAPAROSCOPIC APPOINTMENT PROCEDURE AND STAPLER MISFIRED. THIS WAS THE 2ND RELOAD. THE FIRST RELOAD WAS ETS45 - 3.5MM NO ISSUES WITH FIRING. 2ND RELOAD WAS ETS45 - 2.5MM IT ENGAGED AND STARTED TO FIRE FOR DOCTOR, THEN STOPPED FIRING. DOCTOR #2 STANDING NEXT TO DOCTOR#1 AT TIME OF EVENT.
25-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE CDH29A DIDN'T HAVE THE STAPLE FORMATION WELL.
09-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 25MM, CURVED | Activation, Positioning or Separation Problem | A15 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLER FAILED WHEN ACTIVATING, RELEASING THE LOAD PARTIALLY.
09-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 25MM, CURVED | Activation, Positioning or Separation Problem | A15 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLER FAILED WHEN ACTIVATING, RELEASING THE LOAD PARTIALLY.
09-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 25MM, CURVED | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLER FAILED WHEN ACTIVATING, RELEASING THE LOAD PARTIALLY.
09-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 25MM, CURVED | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLER FAILED WHEN ACTIVATING, RELEASING THE LOAD PARTIALLY.
23-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 29MM, CURVED | Material Twisted/Bent | A040609 | No Consequences Or Impact To Patient | Event Description: CUTTING ISSUE. IT WAS REPORTED THAT DURING THE RECTAL CANCER SURGERY, AFTER FIRE, NOTE THAT THERE IS NO AUDIBLE FEEDBACK. BLADE SHOWED BENT. CUT WASHER WAS NOT CUT COMPLETE. OPERATOR CUT ANASTOMOTIC SITE. CHANGED THE ECS29 TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | PROXIMATE*RL TITAN HEAVY WIRE | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY OF LEFT LUNG, MOST OF THE DEPLOYED STAPLES WERE NOT FORMED PROPERLY AT THE FIRST FIRING ON THE LEFT MAIN BRONCHUS. THOSE WERE BARELY STUCK ON THE TISSUE AND THE DEPLOYED STAPLES WERE U-SHAPED THE ISSUE WAS RECOGNIZED WHEN THE BRONCHIAL TUBE OF THE SPECIMEN SIDE WAS CUT AFTER THE FIRST FIRING. ADDITIONAL SUTURE WAS PERFORMED WITH PDS TO COMPLETE THE CASE, AND NO PROBLEM WAS OBSERVED AT A LEAK TEST. NO LEAK OCCURRED WHEN THE STAPLES WERE DEPLOYED IN U-SHAPED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLING WAS NOT COMPLETE.
26-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ANASTOMOSIS DIDN'T FORM PROPERLY, CIRCULAR STAPLER. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
27-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 25MM, CURVED | Difficult to Remove | A150207 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE RESISTANCE OF REMOVING THE DEVICE FROM THE PATIENT WAS STRONGER THAN USUAL AFTER THE FIRING. THE DEVICE WAS USED BETWEEN THE COLON AND THE RECTUM FOR RECONSTRUCTION. A PART OF THE SEROUS MEMBRANE WAS TORN WHEN REMOVING THE DEVICE SINCE THE DEVICE HAD A DIFFICULTY IN BEING REMOVED. THE ANASTOMOSIS SITE WAS CUT ADDITIONALLY, AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THE SURGEON HAD EXPERIENCE IN USING CDH DEVICES. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 25MM, CURVED | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE RESISTANCE OF REMOVING THE DEVICE FROM THE PATIENT WAS STRONGER THAN USUAL AFTER THE FIRING. THE DEVICE WAS USED BETWEEN THE COLON AND THE RECTUM FOR RECONSTRUCTION. A PART OF THE SEROUS MEMBRANE WAS TORN WHEN REMOVING THE DEVICE SINCE THE DEVICE HAD A DIFFICULTY IN BEING REMOVED. THE ANASTOMOSIS SITE WAS CUT ADDITIONALLY, AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THE SURGEON HAD EXPERIENCE IN USING CDH DEVICES. NO FURTHER INFORMATION IS AVAILABLE.
30-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, THERE WERE THREE MALFORMED STAPLES IN THE ANASTOMOSES. CHD29P WAS FIRED NORMALLY WITH TWO COMPLETE INTACT DONUTS AND NO BUBBLES WERE SEEN IN THE LEAK TEST. SURGEON BELIEVES STAPLE LINES FROM PREVIOUS TRANSECTIONS MAY HAVE BEEN CROSSED. THE CASE WAS SUCCESSFULLY COMPLETED WITH THE DEVICE AND NO PATIENT HARM.
30-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 21MM, CURVED | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A BARIATRIC SURGERY ON (B)(6) 2019 THE DOCTOR HAD A PROBLEM WITH A CIRCULAR STAPLER CDH21A (LOT T40L8K). THE STAPLER FIRED BUT DID NOT CUT. HE COMPLETED THE SURGERY WITH A CHD25A, DESPITE BEING TOO LARGE, NOT HAVING OTHER CIRCULARS AVAILABLE LIKE A COMPETITIVE STAPLER OF THE CORRECT SIZE AND NOT TRUSTING TO USE THE CDH21A OF THE SAME LOT.
19-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, CDH29A MADE A BAD STAPLE LINE.
20-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: STAPLE NOT IN THE TISSUE. IT WAS REPORTED THAT DURING RADICAL RESECTION OF RECTAL CANCER. AFTER FIRED, SOME STAPLE IN THE TISSUE, OTHERS WERE NOT IN THE TISSUE. CUT THE PREVIOUS ANASTOMOSIS SITE AND USE THE SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | POWERED PLUS LONG 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, STAPLES DEPLOYED BUT NOT ALL FORMED. BLUE LOAD BEING USED. THE SURGERY WAS DELAYED 30 MINUTES DUE TO THE REPORTED EVENT. ACTION TAKEN WHEN EVENT OCCURRED WAS ADDITIONAL FIRING TO RECONTOUR GASTRIC POUCH. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
17-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A BAND REVERSAL AND CONVERSION TO A SLEEVE GASTRECTOMY USING A PLEE60A. THE DEVICE FIRED THREE TIMES WITH NO ISSUES A GOLD RELOAD WAS LOADED AND FIRED ON THE SLEEVE PORTION. IT DID STAPLE ON THE SPECIMEN SIDE AND TRANSECTED, BUT IT DID NOT DEPLOY ANY STAPLES ON THE PATIENT SIDE. THE SURGEON HAD TO OVER SEW ON THE PATIENT SIDE. THE SAME STAPLER WAS USED WITH TWO BLUE RELOADS WITH NO ISSUES. WHERE THE GOLD RELOAD MISFIRE, THE PATIENT ENDED UP GETTING A LEAK AND WILL BE RETURNING TO SURGERY TOMORROW.
17-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A BAND REVERSAL AND CONVERSION TO A SLEEVE GASTRECTOMY USING A PLEE60A. THE DEVICE FIRED THREE TIMES WITH NO ISSUES A GOLD RELOAD WAS LOADED AND FIRED ON THE SLEEVE PORTION. IT DID STAPLE ON THE SPECIMEN SIDE AND TRANSECTED, BUT IT DID NOT DEPLOY ANY STAPLES ON THE PATIENT SIDE. THE SURGEON HAD TO OVER SEW ON THE PATIENT SIDE. THE SAME STAPLER WAS USED WITH TWO BLUE RELOADS WITH NO ISSUES. WHERE THE GOLD RELOAD MISFIRE, THE PATIENT ENDED UP GETTING A LEAK AND WILL BE RETURNING TO SURGERY TOMORROW.
18-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Material Frayed | A040605 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY INTERVENTION, ON THE FIFTH SHOT AT THE LAST ACTIVATION, NUMEROUS MALFORMED POINTS APPEARED. THE OUTER "MARGIME" OF THE SUTURE APPEARS FRAYED WITH SIMULTANEOUS BLEEDING. THE SURGERY WAS DELAYED FOR 15 MINUTES, BUT WAS COMPLETED SUCCESSFULLY USING CLIPS. FRAGMENTS WERE GENERATED, BUT THEY WERE EASILY REMOVED WITHOUT ADDITIONAL INTERVENTION. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
18-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY INTERVENTION, ON THE FIFTH SHOT AT THE LAST ACTIVATION, NUMEROUS MALFORMED POINTS APPEARED. THE OUTER "MARGIME" OF THE SUTURE APPEARS FRAYED WITH SIMULTANEOUS BLEEDING. THE SURGERY WAS DELAYED FOR 15 MINUTES, BUT WAS COMPLETED SUCCESSFULLY USING CLIPS. FRAGMENTS WERE GENERATED, BUT THEY WERE EASILY REMOVED WITHOUT ADDITIONAL INTERVENTION. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
18-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Material Frayed | A040605 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY INTERVENTION, ON THE THIRD SHOT IN CORRESPONDENCE OF THE BODY OF THE STOMACH, THERE ARE NUMEROUS MALFORMED POINTS. THE OUTER MARGIN OF THE SUTURE APPEARS FRAYED WITH SIMULTANEOUS BLEEDING. THE PROCEDURE WAS DELAYED FOR 15 MINUTES, BUT WAS SUCCESSFULLY COMPLETED USING CLIPS. FRAGMENTS WERE GENERATED, BUT WERE REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION. NO PATIENT CONSEQUENCES WERE REPORTED.
18-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY INTERVENTION, ON THE THIRD SHOT IN CORRESPONDENCE OF THE BODY OF THE STOMACH, THERE ARE NUMEROUS MALFORMED POINTS. THE OUTER MARGIN OF THE SUTURE APPEARS FRAYED WITH SIMULTANEOUS BLEEDING. THE PROCEDURE WAS DELAYED FOR 15 MINUTES, BUT WAS SUCCESSFULLY COMPLETED USING CLIPS. FRAGMENTS WERE GENERATED, BUT WERE REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION. NO PATIENT CONSEQUENCES WERE REPORTED.
18-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Material Frayed | A040605 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING BYPASS SURGERY, NUMEROUS BADLY FORMED POINTS WERE HIGHLIGHTED IN THE SECOND ACTIVATION DURING THE PACKAGING OF THE GASTRIC POCKET. THE OUTER EDGE OF THE SUTURE ON THE ONE CLOSEST TO THE SCALPEL BLADE APPEARED FRAYED AS IF THE CUT WAS NOT CLEAN, WITH SIMULTANEOUS ABUNDANT BLEEDING. SURGERY WAS DELAYED 15 MINUTES, BUT WAS SUCCESSFULLY COMPLETED WITH A SUTURE RHYME CORRECTION CLIP. FRAGMENTS WERE GENERATED, BUT WERE EASILY REMOVED WITHOUT ADDITIONAL INTERVENTION. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
18-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING BYPASS SURGERY, NUMEROUS BADLY FORMED POINTS WERE HIGHLIGHTED IN THE SECOND ACTIVATION DURING THE PACKAGING OF THE GASTRIC POCKET. THE OUTER EDGE OF THE SUTURE ON THE ONE CLOSEST TO THE SCALPEL BLADE APPEARED FRAYED AS IF THE CUT WAS NOT CLEAN, WITH SIMULTANEOUS ABUNDANT BLEEDING. SURGERY WAS DELAYED 15 MINUTES, BUT WAS SUCCESSFULLY COMPLETED WITH A SUTURE RHYME CORRECTION CLIP. FRAGMENTS WERE GENERATED, BUT WERE EASILY REMOVED WITHOUT ADDITIONAL INTERVENTION. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
17-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | POWERED PLUS COMPACT 45 | Activation Problem | A1501 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY PROCEDURE, THE SURGEON BELIEVES THAT THE STAPLES DID NOT FIRE AND JUST CUT. THE SURGEON HAD TO CONVERT FROM A VATS TO AN OPEN PROCEDURE. AS A RESULT OF THIS THE RECOVERY TIME OF THE PATIENT WAS A LOT LONGER. THE REASON WHY THE SURGEON HAD TO CONVERT IS BECAUSE HE BELIEVES THE STAPLER DID NOT DEPLOY THE STAPLES OVER A TISSUE BUNDLE WHICH CONTAINED A VASCULAR STRUCTURE. CONVERTING TO AN OPEN PROCEDURE ALLOWED HIM BETTER CONTROL OF THE BLEEDING VESSEL. FORTUNATELY, NO BLOOD TRANSFUSIONS WERE NEEDED BUT THE LENGTH OF SURGERY INCREASED. PATIENT HAS MADE A FULL RECOVERY.
30-09-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Therapeutic or Diagnostic Output Failure | A090406 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, THE STAPLER AND RELOAD DID NOT WORK. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
11-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Problem | A05 | Chest Pain | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY SURGERY, 2 DEVICES PLEE60A IMMEDIATELY BLOCKED, DURING THE FIRST INTERLOCKS. THE SURGEON STARTED WITH A PLEE60A AND A GREEN CHARGER GST60G. THE CLAMP IMMEDIATELY BLOCKED AND THE STAPLING DIDN'T START. THEN, THE SURGEON USED ANOTHER CHARGER WITH THE SAME LOT, THE SAME INCIDENT OCCURRED. THEN, THE SURGEON DECIDED TO CHANGE THE STAPLER : HE TOOK ANOTHER PLEE60A AND USED ANOTHER CHARGER GST60G ( IT IS THE THIRD CHARGER OF THIS BOX ) : THE SAME INCIDENT OCCURRED THEN , THE SURGEON THOUGHT THAT THE TISSUES WERE TOO THICK, CHANGED THE CHARGER, HE TOOK A DEVICE GST60T BUT THE STAPLER DIDN'T STAPLE. THEN, THE SURGEON USED A MANUAL STAPLER LONG60A AND THE SURGERY WAS COMPLETED. DESPITE THE 4 SUCCESSIVE COMPRESSION OF THE TISSUES ON THE PATIENT PYLORUS, NO CONSEQUENCE FOR THE PATIENT AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THE FOLLOWING DAY, A REVISION WAS DONE BECAUSE THE PATIENT HAD PAIN ON THE CHEST : THE SURGEON NOTIFIED THAT THE STAPLING REMAINED IN PLACE. A BLUE TEST SHOWED THAT THE MECHANICAL SUTURE WAS HERMETIC. THE SALES REP DIDN'T HAVE OTHER INFORMATION CONCERNING THE REVISION SURGERY BUT THE PATIENT IS WELL.
11-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Problem | A05 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY SURGERY, 2 DEVICES PLEE60A IMMEDIATELY BLOCKED, DURING THE FIRST INTERLOCKS. THE SURGEON STARTED WITH A PLEE60A AND A GREEN CHARGER GST60G. THE CLAMP IMMEDIATELY BLOCKED AND THE STAPLING DIDN'T START. THEN, THE SURGEON USED ANOTHER CHARGER WITH THE SAME LOT, THE SAME INCIDENT OCCURRED. THEN, THE SURGEON DECIDED TO CHANGE THE STAPLER : HE TOOK ANOTHER PLEE60A AND USED ANOTHER CHARGER GST60G ( IT IS THE THIRD CHARGER OF THIS BOX ) : THE SAME INCIDENT OCCURRED THEN , THE SURGEON THOUGHT THAT THE TISSUES WERE TOO THICK, CHANGED THE CHARGER, HE TOOK A DEVICE GST60T BUT THE STAPLER DIDN'T STAPLE. THEN, THE SURGEON USED A MANUAL STAPLER LONG60A AND THE SURGERY WAS COMPLETED. DESPITE THE 4 SUCCESSIVE COMPRESSION OF THE TISSUES ON THE PATIENT PYLORUS, NO CONSEQUENCE FOR THE PATIENT AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THE FOLLOWING DAY, A REVISION WAS DONE BECAUSE THE PATIENT HAD PAIN ON THE CHEST : THE SURGEON NOTIFIED THAT THE STAPLING REMAINED IN PLACE. A BLUE TEST SHOWED THAT THE MECHANICAL SUTURE WAS HERMETIC. THE SALES REP DIDN'T HAVE OTHER INFORMATION CONCERNING THE REVISION SURGERY BUT THE PATIENT IS WELL.
11-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | Chest Pain | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY SURGERY, 2 DEVICES PLEE60A IMMEDIATELY BLOCKED, DURING THE FIRST INTERLOCKS. THE SURGEON STARTED WITH A PLEE60A AND A GREEN CHARGER GST60G. THE CLAMP IMMEDIATELY BLOCKED AND THE STAPLING DIDN'T START. THEN, THE SURGEON USED ANOTHER CHARGER WITH THE SAME LOT, THE SAME INCIDENT OCCURRED. THEN, THE SURGEON DECIDED TO CHANGE THE STAPLER : HE TOOK ANOTHER PLEE60A AND USED ANOTHER CHARGER GST60G ( IT IS THE THIRD CHARGER OF THIS BOX ) : THE SAME INCIDENT OCCURRED THEN , THE SURGEON THOUGHT THAT THE TISSUES WERE TOO THICK, CHANGED THE CHARGER, HE TOOK A DEVICE GST60T BUT THE STAPLER DIDN'T STAPLE. THEN, THE SURGEON USED A MANUAL STAPLER LONG60A AND THE SURGERY WAS COMPLETED. DESPITE THE 4 SUCCESSIVE COMPRESSION OF THE TISSUES ON THE PATIENT PYLORUS, NO CONSEQUENCE FOR THE PATIENT AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THE FOLLOWING DAY, A REVISION WAS DONE BECAUSE THE PATIENT HAD PAIN ON THE CHEST : THE SURGEON NOTIFIED THAT THE STAPLING REMAINED IN PLACE. A BLUE TEST SHOWED THAT THE MECHANICAL SUTURE WAS HERMETIC. THE SALES REP DIDN'T HAVE OTHER INFORMATION CONCERNING THE REVISION SURGERY BUT THE PATIENT IS WELL.
11-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY SURGERY, 2 DEVICES PLEE60A IMMEDIATELY BLOCKED, DURING THE FIRST INTERLOCKS. THE SURGEON STARTED WITH A PLEE60A AND A GREEN CHARGER GST60G. THE CLAMP IMMEDIATELY BLOCKED AND THE STAPLING DIDN'T START. THEN, THE SURGEON USED ANOTHER CHARGER WITH THE SAME LOT, THE SAME INCIDENT OCCURRED. THEN, THE SURGEON DECIDED TO CHANGE THE STAPLER : HE TOOK ANOTHER PLEE60A AND USED ANOTHER CHARGER GST60G ( IT IS THE THIRD CHARGER OF THIS BOX ) : THE SAME INCIDENT OCCURRED THEN , THE SURGEON THOUGHT THAT THE TISSUES WERE TOO THICK, CHANGED THE CHARGER, HE TOOK A DEVICE GST60T BUT THE STAPLER DIDN'T STAPLE. THEN, THE SURGEON USED A MANUAL STAPLER LONG60A AND THE SURGERY WAS COMPLETED. DESPITE THE 4 SUCCESSIVE COMPRESSION OF THE TISSUES ON THE PATIENT PYLORUS, NO CONSEQUENCE FOR THE PATIENT AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THE FOLLOWING DAY, A REVISION WAS DONE BECAUSE THE PATIENT HAD PAIN ON THE CHEST : THE SURGEON NOTIFIED THAT THE STAPLING REMAINED IN PLACE. A BLUE TEST SHOWED THAT THE MECHANICAL SUTURE WAS HERMETIC. THE SALES REP DIDN'T HAVE OTHER INFORMATION CONCERNING THE REVISION SURGERY BUT THE PATIENT IS WELL.
11-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | GOLD RELOAD FOR ECHELON 60 | Activation, Positioning or Separation Problem | A15 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE SINGLE HOLE THORACOSCOPIC LEFT UPPER LOBECTOMY SURGERY, AFTER THE DEVICE WAS FIRED, IT WAS NOTED THAT ONE SIDE OF ANASTOMOTIC SITE WAS WITHOUT ANY STAPLES AND TISSUE WAS OOZING. USED SUTURE TO OVERSEW AND STOP OOZING. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2019 | Malfunction | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Contamination /Decontamination Problem | A18 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE, A FLAKE OF BLUE PLASTIC WAS LEFT ON THE STAPLE LINE. THE STAPLE LINE WAS COMPLETE. THIS WAS ON THE FOURTH FIRING. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
11-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Problem | A05 | Chest Pain | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY SURGERY, 2 DEVICES PLEE60A IMMEDIATELY BLOCKED, DURING THE FIRST INTERLOCKS. THE SURGEON STARTED WITH A PLEE60A AND A GREEN CHARGER GST60G. THE CLAMP IMMEDIATELY BLOCKED AND THE STAPLING DIDN'T START. THEN, THE SURGEON USED ANOTHER CHARGER WITH THE SAME LOT, THE SAME INCIDENT OCCURRED. THEN, THE SURGEON DECIDED TO CHANGE THE STAPLER : HE TOOK ANOTHER PLEE60A AND USED ANOTHER CHARGER GST60G ( IT IS THE THIRD CHARGER OF THIS BOX ) : THE SAME INCIDENT OCCURRED. THEN , THE SURGEON THOUGHT THAT THE TISSUES WERE TOO THICK, CHANGED THE CHARGER, HE TOOK A DEVICE GST60T BUT THE STAPLER DIDN'T STAPLE. THEN, THE SURGEON USED A MANUAL STAPLER LONG60A AND THE SURGERY WAS COMPLETED. DESPITE THE 4 SUCCESSIVE COMPRESSION OF THE TISSUES ON THE PATIENT PYLORUS, NO CONSEQUENCE FOR THE PATIENT AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THE FOLLOWING DAY, A REVISION WAS DONE BECAUSE THE PATIENT HAD PAIN ON THE CHEST : THE SURGEON NOTIFIED THAT THE STAPLING REMAINED IN PLACE. A BLUE TEST SHOWED THAT THE MECHANICAL SUTURE WAS HERMETIC. THE SALES REP DIDN'T HAVE OTHER INFORMATION CONCERNING THE REVISION SURGERY BUT THE PATIENT IS WELL.
11-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Problem | A05 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY SURGERY, 2 DEVICES PLEE60A IMMEDIATELY BLOCKED, DURING THE FIRST INTERLOCKS. THE SURGEON STARTED WITH A PLEE60A AND A GREEN CHARGER GST60G. THE CLAMP IMMEDIATELY BLOCKED AND THE STAPLING DIDN'T START. THEN, THE SURGEON USED ANOTHER CHARGER WITH THE SAME LOT, THE SAME INCIDENT OCCURRED. THEN, THE SURGEON DECIDED TO CHANGE THE STAPLER : HE TOOK ANOTHER PLEE60A AND USED ANOTHER CHARGER GST60G ( IT IS THE THIRD CHARGER OF THIS BOX ) : THE SAME INCIDENT OCCURRED. THEN , THE SURGEON THOUGHT THAT THE TISSUES WERE TOO THICK, CHANGED THE CHARGER, HE TOOK A DEVICE GST60T BUT THE STAPLER DIDN'T STAPLE. THEN, THE SURGEON USED A MANUAL STAPLER LONG60A AND THE SURGERY WAS COMPLETED. DESPITE THE 4 SUCCESSIVE COMPRESSION OF THE TISSUES ON THE PATIENT PYLORUS, NO CONSEQUENCE FOR THE PATIENT AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THE FOLLOWING DAY, A REVISION WAS DONE BECAUSE THE PATIENT HAD PAIN ON THE CHEST : THE SURGEON NOTIFIED THAT THE STAPLING REMAINED IN PLACE. A BLUE TEST SHOWED THAT THE MECHANICAL SUTURE WAS HERMETIC. THE SALES REP DIDN'T HAVE OTHER INFORMATION CONCERNING THE REVISION SURGERY BUT THE PATIENT IS WELL.
11-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | Chest Pain | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY SURGERY, 2 DEVICES PLEE60A IMMEDIATELY BLOCKED, DURING THE FIRST INTERLOCKS. THE SURGEON STARTED WITH A PLEE60A AND A GREEN CHARGER GST60G. THE CLAMP IMMEDIATELY BLOCKED AND THE STAPLING DIDN'T START. THEN, THE SURGEON USED ANOTHER CHARGER WITH THE SAME LOT, THE SAME INCIDENT OCCURRED. THEN, THE SURGEON DECIDED TO CHANGE THE STAPLER : HE TOOK ANOTHER PLEE60A AND USED ANOTHER CHARGER GST60G ( IT IS THE THIRD CHARGER OF THIS BOX ) : THE SAME INCIDENT OCCURRED. THEN , THE SURGEON THOUGHT THAT THE TISSUES WERE TOO THICK, CHANGED THE CHARGER, HE TOOK A DEVICE GST60T BUT THE STAPLER DIDN'T STAPLE. THEN, THE SURGEON USED A MANUAL STAPLER LONG60A AND THE SURGERY WAS COMPLETED. DESPITE THE 4 SUCCESSIVE COMPRESSION OF THE TISSUES ON THE PATIENT PYLORUS, NO CONSEQUENCE FOR THE PATIENT AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THE FOLLOWING DAY, A REVISION WAS DONE BECAUSE THE PATIENT HAD PAIN ON THE CHEST : THE SURGEON NOTIFIED THAT THE STAPLING REMAINED IN PLACE. A BLUE TEST SHOWED THAT THE MECHANICAL SUTURE WAS HERMETIC. THE SALES REP DIDN'T HAVE OTHER INFORMATION CONCERNING THE REVISION SURGERY BUT THE PATIENT IS WELL.
11-12-2019 | Injury | ethicon endo-surgery | 16-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY SURGERY, 2 DEVICES PLEE60A IMMEDIATELY BLOCKED, DURING THE FIRST INTERLOCKS. THE SURGEON STARTED WITH A PLEE60A AND A GREEN CHARGER GST60G. THE CLAMP IMMEDIATELY BLOCKED AND THE STAPLING DIDN'T START. THEN, THE SURGEON USED ANOTHER CHARGER WITH THE SAME LOT, THE SAME INCIDENT OCCURRED. THEN, THE SURGEON DECIDED TO CHANGE THE STAPLER : HE TOOK ANOTHER PLEE60A AND USED ANOTHER CHARGER GST60G ( IT IS THE THIRD CHARGER OF THIS BOX ) : THE SAME INCIDENT OCCURRED. THEN , THE SURGEON THOUGHT THAT THE TISSUES WERE TOO THICK, CHANGED THE CHARGER, HE TOOK A DEVICE GST60T BUT THE STAPLER DIDN'T STAPLE. THEN, THE SURGEON USED A MANUAL STAPLER LONG60A AND THE SURGERY WAS COMPLETED. DESPITE THE 4 SUCCESSIVE COMPRESSION OF THE TISSUES ON THE PATIENT PYLORUS, NO CONSEQUENCE FOR THE PATIENT AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THE FOLLOWING DAY, A REVISION WAS DONE BECAUSE THE PATIENT HAD PAIN ON THE CHEST : THE SURGEON NOTIFIED THAT THE STAPLING REMAINED IN PLACE. A BLUE TEST SHOWED THAT THE MECHANICAL SUTURE WAS HERMETIC. THE SALES REP DIDN'T HAVE OTHER INFORMATION CONCERNING THE REVISION SURGERY BUT THE PATIENT IS WELL.
05-11-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE PATIENT WAS GIVEN THE RIGHT HEMICOLECTOMY SURGERY ON (B)(6) 2019, NO ABNORMAL SITUATION WAS OBSERVED IN THE SURGERY. 8 DAYS AFTER SURGERY, NOTED ABNORMAL LIQUID IN DRAINAGE TUBE, DID THE 2ND SURGERY, FOUND ANASTOMOTIC SITE WAS LEAK, USED THE SUTURE TO OVERSEW. THE PATIENT IS STABLE AND LEFT FROM HOSPITAL.
05-11-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT THE PATIENT WAS GIVEN THE RIGHT HEMICOLECTOMY SURGERY ON (B)(6) 2019, NO ABNORMAL SITUATION WAS OBSERVED IN THE SURGERY. 8 DAYS AFTER SURGERY, NOTED ABNORMAL LIQUID IN DRAINAGE TUBE, DID THE 2ND SURGERY, FOUND ANASTOMOTIC SITE WAS LEAK, USED THE SUTURE TO OVERSEW. THE PATIENT IS STABLE AND LEFT FROM HOSPITAL.
17-11-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE PATIENT WAS GIVEN THE INTESTINAL ADHESIOLYSIS AND OPEN EXPLORATION SURGERY ON (B)(6) 2019, NO ABNORMAL SITUATION WAS OBSERVED IN THE SURGERY. 5 DAYS AFTER SURGERY, NOTED ABNORMAL LIQUID IN DRAINAGE TUBE, DID THE 2ND SURGERY, FOUND ANASTOMOTIC SITE WAS LEAK, USED THE SUTURE TO OVERSEW. THE PATIENT IS STABLE AND LEFT FROM HOSPITAL.
17-11-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT THE PATIENT WAS GIVEN THE INTESTINAL ADHESIOLYSIS AND OPEN EXPLORATION SURGERY ON (B)(6) 2019, NO ABNORMAL SITUATION WAS OBSERVED IN THE SURGERY. 5 DAYS AFTER SURGERY, NOTED ABNORMAL LIQUID IN DRAINAGE TUBE, DID THE 2ND SURGERY, FOUND ANASTOMOTIC SITE WAS LEAK, USED THE SUTURE TO OVERSEW. THE PATIENT IS STABLE AND LEFT FROM HOSPITAL.
17-01-2020 | Malfunction | covidien lp llc north haven | 17-01-2020 | SIGNIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING RESECTION OF GASTRIC SPECIMEN FOR BARIATRIC PROCEDURE IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE REMAINING UNCUT PORTION OF BUTTRESS MATERIAL HAD TO BE CUT IN ORDER TO RELEASE THE STAPLER JAWS FOR REMOVAL. THERE WAS NO PATIENT INJURY.
17-01-2020 | Malfunction | covidien lp llc north haven | 17-01-2020 | SIGNIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING RESECTION OF GASTRIC SPECIMEN FOR BARIATRIC PROCEDURE IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE REMAINING UNCUT PORTION OF BUTTRESS MATERIAL HAD TO BE CUT IN ORDER TO RELEASE THE STAPLER JAWS FOR REMOVAL. THERE WAS NO PATIENT INJURY.
09-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Break | A0401 | No Code Available | Event Description: IT WAS REPORTED THAT A REVISION SURGERY OCCURRED ON (B)(6) 2019 : THE PATIENT WAS OPERATED ON (B)(6) 2019 FOR A LEFT COLECTOMY. DURING THE REVISION, THE SURGEON DISCOVERED THAT THE MECANICAL SUTURE BROKE, A STOMA WAS DONE.
09-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Appropriate Term/Code Not Available |  | No Code Available | Event Description: IT WAS REPORTED THAT A REVISION SURGERY OCCURRED ON (B)(6) 2019 : THE PATIENT WAS OPERATED ON (B)(6) 2019 FOR A LEFT COLECTOMY. DURING THE REVISION, THE SURGEON DISCOVERED THAT THE MECANICAL SUTURE BROKE, A STOMA WAS DONE.
31-12-2019 | Malfunction | covidien lp llc north haven | 17-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE INSTALLING THE DEVICE DURING A LAPAROSCOPIC DISTAL STOMACH SURGERY, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2019 | Malfunction | us surgical puerto rico | 17-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE INSTALLING THE DEVICE DURING A LAPAROSCOPIC DISTAL STOMACH SURGERY, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2020 | Malfunction | covidien lp llc north haven | 17-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SIGMOID COLECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 17-01-2020 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING RESECTION OF GASTRIC SPECIMEN FOR BARIATRIC PROCEDURE IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE REMAINING UNCUT PORTION OF BUTTRESS MATERIAL HAD TO BE CUT IN ORDER TO RELEASE THE STAPLER JAWS FOR REMOVAL. THERE WAS NO PATIENT INJURY.
10-01-2020 | Malfunction | ethicon endo-surgery | 17-01-2020 | PROXIMATE | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: MD WAS USING THE LINEAR STAPLER TO CLOSE BOWL AND THE STAPLER WOULD NOT FIRE.
18-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | Blood Loss | Event Description: IT WAS REPORTED THAT DURING A DIFFICULT ROBOTIC ASSISTED THORACIC SURGERY THE SURGEON ASKED TO OPEN A POWERED ECHELON STAPLER 60MM (PRODUCT CODE: PCEE60A) WHICH, AFTER BEING FIRED SEVERAL TIMES, JAMMED WHILST CLAMPED ON LUNG TISSUE; THE JAWS WOULD NOT OPEN REGARDLESS OF TROUBLESHOOTING ATTEMPTS (INCLUDING MANUAL OVERRIDE). A SECOND STAPLER WAS OPENED AND CLAMPED JUST ABOVE THE ORIGINAL STAPLER TO REMOVE THE LUNG TISSUE SAFELY HOWEVER FOLLOWING THIS, HEAVY BLEEDING OCCURRED THAT COULD ONLY BE SAFELY CONTROLLED BY CONVERSION TO AN EMERGENCY THORACOTOMY.
18-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A DIFFICULT ROBOTIC ASSISTED THORACIC SURGERY THE SURGEON ASKED TO OPEN A POWERED ECHELON STAPLER 60MM (PRODUCT CODE: PCEE60A) WHICH, AFTER BEING FIRED SEVERAL TIMES, JAMMED WHILST CLAMPED ON LUNG TISSUE; THE JAWS WOULD NOT OPEN REGARDLESS OF TROUBLESHOOTING ATTEMPTS (INCLUDING MANUAL OVERRIDE). A SECOND STAPLER WAS OPENED AND CLAMPED JUST ABOVE THE ORIGINAL STAPLER TO REMOVE THE LUNG TISSUE SAFELY HOWEVER FOLLOWING THIS, HEAVY BLEEDING OCCURRED THAT COULD ONLY BE SAFELY CONTROLLED BY CONVERSION TO AN EMERGENCY THORACOTOMY.
03-01-2020 | Injury | us surgical puerto rico | 17-01-2020 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO THE SMALL BOWEL DURING AN OPEN COLECTOMY PROCEDURE, THERE WAS STAPLE LINE BLEEDING. THE STAPLE LINE WAS OVERSEWED AND BOVIED (DISSECTED AND CAUTERIZED USING A BOVIE INSTRUMENT) TO CONTROL THE BLEEDING AND COMPLETE THE CASE.
30-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT IS REPORTED THAT DURING THE RADICAL RESECTION OF RECTAL CANCER SURGERY WHEN SERVING RECTAL TISSUE WITH CS40G. AFTER FIRED, NOTED THAT SOME STAPLES WERE IN THE TISSUE, OTHERS WERE NOT IN THE TISSUE. CHANGED CDH29A STILL THE SAME SITUATION. OSTOMY USED TO COMPLETE SURGERY.
30-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ILS 29MM, CURVED | Activation Problem | A1501 | No Code Available | Event Description: IT IS REPORTED THAT DURING THE RADICAL RESECTION OF RECTAL CANCER SURGERY WHEN SERVING RECTAL TISSUE WITH CS40G. AFTER FIRED, NOTED THAT SOME STAPLES WERE IN THE TISSUE, OTHERS WERE NOT IN THE TISSUE. CHANGED CDH29A STILL THE SAME SITUATION. OSTOMY USED TO COMPLETE SURGERY.
30-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Code Available | Event Description: IT IS REPORTED THAT DURING THE RADICAL RESECTION OF RECTAL CANCER SURGERY WHEN SERVING RECTAL TISSUE WITH CS40G. AFTER FIRED, NOTED THAT SOME STAPLES WERE IN THE TISSUE, OTHERS WERE NOT IN THE TISSUE. CHANGED CDH29A STILL THE SAME SITUATION. OSTOMY USED TO COMPLETE SURGERY.
30-10-2019 | Malfunction | covidien lp llc north haven | 17-01-2020 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE HANDLE BROKE AND WOULD NOT RELEASE THE RELOAD. THE PROCEDURE WAS DONE BY USING ANOTHER DEVICE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-10 (EGIAUSTND) HANDLE: AFC ANALYSIS FINDING CODES USEGATKTIS, US: EGIA HANDLE THICK TISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-20 (SIGRADMT) RELOAD: AFC ANALYSIS FINDING CODES USEGLTKTIS, US: EGIA RELOAD THICK TISSUE.
03-01-2020 | Injury | us surgical puerto rico | 17-01-2020 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO THE SMALL BOWEL DURING AN OPEN COLECTOMY PROCEDURE, THERE WAS STAPLE LINE BLEEDING. THE STAPLE LINE WAS OVERSEWED AND BOVIED (DISSECTED AND CAUTERIZED USING A BOVIE INSTRUMENT) TO CONTROL THE BLEEDING AND COMPLETE THE CASE.
16-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | PWRD ECH FLEX 45MM 44CM SHAFT | Leak/Splash | A0504 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT ON THE 2ND POSTOPERATIVE DAY, THE PATIENT WAS TACHYCARDIC, HAD AN IMAGING EXAM, THE SAME WAS NOT CONCLUSIVE AND RESOLVED TO REINTERVENTION THE PATIENT. WHEN HE ENTERED THE CAVITY, HE NOTICED THAT THERE WAS A LEAK IN HIS STOMACH, CONCLUDING THAT THERE WAS A FISTULA BETWEEN THE 2ND AND 3RD STAPLED GREEN CHARGE AND GOLDEN CHARGE. THE TECHNIQUE WAS SLEEVE. HE PLACED A DRAIN ON THE PATIENT AND WILL FOLLOW THE ENDOSCOPIC TREATMENT.
16-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | PWRD ECH FLEX 45MM 44CM SHAFT | Leak/Splash | A0504 | Fistula | Event Description: IT WAS REPORTED THAT ON THE 2ND POSTOPERATIVE DAY, THE PATIENT WAS TACHYCARDIC, HAD AN IMAGING EXAM, THE SAME WAS NOT CONCLUSIVE AND RESOLVED TO REINTERVENTION THE PATIENT. WHEN HE ENTERED THE CAVITY, HE NOTICED THAT THERE WAS A LEAK IN HIS STOMACH, CONCLUDING THAT THERE WAS A FISTULA BETWEEN THE 2ND AND 3RD STAPLED GREEN CHARGE AND GOLDEN CHARGE. THE TECHNIQUE WAS SLEEVE. HE PLACED A DRAIN ON THE PATIENT AND WILL FOLLOW THE ENDOSCOPIC TREATMENT.
16-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | PWRD ECH FLEX 45MM 44CM SHAFT | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT ON THE 2ND POSTOPERATIVE DAY, THE PATIENT WAS TACHYCARDIC, HAD AN IMAGING EXAM, THE SAME WAS NOT CONCLUSIVE AND RESOLVED TO REINTERVENTION THE PATIENT. WHEN HE ENTERED THE CAVITY, HE NOTICED THAT THERE WAS A LEAK IN HIS STOMACH, CONCLUDING THAT THERE WAS A FISTULA BETWEEN THE 2ND AND 3RD STAPLED GREEN CHARGE AND GOLDEN CHARGE. THE TECHNIQUE WAS SLEEVE. HE PLACED A DRAIN ON THE PATIENT AND WILL FOLLOW THE ENDOSCOPIC TREATMENT.
04-01-2020 | Malfunction | covidien lp llc north haven | 17-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER ATTACHING THE RELOAD INTO THE STAPLER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DOCTOR WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE DOCTOR THEN CHECKED THE STAPLER IF THERE WAS SOMETHING WRONG BUT STILL COULD NOT FIRE THE DEVICE. THE DEVICE WAS REPLACED BY ANOTHER STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2020 | Malfunction | us surgical puerto rico | 17-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER ATTACHING THE RELOAD INTO THE STAPLER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DOCTOR WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE DOCTOR THEN CHECKED THE STAPLER IF THERE WAS SOMETHING WRONG BUT STILL COULD NOT FIRE THE DEVICE. THE DEVICE WAS REPLACED BY ANOTHER STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2020 | Malfunction | us surgical puerto rico | 17-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AFTER ATTACHING THE RELOAD INTO THE STAPLER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DOCTOR WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE DOCTOR THEN CHECKED THE STAPLER IF THERE WAS SOMETHING WRONG BUT STILL COULD NOT FIRE THE DEVICE. THE DEVICE WAS REPLACED BY ANOTHER STAPLER AND RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-10-2019 | Malfunction | us surgical puerto rico | 17-01-2020 | ENDO GIA ULTRA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE HANDLE BROKE AND WOULD NOT RELEASE THE RELOAD. THE PROCEDURE WAS DONE BY USING ANOTHER DEVICE. THERE WAS NO PATIENT INJURY. *MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-10 (EGIAUSTND) HANDLE: AFC ANALYSIS FINDING CODES USEGATKTIS US: EGIA HANDLE THICK TISSUE *MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-20 (SIGRADMT) RELOAD: AFC ANALYSIS FINDING CODES USEGLTKTIS US: EGIA RELOAD THICK TISSUE.
30-10-2019 | Malfunction | us surgical puerto rico | 17-01-2020 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE HANDLE BROKE AND WOULD NOT RELEASE THE RELOAD. THE PROCEDURE WAS DONE BY USING ANOTHER DEVICE. THERE WAS NO PATIENT INJURY. *MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-10 (EGIAUSTND) HANDLE: AFC ANALYSIS FINDING CODES USEGATKTIS US: EGIA HANDLE THICK TISSUE *MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-20 (SIGRADMT) RELOAD: AFC ANALYSIS FINDING CODES USEGLTKTIS US: EGIA RELOAD THICK TISSUE.
28-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ECHELON FLEX 60 LONG | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, THE FAILURE OCCURS DURING STAPLING PERFORMED FOR THE ANASTOMOSIS. WHEN TRYING TO OPEN THE JAWS OF THE STAPLER, THE MECHANISMS OF THE STAPLER DID NOT REACT AND THE DOCTOR REPORTS THAT HE WAS MAKING MOVEMENTS TO TRY TO EXTRACT IT WITH THE JAWS CLOSED, AT THIS TIME THEY PROVIDED AN EXTRA STAPLER. IT DID NOT SIGNIFICANTLY DELAY THE PROCEDURE. RESPONSE TIME BY TECHNICIANS WAS LESS THAN 5 MINUTES. IT DID NOT PUT THE PATIENT'S LIFE AT RISK.
26-10-2019 | Malfunction | covidien lp llc north haven | 17-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC GASTRIC BYPASS, WHILE STAPLING THE STOMACH FUNDUS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE DEVICE SOUNDED LIKE IT BEGAN TO FIRE BUT THEN STOPPED. UPON CHECKING THE RELOAD, NO STAPLES WERE DEPLOYED BUT STAPLE LEGS COULD BE SEEN IN THE APEX OF THE JAWS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC DONOR NEPHRECTOMY, THE KNIFE DID NOT RETURN TO HOME POSITION AFTER THE 2ND FIRING. THE DEVICE COULD BE REMOVED FROM THE PATIENT USING THE KNIFE REVERSE SWITCH. THE DEPLOYED STAPLES WERE FORMED AS INTENDED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Loss of or Failure to Bond | A040102 | Blood Loss | Event Description: IT WAS REPORTED THAT THE DEVICE WAS BEING USED DURING A ROBOTIC DONOR NEPHRECTOMY. WHEN FIRED ON THE RENAL VESSEL, THE VESSEL ON THE PATIENT SIDE DID NOT SEAL AND THE PATIENT LOST 8 LITRES OF BLOOD. THE CASE WAS IMMEDIATELY CONVERTED TO AN EMERGENCY LAPAROTOMY IN ORDER TO STOP THE BLEEDING. IT IS CURRENTLY UNSURE WHETHER THERE WAS MALFORMED STAPLES, OR NO STAPLES WERE DEPLOYED. ON THE KIDNEY SIDE THE STAPLES FORMED PROPERLY.
18-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: IT WAS REPORTED THAT THE DEVICE WAS BEING USED DURING A ROBOTIC DONOR NEPHRECTOMY. WHEN FIRED ON THE RENAL VESSEL, THE VESSEL ON THE PATIENT SIDE DID NOT SEAL AND THE PATIENT LOST 8 LITRES OF BLOOD. THE CASE WAS IMMEDIATELY CONVERTED TO AN EMERGENCY LAPAROTOMY IN ORDER TO STOP THE BLEEDING. IT IS CURRENTLY UNSURE WHETHER THERE WAS MALFORMED STAPLES, OR NO STAPLES WERE DEPLOYED. ON THE KIDNEY SIDE THE STAPLES FORMED PROPERLY.
29-11-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Activation, Positioning or Separation Problem | A15 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THERE WAS A WIDE GAP ON THE JAWS OF THE DEVICE AFTER FIRING. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
19-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED PLUS COMPACT 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LIVER RESECTION, THE STAPLES DIDN'T FORM. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
09-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | High Blood Pressure/ Hypertension | Event Description: IT WAS REPORTED THAT THE DOCTOR PERFORMED A LAPAROSCOPIC RIGHT COLECTOMY ON A PATIENT ON (B)(6) 2019. SHORTLY AFTER THE PROCEDURE, THE DOCTOR IDENTIFIED BLOOD IN THE PATIENT'S STOOL. THE PATIENT BECAME HYPERTENSIVE AND REQUIRED TWO BLOOD TRANSFUSIONS. THE BLEEDING WAS SUBSTANTIAL. THE DOCTOR INDICATED THERE WAS POST OPERATIVE BLEEDING FROM THE ANASTOMOSIS. SPECIFIC AMOUNT OF BLOOD LOSS WAS NOT REPORTED. THE DOCTOR INDICATED THE PATIENT WAS ON BLOOD THINNER MEDICATION BUT WAS UNCERTAIN WHETHER THE MEDICATION WAS TAKEN PRIOR TO SURGERY OR IF IT WAS ADMINISTERED POST OP. THE PATIENT DID RECEIVE A COLONOSCOPY RECENTLY, DATE UNKNOWN, AND THERE WAS NO LONGER BLEEDING IDENTIFIED. THE PATIENT STILL REMAINS ON POST OP IN THE HOSPITAL DAY TEN.
09-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT THE DOCTOR PERFORMED A LAPAROSCOPIC RIGHT COLECTOMY ON A PATIENT ON (B)(6) 2019. SHORTLY AFTER THE PROCEDURE, THE DOCTOR IDENTIFIED BLOOD IN THE PATIENT'S STOOL. THE PATIENT BECAME HYPERTENSIVE AND REQUIRED TWO BLOOD TRANSFUSIONS. THE BLEEDING WAS SUBSTANTIAL. THE DOCTOR INDICATED THERE WAS POST OPERATIVE BLEEDING FROM THE ANASTOMOSIS. SPECIFIC AMOUNT OF BLOOD LOSS WAS NOT REPORTED. THE DOCTOR INDICATED THE PATIENT WAS ON BLOOD THINNER MEDICATION BUT WAS UNCERTAIN WHETHER THE MEDICATION WAS TAKEN PRIOR TO SURGERY OR IF IT WAS ADMINISTERED POST OP. THE PATIENT DID RECEIVE A COLONOSCOPY RECENTLY, DATE UNKNOWN, AND THERE WAS NO LONGER BLEEDING IDENTIFIED. THE PATIENT STILL REMAINS ON POST OP IN THE HOSPITAL DAY TEN.
09-12-2019 | Injury | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT THE DOCTOR PERFORMED A LAPAROSCOPIC RIGHT COLECTOMY ON A PATIENT ON (B)(6) 2019. SHORTLY AFTER THE PROCEDURE, THE DOCTOR IDENTIFIED BLOOD IN THE PATIENT'S STOOL. THE PATIENT BECAME HYPERTENSIVE AND REQUIRED TWO BLOOD TRANSFUSIONS. THE BLEEDING WAS SUBSTANTIAL. THE DOCTOR INDICATED THERE WAS POST OPERATIVE BLEEDING FROM THE ANASTOMOSIS. SPECIFIC AMOUNT OF BLOOD LOSS WAS NOT REPORTED. THE DOCTOR INDICATED THE PATIENT WAS ON BLOOD THINNER MEDICATION BUT WAS UNCERTAIN WHETHER THE MEDICATION WAS TAKEN PRIOR TO SURGERY OR IF IT WAS ADMINISTERED POST OP. THE PATIENT DID RECEIVE A COLONOSCOPY RECENTLY, DATE UNKNOWN, AND THERE WAS NO LONGER BLEEDING IDENTIFIED. THE PATIENT STILL REMAINS ON POST OP IN THE HOSPITAL DAY TEN.
13-11-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A THORAX SURGERY, A PSEE60A WAS USED AND THE GST60D RELOAD BLOCKED THE STAPLER IN THE FIRST CUT. ANOTHER PSEE60A STAPLER WITH GREEN GST60G RELOAD WAS REQUESTED, WHICH ALSO LOCKED. THERE WAS NO DAMAGE TO THE PATIENT.
01-11-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE REP THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER WOULD NOT FIRE. NO PATIENT CONSEQUENCES REPORTED. THERE IS NO ADDITIONAL INFORMATION AVAILABLE AT THIS TIME.
11-11-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY, GUN LOCKED WITH IN CASE. OPENED WITH LOT OF EFFORT. THE SURGERY WAS DELAYED BY 20-MINUTES. USED A NEW GUN TO SUCCESSFULLY COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES.
27-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED PLUS STANDARD 45 | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS FOR RIGHT PARTIAL LUNG RESECTION, THE PSEE45A LOADING GST60T STOPPED ON THE WAY OF THE FIRST FIRING ON THE LUNG PARENCHYMA. THE KNIFE STOPPED ABOUT 15MM AWAY FROM THE PROXIMAL END OF THE CARTRIDGE. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE SINCE THE KNIFE REVERSE SWITCH DID NOT FUNCTION. A PSEE60A WAS USED FOLLOWING THE EVENT. WHEN THE SALES REP CHECKED THE PSEE45A AFTER THE PROCEDURE, IT WAS FOUND THAT A PART OF THE GROOVE OF THE ANVIL JAW WHERE THE KNIFE MOVED THROUGH WAS TURNED UP. THE ISSUE WAS OBSERVED WHERE THE EVENT OCCURRED OF THE KNIFE. THE SURGEON¿S COMMENT IS FOLLOWING. THE TARGET TISSUE WAS STIFF AND THICK. THERE WAS A POSSIBILITY THAT THE DEVICE WAS CLAMPED THE TARGET TISSUE MUCH AT THE PROXIMAL END OF THE JAWS. THE MOTOR SOUND NEARLY STOPPED AND THE KNIFE MANAGED TO MOVE FORWARD AT THE 3RD FIRING FOR COMPLETION THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
27-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED PLUS STANDARD 45 | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS FOR RIGHT PARTIAL LUNG RESECTION, THE PSEE45A LOADING GST60T STOPPED ON THE WAY OF THE FIRST FIRING ON THE LUNG PARENCHYMA. THE KNIFE STOPPED ABOUT 15MM AWAY FROM THE PROXIMAL END OF THE CARTRIDGE. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE SINCE THE KNIFE REVERSE SWITCH DID NOT FUNCTION. A PSEE60A WAS USED FOLLOWING THE EVENT. WHEN THE SALES REP CHECKED THE PSEE45A AFTER THE PROCEDURE, IT WAS FOUND THAT A PART OF THE GROOVE OF THE ANVIL JAW WHERE THE KNIFE MOVED THROUGH WAS TURNED UP. THE ISSUE WAS OBSERVED WHERE THE EVENT OCCURRED OF THE KNIFE. THE SURGEON¿S COMMENT IS FOLLOWING. THE TARGET TISSUE WAS STIFF AND THICK. THERE WAS A POSSIBILITY THAT THE DEVICE WAS CLAMPED THE TARGET TISSUE MUCH AT THE PROXIMAL END OF THE JAWS. THE MOTOR SOUND NEARLY STOPPED AND THE KNIFE MANAGED TO MOVE FORWARD AT THE 3RD FIRING FOR COMPLETION THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
07-10-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED BLEEDING 2 DAYS AFTER THE SURGICAL PROCEDURE, PRESENTING FISTULA. THE PATIENT WAS READMITTED AND REOPERATED ON (B)(6) 2019. IT WAS NECESSARY TO USE A DRAIN AND A NASOENTERIC TUBE. PATIENT HAD TO DO A PARENTERAL DIET. PATIENT IS WELL AND AT HOME.
07-10-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED BLEEDING 2 DAYS AFTER THE SURGICAL PROCEDURE, PRESENTING FISTULA. THE PATIENT WAS READMITTED AND REOPERATED ON (B)(6) 2019. IT WAS NECESSARY TO USE A DRAIN AND A NASOENTERIC TUBE. PATIENT HAD TO DO A PARENTERAL DIET. PATIENT IS WELL AND AT HOME.
07-10-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED BLEEDING 2 DAYS AFTER THE SURGICAL PROCEDURE, PRESENTING FISTULA. THE PATIENT WAS READMITTED AND REOPERATED ON (B)(6) 2019. IT WAS NECESSARY TO USE A DRAIN AND A NASOENTERIC TUBE. PATIENT HAD TO DO A PARENTERAL DIET. PATIENT IS WELL AND AT HOME.
07-10-2019 | Injury | ethicon endo-surgery | 17-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED BLEEDING 2 DAYS AFTER THE SURGICAL PROCEDURE, PRESENTING FISTULA. THE PATIENT WAS READMITTED AND REOPERATED ON (B)(6) 2019. IT WAS NECESSARY TO USE A DRAIN AND A NASOENTERIC TUBE. PATIENT HAD TO DO A PARENTERAL DIET. PATIENT IS WELL AND AT HOME.
27-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS FOR RIGHT PARTIAL LUNG RESECTION, THE PSEE60A LOADING GST60T WAS USED. THE KNIFE STOPPED ABOUT 20MM AWAY FROM THE PROXIMAL END OF THE CARTRIDGE. THE SURGEON USED THE MANUAL OVERRIDE LEVER SINCE THE KNIFE REVERSE SWITCH DID NOT FUNCTION. BUT THE PROXIMAL END PART OF THE MANUAL OVERRIDE LEVER WAS BROKEN WHILE MOVING THE LEVER, AND THE LEVER DID NOT FUNCTION. THEREFORE, THE SURGEON INSERTED A FORCEPS BETWEEN THE JAWS, AND THE JAWS WERE OPENED FORCIBLY TO REMOVE THE DEVICE FROM THE PATIENT. ANOTHER PSEE60A LOADING ANOTHER GST60T WAS USED TO COMPLETE THE CASE. WHEN THE SALES REP CHECKED THE PSEE60A AFTER THE PROCEDURE, IT WAS FOUND THAT A PART OF THE GROOVE OF THE ANVIL JAW WHERE THE KNIFE MOVED THROUGH WAS TURNED UP. THE ISSUE WAS OBSERVED WHERE THE EVENT OCCURRED OF THE KNIFE. THE SURGEON¿S COMMENT IS FOLLOWING. THE TARGET TISSUE WAS STIFF AND THICK. THERE WAS A POSSIBILITY THAT THE DEVICE WAS CLAMPED THE TARGET TISSUE MUCH AT THE PROXIMAL END OF THE JAWS. THE MOTOR SOUND NEARLY STOPPED AND THE KNIFE MANAGED TO MOVE FORWARD AT THE 3RD FIRING FOR COMPLETION THE CASE. THERE WAS NO PROBLEM FOR THE PATIENT WHEN THE SURGEON OPENED THE JAW WITH THE FORCEPS. THE PATIENT IS STABLE AND IS DISCHARGED ON SCHEDULE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
27-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS FOR RIGHT PARTIAL LUNG RESECTION, THE PSEE60A LOADING GST60T WAS USED. THE KNIFE STOPPED ABOUT 20MM AWAY FROM THE PROXIMAL END OF THE CARTRIDGE. THE SURGEON USED THE MANUAL OVERRIDE LEVER SINCE THE KNIFE REVERSE SWITCH DID NOT FUNCTION. BUT THE PROXIMAL END PART OF THE MANUAL OVERRIDE LEVER WAS BROKEN WHILE MOVING THE LEVER, AND THE LEVER DID NOT FUNCTION. THEREFORE, THE SURGEON INSERTED A FORCEPS BETWEEN THE JAWS, AND THE JAWS WERE OPENED FORCIBLY TO REMOVE THE DEVICE FROM THE PATIENT. ANOTHER PSEE60A LOADING ANOTHER GST60T WAS USED TO COMPLETE THE CASE. WHEN THE SALES REP CHECKED THE PSEE60A AFTER THE PROCEDURE, IT WAS FOUND THAT A PART OF THE GROOVE OF THE ANVIL JAW WHERE THE KNIFE MOVED THROUGH WAS TURNED UP. THE ISSUE WAS OBSERVED WHERE THE EVENT OCCURRED OF THE KNIFE. THE SURGEON¿S COMMENT IS FOLLOWING. THE TARGET TISSUE WAS STIFF AND THICK. THERE WAS A POSSIBILITY THAT THE DEVICE WAS CLAMPED THE TARGET TISSUE MUCH AT THE PROXIMAL END OF THE JAWS. THE MOTOR SOUND NEARLY STOPPED AND THE KNIFE MANAGED TO MOVE FORWARD AT THE 3RD FIRING FOR COMPLETION THE CASE. THERE WAS NO PROBLEM FOR THE PATIENT WHEN THE SURGEON OPENED THE JAW WITH THE FORCEPS. THE PATIENT IS STABLE AND IS DISCHARGED ON SCHEDULE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
13-11-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED PLUS COMPACT 45 | Energy Output Problem | A0904 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE HAD BEEN FIRED TWICE ALREADY. WHEN FIRING THE 3RD RELOAD IT WAS LIKE THE BATTERY HAD RUN OUT OF POWER. REVERSE BUTTON DID NOT WORK. MANUAL OVERRIDE WAS USED. THERE WAS NO HARM TO THE PATIENT.
13-11-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED PLUS COMPACT 45 | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE HAD BEEN FIRED TWICE ALREADY. WHEN FIRING THE 3RD RELOAD IT WAS LIKE THE BATTERY HAD RUN OUT OF POWER. REVERSE BUTTON DID NOT WORK. MANUAL OVERRIDE WAS USED. THERE WAS NO HARM TO THE PATIENT.
12-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRO-INTESTINAL PROCEDURE, THE SURGEON WAS ATTEMPTING TO PULSE FIRE BUT THE STAPLER DID A CONTINUOUS FIRE. ONCE THE STAPLER FIRED, THE JAWS WOULD NOT OPEN AND THE MANUAL OVERRIDE HAD TO BE USED. UPON OBSERVATION, ALL THE STAPLES SEEMED FORMED PROPERLY AND THE CUT LINE LOOKED GOOD. A SECOND STAPLER WAS OPENED TO COMPLETE THE CASE. I BELIEVE IT WAS THE SECOND FIRING WHEN THE INCIDENT OCCURRED. NO HARM TO THE PATIENT.
12-11-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: ON THE (B)(6) 2019 THE SURGEONS WERE PERFORMING HEMICOLECTOMY DEXTER. DURING THE PROCEDURE THEY HAVE USED PLEE60AU, LOT T9263A. WHILE FIRING THE PRODUCT STARTED STAPLING AND CUTTING, BUT AFTER A FEW CENTIMETERS STOPPED. THEY MANUALLY RETURNED THE KNIFE INTO THE POSITION SO THEY CAN RELEASE THE TISSUE FROM THE JAW. AFTER THEY USED THE SAME DEVICE, REPEATED THE PROCEDURE AND FINISHED IT. THERE WAS NO ADVERSE EVENT OR DELAY IN THE SURGERY.
01-11-2018 | Injury | ethicon endo-surgery | 17-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE DOCTOR FIRED THE STAPLER, WITH A BLUE RELOAD, NO BUTTRESSING MATERIAL AND THE DOCTOR REPORTED THE STAPLER FELT FUNNY AFTER FIRING. AFTER THE DOCTOR REMOVED THE STAPLER FROM TISSUE, THE DOCTOR NOTICED THERE WERE LOOSE STAPLES ALONG THE STAPLE LINE. THE DOCTOR DECIDED TO CONVERT TO AN OPEN PROCEDURE. THE LOOSE STAPLES WERE REMOVED PRIOR TO THE DOCTOR DECIDING TO CONVERT TO AN OPEN PROCEDURE, TO COMPLETE THE PROCEDURE. THERE WERE NO OTHER CONSEQUENCES REPORTED.
01-11-2018 | Injury | ethicon endo-surgery | 17-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE DOCTOR FIRED THE STAPLER, WITH A BLUE RELOAD, NO BUTTRESSING MATERIAL AND THE DOCTOR REPORTED THE STAPLER FELT FUNNY AFTER FIRING. AFTER THE DOCTOR REMOVED THE STAPLER FROM TISSUE, THE DOCTOR NOTICED THERE WERE LOOSE STAPLES ALONG THE STAPLE LINE. THE DOCTOR DECIDED TO CONVERT TO AN OPEN PROCEDURE. THE LOOSE STAPLES WERE REMOVED PRIOR TO THE DOCTOR DECIDING TO CONVERT TO AN OPEN PROCEDURE, TO COMPLETE THE PROCEDURE. THERE WERE NO OTHER CONSEQUENCES REPORTED.
04-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | GST GOLD RELOAD, 60MM, 6 ROW | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, WHEN THE FIRST SHOT IS MADE, THE METAL PART IS RELEASED AND THE SECOND RECHARGE IS NOT SHOT, A NEW ONE IS PASSED AND IT WORKS, THERE ARE NO FRAGMENTS LEFT INSIDE THE PATIENT. SURGERY WAS DELAYED 5 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED AND NO PATIENT CONSEQUENCES WERE REPORTED.
04-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, WHEN THE FIRST SHOT IS MADE, THE METAL PART IS RELEASED AND THE SECOND RECHARGE IS NOT SHOT, A NEW ONE IS PASSED AND IT WORKS, THERE ARE NO FRAGMENTS LEFT INSIDE THE PATIENT. SURGERY WAS DELAYED 5 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED AND NO PATIENT CONSEQUENCES WERE REPORTED.
04-11-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | FLEX 60 ARTICULATING | Crack | A0404 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, ECHELON 60 CRACKED WHEN THE 6TH RELOAD WAS USED. RELOAD DID NOT FIRE BECAUSE THE STAPLE DID NOT CLOSE.
01-11-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED PLUS COMPACT 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, THE SURGEON HAD CLAMPED THE GUN AROUND LUNG TISSUE AND FIRED THE PRODUCT. AFTER FIRING WAS COMPLETED HE NOTICED THE TISSUE WAS STAPLED BUT NO STAPLES DEPLOYED DURING FIRING. THE SURGEON PUT THE DEVICE ASIDE AND USED ANOTHER PRODUCT TO COMPLETE THE PROCEDURE. NO DAMAGE WAS DONE TO THE PATIENT.
27-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER KEPT DROPPING LOAD CARTRIDGES. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
25-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE SLEEVE GASTRECTOMY SURGERY, WHEN SEVERING STOMACH TISSUE ON THE FIRST FIRING, AFTER THE DEVICE WAS FIRED, IT WAS NOTED THAT STAPLES WERE MALFORMED. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. NOTED THAT BLADE COULD NOT MOVE, BOTH KNIFE REVERSE BUTTON AND MANUAL OVERRIDE LEVER COULD NOT WORK. THE RENAL PEDICLE WAS AMPUTATED IN THE TRADITIONAL WAY AND THEN THE DEVICE WAS FORCIBLY REMOVED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. NOTED THAT BLADE COULD NOT MOVE, BOTH KNIFE REVERSE BUTTON AND MANUAL OVERRIDE LEVER COULD NOT WORK. THE RENAL PEDICLE WAS AMPUTATED IN THE TRADITIONAL WAY AND THEN THE DEVICE WAS FORCIBLY REMOVED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT A SURGEON TRIED TO USE THE ATS45 STAPLER FOR A COLORECTAL PROCEDURE, HOWEVER, THE STAPLER DID NOT FIRE AND A NEW ATS45 STAPLER WAS OPENED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
25-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC LOBECTOMY SURGERY, AFTER THE DEVICE WAS FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. CHANGED TO ANOTHER TO COMPLETE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED PLUS COMPACT 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, NURSE ASSEMBLED CARTRIDGE TO BODY AND SURGEON FIRED TO TARGET TISSUE. KNIFE MOVED FORWARD WELL TO JAW'S DISTAL POINT BUT WHEN COMING BACK, KNIFE WAS STOPPED AND IT WAS NOT MOVED ANYMORE. SURGEON MANUALLY FORCED TO KNIFE BACK. WHEN OPEN THE JAW, SURGEON SAW EACH CLIPS WERE NOT DEPLOYED ON TISSUE WELL ALSO IT WAS NOT MADE FULLY B-SHAPE.
29-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING NEPHRECTOMY SURGERY, AN ECHELON GST WAS APPLIED ON RENAL VESSELS WITH WHITE CARTRIDGE. THE DEVICE STOPPED AT THE MID-FIRING, DID NOT OPEN AFTER REVERSE BUTTON WAS PUSHED, DID NOT OPEN AFTER MANUAL OVERRIDE WAS USED. AFTER DISSECTION THE RENAL VESSELS WERE CLIPPED AND CUT AROUND THE DEVICE. IT WAS PULLED OUT OF THE PATIENT WITH NO OTHER CONSEQUENCES. THERE WAS A SUSPICION THAT THE DEVICE WAS LOCKED ON A PART OF A SUCTION ROD BUT THAT (SUCTION) WAS PULLED OUT AND WAS COMPLETE. THE SURGERY WAS DELAYED 10 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED AND THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Peeled/Delaminated | A040506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT THE PLASTIC ON THE DECK OF THE VASECR35 WAS PEELING UP. NEW RELOAD WAS USED WAS USED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCES REPORTED.
09-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED PLUS COMPACT 45 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT STAPLER BECAME STUCK DURING THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
09-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | POWERED PLUS COMPACT 45 | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT STAPLER BECAME STUCK DURING THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
18-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN USING THE FLEX STAPLER INTRAOPERATIVELY, IT WAS NOTICED THAT ONE OF THE DEVICES TO TRIGGER THE RELOAD WAS DAMAGED. NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
18-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Firing Problem | A0505 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT WHEN USING THE FLEX STAPLER INTRAOPERATIVELY, IT WAS NOTICED THAT ONE OF THE DEVICES TO TRIGGER THE RELOAD WAS DAMAGED. NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
17-12-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, THE TRIGGER WAS PULLED TO FIRE THE STAPLER, BUT THE KNIFE DID NOT MOVE FORWARD AND STUCK IN THE MIDDLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
01-01-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE DEVICE WOULD NOT OPEN AFTER IT WAS INTRODUCED INTO THE ABDOMEN. NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED AT THIS TIME.
01-01-2019 | Malfunction | ethicon endo-surgery | 17-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER KEPT DROPPING LOADED CARTRIDGES. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. NO PATIENT CONSEQUENCES WERE REPORTED.
20-12-2019 | Injury | ethicon | 18-01-2020 | ETHICON SECURESTRAP OPEN | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND THE ABSORBABLE STRAPS WERE USED. IT WAS REPORTED THAT THE SHARP TRACKER EDGES CAUSED PATIENT BLEEDING FROM THE FIXATION POINT, AND THE SURGEON HAD TO RE-OPERATE THE PATIENT.
20-12-2019 | Injury | ethicon | 18-01-2020 | ETHICON SECURESTRAP OPEN | Adverse Event Without Identified Device or Use Problem | A24 | Not Applicable | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND THE ABSORBABLE STRAPS WERE USED. IT WAS REPORTED THAT THE SHARP TRACKER EDGES CAUSED PATIENT BLEEDING FROM THE FIXATION POINT, AND THE SURGEON HAD TO RE-OPERATE THE PATIENT.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | TA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE ATTEMPT TO STAPLE THE SIGMOID COLON AND CLOSE A COMMON ENTEROTOMY ON AN OPEN EXPLORATORY LAPAROTOMY, SIGMOIDECTOMY, THE STAPLER DID NOT CLOSE PROPERLY ALL THE WAY ENOUGH TO MEET THE NECESSARY TISSUE GAP CONTROL CLOSURE. THE STAPLER WAS FIRED AND STAPLES DEPLOYED, BUT NONE OF THE STAPLES WERE FORMED UPON INSPECTION. THE STAPLER WAS RELOADED AND IT WAS DISCOVERED THAT EVEN WITH A NEW RELOAD, STAPLER COULD NOT CLOSE WITHIN 2CM OF THE ANVIL. ANOTHER OPEN STAPLER WAS OPENED AND USED TO STAPLE AND COMPLETE PROCEDURE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | TA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THE ATTEMPT TO STAPLE THE SIGMOID COLON AND CLOSE A COMMON ENTEROTOMY ON AN OPEN EXPLORATORY LAPAROTOMY, SIGMOIDECTOMY, THE STAPLER DID NOT CLOSE PROPERLY ALL THE WAY ENOUGH TO MEET THE NECESSARY TISSUE GAP CONTROL CLOSURE. THE STAPLER WAS FIRED AND STAPLES DEPLOYED, BUT NONE OF THE STAPLES WERE FORMED UPON INSPECTION. THE STAPLER WAS RELOADED AND IT WAS DISCOVERED THAT EVEN WITH A NEW RELOAD, STAPLER COULD NOT CLOSE WITHIN 2CM OF THE ANVIL. ANOTHER OPEN STAPLER WAS OPENED AND USED TO STAPLE AND COMPLETE PROCEDURE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A MINIMALLY INVASIVE LOBECTOMY, WHEN STAPLING ACROSS THE PULMONARY ARTERY, THE STAPLER WAS A FULL 45 DEGREES ARTICULATION, NO ISSUE WITH FIRING BUT THE SURGEON HAD A HARD TIME RETRACTING AND OPENING UP THE STAPLER. THEY WERE EVENTUALLY ABLE TO RETRACT WITHOUT DAMAGE. ON THE NEXT FIRING, THE SURGEON OPENED UP A NEW HANDLE TO FINISH CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A MINIMALLY INVASIVE LOBECTOMY, WHEN STAPLING ACROSS THE PULMONARY ARTERY, THE STAPLER WAS A FULL 45 DEGREES ARTICULATION, NO ISSUE WITH FIRING BUT THE SURGEON HAD A HARD TIME RETRACTING AND OPENING UP THE STAPLER. THEY WERE EVENTUALLY ABLE TO RETRACT WITHOUT DAMAGE. ON THE NEXT FIRING, THE SURGEON OPENED UP A NEW HANDLE TO FINISH CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A MINIMALLY INVASIVE LOBECTOMY, WHEN STAPLING ACROSS THE PULMONARY ARTERY, THE STAPLER WAS A FULL 45 DEGREES ARTICULATION, NO ISSUE WITH FIRING BUT THE SURGEON HAD A HARD TIME RETRACTING AND OPENING UP THE STAPLER. THEY WERE EVENTUALLY ABLE TO RETRACT WITHOUT DAMAGE. ON THE NEXT FIRING, THE SURGEON OPENED UP A NEW HANDLE TO FINISH CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A MINIMALLY INVASIVE LOBECTOMY, WHEN STAPLING ACROSS THE PULMONARY ARTERY, THE STAPLER WAS A FULL 45 DEGREES ARTICULATION, NO ISSUE WITH FIRING BUT THE SURGEON HAD A HARD TIME RETRACTING AND OPENING UP THE STAPLER. THEY WERE EVENTUALLY ABLE TO RETRACT WITHOUT DAMAGE. ON THE NEXT FIRING, THE SURGEON OPENED UP A NEW HANDLE TO FINISH CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A MINIMALLY INVASIVE LOBECTOMY, WHEN STAPLING ACROSS THE PULMONARY ARTERY, THE STAPLER WAS A FULL 45 DEGREES ARTICULATION, NO ISSUE WITH FIRING BUT THE SURGEON HAD A HARD TIME RETRACTING AND OPENING UP THE STAPLER. THEY WERE EVENTUALLY ABLE TO RETRACT WITHOUT DAMAGE. ON THE NEXT FIRING, THE SURGEON OPENED UP A NEW HANDLE TO FINISH CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A MINIMALLY INVASIVE LOBECTOMY, WHEN STAPLING ACROSS THE PULMONARY ARTERY, THE STAPLER WAS A FULL 45 DEGREES ARTICULATION, NO ISSUE WITH FIRING BUT THE SURGEON HAD A HARD TIME RETRACTING AND OPENING UP THE STAPLER. THEY WERE EVENTUALLY ABLE TO RETRACT WITHOUT DAMAGE. ON THE NEXT FIRING, THE SURGEON OPENED UP A NEW HANDLE TO FINISH CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | SIGNIA | Break | A0401 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LIVER LOBECTOMY, NOTHING WAS DISPLAYED ON THE HANDLE. IT SEEMED TO BE THAT THE DEVICE BLACKED OUT, BUT AFTER PERFORMING A POWERED RESET, A TONE WAS HEARD AS USUAL AND IT SEEMED THAT THE DEVICE POWERED ON, BUT NOTHING WAS DISPLAYED. IT WAS FOUND THAT THE GLASS INSIDE THE DISPLAY WAS BROKEN WHEN CHECKING THE DEVICE CAREFULLY. A REPLACEMENT WAS PROVIDED TO RESOLVE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
07-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | SIGNIA | Device Misassembled During Manufacturing /Shipping | A0206 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LIVER LOBECTOMY, NOTHING WAS DISPLAYED ON THE HANDLE. IT SEEMED TO BE THAT THE DEVICE BLACKED OUT, BUT AFTER PERFORMING A POWERED RESET, A TONE WAS HEARD AS USUAL AND IT SEEMED THAT THE DEVICE POWERED ON, BUT NOTHING WAS DISPLAYED. IT WAS FOUND THAT THE GLASS INSIDE THE DISPLAY WAS BROKEN WHEN CHECKING THE DEVICE CAREFULLY. A REPLACEMENT WAS PROVIDED TO RESOLVE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT THE DEVICE LOCKED AND COULD NOT SQUEEZE THE HANDLE. THE SURGEON TRIED TO USE ANOTHER HANDLE HOWEVER, SAME ISSUE OCCURRED. THE SURGEON THEN REPLACED AGAIN THE HANDLE AND WORKED WITH THE ORIGINAL LOADS THAT DID NOT WORK ON THE FIRST TWO HANDLE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION SURGERY, UPON RECTAL RESECTION, THE JAWS COULD NOT BE CLOSED AFTER APPROACHING THE TISSUE. WHEN THE RELOAD WAS REPLACED, THE RESECTION WAS PERFORMED WITHOUT PROBLEM. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, THE ADAPTER WAS CONNECTED TO THE POWER HANDLE AND THERE WAS NO ISSUE. THE SURGEON WAS ALSO ABLE TO DO THE CLAMPING TEST AND IT WAS INDICATED THAT THE DISPLAY WENT GREEN. THE TISSUE WAS THEN CLAMPED AND FIRING WAS ATTEMPTED, BUT THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Injury | covidien lp llc north haven | 20-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT SUPERIOR LOBECTOMY, THE STAPLER WAS ABLE TO FIRE, THERE WAS POOR STAPLE FORMATION IN THE MIDDLE PART OF THE RELOAD. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THEY SUTURED TO FIX THE STAPLE LINE.
07-01-2020 | Injury | covidien lp llc north haven | 20-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT SUPERIOR LOBECTOMY, THE STAPLER WAS ABLE TO FIRE, THERE WAS POOR STAPLE FORMATION IN THE MIDDLE PART OF THE RELOAD. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THEY SUTURED TO FIX THE STAPLE LINE.
07-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC SEGMENTECTOMY AND WEDGE RESECTION, WHILE REMOVING THE LUNG PARENCHYMA, AFTER FIRING, IT WAS NOTED THAT THERE WAS POOR STAPLE FORMATION. SOME OF THE STAPLES ARE NOT FORMED OR ARE POORLY FORMED. SURGEON USED ANOTHER RELOAD TO COMPLETE THE CASE. NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC SEGMENTECTOMY AND WEDGE RESECTION, WHILE REMOVING THE LUNG PARENCHYMA, AFTER FIRING, IT WAS NOTED THAT THERE WAS POOR STAPLE FORMATION. SOME OF THE STAPLES ARE NOT FORMED OR ARE POORLY FORMED. SURGEON USED ANOTHER RELOAD TO COMPLETE THE CASE. NO PATIENT INJURY.
07-01-2020 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Misfire | A050502 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY PROCEDURE, WHILE WORKING ON THE RIGHT COLON, THE DEVICE PARTIALLY FIRED BY 50% THEN IT JAMMED. THE STAPLE LINE WAS INCOMPLETE. STAPLE LINE WAS FIXED BY DOING RE-SECTION AND RE-STAPLED USING A NEW STAPLER. COLON WAS TRANSECTED PROXIMAL TO INITIAL TRANSECTION.
07-01-2020 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY PROCEDURE, WHILE WORKING ON THE RIGHT COLON, THE DEVICE PARTIALLY FIRED BY 50% THEN IT JAMMED. THE STAPLE LINE WAS INCOMPLETE. STAPLE LINE WAS FIXED BY DOING RE-SECTION AND RE-STAPLED USING A NEW STAPLER. COLON WAS TRANSECTED PROXIMAL TO INITIAL TRANSECTION.
07-01-2020 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY PROCEDURE, WHILE WORKING ON THE RIGHT COLON, THE DEVICE PARTIALLY FIRED BY 50% THEN IT JAMMED. THE STAPLE LINE WAS INCOMPLETE. STAPLE LINE WAS FIXED BY DOING RE-SECTION AND RE-STAPLED USING A NEW STAPLER. COLON WAS TRANSECTED PROXIMAL TO INITIAL TRANSECTION.
07-01-2020 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY PROCEDURE, WHILE WORKING ON THE RIGHT COLON, THE DEVICE PARTIALLY FIRED BY 50% THEN IT JAMMED. THE STAPLE LINE WAS INCOMPLETE. STAPLE LINE WAS FIXED BY DOING RE-SECTION AND RE-STAPLED USING A NEW STAPLER. COLON WAS TRANSECTED PROXIMAL TO INITIAL TRANSECTION.
07-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | SIGNIA | Failure to Charge | A070603 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE ON A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY PROCEDURE, THE POWER HANDLE DOES NOT TURN ON AND IT'S DISPLAY WAS SHUT OFF. THE DEVICE DID NOT CHARGE. ANOTHER POWER HANDLE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INVOLVEMENT.
07-01-2020 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY PROCEDURE, WHILE WORKING ON THE RIGHT COLON, THE DEVICE PARTIALLY FIRED BY 50% THEN IT JAMMED. THE STAPLE LINE WAS INCOMPLETE. STAPLE LINE WAS FIXED BY DOING RE-SECTION AND RE-STAPLED USING A NEW STAPLER. COLON WAS TRANSECTED PROXIMAL TO INITIAL TRANSECTION.
07-01-2020 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY PROCEDURE, WHILE WORKING ON THE RIGHT COLON, THE DEVICE PARTIALLY FIRED BY 50% THEN IT JAMMED. THE STAPLE LINE WAS INCOMPLETE. STAPLE LINE WAS FIXED BY DOING RE-SECTION AND RE-STAPLED USING A NEW STAPLER. COLON WAS TRANSECTED PROXIMAL TO INITIAL TRANSECTION.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT THORACOTOMY, RIGHT UPPER LOBE LUNG RESECTION PROCEDURE, THE DEVICE DID NOT FIRE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. TO RESOLVE THE ISSUE THE SURGEON HAD TO RETRACT THE BLACK KNOB AND USED ANOTHER HANDLE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT THORACOTOMY, RIGHT UPPER LOBE LUNG RESECTION PROCEDURE, THE DEVICE DID NOT FIRE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. TO RESOLVE THE ISSUE THE SURGEON HAD TO RETRACT THE BLACK KNOB AND USED ANOTHER HANDLE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY END TO END ANASTOMOSIS, THE HANDLE WAS SQUEEZED BUT FIRING COULD NOT BE PERFORMED AT ALL. UPON CHECKING THE CARTRIDGE, THE STAPLE WAS NOT FOUND TO BE EXPOSED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ANTHRO RESECTION ON A LAPAROSCOPIC GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY END TO END ANASTOMOSIS, THE HANDLE WAS SQUEEZED BUT FIRING COULD NOT BE PERFORMED AT ALL. UPON CHECKING THE CARTRIDGE, THE STAPLE WAS NOT FOUND TO BE EXPOSED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
03-12-2019 | Malfunction | covidien lp llc north haven | 20-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF FIRING DURING A LAPAROSCOPIC VIDEO GASTROPLASTY PROCEDURE WHEN THE DEVICE WAS BEING APPLIED IN THE TISSUE, THE DEVICE LOCKED AND DID NOT FIRE. THE RELOAD WAS WITHDRAWN AND PLACED AGAIN ON THE STAPLER AND THE ERROR CONTINUED. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-04-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, AN ABNORMAL SOUND WAS HEARD ON THE HANDLE. THE DEVICE WAS UNABLE TO ADVANCE AND COULD NOT BE FIRED. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | covidien lp llc north haven | 20-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY PROCEDURE, AT FOURTH FIRING, THE DEVICE WAS UNABLE TO BE FIRED. IT WAS ALSO REPORTED THAT THE SURGEON DID NOT SEE ANY ERROR MESSAGE DISPLAYED ON THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
22-04-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, AN ABNORMAL SOUND WAS HEARD ON THE HANDLE. THE DEVICE WAS UNABLE TO ADVANCE AND COULD NOT BE FIRED. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-04-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, AN ABNORMAL SOUND WAS HEARD ON THE HANDLE. THE DEVICE WAS UNABLE TO ADVANCE AND COULD NOT BE FIRED. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | covidien lp llc north haven | 20-01-2020 | SIGNIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY PROCEDURE, AT FOURTH FIRING, THE DEVICE WAS UNABLE TO BE FIRED. IT WAS ALSO REPORTED THAT THE SURGEON DID NOT SEE ANY ERROR MESSAGE DISPLAYED ON THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | covidien lp llc north haven | 20-01-2020 | SIGNIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY PROCEDURE, AT FOURTH FIRING, THE DEVICE WAS UNABLE TO BE FIRED. IT WAS ALSO REPORTED THAT THE SURGEON DID NOT SEE ANY ERROR MESSAGE DISPLAYED ON THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
09-12-2019 | Malfunction | davol inc sub c.r bard | 20-01-2020 | SORBAFIX ABSORBABLE FIXATION SYSTEM | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT AFTER TWO FASTENERS WERE FIRED FROM THE SORBAFIX ABSORBABLE FIXATION SYSTEM, A PIECE OF FOREIGN MATTER FELL DOWN FROM THE LOCATION OF THE SPANNER (TRIGGER) AND NO MORE FASTENERS COULD BE FIRED. THE SORBAFIX ABSORBABLE FIXATION SYSTEM WAS REPLACED WITH A NEW ONE. THERE WAS NO REPORTED PATIENT INJURY.
09-12-2019 | Malfunction | us surgical puerto rico | 20-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO GASTROPLASTY PROCEDURE, THE PURPLE LOAD FAILED AT THE BEGINNING OF FIRING AND DID NOT PROGRESS. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
09-12-2019 | Malfunction | us surgical puerto rico | 20-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO GASTROPLASTY PROCEDURE, THE PURPLE LOAD FAILED AT THE BEGINNING OF FIRING AND DID NOT PROGRESS. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | us surgical puerto rico | 20-01-2020 | ILA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROXIMAL GASTRECTOMY, WHEN USING THE LOADING UNIT, IT WAS TOO STIFF AND THE KNIFE DID NOT MOVE. THE SURGEON TRIED TO FIRE ON NOTHING WITHOUT TISSUE CLAMPED, BUT THE KNIFE STILL DID NOT MOVE AT ALL. WHEN USING ANOTHER LOADING UNIT, THE KNIFE BENT AND COULD NOT CUT AT ALL HOWEVER, THE STAPLES HAVE BEEN PROPERLY FORMED UNTIL THE END. THEY WERE NOT ABLE TO CUT WITH THE KNIFE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | us surgical puerto rico | 20-01-2020 | ILA | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROXIMAL GASTRECTOMY, WHEN USING THE LOADING UNIT, IT WAS TOO STIFF AND THE KNIFE DID NOT MOVE. THE SURGEON TRIED TO FIRE ON NOTHING WITHOUT TISSUE CLAMPED, BUT THE KNIFE STILL DID NOT MOVE AT ALL. WHEN USING ANOTHER LOADING UNIT, THE KNIFE BENT AND COULD NOT CUT AT ALL HOWEVER, THE STAPLES HAVE BEEN PROPERLY FORMED UNTIL THE END. THEY WERE NOT ABLE TO CUT WITH THE KNIFE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | us surgical puerto rico | 20-01-2020 | ILA | Material Twisted/Bent | A040609 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING PROXIMAL GASTRECTOMY, WHEN USING THE LOADING UNIT, IT WAS TOO STIFF AND THE KNIFE DID NOT MOVE. THE SURGEON TRIED TO FIRE ON NOTHING WITHOUT TISSUE CLAMPED, BUT THE KNIFE STILL DID NOT MOVE AT ALL. WHEN USING ANOTHER LOADING UNIT, THE KNIFE BENT AND COULD NOT CUT AT ALL HOWEVER, THE STAPLES HAVE BEEN PROPERLY FORMED UNTIL THE END. THEY WERE NOT ABLE TO CUT WITH THE KNIFE. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
17-04-2019 | Malfunction | covidien lp llc north haven | 20-01-2020 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN CLOSING THE RELOAD FOR FIRING DURING A VIDEO ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, THE SURGEON WAS UNABLE TO CLOSE THE JAWS OF THE RELOAD ON THE TISSUES EVEN WHEN MADE SEVERAL ASSEMBLE AND DISASSEMBLE OF THE RELOAD ON THE HANDLE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE DUODENUM DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE RELOAD WAS STUCK AND COULD NOT BE FIRED AFTER FIRING 15MM. IT WAS ALSO REPORTED THAT THE RELOAD COULD NOT BE REMOVED FROM THE STAPLER. THE DEVICE WAS REPLACED WITH A NEW INSTRUMENT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE DUODENUM DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE RELOAD WAS STUCK AND COULD NOT BE FIRED AFTER FIRING 15MM. IT WAS ALSO REPORTED THAT THE RELOAD COULD NOT BE REMOVED FROM THE STAPLER. THE DEVICE WAS REPLACED WITH A NEW INSTRUMENT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE DUODENUM DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE RELOAD WAS STUCK AND COULD NOT BE FIRED AFTER FIRING 15MM. IT WAS ALSO REPORTED THAT THE RELOAD COULD NOT BE REMOVED FROM THE STAPLER. THE DEVICE WAS REPLACED WITH A NEW INSTRUMENT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2019 | Malfunction | covidien lp llc north haven | 20-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING STAPLING IN A LAPAROSCOPIC GASTROPLASTY PROCEDURE, THE RELOAD DID NOT FIRE. IT WAS UNLOADED AND RE LOADED BUT THE RELOAD REMAINED LOCKED. THE RELOAD WAS REPLACED AND THE PROCEDURE PROCEEDED NORMALLY. THERE WAS NO PATIENT INJURY.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Fluid/Blood Leak | A050401 | Post Operative Wound Infection | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Mechanics Altered | A0507 | Post Operative Wound Infection | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Fluid/Blood Leak | A050401 | Post Operative Wound Infection | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Mechanics Altered | A0507 | Post Operative Wound Infection | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
31-12-2019 | Injury | us surgical puerto rico | 20-01-2020 | GIA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL CYSTECTOMY ILEAL CONDUIT PROCEDURE, ON THE THIRD FIRING, WHEN THE ANASTOMOSIS WAS MADE, THE BLADE CUT THE BOWEL BUT DID NOT DEPLOY STAPLES. THE SURGEON HAD TO SUTURE THE BOWEL WHICH RESULTED IN COMPLICATIONS DUE TO A LEAK. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE INCISION WAS EXTENDED BY MORE THAN 1 INCH. IT WAS ALSO REPORTED THAT THERE WAS TISSUE DAMAGE AND TISSUE LOSS. THE HOSPITAL STAY WAS EXTENDED AND 5 DAYS AFTER THE SURGERY, THE PATIENT WAS RUSHED TO THE INTENSIVE CARE UNIT WITH A SEVERE LEAK FROM THE ANASTOMOSIS SITE AND INFECTION WHICH ENDED UP WITH THE PATIENT HAVING A SECOND STOMA.
10-03-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. IT WAS NOTED THAT THE HANDLE BROKE. SURGEON USED ANOTHER HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. NO PATIENT INJURY.
10-03-2020 | Malfunction | us surgical puerto rico | 20-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. IT WAS NOTED THAT THE HANDLE BROKE. SURGEON USED ANOTHER HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. NO PATIENT INJURY.
04-01-2020 | Injury | covidien lp llc north haven | 20-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE ON FOURTH FIRING, THERE WAS A STAPLE LINE INCOMPLETE AND POOR STAPLE FORMATION. THE SURGICAL PROCEDURE WAS EXTENDED BY MORE THAN 30 MINUTES. THE SURGEON HAD TO EXCISE THE MISFIRE PART AND RE-STAPLED AGAIN BESIDE IT.
04-01-2020 | Injury | covidien lp llc north haven | 20-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE ON FOURTH FIRING, THERE WAS A STAPLE LINE INCOMPLETE AND POOR STAPLE FORMATION. THE SURGICAL PROCEDURE WAS EXTENDED BY MORE THAN 30 MINUTES. THE SURGEON HAD TO EXCISE THE MISFIRE PART AND RE-STAPLED AGAIN BESIDE IT.
09-12-2019 | Malfunction | covidien lp llc north haven | 20-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO GASTROPLASTY PROCEDURE, THE PURPLE LOAD FAILED AT THE BEGINNING OF FIRING AND DID NOT PROGRESS. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2020 | Injury | covidien lp llc north haven | 20-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING FOURTH FIRING, THERE WAS A STAPLE LINE INCOMPLETE AND POOR STAPLE FORMATION. THE SURGICAL PROCEDURE WAS EXTENDED BY MORE THAN 30 MINUTES. THE SURGEON HAD TO EXCISE THE MISFIRE PART AND RE-STAPLED AGAIN BESIDE IT.
04-01-2020 | Injury | covidien lp llc north haven | 20-01-2020 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, DURING FOURTH FIRING, THERE WAS A STAPLE LINE INCOMPLETE AND POOR STAPLE FORMATION. THE SURGICAL PROCEDURE WAS EXTENDED BY MORE THAN 30 MINUTES. THE SURGEON HAD TO EXCISE THE MISFIRE PART AND RE-STAPLED AGAIN BESIDE IT.
08-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | TRI-STAPLE 2.0 | Failure to Cycle | A050703 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING THE TRANSECTION OF THE RIGHT UPPER LOBE DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS ALSO REPORTED THAT THE SURGEON APPROXIMATED THE TISSUE USING A BLACK RELOAD. THE TISSUE WAS VERY DIFFICULT TO COMPRESS WITHIN THE JAWS OF THE RELOAD, THOUGH THE TISSUE DID NOT APPEAR TO BE ANY THICKER THAN THOSE THE SURGEON HAVE SUCCESSFULLY FIRED IN THE PAST. AFTER COMPRESSING THE TISSUE, THE SURGEON ENTERED FIRING MODE WHERE THE STAPLER SEEMED TO "BLOW A PISTON" IN THE RELOAD BUT THE RELOAD WAS NOT FIRED. THE STAPLER WAS RELOADED WITH ANOTHER BLACK RELOAD, APPROXIMATED THE TISSUE, BUT UPON TRYING TO FIRE, THE TISSUE WAS PUSHED/PRESSED OUT OF THE RELOAD. A NEW STAPLER AND RELOAD WERE OPENED AND USED TO TRANSECT/RESECT THE TISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING THE TRANSECTION OF THE RIGHT UPPER LOBE DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS ALSO REPORTED THAT THE SURGEON APPROXIMATED THE TISSUE USING A BLACK RELOAD. THE TISSUE WAS VERY DIFFICULT TO COMPRESS WITHIN THE JAWS OF THE RELOAD, THOUGH THE TISSUE DID NOT APPEAR TO BE ANY THICKER THAN THOSE THE SURGEON HAVE SUCCESSFULLY FIRED IN THE PAST. AFTER COMPRESSING THE TISSUE, THE SURGEON ENTERED FIRING MODE WHERE THE STAPLER SEEMED TO "BLOW A PISTON" IN THE RELOAD BUT THE RELOAD WAS NOT FIRED. THE STAPLER WAS RELOADED WITH ANOTHER BLACK RELOAD, APPROXIMATED THE TISSUE, BUT UPON TRYING TO FIRE, THE TISSUE WAS PUSHED/PRESSED OUT OF THE RELOAD. A NEW STAPLER AND RELOAD WERE OPENED AND USED TO TRANSECT/RESECT THE TISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | covidien lp llc north haven | 20-01-2020 | TRI-STAPLE 2.0 | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING THE TRANSECTION OF THE RIGHT UPPER LOBE DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS ALSO REPORTED THAT THE SURGEON APPROXIMATED THE TISSUE USING A BLACK RELOAD. THE TISSUE WAS VERY DIFFICULT TO COMPRESS WITHIN THE JAWS OF THE RELOAD, THOUGH THE TISSUE DID NOT APPEAR TO BE ANY THICKER THAN THOSE THE SURGEON HAVE SUCCESSFULLY FIRED IN THE PAST. AFTER COMPRESSING THE TISSUE, THE SURGEON ENTERED FIRING MODE WHERE THE STAPLER SEEMED TO "BLOW A PISTON" IN THE RELOAD BUT THE RELOAD WAS NOT FIRED. THE STAPLER WAS RELOADED WITH ANOTHER BLACK RELOAD, APPROXIMATED THE TISSUE, BUT UPON TRYING TO FIRE, THE TISSUE WAS PUSHED/PRESSED OUT OF THE RELOAD. A NEW STAPLER AND RELOAD WERE OPENED AND USED TO TRANSECT/RESECT THE TISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2019 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING GASTRECTOMY, WHEN PREPARING THE DEVICE TO CUT GASTRIC TUMOR, AFTER LOADING THE SECOND RELOAD, THE RELOAD COULD NOT FULLY CLAMP AND CUT. THE DEVICE WAS STILL UNUSABLE AFTER RELOADING. SURGEON REPLACED WITH A NEW GUN TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO GASTRIC SEPTATION DURING LAPAROSCOPIC GASTROPLASTY, THE DEVICE DID NOT FIRE. THE PRODUCT WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO GASTRIC SEPTATION DURING LAPAROSCOPIC GASTROPLASTY, THE DEVICE DID NOT FIRE. THE PRODUCT WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | MULTIFIRE ENDO GIA | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC APPENDECTOMY, DURING STAPLING, THE STAPLES WERE DEPLOYED BUT THE KNIFE DID NOT CUT. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING THE STOMACH TISSUE DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT IT WAS STATED THAT UPON VISUAL INSPECTION, THE STAPLER APPEARED AS THOUGH THE RELOAD WAS PRE-FIRED. THERE WERE COUPLE OF STAPLES THAT HAD STARTED TO DEPLOY AT THE PROXIMAL END OF THE RELOAD HENCE, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE AS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Expulsion | A010401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING THE STOMACH TISSUE DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT IT WAS STATED THAT UPON VISUAL INSPECTION, THE STAPLER APPEARED AS THOUGH THE RELOAD WAS PRE-FIRED. THERE WERE COUPLE OF STAPLES THAT HAD STARTED TO DEPLOY AT THE PROXIMAL END OF THE RELOAD HENCE, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE AS NO PATIENT INJURY.
08-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL PANCREATECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND FIRING OPERATION WAS PERFORMED, BUT THE DEVICE BECAME IDLED. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THE EVENT OCCURRED DURING THE PROCEDURE BUT THE PRODUCT WAS NOT USED FOR PATIENT.
23-10-2019 | Injury | covidien lp llc north haven | 21-01-2020 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, AFTER FIRING, THE DEVICE'S JAWS LOCKED ON THE TISSUE AND WAS REMOVED BY CUTTING FROM THE TISSUE. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
23-10-2019 | Injury | covidien lp llc north haven | 21-01-2020 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, AFTER FIRING, THE DEVICE'S JAWS LOCKED ON THE TISSUE AND WAS REMOVED BY CUTTING FROM THE TISSUE. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
08-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, ON AN OPEN ILEOSTOMY REPAIR, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING STAPLING AND AFTER CHECKING ALL LIGHTS IN A LAPAROSCOPIC PROCEDURE, EVERYTHING WAS OKAY, BUT THE RELOAD DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RECTAL RESECTION THE PURPLE RELOAD COULD NOT BE FIRED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE COMPLETE THE PROCEDURE. NO PATIENT HARM.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY, AT THE THIRD FIRING, THE DOCTOR CLOSED THE JAWS BUT THEY WERE LOCKED, AND FIRING WAS UNABLE TO BE PERFORMED. EVEN AFTER THE DOCTOR REPEATED SEVERAL TIMES, THE SAME EVENT OCCURRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT HARM.
21-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Information | Event Description: THE DEVICE WAS RETURNED WITHOUT ANY ACCOMPANYING INFORMATION.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LOBECTOMY, AT THE THIRD FIRING, THE DOCTOR CLOSED THE JAWS BUT THEY WERE LOCKED, AND FIRING WAS UNABLE TO BE PERFORMED. EVEN AFTER THE DOCTOR REPEATED SEVERAL TIMES, THE SAME EVENT OCCURRED. THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. NO PATIENT HARM.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, TWO DEVICES DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, TWO DEVICES DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, TWO DEVICES DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, TWO DEVICES DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Injury | us surgical puerto rico | 21-01-2020 | GIA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A DISTAL GASTRECTOMY OPEN PROCEDURE, THE STAPLES DID NOT DEPLOY AND THE TISSUE WAS CUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
18-12-2019 | Injury | ethicon endo-surgery | 21-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT A PATIENT WHO UNDERGONE SEGMENTAL LOBECTOMY BY VIDEO THORACOSCOPY, FAILURES IN THE UNIVERSAL ENDOCRINE STAPLERS OF MEDTRONIC AND ECHELON 60 MM OF JOHNSON & JOHNSON. INITIALLY THE FIRST UNIVERSAL ENDO STAPLER WAS USED WITH TWO PURPLE 60 MM RELOADS WHICH WORKED PERFECTLY, WHEN THE THIRD PURPLE 60 MM RECHARGE PASSES, IT PRESENTS A FAILURE AND THE ENDO GIA STAPLER IS ALSO DAMAGED, FOR THIS REASON A SECOND NEW ENDO GIA STAPLER WITH PURPLE RECHARGE IS PASSED TO THE TABLE, WHICH IS ALSO DAMAGED. SO REFILLS ARE PASSED 60 MM BLACK AND THESE ARE ALSO DAMAGED, SUBSEQUENTLY, THE DOCTOR REQUESTS A 60 MM ECHELON STAPLER WITH 4 GREEN 60 MM REFILLS, WHICH ONLY STAPLED HALFWAY. WHEN REQUESTING MORE REFILLS WAS NOT POSSIBLE DUE TO THEIR LACK OF EXISTENCE, THEREFORE, IN ORDER TO COMPLETE THE PROCEDURE, THE DOCTOR REQUESTS A THIRD ENDO GIA STAPLER WITH 2 BLACK REFILLS, OF WHICH ONLY ONE OF THEM STAPLED AND THE ENDO STAPLER IS DAMAGED AGAIN. IN THE END, THE DOCTOR DECIDES TO CONVERT THORACOSCOPY SURGERY TO OPEN THORACOTOMY.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE DUODENUM DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE RELOAD WAS STUCK AND COULD NOT BE FIRED AFTER FIRING 15MM. IT WAS ALSO REPORTED THAT THE RELOAD COULD NOT BE REMOVED FROM THE STAPLER. BOTH THE RELOAD AND HANDLE WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE DUODENUM DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE RELOAD WAS STUCK AND COULD NOT BE FIRED AFTER FIRING 15MM. IT WAS ALSO REPORTED THAT THE RELOAD COULD NOT BE REMOVED FROM THE STAPLER. BOTH THE RELOAD AND HANDLE WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
04-12-2019 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RECTOSIGMOIDECTOMY, DURING THE FIRST FIRING, THE DEVICE LOCKED AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Injury | covidien medical products | 21-01-2020 | PREMIUM PLUS CEEA | Break | A0401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL COLECTOMY FOLLOWED BY ILEAL POUCH ANASTOMOSIS, BEFORE FIRING, THE ANVIL WAS ENGAGED WITH PURSTRING INTO THE ILEAL POUCH AND THE STAPLER WAS INSERTED IN THE RECTUM. AT THE TIME OF CLOSURE, THE STAPLER WAS UNABLE TO APPROXIMATE AND DID NOT SHOW THE READY TO FIRE INDICATOR, SO THE SURGEON REMOVED THE STAPLER TO SEE THE PROBLEM AND FOUND OUT THAT THE STAPLER HEAD BROKE. SURGEON HAS RESECTED 2-3 CM PART OF COLON AND TAKEN RE PURSE-STRING TO FIRE NEW STAPLER AND THE SURGICAL TIME WAS EXTENDED BY 40 MINUTES.
07-01-2020 | Injury | covidien medical products | 21-01-2020 | PREMIUM PLUS CEEA | Break | A0401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL COLECTOMY FOLLOWED BY ILEAL POUCH ANASTOMOSIS, BEFORE FIRING, THE ANVIL WAS ENGAGED WITH PURSTRING INTO THE ILEAL POUCH AND THE STAPLER WAS INSERTED IN THE RECTUM. AT THE TIME OF CLOSURE, THE STAPLER WAS UNABLE TO APPROXIMATE AND DID NOT SHOW THE READY TO FIRE INDICATOR, SO THE SURGEON REMOVED THE STAPLER TO SEE THE PROBLEM AND FOUND OUT THAT THE STAPLER HEAD BROKE. SURGEON HAS RESECTED 2-3 CM PART OF COLON AND TAKEN RE PURSE-STRING TO FIRE NEW STAPLER AND THE SURGICAL TIME WAS EXTENDED BY 40 MINUTES.
07-01-2020 | Injury | covidien medical products | 21-01-2020 | PREMIUM PLUS CEEA | Device Slipped | A051204 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL COLECTOMY FOLLOWED BY ILEAL POUCH ANASTOMOSIS, BEFORE FIRING, THE ANVIL WAS ENGAGED WITH PURSTRING INTO THE ILEAL POUCH AND THE STAPLER WAS INSERTED IN THE RECTUM. AT THE TIME OF CLOSURE, THE STAPLER WAS UNABLE TO APPROXIMATE AND DID NOT SHOW THE READY TO FIRE INDICATOR, SO THE SURGEON REMOVED THE STAPLER TO SEE THE PROBLEM AND FOUND OUT THAT THE STAPLER HEAD BROKE. SURGEON HAS RESECTED 2-3 CM PART OF COLON AND TAKEN RE PURSE-STRING TO FIRE NEW STAPLER AND THE SURGICAL TIME WAS EXTENDED BY 40 MINUTES.
07-01-2020 | Injury | covidien medical products | 21-01-2020 | PREMIUM PLUS CEEA | Device Slipped | A051204 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL COLECTOMY FOLLOWED BY ILEAL POUCH ANASTOMOSIS, BEFORE FIRING, THE ANVIL WAS ENGAGED WITH PURSTRING INTO THE ILEAL POUCH AND THE STAPLER WAS INSERTED IN THE RECTUM. AT THE TIME OF CLOSURE, THE STAPLER WAS UNABLE TO APPROXIMATE AND DID NOT SHOW THE READY TO FIRE INDICATOR, SO THE SURGEON REMOVED THE STAPLER TO SEE THE PROBLEM AND FOUND OUT THAT THE STAPLER HEAD BROKE. SURGEON HAS RESECTED 2-3 CM PART OF COLON AND TAKEN RE PURSE-STRING TO FIRE NEW STAPLER AND THE SURGICAL TIME WAS EXTENDED BY 40 MINUTES.
04-12-2019 | Malfunction | covidien lp llc north haven | 21-01-2020 | TRI-STAPLE 2.0 | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RECTOSIGMOIDECTOMY, DURING THE FIRST FIRING, THE DEVICE LOCKED AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2019 | Malfunction | ethicon | 21-01-2020 | ETHICON SECURESTRAP | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A BILATERAL INGUINAL HERNIORRHAPHY BY LAPAROSCOPY ON (B)(6) 2019 AND THE ABSORBABLE STRAPS WERE USED. IT WAS REPORTED THAT AT THE TIME OF FIRING THE STAPLES OR HOOKS TO CLOSE THE PERITONEUM, THE DEVICE DID NOT HAVE SUFFICIENT CLAMPING FORCE WHICH DID NOT ALLOW AN APPROPRIATE CAPTURE AND PENETRATION OF THE TISSUE. ANOTHER LIKE DEVICE WAS USED TO SUCCESSFULLY COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED.
26-12-2019 | Injury | ethicon endo-surgery | 21-01-2020 | ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : C-REACTIVE PROTEIN AND PROCALCITONIN FOR THE EARLY DETECTION OF POSTOPERATIVE COMPLICATIONS AFTER SLEEVE GASTRECTOMY: PRELIMINARY STUDY IN 97 PATIENTS. AUTHOR: R. KASSIR ¿ P. BLANC ¿ L. M. BRUNA TIBALBO ¿ C. BRETON ¿ P. LOINTIER. CITATION: SURG ENDOSC (2015) 29:1439¿1444, DOI 10.1007/S00464-014-3821-2, PUBLISHED ONLINE: 27 AUGUST 2014. THIS RETROSPECTIVE STUDY AIMED TO EVALUATE THE INTEREST OF C-REACTIVE PROTEIN (CRP) AND PROCALCITONIN (PCT) ASSAY FOR THE EARLY DETECTION OF GASTRIC FISTULA AFTER SLEEVE GASTRECTOMY AND TO STUDY THE PCT AS AN ADJUNCTIVE MARKER TO THE CRP. BETWEEN JANUARY 1, 2011 AND DECEMBER 31, 2012, 97 PATIENTS UNDERWENT SLEEVE GASTRECTOMY: 18 MALES AND 79 FEMALES WITH A MEAN AGE OF 41 YEARS (21¿68 YEARS). DURING SURGERY, THE GREATER CURVATURE WAS RELEASED FROM THE ANTRUM (4¿6 CM FROM THE PYLORUS) TO THE LEFT CRUS OF THE DIAPHRAGM USING THE ULTRACISION HARMONIC SCALPEL_ (ETHICON ENDOSURGERY, INC., CINCINNATI, OH, USA). AN ECHELON STAPLER, 60 MM GREEN CARTRIDGE (4.8 MM) WAS USED ON THE ANTRUM AND BLUE CARTRIDGES (3.5 MM) WERE USED TO COMPLETE THE DIVISION OF THE REMAINDER OF THE STOMACH (ETHICON ENDOSURGERY, INC., CINCINNATI, OH, USA). THE STAPLE LINE WAS REINFORCED WITH GLUE IN 68 CASES (IFABOND_ GLUE (IFAMEDICAL, (B)(4)) IN 48 CASES AND EVICEL_ GLUE (ETHICON ENDOSURGERY, INC.) IN 20 CASES). THERE WERE 7 PATIENTS WHO HAD POSTOPERATIVE COMPLICATIONS: 1. 23/F HAD GRADE IIIB FISTULAE WHICH WAS TREATED WITH DRAINAGE (LAPAROSCOPY) THEN STENT AND THEN ENTEROCUTANEOUS FISTULA. THIS PATENT HAD STENT REMOVAL AT 6 WEEKS FISTULA CLOSURE AT 6 MONTHS. 2. 33/H HAD GRADE IV FISTULAE WITH PERITONITIS WHICH WAS TREATED WITH DRAINAGE (LAPAROTOMY) THEN STENT WITH 6 WEEKS HEALING TIME. 3. 32/F HAD GRADE III ABSCESS WHICH WAS TREATED WITH STENT BUT HAD FAILED. THIS PATIENT HAD TOTAL GASTRECTOMY 1 YEAR AFTER. 4. 61/F HAD GRADE III ABSCESS WHICH WAS TREATED WITH RADIOLOGICAL DRAIN WHICH HEALED AFTER 1 MONTH. 5. 31/F HAD GRADE II ABSCESSES WHICH WAS TREATED WITH MEDICAL TREATMENT WHICH HEALED AFTER 15 DAYS. 6. 34/H HAD GRADE II FISTULAE WHICH WAS TREATED WITH MEDICAL TREATMENT WHICH HEALED AFTER 15 DAYS. 7. 34/F HAD GRADE II ABSCESSES WHICH WAS TREATED WITH RADIOLOGICAL DRAIN WHICH HEALED AFTER 15 DAYS. IN CONCLUSION, THIS STUDY SHOWS THAT CLINICAL OBSERVATIONS IN ASSOCIATION WITH CRP AND PCT MEASUREMENTS COULD BE OF HELP FOR THE EARLY DETECTION OF SEPTIC COMPLICATIONS AFTER SLEEVE GASTRECTOMY. A PROSPECTIVE STUDY IS IN PROGRESS TO VALIDATE THESE FINDINGS.
26-12-2019 | Injury | ethicon endo-surgery | 21-01-2020 | ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : C-REACTIVE PROTEIN AND PROCALCITONIN FOR THE EARLY DETECTION OF POSTOPERATIVE COMPLICATIONS AFTER SLEEVE GASTRECTOMY: PRELIMINARY STUDY IN 97 PATIENTS. AUTHOR: R. KASSIR ¿ P. BLANC ¿ L. M. BRUNA TIBALBO ¿ C. BRETON ¿ P. LOINTIER. CITATION: SURG ENDOSC (2015) 29:1439¿1444, DOI 10.1007/S00464-014-3821-2, PUBLISHED ONLINE: 27 AUGUST 2014. THIS RETROSPECTIVE STUDY AIMED TO EVALUATE THE INTEREST OF C-REACTIVE PROTEIN (CRP) AND PROCALCITONIN (PCT) ASSAY FOR THE EARLY DETECTION OF GASTRIC FISTULA AFTER SLEEVE GASTRECTOMY AND TO STUDY THE PCT AS AN ADJUNCTIVE MARKER TO THE CRP. BETWEEN JANUARY 1, 2011 AND DECEMBER 31, 2012, 97 PATIENTS UNDERWENT SLEEVE GASTRECTOMY: 18 MALES AND 79 FEMALES WITH A MEAN AGE OF 41 YEARS (21¿68 YEARS). DURING SURGERY, THE GREATER CURVATURE WAS RELEASED FROM THE ANTRUM (4¿6 CM FROM THE PYLORUS) TO THE LEFT CRUS OF THE DIAPHRAGM USING THE ULTRACISION HARMONIC SCALPEL_ (ETHICON ENDOSURGERY, INC., CINCINNATI, OH, USA). AN ECHELON STAPLER, 60 MM GREEN CARTRIDGE (4.8 MM) WAS USED ON THE ANTRUM AND BLUE CARTRIDGES (3.5 MM) WERE USED TO COMPLETE THE DIVISION OF THE REMAINDER OF THE STOMACH (ETHICON ENDOSURGERY, INC., CINCINNATI, OH, USA). THE STAPLE LINE WAS REINFORCED WITH GLUE IN 68 CASES (IFABOND_ GLUE (IFAMEDICAL, (B)(4)) IN 48 CASES AND EVICEL_ GLUE (ETHICON ENDOSURGERY, INC.) IN 20 CASES). THERE WERE 7 PATIENTS WHO HAD POSTOPERATIVE COMPLICATIONS: 1. 23/F HAD GRADE IIIB FISTULAE WHICH WAS TREATED WITH DRAINAGE (LAPAROSCOPY) THEN STENT AND THEN ENTEROCUTANEOUS FISTULA. THIS PATENT HAD STENT REMOVAL AT 6 WEEKS FISTULA CLOSURE AT 6 MONTHS. 2. 33/H HAD GRADE IV FISTULAE WITH PERITONITIS WHICH WAS TREATED WITH DRAINAGE (LAPAROTOMY) THEN STENT WITH 6 WEEKS HEALING TIME. 3. 32/F HAD GRADE III ABSCESS WHICH WAS TREATED WITH STENT BUT HAD FAILED. THIS PATIENT HAD TOTAL GASTRECTOMY 1 YEAR AFTER. 4. 61/F HAD GRADE III ABSCESS WHICH WAS TREATED WITH RADIOLOGICAL DRAIN WHICH HEALED AFTER 1 MONTH. 5. 31/F HAD GRADE II ABSCESSES WHICH WAS TREATED WITH MEDICAL TREATMENT WHICH HEALED AFTER 15 DAYS. 6. 34/H HAD GRADE II FISTULAE WHICH WAS TREATED WITH MEDICAL TREATMENT WHICH HEALED AFTER 15 DAYS. 7. 34/F HAD GRADE II ABSCESSES WHICH WAS TREATED WITH RADIOLOGICAL DRAIN WHICH HEALED AFTER 15 DAYS. IN CONCLUSION, THIS STUDY SHOWS THAT CLINICAL OBSERVATIONS IN ASSOCIATION WITH CRP AND PCT MEASUREMENTS COULD BE OF HELP FOR THE EARLY DETECTION OF SEPTIC COMPLICATIONS AFTER SLEEVE GASTRECTOMY. A PROSPECTIVE STUDY IS IN PROGRESS TO VALIDATE THESE FINDINGS.
26-12-2019 | Injury | ethicon endo-surgery | 21-01-2020 | ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Peritonitis | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : C-REACTIVE PROTEIN AND PROCALCITONIN FOR THE EARLY DETECTION OF POSTOPERATIVE COMPLICATIONS AFTER SLEEVE GASTRECTOMY: PRELIMINARY STUDY IN 97 PATIENTS. AUTHOR: R. KASSIR ¿ P. BLANC ¿ L. M. BRUNA TIBALBO ¿ C. BRETON ¿ P. LOINTIER. CITATION: SURG ENDOSC (2015) 29:1439¿1444, DOI 10.1007/S00464-014-3821-2, PUBLISHED ONLINE: 27 AUGUST 2014. THIS RETROSPECTIVE STUDY AIMED TO EVALUATE THE INTEREST OF C-REACTIVE PROTEIN (CRP) AND PROCALCITONIN (PCT) ASSAY FOR THE EARLY DETECTION OF GASTRIC FISTULA AFTER SLEEVE GASTRECTOMY AND TO STUDY THE PCT AS AN ADJUNCTIVE MARKER TO THE CRP. BETWEEN JANUARY 1, 2011 AND DECEMBER 31, 2012, 97 PATIENTS UNDERWENT SLEEVE GASTRECTOMY: 18 MALES AND 79 FEMALES WITH A MEAN AGE OF 41 YEARS (21¿68 YEARS). DURING SURGERY, THE GREATER CURVATURE WAS RELEASED FROM THE ANTRUM (4¿6 CM FROM THE PYLORUS) TO THE LEFT CRUS OF THE DIAPHRAGM USING THE ULTRACISION HARMONIC SCALPEL_ (ETHICON ENDOSURGERY, INC., CINCINNATI, OH, USA). AN ECHELON STAPLER, 60 MM GREEN CARTRIDGE (4.8 MM) WAS USED ON THE ANTRUM AND BLUE CARTRIDGES (3.5 MM) WERE USED TO COMPLETE THE DIVISION OF THE REMAINDER OF THE STOMACH (ETHICON ENDOSURGERY, INC., CINCINNATI, OH, USA). THE STAPLE LINE WAS REINFORCED WITH GLUE IN 68 CASES (IFABOND_ GLUE (IFAMEDICAL, (B)(4)) IN 48 CASES AND EVICEL_ GLUE (ETHICON ENDOSURGERY, INC.) IN 20 CASES). THERE WERE 7 PATIENTS WHO HAD POSTOPERATIVE COMPLICATIONS: 1. 23/F HAD GRADE IIIB FISTULAE WHICH WAS TREATED WITH DRAINAGE (LAPAROSCOPY) THEN STENT AND THEN ENTEROCUTANEOUS FISTULA. THIS PATENT HAD STENT REMOVAL AT 6 WEEKS FISTULA CLOSURE AT 6 MONTHS. 2. 33/H HAD GRADE IV FISTULAE WITH PERITONITIS WHICH WAS TREATED WITH DRAINAGE (LAPAROTOMY) THEN STENT WITH 6 WEEKS HEALING TIME. 3. 32/F HAD GRADE III ABSCESS WHICH WAS TREATED WITH STENT BUT HAD FAILED. THIS PATIENT HAD TOTAL GASTRECTOMY 1 YEAR AFTER. 4. 61/F HAD GRADE III ABSCESS WHICH WAS TREATED WITH RADIOLOGICAL DRAIN WHICH HEALED AFTER 1 MONTH. 5. 31/F HAD GRADE II ABSCESSES WHICH WAS TREATED WITH MEDICAL TREATMENT WHICH HEALED AFTER 15 DAYS. 6. 34/H HAD GRADE II FISTULAE WHICH WAS TREATED WITH MEDICAL TREATMENT WHICH HEALED AFTER 15 DAYS. 7. 34/F HAD GRADE II ABSCESSES WHICH WAS TREATED WITH RADIOLOGICAL DRAIN WHICH HEALED AFTER 15 DAYS. IN CONCLUSION, THIS STUDY SHOWS THAT CLINICAL OBSERVATIONS IN ASSOCIATION WITH CRP AND PCT MEASUREMENTS COULD BE OF HELP FOR THE EARLY DETECTION OF SEPTIC COMPLICATIONS AFTER SLEEVE GASTRECTOMY. A PROSPECTIVE STUDY IS IN PROGRESS TO VALIDATE THESE FINDINGS.
26-12-2019 | Injury | ethicon endo-surgery | 21-01-2020 | ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : C-REACTIVE PROTEIN AND PROCALCITONIN FOR THE EARLY DETECTION OF POSTOPERATIVE COMPLICATIONS AFTER SLEEVE GASTRECTOMY: PRELIMINARY STUDY IN 97 PATIENTS. AUTHOR: R. KASSIR ¿ P. BLANC ¿ L. M. BRUNA TIBALBO ¿ C. BRETON ¿ P. LOINTIER. CITATION: SURG ENDOSC (2015) 29:1439¿1444, DOI 10.1007/S00464-014-3821-2, PUBLISHED ONLINE: 27 AUGUST 2014. THIS RETROSPECTIVE STUDY AIMED TO EVALUATE THE INTEREST OF C-REACTIVE PROTEIN (CRP) AND PROCALCITONIN (PCT) ASSAY FOR THE EARLY DETECTION OF GASTRIC FISTULA AFTER SLEEVE GASTRECTOMY AND TO STUDY THE PCT AS AN ADJUNCTIVE MARKER TO THE CRP. BETWEEN JANUARY 1, 2011 AND DECEMBER 31, 2012, 97 PATIENTS UNDERWENT SLEEVE GASTRECTOMY: 18 MALES AND 79 FEMALES WITH A MEAN AGE OF 41 YEARS (21¿68 YEARS). DURING SURGERY, THE GREATER CURVATURE WAS RELEASED FROM THE ANTRUM (4¿6 CM FROM THE PYLORUS) TO THE LEFT CRUS OF THE DIAPHRAGM USING THE ULTRACISION HARMONIC SCALPEL_ (ETHICON ENDOSURGERY, INC., CINCINNATI, OH, USA). AN ECHELON STAPLER, 60 MM GREEN CARTRIDGE (4.8 MM) WAS USED ON THE ANTRUM AND BLUE CARTRIDGES (3.5 MM) WERE USED TO COMPLETE THE DIVISION OF THE REMAINDER OF THE STOMACH (ETHICON ENDOSURGERY, INC., CINCINNATI, OH, USA). THE STAPLE LINE WAS REINFORCED WITH GLUE IN 68 CASES (IFABOND_ GLUE (IFAMEDICAL, (B)(4)) IN 48 CASES AND EVICEL_ GLUE (ETHICON ENDOSURGERY, INC.) IN 20 CASES). THERE WERE 7 PATIENTS WHO HAD POSTOPERATIVE COMPLICATIONS: 1. 23/F HAD GRADE IIIB FISTULAE WHICH WAS TREATED WITH DRAINAGE (LAPAROSCOPY) THEN STENT AND THEN ENTEROCUTANEOUS FISTULA. THIS PATENT HAD STENT REMOVAL AT 6 WEEKS FISTULA CLOSURE AT 6 MONTHS. 2. 33/H HAD GRADE IV FISTULAE WITH PERITONITIS WHICH WAS TREATED WITH DRAINAGE (LAPAROTOMY) THEN STENT WITH 6 WEEKS HEALING TIME. 3. 32/F HAD GRADE III ABSCESS WHICH WAS TREATED WITH STENT BUT HAD FAILED. THIS PATIENT HAD TOTAL GASTRECTOMY 1 YEAR AFTER. 4. 61/F HAD GRADE III ABSCESS WHICH WAS TREATED WITH RADIOLOGICAL DRAIN WHICH HEALED AFTER 1 MONTH. 5. 31/F HAD GRADE II ABSCESSES WHICH WAS TREATED WITH MEDICAL TREATMENT WHICH HEALED AFTER 15 DAYS. 6. 34/H HAD GRADE II FISTULAE WHICH WAS TREATED WITH MEDICAL TREATMENT WHICH HEALED AFTER 15 DAYS. 7. 34/F HAD GRADE II ABSCESSES WHICH WAS TREATED WITH RADIOLOGICAL DRAIN WHICH HEALED AFTER 15 DAYS. IN CONCLUSION, THIS STUDY SHOWS THAT CLINICAL OBSERVATIONS IN ASSOCIATION WITH CRP AND PCT MEASUREMENTS COULD BE OF HELP FOR THE EARLY DETECTION OF SEPTIC COMPLICATIONS AFTER SLEEVE GASTRECTOMY. A PROSPECTIVE STUDY IS IN PROGRESS TO VALIDATE THESE FINDINGS.
06-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | EEA | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, AT COLOSTOMY CLOSURE DURING AN OPEN RECTOSIGMOIDECTOMY PROCEDURE, THE STAPLER'S WARHEAD WAS OPEN AND WAS NOT APPROPRIATELY MATING THE STAPLER. THE STAPLER WAS REPLACED TO COMPLETE THE CASE. THE SURGICAL TIME EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AT COLOSTOMY CLOSURE DURING AN OPEN RECTOSIGMOIDECTOMY PROCEDURE, THE STAPLER'S WARHEAD WAS OPEN AND WAS NOT APPROPRIATELY MATING THE STAPLER. THE STAPLER WAS REPLACED TO COMPLETE THE CASE. THE SURGICAL TIME EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING THE TRANSECTION OF THE RIGHT UPPER LOBE DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. IT WAS ALSO REPORTED THAT THE SURGEON APPROXIMATED THE TISSUE USING A BLACK RELOAD. THE TISSUE WAS VERY DIFFICULT TO COMPRESS WITHIN THE JAWS OF THE RELOAD, THOUGH THE TISSUE DID NOT APPEAR TO BE ANY THICKER THAN THOSE THE SURGEON HAVE SUCCESSFULLY FIRED IN THE PAST. AFTER COMPRESSING THE TISSUE, THE SURGEON ENTERED FIRING MODE WHERE THE STAPLER SEEMED TO "BLOW A PISTON" IN THE RELOAD BUT THE RELOAD WAS NOT FIRED. THE STAPLER WAS RELOADED WITH ANOTHER BLACK RELOAD, APPROXIMATED THE TISSUE, BUT UPON TRYING TO FIRE, THE TISSUE WAS PUSHED/PRESSED OUT OF THE RELOAD. A NEW STAPLER AND RELOAD WERE OPENED AND USED TO TRANSECT/RESECT THE TISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Injury | us surgical puerto rico | 21-01-2020 | TA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING FIRING THE DISTAL RECTUM ON OPEN RIGHT HEMICOLECTOMY, THE STAPLER HAD POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE. THE SURGEON HAD TO OVERSEW TO FIXED THE STAPLE LINE. SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM.
07-01-2020 | Injury | us surgical puerto rico | 21-01-2020 | TA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING FIRING THE DISTAL RECTUM ON OPEN RIGHT HEMICOLECTOMY, THE STAPLER HAD POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE. THE SURGEON HAD TO OVERSEW TO FIXED THE STAPLE LINE. SURGICAL TIME WAS EXTENDED MORE THAN 30 MINUTES DUE TO THE PRODUCT PROBLEM.
07-01-2020 | Injury | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLE LINE WAS INCOMPLETE DISTALLY EVEN THOUGH FIRING CYCLE WAS COMPLETED - WHICH WAS FIXED BY USING A SUTURE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | IDRIVE ULTRA | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, BEFORE THE DEVICE WAS SENT FOR THE PROCEDURE, THE PRODUCT'S BLUE PUSH BUTTON LOCKED, THE STAPLER WAS NOT ABLE TO FIRE. THE INSTRUMENT WAS NOT USED. THERE WAS NO PATIENT INVOLVEMENT.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | IDRIVE ULTRA | Difficult to Open or Close | A051104 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, BEFORE THE DEVICE WAS SENT FOR THE PROCEDURE, THE PRODUCT'S BLUE PUSH BUTTON LOCKED, THE STAPLER WAS NOT ABLE TO FIRE. THE INSTRUMENT WAS NOT USED. THERE WAS NO PATIENT INVOLVEMENT.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | UNKNOWN EGIA ADAPTER | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, THE PRODUCT ARRIVED AT THE COMPANY WITH THE BLUE PUSH BUTTON LOCKED, THE STAPLER WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INVOLVEMENT.
07-01-2020 | Malfunction | covidien lp llc north haven | 21-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, THE PRODUCT ARRIVED AT THE COMPANY WITH THE BLUE PUSH BUTTON LOCKED, THE STAPLER WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INVOLVEMENT.
07-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC RADICAL LUNG CANCER SURGERY, ON THE RIGHT UPPER LOBE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY JAWS CANNOT BE CLOSED. THEY REPLACED A NEW RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC RADICAL LUNG CANCER SURGERY, ON THE RIGHT UPPER LOBE, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE DID NOT SQUEEZE. THEY JAWS CANNOT BE CLOSED. THEY REPLACED A NEW RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT PULMONARY SEGMENTATION, THE SURGEON HAD A DIFFICULT TIME TO CLOSE THE STAPLER AFTER TISSUE REPOSITIONING AND GRASPING. HE HEARD A CLICKING NOISE AND OBSERVED THE ARTICULATION BREAKING. THE STAPLE WAS UNABLE TO BE FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | us surgical puerto rico | 21-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT PULMONARY SEGMENTATION, THE SURGEON HAD A DIFFICULT TIME TO CLOSE THE STAPLER AFTER TISSUE REPOSITIONING AND GRASPING. HE HEARD A CLICKING NOISE AND OBSERVED THE ARTICULATION BREAKING. THE STAPLE WAS UNABLE TO BE FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-11-2019 | Malfunction | covidien lp llc north haven | 21-01-2020 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO DISTAL PANCREATECTOMY PROCEDURE, A NEW SHELL WAS USED FOR DEMONSTRATION OF THE DEVICE BUT THE DISPLAY INDICATOR DID NOT CHANGE AT ALL WHICH MADE IT AN UNSUCCESSFUL DEMONSTRATION. THERE WAS NO PATIENT INVOLVEMENT.
30-12-2019 | Malfunction | ethicon endo-surgery | 21-01-2020 | ENDO ILS 29MM, CURVED | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ROBOTIC PROCEDURE THE DEVICE WAS FIRED AND IT DID NOT FIRING SEQUENCE. THE BLADE WOULD NOT RETRACT AND ALL OF THE STAPLES DID NOT FORM. ADDITIONAL STAPLING WAS REQUIRED. IT IS NOT KNOWN IF THE PROCEDURE WAS COMPLETED WITH THIS DEVICE AND THERE WERE NO PATIENT CONSEQUENCES REPORTED.
30-12-2019 | Malfunction | ethicon endo-surgery | 21-01-2020 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ROBOTIC PROCEDURE THE DEVICE WAS FIRED AND IT DID NOT FIRING SEQUENCE. THE BLADE WOULD NOT RETRACT AND ALL OF THE STAPLES DID NOT FORM. ADDITIONAL STAPLING WAS REQUIRED. IT IS NOT KNOWN IF THE PROCEDURE WAS COMPLETED WITH THIS DEVICE AND THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-11-2019 | Malfunction | ethicon endo-surgery | 21-01-2020 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Misfire | A050502 | No Consequences Or Impact To Patient | nan
01-01-2019 | Injury | ethicon endo-surgery | 21-01-2020 | GST BLUE RELOAD, 60MM, 6 ROW | Leak/Splash | A0504 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE PHYSICIAN EXPERIENCED A LEAK POST-OP ON A SLEEVE GASTRECTOMY PATIENT. IT IS UNKNOWN OF WETHER INTRA-OPERATIVE EVERYTHING LOOKED TO BE A SUCCESS. THE LEAK WAS IN THE MIDDLE PORTION OF THE STOMACH. THE PATIENT WAS READMITTED TO THE HOSPITAL.
01-01-2019 | Injury | ethicon endo-surgery | 21-01-2020 | GST BLUE RELOAD, 60MM, 6 ROW | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT THE PHYSICIAN EXPERIENCED A LEAK POST-OP ON A SLEEVE GASTRECTOMY PATIENT. IT IS UNKNOWN OF WETHER INTRA-OPERATIVE EVERYTHING LOOKED TO BE A SUCCESS. THE LEAK WAS IN THE MIDDLE PORTION OF THE STOMACH. THE PATIENT WAS READMITTED TO THE HOSPITAL.
03-01-2020 | Malfunction | ethicon endo-surgery | 21-01-2020 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A COLOSTOMY REVERSAL, ONLY A SMALL AMOUNT OF THE STAPLES CREATED THE ANASTOMOSIS. NO CRUNCH WAS HEARD. CASE COMPLETED BY HAND SEWING THE ANASTOMOSIS. THE CASE WAS EXTENDED BETWEEN AND HOUR AND TWO HOURS. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
03-01-2020 | Malfunction | ethicon endo-surgery | 21-01-2020 | ILS 29MM, CURVED | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, INSTRUMENT DIDN'T WORK PROPERLY. THERE WERE NO PATIENT CONSEQUENCES.
04-01-2020 | Malfunction | ethicon endo-surgery | 21-01-2020 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DOCTOR ATTEMPTED TO CONNECT THE DEVICE TO THE ANVIL, HOWEVER, THE DOCTOR STATED HE COULD NOT SEE THE ORANGE INDICATOR ON THE ANVIL DUE TO AN ABUNDANCE OF TISSUE. WHEN ATTEMPTING TO DIAL DOWN THE ORANGE INDICATOR INTO THE GREEN RANGE, THE DEVICE FAILED. THE DOCTOR REMOVED THE DEVICE AND ATTEMPTED THE PROCEDURE WITH A SECOND LIKE DEVICE. THE SAME ISSUE OCCURRED WITH THE SECOND LIKE CIRCULAR STAPLER. THE DOCTOR REMOVED THE SECOND DEVICE AND USED A CONTOUR STAPLER TO OBTAIN A NEW STAPLE LINE AND LESSEN THE TISSUE MASS. THIRD STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THIS IS ALL THE INFORMATION KNOWN/AVAILABLE REGARDING THIS EVENT.
03-01-2020 | Malfunction | ethicon endo-surgery | 21-01-2020 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Defective Device | A0203 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DOCTOR ATTEMPTED TO CONNECT THE DEVICE TO THE ANVIL, HOWEVER, THE DOCTOR STATED HE COULD NOT SEE THE ORANGE INDICATOR ON THE ANVIL DUE TO AN ABUNDANCE OF TISSUE. WHEN ATTEMPTING TO DIAL DOWN THE ORANGE INDICATOR INTO THE GREEN RANGE, THE DEVICE FAILED. THE DOCTOR REMOVED THE DEVICE AND ATTEMPTED THE PROCEDURE WITH A SECOND LIKE DEVICE. THE SAME ISSUE OCCURRED WITH THE SECOND LIKE CIRCULAR STAPLER. THE DOCTOR REMOVED THE SECOND DEVICE AND USED A CONTOUR STAPLER TO OBTAIN A NEW STAPLE LINE AND LESSEN THE TISSUE MASS. THIRD STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. THIS IS ALL THE INFORMATION KNOWN/AVAILABLE REGARDING THIS EVENT.
01-01-2019 | Malfunction | ethicon endo-surgery | 21-01-2020 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY, AFTER DISSECTION, THE VESSELS WERE RETRACTED BY SILICON SLING AND TRANSECTED BY STAPLER AND RELOAD. AFTER THE TRANSECTION, THE TRANSECTION LINE STARTED OOZING AND BLEEDING "MINORITY." AMOUNT OF BLOOD LOSS WAS NOT MENTIONED BUT REPORTED AS INSIGNIFICANT AND THE PATIENT DID NOT RECEIVE A TRANSFUSION. SEVERAL STAPLES WERE OBSERVED TO BE LOOSE AND FELL OFF. A METAL CLIP WAS APPLIED AND THE BLEEDING STOPPED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED AND NO PATIENT CONSEQUENCES WERE REPORTED.
24-12-2019 | Injury | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE FIRED BUT THE STAPLES WERE LOOSE AND DID NOT FORM CORRECTLY. THE SURGEON PERFORMED MANUAL SUTURING DUE TO THE DEVICE FAILURE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE DEVICE FAILURE LED TO OPEN SURGERY. THE EVENT LEAD TO EXTENDED PATIENT HOSPITALIZATION.
24-12-2019 | Injury | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE FIRED BUT THE STAPLES WERE LOOSE AND DID NOT FORM CORRECTLY. THE SURGEON PERFORMED MANUAL SUTURING DUE TO THE DEVICE FAILURE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE DEVICE FAILURE LED TO OPEN SURGERY. THE EVENT LEAD TO EXTENDED PATIENT HOSPITALIZATION.
30-12-2019 | Malfunction | ethicon endo-surgery | 21-01-2020 | POWERED PLUS STANDARD 45 | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, THE STAPLER WOULD NOT OPEN AFTER IT WAS FIRED. IT WAS NOT ON TISSUE WHEN IT WAS FIRED AND THE REVERSE BUTTON WOULD NOT REVERSE THE BLADE TO HOME POSITION TO OPEN THE JAWS. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED AND THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
30-12-2019 | Malfunction | ethicon endo-surgery | 21-01-2020 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, THE STAPLER WOULD NOT OPEN AFTER IT WAS FIRED. IT WAS NOT ON TISSUE WHEN IT WAS FIRED AND THE REVERSE BUTTON WOULD NOT REVERSE THE BLADE TO HOME POSITION TO OPEN THE JAWS. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED AND THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
24-12-2019 | Injury | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN UNRELATED TO THE DEVICE PROBLEM. THE EVENT LEADS TO EXTENDED PATIENT HOSPITALIZATION.
24-12-2019 | Injury | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN UNRELATED TO THE DEVICE PROBLEM. THE EVENT LEADS TO EXTENDED PATIENT HOSPITALIZATION.
24-12-2019 | Injury | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE FIRED BUT THE STAPLES WERE LOOSE AND DID NOT FORM CORRECTLY. THE SURGEON PERFORMED MANUAL SUTURING DUE TO THE DEVICE FAILURE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE DEVICE FAILURE LED TO OPEN SURGERY. THE EVENT LEAD TO EXTENDED PATIENT HOSPITALIZATION.
24-12-2019 | Injury | us surgical puerto rico | 21-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE FIRED BUT THE STAPLES WERE LOOSE AND DID NOT FORM CORRECTLY. THE SURGEON PERFORMED MANUAL SUTURING DUE TO THE DEVICE FAILURE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE DEVICE FAILURE LED TO OPEN SURGERY. THE EVENT LEAD TO EXTENDED PATIENT HOSPITALIZATION.
25-12-2019 | Malfunction | ethicon endo-surgery | 21-01-2020 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, AFTER THE DEVICE WAS FIRED, IT WAS LOCKED AND COULD NOT OPEN. IT WAS UNKNOWN HOW TO REMOVE THE DEVICE. THE JAWS NEVER OPENED. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-12-2019 | Malfunction | ethicon endo-surgery | 21-01-2020 | ECHELON FLEX 45 STANDARD | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, AFTER THE DEVICE WAS FIRED, IT WAS LOCKED AND COULD NOT OPEN. IT WAS UNKNOWN HOW TO REMOVE THE DEVICE. THE JAWS NEVER OPENED. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2019 | Injury | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN UNRELATED TO THE DEVICE PROBLEM. THE EVENT LEADS TO EXTENDED PATIENT HOSPITALIZATION.
24-12-2019 | Injury | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN UNRELATED TO THE DEVICE PROBLEM. THE EVENT LEADS TO EXTENDED PATIENT HOSPITALIZATION.
24-12-2019 | Injury | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN UNRELATED TO THE DEVICE PROBLEM. THE EVENT LEADS TO EXTENDED PATIENT HOSPITALIZATION.
24-12-2019 | Injury | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN UNRELATED TO THE DEVICE PROBLEM. THE EVENT LEADS TO EXTENDED PATIENT HOSPITALIZATION.
24-12-2019 | Injury | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN UNRELATED TO THE DEVICE PROBLEM. THE EVENT LEADS TO EXTENDED PATIENT HOSPITALIZATION.
24-12-2019 | Injury | covidien lp llc north haven | 21-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE LUNG DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC TO OPEN UNRELATED TO THE DEVICE PROBLEM. THE EVENT LEADS TO EXTENDED PATIENT HOSPITALIZATION.
08-01-2020 | Injury | us surgical puerto rico | 21-01-2020 | EEA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO THE COLON IN A LAPAROSCOPIC SIGMOID RESECTION PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND BUBBLES WERE NOTICED DURING THE LEAK TEST. THE BLACK KNOB WAS HARD TO RETURN. IT WAS ALSO REPORTED THAT MALFORMED STAPLES WERE IDENTIFIED LAPAROSCOPICALLY AND THE DONUTS WERE NOT COMPLETE. THE STAPLE LINE WAS HANDSEWN IN ORDER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
08-01-2020 | Injury | us surgical puerto rico | 21-01-2020 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO THE COLON IN A LAPAROSCOPIC SIGMOID RESECTION PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND BUBBLES WERE NOTICED DURING THE LEAK TEST. THE BLACK KNOB WAS HARD TO RETURN. IT WAS ALSO REPORTED THAT MALFORMED STAPLES WERE IDENTIFIED LAPAROSCOPICALLY AND THE DONUTS WERE NOT COMPLETE. THE STAPLE LINE WAS HANDSEWN IN ORDER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
08-01-2020 | Injury | us surgical puerto rico | 21-01-2020 | EEA | Difficult to Open or Close | A051104 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO THE COLON IN A LAPAROSCOPIC SIGMOID RESECTION PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND BUBBLES WERE NOTICED DURING THE LEAK TEST. THE BLACK KNOB WAS HARD TO RETURN. IT WAS ALSO REPORTED THAT MALFORMED STAPLES WERE IDENTIFIED LAPAROSCOPICALLY AND THE DONUTS WERE NOT COMPLETE. THE STAPLE LINE WAS HANDSEWN IN ORDER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
08-01-2020 | Injury | us surgical puerto rico | 21-01-2020 | EEA | Gas/Air Leak | A050402 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO THE COLON IN A LAPAROSCOPIC SIGMOID RESECTION PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND BUBBLES WERE NOTICED DURING THE LEAK TEST. THE BLACK KNOB WAS HARD TO RETURN. IT WAS ALSO REPORTED THAT MALFORMED STAPLES WERE IDENTIFIED LAPAROSCOPICALLY AND THE DONUTS WERE NOT COMPLETE. THE STAPLE LINE WAS HANDSEWN IN ORDER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM.
01-01-2020 | Malfunction | ethicon endo-surgery | 21-01-2020 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT THE STAPLER IN A SLEEVE GASTRECTOMY CASE DRAGGED (TISSUE SLIPPAGE) WHILE FIRING AND IT CAME OUT OF THE TISSUE CREATING MALFORMED STAPLES, AND IT CUT WITHOUT STAPLING. THE STAPLE LINE WAS INTERRUPTED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
13-12-2019 | Malfunction |  | 22-01-2020 | OPTIFIX AT ABSORBABLE FIXATION SYSTEM WITH ARTICULATING TECHNOLOGY, 30 FASTENERS | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INGUINAL PROCEDURE USING AN OPTIFIX AT, THE TACK SPLIT IN HALF AFTER THE THIRD TACK WAS FIRED. THE OPTIFIX AT WAS ONLY USED IN STRAIGHT MODE WHEN TRYING TO FIXATE A 3DMAX. IT WAS REPORTED THAT IT WAS TACKED OUT LATERAL AND THE SURGEON WAS ABLE TO RETRIEVE THE PIECES FROM THE BODY. SURGEON WAS ABLE TO COMPLETE THE CASE BY SWITCHING TO AN OPTIFIX DEVICE TO OBTAIN FIXATION. THIS IS THE SURGEONS SECOND CASE USING THE DEVICE. AS REPORTED, THERE WAS NO INJURY OR IMPACT TO THE PATIENT.
13-12-2019 | Malfunction |  | 22-01-2020 | OPTIFIX AT ABSORBABLE FIXATION SYSTEM WITH ARTICULATING TECHNOLOGY, 30 FASTENERS | Insufficient Information | E2401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INGUINAL PROCEDURE USING AN OPTIFIX AT, THE TACK SPLIT IN HALF AFTER THE THIRD TACK WAS FIRED. THE OPTIFIX AT WAS ONLY USED IN STRAIGHT MODE WHEN TRYING TO FIXATE A 3DMAX. IT WAS REPORTED THAT IT WAS TACKED OUT LATERAL AND THE SURGEON WAS ABLE TO RETRIEVE THE PIECES FROM THE BODY. SURGEON WAS ABLE TO COMPLETE THE CASE BY SWITCHING TO AN OPTIFIX DEVICE TO OBTAIN FIXATION. THIS IS THE SURGEONS SECOND CASE USING THE DEVICE. AS REPORTED, THERE WAS NO INJURY OR IMPACT TO THE PATIENT.
06-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY PROCEDURE, UPON LAST STAPLING ON GASTRIC TISSUE, THE LOAD HAS LOCKED. THE SURGEON CHANGE THE RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2020 | Injury | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GROUND NODULE PROCEDURE, WHILE DETACHING THE BRONCHUS, THERE WAS POOR STAPLE FORMATION WHEREIN THE STAPLES DID NOT FORM PROPERLY AND IT DID NOT FORM "B" SHAPE. THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY WHICH WAS RESOLVED BY SURGEON MANUALLY SUTURING THE DEVICE AND CHANGING THE RELOAD.
10-12-2019 | Malfunction | davol | 22-01-2020 | OPTIFIX ABSORBABLE FIXATION SYSTEM | Defective Device | A0203 | No Known Impact Or Consequence To Patient | Event Description: THE TACKER TACKED 2-3 TIMES THEN WOULD NOT TACK ANYMORE. THE SPRING LOOKED TO BE STICKING OUT THE TIP OF THE TACKER.
01-11-2019 | Malfunction | ethicon endo-surgery | 22-01-2020 | CONTOUR CURVED CUTTER STAPLER | Defective Component | A0202 | No Consequences Or Impact To Patient | nan
01-11-2019 | Malfunction | ethicon endo-surgery | 22-01-2020 | CONTOUR CURVED CUTTER STAPLER | Misfire | A050502 | No Consequences Or Impact To Patient | nan
09-11-2019 | Malfunction | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A BARIATRIC SURGERY, IT WAS STATED THAT THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD AND STAPLER WAS ABLE TO FIRE. THE RELOAD WAS CORRECTLY ASSEMBLED AND CALIBRATED WHEN FIRING; HOWEVER, IT DID NOT COMPLETE THE STAPLING. IT WAS SAME HANDLE USED ON THE WHOLE PROCEDURE AND ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-05-2019 | Malfunction | us surgical puerto rico | 22-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLON PROCEDURE, THE RELOAD DID NOT FIRE WHICH THE SURGEON RELATES TO THE GREEN BUTTON, THAT SOMEHOW NEGATIVELY AFFECTED THE TRIGGERING OF THE RELOAD. ANOTHER RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2019 | Malfunction | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THEY CHANGED LOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-05-2019 | Malfunction | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A COLON PROCEDURE, THE RELOAD DID NOT FIRE WHICH THE SURGEON RELATES TO THE GREEN BUTTON, THAT SOMEHOW NEGATIVELY AFFECTED THE TRIGGERING OF THE RELOAD. ANOTHER RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | EEA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, IT WAS STATED THAT THERE WAS A POOR STAPLE FORMATION, ONE SIDE OF THE DONUT WAS COMPLETE BUT THE RECTAL SIDE WAS NOT AND IT REQUIRED TO USE MANY STITCHES TO REPAIR. ANASTOMOSIS WAS OVER SEWN TO COMPLETE THE CASE.
09-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, IT WAS STATED THAT THERE WAS A POOR STAPLE FORMATION, ONE SIDE OF THE DONUT WAS COMPLETE BUT THE RECTAL SIDE WAS NOT AND IT REQUIRED TO USE MANY STITCHES TO REPAIR. ANASTOMOSIS WAS OVER SEWN TO COMPLETE THE CASE.
22-01-2020 | Malfunction | us surgical puerto rico | 22-01-2020 | ENDO GIA ULTRA | Mechanical Jam | A0506 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE JAMMED BEFORE USING ON THE PATIENT. THERE WAS NO PATIENT INVOLVEMENT.
22-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | UNKNOWN ENDO GIA SULU | Mechanical Jam | A0506 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE JAMMED BEFORE USING ON THE PATIENT. THERE WAS NO PATIENT INVOLVEMENT.
08-01-2020 | Malfunction | us surgical puerto rico | 22-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RADICAL GASTRIC CANCER, WHILE WORKING ON THE SEVERED GASTRIC TISSUE, THE DEVICE WAS NOT ABLE TO BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
08-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RADICAL GASTRIC CANCER, WHILE WORKING ON THE SEVERED GASTRIC TISSUE, THE DEVICE WAS NOT ABLE TO BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
09-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | SIGNIA | Failure to Charge | A070603 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO LOBECTOMY, THE PRODUCT WAS NOT USED ON THE PATIENT, WHEN SETTING AND TRYING TO CONNECT THE CARTRIDGE, THE SCREEN DISPLAY BECAME DARK, AND THE HANDLE BECAME UNABLE TO WORK AT ALL. RESET BY POWERED APPROACH WAS ATTEMPTED, BUT THE HANDLE COULD NOT BE RESTARTED. IN ADDITION, ALTHOUGH A CARTRIDGE WAS CONNECTED, CALIBRATION WAS NOT PERFORMED, SO A NEW MANUAL HANDLE WAS OPENED AND USED, AND THE OPERATION WAS COMPLETED. AFTER RECEIVING THE CONTACT, THE SALES REPRESENTATIVE VISITED THE OPERATION ROOM AND CHECKED THE REPORTED PRODUCT, BUT IT DID NOT WORK AT ALL. THE HANDLE WAS RETURNED TO CHARGER, BUT IT COULD NOT BE RESET. ALSO, WHEN IT WAS TAKEN OUT FROM THE SHELL, THE HANDLE ITSELF WAS HOT. THERE WAS NO PATIENT INJURY.
09-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | SIGNIA | Overheating of Device | A1005 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO LOBECTOMY, THE PRODUCT WAS NOT USED ON THE PATIENT, WHEN SETTING AND TRYING TO CONNECT THE CARTRIDGE, THE SCREEN DISPLAY BECAME DARK, AND THE HANDLE BECAME UNABLE TO WORK AT ALL. RESET BY POWERED APPROACH WAS ATTEMPTED, BUT THE HANDLE COULD NOT BE RESTARTED. IN ADDITION, ALTHOUGH A CARTRIDGE WAS CONNECTED, CALIBRATION WAS NOT PERFORMED, SO A NEW MANUAL HANDLE WAS OPENED AND USED, AND THE OPERATION WAS COMPLETED. AFTER RECEIVING THE CONTACT, THE SALES REPRESENTATIVE VISITED THE OPERATION ROOM AND CHECKED THE REPORTED PRODUCT, BUT IT DID NOT WORK AT ALL. THE HANDLE WAS RETURNED TO CHARGER, BUT IT COULD NOT BE RESET. ALSO, WHEN IT WAS TAKEN OUT FROM THE SHELL, THE HANDLE ITSELF WAS HOT. THERE WAS NO PATIENT INJURY.
09-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | SIGNIA | Calibration Problem | A08 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO LOBECTOMY, THE PRODUCT WAS NOT USED ON THE PATIENT, WHEN SETTING AND TRYING TO CONNECT THE CARTRIDGE, THE SCREEN DISPLAY BECAME DARK, AND THE HANDLE BECAME UNABLE TO WORK AT ALL. RESET BY POWERED APPROACH WAS ATTEMPTED, BUT THE HANDLE COULD NOT BE RESTARTED. IN ADDITION, ALTHOUGH A CARTRIDGE WAS CONNECTED, CALIBRATION WAS NOT PERFORMED, SO A NEW MANUAL HANDLE WAS OPENED AND USED, AND THE OPERATION WAS COMPLETED. AFTER RECEIVING THE CONTACT, THE SALES REPRESENTATIVE VISITED THE OPERATION ROOM AND CHECKED THE REPORTED PRODUCT, BUT IT DID NOT WORK AT ALL. THE HANDLE WAS RETURNED TO CHARGER, BUT IT COULD NOT BE RESET. ALSO, WHEN IT WAS TAKEN OUT FROM THE SHELL, THE HANDLE ITSELF WAS HOT. THERE WAS NO PATIENT INJURY.
09-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | SIGNIA | Complete Loss of Power | A070801 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO LOBECTOMY, THE PRODUCT WAS NOT USED ON THE PATIENT, WHEN SETTING AND TRYING TO CONNECT THE CARTRIDGE, THE SCREEN DISPLAY BECAME DARK, AND THE HANDLE BECAME UNABLE TO WORK AT ALL. RESET BY POWERED APPROACH WAS ATTEMPTED, BUT THE HANDLE COULD NOT BE RESTARTED. IN ADDITION, ALTHOUGH A CARTRIDGE WAS CONNECTED, CALIBRATION WAS NOT PERFORMED, SO A NEW MANUAL HANDLE WAS OPENED AND USED, AND THE OPERATION WAS COMPLETED. AFTER RECEIVING THE CONTACT, THE SALES REPRESENTATIVE VISITED THE OPERATION ROOM AND CHECKED THE REPORTED PRODUCT, BUT IT DID NOT WORK AT ALL. THE HANDLE WAS RETURNED TO CHARGER, BUT IT COULD NOT BE RESET. ALSO, WHEN IT WAS TAKEN OUT FROM THE SHELL, THE HANDLE ITSELF WAS HOT. THERE WAS NO PATIENT INJURY.
22-01-2020 | Injury | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE WHICH LEAD TO ADDITIONAL OPERATION PERFORMED TO CORRECT THE ISSUE.
22-01-2020 | Injury | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE WHICH LEAD TO ADDITIONAL OPERATION PERFORMED TO CORRECT THE ISSUE.
22-01-2020 | Injury | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE WHICH LEAD TO ADDITIONAL OPERATION PERFORMED TO CORRECT THE ISSUE.
22-01-2020 | Injury | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, THE STAPLER DID NOT FIRE EVEN THOUGH GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED. THE JAWS LOCKED ON TISSUE AND JAWS WOULD NOT OPEN. A CRACKING SOUND WAS ALSO HEARD. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY DUE TO THE REPORTED CONDITION.
22-01-2020 | Injury | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, THE STAPLER DID NOT FIRE EVEN THOUGH GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED. THE JAWS LOCKED ON TISSUE AND JAWS WOULD NOT OPEN. A CRACKING SOUND WAS ALSO HEARD. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY DUE TO THE REPORTED CONDITION.
08-01-2020 | Malfunction | us surgical puerto rico | 22-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, AFTER PREPARING THE DEVICE AND THE SURGEON TRIED TO FIRE, BUT THE DEVICE COULD NOT FIRE EVEN THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. SURGEON OPENED A NEW RELOAD AND RESECTION WAS ABLE TO BE PERFORMED WITHOUT PROBLEM. NO PATIENT INJURY.
22-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, THE STAPLER DID NOT FIRE EVEN THOUGH GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED. THE JAWS LOCKED ON TISSUE AND JAWS WOULD NOT OPEN. A CRACKING SOUND WAS ALSO HEARD. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY DUE TO THE REPORTED CONDITION.
22-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, THE STAPLER DID NOT FIRE EVEN THOUGH GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED. THE JAWS LOCKED ON TISSUE AND JAWS WOULD NOT OPEN. A CRACKING SOUND WAS ALSO HEARD. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY DUE TO THE REPORTED CONDITION.
22-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, THE STAPLER DID NOT FIRE EVEN THOUGH GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED. THE JAWS LOCKED ON TISSUE AND JAWS WOULD NOT OPEN. A CRACKING SOUND WAS ALSO HEARD. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY DUE TO THE REPORTED CONDITION.
22-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PULMONARY LOBECTOMY, POOR STAPLE FORMATION WAS OBSERVED AND STAPLE LINE WAS INCOMPLETE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW RECTAL RESECTION PROCEDURE, WHILE DETACHING THE INTESTINAL CANAL, THE SURGEON ABLE TO FIRE THE DEVICE HOWEVER THE STAPLE FAILED TO FORM PROPER SHAPE ON 2 RELOADS. THE SURGEON THEN USED ANOTHER DEVICE RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | us surgical puerto rico | 22-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO DETACH THE RECTUM DURING LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. REPLACED WITH OTHER PRODUCT. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | covidien lp llc north haven | 22-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO DETACH THE RECTUM DURING LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. REPLACED WITH OTHER PRODUCT. THERE WAS NO PATIENT INJURY.
22-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | GIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE WHICH LEAD TO ADDITIONAL OPERATION PERFORMED TO CORRECT THE ISSUE.
22-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | GIA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE WHICH LEAD TO ADDITIONAL OPERATION PERFORMED TO CORRECT THE ISSUE.
22-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE WHICH LEAD TO ADDITIONAL OPERATION PERFORMED TO CORRECT THE ISSUE.
26-11-2019 | Malfunction | rotation medical | 22-01-2020 | TENDON ANCHORS 8 | Material Separation | A0413 | No Known Impact Or Consequence To Patient | Event Description: IT WAS REPORTED THAT, DURING A ROTATOR CUFF REPAIR SURGERY, SEVERAL TENDON ANCHORS BROKE WHEN BEING IMPLANTED IN THE PATIENT. SOME OF THE ANCHORS WERE RETRIEVED FROM THE PATIENT WITH A GRASPER. SEVEN TENDON ANCHORS, FROM TWO PUCKS, WERE USED TO COMPLETE THE SURGERY, WHICH WAS DELAYED LESS THAN 30 MIN AS CONSEQUENCE OF THE ALLEGATION. THE PATIENT WAS NOT INJURED. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY.
22-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | EEA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE WHICH LEAD TO ADDITIONAL OPERATION PERFORMED TO CORRECT THE ISSUE.
22-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | EEA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE WHICH LEAD TO ADDITIONAL OPERATION PERFORMED TO CORRECT THE ISSUE.
22-01-2020 | Injury | us surgical puerto rico | 22-01-2020 | EEA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE WHICH LEAD TO ADDITIONAL OPERATION PERFORMED TO CORRECT THE ISSUE.
05-12-2019 | Injury | intuitive surgical | 22-01-2020 | SUREFORM | Insufficient Information | E2401 | Injury | Event Description: IT WAS REPORTED THAT SOME WEEKS AFTER COMPLETION OF A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT RETURNED TO THE HOSPITAL DUE TO A HOLE IN THE STAPLE LINE. AS A RESULT, THE PATIENT UNDERWENT A SECOND PROCEDURE TO REPAIR THE HOLE. HOWEVER, AT THIS TIME, THE CAUSE OF THE STAPLE LINE HOLE IS UNKNOWN. AN INTUITIVE SURGICAL, INC. (ISI) CLINICAL SALES REPRESENTATIVE (CSR) WAS CONTACTED BY THE SURGEON TO REPORT THE EVENT. ON (B)(6) 2019, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE, WHEREIN A SUREFORM STAPLER 60 INSTRUMENT AND BLUE SUREFORM 60 RELOADS WERE USED. THERE WERE NO REPORTS OF ANY INTRA-OPERATIVE COMPLICATIONS DURING THIS INITIAL PROCEDURE. SOME WEEKS POSTOPERATIVE, THE PATIENT REPORTED SYMPTOMS OF NAUSEA, AND A CT SCAN WAS PERFORMED. ON (B)(6) 2019, THE PATIENT UNDERWENT A SECOND PROCEDURE, AND A STAPLE LINE HOLE WAS IDENTIFIED AND REPAIRED. THE PATIENT WAS REPORTED TO BE RECOVERING IN THE ICU. THE CSR CONFIRMED THAT THE SUREFORM STAPLER 60 INSTRUMENT WOULD NOT BE RETURNED TO ISI FOR EVALUATION. THE SURGEON "BELIEVES" THE SUREFORM STAPLER 60 INSTRUMENT ALLEGEDLY CAUSED OR CONTRIBUTED TO THE REPORTED INJURY. THERE IS A VIDEO RECORDING OF THE PROCEDURE; HOWEVER, ISI WAS NOT PROVIDED A COPY FOR REVIEW. ISI MADE MULTIPLE FOLLOW-UP ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION. HOWEVER, NO FURTHER DETAILS HAVE BEEN RECEIVED AS OF THE DATE OF THIS REPORT.
06-01-2020 | Malfunction | ethicon endo-surgery | 22-01-2020 | ETHICON ENDOPATH ETS FLEX 45 MM SURGICAL STAPLER | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: PT WAS HAVING AN APPENDECTOMY FOR ACUTE APPENDICITIS. PER OP REPORT, "THE EMAIL BOWEL WAS RETRACTED OUT OF THE RIGHT LOWER QUADRANT AND THE APPENDIX IDENTIFIED. THE ILEUM ENTERING THE CECUM WAS ALSO IDENTIFIED. THE APPENDIX WAS GRASPED AND A WINDOW MADE AT THE BASE OF THE APPENDIX AND CECUM USING A MARYLAND DISSECTOR. A 45 MM BLUE ENDOGIA STAPLER WAS INSERTED. UPON INSERTION IT APPEARED THAT SOME OF THE STAPLES WERE DISLODGED WITHOUT FIRING OF THE STAPLER. IT WAS IMMEDIATELY REMOVED, "THE STAPLES THAT FELL OUT" WERE ABLE TO BE RETRIEVED / REMOVED FROM THE PT'S ABDOMEN AND NO HARM TO PT OCCURRED. FDA SAFETY REPORT ID# (B)(4).
06-01-2020 | Malfunction | ethicon endo-surgery | 22-01-2020 | ETHICON ENDOPATH ETS FLEX 45 MM SURGICAL STAPLER | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: PT WAS HAVING AN APPENDECTOMY FOR ACUTE APPENDICITIS. PER OP REPORT, "THE EMAIL BOWEL WAS RETRACTED OUT OF THE RIGHT LOWER QUADRANT AND THE APPENDIX IDENTIFIED. THE ILEUM ENTERING THE CECUM WAS ALSO IDENTIFIED. THE APPENDIX WAS GRASPED AND A WINDOW MADE AT THE BASE OF THE APPENDIX AND CECUM USING A MARYLAND DISSECTOR. A 45 MM BLUE ENDOGIA STAPLER WAS INSERTED. UPON INSERTION IT APPEARED THAT SOME OF THE STAPLES WERE DISLODGED WITHOUT FIRING OF THE STAPLER. IT WAS IMMEDIATELY REMOVED, "THE STAPLES THAT FELL OUT" WERE ABLE TO BE RETRIEVED / REMOVED FROM THE PT'S ABDOMEN AND NO HARM TO PT OCCURRED. FDA SAFETY REPORT ID# (B)(4).
06-01-2020 | Malfunction | ethicon endo-surgery | 22-01-2020 | ETHICON ENDOPATH ETS FLEX 45 MM SURGICAL STAPLER | Device Dislodged or Dislocated | A051201 | No Consequences Or Impact To Patient | Event Description: PT WAS HAVING AN APPENDECTOMY FOR ACUTE APPENDICITIS. PER OP REPORT, "THE EMAIL BOWEL WAS RETRACTED OUT OF THE RIGHT LOWER QUADRANT AND THE APPENDIX IDENTIFIED. THE ILEUM ENTERING THE CECUM WAS ALSO IDENTIFIED. THE APPENDIX WAS GRASPED AND A WINDOW MADE AT THE BASE OF THE APPENDIX AND CECUM USING A MARYLAND DISSECTOR. A 45 MM BLUE ENDOGIA STAPLER WAS INSERTED. UPON INSERTION IT APPEARED THAT SOME OF THE STAPLES WERE DISLODGED WITHOUT FIRING OF THE STAPLER. IT WAS IMMEDIATELY REMOVED, "THE STAPLES THAT FELL OUT" WERE ABLE TO BE RETRIEVED / REMOVED FROM THE PT'S ABDOMEN AND NO HARM TO PT OCCURRED. FDA SAFETY REPORT ID# (B)(4).
06-01-2020 | Malfunction | ethicon endo-surgery | 22-01-2020 | ETHICON ENDOPATH ETS FLEX 45 MM SURGICAL STAPLER | Device Fell | A051203 | No Consequences Or Impact To Patient | Event Description: PT WAS HAVING AN APPENDECTOMY FOR ACUTE APPENDICITIS. PER OP REPORT, "THE EMAIL BOWEL WAS RETRACTED OUT OF THE RIGHT LOWER QUADRANT AND THE APPENDIX IDENTIFIED. THE ILEUM ENTERING THE CECUM WAS ALSO IDENTIFIED. THE APPENDIX WAS GRASPED AND A WINDOW MADE AT THE BASE OF THE APPENDIX AND CECUM USING A MARYLAND DISSECTOR. A 45 MM BLUE ENDOGIA STAPLER WAS INSERTED. UPON INSERTION IT APPEARED THAT SOME OF THE STAPLES WERE DISLODGED WITHOUT FIRING OF THE STAPLER. IT WAS IMMEDIATELY REMOVED, "THE STAPLES THAT FELL OUT" WERE ABLE TO BE RETRIEVED / REMOVED FROM THE PT'S ABDOMEN AND NO HARM TO PT OCCURRED. FDA SAFETY REPORT ID# (B)(4).
10-01-2020 | Malfunction | ethicon inc / johnson & johnson | 22-01-2020 | ENDOPATH 45MM ARTICULATING LINEAR ETS FLE | Misfire | A050502 | No Known Impact Or Consequence To Patient | Event Description: THE PHYSICIAN WAS PERFORMING A LAPAROSCOPIC APPENDECTOMY WHEN THE STAPLER GUN MISFIRED. THE STAPLER MADE A LOUD CRUNCHING SOUND WHEN THE MISFIRE OCCURRED. THE STAPLER WAS REMOVED FROM THE FIELD AND REPLACED. NO HARM WAS NOTED TO THE PT. FDA SAFETY REPORT ID# (B)(4).
10-01-2020 | Malfunction | ethicon inc / johnson & johnson | 22-01-2020 | ENDOPATH 45MM ARTICULATING LINEAR ETS FLE | Noise, Audible | A0508 | No Known Impact Or Consequence To Patient | Event Description: THE PHYSICIAN WAS PERFORMING A LAPAROSCOPIC APPENDECTOMY WHEN THE STAPLER GUN MISFIRED. THE STAPLER MADE A LOUD CRUNCHING SOUND WHEN THE MISFIRE OCCURRED. THE STAPLER WAS REMOVED FROM THE FIELD AND REPLACED. NO HARM WAS NOTED TO THE PT. FDA SAFETY REPORT ID# (B)(4).
15-01-2020 | Malfunction | bard shnnon | 22-01-2020 | OPTIFIX ABSORBABLE FIXATIONS SYS | Failure to Fire | A050501 | No Known Impact Or Consequence To Patient | Event Description: A PRIMARY UMBILICAL HERNIA REPAIR, PRIMARY REPAIR OF LOWER MIDLINE INCISIONAL HERNIA WITH PLACEMENT OF 8 X 12 CM BARD MESH WAS BEING PERFORMED. THE OPTIFIX ABSORBABLE FIXATION DEVICE WOULD NOT FIRE THE TACKS TO HOLD THE MESH IN PLACE, THE MESH WAS THEN SUTURED IN PLACE. NO HARM TO PATIENT. NO DELAY IN PROCEDURE. PATIENT DC TO HOME THE FOLLOWING DAY. FDA SAFETY REPORT ID # (B)(4).
10-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING ON OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, THE STAPLER PINS DID NOT COME OUT PROPERLY LEADING TO AN IMPROPER ANASTOMOSIS. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
10-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | EEA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING ON OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, THE STAPLER PINS DID NOT COME OUT PROPERLY LEADING TO AN IMPROPER ANASTOMOSIS. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
10-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | EEA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING ON OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, THE STAPLER PINS DID NOT COME OUT PROPERLY LEADING TO AN IMPROPER ANASTOMOSIS. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
10-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING ON OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, THE STAPLER PINS DID NOT COME OUT PROPERLY LEADING TO AN IMPROPER ANASTOMOSIS. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
10-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | EEA | Failure to Fire | A050501 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING ON OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, THE STAPLER PINS DID NOT COME OUT PROPERLY LEADING TO AN IMPROPER ANASTOMOSIS. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
10-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | EEA | Failure to Fire | A050501 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING ON OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, THE STAPLER PINS DID NOT COME OUT PROPERLY LEADING TO AN IMPROPER ANASTOMOSIS. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
10-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | EEA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING ON OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, THE STAPLER PINS DID NOT COME OUT PROPERLY LEADING TO AN IMPROPER ANASTOMOSIS. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
10-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | EEA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN FIRING THE STAPLER DURING ON OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, THE STAPLER PINS DID NOT COME OUT PROPERLY LEADING TO AN IMPROPER ANASTOMOSIS. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE.
08-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Fistula | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC GASTRIC BYPASS OBESITY SURGERY, ON THE STAPLING PHASE, THE STAPLER WAS ABLE TO FIRE. THE PATIENT WAS HOSPITALIZED AND WAS TRANSFERRED TO ANOTHER HOSPITAL TO STOP THE FISTULA. THE FISTULA WAS FOUND ON THE STAPLING LINE. THEY NEEDED TO REDO THE SURGERY TO FIND THE LEAK.
08-01-2020 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC GASTRIC BYPASS OBESITY SURGERY, ON THE STAPLING PHASE, THE STAPLER WAS ABLE TO FIRE. THE PATIENT WAS HOSPITALIZED AND WAS TRANSFERRED TO ANOTHER HOSPITAL TO STOP THE FISTULA. THE FISTULA WAS FOUND ON THE STAPLING LINE. THEY NEEDED TO REDO THE SURGERY TO FIND THE LEAK.
03-01-2020 | Malfunction | us surgical puerto rico | 23-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE TRIGGER REMAINED LOCKED. HANDLE WAS NOT ABLE TO SQUEEZE THE HANDLE SO A DIFFERENT LOADING UNIT WAS USED BUT PROBLEM PERSISTED. ONLY AFTER A NEW STAPLER WAS OPENED THAT THE SURGEON WAS ABLE TO CONTINUE AND FINISHED THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2019 | Malfunction | us surgical puerto rico | 23-01-2020 | EEA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS W/HIATAL HERNIA, THE SURGEON WENT TO ENTER THE ABDOMEN WITH THE STAPLER AND THE HAND PIECE BROKE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2019 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, 1 DAY POSTOPERATIVE FROM A LAPAROSCOPIC BYPASS, THE PATIENT HAD FEVER, TACHYCARDIA AND ABDOMINAL PAIN. AN ASTOMOTIC LEAK WAS CONFIRMED. THE PATIENT HAD AN EARLY COMPLICATION AT D1 FOR A GASTRIC BY PASS - THE BLUE TEST ON THE GASTROJEJUNAL ANASTOMOSIS WAS POSITIVE, WHICH REQUIRED TO RE-DO SURGERY. THIS RESULTED TO SOME TISSUE LOSS AND DAMAGE ON A PART OF THE GASTRIC TU BE. THE EVENT LED TO EXTENDED HOSPITALIZATION WITH A CATHETER FOR 1 MONTH.
11-12-2019 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Fever | Event Description: ACCORDING TO THE REPORTER, 1 DAY POSTOPERATIVE FROM A LAPAROSCOPIC BYPASS, THE PATIENT HAD FEVER, TACHYCARDIA AND ABDOMINAL PAIN. AN ASTOMOTIC LEAK WAS CONFIRMED. THE PATIENT HAD AN EARLY COMPLICATION AT D1 FOR A GASTRIC BY PASS - THE BLUE TEST ON THE GASTROJEJUNAL ANASTOMOSIS WAS POSITIVE, WHICH REQUIRED TO RE-DO SURGERY. THIS RESULTED TO SOME TISSUE LOSS AND DAMAGE ON A PART OF THE GASTRIC TU BE. THE EVENT LED TO EXTENDED HOSPITALIZATION WITH A CATHETER FOR 1 MONTH.
11-12-2019 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Tachycardia | Event Description: ACCORDING TO THE REPORTER, 1 DAY POSTOPERATIVE FROM A LAPAROSCOPIC BYPASS, THE PATIENT HAD FEVER, TACHYCARDIA AND ABDOMINAL PAIN. AN ASTOMOTIC LEAK WAS CONFIRMED. THE PATIENT HAD AN EARLY COMPLICATION AT D1 FOR A GASTRIC BY PASS - THE BLUE TEST ON THE GASTROJEJUNAL ANASTOMOSIS WAS POSITIVE, WHICH REQUIRED TO RE-DO SURGERY. THIS RESULTED TO SOME TISSUE LOSS AND DAMAGE ON A PART OF THE GASTRIC TU BE. THE EVENT LED TO EXTENDED HOSPITALIZATION WITH A CATHETER FOR 1 MONTH.
11-12-2019 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 1 DAY POSTOPERATIVE FROM A LAPAROSCOPIC BYPASS, THE PATIENT HAD FEVER, TACHYCARDIA AND ABDOMINAL PAIN. AN ASTOMOTIC LEAK WAS CONFIRMED. THE PATIENT HAD AN EARLY COMPLICATION AT D1 FOR A GASTRIC BY PASS - THE BLUE TEST ON THE GASTROJEJUNAL ANASTOMOSIS WAS POSITIVE, WHICH REQUIRED TO RE-DO SURGERY. THIS RESULTED TO SOME TISSUE LOSS AND DAMAGE ON A PART OF THE GASTRIC TU BE. THE EVENT LED TO EXTENDED HOSPITALIZATION WITH A CATHETER FOR 1 MONTH.
11-12-2019 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Gas/Air Leak | A050402 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, 1 DAY POSTOPERATIVE FROM A LAPAROSCOPIC BYPASS, THE PATIENT HAD FEVER, TACHYCARDIA AND ABDOMINAL PAIN. AN ASTOMOTIC LEAK WAS CONFIRMED. THE PATIENT HAD AN EARLY COMPLICATION AT D1 FOR A GASTRIC BY PASS - THE BLUE TEST ON THE GASTROJEJUNAL ANASTOMOSIS WAS POSITIVE, WHICH REQUIRED TO RE-DO SURGERY. THIS RESULTED TO SOME TISSUE LOSS AND DAMAGE ON A PART OF THE GASTRIC TU BE. THE EVENT LED TO EXTENDED HOSPITALIZATION WITH A CATHETER FOR 1 MONTH.
11-12-2019 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Gas/Air Leak | A050402 | Fever | Event Description: ACCORDING TO THE REPORTER, 1 DAY POSTOPERATIVE FROM A LAPAROSCOPIC BYPASS, THE PATIENT HAD FEVER, TACHYCARDIA AND ABDOMINAL PAIN. AN ASTOMOTIC LEAK WAS CONFIRMED. THE PATIENT HAD AN EARLY COMPLICATION AT D1 FOR A GASTRIC BY PASS - THE BLUE TEST ON THE GASTROJEJUNAL ANASTOMOSIS WAS POSITIVE, WHICH REQUIRED TO RE-DO SURGERY. THIS RESULTED TO SOME TISSUE LOSS AND DAMAGE ON A PART OF THE GASTRIC TU BE. THE EVENT LED TO EXTENDED HOSPITALIZATION WITH A CATHETER FOR 1 MONTH.
11-12-2019 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Gas/Air Leak | A050402 | Tachycardia | Event Description: ACCORDING TO THE REPORTER, 1 DAY POSTOPERATIVE FROM A LAPAROSCOPIC BYPASS, THE PATIENT HAD FEVER, TACHYCARDIA AND ABDOMINAL PAIN. AN ASTOMOTIC LEAK WAS CONFIRMED. THE PATIENT HAD AN EARLY COMPLICATION AT D1 FOR A GASTRIC BY PASS - THE BLUE TEST ON THE GASTROJEJUNAL ANASTOMOSIS WAS POSITIVE, WHICH REQUIRED TO RE-DO SURGERY. THIS RESULTED TO SOME TISSUE LOSS AND DAMAGE ON A PART OF THE GASTRIC TU BE. THE EVENT LED TO EXTENDED HOSPITALIZATION WITH A CATHETER FOR 1 MONTH.
11-12-2019 | Injury | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Gas/Air Leak | A050402 | No Code Available | Event Description: ACCORDING TO THE REPORTER, 1 DAY POSTOPERATIVE FROM A LAPAROSCOPIC BYPASS, THE PATIENT HAD FEVER, TACHYCARDIA AND ABDOMINAL PAIN. AN ASTOMOTIC LEAK WAS CONFIRMED. THE PATIENT HAD AN EARLY COMPLICATION AT D1 FOR A GASTRIC BY PASS - THE BLUE TEST ON THE GASTROJEJUNAL ANASTOMOSIS WAS POSITIVE, WHICH REQUIRED TO RE-DO SURGERY. THIS RESULTED TO SOME TISSUE LOSS AND DAMAGE ON A PART OF THE GASTRIC TU BE. THE EVENT LED TO EXTENDED HOSPITALIZATION WITH A CATHETER FOR 1 MONTH.
10-01-2020 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 23-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A STRAIGHT AMPUTATION, DURING PROXIMAL SECTION OF THE RECTUM, WHEN THE DEVICE WAS OPENED AND ASSEMBLED, IT DID NOT WORK WHEN TRIGGERED. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. TRIGGER SHOWED A LOT OF RESISTANCE. SURGEON WITHDREW AND CHANGED LOADS AND CONTINUED WITHOUT WORKING. SURGEON REPLACED BY SUBSTITUTING DEVICE. IT TOOK ABOUT 30 MINUTES FOR THE EXCHANGE OF THE RELOADS AND SUBSEQUENT REPLACEMENT OF THE STAPLER. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 23-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A STRAIGHT AMPUTATION, DURING PROXIMAL SECTION OF THE RECTUM, WHEN THE DEVICE WAS OPENED AND ASSEMBLED, IT DID NOT WORK WHEN TRIGGERED. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. TRIGGER SHOWED A LOT OF RESISTANCE. SURGEON WITHDREW AND CHANGED LOADS AND CONTINUED WITHOUT WORKING. SURGEON REPLACED BY SUBSTITUTING DEVICE. IT TOOK ABOUT 30 MINUTES FOR THE EXCHANGE OF THE RELOADS AND SUBSEQUENT REPLACEMENT OF THE STAPLER. THERE WAS NO PATIENT INJURY.
19-12-2019 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Loss of or Failure to Bond | A040102 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE STAPLE DID NOT FORM "B" SHAPE. IT WAS REPORTED THAT THE STAPLE BECAME LOOSE IN THE ABDOMINAL CAVITY AND SEAM DID NOT CLOSED AT ALL. THE LOOSE STAPLES WERE RETRIEVED BY USING SUCTION. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE PROCEDURE AND A NEW STAPLE LINE WAS CREATED. THE PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY
19-12-2019 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE STAPLE DID NOT FORM "B" SHAPE. IT WAS REPORTED THAT THE STAPLE BECAME LOOSE IN THE ABDOMINAL CAVITY AND SEAM DID NOT CLOSED AT ALL. THE LOOSE STAPLES WERE RETRIEVED BY USING SUCTION. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE PROCEDURE AND A NEW STAPLE LINE WAS CREATED. THE PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY
19-12-2019 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE STAPLE DID NOT FORM "B" SHAPE. IT WAS REPORTED THAT THE STAPLE BECAME LOOSE IN THE ABDOMINAL CAVITY AND SEAM DID NOT CLOSED AT ALL. THE LOOSE STAPLES WERE RETRIEVED BY USING SUCTION. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE PROCEDURE AND A NEW STAPLE LINE WAS CREATED. THE PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY
19-12-2019 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE STAPLE DID NOT FORM "B" SHAPE. IT WAS REPORTED THAT THE STAPLE BECAME LOOSE IN THE ABDOMINAL CAVITY AND SEAM DID NOT CLOSED AT ALL. THE LOOSE STAPLES WERE RETRIEVED BY USING SUCTION. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE PROCEDURE AND A NEW STAPLE LINE WAS CREATED. THE PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY
19-12-2019 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE STAPLE DID NOT FORM "B" SHAPE. IT WAS REPORTED THAT THE STAPLE BECAME LOOSE IN THE ABDOMINAL CAVITY AND SEAM DID NOT CLOSED AT ALL. THE LOOSE STAPLES WERE RETRIEVED BY USING SUCTION. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE PROCEDURE AND A NEW STAPLE LINE WAS CREATED. THE PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY
19-12-2019 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE STAPLE DID NOT FORM "B" SHAPE. IT WAS REPORTED THAT THE STAPLE BECAME LOOSE IN THE ABDOMINAL CAVITY AND SEAM DID NOT CLOSED AT ALL. THE LOOSE STAPLES WERE RETRIEVED BY USING SUCTION. A NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE PROCEDURE AND A NEW STAPLE LINE WAS CREATED. THE PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY
10-01-2020 | Malfunction | us surgical puerto rico | 23-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO TISSUE ANASTOMOSIS DURING OPEN COLON LEFT HEMICOLECTOMY, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-01-2019 | Malfunction | covidien lp | 23-01-2020 | COVIDIEN | Difficult to Open or Close | A051104 | No Known Impact Or Consequence To Patient | Event Description: USING AN ENDO GIA STANDARD STAPLER ON THE RIGHT MIDDLE LUNG. STAPLER FIRED AND UNABLE TO OPEN JAWS. REPRESENTATIVE CALLED AND STAPLER AND STAPLES REPLACED. MANUFACTURER RESPONSE FOR COVIDIEN ENDO GIA ULTRA UNIVERSAL STAPLER, COVIDIEN (PER SITE REPORTER). NONE TO DATE.
10-01-2020 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT COLON HEMICOLECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE ON THE CENTRAL AREA AND POOR STAPLE FORMATION USING THE FIRST RELOAD AND HANDLE FOR TISSUE ANASTOMOSIS. ANOTHER RELOAD WAS USED TO REPAIR THE STAPLE LINE BUT THE DEVICE COULD NOT BE FIRED. IT WAS REPORTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A THIRD RELOAD WAS USED TOGETHER WITH A NEW HANDLE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-01-2020 | Malfunction | us surgical puerto rico | 23-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT COLON HEMICOLECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE ON THE CENTRAL AREA AND POOR STAPLE FORMATION USING THE FIRST RELOAD AND HANDLE FOR TISSUE ANASTOMOSIS. ANOTHER RELOAD WAS USED TO REPAIR THE STAPLE LINE BUT THE DEVICE COULD NOT BE FIRED. IT WAS REPORTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A THIRD RELOAD WAS USED TOGETHER WITH A NEW HANDLE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-07-2014 | Injury | us surgical puerto rico | 23-01-2020 | UNKNOWN EEA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A STUDY PERFORMED BETWEEN JUNE 2011 AND JUNE 2012 ABOUT SINGLE INCISION AND REDUCED PORT LAPAROSCOPIC LOW ANTERIOR RESECTION FOR RECTAL CANCER, IN 96 CONSECUTIVE CASES OF SINGLE INCISION LAPAROSCOPIC LOW ANTERIOR RESECTION (SIL¿LAR) AND REDUCED PORT LAR CASES, OF WHICH THE INITIAL APPROACH WAS SIL IN COLORECTAL SURGERY, THE COMPLICATION RATE WAS 20%. THE DEVICE WAS USED FOR THE ANASTOMOSES, WITH THE EXCEPTION OF CASES WHERE TUMOR LOCATION WAS VERY CLOSE TO THE ANUS WHICH WERE HAND-SEWN. THERE WERE SIX CASES OF ANASTOMOSIS LEAKAGE, 1 CASE OF BLEEDING, 1 CASE OF DEEP SURGICAL SITE INFECTION, 5 CASES OF POST-OPERATIVE ILEUS, 1 CASE OF CHYLE LEAKAGE, 1 CASE OF ILEOSTOMY PROLAPSE, 1 CASE OF DELIRIUM, 1 CASE OF PAROXYSMAL SUPRAVENTRICULAR TACHYCARDIA, AND 2 CASES OF URINARY RETENTION.
18-07-2014 | Injury | us surgical puerto rico | 23-01-2020 | UNKNOWN EEA | Loss of or Failure to Bond | A040102 | Unspecified Infection | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A STUDY PERFORMED BETWEEN JUNE 2011 AND JUNE 2012 ABOUT SINGLE INCISION AND REDUCED PORT LAPAROSCOPIC LOW ANTERIOR RESECTION FOR RECTAL CANCER, IN 96 CONSECUTIVE CASES OF SINGLE INCISION LAPAROSCOPIC LOW ANTERIOR RESECTION (SIL¿LAR) AND REDUCED PORT LAR CASES, OF WHICH THE INITIAL APPROACH WAS SIL IN COLORECTAL SURGERY, THE COMPLICATION RATE WAS 20%. THE DEVICE WAS USED FOR THE ANASTOMOSES, WITH THE EXCEPTION OF CASES WHERE TUMOR LOCATION WAS VERY CLOSE TO THE ANUS WHICH WERE HAND-SEWN. THERE WERE SIX CASES OF ANASTOMOSIS LEAKAGE, 1 CASE OF BLEEDING, 1 CASE OF DEEP SURGICAL SITE INFECTION, 5 CASES OF POST-OPERATIVE ILEUS, 1 CASE OF CHYLE LEAKAGE, 1 CASE OF ILEOSTOMY PROLAPSE, 1 CASE OF DELIRIUM, 1 CASE OF PAROXYSMAL SUPRAVENTRICULAR TACHYCARDIA, AND 2 CASES OF URINARY RETENTION.
18-07-2014 | Injury | us surgical puerto rico | 23-01-2020 | UNKNOWN EEA | Loss of or Failure to Bond | A040102 | Tachycardia | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A STUDY PERFORMED BETWEEN JUNE 2011 AND JUNE 2012 ABOUT SINGLE INCISION AND REDUCED PORT LAPAROSCOPIC LOW ANTERIOR RESECTION FOR RECTAL CANCER, IN 96 CONSECUTIVE CASES OF SINGLE INCISION LAPAROSCOPIC LOW ANTERIOR RESECTION (SIL¿LAR) AND REDUCED PORT LAR CASES, OF WHICH THE INITIAL APPROACH WAS SIL IN COLORECTAL SURGERY, THE COMPLICATION RATE WAS 20%. THE DEVICE WAS USED FOR THE ANASTOMOSES, WITH THE EXCEPTION OF CASES WHERE TUMOR LOCATION WAS VERY CLOSE TO THE ANUS WHICH WERE HAND-SEWN. THERE WERE SIX CASES OF ANASTOMOSIS LEAKAGE, 1 CASE OF BLEEDING, 1 CASE OF DEEP SURGICAL SITE INFECTION, 5 CASES OF POST-OPERATIVE ILEUS, 1 CASE OF CHYLE LEAKAGE, 1 CASE OF ILEOSTOMY PROLAPSE, 1 CASE OF DELIRIUM, 1 CASE OF PAROXYSMAL SUPRAVENTRICULAR TACHYCARDIA, AND 2 CASES OF URINARY RETENTION.
18-07-2014 | Injury | us surgical puerto rico | 23-01-2020 | UNKNOWN EEA | Loss of or Failure to Bond | A040102 | Urinary Retention | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A STUDY PERFORMED BETWEEN JUNE 2011 AND JUNE 2012 ABOUT SINGLE INCISION AND REDUCED PORT LAPAROSCOPIC LOW ANTERIOR RESECTION FOR RECTAL CANCER, IN 96 CONSECUTIVE CASES OF SINGLE INCISION LAPAROSCOPIC LOW ANTERIOR RESECTION (SIL¿LAR) AND REDUCED PORT LAR CASES, OF WHICH THE INITIAL APPROACH WAS SIL IN COLORECTAL SURGERY, THE COMPLICATION RATE WAS 20%. THE DEVICE WAS USED FOR THE ANASTOMOSES, WITH THE EXCEPTION OF CASES WHERE TUMOR LOCATION WAS VERY CLOSE TO THE ANUS WHICH WERE HAND-SEWN. THERE WERE SIX CASES OF ANASTOMOSIS LEAKAGE, 1 CASE OF BLEEDING, 1 CASE OF DEEP SURGICAL SITE INFECTION, 5 CASES OF POST-OPERATIVE ILEUS, 1 CASE OF CHYLE LEAKAGE, 1 CASE OF ILEOSTOMY PROLAPSE, 1 CASE OF DELIRIUM, 1 CASE OF PAROXYSMAL SUPRAVENTRICULAR TACHYCARDIA, AND 2 CASES OF URINARY RETENTION.
18-07-2014 | Injury | us surgical puerto rico | 23-01-2020 | UNKNOWN EEA | Loss of or Failure to Bond | A040102 | Prolapse | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A STUDY PERFORMED BETWEEN JUNE 2011 AND JUNE 2012 ABOUT SINGLE INCISION AND REDUCED PORT LAPAROSCOPIC LOW ANTERIOR RESECTION FOR RECTAL CANCER, IN 96 CONSECUTIVE CASES OF SINGLE INCISION LAPAROSCOPIC LOW ANTERIOR RESECTION (SIL¿LAR) AND REDUCED PORT LAR CASES, OF WHICH THE INITIAL APPROACH WAS SIL IN COLORECTAL SURGERY, THE COMPLICATION RATE WAS 20%. THE DEVICE WAS USED FOR THE ANASTOMOSES, WITH THE EXCEPTION OF CASES WHERE TUMOR LOCATION WAS VERY CLOSE TO THE ANUS WHICH WERE HAND-SEWN. THERE WERE SIX CASES OF ANASTOMOSIS LEAKAGE, 1 CASE OF BLEEDING, 1 CASE OF DEEP SURGICAL SITE INFECTION, 5 CASES OF POST-OPERATIVE ILEUS, 1 CASE OF CHYLE LEAKAGE, 1 CASE OF ILEOSTOMY PROLAPSE, 1 CASE OF DELIRIUM, 1 CASE OF PAROXYSMAL SUPRAVENTRICULAR TACHYCARDIA, AND 2 CASES OF URINARY RETENTION.
18-07-2014 | Injury | us surgical puerto rico | 23-01-2020 | UNKNOWN EEA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A STUDY PERFORMED BETWEEN JUNE 2011 AND JUNE 2012 ABOUT SINGLE INCISION AND REDUCED PORT LAPAROSCOPIC LOW ANTERIOR RESECTION FOR RECTAL CANCER, IN 96 CONSECUTIVE CASES OF SINGLE INCISION LAPAROSCOPIC LOW ANTERIOR RESECTION (SIL¿LAR) AND REDUCED PORT LAR CASES, OF WHICH THE INITIAL APPROACH WAS SIL IN COLORECTAL SURGERY, THE COMPLICATION RATE WAS 20%. THE DEVICE WAS USED FOR THE ANASTOMOSES, WITH THE EXCEPTION OF CASES WHERE TUMOR LOCATION WAS VERY CLOSE TO THE ANUS WHICH WERE HAND-SEWN. THERE WERE SIX CASES OF ANASTOMOSIS LEAKAGE, 1 CASE OF BLEEDING, 1 CASE OF DEEP SURGICAL SITE INFECTION, 5 CASES OF POST-OPERATIVE ILEUS, 1 CASE OF CHYLE LEAKAGE, 1 CASE OF ILEOSTOMY PROLAPSE, 1 CASE OF DELIRIUM, 1 CASE OF PAROXYSMAL SUPRAVENTRICULAR TACHYCARDIA, AND 2 CASES OF URINARY RETENTION.
10-01-2020 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT COLON HEMICOLECTOMY PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE ON THE CENTRAL AREA AND POOR STAPLE FORMATION USING THE FIRST RELOAD AND HANDLE FOR TISSUE ANASTOMOSIS. ANOTHER RELOAD WAS USED TO REPAIR THE STAPLE LINE BUT THE DEVICE COULD NOT BE FIRED. IT WAS REPORTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A THIRD RELOAD WAS USED TOGETHER WITH A NEW HANDLE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
03-11-2015 | Injury | us surgical puerto rico | 23-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN JULY 2000 AND AUGUST 2012. PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY (ILE) WERE IDENTIFIED FROM A PROSPECTIVE DATABASE. ANASTOMOTIC LEAKAGE OCCURRED IN 5.6%.
03-11-2015 | Injury | us surgical puerto rico | 23-01-2020 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN JULY 2000 AND AUGUST 2012. PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY (ILE) WERE IDENTIFIED FROM A PROSPECTIVE DATABASE. ANASTOMOTIC LEAKAGE OCCURRED IN 5.6%.
10-01-2020 | Malfunction | us surgical puerto rico | 23-01-2020 | EEA | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, THE ANVIL AND TROCAR WERE CONNECTED WITH SOUND BUT THEY FELL OUT DURING APPROXIMATION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-05-2019 | Malfunction | us surgical puerto rico | 23-01-2020 | ENDO GIA ULTRA | Device Difficult to Setup or Prepare | A2201 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE FIRST FIRING WAS PERFORMED NORMALLY, BUT IN THE SECOND SHOT THE HANDLE WAS LOOSE AND UNFIRMLY TO FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT HARM.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Unspecified Infection | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Blood Loss | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
09-01-2020 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH RESECTION IN GASTRIC BYPASS LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE.THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 23-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING THE RESECTION OF THE PULMONARY ARTERY IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE RELOAD WAS CONNECTED, THEN THE YELLOW SHIPPING WEDGE WAS REMOVED AND THE GREEN FIRING BUTTON WAS PUSHED, BUT THE RELOAD COULD NOT BE FIRED. A NEW HANDLE WAS USED WITH THE SAME RESULT. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. A THIRD HANDLE AND RELOAD WAS USED WHICH FIRED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 23-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING THE RESECTION OF THE PULMONARY ARTERY IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE RELOAD WAS CONNECTED, THEN THE YELLOW SHIPPING WEDGE WAS REMOVED AND THE GREEN FIRING BUTTON WAS PUSHED, BUT THE RELOAD COULD NOT BE FIRED. A NEW HANDLE WAS USED WITH THE SAME RESULT. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. A THIRD HANDLE AND RELOAD WAS USED WHICH FIRED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | covidien lp llc north haven | 23-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING THE RESECTION OF THE PULMONARY ARTERY IN A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE RELOAD WAS CONNECTED, THEN THE YELLOW SHIPPING WEDGE WAS REMOVED AND THE GREEN FIRING BUTTON WAS PUSHED, BUT THE RELOAD COULD NOT BE FIRED. A NEW HANDLE WAS USED WITH THE SAME RESULT. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. A THIRD HANDLE AND RELOAD WAS USED WHICH FIRED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | Unspecified Infection | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | Blood Loss | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | No Code Available | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
18-04-2017 | Injury | covidien lp llc north haven | 23-01-2020 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE LITERATURE, 268 PATIENTS WHO UNDERWENT LAPAROSCOPY-ASSISTED TOTAL GASTRECTOMY (LATG) WITH ROUX-EN-Y RECONSTRUCTION EXPERIENCED POSTOPERATIVE COMPLICATIONS INCLUDING BLEEDING, LEAKAGE, AND STRICTURE. OTHER COMPLICATIONS SUCH AS PULMONARY INFECTION, PLEURAL EFFUSION, DUODENAL STUMP LEAKAGE, ILEUS, ABDOMINAL BLEEDING WERE ALSO REPORTED. 6 CASES OF STRICTURE WERE RESOLVED WITH BALLOON DILATATION AND THREE PATIENTS DEVELOPED ABDOMINAL BLEEDING UNDERWENT REOPERATION. PATIENTS WITH ILEUS AND AN ABDOMINAL HERNIA ALSO UNDERWENT REOPERATION. THE REMAINING PATIENTS WERE CONTROLLED WITH CONSERVATIVE TREATMENT.
05-11-2019 | Malfunction | us surgical puerto rico | 23-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN DIGESTIVE SURGERY, THE STAPLES WERE STUCK AND UNABLE TO STAPLE. ANOTHER STAPLER WAS USED TO FINISH THE CASE. THERE WAS NO PATIENT INJURY.
23-01-2020 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Information | Event Description: THE DEVICE WAS RECEIVED WITH NO ACCOMPANYING INFORMATION.
25-10-2019 | Malfunction | ethicon endo-surgery | 23-01-2020 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN OPEN SIGMOIDECTOMY, THE DEVICE COULD NOT BE FIRED. THE DEVICE WAS USED ON THE COLON. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
09-01-2020 | Malfunction | covidien lp llc north haven | 23-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLES DID NOT FORM PROPERLY. THE INSTRUMENT WAS CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2020 | Malfunction | ethicon endo-surgery | 23-01-2020 | PWRD ECH FLEX 45MM 34CM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS WEDGE, THE SURGEON USED A POWER ECHELON 45 AND EVERYTHING WENT SMOOTHLY FOR THE FIRST FIRING (WITH ECR45G). DURING THE SECOND FIRING, STAPLES WERE FORMED SUCCESSFULLY BUT THE BLADE COULD NOT BE RETRIEVED, HENCE THE JAW COULD NOT BE OPENED. SURGEON PRESSED THE REVERSE BUTTON BUT IT DIDN'T WORK, AND HAD TO USE MANUAL OVERRIDE AND THE BLADE WAS RETRIEVED BACK SUCCESSFULLY AND JAW WAS OPENED. HE COMPLETED THE WEDGE BY OPENING A MANUAL ECHELON AND FIRING A THIRD RELOAD. SURGERY WAS NOT DELAYED DUE TO THE REPORTED EVENT AND WAS SUCCESSFULLY COMPLETED. NO FRAGMENTS WERE GENERATED AND THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
01-01-2020 | Malfunction | ethicon endo-surgery | 23-01-2020 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE SURGERY, THE RELOAD CUT WITHOUT STAPLING. THE SURGEON MANAGED TO CLOSE BY MANUAL SUTURING. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
04-12-2019 | Malfunction | ethicon endo-surgery | 23-01-2020 | GST GOLD RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, WHEN THE SHOT IS MADE, THE STAPLES ARE BADLY FORMED, AND IT IS NECESSARY TO REINFORCE THEM WITH THREAD SUTURES. THE PROCEDURE WAS DELAYED BY 5-MINUTES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
31-12-2019 | Injury | intuitive surgical | 23-01-2020 | SUREFORM 60 | Insufficient Information | E2401 | Blood Loss | Event Description: IT WAS INITIALLY REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT EXPERIENCED ¿EXCESSIVE BLEEDING.¿ THE INTUITIVE SURGICAL, INC. (ISI) TECHNICAL SUPPORT ENGINEER (TSE) WAS CONTACTED AFTER THE PROCEDURE AND THE SYSTEM LOGS WERE REVIEWED. NO RELATED SYSTEM ERRORS WERE FOUND TO HAVE OCCURRED DURING THE SURGICAL PROCEDURE. ON 22-JANUARY-2020, ISI CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE SURGEON STAPLED THE STOMACH TISSUE WITH A SUREFORM 60 STAPLER INSTRUMENT AND AN UNSPECIFIED COLOR SUREFORM 60 RELOAD. WHEN THE STAPLE LINE WAS TRANSECTED AFTER A SUCCESSFUL FIRE, THE PATIENT ALLEGEDLY EXPERIENCED BLEEDING FROM THE STAPLE LINE. THE BLEEDING WAS REPORTED TO BE MORE THAN NORMAL. AS A RESULT, THE SURGEON HAD TO PLACE A CLIP ON THE STAPLE LINE TO CONTROL THE BLEEDING. THE PATIENT DID NOT REQUIRE A BLOOD TRANSFUSION. THE SURGEON DID NOT RECALL WHAT STAPLER FIRE AND SUREFORM 60 RELOAD COLOR WAS USED ON THE STAPLE LINE THAT EXPERIENCED THE BLEEDING. THE SURGEON HAD FIRED A TOTAL OF FIVE SUREFORM 60 RELOADS WITH THE SUREFORM 60 STAPLER INSTRUMENT. THE FIVE SUREFORM 60 RELOADS USED DURING THE PROCEDURE WERE: TWO BLACK, TWO BLUE, AND ONE GREEN. THE SURGEON DID COMPLETE THE PROCEDURE WITH THE SAME SUREFORM 60 STAPLER INSTRUMENT WITH NO FURTHER ISSUES. THE FOLLOWING WERE CONFIRMED; THERE WAS NO BUTTRESS MATERIAL USED, NO TISSUE TENSION OR BUNCHING, AND NO EXPOSURE TO RADIATION OR CHEMOTHERAPY. THERE IS NO VIDEO RECORDING OF THE PROCEDURE FOR REVIEW. THE SUREFORM 60 STAPLER INSTRUMENT AND THE RELOADS HAVE BEEN DISCARDED AND WILL NOT BE RETURNED TO ISI FOR EVALUATION.
02-01-2020 | Injury | ethicon endo-surgery | 23-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: IT WAS REPORTED THAT DURING REVIEW OF A JOURNAL ARTICLE, TITLE: PROXIMAL LEAKAGE AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY: AN ANALYSIS OF PREOPERATIVE AND OPERATIVE PREDICTORS ON 1738 CONSECUTIVE PROCEDURES. AUTHOR/S: GIOVANNI CESANA, STEFANO CIOFFI, RICCARDO GIORGI, ROBERTA VILLA1, MATTEO UCCELLI, FRANCESCA CICCARESE, GIORGIO CASTELLO, BRUNO SCOTTO AND STEFANO OLMI. CITATION: OBES SURG (2018) 28:627¿635; DOI 10.1007/S11695-017-2907-Z; PUBLISHED ONLINE: 24 AUGUST 2017. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO SEARCH FOR PREDICTIVE FACTORS FOR PROXIMAL LEAKAGE AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) IN A LARGE COHORT FROM A SINGLE REFERRAL CENTER. FROM JANUARY 2008 TO OCTOBER 2016, ONE THOUSAND SEVEN HUNDRED AND THIRTY-EIGHT PATIENTS (MEDIAN AGE 41.0 (33.0 TO 49.0) YEARS, MALE N=428, FEMALE N=1310) WHO UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY WERE ANALYZED. LSG WAS PERFORMED USING FOUR TROCARS. THE GREAT CURVATURE WAS DISSECTED AND ISOLATED FROM THE GASTROCOLIC LIGAMENT AND THE SPLENOCOLIC LIGAMENT THROUGH AN ULTRASONIC DEVICE (HARMONIC ACE, ETHICON ENDO-SURGERY INC., (B)(4)) OR A RADIOFREQUENCY SEALER DEVICE (LIGASURE MARYLAND JAW, MEDTRONIC, (B)(4); OR CONMED ALTRUS, CONMED ELECTROSURGERY, (B)(4), USA) OR AN INTEGRATED DEVICE OF BOTH ADVANCED BIPOLAR ENERGY AND ULTRASONIC ENERGY (THUNDERBEAT, OLYMPUS MEDICAL SYSTEM CORP., (B)(4)). THE LINEAR STAPLER USED WERE COVIDIEN ENDO GIA 60 RELOADS WITH OR WITHOUT TRI-STAPLE TECHNOLOGY (MEDTRONIC, (B)(4)) AND ETHICON ECHELON FLEX 60 ENDOPATH STAPLER (ETHICON ENDO-SURGERY INC., (B)(4)). THE CHOICE OF A PARTICULAR CARTRIDGE DEPENDED ON THE TISSUE THICKNESS ACCORDING TO THE INTRA-OPERATIVE PERCEPTION AND EXPERIENCE OF THE SURGEON. IN MOST OF THE LSGS, THE HIGHEST PART OF THE SUTURE WAS REINFORCED WITH AN OVERRIDING SUTURE: PDS 3/0 (MIC55E, PDS II, ETHICON ENDO-CLIP SUTURE, (B)(4), USA) OR POLYSORB 2/0 (ENDOSTITCH SUTURING DEVICE, COVIDIEN, (B)(4)), OR WITH BUTTRESS MATERIAL: BOVINE PERICARDIUM (PERI-STRIP DRY, SYNOVIS LIFE TECHNOLOGIES, INC., (B)(4), USA) OR POLYGLYCOLIC ACID (ENDO GI REINFORCED RELOADED WITH TRI-STAPLE TECHNOLOGY, MEDTRONIC, (B)(4)). COMPLICATIONS INCLUDED LEAKAGE AT THE BOTTOM OF GASTRIC SUTURE (N=3) WHICH WAS EASILY MANAGED BY A LAPAROSCOPIC INTERVENTION WITH SUTURE OF THE GASTRIC DEHISCENCE; LEAKAGE AT THE PROXIMAL PART OF THE GASTRIC SUTURE (N=45), ABDOMINAL PAIN (N=45), FEVER (N=45) AND ABDOMINAL ABSCESS (N=42). IN CONCLUSION, IN THIS LARGE CONSECUTIVE COHORT STUDY OF LSG, PROXIMAL STAPLE LINE REINFORCEMENT (BUTTRESS MATERIAL OR SUTURE) REDUCED THE RISK OF A LEAK.
02-01-2020 | Injury | ethicon endo-surgery | 23-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: IT WAS REPORTED THAT DURING REVIEW OF A JOURNAL ARTICLE, TITLE: PROXIMAL LEAKAGE AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY: AN ANALYSIS OF PREOPERATIVE AND OPERATIVE PREDICTORS ON 1738 CONSECUTIVE PROCEDURES. AUTHOR/S: GIOVANNI CESANA, STEFANO CIOFFI, RICCARDO GIORGI, ROBERTA VILLA1, MATTEO UCCELLI, FRANCESCA CICCARESE, GIORGIO CASTELLO, BRUNO SCOTTO AND STEFANO OLMI. CITATION: OBES SURG (2018) 28:627¿635; DOI 10.1007/S11695-017-2907-Z; PUBLISHED ONLINE: 24 AUGUST 2017. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO SEARCH FOR PREDICTIVE FACTORS FOR PROXIMAL LEAKAGE AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) IN A LARGE COHORT FROM A SINGLE REFERRAL CENTER. FROM JANUARY 2008 TO OCTOBER 2016, ONE THOUSAND SEVEN HUNDRED AND THIRTY-EIGHT PATIENTS (MEDIAN AGE 41.0 (33.0 TO 49.0) YEARS, MALE N=428, FEMALE N=1310) WHO UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY WERE ANALYZED. LSG WAS PERFORMED USING FOUR TROCARS. THE GREAT CURVATURE WAS DISSECTED AND ISOLATED FROM THE GASTROCOLIC LIGAMENT AND THE SPLENOCOLIC LIGAMENT THROUGH AN ULTRASONIC DEVICE (HARMONIC ACE, ETHICON ENDO-SURGERY INC., (B)(4)) OR A RADIOFREQUENCY SEALER DEVICE (LIGASURE MARYLAND JAW, MEDTRONIC, (B)(4); OR CONMED ALTRUS, CONMED ELECTROSURGERY, (B)(4), USA) OR AN INTEGRATED DEVICE OF BOTH ADVANCED BIPOLAR ENERGY AND ULTRASONIC ENERGY (THUNDERBEAT, OLYMPUS MEDICAL SYSTEM CORP., (B)(4)). THE LINEAR STAPLER USED WERE COVIDIEN ENDO GIA 60 RELOADS WITH OR WITHOUT TRI-STAPLE TECHNOLOGY (MEDTRONIC, (B)(4)) AND ETHICON ECHELON FLEX 60 ENDOPATH STAPLER (ETHICON ENDO-SURGERY INC., (B)(4)). THE CHOICE OF A PARTICULAR CARTRIDGE DEPENDED ON THE TISSUE THICKNESS ACCORDING TO THE INTRA-OPERATIVE PERCEPTION AND EXPERIENCE OF THE SURGEON. IN MOST OF THE LSGS, THE HIGHEST PART OF THE SUTURE WAS REINFORCED WITH AN OVERRIDING SUTURE: PDS 3/0 (MIC55E, PDS II, ETHICON ENDO-CLIP SUTURE, (B)(4), USA) OR POLYSORB 2/0 (ENDOSTITCH SUTURING DEVICE, COVIDIEN, (B)(4)), OR WITH BUTTRESS MATERIAL: BOVINE PERICARDIUM (PERI-STRIP DRY, SYNOVIS LIFE TECHNOLOGIES, INC., (B)(4), USA) OR POLYGLYCOLIC ACID (ENDO GI REINFORCED RELOADED WITH TRI-STAPLE TECHNOLOGY, MEDTRONIC, (B)(4)). COMPLICATIONS INCLUDED LEAKAGE AT THE BOTTOM OF GASTRIC SUTURE (N=3) WHICH WAS EASILY MANAGED BY A LAPAROSCOPIC INTERVENTION WITH SUTURE OF THE GASTRIC DEHISCENCE; LEAKAGE AT THE PROXIMAL PART OF THE GASTRIC SUTURE (N=45), ABDOMINAL PAIN (N=45), FEVER (N=45) AND ABDOMINAL ABSCESS (N=42). IN CONCLUSION, IN THIS LARGE CONSECUTIVE COHORT STUDY OF LSG, PROXIMAL STAPLE LINE REINFORCEMENT (BUTTRESS MATERIAL OR SUTURE) REDUCED THE RISK OF A LEAK.
02-01-2020 | Injury | ethicon endo-surgery | 23-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: IT WAS REPORTED THAT DURING REVIEW OF A JOURNAL ARTICLE, TITLE: PROXIMAL LEAKAGE AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY: AN ANALYSIS OF PREOPERATIVE AND OPERATIVE PREDICTORS ON 1738 CONSECUTIVE PROCEDURES. AUTHOR/S: GIOVANNI CESANA, STEFANO CIOFFI, RICCARDO GIORGI, ROBERTA VILLA1, MATTEO UCCELLI, FRANCESCA CICCARESE, GIORGIO CASTELLO, BRUNO SCOTTO AND STEFANO OLMI. CITATION: OBES SURG (2018) 28:627¿635; DOI 10.1007/S11695-017-2907-Z; PUBLISHED ONLINE: 24 AUGUST 2017. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO SEARCH FOR PREDICTIVE FACTORS FOR PROXIMAL LEAKAGE AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) IN A LARGE COHORT FROM A SINGLE REFERRAL CENTER. FROM JANUARY 2008 TO OCTOBER 2016, ONE THOUSAND SEVEN HUNDRED AND THIRTY-EIGHT PATIENTS (MEDIAN AGE 41.0 (33.0 TO 49.0) YEARS, MALE N=428, FEMALE N=1310) WHO UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY WERE ANALYZED. LSG WAS PERFORMED USING FOUR TROCARS. THE GREAT CURVATURE WAS DISSECTED AND ISOLATED FROM THE GASTROCOLIC LIGAMENT AND THE SPLENOCOLIC LIGAMENT THROUGH AN ULTRASONIC DEVICE (HARMONIC ACE, ETHICON ENDO-SURGERY INC., (B)(4)) OR A RADIOFREQUENCY SEALER DEVICE (LIGASURE MARYLAND JAW, MEDTRONIC, (B)(4); OR CONMED ALTRUS, CONMED ELECTROSURGERY, (B)(4), USA) OR AN INTEGRATED DEVICE OF BOTH ADVANCED BIPOLAR ENERGY AND ULTRASONIC ENERGY (THUNDERBEAT, OLYMPUS MEDICAL SYSTEM CORP., (B)(4)). THE LINEAR STAPLER USED WERE COVIDIEN ENDO GIA 60 RELOADS WITH OR WITHOUT TRI-STAPLE TECHNOLOGY (MEDTRONIC, (B)(4)) AND ETHICON ECHELON FLEX 60 ENDOPATH STAPLER (ETHICON ENDO-SURGERY INC., (B)(4)). THE CHOICE OF A PARTICULAR CARTRIDGE DEPENDED ON THE TISSUE THICKNESS ACCORDING TO THE INTRA-OPERATIVE PERCEPTION AND EXPERIENCE OF THE SURGEON. IN MOST OF THE LSGS, THE HIGHEST PART OF THE SUTURE WAS REINFORCED WITH AN OVERRIDING SUTURE: PDS 3/0 (MIC55E, PDS II, ETHICON ENDO-CLIP SUTURE, (B)(4), USA) OR POLYSORB 2/0 (ENDOSTITCH SUTURING DEVICE, COVIDIEN, (B)(4)), OR WITH BUTTRESS MATERIAL: BOVINE PERICARDIUM (PERI-STRIP DRY, SYNOVIS LIFE TECHNOLOGIES, INC., (B)(4), USA) OR POLYGLYCOLIC ACID (ENDO GI REINFORCED RELOADED WITH TRI-STAPLE TECHNOLOGY, MEDTRONIC, (B)(4)). COMPLICATIONS INCLUDED LEAKAGE AT THE BOTTOM OF GASTRIC SUTURE (N=3) WHICH WAS EASILY MANAGED BY A LAPAROSCOPIC INTERVENTION WITH SUTURE OF THE GASTRIC DEHISCENCE; LEAKAGE AT THE PROXIMAL PART OF THE GASTRIC SUTURE (N=45), ABDOMINAL PAIN (N=45), FEVER (N=45) AND ABDOMINAL ABSCESS (N=42). IN CONCLUSION, IN THIS LARGE CONSECUTIVE COHORT STUDY OF LSG, PROXIMAL STAPLE LINE REINFORCEMENT (BUTTRESS MATERIAL OR SUTURE) REDUCED THE RISK OF A LEAK.
02-01-2020 | Injury | ethicon endo-surgery | 23-01-2020 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING REVIEW OF A JOURNAL ARTICLE, TITLE: PROXIMAL LEAKAGE AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY: AN ANALYSIS OF PREOPERATIVE AND OPERATIVE PREDICTORS ON 1738 CONSECUTIVE PROCEDURES. AUTHOR/S: GIOVANNI CESANA, STEFANO CIOFFI, RICCARDO GIORGI, ROBERTA VILLA1, MATTEO UCCELLI, FRANCESCA CICCARESE, GIORGIO CASTELLO, BRUNO SCOTTO AND STEFANO OLMI. CITATION: OBES SURG (2018) 28:627¿635; DOI 10.1007/S11695-017-2907-Z; PUBLISHED ONLINE: 24 AUGUST 2017. THE PURPOSE OF THIS RETROSPECTIVE STUDY WAS TO SEARCH FOR PREDICTIVE FACTORS FOR PROXIMAL LEAKAGE AFTER LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) IN A LARGE COHORT FROM A SINGLE REFERRAL CENTER. FROM JANUARY 2008 TO OCTOBER 2016, ONE THOUSAND SEVEN HUNDRED AND THIRTY-EIGHT PATIENTS (MEDIAN AGE 41.0 (33.0 TO 49.0) YEARS, MALE N=428, FEMALE N=1310) WHO UNDERWENT LAPAROSCOPIC SLEEVE GASTRECTOMY WERE ANALYZED. LSG WAS PERFORMED USING FOUR TROCARS. THE GREAT CURVATURE WAS DISSECTED AND ISOLATED FROM THE GASTROCOLIC LIGAMENT AND THE SPLENOCOLIC LIGAMENT THROUGH AN ULTRASONIC DEVICE (HARMONIC ACE, ETHICON ENDO-SURGERY INC., (B)(4)) OR A RADIOFREQUENCY SEALER DEVICE (LIGASURE MARYLAND JAW, MEDTRONIC, (B)(4); OR CONMED ALTRUS, CONMED ELECTROSURGERY, (B)(4), USA) OR AN INTEGRATED DEVICE OF BOTH ADVANCED BIPOLAR ENERGY AND ULTRASONIC ENERGY (THUNDERBEAT, OLYMPUS MEDICAL SYSTEM CORP., (B)(4)). THE LINEAR STAPLER USED WERE COVIDIEN ENDO GIA 60 RELOADS WITH OR WITHOUT TRI-STAPLE TECHNOLOGY (MEDTRONIC, (B)(4)) AND ETHICON ECHELON FLEX 60 ENDOPATH STAPLER (ETHICON ENDO-SURGERY INC., (B)(4)). THE CHOICE OF A PARTICULAR CARTRIDGE DEPENDED ON THE TISSUE THICKNESS ACCORDING TO THE INTRA-OPERATIVE PERCEPTION AND EXPERIENCE OF THE SURGEON. IN MOST OF THE LSGS, THE HIGHEST PART OF THE SUTURE WAS REINFORCED WITH AN OVERRIDING SUTURE: PDS 3/0 (MIC55E, PDS II, ETHICON ENDO-CLIP SUTURE, (B)(4), USA) OR POLYSORB 2/0 (ENDOSTITCH SUTURING DEVICE, COVIDIEN, (B)(4)), OR WITH BUTTRESS MATERIAL: BOVINE PERICARDIUM (PERI-STRIP DRY, SYNOVIS LIFE TECHNOLOGIES, INC., (B)(4), USA) OR POLYGLYCOLIC ACID (ENDO GI REINFORCED RELOADED WITH TRI-STAPLE TECHNOLOGY, MEDTRONIC, (B)(4)). COMPLICATIONS INCLUDED LEAKAGE AT THE BOTTOM OF GASTRIC SUTURE (N=3) WHICH WAS EASILY MANAGED BY A LAPAROSCOPIC INTERVENTION WITH SUTURE OF THE GASTRIC DEHISCENCE; LEAKAGE AT THE PROXIMAL PART OF THE GASTRIC SUTURE (N=45), ABDOMINAL PAIN (N=45), FEVER (N=45) AND ABDOMINAL ABSCESS (N=42). IN CONCLUSION, IN THIS LARGE CONSECUTIVE COHORT STUDY OF LSG, PROXIMAL STAPLE LINE REINFORCEMENT (BUTTRESS MATERIAL OR SUTURE) REDUCED THE RISK OF A LEAK.
01-01-2020 | Malfunction | ethicon endo-surgery | 23-01-2020 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ANVIL WOULD NOT OPEN AFTER FIRING. SURGERY WAS DELAYED FOR 10 MINUTES AND A NEW DEVICE WAS OPENED. PATIENT CONSEQUENCE WAS NOT REPORTED.
19-11-2019 | Malfunction | us surgical puerto rico | 24-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE DEVICE FIRED, BUT THERE WAS POOR STAPLE FORMATION AT THE DISTAL PART OF THE STAPLE LINE. ANOTHER HANDLE AND RELOAD WAS USED TO RESTAPLE THE LINE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-01-2020 | Malfunction | us surgical puerto rico | 24-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE PULMONARY PARENCHYMA WAS BEING DETACHED DURING A LAPAROSCOPIC RADICAL RESECTION OF PULMONARY CARCINOMA, THE STAPLER WAS ABLE TO FIRE 4.8MM RELOAD. THE SURGEON THEN LOADED A 4.8MM RELOAD BUT THE STAPLER DID NOT FIRE. THE HANDLE WAS REPLACED AND WAS ABLE TO FIRE THE SECOND RELOAD. THERE WAS NO PATIENT INJURY.
11-01-2020 | Injury | us surgical puerto rico | 24-01-2020 | GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPY ASSISTED COLON RESECTION, WHILE RESECTING THE COLON TISSUE IN THE LUMEN, THE STAPLES DID NOT FORM A B SHAPE. THE STAPLE LINE WAS RESECTED AND RESTAPLED WITH ANOTHER DEVICE.
11-01-2020 | Injury | us surgical puerto rico | 24-01-2020 | GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPY ASSISTED COLON RESECTION, WHILE RESECTING THE COLON TISSUE IN THE LUMEN, THE STAPLES DID NOT FORM A B SHAPE. THE STAPLE LINE WAS RESECTED AND RESTAPLED WITH ANOTHER DEVICE.
11-01-2020 | Injury | covidien lp llc north haven | 24-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN DETACHING A PULMONARY PARENCHYMA DURING A LAPAROSCOPIC RADICAL RESECTION OF PULMONARY CARCINOMA, THE STAPLER WAS ABLE TO FIRE HOWEVER THE STAPLE LINE WAS INCOMPLETE DISTALLY. IT WAS ALSO REPORTED THERE WAS A POOR STAPLE FORMATION. ANOTHER RELOAD WAS USED AND THE STAPLE LINE WAS FIXED BY MANUAL SUTURING TO COMPLETE THE CASE.
11-01-2020 | Injury | covidien lp llc north haven | 24-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN DETACHING A PULMONARY PARENCHYMA DURING A LAPAROSCOPIC RADICAL RESECTION OF PULMONARY CARCINOMA, THE STAPLER WAS ABLE TO FIRE HOWEVER THE STAPLE LINE WAS INCOMPLETE DISTALLY. IT WAS ALSO REPORTED THERE WAS A POOR STAPLE FORMATION. ANOTHER RELOAD WAS USED AND THE STAPLE LINE WAS FIXED BY MANUAL SUTURING TO COMPLETE THE CASE.
11-01-2020 | Injury | us surgical puerto rico | 24-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN DETACHING A PULMONARY PARENCHYMA DURING A LAPAROSCOPIC RADICAL RESECTION OF PULMONARY CARCINOMA, THE STAPLER WAS ABLE TO FIRE HOWEVER THE STAPLE LINE WAS INCOMPLETE DISTALLY. IT WAS ALSO REPORTED THERE WAS A POOR STAPLE FORMATION. ANOTHER RELOAD WAS USED AND THE STAPLE LINE WAS FIXED BY MANUAL SUTURING TO COMPLETE THE CASE.
11-01-2020 | Injury | us surgical puerto rico | 24-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING RADICAL RESECTION OF COLORECTAL CANCER, WHEN THE COLON TISSUES WERE DETACHED, THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE DISTALLY THAT WAS FIXED BY HAND SEWING. DEVICE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE.
01-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ILS 29MM, CURVED | Misfire | A050502 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN ANTERIOR RESECTION PROCEDURE, THE SURGEON DESCRIBED THAT HE DID NOT HEAR THE CLICK OF THE WASHER AND THE TACTILE FEEL OF THE HANDLE COMPRESSING FULLY. THE CIRCULAR STAPLER DID NOT FIRE COMPLETELY TO COMPLETE THE ANASTOMOSIS. MISFIRED STAPLES NEEDED TO BE REMOVED AND A COLOSTOMY WAS PERFORMED. PROCEDURE WAS DELAYED 30 MINUTES. PATIENT NOW HAS A PERMANENT COLOSTOMY BAG. THE MISFIRED STAPLES HAD TO BE REMOVED.
01-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN ANTERIOR RESECTION PROCEDURE, THE SURGEON DESCRIBED THAT HE DID NOT HEAR THE CLICK OF THE WASHER AND THE TACTILE FEEL OF THE HANDLE COMPRESSING FULLY. THE CIRCULAR STAPLER DID NOT FIRE COMPLETELY TO COMPLETE THE ANASTOMOSIS. MISFIRED STAPLES NEEDED TO BE REMOVED AND A COLOSTOMY WAS PERFORMED. PROCEDURE WAS DELAYED 30 MINUTES. PATIENT NOW HAS A PERMANENT COLOSTOMY BAG. THE MISFIRED STAPLES HAD TO BE REMOVED.
01-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN ANTERIOR RESECTION PROCEDURE, THE SURGEON DESCRIBED THAT HE DID NOT HEAR THE CLICK OF THE WASHER AND THE TACTILE FEEL OF THE HANDLE COMPRESSING FULLY. THE CIRCULAR STAPLER DID NOT FIRE COMPLETELY TO COMPLETE THE ANASTOMOSIS. MISFIRED STAPLES NEEDED TO BE REMOVED AND A COLOSTOMY WAS PERFORMED. PROCEDURE WAS DELAYED 30 MINUTES. PATIENT NOW HAS A PERMANENT COLOSTOMY BAG. THE MISFIRED STAPLES HAD TO BE REMOVED.
01-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, WHILE USING THE CIRCULAR STAPLER, CHECKED THE DONUTS AND DISCOVERED A HALF OF IT NOT STAPLED. THE DOCTOR COMPLETED THE SURGERY WITH SUTURE, BUT HE CHANGED THE PLAN BECAUSE OF THIS EVENT AND DID A COLOSTOMY RATHER THAN ANASTOMOSIS. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THE PATIENT CONDITION IS STABLE.
02-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT HEALTHCARE PROFESSIONAL HAD JUST SCOPED ONE PATIENT POST LAP ANTERIOR RESECTION WITH BLEEDING ANASTOMOSIS LINE ON POST-OP DAY 3. THERE WERE NO DIFFICULTIES ON FIRING AND BOTH DONUTS COMPLETED.
11-01-2020 | Injury | covidien lp llc north haven | 24-01-2020 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING RADICAL RESECTION OF COLORECTAL CANCER, WHEN THE COLON TISSUES WERE DETACHED, THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE DISTALLY THAT WAS FIXED BY HAND SEWING. DEVICE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE.
11-01-2020 | Injury | covidien lp llc north haven | 24-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING RADICAL RESECTION OF COLORECTAL CANCER, WHEN THE COLON TISSUES WERE DETACHED, THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE DISTALLY THAT WAS FIXED BY HAND SEWING. DEVICE WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE.
08-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | POWERED PLUS COMPACT 45 | Gas/Air Leak | A050402 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS, THERE WERE FIVE AIR LEAKS IN NOVEMBER AND SEVEN IN DECEMBER. PATIENT READMITTED WITH A PNEUMOTHORAX ON POD#10.
20-12-2019 | Injury | ethicon endo-surgery | 24-01-2020 | POWERED PLUS COMPACT 45 | Gas/Air Leak | A050402 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS,THERE WERE FIVE AIR LEAKS IN NOVEMBER AND SEVEN IN DECEMBER. (B)(6) 2019 INCREASING SUBCUTANEOUS EMPHYSEMA OVERLYING THE LEFT NECK AND CHEST WALL. THERE IS A 1.9 CM LEFT APICAL PNEUMOTHORAX. THERE IS ALSO A 7 MM LEFT BASILAR PNEUMO. CHEST TUBE ULTIMATED REMOVED AND DISCHARGED TO HOME ON (B)(6) 2019.
10-01-2020 | Malfunction | covidien lp llc north haven | 24-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE PULMONARY PARENCHYMA WAS BEING DETACHED DURING A LAPAROSCOPIC RADICAL RESECTION OF PULMONARY CARCINOMA, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. THE DEVICE WAS REPLACED. THERE WAS NO PATIENT INJURY.
03-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ETHICON CURVED INTRALUMINAL STAPLER | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: DURING ANASTOMOSIS SURGICAL PROCEDURE, RN GAVE MANUAL ETHICON 29 TO THE FIELD. DR. USED STAPLER ON PATIENT. STAPLER ONLY CUT AND DID NOT STAPLE. STAPLES WERE SEEN IN THE ABDOMEN. MFG REPS, CALLED TO COME INTO THE ROOM. DURING CONVERSATION BETWEEN REPS AND SURGEON, DR. SAID THOUGHT STAPLER WAS ELECTRIC. STAPLER AND BOX SEQUESTERED AND GIVEN TO MANAGER TO BE SENT TO BIOMEDICAL FOR EVALUATION. MANUFACTURER RESPONSE FOR CURVED INTRALUMINAL STAPLER (ILS), ETHICON (PER SITE REPORTER). FOLLOW UP PENDING.
03-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ETHICON CURVED INTRALUMINAL STAPLER | Activation, Positioning or Separation Problem | A15 | No Known Impact Or Consequence To Patient | Event Description: DURING ANASTOMOSIS SURGICAL PROCEDURE, RN GAVE MANUAL ETHICON 29 TO THE FIELD. DR. USED STAPLER ON PATIENT. STAPLER ONLY CUT AND DID NOT STAPLE. STAPLES WERE SEEN IN THE ABDOMEN. MFG REPS, CALLED TO COME INTO THE ROOM. DURING CONVERSATION BETWEEN REPS AND SURGEON, DR. SAID THOUGHT STAPLER WAS ELECTRIC. STAPLER AND BOX SEQUESTERED AND GIVEN TO MANAGER TO BE SENT TO BIOMEDICAL FOR EVALUATION. MANUFACTURER RESPONSE FOR CURVED INTRALUMINAL STAPLER (ILS), ETHICON (PER SITE REPORTER). FOLLOW UP PENDING.
07-12-2019 | Injury | ethicon endo-surgery | 24-01-2020 | POWERED PLUS COMPACT 45 | Leak/Splash | A0504 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS,THERE WERE FIVE AIR LEAKS IN (B)(6) AND (B)(6). (B)(6) 2019 PATIENT READMITTED WITH HYDROPNEUMO FOR CHEST TUBE S/P R MID AND LOWER LOBECTOMY. TREATED WITH CHEST TUBE TO PNEUMOSTAT.
01-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ILS 29MM, CURVED | Misfire | A050502 | Pain | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY PROCEDURE, THE CUSTOMER STATED THE STAPLER MISFIRED AND DIDN'T COMPLETELY FORM STAPLES. AFTER FIRING AND REMOVING THE STAPLER THEY DISCOVERED ONLY ONE ¿DOUGHNUT¿ OF TISSUE. ONE RING OF STAPLES HAD BEEN FIRED AND FREE, UNFORMED STAPLES WERE DEPLOYED AND THEN COLLECTED IN A CUP. A SECOND DOCTOR WAS CALLED IN TO HELP COMPLETE THE PROCEDURE. THEY WERE ABLE TO COMPLETE THE PROCEDURE BY HAND-SEWING. ¿THE EXTENT OF ISSUE/ INJURY TO THE PATIENT WAS EXTENDED ANESTHESIA TIME (DOUBLE THE ORIGINALLY SCHEDULED TIME). THE PATIENT DID REQUIRE AN ILEOSTOMY. HOWEVER, THIS WAS ALWAYS A POSSIBILITY BUT WHEN THE STAPLER MALFUNCTIONED IT INCREASED THE LIKELIHOOD OF THE ILEOSTOMY TO 100%. PATIENT HAS BEEN DISCHARGED BUT CONTINUES TO HAVE PAIN BUT OTHERWISE, THE PATIENT IS DOING FINE.¿
01-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ILS 29MM, CURVED | Misfire | A050502 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY PROCEDURE, THE CUSTOMER STATED THE STAPLER MISFIRED AND DIDN'T COMPLETELY FORM STAPLES. AFTER FIRING AND REMOVING THE STAPLER THEY DISCOVERED ONLY ONE ¿DOUGHNUT¿ OF TISSUE. ONE RING OF STAPLES HAD BEEN FIRED AND FREE, UNFORMED STAPLES WERE DEPLOYED AND THEN COLLECTED IN A CUP. A SECOND DOCTOR WAS CALLED IN TO HELP COMPLETE THE PROCEDURE. THEY WERE ABLE TO COMPLETE THE PROCEDURE BY HAND-SEWING. ¿THE EXTENT OF ISSUE/ INJURY TO THE PATIENT WAS EXTENDED ANESTHESIA TIME (DOUBLE THE ORIGINALLY SCHEDULED TIME). THE PATIENT DID REQUIRE AN ILEOSTOMY. HOWEVER, THIS WAS ALWAYS A POSSIBILITY BUT WHEN THE STAPLER MALFUNCTIONED IT INCREASED THE LIKELIHOOD OF THE ILEOSTOMY TO 100%. PATIENT HAS BEEN DISCHARGED BUT CONTINUES TO HAVE PAIN BUT OTHERWISE, THE PATIENT IS DOING FINE.¿
01-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | Pain | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY PROCEDURE, THE CUSTOMER STATED THE STAPLER MISFIRED AND DIDN'T COMPLETELY FORM STAPLES. AFTER FIRING AND REMOVING THE STAPLER THEY DISCOVERED ONLY ONE ¿DOUGHNUT¿ OF TISSUE. ONE RING OF STAPLES HAD BEEN FIRED AND FREE, UNFORMED STAPLES WERE DEPLOYED AND THEN COLLECTED IN A CUP. A SECOND DOCTOR WAS CALLED IN TO HELP COMPLETE THE PROCEDURE. THEY WERE ABLE TO COMPLETE THE PROCEDURE BY HAND-SEWING. ¿THE EXTENT OF ISSUE/ INJURY TO THE PATIENT WAS EXTENDED ANESTHESIA TIME (DOUBLE THE ORIGINALLY SCHEDULED TIME). THE PATIENT DID REQUIRE AN ILEOSTOMY. HOWEVER, THIS WAS ALWAYS A POSSIBILITY BUT WHEN THE STAPLER MALFUNCTIONED IT INCREASED THE LIKELIHOOD OF THE ILEOSTOMY TO 100%. PATIENT HAS BEEN DISCHARGED BUT CONTINUES TO HAVE PAIN BUT OTHERWISE, THE PATIENT IS DOING FINE.¿
01-01-2020 | Injury | ethicon endo-surgery | 24-01-2020 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY PROCEDURE, THE CUSTOMER STATED THE STAPLER MISFIRED AND DIDN'T COMPLETELY FORM STAPLES. AFTER FIRING AND REMOVING THE STAPLER THEY DISCOVERED ONLY ONE ¿DOUGHNUT¿ OF TISSUE. ONE RING OF STAPLES HAD BEEN FIRED AND FREE, UNFORMED STAPLES WERE DEPLOYED AND THEN COLLECTED IN A CUP. A SECOND DOCTOR WAS CALLED IN TO HELP COMPLETE THE PROCEDURE. THEY WERE ABLE TO COMPLETE THE PROCEDURE BY HAND-SEWING. ¿THE EXTENT OF ISSUE/ INJURY TO THE PATIENT WAS EXTENDED ANESTHESIA TIME (DOUBLE THE ORIGINALLY SCHEDULED TIME). THE PATIENT DID REQUIRE AN ILEOSTOMY. HOWEVER, THIS WAS ALWAYS A POSSIBILITY BUT WHEN THE STAPLER MALFUNCTIONED IT INCREASED THE LIKELIHOOD OF THE ILEOSTOMY TO 100%. PATIENT HAS BEEN DISCHARGED BUT CONTINUES TO HAVE PAIN BUT OTHERWISE, THE PATIENT IS DOING FINE.¿
12-06-2019 | Malfunction | us surgical puerto rico | 24-01-2020 | EEA | Component Missing | A020602 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A GENERAL LAPAROSCOPIC SURGERY, IT WAS IDENTIFIED THAT THE WARHEAD COMPONENT WAS NOT INCLUDED IN THE PACKAGE. ANOTHER DEVICE FROM THE SAME LOT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-11-2019 | Injury | ethicon endo-surgery | 24-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS,THERE WERE FIVE AIR LEAKS IN (B)(6) AND SEVEN IN (B)(6). ON (B)(6) 2019, PATIENT EXPERIENCED 2 MM LEFT LATERAL PNEUMOTHORAX. (B)(6) CHEST TUBE REMOVED AND PATIENT DISCHARGED TO HOME.
30-12-2019 | Malfunction | covidien lp llc north haven | 24-01-2020 | SIGNIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE RELOAD ONLY FIRED FOR THE FIRST PART AND AUTOMATICALLY RECALLED THE BLADE SINCE THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. A DIFFERENT RELOAD WAS USED AND THE SAME ISSUE OCCURRED. IT WAS ALSO STATED THAT FOR THE BOTH RELOADS THERE WAS A POOR STAPLE FORMATION. A DIFFERENT BRAND WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Malfunction | covidien lp llc north haven | 24-01-2020 | SIGNIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE RELOAD ONLY FIRED FOR THE FIRST PART AND AUTOMATICALLY RECALLED THE BLADE SINCE THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. A DIFFERENT RELOAD WAS USED AND THE SAME ISSUE OCCURRED. IT WAS ALSO STATED THAT FOR THE BOTH RELOADS THERE WAS A POOR STAPLE FORMATION. A DIFFERENT BRAND WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2019 | Injury | ethicon endo-surgery | 24-01-2020 | POWERED PLUS COMPACT 45 | Gas/Air Leak | A050402 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS, THERE WERE FIVE AIR LEAKS IN (B)(6) AND SEVEN IN (B)(6). ON (B)(6) 2019 SUBSTANTIAL AMOUNT OF SUB Q AIR AND AIR LEAK NOTED. 28FR CHEST TUBE INSERTED. DISCHARGED HOME POD#8.
26-11-2019 | Malfunction | ethicon endo-surgery | 24-01-2020 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A VATS PROCEDURE, DURING THE THIRD FIRING IT FIRED PARTIALLY AND STOPPED. THE KNIFE WOULD NOT REVERSE SO THE MANUAL OVERRIDE WAS USED TO REMOVE FROM TISSUE. A SIMILAR DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES. ONE DEVICE WILL BE RETURNED.
23-11-2019 | Injury | ethicon endo-surgery | 24-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS,THERE WERE FIVE AIR LEAKS IN (B)(6) AND SEVEN IN (B)(6). ON (B)(6) A 1.5 CM RIGHT APICAL PNEUMOTHORAX AND MILD RIGHT LATERAL CHEST WALL SUBCUTANEOUS EMPHYSEMA. SENT HOME WITH A CHEST TUBE AND PNEUMOSTATIC IN PLACE, AND THE TUBE WAS SUBSEQUENTLY REMOVED ON (B)(6) 2019.
25-11-2019 | Injury | ethicon endo-surgery | 24-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Gas/Air Leak | A050402 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS, THERE WERE FIVE AIR LEAKS IN NOVEMBER AND SEVEN IN DECEMBER. ON 11/25 PATIENT READMITTED FOR PNEUMOTHORAX TREATED W/ CHEST TUBE AND DISCHARGED ON 11/27. INTERMITTENT READMISSION FOR CHF LIKE SYMPTOMS.
04-12-2019 | Malfunction | ethicon endo-surgery | 24-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PARTIAL COLECTOMY, THE GRAY HANDLE BROKE OFF WHILE FIRING THE STAPLER, UNKNOWN FIRING. IT WAS NOT REPORTED IF THE BROKEN HANDLE WAS SAVED. A SECOND LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
25-11-2019 | Malfunction | ethicon endo-surgery | 24-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WOULD NOT ENGAGE. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. CASE WAS DELAYED FIVE MINUTES. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
21-11-2019 | Malfunction | ethicon endo-surgery | 24-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE LAST FIRING OF THE STAPLER, WITH BLACK RELOAD, PERISTRIPS BUTTRESSING MATERIAL, THE DOCTOR TRIED TO CLAMP THE STAPLER ON THE STOMACH AND THE JAWS OF THE STAPLER WOULDN'T CLOSE. A SECOND LIKE STAPLER AND RELOAD WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
11-07-2019 | Malfunction | ethicon endo-surgery | 24-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Material Protrusion/Extrusion | A0411 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBAR SURGERY, NOTED SOME STAPLES PROTRUDING BEFORE USED ON THE PATIENT. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Loss of or Failure to Bond | A040102 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA ULTRA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA ULTRA | Loss of or Failure to Bond | A040102 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA ULTRA | Loss of or Failure to Bond | A040102 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC LEFT PNEUMONECTOMY PROCEDURE, AFTER COMPLETING THE OPERATION, THE SURGEON CLOSED THE WOUND. IT WAS REPORTED THAT BEFORE TRANSFERRING THE PATIENT TO THE RECOVERY ROOM, THE ANESTHETIC DOCTOR SAW THE BLOOD PRESSURE OF THE PATIENT TO DECREASE. THE PATIENT UNDERWENT ANOTHER OPERATION WHEREIN THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO OPEN DUE TO THE DEVICE PROBLEM. THE SURGEON EXTENDED THE INCISION AND WATCHED THE THORACIC CAVITY, IT WAS SEEN THAT THERE WERE LOTS OF BLOOD CLOTS (MORE THAN 1 KG). IT WAS ALSO NOTED THAT THERE WAS A HOLE AT THE POSITION OF STAPLE LINE, IN THE PROCEDURE, THERE WAS NO PROBLEM WITH THE FIRING AND THERE WAS NO STAPLE LINE BLEEDING. THE SURGEON USED HIS FINGER TO BLOCK THE HOLE FIRST THEN LIGATED THROUGH SUTURES. THE PATIENT NEEDED (2 LITERS) FOR BLOOD TRANSFUSION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN ONE HOUR. IT WAS ALSO REPORTED THAT THE PATIENT HAD EXTENDED HOSPITALIZATION FOR 5 DAYS.
11-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING THE DUODENUM DURING LAPAROSCOPIC BARIATRIC SURGERY, THE DEVICE WAS ABLE TO FIRE, HOWEVER, POOR STAPLE FORMATION WAS OBSERVED. IT WAS NOTED THAT THE STAPLE DID NOT FORM PROPERLY, AND SOME DID NOT FORM. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WAS REMOVED BY EXCESSIVE FORCE. ANOTHER RELOAD WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY FOR RADICAL COLON CANCER, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY LAPAROSCOPIC PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE HENCE THE RELOAD DID NOT FIRE. THERE WAS NO PROBLEM WITH THE STAPLER. SAME HANDLE WAS USED AND THE RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, ON THE TREATMENT OF PULMONARY VEINS AND PULMONARY PARENCHYMA, IT WAS IMPOSSIBLE TO UNLOAD THE TWO RELOADS IN THE NORMAL WAY, SO THE NURSE INCREASED THE STRENGTH TO REMOVE THE STAPLES, CAUSING THE RELOAD TO BE BROKEN AND TWO PIECES DROPPED NOT IN THE PATIENT¿S CAVITY. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED THE HANDLE AND RELOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA | Human-Device Interface Problem | A22 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, ON THE TREATMENT OF PULMONARY VEINS AND PULMONARY PARENCHYMA, IT WAS IMPOSSIBLE TO UNLOAD THE TWO RELOADS IN THE NORMAL WAY, SO THE NURSE INCREASED THE STRENGTH TO REMOVE THE STAPLES, CAUSING THE RELOAD TO BE BROKEN AND TWO PIECES DROPPED NOT IN THE PATIENT¿S CAVITY. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED THE HANDLE AND RELOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, ON THE TREATMENT OF PULMONARY VEINS AND PULMONARY PARENCHYMA, IT WAS IMPOSSIBLE TO UNLOAD THE TWO RELOADS IN THE NORMAL WAY, SO THE NURSE INCREASED THE STRENGTH TO REMOVE THE STAPLES, CAUSING THE RELOAD TO BE BROKEN AND TWO PIECES DROPPED NOT IN THE PATIENT¿S CAVITY. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED THE HANDLE AND RELOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, ON THE TREATMENT OF PULMONARY VEINS AND PULMONARY PARENCHYMA, IT WAS IMPOSSIBLE TO UNLOAD THE TWO RELOADS IN THE NORMAL WAY, SO THE NURSE INCREASED THE STRENGTH TO REMOVE THE STAPLES, CAUSING THE RELOAD TO BE BROKEN AND TWO PIECES DROPPED NOT IN THE PATIENT¿S CAVITY. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED THE HANDLE AND RELOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, ON THE TREATMENT OF PULMONARY VEINS AND PULMONARY PARENCHYMA, IT WAS IMPOSSIBLE TO UNLOAD THE TWO RELOADS IN THE NORMAL WAY, SO THE NURSE INCREASED THE STRENGTH TO REMOVE THE STAPLES, CAUSING THE RELOAD TO BE BROKEN AND TWO PIECES DROPPED NOT IN THE PATIENT¿S CAVITY. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED THE HANDLE AND RELOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A PNEUMONECTOMY LAPAROSCOPIC PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS FIXED USING CLIP AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A PNEUMONECTOMY LAPAROSCOPIC PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS FIXED USING CLIP AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A PNEUMONECTOMY LAPAROSCOPIC PROCEDURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS FIXED USING CLIP AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, ON THE TREATMENT OF PULMONARY VEINS AND PULMONARY PARENCHYMA, IT WAS IMPOSSIBLE TO UNLOAD THE TWO RELOADS IN THE NORMAL WAY, SO THE NURSE INCREASED THE STRENGTH TO REMOVE THE STAPLES, CAUSING THE RELOAD TO BE BROKEN AND TWO PIECES DROPPED NOT IN THE PATIENT¿S CAVITY. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED THE HANDLE AND RELOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, ON THE TREATMENT OF PULMONARY VEINS AND PULMONARY PARENCHYMA, IT WAS IMPOSSIBLE TO UNLOAD THE TWO RELOADS IN THE NORMAL WAY, SO THE NURSE INCREASED THE STRENGTH TO REMOVE THE STAPLES, CAUSING THE RELOAD TO BE BROKEN AND TWO PIECES DROPPED NOT IN THE PATIENT¿S CAVITY. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED THE HANDLE AND RELOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, ON THE TREATMENT OF PULMONARY VEINS AND PULMONARY PARENCHYMA, IT WAS IMPOSSIBLE TO UNLOAD THE TWO RELOADS IN THE NORMAL WAY, SO THE NURSE INCREASED THE STRENGTH TO REMOVE THE STAPLES, CAUSING THE RELOAD TO BE BROKEN AND TWO PIECES DROPPED NOT IN THE PATIENT¿S CAVITY. THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THEY REPLACED THE HANDLE AND RELOADS TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | GREEN RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF UPPER LEFT LUNG CANCER SURGERY, WHEN SEVERING LOBAR TISSUE, AFTER THE DEVICE WAS FIRED, IT WAS NOTED THAT SOME STAPLES MALFORMED. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO TISSUE DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE DID NOT SPRING BACK, AND THE DEVICE WOULD NOT FIRE. THE SURGEON REMOVED THE RELOAD AND NEW RELOAD WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC THORASCOPIC PNEUMONECTOMY, WHILE IN THE BRONCHUS SECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | WHITE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF UPPER LEFT LUNG CANCER SURGERY, WHEN SEVERING VASCULAR TISSUE, AFTER THE DEVICE WAS FIRED, IT WAS NOTED THAT SOME STAPLES MALFORMED. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | WHITE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF UPPER LEFT LUNG CANCER SURGERY, WHEN SEVERING VASCULAR TISSUE, AFTER THE DEVICE WAS FIRED, IT WAS NOTED THAT SOME STAPLES MALFORMED. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, HENCE THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
27-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ABSORBATACK | Mechanics Altered | A0507 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE DEVICE WAS ABLE TO FIRE ONLY THREE TACKS. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LUNG TRANSPLANT, THE POWERED VASCULAR STAPLER DEPLOYED STAPLES, BUT THE KNIFE DID NOT REVERSE. THE SURGEON USED THE MANUAL OVERRIDE. NO HARM TO PATIENT REPORTED.
13-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY FOR RADICAL COLON CANCER, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LUNG TRANSPLANT, THE POWERED VASCULAR STAPLER DEPLOYED STAPLES, BUT THE KNIFE DID NOT REVERSE. THE SURGEON USED THE MANUAL OVERRIDE. NO HARM TO PATIENT REPORTED.
27-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
27-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
27-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | GIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
27-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | GIA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
04-01-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED BY THE SALES REP THAT THE PVE35A STAPLER'S JAWS WERE STUCK CLOSED AND THE SURGEON WAS UNABLE TO OPEN IT TO RELOAD. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
25-10-2019 | Injury | us surgical puerto rico | 27-01-2020 | TA | Mechanics Altered | A0507 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, WHEN FIRING THE STAPLER DURING AN OPEN BOWEL CASE, THE SURGEON PERFORMED FULL AND COMPLETE SQUEEZES OF THE HANDLE BUT STAPLES DID NOT DEPLOY. IT WAS STATED THAT THE SURGEON USED TWO STAPLERS FROM THE SAME BOX AND NEITHER FIRED ANY STAPLES. THE ACCOUNT WAS ABLE TO CUT THE TISSUE THINKING THAT THE STAPLES HAD DEPLOYED. THE SURGEON THEN HAD TO HAND SEW. THE EVENT RESULTED TO UNANTICIPATED TISSUE LOSS.
25-10-2019 | Injury | us surgical puerto rico | 27-01-2020 | TA | Mechanics Altered | A0507 | No Code Available | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, WHEN FIRING THE STAPLER DURING AN OPEN BOWEL CASE, THE SURGEON PERFORMED FULL AND COMPLETE SQUEEZES OF THE HANDLE BUT STAPLES DID NOT DEPLOY. IT WAS STATED THAT THE SURGEON USED TWO STAPLERS FROM THE SAME BOX AND NEITHER FIRED ANY STAPLES. THE ACCOUNT WAS ABLE TO CUT THE TISSUE THINKING THAT THE STAPLES HAD DEPLOYED. THE SURGEON THEN HAD TO HAND SEW. THE EVENT RESULTED TO UNANTICIPATED TISSUE LOSS.
27-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
27-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
07-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO STOMACH DURING A LAPAROSCOPIC GASTRIC SLEEVE, IT WAS STATED THAT THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE HENCE THE DEVICE DID NOT FIRE. IT WAS THE SAME HANDLE USED ON THE WHOLE PROCEDURE AND ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-09-2019 | Malfunction | covidien lp | 27-01-2020 | GIA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: GIA 100MM-3.5 REF#GIA10038S, LOT#P7G1150KLX, SERIAL # (B)(4), BROKE WHEN ATTENDING USED IT ON THE PATIENT, THE STAPLE FIRED WITHOUT RELEASING THE STAPLES. NO HARM TO THE PATIENT, ATTENDING AWARE, MANAGEMENT AWARE, STAPLER SENT TO STERILE PROCESSING AND DISTRIBUTION. MANUFACTURER RESPONSE FOR STAPLER, GIA (PER SITE REPORTER). NONE TO DATE.
27-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
27-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
27-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
27-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Fluid/Blood Leak | A050401 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, AFTER LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, PATIENT SHOWED AN ANASTOMOTIC LEAKAGE AND THERE WAS TISSUE DAMAGE.
10-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A ROBOTIC LAPAROSCOPIC PROSTATECTOMY, THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE. THE JAWS WERE REMOVED BY RETRACTING THE KNOB ON THE TOP OF THE CARTRIDGE SIDE OF THE LOADING UNIT. THERE WAS NO PATIENT INJURY.
10-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A ROBOTIC LAPAROSCOPIC PROSTATECTOMY, THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE. THE JAWS WERE REMOVED BY RETRACTING THE KNOB ON THE TOP OF THE CARTRIDGE SIDE OF THE LOADING UNIT. THERE WAS NO PATIENT INJURY.
10-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A ROBOTIC LAPAROSCOPIC PROSTATECTOMY, THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE. THE JAWS WERE REMOVED BY RETRACTING THE KNOB ON THE TOP OF THE CARTRIDGE SIDE OF THE LOADING UNIT. THERE WAS NO PATIENT INJURY.
10-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, ON A ROBOTIC LAPAROSCOPIC PROSTATECTOMY, THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE. THE JAWS WERE REMOVED BY RETRACTING THE KNOB ON THE TOP OF THE CARTRIDGE SIDE OF THE LOADING UNIT. THERE WAS NO PATIENT INJURY.
09-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STAPLING THE JAWS OF THE DEVICE LOCKED ON TISSUE, AND WAS REMOVED BY RETRACTING THE KNIFE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. NO PATIENT HARM.
12-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN DETACHING A LUNG PARENCHYMA DURING A LAPAROSCOPIC RADICAL LUNG CANCER SURGERY, THE STAPLE LINE WAS INCOMPLETE DISTALLY. IT WAS ALSO REPORTED THAT SOME OF THE STAPLES DIDN'T FORM PROPERLY. IT WAS THE SAME HANDLE USED ON THE WHOLE PROCEDURE AND ANOTHER RELOAD WAS USED. THE STAPLE LINE WAS FIXED BY MANUAL SUTURING TO COMPLETE THE CASE.
12-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN DETACHING A LUNG PARENCHYMA DURING A LAPAROSCOPIC RADICAL LUNG CANCER SURGERY, THE STAPLE LINE WAS INCOMPLETE DISTALLY. IT WAS ALSO REPORTED THAT SOME OF THE STAPLES DIDN'T FORM PROPERLY. IT WAS THE SAME HANDLE USED ON THE WHOLE PROCEDURE AND ANOTHER RELOAD WAS USED. THE STAPLE LINE WAS FIXED BY MANUAL SUTURING TO COMPLETE THE CASE.
12-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL LUNG CANCER SURGERY, WHEN DETACHING THE RIGHT PULMONARY VEIN, THE STAPLER WAS ABLE TO FIRE, THERE WAS POOR STAPLE FORMATION AND SOME DID NOT FORM. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THEY USED A CLAMP TO FIX THE STAPLE LINE. THEY REPLACED WITH A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL LUNG CANCER SURGERY, WHEN DETACHING THE RIGHT PULMONARY VEIN, THE STAPLER WAS ABLE TO FIRE, THERE WAS POOR STAPLE FORMATION AND SOME DID NOT FORM. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THEY USED A CLAMP TO FIX THE STAPLE LINE. THEY REPLACED WITH A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A RIGHT OPEN RADICAL NEPHRECTOMY, THE JAWS OF THE STAPLER WOULDN'T OPEN AFTER BEING COMPRESSED ON THE TISSUE. ANOTHER LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. IT WAS NOT REPORTED IF THE DEVICE HAD FIRED WHEN THE ISSUE OCCURRED OR HOW THE DEVICE WAS REMOVED FROM TISSUE WHEN THE ISSUE OCCURRED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL LUNG CANCER SURGERY, WHEN DETACHING THE RIGHT PULMONARY VEIN, THE STAPLER WAS ABLE TO FIRE, THERE WAS POOR STAPLE FORMATION AND SOME DID NOT FORM. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THEY USED A CLAMP TO FIX THE STAPLE LINE. THEY REPLACED WITH A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A RIGHT OPEN RADICAL NEPHRECTOMY, THE JAWS OF THE STAPLER WOULDN'T OPEN AFTER BEING COMPRESSED ON THE TISSUE. ANOTHER LIKE STAPLER WAS USED TO COMPLETE THE PROCEDURE. IT WAS NOT REPORTED IF THE DEVICE HAD FIRED WHEN THE ISSUE OCCURRED OR HOW THE DEVICE WAS REMOVED FROM TISSUE WHEN THE ISSUE OCCURRED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
24-12-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | BLUE RELOAD FOR ECHELON 60 | Loose or Intermittent Connection | A1205 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC SURGERY FOR ESOPHAGEAL CANCER, WHEN SEVERING GASTRIC TISSUE, TRIED TO FIRE ECR60B, NOTED THE CARTRIDGE PAN WAS LOOSE. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2019 | Injury | ethicon endo-surgery | 27-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Gas/Air Leak | A050402 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS, THERE WERE FIVE AIR LEAKS IN (B)(6) AND SEVEN IN (B)(6). ON (B)(6) 2019 THE PATIENT RETURNED TO THE OPERATING ROOM. INTRAOPERATIVE FINDINGS-UNREMARKABLE FIBEROPTIC BRONCHOSCOPY WITH AN INTACT RIGHT UPPER LOBE BRONCHIAL CLOSURE. MULTIPLE PARENCHYMAL AIR LEAKS OF THE RIGHT LOWER LOBE. THE REMAINING LUNG, ALTHOUGH THE INFERIOR PULMONARY LIGAMENT HAD BEEN LYSED AND THE PLEURA SURROUNDING THE HILUM HAD BEEN DISSECTED DID NOT COMPLETELY FILL THE RIGHT PLEURAL SPACE.
23-12-2019 | Injury | ethicon endo-surgery | 27-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Gas/Air Leak | A050402 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS, THERE WERE FIVE AIR LEAKS IN (B)(6) AND SEVEN IN (B)(6). ON (B)(6) 2019 THE PATIENT WAS DISCHARGED TO HOME. ON (B)(6) 2019 DEVELOPED SUBCUTANEOUS EMPHYSEMA. A CT SCAN WAS NEGATIVE FOR PNEUMO. NO ADDITIONAL TREATMENT NEEDED.
07-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Retraction Problem | A0510 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE KNIFE DID NOT RETURN TO HOME POSITION AT THE 4TH FIRING. THE DEVICE WAS USED FOR THE DELTA ANASTOMOSIS ON THE STOMACH. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. THE CARTRIDGE COLOR WAS BLUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-12-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Material Twisted/Bent | A040609 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER CARTRIDGE DID NOT FIRE COMPLETELY. THE SURGEON STATED THAT THE STAPLER WAS USED ON A NORMALLY INFLAMED APPENDIX AND THAT PLACEMENT AND CLOSING OF THE STAPLER WORKED JUST FINE. HOWEVER, WHEN FIRING THE STAPLER, IMMEDIATELY A VERY LOUD CRACKING WAS HEARD. THE FIRE LEVER COULD NOT BE MOVED TO A COMPLETE FIRE POSITION, I.E. ONLY HALF OF THE CARTRIDGE WAS FIRED. UPON RE-OPENING THE STAPLER, THE SURGEON IDENTIFIED THAT THE FIRED STAPLES WERE MOSTLY BENT AND EITHER FLOATED ON THE SURFACE OF THE CARTRIDGE OR HAD FALLEN INTO THE ABDOMEN. REMOVAL OF THE STAPLER WORKED FINE. THE SURGEON USED A COMPETITIVE DEVICE TO COMPLETE THE OPERATION. SURGERY WAS DELAYED FOR AN UNKNOWN AMOUNT OF TIME, BUT WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
23-12-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Noise, Audible | A0508 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER CARTRIDGE DID NOT FIRE COMPLETELY. THE SURGEON STATED THAT THE STAPLER WAS USED ON A NORMALLY INFLAMED APPENDIX AND THAT PLACEMENT AND CLOSING OF THE STAPLER WORKED JUST FINE. HOWEVER, WHEN FIRING THE STAPLER, IMMEDIATELY A VERY LOUD CRACKING WAS HEARD. THE FIRE LEVER COULD NOT BE MOVED TO A COMPLETE FIRE POSITION, I.E. ONLY HALF OF THE CARTRIDGE WAS FIRED. UPON RE-OPENING THE STAPLER, THE SURGEON IDENTIFIED THAT THE FIRED STAPLES WERE MOSTLY BENT AND EITHER FLOATED ON THE SURFACE OF THE CARTRIDGE OR HAD FALLEN INTO THE ABDOMEN. REMOVAL OF THE STAPLER WORKED FINE. THE SURGEON USED A COMPETITIVE DEVICE TO COMPLETE THE OPERATION. SURGERY WAS DELAYED FOR AN UNKNOWN AMOUNT OF TIME, BUT WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON CANCER SURGERY, WHILE PREPARING TO REMOVE THE COLON, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2019 | Injury | ethicon endo-surgery | 27-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS, THERE WERE FIVE AIR LEAKS IN (B)(6) AND SEVEN IN (B)(6). POD 1 CHEST TUBE WAS SIGNIFICANT FOR AIR LEAK AND LEFT TO SUCTION. POD 2 PLACED TO WATER SEAL TRIAL WITH INCREASED SUBCUTANEOUS EMPHYSEMA AND PLACED BACK TO SUCTION. POD 3 HE WAS NEGATIVE FOR AIR LEAK AND PLACED TO WATERSEAL AGAIN. HE TOLERATED WATERSEAL TRIAL OVERNIGHT AND WAS CLAMPED THE FOLLOWING MORNING. FOLLOW UP CXR WAS CONTINUED TO BE NEGATIVE FOR PNEUMOTHORAX. POD 4 CHEST TUBE WAS DISCONTINUED WITHOUT DIFFICULTY. FOLLOW UP CXR WAS CONCERNING FOR PNEUMOTHORAX REQUIRING FURTHER EVALUATION WITH CHEST CT. CHEST CT REVIEWED WITH DR. CANE AND PATIENT READY FOR DISCHARGE. TODAY, PATIENT WAS VOIDING, AMBULATING, TOLERATING PO, ON ROOM AIR AND OTHERWISE READY FOR DISCHARGE TO HOME ON (B)(6) 2019.
09-04-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO MEDIUM AND THICK TISSUE DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE WAS UNABLE TO SQUEEZE THE HANDLE, HENCE, THE DEVICE DID NOT FIRE. IT WAS ALSO REPORTED THAT THE CARTRIDGES WERE STUCK DURING THE FIRING PROCESS AND WERE TRIED WITH NEW STAPLERS BUT THE PROBLEM WENT ON. A NEW CARTRIDGE AND A NEW STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-04-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO MEDIUM AND THICK TISSUE DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE WAS UNABLE TO SQUEEZE THE HANDLE, HENCE, THE DEVICE DID NOT FIRE. IT WAS ALSO REPORTED THAT THE CARTRIDGES WERE STUCK DURING THE FIRING PROCESS AND WERE TRIED WITH NEW STAPLERS BUT THE PROBLEM WENT ON. A NEW CARTRIDGE AND A NEW STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE THE DEVICE WAS BEING APPLIED TO MEDIUM AND THICK TISSUE DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE WAS UNABLE TO SQUEEZE THE HANDLE, HENCE, THE DEVICE DID NOT FIRE. IT WAS ALSO REPORTED THAT THE CARTRIDGES WERE STUCK DURING THE FIRING PROCESS AND WERE TRIED WITH NEW STAPLERS BUT THE PROBLEM WENT ON. A NEW CARTRIDGE AND A NEW STAPLER WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING THE SURGERY OF TUMOR EXCISION, DURING THE GASTROINTESTINAL ANASTOMOSIS, AFTER THE DEVICE WAS FIRED, IT WAS NOTED THAT THE ANASTOMOTIC SITE WAS INCOMPLETE. SUTURE WAS USED TO OVERSEW. THE SURGERY DELAYED AND UNKNOWN AMOUNT OF TIME. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
13-10-2019 | Injury | ethicon endo-surgery | 27-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Gas/Air Leak | A050402 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS, THERE WERE FIVE AIR LEAKS IN NOVEMBER AND SEVEN IN DECEMBER. (B)(6) 2019 PATIENT RETURNED TO OR POD#9 BRONCHOSCOPY, LEFT THORACOTOMY, HYLUM MOBILIZATION, PNEUMOSTASIS, DIAGPHRAGMATIC MOBILIZATION, INTERCOSTAL BLOCK FOR PERSISTENT AIR LEAK AND PNEUMOTHORAX. 10/17 CT. PATIENT DISCHARGED TO HOME ON (B)(6) 2019.
12-09-2018 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON PROCEDURE, THE STAPLING WAS NOT EFFECTIVE DESPITE APPLYING PRESSURE. THE STAPLE LINE WAS INCOMPLETE. THERE WAS TISSUE LOSS AND TISSUE DETERIORATION IN THE STAPLING AREA. THERE WAS PROLONGED ANESTHESIA TIME AND THE OPERATING TIME WAS EXTENDED BY MORE THAN 30 MINUTES. THE HOSPITAL STAY WAS ALSO EXTENDED. A NEW STAPLER WAS USED TO COMPLETE THE CASE.
12-09-2018 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON PROCEDURE, THE STAPLING WAS NOT EFFECTIVE DESPITE APPLYING PRESSURE. THE STAPLE LINE WAS INCOMPLETE. THERE WAS TISSUE LOSS AND TISSUE DETERIORATION IN THE STAPLING AREA. THERE WAS PROLONGED ANESTHESIA TIME AND THE OPERATING TIME WAS EXTENDED BY MORE THAN 30 MINUTES. THE HOSPITAL STAY WAS ALSO EXTENDED. A NEW STAPLER WAS USED TO COMPLETE THE CASE.
13-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE CLOSING THE ENTEROTOMY AFTER SIDE-BY-SIDE ANASTOMOSIS OF BOWEL IN AN OPEN BOWEL RESECTION PROCEDURE, THERE WAS BLEEDING AT THE STAPLE LINE. THE SURGEON USED SUTURE TO OVERSEW THE STAPLE LINE.
13-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE CLOSING THE ENTEROTOMY AFTER SIDE-BY-SIDE ANASTOMOSIS OF BOWEL IN AN OPEN BOWEL RESECTION PROCEDURE, THERE WAS BLEEDING AT THE STAPLE LINE. THE SURGEON USED SUTURE TO OVERSEW THE STAPLE LINE.
13-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHILE CLOSING THE ENTEROTOMY AFTER SIDE-BY-SIDE ANASTOMOSIS OF BOWEL IN AN OPEN BOWEL RESECTION PROCEDURE, THERE WAS BLEEDING AT THE STAPLE LINE. THE SURGEON USED SUTURE TO OVERSEW THE STAPLE LINE.
09-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Failure to Fire | A050501 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, AFTER STAPLER WAS FIRED AND REMOVED FROM FIELD, THE ANASTOMOSIS FELL APART. UPON INSPECTION, THE STAPLER DONUTS WERE OBSERVED TO BE COMPLETE BY THE ATTENDING SURGEON. REQUIRED ADDITIONAL ANASTOMOSIS.
09-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, AFTER STAPLER WAS FIRED AND REMOVED FROM FIELD, THE ANASTOMOSIS FELL APART. UPON INSPECTION, THE STAPLER DONUTS WERE OBSERVED TO BE COMPLETE BY THE ATTENDING SURGEON. REQUIRED ADDITIONAL ANASTOMOSIS.
09-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, AFTER STAPLER WAS FIRED AND REMOVED FROM FIELD, THE ANASTOMOSIS FELL APART. UPON INSPECTION, THE STAPLER DONUTS WERE OBSERVED TO BE COMPLETE BY THE ATTENDING SURGEON. REQUIRED ADDITIONAL ANASTOMOSIS.
09-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, AFTER STAPLER WAS FIRED AND REMOVED FROM FIELD, THE ANASTOMOSIS FELL APART. UPON INSPECTION, THE STAPLER DONUTS WERE OBSERVED TO BE COMPLETE BY THE ATTENDING SURGEON. REQUIRED ADDITIONAL ANASTOMOSIS.
09-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, AFTER STAPLER WAS FIRED AND REMOVED FROM FIELD, THE ANASTOMOSIS FELL APART. UPON INSPECTION, THE STAPLER DONUTS WERE OBSERVED TO BE COMPLETE BY THE ATTENDING SURGEON. REQUIRED ADDITIONAL ANASTOMOSIS.
09-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, AFTER STAPLER WAS FIRED AND REMOVED FROM FIELD, THE ANASTOMOSIS FELL APART. UPON INSPECTION, THE STAPLER DONUTS WERE OBSERVED TO BE COMPLETE BY THE ATTENDING SURGEON. REQUIRED ADDITIONAL ANASTOMOSIS.
17-12-2014 | Injury | covidien lp llc north haven | 27-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE CASE, THERE WERE FORTY-ONE PATIENTS WITH ESOPHAGEAL CANCER IN TWO GROUPS UNDERWENT TOTALLY MINIMALLY INVASIVE IVOR-LEWIS ESOPHAGECTOMY (MIIE) WITH SINGLE-INCISION VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS). TWO ANASTOMOSIS LEAKAGES WERE IDENTIFIED AT POST-OPERATIVE DATE 12 AND 14. PATIENTS WERE FOLLOWED AT 3- AND 4-MONTH INTERVALS AND AT 6-MONTH INTERVALS AFTER THE SECOND YEAR. MAJOR MORBIDITIES, INCLUDING INTESTINAL OBSTRUCTION AND ANASTOMOTIC LEAKAGE, OCCURRED IN THREE PATIENTS AFTER DISCHARGE IN GROUP 1 WHILE ONE PATIENT DEVELOPED LATE-STAGE ANASTOMOTIC LEAKAGE IN GROUP 2.
17-12-2014 | Injury | covidien lp llc north haven | 27-01-2020 | UNKNOWN ENDO GIA SULU | Fluid/Blood Leak | A050401 | Obstruction/Occlusion | Event Description: ACCORDING TO THE LITERATURE CASE, THERE WERE FORTY-ONE PATIENTS WITH ESOPHAGEAL CANCER IN TWO GROUPS UNDERWENT TOTALLY MINIMALLY INVASIVE IVOR-LEWIS ESOPHAGECTOMY (MIIE) WITH SINGLE-INCISION VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS). TWO ANASTOMOSIS LEAKAGES WERE IDENTIFIED AT POST-OPERATIVE DATE 12 AND 14. PATIENTS WERE FOLLOWED AT 3- AND 4-MONTH INTERVALS AND AT 6-MONTH INTERVALS AFTER THE SECOND YEAR. MAJOR MORBIDITIES, INCLUDING INTESTINAL OBSTRUCTION AND ANASTOMOTIC LEAKAGE, OCCURRED IN THREE PATIENTS AFTER DISCHARGE IN GROUP 1 WHILE ONE PATIENT DEVELOPED LATE-STAGE ANASTOMOTIC LEAKAGE IN GROUP 2.
17-12-2014 | Injury | us surgical puerto rico | 27-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE CASE, THERE WERE FORTY-ONE PATIENTS WITH ESOPHAGEAL CANCER IN TWO GROUPS UNDERWENT TOTALLY MINIMALLY INVASIVE IVOR-LEWIS ESOPHAGECTOMY (MIIE) WITH SINGLE-INCISION VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS). TWO ANASTOMOSIS LEAKAGES WERE IDENTIFIED AT POST-OPERATIVE DATE 12 AND 14. PATIENTS WERE FOLLOWED AT 3- AND 4-MONTH INTERVALS AND AT 6-MONTH INTERVALS AFTER THE SECOND YEAR. MAJOR MORBIDITIES, INCLUDING INTESTINAL OBSTRUCTION AND ANASTOMOTIC LEAKAGE, OCCURRED IN THREE PATIENTS AFTER DISCHARGE IN GROUP 1 WHILE ONE PATIENT DEVELOPED LATE-STAGE ANASTOMOTIC LEAKAGE IN GROUP 2.
17-12-2014 | Injury | us surgical puerto rico | 27-01-2020 | UNKNOWN EEA | Fluid/Blood Leak | A050401 | Obstruction/Occlusion | Event Description: ACCORDING TO THE LITERATURE CASE, THERE WERE FORTY-ONE PATIENTS WITH ESOPHAGEAL CANCER IN TWO GROUPS UNDERWENT TOTALLY MINIMALLY INVASIVE IVOR-LEWIS ESOPHAGECTOMY (MIIE) WITH SINGLE-INCISION VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS). TWO ANASTOMOSIS LEAKAGES WERE IDENTIFIED AT POST-OPERATIVE DATE 12 AND 14. PATIENTS WERE FOLLOWED AT 3- AND 4-MONTH INTERVALS AND AT 6-MONTH INTERVALS AFTER THE SECOND YEAR. MAJOR MORBIDITIES, INCLUDING INTESTINAL OBSTRUCTION AND ANASTOMOTIC LEAKAGE, OCCURRED IN THREE PATIENTS AFTER DISCHARGE IN GROUP 1 WHILE ONE PATIENT DEVELOPED LATE-STAGE ANASTOMOTIC LEAKAGE IN GROUP 2.
13-01-2020 | Injury | us surgical puerto rico | 27-01-2020 | ENDO GIA ULTRA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY NODULECTOMY, WHILE REMOVING THE PULMONARY PARENCHYMA, THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS FIXED BY HAND SEWING. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
13-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY NODULECTOMY, WHILE REMOVING THE PULMONARY PARENCHYMA, THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS FIXED BY HAND SEWING. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
13-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY NODULECTOMY, WHILE REMOVING THE PULMONARY PARENCHYMA, THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS FIXED BY HAND SEWING. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Loose or Intermittent Connection | A1205 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, PRODUCT PRESENTED PROBLEM. IT HAPPENED DURING THE USE OF THE LAST CARTRIDGE, AFTER THE COMPRESSION OF THE TISSUE, AT THE TIME OF FIRING THE TISSUE SLIPPED LEAVING THE STAPLES LOOSE AND CUTTING PART OF THE TISSUE. HOWEVER, THE KNIFE BLADE WAS COMPLETELY DISPLACED, BUT SINCE THE TISSUE WAS COMPRESSED AND DISPLACED, ONLY THE TISSUE WAS PARTIALLY CUT. THE STAPLES WERE FORMED NOT ENTIRELY BUT MOST OF THEM BEING LOOSE. AT THE TIME THE BLADE WAS DISPLACED, THE TISSUE MOVED TOWARDS THE DISTAL PART OF THE STAPLER GIVING THE IMPRESSION OF BEING REPLACED, WHEN THE SURGEON RELEASES THE TISSUE IT WAS PARTIALLY CUT AND STAPLED. THERE WAS ONLY A DELAY DURING THE PROCEDURE WITHOUT CAUSING GREATER COMPLICATION OR CONSEQUENCE FOR THE PATIENT, SINCE THE DOCTOR RECEIVED ANOTHER CARTRIDGE BEING ABLE TO COMPLETE THE "GRAPEFRUIT" LINE AND REINFORCING WITH STITCHES. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
06-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, PRODUCT PRESENTED PROBLEM. IT HAPPENED DURING THE USE OF THE LAST CARTRIDGE, AFTER THE COMPRESSION OF THE TISSUE, AT THE TIME OF FIRING THE TISSUE SLIPPED LEAVING THE STAPLES LOOSE AND CUTTING PART OF THE TISSUE. HOWEVER, THE KNIFE BLADE WAS COMPLETELY DISPLACED, BUT SINCE THE TISSUE WAS COMPRESSED AND DISPLACED, ONLY THE TISSUE WAS PARTIALLY CUT. THE STAPLES WERE FORMED NOT ENTIRELY BUT MOST OF THEM BEING LOOSE. AT THE TIME THE BLADE WAS DISPLACED, THE TISSUE MOVED TOWARDS THE DISTAL PART OF THE STAPLER GIVING THE IMPRESSION OF BEING REPLACED, WHEN THE SURGEON RELEASES THE TISSUE IT WAS PARTIALLY CUT AND STAPLED. THERE WAS ONLY A DELAY DURING THE PROCEDURE WITHOUT CAUSING GREATER COMPLICATION OR CONSEQUENCE FOR THE PATIENT, SINCE THE DOCTOR RECEIVED ANOTHER CARTRIDGE BEING ABLE TO COMPLETE THE "GRAPEFRUIT" LINE AND REINFORCING WITH STITCHES. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2019 | Injury | ethicon endo-surgery | 27-01-2020 | POWERED PLUS COMPACT 45 | Adverse Event Without Identified Device or Use Problem | A24 | No Code Available | Event Description: IT WAS REPORTED THAT DURING LOBECTOMY AND WEDGE RESECTIONS OVER THE LAST TWO MONTHS, THERE WERE FIVE AIR LEAKS IN NOVEMBER AND SEVEN IN DECEMBER. THE PATIENT EXPERIENCED A RIGHT APICAL PNEUMOTHORAX.
13-12-2019 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | Blood Loss | Event Description: IT WAS REPORTED THAT THE SURGEON WAS USING A POWERED VASCULAR STAPLER WITH A VASCULAR CARTRIDGE ON THE PULMONARY ARTERY. HE CLOSED THE DEVICE ON THE PA AND WENT TO FIRE, HE DID NOTICE THAT THERE FELT A LITTLE RESISTANCE WHEN HE ENGAGED THE RED FIRING BUTTON, BUT NOT TO CLOSE THE DEVICE. HE WENT TO FIRE BUT NOTHING HAPPENED, BUT THEN WHEN HE WENT TO OPEN THE DEVICE IT WOULD NOT DO SO AT ALL. HE DID NOT HEAR THE MOTOR ENGAGE AT ALL, WHICH HE THOUGHT WAS STRANGE AS IT NORMALLY ONLY JAMS WHEN FIRED. THE PRESSED THE REVERSE KNIFE BLADE BUTTON BUT NOTHING HAPPENED. HE THEN TOOK THE BATTERY OUT OF THE PSEE45A HE WAS ALREADY USING ON THE LUNG IN THE KIT AND TOOK THE BATTERY OUT OF THE PVE35A AND REPLACED WITH THE OTHER AND THEN MANAGED TO REVERSE AND GET THE GUN OPEN, BUT THIS TORE THE PA AND SO HE HAD TO CLAMP WITH THE GUN TO CONTROL THE BLEEDING. HE THEN HAD TO OPEN THE PATIENT, WITH A THORACOTOMY TO CONTROL THE BLEEDING AND COMPLETE THE PROCEDURE. NO TRANSFUSION NEEDED, BUT BLOOD WAS ORDERED IN CASE. HAD TO USE THE VASCULAR CLAMPS TO CONTROL IN THE END AND ANOTHER VASCULAR GUN TO COMPLETE. DATA IS ALL NOTED ON POST OP NOTES AND A DATEX WILL BE COMPLETED FOR THIS. 800ML OF BLOOD LOSS AND ABOUT FIFTEEN MINUTES EXTRA ON THE TABLE ACCORDING TO THE SCRUB NURSE. COMMENTS FROM SCRUB TEAM WAS IT WAS AN OOZY CASE AND THE PA WAS ALREADY QUITE OOZY AND ANATOMY WAS STRANGE. I HAVE SPOKEN WITH THE CONSULTANT TODAY AND THE PATIENT IS FINE.
13-12-2019 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON WAS USING A POWERED VASCULAR STAPLER WITH A VASCULAR CARTRIDGE ON THE PULMONARY ARTERY. HE CLOSED THE DEVICE ON THE PA AND WENT TO FIRE, HE DID NOTICE THAT THERE FELT A LITTLE RESISTANCE WHEN HE ENGAGED THE RED FIRING BUTTON, BUT NOT TO CLOSE THE DEVICE. HE WENT TO FIRE BUT NOTHING HAPPENED, BUT THEN WHEN HE WENT TO OPEN THE DEVICE IT WOULD NOT DO SO AT ALL. HE DID NOT HEAR THE MOTOR ENGAGE AT ALL, WHICH HE THOUGHT WAS STRANGE AS IT NORMALLY ONLY JAMS WHEN FIRED. THE PRESSED THE REVERSE KNIFE BLADE BUTTON BUT NOTHING HAPPENED. HE THEN TOOK THE BATTERY OUT OF THE PSEE45A HE WAS ALREADY USING ON THE LUNG IN THE KIT AND TOOK THE BATTERY OUT OF THE PVE35A AND REPLACED WITH THE OTHER AND THEN MANAGED TO REVERSE AND GET THE GUN OPEN, BUT THIS TORE THE PA AND SO HE HAD TO CLAMP WITH THE GUN TO CONTROL THE BLEEDING. HE THEN HAD TO OPEN THE PATIENT, WITH A THORACOTOMY TO CONTROL THE BLEEDING AND COMPLETE THE PROCEDURE. NO TRANSFUSION NEEDED, BUT BLOOD WAS ORDERED IN CASE. HAD TO USE THE VASCULAR CLAMPS TO CONTROL IN THE END AND ANOTHER VASCULAR GUN TO COMPLETE. DATA IS ALL NOTED ON POST OP NOTES AND A DATEX WILL BE COMPLETED FOR THIS. 800ML OF BLOOD LOSS AND ABOUT FIFTEEN MINUTES EXTRA ON THE TABLE ACCORDING TO THE SCRUB NURSE. COMMENTS FROM SCRUB TEAM WAS IT WAS AN OOZY CASE AND THE PA WAS ALREADY QUITE OOZY AND ANATOMY WAS STRANGE. I HAVE SPOKEN WITH THE CONSULTANT TODAY AND THE PATIENT IS FINE.
13-12-2019 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Battery Problem | A0705 | Blood Loss | Event Description: IT WAS REPORTED THAT THE SURGEON WAS USING A POWERED VASCULAR STAPLER WITH A VASCULAR CARTRIDGE ON THE PULMONARY ARTERY. HE CLOSED THE DEVICE ON THE PA AND WENT TO FIRE, HE DID NOTICE THAT THERE FELT A LITTLE RESISTANCE WHEN HE ENGAGED THE RED FIRING BUTTON, BUT NOT TO CLOSE THE DEVICE. HE WENT TO FIRE BUT NOTHING HAPPENED, BUT THEN WHEN HE WENT TO OPEN THE DEVICE IT WOULD NOT DO SO AT ALL. HE DID NOT HEAR THE MOTOR ENGAGE AT ALL, WHICH HE THOUGHT WAS STRANGE AS IT NORMALLY ONLY JAMS WHEN FIRED. THE PRESSED THE REVERSE KNIFE BLADE BUTTON BUT NOTHING HAPPENED. HE THEN TOOK THE BATTERY OUT OF THE PSEE45A HE WAS ALREADY USING ON THE LUNG IN THE KIT AND TOOK THE BATTERY OUT OF THE PVE35A AND REPLACED WITH THE OTHER AND THEN MANAGED TO REVERSE AND GET THE GUN OPEN, BUT THIS TORE THE PA AND SO HE HAD TO CLAMP WITH THE GUN TO CONTROL THE BLEEDING. HE THEN HAD TO OPEN THE PATIENT, WITH A THORACOTOMY TO CONTROL THE BLEEDING AND COMPLETE THE PROCEDURE. NO TRANSFUSION NEEDED, BUT BLOOD WAS ORDERED IN CASE. HAD TO USE THE VASCULAR CLAMPS TO CONTROL IN THE END AND ANOTHER VASCULAR GUN TO COMPLETE. DATA IS ALL NOTED ON POST OP NOTES AND A DATEX WILL BE COMPLETED FOR THIS. 800ML OF BLOOD LOSS AND ABOUT FIFTEEN MINUTES EXTRA ON THE TABLE ACCORDING TO THE SCRUB NURSE. COMMENTS FROM SCRUB TEAM WAS IT WAS AN OOZY CASE AND THE PA WAS ALREADY QUITE OOZY AND ANATOMY WAS STRANGE. I HAVE SPOKEN WITH THE CONSULTANT TODAY AND THE PATIENT IS FINE.
13-12-2019 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Battery Problem | A0705 | No Code Available | Event Description: IT WAS REPORTED THAT THE SURGEON WAS USING A POWERED VASCULAR STAPLER WITH A VASCULAR CARTRIDGE ON THE PULMONARY ARTERY. HE CLOSED THE DEVICE ON THE PA AND WENT TO FIRE, HE DID NOTICE THAT THERE FELT A LITTLE RESISTANCE WHEN HE ENGAGED THE RED FIRING BUTTON, BUT NOT TO CLOSE THE DEVICE. HE WENT TO FIRE BUT NOTHING HAPPENED, BUT THEN WHEN HE WENT TO OPEN THE DEVICE IT WOULD NOT DO SO AT ALL. HE DID NOT HEAR THE MOTOR ENGAGE AT ALL, WHICH HE THOUGHT WAS STRANGE AS IT NORMALLY ONLY JAMS WHEN FIRED. THE PRESSED THE REVERSE KNIFE BLADE BUTTON BUT NOTHING HAPPENED. HE THEN TOOK THE BATTERY OUT OF THE PSEE45A HE WAS ALREADY USING ON THE LUNG IN THE KIT AND TOOK THE BATTERY OUT OF THE PVE35A AND REPLACED WITH THE OTHER AND THEN MANAGED TO REVERSE AND GET THE GUN OPEN, BUT THIS TORE THE PA AND SO HE HAD TO CLAMP WITH THE GUN TO CONTROL THE BLEEDING. HE THEN HAD TO OPEN THE PATIENT, WITH A THORACOTOMY TO CONTROL THE BLEEDING AND COMPLETE THE PROCEDURE. NO TRANSFUSION NEEDED, BUT BLOOD WAS ORDERED IN CASE. HAD TO USE THE VASCULAR CLAMPS TO CONTROL IN THE END AND ANOTHER VASCULAR GUN TO COMPLETE. DATA IS ALL NOTED ON POST OP NOTES AND A DATEX WILL BE COMPLETED FOR THIS. 800ML OF BLOOD LOSS AND ABOUT FIFTEEN MINUTES EXTRA ON THE TABLE ACCORDING TO THE SCRUB NURSE. COMMENTS FROM SCRUB TEAM WAS IT WAS AN OOZY CASE AND THE PA WAS ALREADY QUITE OOZY AND ANATOMY WAS STRANGE. I HAVE SPOKEN WITH THE CONSULTANT TODAY AND THE PATIENT IS FINE.
13-11-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | FLEX 60 ARTICULATING | Mechanical Problem | A05 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SPLENECTOMY, THE DEVICE LOCKED OUT AND THE MANUAL OVER RIDE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
16-10-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DOCTOR PUT A CARTRIDGE INTO THE JAWS OF THE STAPLER, CLOSED THE JAWS OF THE STAPLER AND THE JAWS WOULDN'T REOPEN AFTER CLOSING. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
19-11-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Problem | A1501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AT THE FIRST FIRING, THE DEVICE DID NOT MOVE AT ALL WHEN THE DEVICE WAS USED ON THE RECTUM. THE TARGET TISSUE WAS SWOLLEN INTENSELY. A COMPETITIVE DEVICE WITH BLACK CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEPLOYED STAPLES WERE NOT FORMED AS INTENDED AFTER FIRING THE DEVICE. THE DEVICE WAS USED AT AN ATRIAL APPENDAGE. THE BLEEDING OCCURRED BY MALFORMED STAPLING, BUT THE SURGEON STOPPED THE BLEEDING AND CLOSED WITH THE REPLACEMENT; ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE IS NO PROBLEM FOR THE PATIENT¿S CONDITION. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | EEA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING A CORRECTION OF RECTAL PROLAPSE AND HEMORRHOIDECTOMY PROCEDURE, THERE WAS A POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE. ANOTHER STAPLER WAS USED TO COMPLETE THE PROCEDURE AND TO RESOLVE THE ISSUE THE DEVICE WAS RESECTED AND R E-STAPLED.
10-01-2020 | Injury | covidien lp llc north haven | 27-01-2020 | EEA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A CORRECTION OF RECTAL PROLAPSE AND HEMORRHOIDECTOMY PROCEDURE, THERE WAS A POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE. ANOTHER STAPLER WAS USED TO COMPLETE THE PROCEDURE AND TO RESOLVE THE ISSUE THE DEVICE WAS RESECTED AND R E-STAPLED.
20-11-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER RELEASED STAPLES, HOWEVER, TISSUE HADN'T BEEN CUT. A SECOND DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
08-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON SQUEEZED THE HANDLE, IT STAYED LOCKED, HENCE THE DEVICE DID NOT FIRE. THE SURGEON USED A SECOND STAPLER AND WHEN HE SQUEEZED THE HANDLE, IT STAYED LOCKED TOO. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON SQUEEZED THE HANDLE, IT STAYED LOCKED, HENCE THE DEVICE DID NOT FIRE. THE SURGEON USED A SECOND STAPLER AND WHEN HE SQUEEZED THE HANDLE, IT STAYED LOCKED TOO. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE SURGEON USED AN ETS45. AFTER THE STAPLER WAS FIRED, IT CUT BUT DID NOT STAPLE. SURGEON HAD TO HAND SEW. IT IS UNKNOWN IF THERE WERE ANY ADVERSE CONSEQUENCES TO THE PATIENT, HOWEVER, NONE WERE REPORTED. THIS IS ALL THE DETAIL THE SALES REP WAS GIVEN BY THE CUSTOMER.
08-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON SQUEEZED THE HANDLE, IT STAYED LOCKED, HENCE THE DEVICE DID NOT FIRE. THE SURGEON USED A SECOND STAPLER AND WHEN HE SQUEEZED THE HANDLE, IT STAYED LOCKED TOO. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | GST BLUE RELOAD, 60MM, 6 ROW | Break | A0401 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING A SLEEVE PROCEDURE AFTER GST60B FIRING, THE RELOAD LEFT A PIECE OF PLASTIC OVER THE STAPLER JAWS. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
10-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LEFT COLECTOMY PROCEDURE, WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THEY WERE ABLE TO PUSH THE GREEN BUTTON BUT THE HANDLE COULD NOT BE SQUEEZE AND THE DEVICE COULD NOT BE FIRED.. THE SURGEON THEN USED ANOTHER DEVICE RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-11-2019 | Malfunction | ethicon endo-surgery | 27-01-2020 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER WOULDN'T OPEN. THE DOCTOR TRIED OPENING THE STAPLER USING THE REVERSE BUTTON, SEVERAL TIMES, AS WELL AS USING THE MANUAL OVERRIDE LEVER. IT WAS NOT REPORTED HOW THE JAWS WERE OPENED OR IF THE DEVICE WAS STUCK ON TISSUE WHEN THE DEVICE WOULDN'T OPEN. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-01-2020 | Malfunction | ethicon endo-surgery | 27-01-2020 | ECHELON FLEX 45 COMPACT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE COULD NOT FIRE OR OPEN THE MOUTH DURING THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME.
02-01-2020 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON ENDOPATH STAPLER | Failure to Form Staple | A050704 | Chest Pain | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : INCIDENTAL HAEMOTHORAX AFTER SUBLOBAR RESECTION: DID STAPLE LINE SCRATCH CHEST WALL? AUTHOR : HIDEKI NEGISHI, HIROYOSHI TSUBOCHI, MITSURU MAKI, AND SHUNSUKE ENDO; CITATION: JOURNAL OF SURGICAL CASE REPORTS (2019);12, 1¿2. DOI: 10.1093/JSCR/RJZ276. THE AUTHORS REPORTED A CASE OF A LATE-ONSET HAEMOTHORAX CAUSED BY A STAPLE THAT DIVIDED LUNG PARENCHYMA DURING PULMONARY SEGMENTECTOMY. A (B)(6) MALE PATIENT UNDERWENT RIGHT APICAL AND POSTERIOR SEGMENTECTOMIES VIA VIDEO-ASSISTED THORACIC SURGERY (VATS) FOR EARLY-STAGE LUNG CANCER. THE BRONCHUS AND PULMONARY PARENCHYMA WERE DIVIDED BY ENDOSCOPIC STAPLES (ECHELON®; ETHICON). ON 32ND DAY AFTER DISCHARGE, THE PATIENT WAS TRANSPORTED BY AMBULANCE OWING TO A SUDDEN ONSET OF CHEST PAIN AND DYSPNEA. THE PATIENT DEVELOPED HYPOTENSION AND TACHYCARDIA WITH SYSTOLIC BLOOD PRESSURE OF 50MMHG AND HEART RATE OF 114 BEATS/MIN. THE PATIENT WAS DIAGNOSED WITH HAEMOTHORAX AND AN EMERGENCY OPERATION WAS PERFORMED. A MASSIVE BLOOD COAGULATION IN THE RIGHT THORACIC CAVITY WAS REMOVED. A CONTINUOUS ACTIVE BLEEDING FROM THE INTERCOSTAL ARTERY WAS FOUND NEAR THE FIFTH THORACIC VERTEBRA IN THE VICINITY OF A STUMP OF THE BRONCHUS DIVIDED BY A STAPLE. A THORACOTOMY THROUGH THE FIFTH INTERCOSTAL SPACE WAS THEREFORE PERFORMED TO ACHIEVE HAEMOSTASIS, AND THE BLEEDING WAS STOPPED BY SUTURING. THE BLEEDING POINT WAS COVERED WITH A COLLAGEN-FIBRIN PATCH. THE PATIENT PROGRESSED SATISFACTORILY AFTER THE REOPERATION AND WAS DISCHARGED ON NINTH DAY AFTER REOPERATION. THE POSTERIOR PART OF THE STAPLE LINE OF THE LUNG PARENCHYMA ADJACENT TO THE BLEEDING POINT UNDER FULL INFLATION WAS PARTLY EXPOSED, WHEREAS THE BRONCHIAL STUMP WAS COVERED WITH LUNG. THEREFORE, THE SCRATCHING CREATED BY THE STAPLES WAS ASSUMED TO CAUSE HAEMOTHORAX. IN CONCLUSION, SURGEONS SHOULD RECOGNIZE THE POTENTIAL RISK FOR STAPLES TO SCRATCH ADJACENT ORGANS, AND THE SHARP EDGE OF THE STAPLE LINE OR UNSTABLE STAPLES SHOULD BE TRIMMED DURING VATS TO AVOID STAPLE-ASSOCIATED HAEMOTHORAX.
02-01-2020 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON ENDOPATH STAPLER | Failure to Form Staple | A050704 | Dyspnea | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : INCIDENTAL HAEMOTHORAX AFTER SUBLOBAR RESECTION: DID STAPLE LINE SCRATCH CHEST WALL? AUTHOR : HIDEKI NEGISHI, HIROYOSHI TSUBOCHI, MITSURU MAKI, AND SHUNSUKE ENDO; CITATION: JOURNAL OF SURGICAL CASE REPORTS (2019);12, 1¿2. DOI: 10.1093/JSCR/RJZ276. THE AUTHORS REPORTED A CASE OF A LATE-ONSET HAEMOTHORAX CAUSED BY A STAPLE THAT DIVIDED LUNG PARENCHYMA DURING PULMONARY SEGMENTECTOMY. A (B)(6) MALE PATIENT UNDERWENT RIGHT APICAL AND POSTERIOR SEGMENTECTOMIES VIA VIDEO-ASSISTED THORACIC SURGERY (VATS) FOR EARLY-STAGE LUNG CANCER. THE BRONCHUS AND PULMONARY PARENCHYMA WERE DIVIDED BY ENDOSCOPIC STAPLES (ECHELON®; ETHICON). ON 32ND DAY AFTER DISCHARGE, THE PATIENT WAS TRANSPORTED BY AMBULANCE OWING TO A SUDDEN ONSET OF CHEST PAIN AND DYSPNEA. THE PATIENT DEVELOPED HYPOTENSION AND TACHYCARDIA WITH SYSTOLIC BLOOD PRESSURE OF 50MMHG AND HEART RATE OF 114 BEATS/MIN. THE PATIENT WAS DIAGNOSED WITH HAEMOTHORAX AND AN EMERGENCY OPERATION WAS PERFORMED. A MASSIVE BLOOD COAGULATION IN THE RIGHT THORACIC CAVITY WAS REMOVED. A CONTINUOUS ACTIVE BLEEDING FROM THE INTERCOSTAL ARTERY WAS FOUND NEAR THE FIFTH THORACIC VERTEBRA IN THE VICINITY OF A STUMP OF THE BRONCHUS DIVIDED BY A STAPLE. A THORACOTOMY THROUGH THE FIFTH INTERCOSTAL SPACE WAS THEREFORE PERFORMED TO ACHIEVE HAEMOSTASIS, AND THE BLEEDING WAS STOPPED BY SUTURING. THE BLEEDING POINT WAS COVERED WITH A COLLAGEN-FIBRIN PATCH. THE PATIENT PROGRESSED SATISFACTORILY AFTER THE REOPERATION AND WAS DISCHARGED ON NINTH DAY AFTER REOPERATION. THE POSTERIOR PART OF THE STAPLE LINE OF THE LUNG PARENCHYMA ADJACENT TO THE BLEEDING POINT UNDER FULL INFLATION WAS PARTLY EXPOSED, WHEREAS THE BRONCHIAL STUMP WAS COVERED WITH LUNG. THEREFORE, THE SCRATCHING CREATED BY THE STAPLES WAS ASSUMED TO CAUSE HAEMOTHORAX. IN CONCLUSION, SURGEONS SHOULD RECOGNIZE THE POTENTIAL RISK FOR STAPLES TO SCRATCH ADJACENT ORGANS, AND THE SHARP EDGE OF THE STAPLE LINE OR UNSTABLE STAPLES SHOULD BE TRIMMED DURING VATS TO AVOID STAPLE-ASSOCIATED HAEMOTHORAX.
02-01-2020 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON ENDOPATH STAPLER | Failure to Form Staple | A050704 | Low Blood Pressure/ Hypotension | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : INCIDENTAL HAEMOTHORAX AFTER SUBLOBAR RESECTION: DID STAPLE LINE SCRATCH CHEST WALL? AUTHOR : HIDEKI NEGISHI, HIROYOSHI TSUBOCHI, MITSURU MAKI, AND SHUNSUKE ENDO; CITATION: JOURNAL OF SURGICAL CASE REPORTS (2019);12, 1¿2. DOI: 10.1093/JSCR/RJZ276. THE AUTHORS REPORTED A CASE OF A LATE-ONSET HAEMOTHORAX CAUSED BY A STAPLE THAT DIVIDED LUNG PARENCHYMA DURING PULMONARY SEGMENTECTOMY. A (B)(6) MALE PATIENT UNDERWENT RIGHT APICAL AND POSTERIOR SEGMENTECTOMIES VIA VIDEO-ASSISTED THORACIC SURGERY (VATS) FOR EARLY-STAGE LUNG CANCER. THE BRONCHUS AND PULMONARY PARENCHYMA WERE DIVIDED BY ENDOSCOPIC STAPLES (ECHELON®; ETHICON). ON 32ND DAY AFTER DISCHARGE, THE PATIENT WAS TRANSPORTED BY AMBULANCE OWING TO A SUDDEN ONSET OF CHEST PAIN AND DYSPNEA. THE PATIENT DEVELOPED HYPOTENSION AND TACHYCARDIA WITH SYSTOLIC BLOOD PRESSURE OF 50MMHG AND HEART RATE OF 114 BEATS/MIN. THE PATIENT WAS DIAGNOSED WITH HAEMOTHORAX AND AN EMERGENCY OPERATION WAS PERFORMED. A MASSIVE BLOOD COAGULATION IN THE RIGHT THORACIC CAVITY WAS REMOVED. A CONTINUOUS ACTIVE BLEEDING FROM THE INTERCOSTAL ARTERY WAS FOUND NEAR THE FIFTH THORACIC VERTEBRA IN THE VICINITY OF A STUMP OF THE BRONCHUS DIVIDED BY A STAPLE. A THORACOTOMY THROUGH THE FIFTH INTERCOSTAL SPACE WAS THEREFORE PERFORMED TO ACHIEVE HAEMOSTASIS, AND THE BLEEDING WAS STOPPED BY SUTURING. THE BLEEDING POINT WAS COVERED WITH A COLLAGEN-FIBRIN PATCH. THE PATIENT PROGRESSED SATISFACTORILY AFTER THE REOPERATION AND WAS DISCHARGED ON NINTH DAY AFTER REOPERATION. THE POSTERIOR PART OF THE STAPLE LINE OF THE LUNG PARENCHYMA ADJACENT TO THE BLEEDING POINT UNDER FULL INFLATION WAS PARTLY EXPOSED, WHEREAS THE BRONCHIAL STUMP WAS COVERED WITH LUNG. THEREFORE, THE SCRATCHING CREATED BY THE STAPLES WAS ASSUMED TO CAUSE HAEMOTHORAX. IN CONCLUSION, SURGEONS SHOULD RECOGNIZE THE POTENTIAL RISK FOR STAPLES TO SCRATCH ADJACENT ORGANS, AND THE SHARP EDGE OF THE STAPLE LINE OR UNSTABLE STAPLES SHOULD BE TRIMMED DURING VATS TO AVOID STAPLE-ASSOCIATED HAEMOTHORAX.
02-01-2020 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON ENDOPATH STAPLER | Failure to Form Staple | A050704 | Tachycardia | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : INCIDENTAL HAEMOTHORAX AFTER SUBLOBAR RESECTION: DID STAPLE LINE SCRATCH CHEST WALL? AUTHOR : HIDEKI NEGISHI, HIROYOSHI TSUBOCHI, MITSURU MAKI, AND SHUNSUKE ENDO; CITATION: JOURNAL OF SURGICAL CASE REPORTS (2019);12, 1¿2. DOI: 10.1093/JSCR/RJZ276. THE AUTHORS REPORTED A CASE OF A LATE-ONSET HAEMOTHORAX CAUSED BY A STAPLE THAT DIVIDED LUNG PARENCHYMA DURING PULMONARY SEGMENTECTOMY. A (B)(6) MALE PATIENT UNDERWENT RIGHT APICAL AND POSTERIOR SEGMENTECTOMIES VIA VIDEO-ASSISTED THORACIC SURGERY (VATS) FOR EARLY-STAGE LUNG CANCER. THE BRONCHUS AND PULMONARY PARENCHYMA WERE DIVIDED BY ENDOSCOPIC STAPLES (ECHELON®; ETHICON). ON 32ND DAY AFTER DISCHARGE, THE PATIENT WAS TRANSPORTED BY AMBULANCE OWING TO A SUDDEN ONSET OF CHEST PAIN AND DYSPNEA. THE PATIENT DEVELOPED HYPOTENSION AND TACHYCARDIA WITH SYSTOLIC BLOOD PRESSURE OF 50MMHG AND HEART RATE OF 114 BEATS/MIN. THE PATIENT WAS DIAGNOSED WITH HAEMOTHORAX AND AN EMERGENCY OPERATION WAS PERFORMED. A MASSIVE BLOOD COAGULATION IN THE RIGHT THORACIC CAVITY WAS REMOVED. A CONTINUOUS ACTIVE BLEEDING FROM THE INTERCOSTAL ARTERY WAS FOUND NEAR THE FIFTH THORACIC VERTEBRA IN THE VICINITY OF A STUMP OF THE BRONCHUS DIVIDED BY A STAPLE. A THORACOTOMY THROUGH THE FIFTH INTERCOSTAL SPACE WAS THEREFORE PERFORMED TO ACHIEVE HAEMOSTASIS, AND THE BLEEDING WAS STOPPED BY SUTURING. THE BLEEDING POINT WAS COVERED WITH A COLLAGEN-FIBRIN PATCH. THE PATIENT PROGRESSED SATISFACTORILY AFTER THE REOPERATION AND WAS DISCHARGED ON NINTH DAY AFTER REOPERATION. THE POSTERIOR PART OF THE STAPLE LINE OF THE LUNG PARENCHYMA ADJACENT TO THE BLEEDING POINT UNDER FULL INFLATION WAS PARTLY EXPOSED, WHEREAS THE BRONCHIAL STUMP WAS COVERED WITH LUNG. THEREFORE, THE SCRATCHING CREATED BY THE STAPLES WAS ASSUMED TO CAUSE HAEMOTHORAX. IN CONCLUSION, SURGEONS SHOULD RECOGNIZE THE POTENTIAL RISK FOR STAPLES TO SCRATCH ADJACENT ORGANS, AND THE SHARP EDGE OF THE STAPLE LINE OR UNSTABLE STAPLES SHOULD BE TRIMMED DURING VATS TO AVOID STAPLE-ASSOCIATED HAEMOTHORAX.
02-01-2020 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON ENDOPATH STAPLER | Failure to Form Staple | A050704 | Blood Loss | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : INCIDENTAL HAEMOTHORAX AFTER SUBLOBAR RESECTION: DID STAPLE LINE SCRATCH CHEST WALL? AUTHOR : HIDEKI NEGISHI, HIROYOSHI TSUBOCHI, MITSURU MAKI, AND SHUNSUKE ENDO; CITATION: JOURNAL OF SURGICAL CASE REPORTS (2019);12, 1¿2. DOI: 10.1093/JSCR/RJZ276. THE AUTHORS REPORTED A CASE OF A LATE-ONSET HAEMOTHORAX CAUSED BY A STAPLE THAT DIVIDED LUNG PARENCHYMA DURING PULMONARY SEGMENTECTOMY. A (B)(6) MALE PATIENT UNDERWENT RIGHT APICAL AND POSTERIOR SEGMENTECTOMIES VIA VIDEO-ASSISTED THORACIC SURGERY (VATS) FOR EARLY-STAGE LUNG CANCER. THE BRONCHUS AND PULMONARY PARENCHYMA WERE DIVIDED BY ENDOSCOPIC STAPLES (ECHELON®; ETHICON). ON 32ND DAY AFTER DISCHARGE, THE PATIENT WAS TRANSPORTED BY AMBULANCE OWING TO A SUDDEN ONSET OF CHEST PAIN AND DYSPNEA. THE PATIENT DEVELOPED HYPOTENSION AND TACHYCARDIA WITH SYSTOLIC BLOOD PRESSURE OF 50MMHG AND HEART RATE OF 114 BEATS/MIN. THE PATIENT WAS DIAGNOSED WITH HAEMOTHORAX AND AN EMERGENCY OPERATION WAS PERFORMED. A MASSIVE BLOOD COAGULATION IN THE RIGHT THORACIC CAVITY WAS REMOVED. A CONTINUOUS ACTIVE BLEEDING FROM THE INTERCOSTAL ARTERY WAS FOUND NEAR THE FIFTH THORACIC VERTEBRA IN THE VICINITY OF A STUMP OF THE BRONCHUS DIVIDED BY A STAPLE. A THORACOTOMY THROUGH THE FIFTH INTERCOSTAL SPACE WAS THEREFORE PERFORMED TO ACHIEVE HAEMOSTASIS, AND THE BLEEDING WAS STOPPED BY SUTURING. THE BLEEDING POINT WAS COVERED WITH A COLLAGEN-FIBRIN PATCH. THE PATIENT PROGRESSED SATISFACTORILY AFTER THE REOPERATION AND WAS DISCHARGED ON NINTH DAY AFTER REOPERATION. THE POSTERIOR PART OF THE STAPLE LINE OF THE LUNG PARENCHYMA ADJACENT TO THE BLEEDING POINT UNDER FULL INFLATION WAS PARTLY EXPOSED, WHEREAS THE BRONCHIAL STUMP WAS COVERED WITH LUNG. THEREFORE, THE SCRATCHING CREATED BY THE STAPLES WAS ASSUMED TO CAUSE HAEMOTHORAX. IN CONCLUSION, SURGEONS SHOULD RECOGNIZE THE POTENTIAL RISK FOR STAPLES TO SCRATCH ADJACENT ORGANS, AND THE SHARP EDGE OF THE STAPLE LINE OR UNSTABLE STAPLES SHOULD BE TRIMMED DURING VATS TO AVOID STAPLE-ASSOCIATED HAEMOTHORAX.
02-01-2020 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON ENDOPATH STAPLER | Failure to Form Staple | A050704 | Not Applicable | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : INCIDENTAL HAEMOTHORAX AFTER SUBLOBAR RESECTION: DID STAPLE LINE SCRATCH CHEST WALL? AUTHOR : HIDEKI NEGISHI, HIROYOSHI TSUBOCHI, MITSURU MAKI, AND SHUNSUKE ENDO; CITATION: JOURNAL OF SURGICAL CASE REPORTS (2019);12, 1¿2. DOI: 10.1093/JSCR/RJZ276. THE AUTHORS REPORTED A CASE OF A LATE-ONSET HAEMOTHORAX CAUSED BY A STAPLE THAT DIVIDED LUNG PARENCHYMA DURING PULMONARY SEGMENTECTOMY. A (B)(6) MALE PATIENT UNDERWENT RIGHT APICAL AND POSTERIOR SEGMENTECTOMIES VIA VIDEO-ASSISTED THORACIC SURGERY (VATS) FOR EARLY-STAGE LUNG CANCER. THE BRONCHUS AND PULMONARY PARENCHYMA WERE DIVIDED BY ENDOSCOPIC STAPLES (ECHELON®; ETHICON). ON 32ND DAY AFTER DISCHARGE, THE PATIENT WAS TRANSPORTED BY AMBULANCE OWING TO A SUDDEN ONSET OF CHEST PAIN AND DYSPNEA. THE PATIENT DEVELOPED HYPOTENSION AND TACHYCARDIA WITH SYSTOLIC BLOOD PRESSURE OF 50MMHG AND HEART RATE OF 114 BEATS/MIN. THE PATIENT WAS DIAGNOSED WITH HAEMOTHORAX AND AN EMERGENCY OPERATION WAS PERFORMED. A MASSIVE BLOOD COAGULATION IN THE RIGHT THORACIC CAVITY WAS REMOVED. A CONTINUOUS ACTIVE BLEEDING FROM THE INTERCOSTAL ARTERY WAS FOUND NEAR THE FIFTH THORACIC VERTEBRA IN THE VICINITY OF A STUMP OF THE BRONCHUS DIVIDED BY A STAPLE. A THORACOTOMY THROUGH THE FIFTH INTERCOSTAL SPACE WAS THEREFORE PERFORMED TO ACHIEVE HAEMOSTASIS, AND THE BLEEDING WAS STOPPED BY SUTURING. THE BLEEDING POINT WAS COVERED WITH A COLLAGEN-FIBRIN PATCH. THE PATIENT PROGRESSED SATISFACTORILY AFTER THE REOPERATION AND WAS DISCHARGED ON NINTH DAY AFTER REOPERATION. THE POSTERIOR PART OF THE STAPLE LINE OF THE LUNG PARENCHYMA ADJACENT TO THE BLEEDING POINT UNDER FULL INFLATION WAS PARTLY EXPOSED, WHEREAS THE BRONCHIAL STUMP WAS COVERED WITH LUNG. THEREFORE, THE SCRATCHING CREATED BY THE STAPLES WAS ASSUMED TO CAUSE HAEMOTHORAX. IN CONCLUSION, SURGEONS SHOULD RECOGNIZE THE POTENTIAL RISK FOR STAPLES TO SCRATCH ADJACENT ORGANS, AND THE SHARP EDGE OF THE STAPLE LINE OR UNSTABLE STAPLES SHOULD BE TRIMMED DURING VATS TO AVOID STAPLE-ASSOCIATED HAEMOTHORAX.
02-01-2020 | Injury | ethicon endo-surgery | 27-01-2020 | ECHELON ENDOPATH STAPLER | Failure to Form Staple | A050704 | No Code Available | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE TITLE : INCIDENTAL HAEMOTHORAX AFTER SUBLOBAR RESECTION: DID STAPLE LINE SCRATCH CHEST WALL? AUTHOR : HIDEKI NEGISHI, HIROYOSHI TSUBOCHI, MITSURU MAKI, AND SHUNSUKE ENDO; CITATION: JOURNAL OF SURGICAL CASE REPORTS (2019);12, 1¿2. DOI: 10.1093/JSCR/RJZ276. THE AUTHORS REPORTED A CASE OF A LATE-ONSET HAEMOTHORAX CAUSED BY A STAPLE THAT DIVIDED LUNG PARENCHYMA DURING PULMONARY SEGMENTECTOMY. A (B)(6) MALE PATIENT UNDERWENT RIGHT APICAL AND POSTERIOR SEGMENTECTOMIES VIA VIDEO-ASSISTED THORACIC SURGERY (VATS) FOR EARLY-STAGE LUNG CANCER. THE BRONCHUS AND PULMONARY PARENCHYMA WERE DIVIDED BY ENDOSCOPIC STAPLES (ECHELON®; ETHICON). ON 32ND DAY AFTER DISCHARGE, THE PATIENT WAS TRANSPORTED BY AMBULANCE OWING TO A SUDDEN ONSET OF CHEST PAIN AND DYSPNEA. THE PATIENT DEVELOPED HYPOTENSION AND TACHYCARDIA WITH SYSTOLIC BLOOD PRESSURE OF 50MMHG AND HEART RATE OF 114 BEATS/MIN. THE PATIENT WAS DIAGNOSED WITH HAEMOTHORAX AND AN EMERGENCY OPERATION WAS PERFORMED. A MASSIVE BLOOD COAGULATION IN THE RIGHT THORACIC CAVITY WAS REMOVED. A CONTINUOUS ACTIVE BLEEDING FROM THE INTERCOSTAL ARTERY WAS FOUND NEAR THE FIFTH THORACIC VERTEBRA IN THE VICINITY OF A STUMP OF THE BRONCHUS DIVIDED BY A STAPLE. A THORACOTOMY THROUGH THE FIFTH INTERCOSTAL SPACE WAS THEREFORE PERFORMED TO ACHIEVE HAEMOSTASIS, AND THE BLEEDING WAS STOPPED BY SUTURING. THE BLEEDING POINT WAS COVERED WITH A COLLAGEN-FIBRIN PATCH. THE PATIENT PROGRESSED SATISFACTORILY AFTER THE REOPERATION AND WAS DISCHARGED ON NINTH DAY AFTER REOPERATION. THE POSTERIOR PART OF THE STAPLE LINE OF THE LUNG PARENCHYMA ADJACENT TO THE BLEEDING POINT UNDER FULL INFLATION WAS PARTLY EXPOSED, WHEREAS THE BRONCHIAL STUMP WAS COVERED WITH LUNG. THEREFORE, THE SCRATCHING CREATED BY THE STAPLES WAS ASSUMED TO CAUSE HAEMOTHORAX. IN CONCLUSION, SURGEONS SHOULD RECOGNIZE THE POTENTIAL RISK FOR STAPLES TO SCRATCH ADJACENT ORGANS, AND THE SHARP EDGE OF THE STAPLE LINE OR UNSTABLE STAPLES SHOULD BE TRIMMED DURING VATS TO AVOID STAPLE-ASSOCIATED HAEMOTHORAX.
11-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, DURING FIRST GASTRIC TRANSECTION, SEVERAL OF THE PROXIMAL STAPLES WERE PUSHED OUT OF THE DEVICE BUT DID NOT HOLD THE TISSUE AND SOME DID NOT FORM B SHAPE. THEREFORE, THERE WAS INCOMPLETE STAPLE LINE AT THE PROXIMAL END. SURGEON USED THE SAME HANDLE WITH A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, DURING FIRST GASTRIC TRANSECTION, SEVERAL OF THE PROXIMAL STAPLES WERE PUSHED OUT OF THE DEVICE BUT DID NOT HOLD THE TISSUE AND SOME DID NOT FORM B SHAPE. THEREFORE, THERE WAS INCOMPLETE STAPLE LINE AT THE PROXIMAL END. SURGEON USED THE SAME HANDLE WITH A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | IDRIVE ULTRA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING THE STOMACH DURING A LAPAROSCOPIC GASTRIC BYPASS, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD AND WAS BALE TO FIRE BUT HAD AN UNLOADING PROBLEM. IT WAS ALSO STATED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE ADAPTER HAS BEEN DISCONNECTED AND RECONNECTED IN ORDER TO OPEN THE JAWS AND RETRIEVE THE DEVICE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | UNKNOWN EGIA ADAPTER | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING THE STOMACH DURING A LAPAROSCOPIC GASTRIC BYPASS, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD AND WAS BALE TO FIRE BUT HAD AN UNLOADING PROBLEM. IT WAS ALSO STATED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE ADAPTER HAS BEEN DISCONNECTED AND RECONNECTED IN ORDER TO OPEN THE JAWS AND RETRIEVE THE DEVICE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN STAPLING THE STOMACH DURING A LAPAROSCOPIC GASTRIC BYPASS, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD AND WAS BALE TO FIRE BUT HAD AN UNLOADING PROBLEM. IT WAS ALSO STATED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE ADAPTER HAS BEEN DISCONNECTED AND RECONNECTED IN ORDER TO OPEN THE JAWS AND RETRIEVE THE DEVICE. THERE WAS NO PATIENT INJURY.
07-01-2020 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING THORACIC LOBECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE LOCKED AND COULD NOT SQUEEZE THE HANDLE. THE SURGEON TRIED TO USE ANOTHER HANDLE, HOWEVER, SAME ISSUE OCCURRED. THE SURGEON THEN REPLACES AGAIN THE HANDLE AND WORKED WITH THE ORIGINAL LOADS THAT DID NOT WORK ON THE FIRST TWO HANDLE. THERE WAS NO PATIENT INJURY.
23-12-2019 | Malfunction | us surgical puerto rico | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN ATTEMPTING TO FIRE INTRA-OPERATIVELY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE COULD NOT BE SQUEEZED. THE RELOAD WAS REPLACED TO CONTINUE TO CASE. THE EVENT OCCURRED IN USE FOR PATIENT. THERE WAS NO PATIENT INJURY.
14-01-2020 | Malfunction | covidien lp llc north haven | 27-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON SAID THE RELOAD LOCKED UP AND IT DID NOT WORK PROPERLY. THEY REMOVED THE STAPLER WITH THE RELOAD FROM THE TROCAR AND REPLACED THE RELOAD. THERE WAS NO PATIENT INJURY.
09-01-2020 | Malfunction | covidien lp llc north haven | 28-01-2020 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING STAPLING OF THE STOMACH IN A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE SURGEON ATTEMPTED TO FIRE THE RELOAD AFTER CLAMPING DOWN ON TISSUE BUT THE STAPLER WOULDN'T FIRE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-04-2020 | Malfunction | us surgical puerto rico | 28-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE SURGEON WAS ABLE TO FULLY SQUEEZE THE HANDLE HOWEVER, THE STAPLER DID NOT FIRE. THERE WAS NO PATIENT INJURY.
14-01-2020 | Malfunction | us surgical puerto rico | 28-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN DOING THE FIRST HEPATIC PORTAL FOR RIGHT PORTAL VEIN DISSECTION DURING A LAPAROSCOPIC RIGHT HEPATECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS STATED THAT TWO INSTRUMENTS HAD THE SAME MALFUNCTION DURING THE SAME EVENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | ENDO GIA | Difficult or Delayed Positioning | A150203 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY END TO END ANASTOMOSIS, AN OPEN ILEOCECAL RESECTION WAS PERFORMED. THE RECONSTRUCTION WAS FOUR FIRINGS FUNCTIONAL PROCEDURE. AT THE FOURTH FIRING, WHEN CLOSING THE INSERTION SLOT, THE THREE ROWS OF STAPLES WHICH WAS ONLY NEAR THE CENTER OF THE STAPLE LINE DID NOT CATCH AND WERE IN AN OPEN STATE. ONLY NEAR THE CENTER WAS IT FELT HARD TO FIRE. NO TRACE OF STAPLE CAME OUT. THE PROCEDURE WAS COMPLETED WITH MANUAL SUTURING. THE SURGICAL TIME WAS EXTENDED BY LESS THAN THIRTY MINUTES.
13-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | ENDO GIA | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY END TO END ANASTOMOSIS, AN OPEN ILEOCECAL RESECTION WAS PERFORMED. THE RECONSTRUCTION WAS FOUR FIRINGS FUNCTIONAL PROCEDURE. AT THE FOURTH FIRING, WHEN CLOSING THE INSERTION SLOT, THE THREE ROWS OF STAPLES WHICH WAS ONLY NEAR THE CENTER OF THE STAPLE LINE DID NOT CATCH AND WERE IN AN OPEN STATE. ONLY NEAR THE CENTER WAS IT FELT HARD TO FIRE. NO TRACE OF STAPLE CAME OUT. THE PROCEDURE WAS COMPLETED WITH MANUAL SUTURING. THE SURGICAL TIME WAS EXTENDED BY LESS THAN THIRTY MINUTES.
13-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY END TO END ANASTOMOSIS, AN OPEN ILEOCECAL RESECTION WAS PERFORMED. THE RECONSTRUCTION WAS FOUR FIRINGS FUNCTIONAL PROCEDURE. AT THE FOURTH FIRING, WHEN CLOSING THE INSERTION SLOT, THE THREE ROWS OF STAPLES WHICH WAS ONLY NEAR THE CENTER OF THE STAPLE LINE DID NOT CATCH AND WERE IN AN OPEN STATE. ONLY NEAR THE CENTER WAS IT FELT HARD TO FIRE. NO TRACE OF STAPLE CAME OUT. THE PROCEDURE WAS COMPLETED WITH MANUAL SUTURING. THE SURGICAL TIME WAS EXTENDED BY LESS THAN THIRTY MINUTES.
13-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Difficult or Delayed Positioning | A150203 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY END TO END ANASTOMOSIS, AN OPEN ILEOCECAL RESECTION WAS PERFORMED. THE RECONSTRUCTION WAS FOUR FIRINGS FUNCTIONAL PROCEDURE. AT THE FOURTH FIRING, WHEN CLOSING THE INSERTION SLOT, THE THREE ROWS OF STAPLES WHICH WAS ONLY NEAR THE CENTER OF THE STAPLE LINE DID NOT CATCH AND WERE IN AN OPEN STATE. ONLY NEAR THE CENTER WAS IT FELT HARD TO FIRE. NO TRACE OF STAPLE CAME OUT. THE PROCEDURE WAS COMPLETED WITH MANUAL SUTURING. THE SURGICAL TIME WAS EXTENDED BY LESS THAN THIRTY MINUTES.
13-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY END TO END ANASTOMOSIS, AN OPEN ILEOCECAL RESECTION WAS PERFORMED. THE RECONSTRUCTION WAS FOUR FIRINGS FUNCTIONAL PROCEDURE. AT THE FOURTH FIRING, WHEN CLOSING THE INSERTION SLOT, THE THREE ROWS OF STAPLES WHICH WAS ONLY NEAR THE CENTER OF THE STAPLE LINE DID NOT CATCH AND WERE IN AN OPEN STATE. ONLY NEAR THE CENTER WAS IT FELT HARD TO FIRE. NO TRACE OF STAPLE CAME OUT. THE PROCEDURE WAS COMPLETED WITH MANUAL SUTURING. THE SURGICAL TIME WAS EXTENDED BY LESS THAN THIRTY MINUTES.
14-01-2020 | Malfunction | covidien lp llc north haven | 28-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN DOING THE FIRST HEPATIC PORTAL FOR RIGHT PORTAL VEIN DISSECTION DURING A LAPAROSCOPIC RIGHT HEPATECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS STATED THAT TWO INSTRUMENTS HAD THE SAME MALFUNCTION DURING THE SAME EVENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-03-2020 | Malfunction | us surgical puerto rico | 28-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING OPEN HYPERTHERMIC INTRAPERITONEAL CHEMOTHERAPY PROCEDURE, UPON TRANSACTION OF BOWEL, THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AND THE DEVICE DID NOT FIRE. A NEW DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, THE STAPLER WAS FIRED AND STAPLES WERE RELEASED FROM THE INSTRUMENT BUT B SHAPE FORMATION WAS NOT ACHIEVED AND IT DID NOT CLAMP AND FIX INTO THE TISSUE. A NEW GREEN LOAD WAS USED AND THE SAME THING HAPPENED. THE STAPLE LINE WAS RESECTED AND A NEW STAPLER WAS OPENED WHICH WORKED FINE TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THIS PROBLEM.
09-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, THE STAPLER WAS FIRED AND STAPLES WERE RELEASED FROM THE INSTRUMENT BUT B SHAPE FORMATION WAS NOT ACHIEVED AND IT DID NOT CLAMP AND FIX INTO THE TISSUE. A NEW GREEN LOAD WAS USED AND THE SAME THING HAPPENED. THE STAPLE LINE WAS RESECTED AND A NEW STAPLER WAS OPENED WHICH WORKED FINE TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THIS PROBLEM.
14-01-2020 | Malfunction | covidien lp llc north haven | 28-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO DETACH THE VEIN IN THE MIDDLE LOBE OF THE RIGHT LUNG DURING LAPAROSCOPIC RADICAL RESECTION OF PULMONARY CARCINOMA, THE DEVICE WAS ABLE TO FIRE, HOWEVER, POOR STAPLE FORMATION WAS OBSERVED, AND PART OF NO STAPLE WAS SHAPED. IT WAS ALSO NOTED THAT THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE SURGEON USED A CLIP TO STOP THE BLEEDING AND THEN A NEW RELOAD TO TREAT THE BLOOD VESSELS, AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
14-01-2020 | Malfunction | covidien lp llc north haven | 28-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO DETACH THE VEIN IN THE MIDDLE LOBE OF THE RIGHT LUNG DURING LAPAROSCOPIC RADICAL RESECTION OF PULMONARY CARCINOMA, THE DEVICE WAS ABLE TO FIRE, HOWEVER, POOR STAPLE FORMATION WAS OBSERVED, AND PART OF NO STAPLE WAS SHAPED. IT WAS ALSO NOTED THAT THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE SURGEON USED A CLIP TO STOP THE BLEEDING AND THEN A NEW RELOAD TO TREAT THE BLOOD VESSELS, AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
09-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | GIA | Failure to Form Staple | A050704 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, THE STAPLER WAS FIRED AND STAPLES WERE RELEASED FROM THE INSTRUMENT BUT B SHAPE FORMATION WAS NOT ACHIEVED AND IT DID NOT CLAMP AND FIX INTO THE TISSUE. A NEW GREEN LOAD WAS USED AND THE SAME THING HAPPENED. THE STAPLE LINE WAS RESECTED AND A NEW STAPLER WAS OPENED WHICH WORKED FINE TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THIS PROBLEM.
09-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | GIA | Failure to Form Staple | A050704 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, THE STAPLER WAS FIRED AND STAPLES WERE RELEASED FROM THE INSTRUMENT BUT B SHAPE FORMATION WAS NOT ACHIEVED AND IT DID NOT CLAMP AND FIX INTO THE TISSUE. A NEW GREEN LOAD WAS USED AND THE SAME THING HAPPENED. THE STAPLE LINE WAS RESECTED AND A NEW STAPLER WAS OPENED WHICH WORKED FINE TO COMPLETE THE CASE. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT OF THIS PROBLEM.
13-01-2020 | Injury | covidien lp llc north haven | 28-01-2020 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING OPEN LIVER RESECTION, THE DEVICE LOCKED ON TISSUE EVEN AFTER TRYING TO PULL BACK BLACK KNOBS AND R ETRACTING THE SLIDE. THE DEVICE WOULD ALSO NOT UNLOAD. THIS RESULTED TO TISSUE LOSS AND DAMAGE SINCE SURGEON HAD TO CUT OUT RELOAD THAT WAS LOCKED ON TISSUE. TWO VASCULAR CLAMPS WERE PLACED ON EACH SIDE OF THE STAPLER AND THEN USED A SCALPEL TO CUT IN BETWEEN THE STAPLER AND THE CLAMP. THERE WAS LITTLE BLEEDING WHICH WAS MANAGED WITH SUTURE.
13-01-2020 | Injury | covidien lp llc north haven | 28-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING OPEN LIVER RESECTION, THE DEVICE LOCKED ON TISSUE EVEN AFTER TRYING TO PULL BACK BLACK KNOBS AND RETRACTING THE SLIDE. THE DEVICE WOULD ALSO NOT UNLOAD. THIS RESULTED TO TISSUE LOSS AND DAMAGE SINCE SURGEON HAD TO CUT OUT RELOAD THAT WAS LOCKED ON TISSUE. TWO VASCULAR CLAMPS WERE PLACED ON EACH SIDE OF THE STAPLER AND THEN USED A SCALPEL TO CUT IN BETWEEN THE STAPLER AND THE CLAMP. THERE WAS LITTLE BLEEDING WHICH WAS MANAGED WITH SUTURE.
13-01-2020 | Malfunction | covidien lp llc north haven | 28-01-2020 | ENDO GIA | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING RESECTION OF THE COLON ON AN OPEN RIGHT COLECTOMY, THE STAPLER WAS ABLE TO FIRE HOWEVER SOME STAPLES DID NOT FORM. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-01-2020 | Malfunction | covidien lp llc north haven | 28-01-2020 | ENDO GIA | Misfire | A050502 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, THE DISTAL STAPLE LINE WAS INCOMPLETE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-01-2020 | Malfunction | covidien lp llc north haven | 28-01-2020 | SIGNIA | Complete Loss of Power | A070801 | No Patient Involvement | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PARTIAL ESOPHAGECTOMY SURGERY, THE NURSE REMOVED THE POWER HANDLE FROM THE CHARGER. THE HANDLE WAS SHAKEN, THE POWER TURNED ON AND THE DISPLAY OF THE COMPANY NAME APPEARED. HOWEVER AFTER THAT, THE POWER SHUT DOWN AND THERE WAS NO RESPONSE. THERE WAS NO RESPONSE EVEN THOUGH THE GREEN BUTTON WAS PRESSED. THEN THE HANDLE WAS RETURNED BACK TO THE CHARGER FOR ABOUT 30 MINUTES AND THEN REMOVED THE HANDLE BUT THE SCREEN REMAINED DARK. THE SHELL WAS ATTACHED AND CONNECTED THE ADAPTER, BUT THERE WAS NO CALIBRATION TONE, AND THERE WAS NO RESPONSE. IN ADDITION, THERE WAS NO TRACE OF BEING DROPPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
14-01-2020 | Malfunction | us surgical puerto rico | 28-01-2020 | EEA | Break | A0401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN CREATING THE POUCH ON A LAPAROSCOPIC ROUX EN Y PROCEDURE, THE TIP OF THE DEVICE BROKE AND FELL INTO THE PATIENT¿S CAVITY. IT WAS STATED THAT THE PIECE WAS RETRIEVED BY IRRIGATION. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
14-01-2020 | Malfunction | us surgical puerto rico | 28-01-2020 | EEA | Break | A0401 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN CREATING THE POUCH ON A LAPAROSCOPIC ROUX EN Y PROCEDURE, THE TIP OF THE DEVICE BROKE AND FELL INTO THE PATIENT¿S CAVITY. IT WAS STATED THAT THE PIECE WAS RETRIEVED BY IRRIGATION. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
14-01-2020 | Malfunction | us surgical puerto rico | 28-01-2020 | EEA | Detachment of Device or Device Component | A0501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, WHEN CREATING THE POUCH ON A LAPAROSCOPIC ROUX EN Y PROCEDURE, THE TIP OF THE DEVICE BROKE AND FELL INTO THE PATIENT¿S CAVITY. IT WAS STATED THAT THE PIECE WAS RETRIEVED BY IRRIGATION. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
14-01-2020 | Malfunction | us surgical puerto rico | 28-01-2020 | EEA | Detachment of Device or Device Component | A0501 | No Code Available | Event Description: ACCORDING TO THE REPORTER, WHEN CREATING THE POUCH ON A LAPAROSCOPIC ROUX EN Y PROCEDURE, THE TIP OF THE DEVICE BROKE AND FELL INTO THE PATIENT¿S CAVITY. IT WAS STATED THAT THE PIECE WAS RETRIEVED BY IRRIGATION. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
10-01-2020 | Injury | covidien lp llc north haven | 28-01-2020 | ENDO GIA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING RESECTION OF LESIONS IN A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, THE KNIFE ADVANCED TO THE TIP AND RESECTION WAS COMPLETED, BUT THE KNIFE COULD NOT BE PULLED BACK. NO MATTER WHAT WAS DONE, THE JAWS COULD NOT BE RELEASED AND ADDITIONAL RESECTION WAS PERFORMED TO COMPLETE THE SEPARATION. NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE.
10-01-2020 | Injury | covidien lp llc north haven | 28-01-2020 | ENDO GIA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING RESECTION OF LESIONS IN A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, THE KNIFE ADVANCED TO THE TIP AND RESECTION WAS COMPLETED, BUT THE KNIFE COULD NOT BE PULLED BACK. NO MATTER WHAT WAS DONE, THE JAWS COULD NOT BE RELEASED AND ADDITIONAL RESECTION WAS PERFORMED TO COMPLETE THE SEPARATION. NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE.
10-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Damage | Event Description: ACCORDING TO THE REPORTER, DURING RESECTION OF LESIONS IN A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, THE KNIFE ADVANCED TO THE TIP AND RESECTION WAS COMPLETED, BUT THE KNIFE COULD NOT BE PULLED BACK. NO MATTER WHAT WAS DONE, THE JAWS COULD NOT BE RELEASED AND ADDITIONAL RESECTION WAS PERFORMED TO COMPLETE THE SEPARATION. NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE.
10-01-2020 | Injury | us surgical puerto rico | 28-01-2020 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Code Available | Event Description: ACCORDING TO THE REPORTER, DURING RESECTION OF LESIONS IN A COLECTOMY FUNCTIONAL END TO END ANASTOMOSIS PROCEDURE, THE KNIFE ADVANCED TO THE TIP AND RESECTION WAS COMPLETED, BUT THE KNIFE COULD NOT BE PULLED BACK. NO MATTER WHAT WAS DONE, THE JAWS COULD NOT BE RELEASED AND ADDITIONAL RESECTION WAS PERFORMED TO COMPLETE THE SEPARATION. NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE.
14-01-2020 | Malfunction | us surgical puerto rico | 28-01-2020 | PREMIUM PLUS CEEA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF GASTRIC PYLORUS AND ILEUM IN A PANCREATICODUODENECTOMY PROCEDURE, THE GREEN BAR WAS VISIBLE IN THE INDICATOR WINDOW OF CIRCULAR STAPLER BUT DID NOT CUT AND FIRE. ANOTHER CIRCULAR STAPLER WAS USED TO COMPLETE THE CASE. WHEN STOMACH AND PYLORUS WERE DISCONNECTED, OPEN STAPLER WAS ALSO USED BUT BLACK PUSHER THUMB PIECE COULD NOT BE MOVE AND IT DID NOT FIRE EITHER. NEW OPEN STAPLER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-01-2020 | Malfunction | us surgical puerto rico | 28-01-2020 | GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF GASTRIC PYLORUS AND ILEUM IN A PANCREATICODUODENECTOMY PROCEDURE, THE GREEN BAR WAS VISIBLE IN THE INDICATOR WINDOW OF CIRCULAR STAPLER BUT DID NOT CUT AND FIRE. ANOTHER CIRCULAR STAPLER WAS USED TO COMPLETE THE CASE. WHEN STOMACH AND PYLORUS WERE DISCONNECTED, OPEN STAPLER WAS ALSO USED BUT BLACK PUSHER THUMB PIECE COULD NOT BE MOVE AND IT DID NOT FIRE EITHER. NEW OPEN STAPLER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-09-2019 | Malfunction | covidien lp llc north haven | 28-01-2020 | ENDO GIA | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. STAPLES AT THE PROXIMAL END WERE FOUND TO BE SLIGHTLY FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-09-2019 | Malfunction | covidien lp llc north haven | 28-01-2020 | ENDO GIA | Expulsion | A010401 | No Consequences Or Impact To Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. STAPLES AT THE PROXIMAL END WERE FOUND TO BE SLIGHTLY FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOWER ANTERIOR RESECTION, THE JAWS COMPLETELY CLOSED (GREEN BUTTON WAS ILLUMINATING), BUT THE SURGEON NOTICED THAT THE JAWS DID NOT CLOSE FIRMLY. NEW RELOAD WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-11-2020 | Malfunction | ethicon | 04-01-2021 | ETHICON SECURESTRAP | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC HERNIORRHAPHY ON (B)(6) 2020 AND THE ABSORBABLE STRAPS WERE USED. IT WAS REPORTED THAT THE CLAMP FIXING DEVICE DID NOT TRIGGER THE 12 RELOADS, IT STUCK DURING HANDLING AND THE STAPLES USED DID NOT KEEP FIXED AS THEY SHOULD. A LIKE DEVICE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED.
25-11-2020 | Injury | us surgical puerto rico | 04-01-2021 | EEA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RECTAL RESECTION, AFTER ANASTOMOSIS, THE RING OF THE ANAL SIDE WAS NOT FOUND. LEAK TEST REVEALED LEAKAGE FROM THE RIGHT HALF SIDE. ANASTOMOTIC INSUFFICIENCY WAS FOUND, SO SUTURE WAS PERFORMED ADDITIONALLY AND STOMA CREATION WAS PERFORMED DUE TO THE DEVICE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
25-11-2020 | Injury | us surgical puerto rico | 04-01-2021 | EEA | Misfire | A050502 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RECTAL RESECTION, AFTER ANASTOMOSIS, THE RING OF THE ANAL SIDE WAS NOT FOUND. LEAK TEST REVEALED LEAKAGE FROM THE RIGHT HALF SIDE. ANASTOMOTIC INSUFFICIENCY WAS FOUND, SO SUTURE WAS PERFORMED ADDITIONALLY AND STOMA CREATION WAS PERFORMED DUE TO THE DEVICE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
25-11-2020 | Injury | us surgical puerto rico | 04-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RECTAL RESECTION, AFTER ANASTOMOSIS, THE RING OF THE ANAL SIDE WAS NOT FOUND. LEAK TEST REVEALED LEAKAGE FROM THE RIGHT HALF SIDE. ANASTOMOTIC INSUFFICIENCY WAS FOUND, SO SUTURE WAS PERFORMED ADDITIONALLY AND STOMA CREATION WAS PERFORMED DUE TO THE DEVICE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
25-11-2020 | Injury | us surgical puerto rico | 04-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RECTAL RESECTION, AFTER ANASTOMOSIS, THE RING OF THE ANAL SIDE WAS NOT FOUND. LEAK TEST REVEALED LEAKAGE FROM THE RIGHT HALF SIDE. ANASTOMOTIC INSUFFICIENCY WAS FOUND, SO SUTURE WAS PERFORMED ADDITIONALLY AND STOMA CREATION WAS PERFORMED DUE TO THE DEVICE ISSUE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
04-12-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE SCRUB NURSE TRIED TO TEST THE DEVICE FOR SIGMOID COLON TRANSECTION DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE JAW OF THE DEVICE WAS NOT CLOSED PERFECTLY. THEY TRIED TO CLOSE AND OPEN SEVERAL TIMES, BUT THE JAW HAD THE SAME STATUS: IT WOULD NOT CLOSE PERFECTLY. A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
08-12-2020 | Injury | us surgical puerto rico | 04-01-2021 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY PROCEDURE, AFTER THE TISSUE WAS CLOSED , IT COULD NOT BE OPENED. THE SURGEON THEN EXPANDED THE TISSUE CUTTING AREA FOR MORE THAN 1 INCH, FIRED ANOTHER DEVICE AND REPLACED THE OLD PRODUCT. THE SURGICAL PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES. AND THERE WAS TISSUE DAMAGE, TISSUE TRAUMA AND TISSUE LOSS DUE TO THE ISSUE. IT WAS ALSO REPORTED THAT MOST STAPLES WERE NOT FORMED PROPERLY.
19-02-2019 | Malfunction | us surgical puerto rico | 04-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LUNG CANCER SURGERY, WHILE THE DEVICE WAS BEING APPLIED IN THE LUNG TISSUE, THE DEVICE COULD NOT FIRE NORMALLY. IT WAS STATED THAT THE TISSUE WAS TOO THICK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-02-2019 | Malfunction | us surgical puerto rico | 04-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LUNG CANCER SURGERY, WHILE THE DEVICE WAS BEING APPLIED IN THE LUNG TISSUE, THE DEVICE COULD NOT FIRE NORMALLY. IT WAS STATED THAT THE TISSUE WAS TOO THICK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Injury | us surgical puerto rico | 04-01-2021 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY PROCEDURE, AFTER THE TISSUE WAS CLOSED , IT COULD NOT BE OPENED. THE SURGEON THEN EXPANDED THE TISSUE CUTTING AREA FOR MORE THAN 1 INCH, FIRED ANOTHER DEVICE AND REPLACED THE OLD PRODUCT. THE SURGICAL PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES. AND THERE WAS TISSUE DAMAGE, TISSUE TRAUMA AND TISSUE LOSS DUE TO THE ISSUE. IT WAS ALSO REPORTED THAT MOST STAPLES WERE NOT FORMED PROPERLY.
08-12-2020 | Injury | us surgical puerto rico | 04-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY PROCEDURE, AFTER THE TISSUE WAS CLOSED , IT COULD NOT BE OPENED. THE SURGEON THEN EXPANDED THE TISSUE CUTTING AREA FOR MORE THAN 1 INCH, FIRED ANOTHER DEVICE AND REPLACED THE OLD PRODUCT. THE SURGICAL PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES. AND THERE WAS TISSUE DAMAGE, TISSUE TRAUMA AND TISSUE LOSS DUE TO THE ISSUE. IT WAS ALSO REPORTED THAT MOST STAPLES WERE NOT FORMED PROPERLY.
19-02-2019 | Malfunction | covidien lp llc north haven | 04-01-2021 | UNKNOWN EGIA TRI STAPLE | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LUNG CANCER SURGERY, WHILE THE DEVICE WAS BEING APPLIED IN THE LUNG TISSUE, THE DEVICE COULD NOT FIRE NORMALLY. IT WAS STATED THAT THE TISSUE WAS TOO THICK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-02-2019 | Malfunction | covidien lp llc north haven | 04-01-2021 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LUNG CANCER SURGERY, WHILE THE DEVICE WAS BEING APPLIED IN THE LUNG TISSUE, THE DEVICE COULD NOT FIRE NORMALLY. IT WAS STATED THAT THE TISSUE WAS TOO THICK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | covidien lp llc north haven | 04-01-2021 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, STAPLER FIRED AND THEN STOPPED FIRING. BOTH LOADING UNIT AND STAPLER WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | covidien lp llc north haven | 04-01-2021 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, STAPLER FIRED AND THEN STOPPED FIRING. BOTH LOADING UNIT AND STAPLER WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
01-12-2020 | Injury | us surgical puerto rico | 04-01-2021 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, ON A LUNG RESECTION DURING AN OPEN LUNG LOBECTOMY, THE PATIENT¿S COAGULATION PARAMETERS AND BLOOD WORK WERE NORMAL. THE PATIENT HAD A ROUTINE THORACOTOMY AND LUNG LOBECTOMY WITH MINIMAL HEMORRHAGE. A CHEST TUBE WAS PLACED TO REMOVE AIR/FLUID PRIOR TO CLOSING THE THORACIC CAVITY. OVER THE NEXT 12 HOURS, THEY CONTINUED TO REMOVE BLOOD FROM THE PATIENT¿S THORACOTOMY TUBE, INDICATING PERSISTENT HEMORRHAGE. THE PATIENT WAS GIVEN CRYOPRECIPITATE AND MEDIATION WHICH DID NOT SLOW DOWN THE HEMORRHAGE. ON DECEMBER 2, THE PATIENT WAS GIVEN ADDITIONAL VIALS OF CRYOPRECIPITATE, A TRANSFUSION OF PACKED RED CELLS AND MULTIPLE UNITS OF FRESH FROZEN PLASMA AND WAS TAKEN FOR A REPEAT EXPLORATORY THORACOTOMY. THE LUNG LOBECTOMY STUMP HAD A LARGE CLOT OVER THE STAPLE LINE, ALLUDING TO IT BEING THE SOURCE OF CONTINUED HEMORRHAGE AND A MODERATE AMOUNT FREE BLOOD IN THE CHEST CAVITY. IT WAS STATED THAT NO OTHER SOURCE OF HEMORRHAGE WAS IDENTIFIED. MULTIPLE LARGE CLIPS AND SEVERAL SMALL LIGATURES WERE PLACED OVER STAPLE LINE AND PRODUCT WAS PLACED OVER THE SITE. AFTER ADDITIONAL CLIPS WERE PLACED OVER THE SURGICAL SITE, NO FURTHER HEMORRHAGE WAS IDENTIFIED. THE PATIENT RECOVERED UNEVENTFULLY. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES OR MORE AND PATIENT¿S EXTENDED HOSPITALIZATION TO TWO DAYS.
08-12-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCTECTOMY, THE SURGEON WAS ABLE TO PRESS THE FIRING BUTTON, BUT THE FIRING STOPPED IMMEDIATELY AFTER STARTING THE RESECTION. ONLY A BIT OF STAPLE WAS FOUND TO BE EXPOSED AT THE PROXIMAL END. IN ADDITION, THE DEVICE'S JAWS LOCKED ON TISSUE. THE SURGEON THEN USED ANOTHER DEVICE RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCTECTOMY, THE SURGEON WAS ABLE TO PRESS THE FIRING BUTTON, BUT THE FIRING STOPPED IMMEDIATELY AFTER STARTING THE RESECTION. ONLY A BIT OF STAPLE WAS FOUND TO BE EXPOSED AT THE PROXIMAL END. IN ADDITION, THE DEVICE'S JAWS LOCKED ON TISSUE. THE SURGEON THEN USED ANOTHER DEVICE RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2020 | Malfunction | us surgical puerto rico | 04-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC PROCEDURE, WHILE IN THE COLON, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | celestica de monterrey s.a de c.v | 04-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED IN PULMONARY LOBE DURING LAPAROSCOPIC LOBECTOMY, THE DEVICE WAS ABLE TO FIRE; HOWEVER, THERE WAS A POOR STAPLE FORMATION. THE SAME EVENT OCCURRED ON ANOTHER RELOAD. THE SURGEON HAD TO REPLACE IT WITH ANOTHER RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC PROCEDURE, WHILE IN THE COLON, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-11-2020 | Malfunction | us surgical puerto rico | 04-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS UNABLE TO HOLD THE CLAMP IN A STRAIGHT POSITION . ORIENTATION IS ONLY LEFT OR RIGHT BUT NOT IN THE STRAIGHT POSITION. THE SURGEON COULD NOT STAPLE IN THE STRAIGHT POSITION. IT REMAINS IN THE ORIENTED POSITION EVEN WHEN THE ORIENTATION KNOB IS IN THE STRAIGHT POSITION. THEY TOOK A NEW DEVICE TO FINISH AND EVERYTHING WAS FINE. NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED IN PULMONARY LOBE DURING LAPAROSCOPIC LOBECTOMY, THE DEVICE WAS ABLE TO FIRE; HOWEVER, THERE WAS A POOR STAPLE FORMATION. THE SAME EVENT OCCURRED ON ANOTHER RELOAD. THE SURGEON HAD TO REPLACE IT WITH ANOTHER RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-11-2020 | Malfunction | us surgical puerto rico | 04-01-2021 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SIGMOID RESECTION, DURING THE TRANSECTION OF THE SIGMOID COLON, A PIECE OF EITHER THE STAPLER OR A RELOAD FELL OFF ON THE BACK TABLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | ECHELON FLEX | Mechanical Problem | A05 | Insufficient Information | Event Description: ECHELON FLEX POWER STAPLER ENGAGED AND FAILED TO RELEASE WHEN PHYSICIAN STAPLING BOWEL POST MASS REMOVAL. PHYSICIAN HAD TO USE THE EMERGENCY RELEASE TO REMOVE THE STAPLER FROM THE PATIENT'S BOWEL. STAPLER WAS REMOVED FROM STERILE FIELD AND PLACED IN BIOHAZARD BAG.
07-12-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE USING THE DEVICE ON THE RECTUM IN A RADICAL RESECTION OF COLORECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. SURGEON WAS NOT ABLE TO FIRE THE DEVICE. SURGEON USED ANOTHER STAPLE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
11-11-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | PROXIMATE*RL LINEAR STAPLER | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, OPENED STERILE TX60G TO USE STAPLER ON RECTUM IN ANTERIOR RESECTION PROCEDURE. SURGEON FOUND THAT THE RETAINING PIN COULD NOT PUSH TO CLOSE IT, THE RETAINING PIN BOUNCE BACK TO THE ORIGINAL POSITION. SURGEON TRIED TO CLOSE THE CLOSING HANDLE BUT FAILED TO CLOSE. ENDED OPEN COMPETITIVE DEVICE BECAUSE THAT IS THE LAST PIECE OF AVAILABLE TX60G IN OT. THE STAFF NURSE PASS THE TX60G TO ME AFTER WASH. AFTER TRYING FOR A FEW TIMES, THE RETAINING PIN & CLOSING HANDLE CAN CLOSE AND ABLE TO FIRE. THEREFORE, YOU MAY SEE SOME PART OF THE STAPLER HAS STAPLER. THE SURGERY WAS DELAYED BY 15-MINUTES AND SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
07-12-2020 | Injury | ethicon endo-surgery | 04-01-2021 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT THE CLIPS OF THE RELOAD ECR60 OPENED RIGHT AFTER THE STAPLING (INTRA-OPERATIVE) IN A BARIATRIC SURGERY THAT HAPPENED ON (B)(6) 2020. THE SURGEON HAD TO MANUALLY SUTURE IN ORDER TO STOP THE BLEEDING AND COMPLETE THE ANASTOMOSIS. THE SURGERY WAS SUCCESSFULLY COMPLETED, BUT THE PATIENT IS STILL UNDER OBSERVATION.
10-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, STAPLER WAS IN ¿GREEN ZONE¿ WITH CORRECT FINGER TIGHT STAPLE HEIGHT SELECTED AND WOULD NOT FIRE. BATTERY WAS REMOVED AND REINSERTED, STAPLER WOULD NOT FIRE. NEW BATTERY WAS INSERTED FROM ANOTHER NEW STAPLER, GREEN TICK WAS ILLUMINATED, AND STAPLER WOULD NOT FIRE. SURGEON WAS ASKED IF THE ANVIL HAD BEEN ATTACHED CORRECTLY AND IF ORANGE MARK COULD BE SEEN. SURGEON COMMENTED THAT THEY HAD HEARD AND FELT ANVIL CONNECT. A NEW STAPLER WAS CONNECTED TO THE EXISTING ANVIL AND FIRED CORRECTLY WITH COMPLETE DONUTS. THE PROCEDURE WAS COMPLETED SUCCESSFULLY BUT DELAYED BY 15-MINUTES. NO PATIENT CONSEQUENCES REPORTED.
04-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY THE FIRST DEICE WAS USED THREE TIMES WITH NO ISSUES. ON THE FORTH RELOAD IT WOULD NOT FIRE AND WOULD NOT OPEN. THE MANUAL OVER RIDE HAD TO BE USE TO OPEN IT. THE SECOND DEVICE WAS FIRED AND THERE WAS A "HORRIBLE" NOISE AND DID NOT WORK. A THIRD DEVICE WAS TRIED AND IT MISFIRED. THE FOURTH DEVICE FIRED TWICE WITH NO ISSUES AND ON THE NEXT FIRE IT MISFIRED. THE LAST STAPLER WAS TRIED AGAIN WITH THE LAST RELOAD AND IT WORKED. THE FIRST STAPLER LOCKED ON TISSUE AND MANUAL OVERRIDE WAS USED TO RELEASE IT. THE SECOND DEVICE COULD NOT FULLY RECEIVE STAPLER LOADS. "SEVERAL OF US IN ROOM TRIED." THIRD DEVICE MISFIRED ON PATIENT'S STOMACH, AFTER FIRING, THE SPECIMEN SIDE HAD UNRAVELED WITH SPILLAGE OF GASTRIC CONTENTS. THE SECOND STAPLER WOULD NOT RECEIVE THE RELOAD WHEN SEATED INTO THE STAPLER CHANNEL. IT WAS THE THIRD STAPLER THAT THE SURGEON HEARD A METALLIC SOUND AND THE SURGEON ¿FELT LIKE SOMETHING HAPPENED TO THE BLADE¿. ON THE SPECIMEN SIDE OF THE SLEEVE UNRAVELED ON AT THE STAPLE LINE. FIRST DEVICE: THE POWER FAILED THERE WAS NO BATTERY POWER. IT DIDN'T' MOVE FORWARD OR BACK. WE HAD TO MANUALLY TAKE IT OFF TISSUE. SECOND DEVICE: AT THE SECOND FIRING OF THE SECOND DEVICE IT STARTED MAKING NOISES. IT MADE LOUD CLICKING NOISES SO WE GOT RID OF THAT ONE TOO. THIRD DEVICE: THE THIRD STAPLER "FLAT OUT" FAILED ON THE RIGHT SIDE THE STAPLES DIDN'T "CURL" AT ALL. FOURTH DEVICE: THERE WERE NO ISSUES WITH THE FOURTH DEVICE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-12-2020 | Malfunction | ethicon | 04-01-2021 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC UMBILICAL HERNIA REPAIR ON (B)(6) 2020 AND THE ABSORBABLE STRAPS WERE USED. IT WAS REPORTED THAT THE TACKS IN THE DEVICE DIDN'T HAVE ANY BARBS IN THEM AND WEREN'T ABLE TO PLACE THE MESH PROPERLY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED.
12-11-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING GENERAL SURGERY, THE DEVICE DID NOT FIRE AFTER BEING RELOADED. SURGEON WAS NOT WILLING TO RELOAD SO NEW DEVICE WAS OPENED TO COMPLETE CASE. NO PATIENT COMPLICATIONS.
01-10-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS ON THE CHARGER PRE-OPERATIVELY, THE DISPLAY WAS BLACK AND THE LIGHTSHOW WERE BLUE AND RED. THE TWO SILICONE BUTTONS WERE PRESSED BUT THERE WAS NO REACTION. THE SURGEON USED ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
01-10-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS ON THE CHARGER PRE-OPERATIVELY, THE DISPLAY WAS BLACK AND THE LIGHTSHOW WERE BLUE AND RED. THE TWO SILICONE BUTTONS WERE PRESSED BUT THERE WAS NO REACTION. THE SURGEON USED ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
01-10-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS ON THE CHARGER PRE-OPERATIVELY, THE DISPLAY WAS BLACK AND THE LIGHTSHOW WERE BLUE AND RED. THE TWO SILICONE BUTTONS WERE PRESSED BUT THERE WAS NO REACTION. THE SURGEON USED ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
01-10-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS ON THE CHARGER PRE-OPERATIVELY, THE DISPLAY WAS BLACK AND THE LIGHTSHOW WERE BLUE AND RED. THE TWO SILICONE BUTTONS WERE PRESSED BUT THERE WAS NO REACTION. THE SURGEON USED ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
08-12-2020 | Injury | ethicon endo-surgery | 04-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT IN SLEEVE GASTRECTOMY WITH HIATAL HERNIA PROCEDURE, ON THE THIRD FIRING WITH A BLUE RELOAD USING PLEE60A AND ECH60R, THE STAPLE LINE REINFORCEMENT WAS PUSHED FORWARD INTO A BUNCH AS THE KNIFE BLADE ADVANCED TO THE DISTAL TIP. THIS CAUSED A MISFIRE AS THE STAPLE LEGS DID NOT PROPERLY FORM. THE ECH60R WAS NOT CUT/TRANSECTED FROM THE KNIFE BLADE. THE SURGEON DID NOT USE ANY MORE ECH60R AFTER MISHAP AND CONTROLLED THE BLEEDING BY FIRING ANOTHER RELOAD, OVERSOWING THE GREATER CURVATURE UP TO FUNDUS, USING A CLIP APPLIER, AND VISTASEAL. SAME STAPLER WAS USED FOR THE REST OF THE CASE AND ALL RELOADS USED POST MISFIRE WERE NAKED (NOT WITH SLR). THE STAPLER, RELOADS, AND SLR APPLICATORS WILL BE SUBMITTED TO BE FURTHER INSPECTED. THE CASE WAS DELAYED BY THIRTY MINUTES BUT COMPLETED SUCCESSFULLY. THE EVENT OCCURRED ON THE 3RD FIRING USING A BLUE RELOAD. AS THE KNIFE ADVANCED, IT LOOKED LIKE IT WAS PUSHING THE TISSUE OUT AND THE MOTOR SOUND SLOWED DOWN. WHEN THE JAWS WERE OPENED, SOME BLEEDING WAS NOTICED, AND NO STAPLES FORMED. TO CORRECT ISSUE, A GOLD RELOAD WAS USED MEDIAL TO DEFECT WITH NO SLR. THE SAME DEVICE WAS USED FOR THE REST OF THE CASE WITH NO SLR. A TOTAL OF 3 FIRINGS USING SLR WERE COMPLETED. CLIPS AND OVER-SEWING WERE USED. THE PATIENT DID DEVELOP A POST-OP LEAK AT SITE OF MISFIRE AND REQUIRED A REOP AND WAS IN ICU. SURGEON TECHNIQUE- FIRES ANVIL UP, CONTINUOUS FIRING STROKE. HE IS CONSISTENT ABOUT CROSSING STAPLE LINES. WAS IN APEX DURING FIRING WITH ISSUE. FIRED DEVICE APPROXIMATELY 1-2 MINUTES AFTER SLR WAS LOADED. SWISHING OF DEVICE IN DEEP PLASTIC BASIN AND DRYING PROPERLY. NORMAL FOLLOWING OF BOUGIE AND USES 5-6 RELOADS PER SLEEVE.
08-12-2020 | Injury | ethicon endo-surgery | 04-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Cut | A050702 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT IN SLEEVE GASTRECTOMY WITH HIATAL HERNIA PROCEDURE, ON THE THIRD FIRING WITH A BLUE RELOAD USING PLEE60A AND ECH60R, THE STAPLE LINE REINFORCEMENT WAS PUSHED FORWARD INTO A BUNCH AS THE KNIFE BLADE ADVANCED TO THE DISTAL TIP. THIS CAUSED A MISFIRE AS THE STAPLE LEGS DID NOT PROPERLY FORM. THE ECH60R WAS NOT CUT/TRANSECTED FROM THE KNIFE BLADE. THE SURGEON DID NOT USE ANY MORE ECH60R AFTER MISHAP AND CONTROLLED THE BLEEDING BY FIRING ANOTHER RELOAD, OVERSOWING THE GREATER CURVATURE UP TO FUNDUS, USING A CLIP APPLIER, AND VISTASEAL. SAME STAPLER WAS USED FOR THE REST OF THE CASE AND ALL RELOADS USED POST MISFIRE WERE NAKED (NOT WITH SLR). THE STAPLER, RELOADS, AND SLR APPLICATORS WILL BE SUBMITTED TO BE FURTHER INSPECTED. THE CASE WAS DELAYED BY THIRTY MINUTES BUT COMPLETED SUCCESSFULLY. THE EVENT OCCURRED ON THE 3RD FIRING USING A BLUE RELOAD. AS THE KNIFE ADVANCED, IT LOOKED LIKE IT WAS PUSHING THE TISSUE OUT AND THE MOTOR SOUND SLOWED DOWN. WHEN THE JAWS WERE OPENED, SOME BLEEDING WAS NOTICED, AND NO STAPLES FORMED. TO CORRECT ISSUE, A GOLD RELOAD WAS USED MEDIAL TO DEFECT WITH NO SLR. THE SAME DEVICE WAS USED FOR THE REST OF THE CASE WITH NO SLR. A TOTAL OF 3 FIRINGS USING SLR WERE COMPLETED. CLIPS AND OVER-SEWING WERE USED. THE PATIENT DID DEVELOP A POST-OP LEAK AT SITE OF MISFIRE AND REQUIRED A REOP AND WAS IN ICU. SURGEON TECHNIQUE- FIRES ANVIL UP, CONTINUOUS FIRING STROKE. HE IS CONSISTENT ABOUT CROSSING STAPLE LINES. WAS IN APEX DURING FIRING WITH ISSUE. FIRED DEVICE APPROXIMATELY 1-2 MINUTES AFTER SLR WAS LOADED. SWISHING OF DEVICE IN DEEP PLASTIC BASIN AND DRYING PROPERLY. NORMAL FOLLOWING OF BOUGIE AND USES 5-6 RELOADS PER SLEEVE.
08-12-2020 | Injury | ethicon endo-surgery | 04-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT IN SLEEVE GASTRECTOMY WITH HIATAL HERNIA PROCEDURE, ON THE THIRD FIRING WITH A BLUE RELOAD USING PLEE60A AND ECH60R, THE STAPLE LINE REINFORCEMENT WAS PUSHED FORWARD INTO A BUNCH AS THE KNIFE BLADE ADVANCED TO THE DISTAL TIP. THIS CAUSED A MISFIRE AS THE STAPLE LEGS DID NOT PROPERLY FORM. THE ECH60R WAS NOT CUT/TRANSECTED FROM THE KNIFE BLADE. THE SURGEON DID NOT USE ANY MORE ECH60R AFTER MISHAP AND CONTROLLED THE BLEEDING BY FIRING ANOTHER RELOAD, OVERSOWING THE GREATER CURVATURE UP TO FUNDUS, USING A CLIP APPLIER, AND VISTASEAL. SAME STAPLER WAS USED FOR THE REST OF THE CASE AND ALL RELOADS USED POST MISFIRE WERE NAKED (NOT WITH SLR). THE STAPLER, RELOADS, AND SLR APPLICATORS WILL BE SUBMITTED TO BE FURTHER INSPECTED. THE CASE WAS DELAYED BY THIRTY MINUTES BUT COMPLETED SUCCESSFULLY. THE EVENT OCCURRED ON THE 3RD FIRING USING A BLUE RELOAD. AS THE KNIFE ADVANCED, IT LOOKED LIKE IT WAS PUSHING THE TISSUE OUT AND THE MOTOR SOUND SLOWED DOWN. WHEN THE JAWS WERE OPENED, SOME BLEEDING WAS NOTICED, AND NO STAPLES FORMED. TO CORRECT ISSUE, A GOLD RELOAD WAS USED MEDIAL TO DEFECT WITH NO SLR. THE SAME DEVICE WAS USED FOR THE REST OF THE CASE WITH NO SLR. A TOTAL OF 3 FIRINGS USING SLR WERE COMPLETED. CLIPS AND OVER-SEWING WERE USED. THE PATIENT DID DEVELOP A POST-OP LEAK AT SITE OF MISFIRE AND REQUIRED A REOP AND WAS IN ICU. SURGEON TECHNIQUE- FIRES ANVIL UP, CONTINUOUS FIRING STROKE. HE IS CONSISTENT ABOUT CROSSING STAPLE LINES. WAS IN APEX DURING FIRING WITH ISSUE. FIRED DEVICE APPROXIMATELY 1-2 MINUTES AFTER SLR WAS LOADED. SWISHING OF DEVICE IN DEEP PLASTIC BASIN AND DRYING PROPERLY. NORMAL FOLLOWING OF BOUGIE AND USES 5-6 RELOADS PER SLEEVE.
08-12-2020 | Injury | ethicon endo-surgery | 04-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT IN SLEEVE GASTRECTOMY WITH HIATAL HERNIA PROCEDURE, ON THE THIRD FIRING WITH A BLUE RELOAD USING PLEE60A AND ECH60R, THE STAPLE LINE REINFORCEMENT WAS PUSHED FORWARD INTO A BUNCH AS THE KNIFE BLADE ADVANCED TO THE DISTAL TIP. THIS CAUSED A MISFIRE AS THE STAPLE LEGS DID NOT PROPERLY FORM. THE ECH60R WAS NOT CUT/TRANSECTED FROM THE KNIFE BLADE. THE SURGEON DID NOT USE ANY MORE ECH60R AFTER MISHAP AND CONTROLLED THE BLEEDING BY FIRING ANOTHER RELOAD, OVERSOWING THE GREATER CURVATURE UP TO FUNDUS, USING A CLIP APPLIER, AND VISTASEAL. SAME STAPLER WAS USED FOR THE REST OF THE CASE AND ALL RELOADS USED POST MISFIRE WERE NAKED (NOT WITH SLR). THE STAPLER, RELOADS, AND SLR APPLICATORS WILL BE SUBMITTED TO BE FURTHER INSPECTED. THE CASE WAS DELAYED BY THIRTY MINUTES BUT COMPLETED SUCCESSFULLY. THE EVENT OCCURRED ON THE 3RD FIRING USING A BLUE RELOAD. AS THE KNIFE ADVANCED, IT LOOKED LIKE IT WAS PUSHING THE TISSUE OUT AND THE MOTOR SOUND SLOWED DOWN. WHEN THE JAWS WERE OPENED, SOME BLEEDING WAS NOTICED, AND NO STAPLES FORMED. TO CORRECT ISSUE, A GOLD RELOAD WAS USED MEDIAL TO DEFECT WITH NO SLR. THE SAME DEVICE WAS USED FOR THE REST OF THE CASE WITH NO SLR. A TOTAL OF 3 FIRINGS USING SLR WERE COMPLETED. CLIPS AND OVER-SEWING WERE USED. THE PATIENT DID DEVELOP A POST-OP LEAK AT SITE OF MISFIRE AND REQUIRED A REOP AND WAS IN ICU. SURGEON TECHNIQUE- FIRES ANVIL UP, CONTINUOUS FIRING STROKE. HE IS CONSISTENT ABOUT CROSSING STAPLE LINES. WAS IN APEX DURING FIRING WITH ISSUE. FIRED DEVICE APPROXIMATELY 1-2 MINUTES AFTER SLR WAS LOADED. SWISHING OF DEVICE IN DEEP PLASTIC BASIN AND DRYING PROPERLY. NORMAL FOLLOWING OF BOUGIE AND USES 5-6 RELOADS PER SLEEVE.
08-06-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN ATTEMPTING TO SET UP, THE DISPLAY WAS FOUND TO BE BLACKOUT SO THE PRODUCT COULD NOT BE USED. THE DEVICE WAS UNRESPONSIVE WHEN REMOVED FROM THE CHARGER. ANOTHER HANDLE WAS USED TO DEAL WITH THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
08-06-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN ATTEMPTING TO SET UP, THE DISPLAY WAS FOUND TO BE BLACKOUT SO THE PRODUCT COULD NOT BE USED. THE DEVICE WAS UNRESPONSIVE WHEN REMOVED FROM THE CHARGER. ANOTHER HANDLE WAS USED TO DEAL WITH THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
08-06-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN ATTEMPTING TO SET UP, THE DISPLAY WAS FOUND TO BE BLACKOUT SO THE PRODUCT COULD NOT BE USED. THE DEVICE WAS UNRESPONSIVE WHEN REMOVED FROM THE CHARGER. ANOTHER HANDLE WAS USED TO DEAL WITH THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
05-06-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, WHEN THE DEVICE WAS CHECKED, IT WAS BLACKED OUT AND UNABLE TO BE USED. THE DEVICE WAS TOTALLY UNRESPONSIVE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
05-06-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, WHEN THE DEVICE WAS CHECKED, IT WAS BLACKED OUT AND UNABLE TO BE USED. THE DEVICE WAS TOTALLY UNRESPONSIVE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
11-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | PROXIMATE*75MM LINEAR CUTTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DIVIDED A THICKENED BIT OF SMALL BOWEL WITH TLC TO FIND THAT IT HAD ONLY STAPLED ONE HALF AND NOT THE OTHER. NO STAPLES WERE PRESENT IN ONE HALF OF THE RELOAD. THE SURGEON HAD TO REPAIR THE HOLE WITH SUTURES. NO PATIENT CONSEQUENCE.
11-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | PROXIMATE*75MM LINEAR CUTTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DIVIDED A THICKENED BIT OF SMALL BOWEL WITH TLC TO FIND THAT IT HAD ONLY STAPLED ONE HALF AND NOT THE OTHER. NO STAPLES WERE PRESENT IN ONE HALF OF THE RELOAD. THE SURGEON HAD TO REPAIR THE HOLE WITH SUTURES. THERE WERE NO PATIENT CONSEQUENCES.
25-10-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RECTAL SURGERY, COULD NOT FIRE WHEN SEVERING RECTUM TISSUE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-02-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE ON A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY PROCEDURE, THEY PERFORMED A START UP OF THE POWER HANDLE, BUT A HANDLE ERROR OCCURRED. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
03-02-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE ON A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY PROCEDURE, THEY PERFORMED A START UP OF THE POWER HANDLE, BUT A HANDLE ERROR OCCURRED. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
03-02-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE ON A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY PROCEDURE, THEY PERFORMED A START UP OF THE POWER HANDLE, BUT A HANDLE ERROR OCCURRED. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
01-09-2020 | Injury | us surgical puerto rico | 04-01-2021 | GIA | Misfire | A050502 | Bowel Perforation | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT COLECTOMY, THE DEVICE MISFIRED WHICH CAUSING A BOWEL LEAK.
25-06-2020 | Malfunction | us surgical puerto rico | 04-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. A NEW PRODUCT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2020 | Malfunction | us surgical puerto rico | 04-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, UPON TRANSACTION OF FISSURE, THE DEVICE DID NOT FIRE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. RELOAD AND STAPLER WERE CHANGE TO ANOTHER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE CARDIAC PATIENT HAD PREVIOUS CARDIAC PROCEDURES. DURING DISSECTION OF ADHESIONS ON THE LUNG, THE SURGEON CALLED FOR A POWERED ECHELON 45MM TO EXCISE A WEDGE OF LUNG TISSUE. GST45B WAS COLOUR RELOAD. 2 ROWS OF STAPLING WAS COMPLETED. A 3RD RELOAD WAS REQUESTED, AND THE STAPLER WAS FIRED. SURGEON COMMENTED THAT THE STAPLER HADN'T FIRED, AND SOMEWHAT STUCK. THEY COULD NOT OPEN THE JAWS, AND THE KNIFE REVERSE BUTTON, DIDN'T ALLOW THE JAWS TO OPEN. THEY CALLED FOR ANOTHER THORACIC SURGEON, WITH FAMILIAR WITH ECHELON. THEY MOVED STRAIGHT TO BAIL OUT STEPS. WOUND THE BLACK LEVER BACK AND FORTH. STAPLER WOULD NOT OPEN. SURGEONS CALL FOR A 2ND STAPLER, LOADED A GST45B, AND STAPLED BELOW THE STUCK STAPLER WHICH WORKED. THEN THEY REMOVED THE STUCK STAPLER, AND DUVAL INSTRUMENT AND THE WEDGE OF TISSUE. ONE OF THE SURGEONS INVESTIGATED THE JAW OF THE STAPLER THAT WOULD NOT OPEN TO FIND, THAT THE INSTRUMENT WAS STUCK IN THE JAW. ONCE THEY REMOVED THE INSTRUMENT, THE STAPLER OPENED, AS NORMAL. PROCEDURE CONTINUED WITH NO FURTHER ISSUES WITH STAPLING.
09-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE CARDIAC PATIENT HAD PREVIOUS CARDIAC PROCEDURES. DURING DISSECTION OF ADHESIONS ON THE LUNG, THE SURGEON CALLED FOR A POWERED ECHELON 45MM TO EXCISE A WEDGE OF LUNG TISSUE. GST45B WAS COLOUR RELOAD. 2 ROWS OF STAPLING WAS COMPLETED. A 3RD RELOAD WAS REQUESTED, AND THE STAPLER WAS FIRED. SURGEON COMMENTED THAT THE STAPLER HADN'T FIRED, AND SOMEWHAT STUCK. THEY COULD NOT OPEN THE JAWS, AND THE KNIFE REVERSE BUTTON, DIDN'T ALLOW THE JAWS TO OPEN. THEY CALLED FOR ANOTHER THORACIC SURGEON, WITH FAMILIAR WITH ECHELON. THEY MOVED STRAIGHT TO BAIL OUT STEPS. WOUND THE BLACK LEVER BACK AND FORTH. STAPLER WOULD NOT OPEN. SURGEONS CALL FOR A 2ND STAPLER, LOADED A GST45B, AND STAPLED BELOW THE STUCK STAPLER WHICH WORKED. THEN THEY REMOVED THE STUCK STAPLER, AND DUVAL INSTRUMENT AND THE WEDGE OF TISSUE. ONE OF THE SURGEONS INVESTIGATED THE JAW OF THE STAPLER THAT WOULD NOT OPEN TO FIND, THAT THE INSTRUMENT WAS STUCK IN THE JAW. ONCE THEY REMOVED THE INSTRUMENT, THE STAPLER OPENED, AS NORMAL. PROCEDURE CONTINUED WITH NO FURTHER ISSUES WITH STAPLING.
09-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | POWERED PLUS STANDARD 45 | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE CARDIAC PATIENT HAD PREVIOUS CARDIAC PROCEDURES. DURING DISSECTION OF ADHESIONS ON THE LUNG, THE SURGEON CALLED FOR A POWERED ECHELON 45MM TO EXCISE A WEDGE OF LUNG TISSUE. GST45B WAS COLOUR RELOAD. 2 ROWS OF STAPLING WAS COMPLETED. A 3RD RELOAD WAS REQUESTED, AND THE STAPLER WAS FIRED. SURGEON COMMENTED THAT THE STAPLER HADN'T FIRED, AND SOMEWHAT STUCK. THEY COULD NOT OPEN THE JAWS, AND THE KNIFE REVERSE BUTTON, DIDN'T ALLOW THE JAWS TO OPEN. THEY CALLED FOR ANOTHER THORACIC SURGEON, WITH FAMILIAR WITH ECHELON. THEY MOVED STRAIGHT TO BAIL OUT STEPS. WOUND THE BLACK LEVER BACK AND FORTH. STAPLER WOULD NOT OPEN. SURGEONS CALL FOR A 2ND STAPLER, LOADED A GST45B, AND STAPLED BELOW THE STUCK STAPLER WHICH WORKED. THEN THEY REMOVED THE STUCK STAPLER, AND DUVAL INSTRUMENT AND THE WEDGE OF TISSUE. ONE OF THE SURGEONS INVESTIGATED THE JAW OF THE STAPLER THAT WOULD NOT OPEN TO FIND, THAT THE INSTRUMENT WAS STUCK IN THE JAW. ONCE THEY REMOVED THE INSTRUMENT, THE STAPLER OPENED, AS NORMAL. PROCEDURE CONTINUED WITH NO FURTHER ISSUES WITH STAPLING.
09-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | POWERED PLUS STANDARD 45 | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE CARDIAC PATIENT HAD PREVIOUS CARDIAC PROCEDURES. DURING DISSECTION OF ADHESIONS ON THE LUNG, THE SURGEON CALLED FOR A POWERED ECHELON 45MM TO EXCISE A WEDGE OF LUNG TISSUE. GST45B WAS COLOUR RELOAD. 2 ROWS OF STAPLING WAS COMPLETED. A 3RD RELOAD WAS REQUESTED, AND THE STAPLER WAS FIRED. SURGEON COMMENTED THAT THE STAPLER HADN'T FIRED, AND SOMEWHAT STUCK. THEY COULD NOT OPEN THE JAWS, AND THE KNIFE REVERSE BUTTON, DIDN'T ALLOW THE JAWS TO OPEN. THEY CALLED FOR ANOTHER THORACIC SURGEON, WITH FAMILIAR WITH ECHELON. THEY MOVED STRAIGHT TO BAIL OUT STEPS. WOUND THE BLACK LEVER BACK AND FORTH. STAPLER WOULD NOT OPEN. SURGEONS CALL FOR A 2ND STAPLER, LOADED A GST45B, AND STAPLED BELOW THE STUCK STAPLER WHICH WORKED. THEN THEY REMOVED THE STUCK STAPLER, AND DUVAL INSTRUMENT AND THE WEDGE OF TISSUE. ONE OF THE SURGEONS INVESTIGATED THE JAW OF THE STAPLER THAT WOULD NOT OPEN TO FIND, THAT THE INSTRUMENT WAS STUCK IN THE JAW. ONCE THEY REMOVED THE INSTRUMENT, THE STAPLER OPENED, AS NORMAL. PROCEDURE CONTINUED WITH NO FURTHER ISSUES WITH STAPLING.
07-12-2020 | Malfunction | covidien lp llc north haven | 04-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, UPON TRANSACTION OF FISSURE, THE DEVICE DID NOT FIRE. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. RELOAD AND STAPLER WERE CHANGE TO ANOTHER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | intuitive surgical | 04-01-2021 | SUREFORM | Failure to Form Staple | A050704 | No Known Impact Or Consequence To Patient | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
29-12-2020 | Malfunction | ethicon endo-surgery | 04-01-2021 | ETHICON ECHELON FLEX 60 | Therapeutic or Diagnostic Output Failure | A090406 | Insufficient Information | Event Description: STAPLER STOPPED WORKING MIDWAY THROUGH STAPLE LINE. STAPLE LINE FAILED. SURGERY CONVERTED TO OPEN. FDA SAFETY REPORT ID # (B)(4).
08-12-2020 | Injury | covidien lp llc north haven | 05-01-2021 | RELIAMAX | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A HEMICOLECTOMY, THERE WAS POOR STAPLE FORMATION, THE STAPLE LINE WAS INCOMPLETE, THE USER OVERSEWED TO FIX THE ISSUE. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE.
07-12-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | TA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE.
07-12-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE.
07-12-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | TA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE.
27-11-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS UNABLE TO HOLD THE CLAMP IN A STRAIGHT POSITION. ORIENTATION IS ONLY LEFT OR RIGHT BUT NOT IN THE STRAIGHT POSITION. THE SURGEON COULD NOT STAPLE IN THE STRAIGHT POSITION. IT REMAINS IN THE ORIENTED POSITION EVEN WHEN THE ORIENTATION KNOB IS IN THE STRAIGHT POSITION. THEY TOOK A NEW DEVICE TO FINISH AND EVERYTHING WAS FINE. NO PATIENT INJURY.
07-12-2020 | Injury | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE DISSECTION OF THE STOMACH, THERE WAS POOR STAPLE FORMATION. THE STAPLES WERE JERKY, AND THE LINEAR STAPLES WERE NOT STRAIGHT. THE TISSUE WAS DAMAGED AS A RESULT AND THERE WAS A LITTLE BLEEDING. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
07-12-2020 | Injury | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE DISSECTION OF THE STOMACH, THERE WAS POOR STAPLE FORMATION. THE STAPLES WERE JERKY, AND THE LINEAR STAPLES WERE NOT STRAIGHT. THE TISSUE WAS DAMAGED AS A RESULT AND THERE WAS A LITTLE BLEEDING. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
02-02-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, ON A SLEEVE GASTRECTOMY, DURING THE DISSECTION OF THE STOMACH, THE STAPLES WERE JERKY, AND THE LINEAR STA PLES WERE NOT STRAIGHT. THERE WAS TISSUE DAMAGE AND A LITTLE BLEEDING ALONG THE LINE OF THE SLEEVE. ANOTHER RELOAD FROM ANOTHER LOT WAS USED TO COMPLETE THE CASE.
02-02-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A SLEEVE GASTRECTOMY, DURING THE DISSECTION OF THE STOMACH, THE STAPLES WERE JERKY, AND THE LINEAR STA PLES WERE NOT STRAIGHT. THERE WAS TISSUE DAMAGE AND A LITTLE BLEEDING ALONG THE LINE OF THE SLEEVE. ANOTHER RELOAD FROM ANOTHER LOT WAS USED TO COMPLETE THE CASE.
07-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Break | A0401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY, THE FIRST HANDLE THAT WAS USED FIRED WITH PARTIAL FIRING IN THE PANCREAS WITH SHREDDED STAPLES AND TISSUE AND MADE A CRACKING SOUND AS IT REACHED ITS BREAKING POINT. A NEW STAPLER WAS OPENED AND MORE DEBULKING OF PANCREAS WAS DONE TO ALLOW FOR FIRING, IT PERFORMED BETTER BUT HAD POORLY FORMED STAPLES. THE SURGICAL PROCEDURE WAS CONVERTED FROM A LAPAROSCOPIC TO AN OPEN PROCEDURE AND A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THE INCISION WAS EXTENDED BY MORE THAN 1 INCH.
07-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY, THE FIRST HANDLE THAT WAS USED FIRED WITH PARTIAL FIRING IN THE PANCREAS WITH SHREDDED STAPLES AND TISSUE AND MADE A CRACKING SOUND AS IT REACHED ITS BREAKING POINT. A NEW STAPLER WAS OPENED AND MORE DEBULKING OF PANCREAS WAS DONE TO ALLOW FOR FIRING, IT PERFORMED BETTER BUT HAD POORLY FORMED STAPLES. THE SURGICAL PROCEDURE WAS CONVERTED FROM A LAPAROSCOPIC TO AN OPEN PROCEDURE AND A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THE INCISION WAS EXTENDED BY MORE THAN 1 INCH.
07-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY, THE FIRST HANDLE THAT WAS USED FIRED WITH PARTIAL FIRING IN THE PANCREAS WITH SHREDDED STAPLES AND TISSUE AND MADE A CRACKING SOUND AS IT REACHED ITS BREAKING POINT. A NEW STAPLER WAS OPENED AND MORE DEBULKING OF PANCREAS WAS DONE TO ALLOW FOR FIRING, IT PERFORMED BETTER BUT HAD POORLY FORMED STAPLES. THE SURGICAL PROCEDURE WAS CONVERTED FROM A LAPAROSCOPIC TO AN OPEN PROCEDURE AND A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THE INCISION WAS EXTENDED BY MORE THAN 1 INCH.
07-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Noise, Audible | A0508 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY, THE FIRST HANDLE THAT WAS USED FIRED WITH PARTIAL FIRING IN THE PANCREAS WITH SHREDDED STAPLES AND TISSUE AND MADE A CRACKING SOUND AS IT REACHED ITS BREAKING POINT. A NEW STAPLER WAS OPENED AND MORE DEBULKING OF PANCREAS WAS DONE TO ALLOW FOR FIRING, IT PERFORMED BETTER BUT HAD POORLY FORMED STAPLES. THE SURGICAL PROCEDURE WAS CONVERTED FROM A LAPAROSCOPIC TO AN OPEN PROCEDURE AND A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THE INCISION WAS EXTENDED BY MORE THAN 1 INCH.
07-12-2020 | Injury | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY PROCEDURE, THE DEVICE WOULD NOT FIRE USING THE FIRST HANDLE. A SECOND HANDLE WAS USED AND THE STAPLER FIRED BUT THERE WAS STAPLE LINE INCOMPLETE AND POOR STAPLE FORMATION. IT WAS ALSO REPORTED THAT THERE WAS CRACKLING WITH THE HANDLE. THE SURGICAL PROCEDURE WAS CONVERTED FROM A LAPAROSCOPIC TO AN OPEN PROCEDURE AND A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE.
07-12-2020 | Injury | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY PROCEDURE, THE DEVICE WOULD NOT FIRE USING THE FIRST HANDLE. A SECOND HANDLE WAS USED AND THE STAPLER FIRED BUT THERE WAS STAPLE LINE INCOMPLETE AND POOR STAPLE FORMATION. IT WAS ALSO REPORTED THAT THERE WAS CRACKLING WITH THE HANDLE. THE SURGICAL PROCEDURE WAS CONVERTED FROM A LAPAROSCOPIC TO AN OPEN PROCEDURE AND A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE.
07-12-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING INITIAL FIRING ACROSS BASE OF APPENDIX ON A LAPAROSCOPIC APPENDECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. THE SURGEON AND SALES REP LOOKED AT RELOAD. STAPLES ARE VISIBLE IN THE RELOAD. SURGEON HAD PRE-FIRED OR BUMPED THE BLACK RING HANDLE CAUSING THE RELOAD TO PREMATURELY FIRE. ANOTHER RELOAD WAS OPENED AND USED ON THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING INITIAL FIRING ACROSS BASE OF APPENDIX ON A LAPAROSCOPIC APPENDECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. THE SURGEON AND SALES REP LOOKED AT RELOAD. STAPLES ARE VISIBLE IN THE RELOAD. SURGEON HAD PRE-FIRED OR BUMPED THE BLACK RING HANDLE CAUSING THE RELOAD TO PREMATURELY FIRE. ANOTHER RELOAD WAS OPENED AND USED ON THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, ON AN INCISION AND ANASTOMOSIS OF PANCREAS DURING A LAPAROSCOPIC PANCREATECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE COULD NOT BE SQUEEZED. THE STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, ON AN INCISION AND ANASTOMOSIS OF PANCREAS DURING A LAPAROSCOPIC PANCREATECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE COULD NOT BE SQUEEZED. THE STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
09-07-2020 | Injury | covidien lp llc north haven | 05-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED MAY 2018 TO AUGUST 2019, 123 PATIENTS UNDERWENT SIGMOIDECTOMY, HIGH ANTERIOR RESECTION, LOW ANTERIOR RESECTION, OR ULTRA-LOW ANTERIOR RESECTION. LASER SPECKLE FLOWGRAPHY WAS USED IN 36 PATIENTS AND 87 PATIENTS WERE IN THE CONTROL GROUP. IN 96 PATIENTS, THE POWERED STAPLER WAS USED TO TRANSECT THE DISTAL COLON OR THE RECTUM IN A DOUBLE STAPLE TECHNIQUE FOR ANASTOMOSIS. HAND-SEWN ANASTOMOSIS WAS USED IN 27 PATIENTS. POST-OPERATIVE COMPLICATIONS INCLUDED: 16 ANASTOMOTIC LEAKS, 2 PATIENTS REQUIRING REOPERATION, INTRA-OPERATIVE BLOOD LOSS WITH 7 PATIENTS REQUIRING TRANSFUSIONS.
09-07-2020 | Injury | covidien lp llc north haven | 05-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED MAY 2018 TO AUGUST 2019, 123 PATIENTS UNDERWENT SIGMOIDECTOMY, HIGH ANTERIOR RESECTION, LOW ANTERIOR RESECTION, OR ULTRA-LOW ANTERIOR RESECTION. LASER SPECKLE FLOWGRAPHY WAS USED IN 36 PATIENTS AND 87 PATIENTS WERE IN THE CONTROL GROUP. IN 96 PATIENTS, THE POWERED STAPLER WAS USED TO TRANSECT THE DISTAL COLON OR THE RECTUM IN A DOUBLE STAPLE TECHNIQUE FOR ANASTOMOSIS. HAND-SEWN ANASTOMOSIS WAS USED IN 27 PATIENTS. POST-OPERATIVE COMPLICATIONS INCLUDED: 16 ANASTOMOTIC LEAKS, 2 PATIENTS REQUIRING REOPERATION, INTRA-OPERATIVE BLOOD LOSS WITH 7 PATIENTS REQUIRING TRANSFUSIONS.
09-07-2020 | Injury | covidien lp llc north haven | 05-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED MAY 2018 TO AUGUST 2019, 123 PATIENTS UNDERWENT SIGMOIDECTOMY, HIGH ANTERIOR RESECTION, LOW ANTERIOR RESECTION, OR ULTRA-LOW ANTERIOR RESECTION. LASER SPECKLE FLOWGRAPHY WAS USED IN 36 PATIENTS AND 87 PATIENTS WERE IN THE CONTROL GROUP. IN 96 PATIENTS, THE POWERED STAPLER WAS USED TO TRANSECT THE DISTAL COLON OR THE RECTUM IN A DOUBLE STAPLE TECHNIQUE FOR ANASTOMOSIS. HAND-SEWN ANASTOMOSIS WAS USED IN 27 PATIENTS. POST-OPERATIVE COMPLICATIONS INCLUDED: 16 ANASTOMOTIC LEAKS, 2 PATIENTS REQUIRING REOPERATION, INTRA-OPERATIVE BLOOD LOSS WITH 7 PATIENTS REQUIRING TRANSFUSIONS.
06-10-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE POWER HANDLE DID NOT START WHEN REMOVED FROM THE CHARGER. THE CHARGING INDICATOR WAS REPEATEDLY ILLUMINATING, INDICATING THAT THE CHARGING WAS COMPLETED. THE HANDLE WAS PLACED ON THE POWER SHELL AND TRIED TO PREPARE, BUT THE DEVICE DID NOT RESPOND AT ALL. REPEATEDLY REMOVING AND REINSERTING THE HANDLE TO THE BATTERY CHARGER TO RECOGNIZE THE TROUBLE, THE CHARGING INDICATOR WORKED USUALLY, BUT THE DEVICE DID NOT START-UP. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES TO BE VENTED.
06-10-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE POWER HANDLE DID NOT START WHEN REMOVED FROM THE CHARGER. THE CHARGING INDICATOR WAS REPEATEDLY ILLUMINATING, INDICATING THAT THE CHARGING WAS COMPLETED. THE HANDLE WAS PLACED ON THE POWER SHELL AND TRIED TO PREPARE, BUT THE DEVICE DID NOT RESPOND AT ALL. REPEATEDLY REMOVING AND REINSERTING THE HANDLE TO THE BATTERY CHARGER TO RECOGNIZE THE TROUBLE, THE CHARGING INDICATOR WORKED USUALLY, BUT THE DEVICE DID NOT START-UP. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES TO BE VENTED.
09-12-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
11-06-2020 | Injury | us surgical puerto rico | 05-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER THE RECTUM WAS DILATED AND STAPLER WAS PASSED DOWN INTO APPROPRIATE POSITION, THE SPIKE WAS DEPLOYED AND CONNECTED TO THE ANVIL AND WAS THEN CLOSED AND FIRED. UPON REMOVAL OF THE STAPLER, THE ANASTOMOSIS FELL APART. SURGEON COMPLETED THE PROCEDURE WITH THE USE OF ANOTHER REPLACEMENT STAPLER.
11-06-2020 | Injury | us surgical puerto rico | 05-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER THE RECTUM WAS DILATED AND STAPLER WAS PASSED DOWN INTO APPROPRIATE POSITION, THE SPIKE WAS DEPLOYED AND CONNECTED TO THE ANVIL AND WAS THEN CLOSED AND FIRED. UPON REMOVAL OF THE STAPLER, THE ANASTOMOSIS FELL APART. SURGEON COMPLETED THE PROCEDURE WITH THE USE OF ANOTHER REPLACEMENT STAPLER.
09-12-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE COMPONENT BROKE OFF FROM THE RELOAD, AND A PURPLE FRAGMENT WAS FOUND IN THE PACKAGE PRIOR TO OPENING IT. THE DEFECTIVE PRODUCT WAS REPLACED WITH A NEW PRODUCT THAT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA | Device Contamination with Chemical or Other Material | A180104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE COMPONENT BROKE OFF FROM THE RELOAD, AND A PURPLE FRAGMENT WAS FOUND IN THE PACKAGE PRIOR TO OPENING IT. THE DEFECTIVE PRODUCT WAS REPLACED WITH A NEW PRODUCT THAT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN ABDOMINAL THORACOTOMY-PULMONARY RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-11-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SIGMOID RESECTION, DURING THE TRANSECTION OF THE SIGMOID COLON, A PIECE OF EITHER THE STAPLER OR A RELOAD FELL OFF ON THE BACK TABLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | TA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN EXPLORATORY LAPAROTOMY, TUMOR DEBULKING, TOTAL ABDOMINAL HYSTERECTOMY WITH BILATERAL SALPINGO-OOPHORECTOMY, PARTIAL COLECTOMY, THERE WERE TWO FIRINGS ON THE RECTUM AND ONE ON THE TRANSVERSE COLON, THE TWO RELOADS AND STAPLER PARTIALLY FIRED, THERE WAS POOR STAPLE FORMATION. THE USER OVERSEW THE STAPLE LINE TO COMPLETE THE CASE.
09-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | TA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN EXPLORATORY LAPAROTOMY, TUMOR DEBULKING, TOTAL ABDOMINAL HYSTERECTOMY WITH BILATERAL SALPINGO-OOPHORECTOMY, PARTIAL COLECTOMY, THERE WERE TWO FIRINGS ON THE RECTUM AND ONE ON THE TRANSVERSE COLON, THE TWO RELOADS AND STAPLER PARTIALLY FIRED, THERE WAS POOR STAPLE FORMATION. THE USER OVERSEW THE STAPLE LINE TO COMPLETE THE CASE.
09-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | TA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN EXPLORATORY LAPAROTOMY, TUMOR DEBULKING, TOTAL ABDOMINAL HYSTERECTOMY WITH BILATERAL SALPINGO-OOPHORECTOMY, PARTIAL COLECTOMY, THERE WERE TWO FIRINGS ON THE RECTUM AND ONE ON THE TRANSVERSE COLON, THE TWO RELOADS AND STAPLER PARTIALLY FIRED, THERE WAS POOR STAPLE FORMATION. THE USER OVERSEW THE STAPLE LINE TO COMPLETE THE CASE.
09-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | TA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN EXPLORATORY LAPAROTOMY, TUMOR DEBULKING, TOTAL ABDOMINAL HYSTERECTOMY WITH BILATERAL SALPINGO-OOPHORECTOMY, PARTIAL COLECTOMY, THERE WERE TWO FIRINGS ON THE RECTUM AND ONE ON THE TRANSVERSE COLON, THE TWO RELOADS AND STAPLER PARTIALLY FIRED, THERE WAS POOR STAPLE FORMATION. THE USER OVERSEW THE STAPLE LINE TO COMPLETE THE CASE.
09-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | TA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN EXPLORATORY LAPAROTOMY, TUMOR DEBULKING, TOTAL ABDOMINAL HYSTERECTOMY WITH BILATERAL SALPINGO-OOPHORECTOMY, PARTIAL COLECTOMY, THERE WERE TWO FIRINGS ON THE RECTUM AND ONE ON THE TRANSVERSE COLON, THE TWO RELOADS AND STAPLER PARTIALLY FIRED, THERE WAS POOR STAPLE FORMATION. THE USER OVERSEW THE STAPLE LINE TO COMPLETE THE CASE.
09-12-2020 | Injury | us surgical puerto rico | 05-01-2021 | TA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN EXPLORATORY LAPAROTOMY, TUMOR DEBULKING, TOTAL ABDOMINAL HYSTERECTOMY WITH BILATERAL SALPINGO-OOPHORECTOMY, PARTIAL COLECTOMY, THERE WERE TWO FIRINGS ON THE RECTUM AND ONE ON THE TRANSVERSE COLON, THE TWO RELOADS AND STAPLER PARTIALLY FIRED, THERE WAS POOR STAPLE FORMATION. THE USER OVERSEW THE STAPLE LINE TO COMPLETE THE CASE.
01-03-2021 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THE HANDLE AND LOADING UNIT STICKS, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-03-2021 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THE HANDLE AND LOADING UNIT STICKS, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-02-2021 | Malfunction | us surgical puerto rico | 05-01-2021 | GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED. REPLACED THE DEVICES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-02-2021 | Malfunction | us surgical puerto rico | 05-01-2021 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED. REPLACED THE DEVICES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-01-2021 | Malfunction | us surgical puerto rico | 05-01-2021 | UNKNOWN GIA PRODUCT | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED.
09-10-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE WAS FULLY CHARGED (BATTERY CHARGER INDICATOR WAS FOUND TO BE ILLUMINATING GREEN) AND INSTALLED INTO THE POWER SHELL, BUT WAS UNUSABLE DUE TO THE STATED ERROR DISPLAY (THE ERROR MARK FOR A YELLOW TRIANGLE AND THE ERROR MARK FOR THE BATTERY INDICATOR IN RED WITH A SLASH WERE DISPLAYED). THE SAME SITUATION OCCURRED EVEN AFTER RESTARTING WITH A POWER APPROACH (BOTH PRESSES OF THE GREEN BUTTON). AFTER COLLECTING THE DEVICE, AN ATTEMPT WAS MADE TO DUPLICATE THE EVENT AGAIN, BUT IT WAS FOUND THAT IT COULD NOT BE CHARGED PROPERLY (USING THE SALES DEMO CHARGER). ALTHOUGH THE DEVICE WAS COMPLETELY UNCHARGED, IT SWITCHED TO ILLUMINATING GREEN AFTER ABOUT 3 MINUTES; AFTERWARDS, IT DID NOT RESTART DESPITE REMOVING IT FROM THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 -MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
09-10-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE WAS FULLY CHARGED (BATTERY CHARGER INDICATOR WAS FOUND TO BE ILLUMINATING GREEN) AND INSTALLED INTO THE POWER SHELL, BUT WAS UNUSABLE DUE TO THE STATED ERROR DISPLAY (THE ERROR MARK FOR A YELLOW TRIANGLE AND THE ERROR MARK FOR THE BATTERY INDICATOR IN RED WITH A SLASH WERE DISPLAYED). THE SAME SITUATION OCCURRED EVEN AFTER RESTARTING WITH A POWER APPROACH (BOTH PRESSES OF THE GREEN BUTTON). AFTER COLLECTING THE DEVICE, AN ATTEMPT WAS MADE TO DUPLICATE THE EVENT AGAIN, BUT IT WAS FOUND THAT IT COULD NOT BE CHARGED PROPERLY (USING THE SALES DEMO CHARGER). ALTHOUGH THE DEVICE WAS COMPLETELY UNCHARGED, IT SWITCHED TO ILLUMINATING GREEN AFTER ABOUT 3 MINUTES; AFTERWARDS, IT DID NOT RESTART DESPITE REMOVING IT FROM THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 -MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
20-05-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE HAD NO VISUALS. TROUBLESHOOTING WAS DONE FOR THE HANDLE THAT HAD NO VISUALS, BUT THE HANDLE STILL HAD NO POWER. THE SALES REPRESENTATIVE WAS NOTIFIED TO RETURN THE HANDLE WITH NO VISUALS AND A DEMONSTRATION DEVICE WAS REQUESTED. HOWEVER, THE HANDLE THAT WAS RETURNED WAS NOT A DEMONSTRATION HANDLE AND STILL INDICATED TO BE A PROPERTY OF THE FACILITY. IT WAS ALSO OUT OF WARRANTY. THERE WAS A CHANCE THAT THE FACILITY HAD A HANDLE ALLOCATED TO A SALES REPRESENTATIVE AND THE HANDLES GOT MIXED UP. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE BATTERY WAS INVESTIGATED TO DETERMINE THE CAUSE OF THE INABILITY TO CHARGE AND CUV PERMANENT FAILURE CONDITION. USING A POWER SUPPLY, VOLTAGE WAS APPLIED TO THE BATTERY CELLS DIRECTLY, HOWEVER NO CURRENT WOULD ENTER THE CELLS. THIS SHOWS THAT ONE OR BOTH OF THE CELLS HAD AN OPEN CURRENT INTERRUPT DEVICE (CID). AN OPEN CELL CID WOULD PREVENT THE HANDLE FROM RECEIVING CHARGE OR POWERING ON, AND LEAD TO THE CUV CONDITION, AS THE VOLTAGE POTENTIAL IN THE CELL WOULD GO CLOSE TO ZERO.
20-05-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE HAD NO VISUALS. TROUBLESHOOTING WAS DONE FOR THE HANDLE THAT HAD NO VISUALS, BUT THE HANDLE STILL HAD NO POWER. THE SALES REPRESENTATIVE WAS NOTIFIED TO RETURN THE HANDLE WITH NO VISUALS AND A DEMONSTRATION DEVICE WAS REQUESTED. HOWEVER, THE HANDLE THAT WAS RETURNED WAS NOT A DEMONSTRATION HANDLE AND STILL INDICATED TO BE A PROPERTY OF THE FACILITY. IT WAS ALSO OUT OF WARRANTY. THERE WAS A CHANCE THAT THE FACILITY HAD A HANDLE ALLOCATED TO A SALES REPRESENTATIVE AND THE HANDLES GOT MIXED UP. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE BATTERY WAS INVESTIGATED TO DETERMINE THE CAUSE OF THE INABILITY TO CHARGE AND CUV PERMANENT FAILURE CONDITION. USING A POWER SUPPLY, VOLTAGE WAS APPLIED TO THE BATTERY CELLS DIRECTLY, HOWEVER NO CURRENT WOULD ENTER THE CELLS. THIS SHOWS THAT ONE OR BOTH OF THE CELLS HAD AN OPEN CURRENT INTERRUPT DEVICE (CID). AN OPEN CELL CID WOULD PREVENT THE HANDLE FROM RECEIVING CHARGE OR POWERING ON, AND LEAD TO THE CUV CONDITION, AS THE VOLTAGE POTENTIAL IN THE CELL WOULD GO CLOSE TO ZERO.
09-12-2020 | Malfunction | celestica de monterrey s.a de c.v | 05-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC WHIPPLE PROCEDURE, DURING DUODENUM RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THE SAME ISSUE OCCURRED WITH ANOTHER RELOAD. A MANUAL STAPLER AND A NEW RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | IDRIVE ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC WHIPPLE PROCEDURE, DURING DUODENUM RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THE SAME ISSUE OCCURRED WITH ANOTHER RELOAD. A MANUAL STAPLER AND A NEW RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Injury | covidien lp llc north haven | 05-01-2021 | SIGNIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON THE FOURTH FIRING IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS AN INCOMPLETE STAPLE LINE ON THE CENTRAL PART OF THE ANASTOMOSIS. IT WAS ALSO NOTED THE STAPLES WERE OPEN AND WAS NOT ABLE TO FORM. IT WAS ALSO REPORTED THAT THE TISSUE WAS SCRUNCHED UP RESULTING IN HOLE IN STOMACH. SURGEON SUTURED THE TISSUE TO CLOSE THE ANASTOMOSIS. THE PROCEDURE WAS EXTENDED BY 4 HOURS DUE TO DEVICE PROBLEM. PATIENT HAD AN EXTENDED HOSPITAL STAY AND IS CURRENTLY STAYING AT THE HOSPITAL AS OF THIS REPORT.
09-12-2020 | Injury | covidien lp llc north haven | 05-01-2021 | SIGNIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON THE FOURTH FIRING IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS AN INCOMPLETE STAPLE LINE ON THE CENTRAL PART OF THE ANASTOMOSIS. IT WAS ALSO NOTED THE STAPLES WERE OPEN AND WAS NOT ABLE TO FORM. IT WAS ALSO REPORTED THAT THE TISSUE WAS SCRUNCHED UP RESULTING IN HOLE IN STOMACH. SURGEON SUTURED THE TISSUE TO CLOSE THE ANASTOMOSIS. THE PROCEDURE WAS EXTENDED BY 4 HOURS DUE TO DEVICE PROBLEM. PATIENT HAD AN EXTENDED HOSPITAL STAY AND IS CURRENTLY STAYING AT THE HOSPITAL AS OF THIS REPORT.
09-12-2020 | Injury | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON THE FOURTH FIRING IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS AN INCOMPLETE STAPLE LINE ON THE CENTRAL PART OF THE ANASTOMOSIS. IT WAS ALSO NOTED THE STAPLES WERE OPEN AND WAS NOT ABLE TO FORM. IT WAS ALSO REPORTED THAT THE TISSUE WAS SCRUNCHED UP RESULTING IN HOLE IN STOMACH. SURGEON SUTURED THE TISSUE TO CLOSE THE ANASTOMOSIS. THE PROCEDURE WAS EXTENDED BY 4 HOURS DUE TO DEVICE PROBLEM. PATIENT HAD AN EXTENDED HOSPITAL STAY AND IS CURRENTLY STAYING AT THE HOSPITAL AS OF THIS REPORT.
05-11-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, WHILE STAPLING THE LUNG, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-11-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, WHILE STAPLING THE LUNG, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-11-2020 | Malfunction | ethicon endo-surgery | 05-01-2021 | POWERED PLUS LONG 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: SEE ATTACHED USER FACILITY MEDWATCH FORM, MW# (B)(4).
01-11-2020 | Malfunction | ethicon endo-surgery | 05-01-2021 | POWERED PLUS LONG 45 | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: SEE ATTACHED USER FACILITY MEDWATCH FORM, MW# (B)(4).
01-11-2020 | Malfunction | ethicon endo-surgery | 05-01-2021 | POWERED PLUS LONG 45 | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: SEE ATTACHED USER FACILITY MEDWATCH FORM, MW# (B)(4).
18-10-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY, THE STAPLES BURST OUT WHEN FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-10-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY, THE STAPLES BURST OUT WHEN FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-11-2020 | Malfunction | davis & geck caribe | 05-01-2021 | UNKNOWN ENDO STITCH SULU | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE WAS NOT PASSING THE NEEDLE WELL. WHEN THE SALES REP EXAMINED THE DEVICE, IT WAS FOUND THAT A NEEDLE WAS BROKEN OFF IN THE JAWS OF THE INSTRUMENT. THERE WAS NO PATIENT INJURY.
28-11-2020 | Malfunction | davis & geck caribe | 05-01-2021 | UNKNOWN ENDO STITCH SULU | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE WAS NOT PASSING THE NEEDLE WELL. WHEN THE SALES REP EXAMINED THE DEVICE, IT WAS FOUND THAT A NEEDLE WAS BROKEN OFF IN THE JAWS OF THE INSTRUMENT. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION PROCEDURE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. 2ND RELOAD WAS USED HOWEVER SAME ISSUE OCCURRED IT STILL COULD NOT BE FIRE. THE THIRD RELOAD WAS USED AND ABLE TO BE FIRED NORMALLY. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION PROCEDURE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. 2ND RELOAD WAS USED HOWEVER SAME ISSUE OCCURRED IT STILL COULD NOT BE FIRE. THE THIRD RELOAD WAS USED AND ABLE TO BE FIRED NORMALLY. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, HANDLE WAS BLOCKED. GREEN BUTTON WAS PRESSED BUT DEVICE WOULD NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, HANDLE WAS BLOCKED. GREEN BUTTON WAS PRESSED BUT DEVICE WOULD NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-11-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS UNABLE TO FIRE, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
30-11-2020 | Malfunction | us surgical puerto rico | 05-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS UNABLE TO FIRE, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
02-12-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE, WHEN CLOSING THE RELOAD ON THE STOMACH TISSUE, THE JAWS WOULD NOT CLOSE. THE DEVICE WOULD NOT FIRE EVEN THOUGH GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED. ANOTHER RELOAD WAS USED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | covidien lp llc north haven | 05-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE, WHEN CLOSING THE RELOAD ON THE STOMACH TISSUE, THE JAWS WOULD NOT CLOSE. THE DEVICE WOULD NOT FIRE EVEN THOUGH GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED. ANOTHER RELOAD WAS USED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | smith & nephew | 05-01-2021 | TENDON ANCHORS 8 | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A REGENETEN CASE, A NEW BOX OF TENDON STAPLES WAS OPENED; HOWEVER, THE STERILE PACKAGING SEEMED AS THOUGH IT HAD NEVER BEEN SEALED. THE MALFUNCTION WAS SOLVED WITH NO DELAYS OR PATIENT HARM USING A BACK UP DEVICE. ALL AVAILABLE INFORMATION HAS BEEN DISCLOSED. IF ADDITIONAL INFORMATION SHOULD BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY.
27-10-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WOULD NOT CHARGE. THERE WAS NO PATIENT INVOLVEMENT.
27-10-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WOULD NOT CHARGE. THERE WAS NO PATIENT INVOLVEMENT.
01-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE USING THE DEVICE ON THE RIGHT LIVER PEDICLE IN A RIGHT LIVER TUMOR RESECTION, AFTER THE RELOAD AND HANDLE WERE CONNECTED, THE SURGEON CLAMPED THE DEVICE AT THE RIGHT HEPATIC PEDICLE AND ADJUSTED THE ANGLE, THEN PULLED THE BLACK RETURN BUTTON, AND THE JAWS OF THE RELOAD CANNOT BE OPENED. SURGEON REPLACED THE RELOAD AND HANDLE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
01-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE USING THE DEVICE ON THE RIGHT LIVER PEDICLE IN A RIGHT LIVER TUMOR RESECTION, AFTER THE RELOAD AND HANDLE WERE CONNECTED, THE SURGEON CLAMPED THE DEVICE AT THE RIGHT HEPATIC PEDICLE AND ADJUSTED THE ANGLE, THEN PULLED THE BLACK RETURN BUTTON, AND THE JAWS OF THE RELOAD CANNOT BE OPENED. SURGEON REPLACED THE RELOAD AND HANDLE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
09-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | RELIATACK | Difficult or Delayed Positioning | A150203 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, WHEN FIXING THE HERNIA MESH, THERE WAS PROBLEM WITH THE CARTRIDGE ATTACHMENT, AND THE FIRING WAS FELT TO BE HARD WHEN TACKING. A NEW TACKER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, WHEN FIXING THE HERNIA MESH, THERE WAS PROBLEM WITH THE CARTRIDGE ATTACHMENT, AND THE FIRING WAS FELT TO BE HARD WHEN TACKING. A NEW TACKER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-10-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC PANCREAS RESECTION, THE DEVICE FIRED, BUT THE SURGEON FELT THAT THE HANDLE BROKE. THERE WAS A BREAKING-LIKE SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INSTRUMENT FIRING KNOBS WERE RETRACTED, AND THE ARTICULATION LEVER WAS IN NEUTRAL POSITION. VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
07-10-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC PANCREAS RESECTION, THE DEVICE FIRED, BUT THE SURGEON FELT THAT THE HANDLE BROKE. THERE WAS A BREAKING-LIKE SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INSTRUMENT FIRING KNOBS WERE RETRACTED, AND THE ARTICULATION LEVER WAS IN NEUTRAL POSITION. VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
20-10-2020 | Injury | covidien lp llc north haven | 06-01-2021 | UNKNOWN TRI STAPLE PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN FEBRUARY 2010 AND JUNE 2015, A RETROSPECTIVE STUDY ANALYZED 358 PATIENTS WHO UNDERWENT BARIATRIC PROCEDURES. THE PURPOSE OF THE STUDY WAS TO EVALUATE THE LONG-TERM WEIGHT LOSS RESULTS, REMISSION OF COMORBIDITIES AND NUTRITIONAL DEFICIENCIES. IN ALL CASES, THREE-STAPLE CARTRIDGES WERE USED TO SUTURE THE GASTRIC SAC. POSTOPERATIVE COMPLICATIONS INCLUDED: TWO STAPLE LINE LEAKS AFTER SLEEVE GASTRECTOMY, ONE GASTRO-JEJUNAL ANASTOMOTIC LEAK, ONE HEMOPERITONEUM AFTER ROUX-EN-Y GASTRIC BYPASS AND ONE ANASTOMOTIC LEAK THEN TWO CASES OF INTRAPERITONEAL BLEEDINGS AFTER ONE-ANASTOMOSIS GASTRIC BYPASS. ARTICLE:LONG-TERM WEIGHT LOSS RESULTS, REMISSION OF COMORBIDITIES AND NUTRITIONAL DEFICIENCIES OF SLEEVE GASTRECTOMY (SG), ROUX-EN-Y GASTRIC BYPASS (RYGB) AND ONE-ANASTOMOSIS GASTRIC BYPASS (OAGB) ON TYPE 2 DIABETIC (T2D) PATIENTS AUTHOR: JOSE-MARIA JIMENEZ YEAR: 2020 PUBLICATION: INT. J. ENVIRON. RES. PUBLIC HEALTH 2020, 17, 764.
20-10-2020 | Injury | covidien lp llc north haven | 06-01-2021 | UNKNOWN TRI STAPLE PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN FEBRUARY 2010 AND JUNE 2015, A RETROSPECTIVE STUDY ANALYZED 358 PATIENTS WHO UNDERWENT BARIATRIC PROCEDURES. THE PURPOSE OF THE STUDY WAS TO EVALUATE THE LONG-TERM WEIGHT LOSS RESULTS, REMISSION OF COMORBIDITIES AND NUTRITIONAL DEFICIENCIES. IN ALL CASES, THREE-STAPLE CARTRIDGES WERE USED TO SUTURE THE GASTRIC SAC. POSTOPERATIVE COMPLICATIONS INCLUDED: TWO STAPLE LINE LEAKS AFTER SLEEVE GASTRECTOMY, ONE GASTRO-JEJUNAL ANASTOMOTIC LEAK, ONE HEMOPERITONEUM AFTER ROUX-EN-Y GASTRIC BYPASS AND ONE ANASTOMOTIC LEAK THEN TWO CASES OF INTRAPERITONEAL BLEEDINGS AFTER ONE-ANASTOMOSIS GASTRIC BYPASS. ARTICLE:LONG-TERM WEIGHT LOSS RESULTS, REMISSION OF COMORBIDITIES AND NUTRITIONAL DEFICIENCIES OF SLEEVE GASTRECTOMY (SG), ROUX-EN-Y GASTRIC BYPASS (RYGB) AND ONE-ANASTOMOSIS GASTRIC BYPASS (OAGB) ON TYPE 2 DIABETIC (T2D) PATIENTS AUTHOR: JOSE-MARIA JIMENEZ YEAR: 2020 PUBLICATION: INT. J. ENVIRON. RES. PUBLIC HEALTH 2020, 17, 764.
09-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AT PARENCHYMAL RESECTION DURING AN OPEN DISTAL GASTRECTOMY, THE DEVICE DID NOT FIRE. THE SAME EVENT OCCURRED ON TWO OTHER LOADING UNITS. THE NEW PRODUCT WAS TAKEN AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE LOAD CAME OUT OF ITS PACKAGING PARTIALLY FIRED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
17-01-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING STAPLING IN A BARIATRIC SLEEVE SURGERY/BYPASS, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE WAS NOT ABLE TO FIRE. TWO HANDLES WERE USED AND HAD THE SAME ISSUE. SURGEON USED ANOTHER HANDLE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Overheating of Device | A1005 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, EVEN THOUGH THE HANDLE WAS FULLY CHARGED, 1-2 HOURS AFTER THE SECOND FIRING THE THIRD FIRING WAS ATTEMPTED TO BE PERFORMED, IT WAS OUT OF CHARGE AND WAS UNABLE TO BE FIRED. THE HANDLE WAS SET ON THE CHARGER AND WAS HOTTER THAN USUAL. ANOTHER HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, EVEN THOUGH THE HANDLE WAS FULLY CHARGED, 1-2 HOURS AFTER THE SECOND FIRING THE THIRD FIRING WAS ATTEMPTED TO BE PERFORMED, IT WAS OUT OF CHARGE AND WAS UNABLE TO BE FIRED. THE HANDLE WAS SET ON THE CHARGER AND WAS HOTTER THAN USUAL. ANOTHER HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING OF THE HANDLE AND LOADING THE STAPLES IN A GASTRIC ANASTOMOSIS, THE HANDLE COULDN'T BE FIRED AFTER LOADING THE STAPLES, AND THE SAME EVENT OCCURRED AFTER REPLACING THE STAPLES. THE OPERATION WAS COMPLETED AFTER REPLACING THE HANDLE. NO PATIENT INJURY.
09-12-2020 | Injury | us surgical puerto rico | 06-01-2021 | GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LEFT COLECTOMY, ON THE THIRD FIRING OF THE STAPLER, THE JAWS LOCKED ON THE TISSUE AFTER THE STAPLER WAS FIRED. THE SURGEON PULLED THE BLACK TOGGLES BACK AND PRESSED THE OPEN BUTTON, BUT THE STAPLER DID NOT OPEN. AFTER SEVERAL ATTEMPTS AND MINUTES, THE STAPLER DID OPEN. THE SURGEON OPENED A NEW STAPLER AND TRANSECTED THE STAPLE LINE ON THE PATIENT SIDE BECAUSE THE STAPLE LINE WAS DAMAGED DUE TO THE FACT THAT THE STAPLER WAS LOCKED ON TISSUE FOR SEVERAL MINUTES. UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE WERE NOTED.
09-12-2020 | Injury | us surgical puerto rico | 06-01-2021 | GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LEFT COLECTOMY, ON THE THIRD FIRING OF THE STAPLER, THE JAWS LOCKED ON THE TISSUE AFTER THE STAPLER WAS FIRED. THE SURGEON PULLED THE BLACK TOGGLES BACK AND PRESSED THE OPEN BUTTON, BUT THE STAPLER DID NOT OPEN. AFTER SEVERAL ATTEMPTS AND MINUTES, THE STAPLER DID OPEN. THE SURGEON OPENED A NEW STAPLER AND TRANSECTED THE STAPLE LINE ON THE PATIENT SIDE BECAUSE THE STAPLE LINE WAS DAMAGED DUE TO THE FACT THAT THE STAPLER WAS LOCKED ON TISSUE FOR SEVERAL MINUTES. UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE WERE NOTED.
10-12-2020 | Malfunction | celestica de monterrey s.a de c.v | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO ANASTOMOSES DUODENUM AND STOMACH ON A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE DEVICE DID NOT FIRE. THE SAME EVENT OCCURRED TWICE ON THE RELOAD. A MANUAL HANDLE AND ANOTHER RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-01-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEVICE USAGE IN A RECTAL CANCER, THE BOWEL COULD NOT BE CLOSED DUE TO THE DEVICE FAILURE. THE DEVICE WAS NOT ABLE TO FIRE. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
10-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | IDRIVE ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO ANASTOMOSES DUODENUM AND STOMACH ON A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE DEVICE DID NOT FIRE. THE SAME EVENT OCCURRED TWICE ON THE RELOAD. A MANUAL HANDLE AND ANOTHER RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | UNKNOWN EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED TO ANASTOMOSES DUODENUM AND STOMACH ON A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE DEVICE DID NOT FIRE. THE SAME EVENT OCCURRED TWICE ON THE RELOAD. A MANUAL HANDLE AND ANOTHER RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AT PARENCHYMAL RESECTION DURING AN OPEN DISTAL GASTRECTOMY, THE DEVICE DID NOT FIRE. THE SAME EVENT OCCURRED ON TWO OTHER LOADING UNITS. THE NEW PRODUCT WAS TAKEN AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AT PARENCHYMAL RESECTION DURING AN OPEN DISTAL GASTRECTOMY, THE DEVICE DID NOT FIRE. THE SAME EVENT OCCURRED ON TWO OTHER LOADING UNITS. THE NEW PRODUCT WAS TAKEN AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Injury | us surgical puerto rico | 06-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, AFTER DEEP ENDOMETRIOSIS/ INFILTRATING, PATIENT LOSS 4 HEMOGLOBIN POINTS. PATIENT WAS BLEEDING AND BLOOD LOSS WAS MORE THAN 500CC THAT REQUIRED BLOOD TRANSFUSION. PATIENT HOSPITAL STAY WAS ALSO EXTENDED.
09-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | UNKNOWN GIA DST PRODUCT | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE LOAD CAME OUT OF ITS PACKAGING PARTIALLY FIRED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, WHEN ON STOMACH, THERE WAS POOR STAPLE FORMATION. STAPLE GOT STUCK IN THE LOAD AND DID NOT COME OUT. THE SURGEON REPLACED THE RELOAD WITH ANOTHER ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-10-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE CHARGING PRE-OPERATIVELY, THE DEVICE FAILED TO CHARGE. A DEMONSTRATION DEVICE WAS LENT OUT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
15-10-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE CHARGING PRE-OPERATIVELY, THE DEVICE FAILED TO CHARGE. A DEMONSTRATION DEVICE WAS LENT OUT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
10-01-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL CANCER CLOSED BOWEL, THE BOWEL COULD NOT BE CLOSED BECAUSE THE DEVICE WAS NOT ABLE TO BE FIRED. SURGEON REPLACED THE HANDLE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-10-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE HANDLE WAS TRIED IN 2 OF THE EXISTING CHARGING STATIONS BUT THE LED TURNED RED IMMEDIATELY. IT WAS ALSO NOT POSSIBLE RESETTING THE DEVICE SINCE THE DISPLAY WAS BLACK. BOTH CHARGERS WORK WITH A NEW POWERED HANDLE WITHOUT PROBLEMS. THERE WAS NO PATIENT INVOLVEMENT.
14-10-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE HANDLE WAS TRIED IN 2 OF THE EXISTING CHARGING STATIONS BUT THE LED TURNED RED IMMEDIATELY. IT WAS ALSO NOT POSSIBLE RESETTING THE DEVICE SINCE THE DISPLAY WAS BLACK. BOTH CHARGERS WORK WITH A NEW POWERED HANDLE WITHOUT PROBLEMS. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE BRONCHUS OF THE UPPER LOBE OF THE RIGHT LUNG ON A VIDEO ASSISTED THORACOSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE COULD NOT BE FIRED. ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOW RECTAL SURGERY, AFTER SUCCESSFUL FIRING, WHEN THE RELOAD WAS CHANGED, THE DEVICE WAS NOT ABLE TO BE FIRED. SURGEON REPLACED THE RELOAD AND WAS ABLE TO FIRE SUCCESSFULLY. NO PATIENT INJURY.
25-08-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A PROXIMAL PANCREATICODUODENECTOMY, DURING THE FIRST FIRING, WHEN THE HANDLE WAS INSERTED INTO THE POWER SHELL, THE HANDLE DISPLAYED 0 LIVES LEFT AND BECAME UNUSABLE. THE DEVICE WAS RESET BUT THE SAME ISSUE OCCURRED. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THE HANDLE WAS SET TO THE CHARGER AGAIN AND TRIED TO BE CHARGED, BUT THE SCREEN DID NOT START UP AT ALL. THE CHARGER WORKED NORMALLY WITH OTHER HANDLES. THERE WAS NO PATIENT INJURY.
25-08-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Premature End-of-Life Indicator | A160602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A PROXIMAL PANCREATICODUODENECTOMY, DURING THE FIRST FIRING, WHEN THE HANDLE WAS INSERTED INTO THE POWER SHELL, THE HANDLE DISPLAYED 0 LIVES LEFT AND BECAME UNUSABLE. THE DEVICE WAS RESET BUT THE SAME ISSUE OCCURRED. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THE HANDLE WAS SET TO THE CHARGER AGAIN AND TRIED TO BE CHARGED, BUT THE SCREEN DID NOT START UP AT ALL. THE CHARGER WORKED NORMALLY WITH OTHER HANDLES. THERE WAS NO PATIENT INJURY.
25-08-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A PROXIMAL PANCREATICODUODENECTOMY, DURING THE FIRST FIRING, WHEN THE HANDLE WAS INSERTED INTO THE POWER SHELL, THE HANDLE DISPLAYED 0 LIVES LEFT AND BECAME UNUSABLE. THE DEVICE WAS RESET BUT THE SAME ISSUE OCCURRED. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THE HANDLE WAS SET TO THE CHARGER AGAIN AND TRIED TO BE CHARGED, BUT THE SCREEN DID NOT START UP AT ALL. THE CHARGER WORKED NORMALLY WITH OTHER HANDLES. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC PROCEDURE, AFTER A SUCCESSFUL FIRING, AFTER RELOADING, THE DEVICE WOULD NOT FIRE. SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
10-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE, UPON GASTRIC SLEEVE, FIRING ON STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. A NEW RELOAD WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Injury | us surgical puerto rico | 06-01-2021 | EEA | Loss of or Failure to Bond | A040102 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, WHEN ON RECTAL ANASTOMOSES, THE ANVIL WOULD NOT ENGAGE ONTO THE DEVICE CENTER ROD. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. THE SURGEON HAD TO REINSERT A NEW STAPLER AND RESECT THE ORIGINAL HOLE MADE BY THE CENTER ROD OF THE FIRST DEVICE. THERE WAS A 45 MINUTE DELAY ON THE OPERATION.
09-12-2020 | Injury | us surgical puerto rico | 06-01-2021 | EEA | Loss of or Failure to Bond | A040102 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, WHEN ON RECTAL ANASTOMOSES, THE ANVIL WOULD NOT ENGAGE ONTO THE DEVICE CENTER ROD. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. THE SURGEON HAD TO REINSERT A NEW STAPLER AND RESECT THE ORIGINAL HOLE MADE BY THE CENTER ROD OF THE FIRST DEVICE. THERE WAS A 45 MINUTE DELAY ON THE OPERATION.
02-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
08-03-2019 | Injury | covidien lp llc north haven | 06-01-2021 | UNKNOWN ENDO STITCH | Component or Accessory Incompatibility | A1701 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NISSEN PROCEDURE, WHILE THE DEVICE WAS BEING USED TO CLOSE THE HIATAL HERNIA AND WRAP OF STOMACH, THE NEEDLE FELL OUT ONCE AND DISENGAGED INTO THE PATIENT'S CAVITY. AN X-RAY WAS PERFORMED TO LOCATE THE COMPONENT AND WAS RETRIEVED BY A GRASPER. THE SURGEON CHANGED THE HANDLE THREE TIMES BECAUSE IT WAS VERY DIFFICULTY TO LOAD AND UNLOAD.
08-03-2019 | Injury | covidien lp llc north haven | 06-01-2021 | UNKNOWN ENDO STITCH | Detachment of Device or Device Component | A0501 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NISSEN PROCEDURE, WHILE THE DEVICE WAS BEING USED TO CLOSE THE HIATAL HERNIA AND WRAP OF STOMACH, THE NEEDLE FELL OUT ONCE AND DISENGAGED INTO THE PATIENT'S CAVITY. AN X-RAY WAS PERFORMED TO LOCATE THE COMPONENT AND WAS RETRIEVED BY A GRASPER. THE SURGEON CHANGED THE HANDLE THREE TIMES BECAUSE IT WAS VERY DIFFICULTY TO LOAD AND UNLOAD.
08-03-2019 | Injury | covidien lp llc north haven | 06-01-2021 | UNKNOWN ENDO STITCH | Human-Device Interface Problem | A22 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NISSEN PROCEDURE, WHILE THE DEVICE WAS BEING USED TO CLOSE THE HIATAL HERNIA AND WRAP OF STOMACH, THE NEEDLE FELL OUT ONCE AND DISENGAGED INTO THE PATIENT'S CAVITY. AN X-RAY WAS PERFORMED TO LOCATE THE COMPONENT AND WAS RETRIEVED BY A GRASPER. THE SURGEON CHANGED THE HANDLE THREE TIMES BECAUSE IT WAS VERY DIFFICULTY TO LOAD AND UNLOAD.
02-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
02-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
30-11-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
30-11-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
07-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY, THE DEVICE'S RELOAD COULD NOT BE FIRED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-11-2020 | Injury | us surgical puerto rico | 06-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS, THE DEVICE HAD AN UNKNOWN MALFUNCTION THAT RESULTED TO A BLOOD LOSS OF MORE THAN 500 CC, HEMOGLOBIN 13 TO 7 G/DL. MEDICAL TREATMENT WAS NEEDED TO RESOLVE THE ISSUE. THE EVENT RESULTED TO AN EXTENDED HOSPITALIZATION.
27-11-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE KNIFE DID NOT ADVANCE AND DID NOT CUT THE TISSUE.
26-11-2020 | Injury | us surgical puerto rico | 06-01-2021 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS, THE DEVICE HAD AN UNKNOWN MALFUNCTION THAT RESULTED TO A BLOOD LOSS OF MORE THAN 500CC, HEMOGLOBIN 13 TO 7 G/DL. MEDICAL TREATMENT WAS NEEDED TO RESOLVE THE ISSUE. THE EVENT RESULTED TO ANEXTENDED HOSPITALIZATION.
27-11-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WAS DISCOVERED TO BE PRE-FIRED AFTER FIRING WAS ATTEMPTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
27-11-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WAS DISCOVERED TO BE PRE-FIRED AFTER FIRING WAS ATTEMPTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
27-11-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WAS DISCOVERED TO BE PRE-FIRED AFTER FIRING WAS ATTEMPTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
27-11-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Expulsion | A010401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WAS DISCOVERED TO BE PRE-FIRED AFTER FIRING WAS ATTEMPTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
27-11-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WAS DISCOVERED TO BE PRE-FIRED AFTER FIRING WAS ATTEMPTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
07-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN COLECTOMY, THE DEVICE DID NOT FIRE. IT WOULD NOT CLOSE AROUND BOWEL WHEN TRIED TO FIRE. STAPLER GOT STUCK AND DID NOT WORK WITH THE ORIGINAL RELOAD IN THE DEVICE. ANOTHER STAPLER WAS USED AND THE PROCEDURE CONTINUED. THERE WAS NO PATIENT INJURY.
27-11-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WAS DISCOVERED TO BE PRE-FIRED AFTER FIRING WAS ATTEMPTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
10-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN REPLACING THE SUBSTITUTE BLADE, AN UNEXPERIENCED NURSE WAS NOT ABLE TO REMOVE THE DEVICE WELL AND THE PARTS CAME OFF. A SECOND PRODUCT WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | GIA | Difficult to Open or Remove Packaging Material | A020501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN REPLACING THE SUBSTITUTE BLADE, AN UNEXPERIENCED NURSE WAS NOT ABLE TO REMOVE THE DEVICE WELL AND THE PARTS CAME OFF. A SECOND PRODUCT WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | EEA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE BROKE AND WOULD NOT ACTIVATE.
08-12-2020 | Malfunction | us surgical puerto rico | 06-01-2021 | EEA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE BROKE AND WOULD NOT ACTIVATE.
10-12-2020 | Malfunction | covidien lp llc north haven | 06-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON TRANSECTING THE STOMACH, THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE USER FIRED A NEW RELOAD AND USED THE SAME HANDLE TO COMPLETE STAPLE LINE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | ethicon endo-surgery | 06-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACIC PROCEDURE THE DEVICE WAS PUT DOWN THE TROCAR AND ARTICULATED. THE JAWS WOULD NOT OPEN. THE DEVICE WAS REMOVED AND REPLACED WITH A LIKE DEVICE WITH NO PATIENT CONSEQUENCES.
01-01-2020 | Malfunction | ethicon endo-surgery | 06-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE RELOAD DOES NOT FIT ON THE DEVICE, DEFECTIVE TIP OF THE CLAMP. UNUSABLE DEVICE.
08-12-2020 | Injury | intuitive surgical | 06-01-2021 | SUREFORM | Adverse Event Without Identified Device or Use Problem | A24 | Blood Loss | Event Description: IT WAS REPORTED THAT AFTER SEVERAL DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURES, BLEEDING FROM THE STOMACH WAS IDENTIFIED FROM THE STAPLE LINES CREATED USING A SUREFORM STAPLER 60 INSTRUMENT. THE SURGEON USES WHITE RELOADS FOR THE ILEUM, THEN GREEN AND BLUE FOR THE GASTRIC POUCH. THE TARGET TISSUE WAS NOT UNDER TENSION, NOR WAS THE INSTRUMENT. A DA VINCI REPRESENTATIVE WAS NOTED TO HAVE BEEN PRESENT DURING THE PROCEDURES, DURING WHICH THERE WERE NO ISSUES WITH THE SUREFORM STAPLER 60 INSTRUMENTS. ON 17-DEC-2020, INTUITIVE SURGICAL (ISI) CONTACTED THE SITE AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THIS EVENT: THE LAST OCCURRENCE OF THIS ISSUE WAS ON (B)(6) 2020, AND THE FOLLOWING ADDITIONAL INFORMATION WAS OBTAINED PERTAINING TO THIS PROCEDURE. THE INSTRUMENTS AND RELOADS USED WERE INSPECTED PRIOR TO USE; NOTHING OUT OF THE ORDINARY WAS OBSERVED. THE BLEEDING WAS IDENTIFIED 48 TO 72 HOURS POST-OPERATIVELY. THE AMOUNT OF BLOOD LOSS INCURRED WAS NOT SPECIFIED; HOWEVER, NO BLOOD TRANSFUSIONS WERE REQUIRED TO BE ADMINISTERED. THE SURGEON PERFORMED A LEAK TEST. PER THE SURGEON, THE BLEEDING ORIGINATED FROM THE INSIDE OF THE STOMACH AND THE BOWEL, NOT OUTSIDE. THE STAPLE LINES WERE INSPECTED AND INTACT AFTER EACH RELOAD FIRE. THE FOLLOWING IS UNKNOWN: WHICH RELOAD COLOR OR STAPLER FIRE WAS INVOLVED WITH THE ALLEGED ISSUE, IF ANY ADDITIONAL PROCEDURES WERE CONDUCTED THAT WOULD VISUALIZE THE STAPLE LINE AND CONFIRM IF THE BLEEDING WAS FROM THE STAPLE LINE. THERE WERE NO ERROR MESSAGES WHEN THE ISSUE OCCURRED, BUT A PAUSE FOR TISSUE COMPRESSION WAS NOTED. NO BLEEDING WAS OBSERVED FROM THE STAPLE LINE AFTER FIRING. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. NO BUTTRESS MATERIAL WAS USED. THERE WERE NO MALFORMED STAPLES. THE SURGEON DID NOT EXPERIENCE ANY CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD. THE TISSUE WAS NOT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY PRIOR TO THE PROCEDURE. THERE WAS NO TISSUE BUNCHING. IN ORDER TO ADDRESS THE LEAK, UNSPECIFIED CONSERVATIVE TREATMENT WITHOUT SURGERY WAS PERFORMED. THE PATIENT WAS REPORTED TO BE DOING FINE BUT HAD TO STAY LONGER IN THE HOSPITAL. THE SUREFORM STAPLER 60 INSTRUMENTS AND RELOADS ARE UNABLE TO BE RETURNED TO ISI FOR EVALUATION. PHOTOGRAPHIC IMAGES OF THE DEVICES OR VIDEO RECORDING OF THE PROCEDURES ARE NOT AVAILABLE FOR REVIEW. FURTHER INFORMATION HAS BEEN REQUESTED FROM THE SITE. HOWEVER, NO ADDITIONAL DETAILS HAVE BEEN RECEIVED AS OF THE DATE OF THIS REPORT.
09-12-2020 | Malfunction | ethicon endo-surgery | 06-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHILE CUTTING THE LUNG TISSUE USING THE GST60 BLACK CARTRIDGE, THERE WAS A PHENOMENON THAT THE BLADE COULD NOT COME BACK. FOR THIS REASON, THE CARTRIDGE AND BODY HAD TO BE FORCIBLY SEPARATED FROM THE TISSUE. SURGERY DELAYED 40 MINUTES.
09-12-2020 | Malfunction | ethicon endo-surgery | 06-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHILE CUTTING THE LUNG TISSUE USING THE GST60 BLACK CARTRIDGE, THERE WAS A PHENOMENON THAT THE BLADE COULD NOT COME BACK. FOR THIS REASON, THE CARTRIDGE AND BODY HAD TO BE FORCIBLY SEPARATED FROM THE TISSUE. SURGERY DELAYED 40 MINUTES.
27-10-2020 | Malfunction | ethicon endo-surgery | 06-01-2021 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT HAS BEEN INFORMED THAT 4 DEVICES WERE LOCKED DURING USE AND THE STAPLES DID NOT FIRE. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
09-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN DOING ANASTOMOSIS ON AN OPEN RIGHT HEMICOLECTOMY, THE STAPLER FIRED WITH THE RELOAD INCLUDED. WHEN THE SECOND RELOAD WAS PLACED TO PERFORM THE ANASTOMOSIS, THE STAPLER WAS STUCK. ANOTHER RELOAD WAS OPENED BECAUSE THEY THOUGHT THAT THE PROBLEM WAS IN THE RELOAD, BUT THE STAPLER WAS STILL JAMMED. ANOTHER STAPLER WITH THE SAME LOT WAS OPENED AND THE SURGERY WAS COMPLETED. THEY USED ONE OF THE OPENED RELOADS. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE DID NOT FIRE; THE KNIFE DID NOT GO THROUGH AND CARTRIDGE STOPPED.
08-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE DID NOT FIRE; THE KNIFE DID NOT GO THROUGH AND CARTRIDGE STOPPED.
15-11-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE BROKE. THE KNIFE STOPPED IN THE MIDDLE OF CARTRIDGE.
10-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN TRYING TO FIRE ON THE LUNG PARENCHYMA DURING AN INTERLOBAR RESECTION ON A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE LOCKED. IT WAS ALSO STATED THAT THE JAWS LOCKED ON TISSUE AND WAS REMOVED BY PULLING THE BLACK RETURN KNOB. A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN TRYING TO FIRE ON THE LUNG PARENCHYMA DURING AN INTERLOBAR RESECTION ON A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE LOCKED. IT WAS ALSO STATED THAT THE JAWS LOCKED ON TISSUE AND WAS REMOVED BY PULLING THE BLACK RETURN KNOB. A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
11-12-2020 | Injury | us surgical puerto rico | 07-01-2021 | EEA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW COLON SURGERY, AFTER ANASTOMOSIS OF THE RECTUM, THE ANASTOMOSIS WAS FOUND TO BE INCOMPLETE. THE ANASTOMOSIS "DOUGHNUT" WAS INCOMPLETE, AND AFTER THE OPERATION THE DOCTOR SUTURED AGAIN THE DETACHED INTESTINE.
11-12-2020 | Injury | us surgical puerto rico | 07-01-2021 | EEA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW COLON SURGERY, AFTER ANASTOMOSIS OF THE RECTUM, THE ANASTOMOSIS WAS FOUND TO BE INCOMPLETE. THE ANASTOMOSIS "DOUGHNUT" WAS INCOMPLETE, AND AFTER THE OPERATION THE DOCTOR SUTURED AGAIN THE DETACHED INTESTINE.
10-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS DID NOT OPEN BEFORE USE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON PROCEDURE THE SURGEON FELT THE STAPLER RELOAD MOVE A LITTLE AND STOPPED. THE RELOAD WAS REPLACE. THE DEVICE WAS INSERTED INTO THE TROCAR AND THE SURGEON HAD AN ISSUE WHERE THE REVERSE BUTTON HAD TO BE USED. THIS DID NOT WORK, AFTER WHICH THE MANUAL OVERRIDE WAS USED TO REVERSE THE BLADE AND THE DEVICE WAS REMOVED. A METAL PIECE WAS STICKING OUT OF THE DISTAL END AND A MISFIRE OCCURRED. THE PROCEDURE WAS COMPLETED WITH A COMPETITIVE DEVICE WITH NO PATIENT CONSEQUENCES.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON PROCEDURE THE SURGEON FELT THE STAPLER RELOAD MOVE A LITTLE AND STOPPED. THE RELOAD WAS REPLACE. THE DEVICE WAS INSERTED INTO THE TROCAR AND THE SURGEON HAD AN ISSUE WHERE THE REVERSE BUTTON HAD TO BE USED. THIS DID NOT WORK, AFTER WHICH THE MANUAL OVERRIDE WAS USED TO REVERSE THE BLADE AND THE DEVICE WAS REMOVED. A METAL PIECE WAS STICKING OUT OF THE DISTAL END AND A MISFIRE OCCURRED. THE PROCEDURE WAS COMPLETED WITH A COMPETITIVE DEVICE WITH NO PATIENT CONSEQUENCES.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON PROCEDURE THE SURGEON FELT THE STAPLER RELOAD MOVE A LITTLE AND STOPPED. THE RELOAD WAS REPLACE. THE DEVICE WAS INSERTED INTO THE TROCAR AND THE SURGEON HAD AN ISSUE WHERE THE REVERSE BUTTON HAD TO BE USED. THIS DID NOT WORK, AFTER WHICH THE MANUAL OVERRIDE WAS USED TO REVERSE THE BLADE AND THE DEVICE WAS REMOVED. A METAL PIECE WAS STICKING OUT OF THE DISTAL END AND A MISFIRE OCCURRED. THE PROCEDURE WAS COMPLETED WITH A COMPETITIVE DEVICE WITH NO PATIENT CONSEQUENCES.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON PROCEDURE THE SURGEON FELT THE STAPLER RELOAD MOVE A LITTLE AND STOPPED. THE RELOAD WAS REPLACE. THE DEVICE WAS INSERTED INTO THE TROCAR AND THE SURGEON HAD AN ISSUE WHERE THE REVERSE BUTTON HAD TO BE USED. THIS DID NOT WORK, AFTER WHICH THE MANUAL OVERRIDE WAS USED TO REVERSE THE BLADE AND THE DEVICE WAS REMOVED. A METAL PIECE WAS STICKING OUT OF THE DISTAL END AND A MISFIRE OCCURRED. THE PROCEDURE WAS COMPLETED WITH A COMPETITIVE DEVICE WITH NO PATIENT CONSEQUENCES.
09-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | UNKNOWN GIA PRODUCT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN DOING ANASTOMOSIS ON AN OPEN RIGHT HEMICOLECTOMY, THE STAPLER FIRED WITH THE RELOAD INCLUDED. WHEN THE SECOND RELOAD WAS PLACED TO PERFORM THE ANASTOMOSIS, THE STAPLER WAS STUCK. ANOTHER RELOAD WAS OPENED BECAUSE THEY THOUGHT THAT THE PROBLEM WAS IN THE RELOAD, BUT THE STAPLER WAS STILL JAMMED. ANOTHER STAPLER WITH THE SAME LOT WAS OPENED AND THE SURGERY WAS COMPLETED. THEY USED ONE OF THE OPENED RELOADS. THERE WAS NO PATIENT INJURY.
02-11-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ECHELON FLEX 45 STANDARD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VATS RIGHT UPPER LOBE RESECTION. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES, AND THE ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. CHANGED TO ANOTHER RELOAD GOT THE SAME ISSUE. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2010 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROUX-EN-Y PROCEDURE, THE PLEE60A STAPLER - LOT U958E WOULD NOT RELEASE AFTER IT WAS FIRED. IT WAS USED ON A PATIENT. THE CASE WAS DELAYED BY FIFTEEN MINUTES BUT COMPLETED WITH NO PATIENT CONSEQUENCES.
09-12-2019 | Malfunction | ethicon endo-surgery | 07-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A KIDNEY RESECTION THE BLADE IN RELOAD CAME INTO 2/3 OF ITS SIZE AND THEN STOPPED. THEY PULLED ETS OUT AND CHANGED FOR ANOTHER ETS (IT WORKED FINE). AFTER THE OPERATION, THEY TRIED TO USE RELOAD IN THE SECOND STAPLER AND IT WORKED FINE (CAME THE FINAL 1/3 TO DISTAL END). THERE WERE NO PATIENT CONSEQUENCES.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | No Apparent Adverse Event | A25 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | No Apparent Adverse Event | A25 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | No Apparent Adverse Event | A25 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL TITANIUM (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
10-12-2020 | Injury | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Patient Device Interaction Problem | A01 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, BLEEDING, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND REOPERATION.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, THE STAPLER FIRED BUT THEN WOULDN'T REVERSE. SCISSORS WERE USED TO CUT THE DEVICE OUT. ANOTHER STAPLER WAS USED AND WORKED. THERE WAS NO KNOWN PATIENT CONSEQUENCE.
01-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ECHELON FLEX 45 STANDARD | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A MINIMALLY INVASIVE TOTAL NEPHRECTOMY, THE SURGEON CHOSE TO USE THE ECHELON FLEX. AFTER THE LAST SHOOTING, THERE WAS NO RETURN OF THE KNIFE AND THE JAWS WERE BLOCKED AND CLOSED. EVEN AFTER THE TROUBLE SHOOTING MEASURES, IT WAS NOT POSSIBLE TO REMOVE THE EC45A FROM THE PATIENT. AS A RESULT, THE INSTRUMENT WAS REMOVED BY CUTTING THE NEAR-BY TISSUE. NO PATIENT CONSEQUENCE REPORTED.
01-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A MINIMALLY INVASIVE TOTAL NEPHRECTOMY, THE SURGEON CHOSE TO USE THE ECHELON FLEX. AFTER THE LAST SHOOTING, THERE WAS NO RETURN OF THE KNIFE AND THE JAWS WERE BLOCKED AND CLOSED. EVEN AFTER THE TROUBLE SHOOTING MEASURES, IT WAS NOT POSSIBLE TO REMOVE THE EC45A FROM THE PATIENT. AS A RESULT, THE INSTRUMENT WAS REMOVED BY CUTTING THE NEAR-BY TISSUE. NO PATIENT CONSEQUENCE REPORTED.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, THE STAPLER FIRED BUT THEN WOULDN'T REVERSE. SCISSORS WERE USED TO CUT THE DEVICE OUT. ANOTHER STAPLER WAS USED AND WORKED. THERE WAS NO KNOWN PATIENT CONSEQUENCE.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, THE STAPLER FIRED BUT THEN WOULDN'T REVERSE. SCISSORS WERE USED TO CUT THE DEVICE OUT. ANOTHER STAPLER WAS USED AND WORKED. THERE WAS NO KNOWN PATIENT CONSEQUENCE.
03-11-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GYN PROCEDURE THE FIRST DEVICE MALFUNCTIONED. A SECOND LIKE WAS TRIED AND IT ALSO FAILED. THE PROCEDURE WAS COMPLETED WITH A POWERED STAPLER WITH NO PATIENT CONSEQUENCES. THERE IS NO ADDITIONAL INFORMATION.
16-11-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC LEFT UPPER LOBECTOMY, COULD NOT FIRE WHEN SERVING THE LUNG TISSUE WITH EC45A + ECR45G, RELOAD ANOTHER ECR45G RELOAD STILL GOT THE SAME ISSUE. CHANGED TO ANOTHER ECR45G RELOAD TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE SURGERY, THE INCONVENIENCE THAT OCCURRED WAS THAT THE STAPLE CLOSURE WAS NOT FORMED CORRECTLY WITH ECHELON FLEX CODE PSEE60A AND RELOADS GST60B. THE SURGEON HAD TO REINFORCE THE ENTIRE STAPLE LINE WITH CLIPS. THERE WAS NO PROBLEM WITH THE PATIENT, WHO IS ALREADY DISCHARGED.
07-11-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC DISTAL GASTRECTOMY SURGERY, COULD NOT FIRE WHEN SEVERING DUODENUM TISSUE ON THE 1ST FIRING. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE TWO-PORT VIDEO-ASSISTED THORACOSCOPIC RIGHT UPPER LUNG RESECTION SURGERY, WHEN SEVERING PULMONARY FISSURE TISSUE, AFTER FIRED WITH ECR60B, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE TWO-PORT VIDEO-ASSISTED THORACOSCOPIC RIGHT UPPER LUNG RESECTION SURGERY, WHEN SEVERING PULMONARY FISSURE TISSUE, AFTER FIRED WITH ECR60B, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | WHITE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VATS SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPES (WITH STRAIGHT LEG). THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | WHITE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VATS SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPES (WITH STRAIGHT LEG). THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, THE KNIFE DID NOT RETURN TO HOME POSITION, OR THE JAWS DID NOT OPEN WHEN THE DEVICE WAS FIRED OUTSIDE THE PATIENT AND THE KNIFE REVERSE SWITCH WAS USED BEFORE THE 1ST FIRING. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, THE KNIFE DID NOT RETURN TO HOME POSITION, OR THE JAWS DID NOT OPEN WHEN THE DEVICE WAS FIRED OUTSIDE THE PATIENT AND THE KNIFE REVERSE SWITCH WAS USED BEFORE THE 1ST FIRING. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-10-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC SURGERY, THE STAPLER WAS POSITIONED AND RESECTION STARTED WITH A PURPLE 45 MM LOAD, AFTER A PURPLE 60 MM LOAD "FAILED" BECAUSE THE STAPLER HAD BROKEN AND WAS NOT ABLE TO FIRE. IT WAS NOT "FINCING," THEN ANOTHER STAPLER WAS OPENED AND ANOTHER PURPLE 60 MM LOAD WAS USED, WHICH WORKED CORRECTLY, BUT A PURPLE 45 MM LOAD OPENED THE STAPLES CAUSING ACTIVE BLEEDING, THEN THE RESECTION WAS TRIED TO BE CONTINUED WITH ANOTHER 45 MM LOAD, BUT IT DIDN'T FIRE, AS THERE WERE NO MORE 60 MM LOADS AND BLEEDING WAS OCCURRING, ANOTHER STAPLER WAS OPENED AND IT WORKED CORRECTLY, PERFORMING GOOD HEMOSTASIS AND ENDING THE RESECTION WITH TWO 60 MM GREEN LOADS. THERE WAS NO PATIENT INJURY.
11-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, WHEN THE USER SEVERED THE RECTUM, THE RELOAD FIRED ONLY ONCE, ON THE 2ND TIME, IT ONLY FORWARDED A SMALL STEP WHEN FIRED, AND THE JAWS COULD NOT BE CLOSED COMPLETELY, THE CLAMPING FORCE WAS WEAKENED WHEN THE JAWS WERE CLOSED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE OPERATOR WAS UNABLE TO UNLOAD THE RELOAD SMOOTHLY AND WAS ABLE TO REMOVE AFTER THREE ATTEMPTS. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, WHEN THE USER SEVERED THE RECTUM, THE RELOAD FIRED ONLY ONCE, ON THE 2ND TIME, IT ONLY FORWARDED A SMALL STEP WHEN FIRED, AND THE JAWS COULD NOT BE CLOSED COMPLETELY, THE CLAMPING FORCE WAS WEAKENED WHEN THE JAWS WERE CLOSED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE OPERATOR WAS UNABLE TO UNLOAD THE RELOAD SMOOTHLY AND WAS ABLE TO REMOVE AFTER THREE ATTEMPTS. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, WHEN THE USER SEVERED THE RECTUM, THE RELOAD FIRED ONLY ONCE, ON THE 2ND TIME, IT ONLY FORWARDED A SMALL STEP WHEN FIRED, AND THE JAWS COULD NOT BE CLOSED COMPLETELY, THE CLAMPING FORCE WAS WEAKENED WHEN THE JAWS WERE CLOSED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE OPERATOR WAS UNABLE TO UNLOAD THE RELOAD SMOOTHLY AND WAS ABLE TO REMOVE AFTER THREE ATTEMPTS. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-11-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC HEPATECTOMY, WHEN SEVERING LIVER TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). NEW RELOAD STILL GOT THE SAME ISSUE. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-12-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, ON THE LUNG, THE SURGEON WANTED TO PRESS THE GREEN BUTTON ON THE RIGHT BUT IT DID NOT WORK FOR THE FIRING MODE. SO THE USER TRIED TO REPEAT 5 TIMES, IT ALSO DID NOT WORK. AFTER THAT, THE HANDLE WAS SUDDENLY ABLE TO ENTER FIRING MODE. THE USER WAS READY TO FIRE THE TISSUE BUT IT WOULD NOT FIRE. THE POWERED HANDLE SET WAS CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, ON THE LUNG, THE SURGEON WANTED TO PRESS THE GREEN BUTTON ON THE RIGHT BUT IT DID NOT WORK FOR THE FIRING MODE. SO THE USER TRIED TO REPEAT 5 TIMES, IT ALSO DID NOT WORK. AFTER THAT, THE HANDLE WAS SUDDENLY ABLE TO ENTER FIRING MODE. THE USER WAS READY TO FIRE THE TISSUE BUT IT WOULD NOT FIRE. THE POWERED HANDLE SET WAS CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | SIGNIA | Material Split, Cut or Torn | A0414 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, ON THE LUNG, THE SURGEON WANTED TO PRESS THE GREEN BUTTON ON THE RIGHT BUT IT DID NOT WORK FOR THE FIRING MODE. SO THE USER TRIED TO REPEAT 5 TIMES, IT ALSO DID NOT WORK. AFTER THAT, THE HANDLE WAS SUDDENLY ABLE TO ENTER FIRING MODE. THE USER WAS READY TO FIRE THE TISSUE BUT IT WOULD NOT FIRE. THE POWERED HANDLE SET WAS CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-07-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN DOING LUNG LOBE ANASTOMOSIS ON A LAPAROSCOPIC WEDGE RESECTION, THE STAPLER WAS ABLE TO FIRE BUT THE STAPLE BURST. ANOTHER DEVICE WAS USED. THE EVENT RESULTED TO A TEMPORARY INJURY AND SUTURED MANUALLY. THE SURGICAL TIME WAS ALSO EXTENDED TO 30 MINUTES OR MORE. THE PATIENT STAYED AT THE HOSPITAL FOR OBSERVATION.
10-12-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, ON THE LUNG, THE SURGEON WANTED TO PRESS THE GREEN BUTTON ON THE RIGHT BUT IT DID NOT WORK FOR THE FIRING MODE. SO THE USER TRIED TO REPEAT 5 TIMES, IT ALSO DID NOT WORK. AFTER THAT, THE HANDLE WAS SUDDENLY ABLE TO ENTER FIRING MODE. THE USER WAS READY TO FIRE THE TISSUE BUT IT WOULD NOT FIRE. THE POWERED HANDLE SET WAS CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE GASTROPLASTY (SLEEVE) PROCEDURE, THE RELOAD DID NOT PERFORM SATISFACTORILY IN THE MOMENT OF STAPLING WHICH WAS PERFORMED CORRECTLY. THE SURGEON HAD TO MANUALLY SUTURE THE ENTIRE LINE OF STAPLES AND IN PARTICULAR, THE PART WHERE THE RELOAD STAPLED WITHOUT BEING SUCCESSFUL. THE PRODUCT DID NOT CAUSE DAMAGE TO THE PATIENT.
02-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC PANCREATECTOMY, AFTER LOADING THE DEVICE, THE OPERATOR NOTICED THAT THE RELOAD WAS NOT FULLY CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
02-12-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC PANCREATECTOMY, AFTER LOADING THE DEVICE, THE OPERATOR NOTICED THAT THE RELOAD WAS NOT FULLY CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
14-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC RECTECTOMY SURGERY, WHEN SEVERING RECTUM TISSUE ON THE 2ND FIRING, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO THE NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-09-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD PARTIALLY FIRED.
07-09-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD PARTIALLY FIRED.
07-09-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | EEA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD PARTIALLY FIRED.
07-09-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | EEA | Material Twisted/Bent | A040609 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD PARTIALLY FIRED.
07-09-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | EEA | Material Twisted/Bent | A040609 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD PARTIALLY FIRED.
07-09-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | EEA | Material Twisted/Bent | A040609 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD PARTIALLY FIRED.
11-12-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC WEDGE RESECTION, THE STAPLER WAS ABLE TO FIRE, NAILS BURST WHEN THE LUNG TISSUE WAS CUT, THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER PURPLE NAIL WAS USED ON THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC WEDGE RESECTION, THE STAPLER WAS ABLE TO FIRE, NAILS BURST WHEN THE LUNG TISSUE WAS CUT, THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER PURPLE NAIL WAS USED ON THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | PWRD ECH FLEX 60MM 28CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC SURGERY FOR LUNG CANCER SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN THE KNIFE AUTOMATICALLY. THE KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. CHANGED TO ANOTHER PCE60A + ECR60W AND COULD NOT FIRE. ANOTHER BRAND DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | PWRD ECH FLEX 60MM 28CM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC SURGERY FOR LUNG CANCER SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN THE KNIFE AUTOMATICALLY. THE KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. CHANGED TO ANOTHER PCE60A + ECR60W AND COULD NOT FIRE. ANOTHER BRAND DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | PWRD ECH FLEX 60MM 28CM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC SURGERY FOR LUNG CANCER SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN THE KNIFE AUTOMATICALLY. THE KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. CHANGED TO ANOTHER PCE60A + ECR60W AND COULD NOT FIRE. ANOTHER BRAND DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-09-2020 | Injury | davis & geck caribe | 07-01-2021 | GIA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY PROCEDURE, THE STAPLER STOPPED FIRING AND BROKE OFF INSIDE THE PATIENT'S ABDOMEN. THE COMPONENT WAS RETRIEVE. SURGEON HAD TO REDO THE SUTURING BY HAND AND TOOK THE SURGEON 2 HOURS TO FINISH.
04-09-2020 | Injury | davis & geck caribe | 07-01-2021 | GIA | Detachment of Device or Device Component | A0501 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY PROCEDURE, THE STAPLER STOPPED FIRING AND BROKE OFF INSIDE THE PATIENT'S ABDOMEN. THE COMPONENT WAS RETRIEVE. SURGEON HAD TO REDO THE SUTURING BY HAND AND TOOK THE SURGEON 2 HOURS TO FINISH.
07-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED PLUS COMPACT 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY, THE KNIFE DID NOT MOVE FORWARD AT THE FIRING ON THE LUNG. THE MANUAL OVERRIDE LEVER WAS USED, BUT THE KNIFE DID NOT RETURN TO THE HOME POSITION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED PLUS COMPACT 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY, THE KNIFE DID NOT MOVE FORWARD AT THE FIRING ON THE LUNG. THE MANUAL OVERRIDE LEVER WAS USED, BUT THE KNIFE DID NOT RETURN TO THE HOME POSITION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED PLUS COMPACT 45 | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY, THE KNIFE DID NOT MOVE FORWARD AT THE FIRING ON THE LUNG. THE MANUAL OVERRIDE LEVER WAS USED, BUT THE KNIFE DID NOT RETURN TO THE HOME POSITION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED PLUS COMPACT 45 | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY, THE KNIFE DID NOT MOVE FORWARD AT THE FIRING ON THE LUNG. THE MANUAL OVERRIDE LEVER WAS USED, BUT THE KNIFE DID NOT RETURN TO THE HOME POSITION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
28-11-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN EXPLORATORY LAPAROTOMY WITH A RESECTION OF THE TRANSVERSE COLON PROCEDURE THAT SURGEON WENT TO STAPLE THE DISTAL COLON. SURGEON THEN CUT THE COLON WITH A SCALPEL. WHEN THE SURGEON OPENED THE STAPLER THERE WERE NO STAPLES AND THE COLON WAS OPEN. ANOTHER STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
07-12-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, THE FIRING STOPPED HALFWAY WHEN RESECTING THE STOMACH. ADDITIONAL STAPLING WAS PERFORMED WITH NEW RELOAD. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | WHITE RELOAD FOR ECHELON 60 | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THAT THE STERILE PACKING WAS NOT COMPLETELY SEALED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
07-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, THE FIRING STOPPED HALFWAY WHEN RESECTING THE STOMACH. ADDITIONAL STAPLING WAS PERFORMED WITH NEW RELOAD. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | GREEN RELOAD FOR ECHELON 60 | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THAT THE STERILE PACKING WAS NOT COMPLETELY SEALED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-12-2020 | Malfunction | us surgical puerto rico | 07-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENTERECTOMY, THE STAPLER LOCKED, DID NOT FIRE THE LOAD. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | GST WHITE RELOAD, 60MM, 6 ROW | Use of Device Problem | A23 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TESTING IN THE LAB THAT THERE WAS A DOUBLE-LOADED STAPLE IN THE RELOAD. ALL STAPLES FORMED PERFECTLY. THERE WAS NO PATIENT INVOLVEMENT. RELOAD WAS SCRAPPED. TEST SKIN IS AVAILABLE.
03-11-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GYN PROCEDURE THE FIRST DEVICE MALFUNCTIONED. A SECOND LIKE WAS TRIED AND IT ALSO FAILED. THE PROCEDURE WAS COMPLETED WITH A POWERED STAPLER WITH NO PATIENT CONSEQUENCES. THERE IS NO ADDITIONAL INFORMATION.
11-12-2020 | Malfunction | celestica de monterrey s.a de c.v | 07-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LUNG RESECTION FOR LAPAROSCOPIC WEDGE SURGERY, POOR STAPLE FORMATION WAS OBSERVED ON THE FRONT PART OF FOUR RELOADS USED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE; THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO SHEARED FIRING RACK TEETH DAMAGE.
01-01-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | GREEN RELOAD FOR ECHELON 60 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE RELOAD BROKE IN THE STAPLER. SURGERY WAS DELAYED BY THIRTY MINUTES HOWEVER SURGERY WAS SUCCESSFULLY COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
03-11-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GYN PROCEDURE THE FIRST DEVICE MALFUNCTIONED. A SECOND LIKE WAS TRIED AND IT ALSO FAILED. THE PROCEDURE WAS COMPLETED WITH A POWERED STAPLER WITH NO PATIENT CONSEQUENCES. THERE IS NO ADDITIONAL INFORMATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, WHEN OPENING A 60 MM POWERED STAPLER PACKAGE, THE NURSE FOUND A 45 MM STAPLER DESPITE THE COVERAGE SHOWING A 60 MM ON IT. THEY CHANGED THE STAPLER AND THE SLEEVE PROCEDURE CONTINUED WITHOUT ANY COMPLICATIONS FOR THE PATIENT. THERE WAS NO PATIENT CONSEQUENCE.
11-09-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING LAPAROSCOPIC RESECTION OF LOW RECTAL CANCER SURGERY WHEN SEVERING RECTUM TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO HOME POSITION. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. IT IS UNKNOWN IF THE KNIFE REVERSE SWITCH WAS USED. THE FIRING WAS COMPLETED PROPERLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO HOME POSITION. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. IT IS UNKNOWN IF THE KNIFE REVERSE SWITCH WAS USED. THE FIRING WAS COMPLETED PROPERLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO HOME POSITION AT THE SECOND FIRING. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2020 | Malfunction | ethicon endo-surgery | 07-01-2021 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO HOME POSITION AT THE SECOND FIRING. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
28-09-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE ADAPTER STATED THAT THE RELOAD WAS INCOMPATIBLE WITH THE TISSUE IT WAS BEING APPLIED TO. THE DEVICE FIRST INDICATED THAT THE RELOAD WAS GOOD THEN FAILED STATING THAT THE RELOAD WAS INADEQUATE, GOING FROM 1 TO 3. ANOTHER ADAPTER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-09-2020 | Malfunction | covidien lp llc north haven | 07-01-2021 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE ADAPTER STATED THAT THE RELOAD WAS INCOMPATIBLE WITH THE TISSUE IT WAS BEING APPLIED TO. THE DEVICE FIRST INDICATED THAT THE RELOAD WAS GOOD THEN FAILED STATING THAT THE RELOAD WAS INADEQUATE, GOING FROM 1 TO 3. ANOTHER ADAPTER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-03-2018 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE HANDLE AND RELOAD WERE SEPARATED WHICH CAUSED THE DEVICE TO NOT FIRE.
21-03-2018 | Malfunction | covidien lp llc north haven | 08-01-2021 | UNKNOWN ENDO GIA SULU | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE HANDLE AND RELOAD WERE SEPARATED WHICH CAUSED THE DEVICE TO NOT FIRE.
03-11-2020 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE DEVICE HAD AN UNKNOWN FAILURE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH ON THE FIRING RACK WHICH MAY OCCUR WHEN FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE OR FIRING WITH AN OBSTACLE INCORPORATED IN THE JAWS. IN EITHER OF THESE CIRCUMSTANCES, IT WILL BECOME INCREASINGLY DIFFICULT TO ACTUATE THE FIRING HANDLE AND THE INSTRUMENT RETURN KNOBS WILL BE DIFFICULT TO RETRACT. IN ADDITION, STAPLES MAY NOT FORM PROPERLY, AND TISSUE MAY NOT BE FULLY TRANSECTED.
03-11-2020 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE DEVICE HAD AN UNKNOWN FAILURE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH ON THE FIRING RACK WHICH MAY OCCUR WHEN FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE OR FIRING WITH AN OBSTACLE INCORPORATED IN THE JAWS. IN EITHER OF THESE CIRCUMSTANCES, IT WILL BECOME INCREASINGLY DIFFICULT TO ACTUATE THE FIRING HANDLE AND THE INSTRUMENT RETURN KNOBS WILL BE DIFFICULT TO RETRACT. IN ADDITION, STAPLES MAY NOT FORM PROPERLY, AND TISSUE MAY NOT BE FULLY TRANSECTED.
16-11-2020 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SIX RELOADS' JAWS WOULD NOT CLOSE AND OPEN. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-10-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY TWO DEVICES LOCKED ON TISSUE AND HAD TO BE MANUALLY OVERRIDDEN TO REMOVE. A THIRD DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
01-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | PROXIMATE*RELOAD THICK TISSUE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY, "LINEAR RELOAD THAT IS PASSED TO BE USED IN INTESTINAL RESECTION PLUS ANASTOMOSIS OF THE DRA IS ASSEMBLED ON THE TLC 75 STAPLER AND WHEN IT IS FIRED, THE STAPLER STICKS AND WHEN TRYING TO UNLOCK BECAUSE IT WAS HOLDING THE TISSUE, THE REFILL SHOOT. IT DOES NOT HARM THE PATIENT." THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
01-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | PROXIMATE*RELOAD THICK TISSUE | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY, "LINEAR RELOAD THAT IS PASSED TO BE USED IN INTESTINAL RESECTION PLUS ANASTOMOSIS OF THE DRA IS ASSEMBLED ON THE TLC 75 STAPLER AND WHEN IT IS FIRED, THE STAPLER STICKS AND WHEN TRYING TO UNLOCK BECAUSE IT WAS HOLDING THE TISSUE, THE REFILL SHOOT. IT DOES NOT HARM THE PATIENT." THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
13-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF SIGMOID COLON CANCER SURGERY, WHEN SEVERING COLON TISSUE, AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE( WITH STRAIGHT LEG). CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE DISTAL GASTRECTOMY SURGERY, WHEN SEVERING GASTROJEJUNUM TISSUE , AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE( WITH STRAIGHT LEG). CHANGED THE NEW ONE TO COMPLETE SURGERY ( USED SUTURE TO OVERSEW). THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-10-2020 | Malfunction | covidien lp llc north haven | 08-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, ON PREPARATION, THERE WERE TWO HANDLES AND TWO CHARGERS, THERE WAS A TRACE THAT THE MALFUNCTION SEEMED TO BE CAUSED BY LIQUID LEAKAGE FROM THE HANDLE TO THE CHARGER SIDE. IT WAS UNKNOWN WHETHER THIS WAS LEAKAGE FROM THE POWER HANDLE, OR WHETHER IT WAS A SHORT DUE TO RETURNING THE POWER HANDLE TO THE CHARGER BEFORE ALCOHOL COMPLETELY EVAPORATED AFTER THE POWER HANDLE WAS CLEANED WITH ALCOHOL SWAB. BOTH THE POWER HANDLES MALFUNCTIONED AND BECAME UNABLE TO BE USED. THE FIRST HANDLE DISPLAY SCREEN DID NOT START UP COMPLETELY. THE SECOND HANDLE DISPLAY SCREEN STARTED UP AND WAS BRIGHT WHEN THE POWER HANDLE WAS PLACED ON THE CHARGER, BUT AS SOON AS IT WAS REMOVED FROM THE CHARGER, THE DISPLAY SCREEN BECAME COMPLETELY DARK AND STOPPED OPERATING. A PRODUCT FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
06-10-2020 | Malfunction | covidien lp llc north haven | 08-01-2021 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, ON PREPARATION, THERE WERE TWO HANDLES AND TWO CHARGERS, THERE WAS A TRACE THAT THE MALFUNCTION SEEMED TO BE CAUSED BY LIQUID LEAKAGE FROM THE HANDLE TO THE CHARGER SIDE. IT WAS UNKNOWN WHETHER THIS WAS LEAKAGE FROM THE POWER HANDLE, OR WHETHER IT WAS A SHORT DUE TO RETURNING THE POWER HANDLE TO THE CHARGER BEFORE ALCOHOL COMPLETELY EVAPORATED AFTER THE POWER HANDLE WAS CLEANED WITH ALCOHOL SWAB. BOTH THE POWER HANDLES MALFUNCTIONED AND BECAME UNABLE TO BE USED. THE FIRST HANDLE DISPLAY SCREEN DID NOT START UP COMPLETELY. THE SECOND HANDLE DISPLAY SCREEN STARTED UP AND WAS BRIGHT WHEN THE POWER HANDLE WAS PLACED ON THE CHARGER, BUT AS SOON AS IT WAS REMOVED FROM THE CHARGER, THE DISPLAY SCREEN BECAME COMPLETELY DARK AND STOPPED OPERATING. A PRODUCT FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
06-10-2020 | Malfunction | covidien lp llc north haven | 08-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, ON PREPARATION, THERE WERE TWO HANDLES AND TWO CHARGERS, THERE WAS A TRACE THAT THE MALFUNCTION SEEMED TO BE CAUSED BY LIQUID LEAKAGE FROM THE HANDLE TO THE CHARGER SIDE. IT WAS UNKNOWN WHETHER THIS WAS LEAKAGE FROM THE POWER HANDLE, OR WHETHER IT WAS A SHORT DUE TO RETURNING THE POWER HANDLE TO THE CHARGER BEFORE ALCOHOL COMPLETELY EVAPORATED AFTER THE POWER HANDLE WAS CLEANED WITH ALCOHOL SWAB. BOTH THE POWER HANDLES MALFUNCTIONED AND BECAME UNABLE TO BE USED. THE FIRST HANDLE DISPLAY SCREEN DID NOT START UP COMPLETELY. THE SECOND HANDLE DISPLAY SCREEN STARTED UP AND WAS BRIGHT WHEN THE POWER HANDLE WAS PLACED ON THE CHARGER, BUT AS SOON AS IT WAS REMOVED FROM THE CHARGER, THE DISPLAY SCREEN BECAME COMPLETELY DARK AND STOPPED OPERATING. A PRODUCT FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE DEVICE DID NOT FUNCTION ALTHOUGH THE RED SAFETY WAS RELEASED AND THE FIRING TRIGGER WAS PULLED. IT WAS CHECKED IF THE ANVIL WAS ATTACHED TO THE TROCAR PROPERLY AND IF THE BATTERY WAS INSERTED PROPERLY, BUT THE DEVICE COULD NOT BE FIRED. THE DEVICE WAS USED BETWEEN THE COLON AND THE RECTUM. WHEN THE DEVICE WAS FIRED OUTSIDE THE PATIENT FOR FUNCTIONALITY, IT WAS FIRED PROPERLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS NOT CONVERTED TO OPEN. THE PATIENT IS STABLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2020 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DISCONNECTING THE LUNG PARENCHYMA IN A THORACOSCOPIC LEFT LOWER LUNG RESECTION, IT WAS NOTED THAT THERE WAS POOR STAPLE FORMATION. THE STAPLER WON'T DO A PROPER B SHAPE. SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
17-11-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED PLUS STANDARD 45 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WAS AN UNMATCHED SIZE BETWEEN PACKAGING AND PRODUCT. 60 MM ON PACKAGING BUT ACTUAL PRODUCT SIZE WAS 45 MM. NO PATIENT CONSEQUENCE REPORTED.
14-12-2020 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY, WHEN RESECTING LUNG, THE DEVICE WAS UNABLE FIRE. GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-11-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED PLUS LONG 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOT ASSISTED PROSTATECTOMY PROCEDURE THAT FROM THE ANCILLARY PORT ON THE FIRST ATTEMPT TO FIRE WITH A WHITE RELOAD THE DEVICE MADE A ¿WEIRD¿ SOUND AND INCHED FORWARD BUT STOPPED. REVERSE WAS ATTEMPTED BUT MANUAL OVERRIDE WAS USED TO REMOVE THE DEVICE FROM THE PATIENT. UNKNOWN IF THE DEVICE DEPLOYED ANY STAPLES OR CUT TISSUE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
02-11-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED PLUS LONG 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOT ASSISTED PROSTATECTOMY PROCEDURE THAT FROM THE ANCILLARY PORT ON THE FIRST ATTEMPT TO FIRE WITH A WHITE RELOAD THE DEVICE MADE A ¿WEIRD¿ SOUND AND INCHED FORWARD BUT STOPPED. REVERSE WAS ATTEMPTED BUT MANUAL OVERRIDE WAS USED TO REMOVE THE DEVICE FROM THE PATIENT. UNKNOWN IF THE DEVICE DEPLOYED ANY STAPLES OR CUT TISSUE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT BE TRIGGERED. ANOTHER DEVICE WAS OPENED AND THE SEWN-IN HEAD REMAINED IN THE PATIENT. WITH THE SECOND DEVICE THE OP WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES. PROCEDURE: COLORECTAL SURGERY
14-12-2020 | Malfunction | covidien lp llc north haven | 08-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY, WHEN RESECTING LUNG, THE DEVICE WAS UNABLE FIRE - GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-10-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED PLUS STANDARD 45 | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT IT WAS IMPOSSIBLE TO INSERT A RELOAD AFTER 2 USAGES. NO CONSEQUENCES TO THE PATIENT AND NO POST-OP CHANGE.
11-12-2020 | Malfunction | covidien lp llc north haven | 08-01-2021 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE LOAD HAD BEEN CAUGHT. SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE.
12-12-2020 | Malfunction | covidien lp llc north haven | 08-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING THE ANTRUM DURING A LAPAROSCOPIC PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. OPEN STAPLES WERE ALSO NOTICED AT THE DISTAL PART. ANOTHER RELOAD WAS USED TO FIX THE STAPLE LINE AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
12-12-2020 | Malfunction | covidien lp llc north haven | 08-01-2021 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING THE ANTRUM DURING A LAPAROSCOPIC PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. OPEN STAPLES WERE ALSO NOTICED AT THE DISTAL PART. ANOTHER RELOAD WAS USED TO FIX THE STAPLE LINE AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
12-12-2020 | Malfunction | covidien lp llc north haven | 08-01-2021 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING THE ANTRUM DURING A LAPAROSCOPIC PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. OPEN STAPLES WERE ALSO NOTICED AT THE DISTAL PART. ANOTHER RELOAD WAS USED TO FIX THE STAPLE LINE AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
12-12-2020 | Malfunction | covidien lp llc north haven | 08-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING THE ANTRUM DURING A LAPAROSCOPIC PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. OPEN STAPLES WERE ALSO NOTICED AT THE DISTAL PART. ANOTHER RELOAD WAS USED TO FIX THE STAPLE LINE AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | No Apparent Adverse Event | A25 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | No Apparent Adverse Event | A25 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | No Apparent Adverse Event | A25 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Pleural Effusion | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | UNKNOWN MM NEW TLC | Patient Device Interaction Problem | A01 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD 9 & 10 CATEGORIZATION EXPERIENCED THE FOLLOWING WITH CORRESPONDING INTERVENTION. PATIENT HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK, ANASTOMOTIC LEAK, BLEEDING REQUIRING BLOOD TRANSFUSION, PLEURAL EFFUSION, PNEUMONIA, SEPSIS, SURGICAL SIT INFECTION AND 30-DAY REOPERATION.
03-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING UNKNOWN SURGERY, WHEN SEVERING INTESTINE TISSUE, AFTER FIRED, NOTED THE DISTAL STAPLES MALFORMED . THE SURGEON NOTED ONE STAPLE PROTRUDING. THE DOCTOR STITCHED AND REINFORCED BY HAND. THE PATIENT WAS CHECKED ON (B)(6) AND NO ANASTOMOTIC LEAKAGE OCCURRED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. THERE WERE NO PATIENT CONSEQUENCES OR CHANGE TO THE POST-OPERATIVE CARE OF THE PATIENT.
14-12-2020 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING USED FOR GASTRIC POUCH CREATION DURING A LAPAROSCOPIC BYPASS REVISION, NO STAPLES WERE FIRED. IT WAS STATED THAT WHEN CLAMPING, THE STAPLER DID NOT CLOSE COMPLETELY AS IT NORMALLY WOULD. THE JAWS DID NOT CLOSE AS TIGHTLY AS NORMAL BUT DID CLOSE. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING USED FOR GASTRIC POUCH CREATION DURING A LAPAROSCOPIC BYPASS REVISION, NO STAPLES WERE FIRED. IT WAS STATED THAT WHEN CLAMPING, THE STAPLER DID NOT CLOSE COMPLETELY AS IT NORMALLY WOULD. THE JAWS DID NOT CLOSE AS TIGHTLY AS NORMAL BUT DID CLOSE. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING USED FOR GASTRIC POUCH CREATION DURING A LAPAROSCOPIC BYPASS REVISION, NO STAPLES WERE FIRED. IT WAS STATED THAT WHEN CLAMPING, THE STAPLER DID NOT CLOSE COMPLETELY AS IT NORMALLY WOULD. THE JAWS DID NOT CLOSE AS TIGHTLY AS NORMAL BUT DID CLOSE. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 08-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING USED FOR GASTRIC POUCH CREATION DURING A LAPAROSCOPIC BYPASS REVISION, NO STAPLES WERE FIRED. IT WAS STATED THAT WHEN CLAMPING, THE STAPLER DID NOT CLOSE COMPLETELY AS IT NORMALLY WOULD. THE JAWS DID NOT CLOSE AS TIGHTLY AS NORMAL BUT DID CLOSE. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY AND SMALL BOWEL RESECTION, THE KNIFE MOVED ALL THE WAY BUT STOPPED ON THE WAY BACK TO HOME POSITION AT THE 2ND FIRING ON THE COLON. THE KNIFE REVERSE SWITCH AND THE MANUAL OVERRIDE LEVER WERE USED, BUT THE KNIFE DID NOT RETURN. THE TARGET TISSUE WAS NOT THICK, OR THE JAW DID NOT CLAMP THE CLIP. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THERE WAS NO NEED ADDITIONAL INCISION AND CONVERTING TO OPEN SURGERY. THERE WAS NO BLEEDING AND LEAKAGE FOR THE PATIENT.
11-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY AND SMALL BOWEL RESECTION, THE KNIFE MOVED ALL THE WAY BUT STOPPED ON THE WAY BACK TO HOME POSITION AT THE 2ND FIRING ON THE COLON. THE KNIFE REVERSE SWITCH AND THE MANUAL OVERRIDE LEVER WERE USED, BUT THE KNIFE DID NOT RETURN. THE TARGET TISSUE WAS NOT THICK, OR THE JAW DID NOT CLAMP THE CLIP. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THERE WAS NO NEED ADDITIONAL INCISION AND CONVERTING TO OPEN SURGERY. THERE WAS NO BLEEDING AND LEAKAGE FOR THE PATIENT.
11-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY AND SMALL BOWEL RESECTION, THE KNIFE MOVED ALL THE WAY BUT STOPPED ON THE WAY BACK TO HOME POSITION AT THE 2ND FIRING ON THE COLON. THE KNIFE REVERSE SWITCH AND THE MANUAL OVERRIDE LEVER WERE USED, BUT THE KNIFE DID NOT RETURN. THE TARGET TISSUE WAS NOT THICK, OR THE JAW DID NOT CLAMP THE CLIP. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THERE WAS NO NEED ADDITIONAL INCISION AND CONVERTING TO OPEN SURGERY. THERE WAS NO BLEEDING AND LEAKAGE FOR THE PATIENT.
11-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY AND SMALL BOWEL RESECTION, THE KNIFE MOVED ALL THE WAY BUT STOPPED ON THE WAY BACK TO HOME POSITION AT THE 2ND FIRING ON THE COLON. THE KNIFE REVERSE SWITCH AND THE MANUAL OVERRIDE LEVER WERE USED, BUT THE KNIFE DID NOT RETURN. THE TARGET TISSUE WAS NOT THICK, OR THE JAW DID NOT CLAMP THE CLIP. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THERE WAS NO NEED ADDITIONAL INCISION AND CONVERTING TO OPEN SURGERY. THERE WAS NO BLEEDING AND LEAKAGE FOR THE PATIENT.
10-08-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RIGHT HEMICOLECTOMY SURGERY, WHEN SEVERING COLON TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE( WITH STRAIGHT LEG). CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-10-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC TOTAL GASTRECTOMY SURGERY, WHEN SEVERING STOMACH TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY ON THE LAST FIRING THERE WAS A STAPLE LEG 'STICKING UP'. THE SURGEON OVERSEWED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE CIRCULAR STAPLER DID NOT FIRE THE RELOAD.
01-01-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORAX SEGMENT RESECTION, AFTER USING THE SECOND RELOAD, THE INSTRUMENT COULD NOT BE OPENED EITHER ELECTRICALLY OR MANUALLY. THE TISSUE COULD BE REMOVED FROM THE INSTRUMENT BY MANIPULATION WITH THE GRASPING FORCEPS. THE INSTRUMENT REMAINS IN THE CLOSED POSITION. THE DEVICE WAS CLEANED AND WIPED IN SOLUTION BEFORE EACH USE OF ANOTHER RELOAD. SURGERY WAS COMPLETED WITH A NEW STAPLER. WITHOUT COMPLICATIONS. THERE WAS NO PATIENT CONSEQUENCE.
01-01-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED PLUS STANDARD 45 | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORAX SEGMENT RESECTION, AFTER USING THE SECOND RELOAD, THE INSTRUMENT COULD NOT BE OPENED EITHER ELECTRICALLY OR MANUALLY. THE TISSUE COULD BE REMOVED FROM THE INSTRUMENT BY MANIPULATION WITH THE GRASPING FORCEPS. THE INSTRUMENT REMAINS IN THE CLOSED POSITION. THE DEVICE WAS CLEANED AND WIPED IN SOLUTION BEFORE EACH USE OF ANOTHER RELOAD. SURGERY WAS COMPLETED WITH A NEW STAPLER. WITHOUT COMPLICATIONS. THERE WAS NO PATIENT CONSEQUENCE.
01-01-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED PLUS STANDARD 45 | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORAX SEGMENT RESECTION, AFTER USING THE SECOND RELOAD, THE INSTRUMENT COULD NOT BE OPENED EITHER ELECTRICALLY OR MANUALLY. THE TISSUE COULD BE REMOVED FROM THE INSTRUMENT BY MANIPULATION WITH THE GRASPING FORCEPS. THE INSTRUMENT REMAINS IN THE CLOSED POSITION. THE DEVICE WAS CLEANED AND WIPED IN SOLUTION BEFORE EACH USE OF ANOTHER RELOAD. SURGERY WAS COMPLETED WITH A NEW STAPLER. WITHOUT COMPLICATIONS. THERE WAS NO PATIENT CONSEQUENCE.
15-07-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: THE DEVICE GOT JAMMED ON THE LUNG PARENCHYMA WHICH HAD NOT BEEN STAPLED. USE OF THE DEVICE'S URGENCY SYSTEM.
15-07-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: THE DEVICE GOT JAMMED ON THE LUNG PARENCHYMA WHICH HAD NOT BEEN STAPLED. USE OF THE DEVICE'S URGENCY SYSTEM.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PULMONARY SEGMENTECTOMY, THE KNIFE MOVED ALL THE WAY BUT STOPPED MOVING BACK TO HOME POSITION AT THE 4TH FIRING. THE KNIFE REVERSE SWITCH WAS NOT USED, BUT THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PULMONARY SEGMENTECTOMY, THE KNIFE MOVED ALL THE WAY BUT STOPPED MOVING BACK TO HOME POSITION AT THE 4TH FIRING. THE KNIFE REVERSE SWITCH WAS NOT USED, BUT THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | GRAY RELOAD FOR 60MM ECHELON | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A NEPHRECTOMY PROCEDURE, UPON OPENING THE RELOAD, THE METAL TRAY ON THE RELOAD WAS NOTED TO BE DISLODGED FROM THE PLASTIC CARTRIDGE. A NEW RELOAD WAS OPENED. THE PROCEDURE WAS COMPLETED WITH NO PATIENT CONSEQUENCES.
15-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE KNIFE MOVED ALL THE WAY BUT STOPPED MOVING BACK TO HOME POSITION AT THE FIRST FIRING OF FEEA. THE DEVICE WAS USED ON THE JEJUNUM. THE CARTRIDGE COLOR WAS BLUE. THE KNIFE WAS RETURNED BY KNIFE REVERSE SWITCH. THE DEVICE CUT THE TISSUE AND THE STAPLES FORMED PROPERLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
08-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DUODENAL SWITCH PROCEDURE, THE SURGEON WAS GOING TO USE THE POWERED ECHELON FOR A FIRST APPLICATION ON THE DUODENUM WITH A BLUE RELOAD. HE SAYS HE PRESSED THE ACTIVATION BUTTON AND NOTHING HAPPENED LIKE THE BATTERY WAS UNCHARGED. SURGEON WAS ASKED IF THE BLADE HAD ENGAGED ITSELF AND BLOCKED AND HE SAID NO, THE BATTERY WAS JUST DEAD. SURGEON SAID ONCE IT WAS REMOVED FROM THE PATIENT, THEY "PLAYED WITH DEVICE" AND TRIED ANYTHING JUST TO SEE IF THEY COULD UNDERSTAND WHAT WAS GOING ON AND THEN SOMEONE PLAYED WITH THE MANUAL OVERRIDE. SURGERY WAS DELAYED BY FIVE MINUTES, ANOTHER DEVICE WAS USED, AND PROCEDURE WAS SUCCESSFULLY COMPLETED. THE PERSON WHO CLEANED DEVICE AFTER THE CASE MENTIONED THAT THE BATTERY WAS VERY HOT. THERE WAS NO PATIENT CONSEQUENCE.
08-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Excessive Heating | A1002 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DUODENAL SWITCH PROCEDURE, THE SURGEON WAS GOING TO USE THE POWERED ECHELON FOR A FIRST APPLICATION ON THE DUODENUM WITH A BLUE RELOAD. HE SAYS HE PRESSED THE ACTIVATION BUTTON AND NOTHING HAPPENED LIKE THE BATTERY WAS UNCHARGED. SURGEON WAS ASKED IF THE BLADE HAD ENGAGED ITSELF AND BLOCKED AND HE SAID NO, THE BATTERY WAS JUST DEAD. SURGEON SAID ONCE IT WAS REMOVED FROM THE PATIENT, THEY "PLAYED WITH DEVICE" AND TRIED ANYTHING JUST TO SEE IF THEY COULD UNDERSTAND WHAT WAS GOING ON AND THEN SOMEONE PLAYED WITH THE MANUAL OVERRIDE. SURGERY WAS DELAYED BY FIVE MINUTES, ANOTHER DEVICE WAS USED, AND PROCEDURE WAS SUCCESSFULLY COMPLETED. THE PERSON WHO CLEANED DEVICE AFTER THE CASE MENTIONED THAT THE BATTERY WAS VERY HOT. THERE WAS NO PATIENT CONSEQUENCE.
01-01-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | ECHELON FLEX 45 COMPACT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SC45A WAS DIFFICULT TO USE RELEASE BUTTON . THE DOCTOR WORRIED THIS CONDITION WILL DAMAGE THE PATIENT, SO CHANGED ANOTHER ONE. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
15-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, "WHEN ASSEMBLING THE ECHELON STAPLER WITH THE RELOAD TO THE JAW, IT PRESENTED DIFFICULTY ON THE ASSEMBLING." THE RELOAD WAS CHANGED AND THE STAPLING ON THE TISSUE WAS PERFORMED WITHOUT DIFFICULTIES. IN THE SECOND RELOAD PLACEMENT, AGAIN THERE IS DIFFICULTY IN THE ASSEMBLY AND THE STAPLING DRAGGED THE TISSUE, LEAVING THE EDGES BLEEDING. THERE WAS A FIRST LINE OF POORLY FORMED STAPLES AFTER CUTTING. BLEEDING WAS CONTROLLED WITH LT300 CLIP. IT WAS UNKNOWN HOW MUCH BLEEDING OCCURRED BUT NO TRANSFUSION WAS REQUIRED AND THERE WAS NO CHANGE TO POST OP CARE OF PATIENT. THERE WAS NO PATIENT CONSEQUENCE.
15-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, "WHEN ASSEMBLING THE ECHELON STAPLER WITH THE RELOAD TO THE JAW, IT PRESENTED DIFFICULTY ON THE ASSEMBLING." THE RELOAD WAS CHANGED AND THE STAPLING ON THE TISSUE WAS PERFORMED WITHOUT DIFFICULTIES. IN THE SECOND RELOAD PLACEMENT, AGAIN THERE IS DIFFICULTY IN THE ASSEMBLY AND THE STAPLING DRAGGED THE TISSUE, LEAVING THE EDGES BLEEDING. THERE WAS A FIRST LINE OF POORLY FORMED STAPLES AFTER CUTTING. BLEEDING WAS CONTROLLED WITH LT300 CLIP. IT WAS UNKNOWN HOW MUCH BLEEDING OCCURRED BUT NO TRANSFUSION WAS REQUIRED AND THERE WAS NO CHANGE TO POST OP CARE OF PATIENT. THERE WAS NO PATIENT CONSEQUENCE.
15-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, "WHEN ASSEMBLING THE ECHELON STAPLER WITH THE RELOAD TO THE JAW, IT PRESENTED DIFFICULTY ON THE ASSEMBLING." THE RELOAD WAS CHANGED AND THE STAPLING ON THE TISSUE WAS PERFORMED WITHOUT DIFFICULTIES. IN THE SECOND RELOAD PLACEMENT, AGAIN THERE IS DIFFICULTY IN THE ASSEMBLY AND THE STAPLING DRAGGED THE TISSUE, LEAVING THE EDGES BLEEDING. THERE WAS A FIRST LINE OF POORLY FORMED STAPLES AFTER CUTTING. BLEEDING WAS CONTROLLED WITH LT300 CLIP. IT WAS UNKNOWN HOW MUCH BLEEDING OCCURRED BUT NO TRANSFUSION WAS REQUIRED AND THERE WAS NO CHANGE TO POST OP CARE OF PATIENT. THERE WAS NO PATIENT CONSEQUENCE.
15-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, "WHEN ASSEMBLING THE ECHELON STAPLER WITH THE RELOAD TO THE JAW, IT PRESENTED DIFFICULTY ON THE ASSEMBLING." THE RELOAD WAS CHANGED AND THE STAPLING ON THE TISSUE WAS PERFORMED WITHOUT DIFFICULTIES. IN THE SECOND RELOAD PLACEMENT, AGAIN THERE IS DIFFICULTY IN THE ASSEMBLY AND THE STAPLING DRAGGED THE TISSUE, LEAVING THE EDGES BLEEDING. THERE WAS A FIRST LINE OF POORLY FORMED STAPLES AFTER CUTTING. BLEEDING WAS CONTROLLED WITH LT300 CLIP. IT WAS UNKNOWN HOW MUCH BLEEDING OCCURRED BUT NO TRANSFUSION WAS REQUIRED AND THERE WAS NO CHANGE TO POST OP CARE OF PATIENT. THERE WAS NO PATIENT CONSEQUENCE.
17-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | PWRD ECH FLEX 45MM 28CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LEFT UPPER LOBE SURGERY/LOBECTOMY, WHEN SEVERING VESSELS/ TISSUE, AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | PWRD ECH FLEX 45MM 28CM SHAFT | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AND UNKNOWN PROCEDURE, NOTED THAT THE FIRING TRIGGER WAS LOOSENED. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE CLAMP CONTOUR DIDN'T WORK IN THE OPERATING ROOM. ANOTHER CLAMP CONTOUR WAS USED TO SOLVE THE ISSUE. THE DEVICE WAS USED IN USUAL CONDITIONS OF USE. THE STAPLES WERE NOT RELEASED AFTER THE PRESSURE ON THE DEDICATED HANDLE. AFTER A VISUAL INSPECTION, THE STAPLES ARE VISIBLE ON THE CARTRIDGE; THE CLOSED HANDLE WAS BROKEN BY THE SURGEON BECAUSE HE WAS NOT ABLE TO RETRACT THE TOP CARTRIDGE. NO PATIENT CONSEQUENCE REPORTED.
01-01-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | CONTOUR CURVED CUTTER STAPLER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE CLAMP CONTOUR DIDN'T WORK IN THE OPERATING ROOM. ANOTHER CLAMP CONTOUR WAS USED TO SOLVE THE ISSUE. THE DEVICE WAS USED IN USUAL CONDITIONS OF USE. THE STAPLES WERE NOT RELEASED AFTER THE PRESSURE ON THE DEDICATED HANDLE. AFTER A VISUAL INSPECTION, THE STAPLES ARE VISIBLE ON THE CARTRIDGE; THE CLOSED HANDLE WAS BROKEN BY THE SURGEON BECAUSE HE WAS NOT ABLE TO RETRACT THE TOP CARTRIDGE. NO PATIENT CONSEQUENCE REPORTED.
01-01-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | CONTOUR CURVED CUTTER STAPLER | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE CLAMP CONTOUR DIDN'T WORK IN THE OPERATING ROOM. ANOTHER CLAMP CONTOUR WAS USED TO SOLVE THE ISSUE. THE DEVICE WAS USED IN USUAL CONDITIONS OF USE. THE STAPLES WERE NOT RELEASED AFTER THE PRESSURE ON THE DEDICATED HANDLE. AFTER A VISUAL INSPECTION, THE STAPLES ARE VISIBLE ON THE CARTRIDGE; THE CLOSED HANDLE WAS BROKEN BY THE SURGEON BECAUSE HE WAS NOT ABLE TO RETRACT THE TOP CARTRIDGE. NO PATIENT CONSEQUENCE REPORTED.
15-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN SMALL BOWEL RESECTION, THE KNIFE MOVED ALL THE WAY BUT STOPPED MOVING BACK TO HOME POSITION AND THE JAW DID NOT OPEN AT THE EIGHTH FIRING. THE KNIFE REVERSE SWITCH DID NOT FUNCTION. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. NO ADDITIONAL TREATMENT WAS REQUIRED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN SMALL BOWEL RESECTION, THE KNIFE MOVED ALL THE WAY BUT STOPPED MOVING BACK TO HOME POSITION AND THE JAW DID NOT OPEN AT THE EIGHTH FIRING. THE KNIFE REVERSE SWITCH DID NOT FUNCTION. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. NO ADDITIONAL TREATMENT WAS REQUIRED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN SMALL BOWEL RESECTION, THE KNIFE MOVED ALL THE WAY BUT STOPPED MOVING BACK TO HOME POSITION AND THE JAW DID NOT OPEN AT THE EIGHTH FIRING. THE KNIFE REVERSE SWITCH DID NOT FUNCTION. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. NO ADDITIONAL TREATMENT WAS REQUIRED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | BLUE RELOAD FOR ECHELON 60 | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AND UNKNOWN PROCEDURE, AFTER OPENED THE PACKING, FOUND THE CARTRIDGE PAN WAS LOOSE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | PWRD ECH FLEX 45MM 28CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOBECTOMY SURGERY, WHEN SEVERING LUNG TISSUE AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO ANOTHER ECR45B, THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2020 | Malfunction | ethicon endo-surgery | 08-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY CUSTOMER THAT DURING AN UNKNOWN PROCEDURE, THE POWERED VASCULAR STAPLER, PVE35A, WOULD NOT OPEN "AFTER PLACING CARTRIDGE IN." THE DEVICE WAS NOT LOCKED ON TISSUE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
16-12-2020 | Injury | intuitive surgical | 10-01-2021 | ENDOWRIST | Adverse Event Without Identified Device or Use Problem | A24 | Injury | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
16-12-2020 | Injury | intuitive surgical | 10-01-2021 | ENDOWRIST | Insufficient Information | E2401 | Injury | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
04-09-2020 | Injury | covidien lp llc north haven | 11-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE HAD AN UNKNOWN ISSUE. THE PATIENT NEEDED SURGICAL RE-INTERVENTION TO RESOLVE THE ISSUE.
17-10-2020 | Injury | covidien lp llc north haven | 11-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, PATIENTS HAD TO BE RE-OPERATED AFTER USING THE DEVICE. THERE WAS NO PATIENT INJURY.
30-01-2020 | Injury | covidien lp llc north haven | 11-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, (YEAR JANUARY 2019 TO JULY 2019) THIS STUDY AIMED TO ASSESS THE OUTCOME OF A SINGLE-PORT ROBOTIC PARTIAL AND RADICAL NEPHRECTOMIES FOR RENAL CORTICAL TUMORS: INITIAL CLINICAL EXPERIENCE. THE ENDO-GIA STAPLER WAS USED TO STAPLE ACROSS THE RENAL VEIN AND A SEPARATE LOAD WAS USED TO STAPLE ACROSS THE RENAL ARTERY. THE SPECIMEN WAS EXTRACTED AFTER PLACING IN AN ENDOBAG. THERE WERE A TOTAL OF 16 PATIENTS AND THE FOLLOWING POST-OPERATIVE COMPLICATIONS WERE REPORTED: BLOOD LOSS UP TO 800 ML WITH ONE PATIENT REQUIRING BOTH A BLOOD TRANSFUSION ANTIBIOTICS. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED: TWO PATIENTS REQUIRING ANTIBIOTICS FOR POSSIBLE SEPSIS (CULTURES WERE NEGATIVE), ONE CONVERSION TO OPEN PROCEDURE, TWO POST-OPERATIVE COMPLICATIONS RELATED TO END-STAGE RENAL DISEASE.
14-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROINTESTINAL TRACT RESECTION, THE DEVICE WOULD NOT FIRE. THE GREEN BUTTON COULD NOT BE PRESSED NOR HANDLE COULD BE SQUEEZED. A TRACE OF PREFIRE WAS OBSERVED. JAWS LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROINTESTINAL TRACT RESECTION, THE DEVICE WOULD NOT FIRE. THE GREEN BUTTON COULD NOT BE PRESSED NOR HANDLE COULD BE SQUEEZED. A TRACE OF PREFIRE WAS OBSERVED. JAWS LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROINTESTINAL TRACT RESECTION, THE DEVICE WOULD NOT FIRE. THE GREEN BUTTON COULD NOT BE PRESSED NOR HANDLE COULD BE SQUEEZED. A TRACE OF PREFIRE WAS OBSERVED. JAWS LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Injury | celestica de monterrey s.a de c.v | 11-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF LUNG CANCER, WHILE DETACHING THE LUNG PARENCHYMA, THE STAPLER WAS FIRED BUT HAD AN INCOMPLETE STAPLE LINE ON THE DISTAL PART. IT WAS ALSO NOTED THAT SOME OF THE STAPLES DID NOT FORM. MANUAL SUTURING WAS DONE TO COMPLETE THE CASE.
14-12-2020 | Injury | celestica de monterrey s.a de c.v | 11-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF LUNG CANCER, WHILE DETACHING THE LUNG PARENCHYMA, THE STAPLER WAS FIRED BUT HAD AN INCOMPLETE STAPLE LINE ON THE DISTAL PART. IT WAS ALSO NOTED THAT SOME OF THE STAPLES DID NOT FORM. MANUAL SUTURING WAS DONE TO COMPLETE THE CASE.
08-09-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | CIRCULAR STAPLER | Output Problem | A09 | Appropriate Term / Code Not Available | Event Description: LEAK TESTED WAS PERFORMED AND POSITIVE FOR LEAK, TRIED TO REPAIR THE LEAK WITH SUTURE BUT CANNOT DO IT SO DOCTOR HAD TO DO A TEMP LOOP ILEOSTOMY. MANUFACTURER RESPONSE FOR ETHICON CURVED CDH25, (BRAND NOT PROVIDED) (PER SITE REPORTER). PRODUCT PICKED UP BY REP FOR EVALUATION.
08-09-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | CIRCULAR STAPLER | Therapeutic or Diagnostic Output Failure | A090406 | Failure to Anastomose | Event Description: 25MM CURVED EEA STAPLER WAS USED AND AFTER DEPLOYING IT, SURGEON HAS LEAK AROUND THE ANASTOMOSIS, COULD NOT HAVE A REASSURING LEAK TEST (HAD MULTIPLE BUBBLES), AND HAD TO END ILEOSTOMY BECAUSE OF THE LEAK. TRIED TO SUTURE BUT IT WAS NOT WORKING AND THE LEAK COULD NOT BE FIXED SO HAD TO DO END ILEOSTOMY. MANUFACTURER RESPONSE FOR ETHICON CURVED STAPLER, (BRAND NOT PROVIDED) (PER SITE REPORTER): PRODUCT GIVEN TO REP FOR EVALUATION.
01-01-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AFTER THE SHOT, THE STAPLER DID NOT FORM THE COMPLETE ANASTOMOSIS. THERE WERE NO PATIENT CONSEQUENCES.
14-12-2020 | Injury | us surgical puerto rico | 11-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF LUNG CANCER, WHILE DETACHING THE LUNG PARENCHYMA, THE STAPLER WAS FIRED BUT HAD AN INCOMPLETE STAPLE LINE ON THE DISTAL PART. IT WAS ALSO NOTED THAT SOME OF THE STAPLES DID NOT FORM. MANUAL SUTURING WAS DONE TO COMPLETE THE CASE.
05-01-2021 | Malfunction | ethicon endo-surgery | 11-01-2021 | ENDOPATH | Physical Resistance/Sticking | A0509 | No Clinical Signs, Symptoms or Conditions | Event Description: ELDERLY MALE WITH HISTORY OF HYPERTENSION, DIABETES AND MALIGNANT NEOPLASM OF RECTOSIGMOID COLON. IN OR FOR LAPAROSCOPIC ASSISTED LOW ANTERIOR RECTOSIGMOID RESECTION. ETHICON ENDOPATH 45MM FLEX STAPLER USED AND GOT STUCK ON THE 3RD FIRE. SURGEON WAS THE ONE USING THE PRODUCT IN OR; STAPLER WAS UNABLE TO BE ENGAGED. MINIMAL DELAY IN PROCEDURE UNTIL NEW PRODUCT OBTAINED. NO KNOWN HARM TO PATIENT.
11-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ANTERIOR RESECTION, STAPLE MALFORMATION. THERE WAS NO DELAY AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
30-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | ECHELON FLEX | Failure to Fire | A050501 | Insufficient Information | Event Description: ELDERLY MALE WITH HISTORY OF DIABETES, HYPERTENSION, CORONARY ARTERY DISEASE AND LUNG CANCER. SURGERY WAS A THORACOSCOPY, LEFT UPPER LOBECTOMY- WHEN CLOSING, THE ETHICON POWER PLUS ARTICULATING ENDOSCOPIC LINEAR CUTTER WOULD NOT FIRE 60 BLUE STAPLE. NO KNOWN HARM TO PATIENT, SLIGHT DELAY RETRIEVING ANOTHER PRODUCT.
28-11-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLE LINE OPENED.
14-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN TRANSECTING THE STOMACH ON A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, AFTER PUSHING THE GREEN BUTTON TO READY THE FIRING, THE TOGGLE BUTTON WAS PUSHED DOWN. HOWEVER, THE BLADE OF THE RELOAD STOPPED AS SOON AS THE TOGGLE BUTTON WAS PUSHED DOWN. THE BUTTON WAS PUSHED AGAIN SEVERAL TIMES BUT IT DID NOT MOVE AT ALL. THE BUTTON WAS PULLED UP TO RETRACT THE BLADE AND THE SCRUB NURSE GOT THE DEVICE AND DISCONNECTED THE RELOAD FROM THE ADAPTER. THE CIRCULATING NURSE PREPARED A NEW RELOAD AND THE SCRUB NURSE CONNECTED A NEW RELOAD WITH THE USED ADAPTER. THE OPERATION WAS THEN ABLE TO COMPLETE THE TRANSECTION. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | EEA ORVIL | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC REVISION OF ROUX EN Y, THE AESTHESIA WORKER PASSED THE ORVIL DOWN THE ESOPHAGUS WITH THE WRONG END FIRST, CAUSING THE ORVIL ANVIL TO BE DISCONNECTED FROM THE TUBING. SURGICAL TIME EXTENDED 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC REVISION OF ROUX EN Y, THE AESTHESIA WORKER PASSED THE ORVIL DOWN THE ESOPHAGUS WITH THE WRONG END FIRST, CAUSING THE ORVIL ANVIL TO BE DISCONNECTED FROM THE TUBING. SURGICAL TIME EXTENDED 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE I-BEAM GOT STUCK AND STOPPED MOVING 2 TIMES. THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE. SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
16-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC LOW RECTECTOMY SURGERY, AFTER FIRED, NOTED THAT THE TISSUE WAS NOT CUT, BUT THERE WERE SOME STAPLERS DEPLOYED ON THE TISSUE, THEN CHANGED ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
08-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER THE COLON WAS FIRST RESECTION DURING AN OPEN RIGHT HEMICOLECTOMY, THE BLACK RETURN KNOB COULD NOT BE PULLED BACK TO THE END AND THE JAWS COULD NOT BE USED. THE HANDLE AND RELOAD COULD NOT BE DISCONNECTED, SO RELOAD WAS REPLACED WITH A NEW HANDLE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER THE COLON WAS FIRST RESECTION DURING AN OPEN RIGHT HEMICOLECTOMY, THE BLACK RETURN KNOB COULD NOT BE PULLED BACK TO THE END AND THE JAWS COULD NOT BE USED. THE HANDLE AND RELOAD COULD NOT BE DISCONNECTED, SO RELOAD WAS REPLACED WITH A NEW HANDLE. THERE WAS NO PATIENT INJURY.
30-12-2019 | Injury | covidien lp llc north haven | 11-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS STUCK SO THE SURGEON SUTURED BY HAND.
30-12-2019 | Injury | us surgical puerto rico | 11-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS STUCK SO THE SURGEON SUTURED BY HAND.
15-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | PROXIMATE*RL LINEAR STAPLER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT RESECTING BOWEL AND WHEN USING THE STAPLER, THE DEVICE WAS CLAMPED AROUND THE BOWEL AND BOTH THE CLOSING HANDLE AND FIRING HANDLE WERE PRESSED SIMULTANEOUSLY WHICH CAUSE THE DEVICE TO LOCK. THE GREY RELEASE BUTTON WOULD NOT WORK.
08-12-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER THE COLON WAS FIRST RESECTION DURING AN OPEN RIGHT HEMICOLECTOMY, THE BLACK RETURN KNOB COULD NOT BE PULLED BACK TO THE END AND THE JAWS COULD NOT BE USED. THE HANDLE AND RELOAD COULD NOT BE DISCONNECTED, SO RELOAD WAS REPLACED WITH A NEW HANDLE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER THE COLON WAS FIRST RESECTION DURING AN OPEN RIGHT HEMICOLECTOMY, THE BLACK RETURN KNOB COULD NOT BE PULLED BACK TO THE END AND THE JAWS COULD NOT BE USED. THE HANDLE AND RELOAD COULD NOT BE DISCONNECTED, SO RELOAD WAS REPLACED WITH A NEW HANDLE. THERE WAS NO PATIENT INJURY.
29-11-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON LAPAROSCOPIC CHOLECYSTECTOMY, WHILE BEING APPLIED ACROSS THE CYSTIC DUCT, THE SURGEON COULD NOT FIRE THE DEVICE ONCE PLACED AROUND THE TISSUE. IT WAS STATED THAT IT SEEMED THAT RELOAD HAD ALREADY BEEN FIRED AND THE HANDLE WAS LOCKED OUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-11-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | TA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON LAPAROSCOPIC CHOLECYSTECTOMY, WHILE BEING APPLIED ACROSS THE CYSTIC DUCT, THE SURGEON COULD NOT FIRE THE DEVICE ONCE PLACED AROUND THE TISSUE. IT WAS STATED THAT IT SEEMED THAT RELOAD HAD ALREADY BEEN FIRED AND THE HANDLE WAS LOCKED OUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Injury | us surgical puerto rico | 11-01-2021 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE DEVICE ABLE TO BE FIRED HOWEVER MUCOUS MEMBRANE WAS CAUGHT IN THE CLAMP COVER, THE MUCOUS MEMBRANE ENTERED THE GROOVE THAT THE CLAMP COVER WENT THROUGH, THE KNIFE RETURNED TO THE INITIAL POSITION WITHOUT ANY PROBLEMS AND THE JAWS WERE OPENED, BUT THE MUCOUS MEMBRANE ENTERED INTO THE GROOVE, WHILE REMOVING THE CARTRIDGE, IT GOT CAUGHT AND THE MUCOUS MEMBRANE GOT DAMAGED. TO RESOLVE THE ISSUE THE SURGEON HAD TO SUTURE THE TISSUE. THERE WERE TISSUE DAMAGE AND TISSUE LOSS DUE TO THE DEVICE ISSUE.
09-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTAL LAPAROSCOPY, THE SURGEON DISCOVERED THAT THE CLAMP USED FOR ANASTOMOSIS ONLY STAPLED IN THE CENTER AND NOT AT THE EXTREMITIES. THE SURGEON HAD TO REOPEN THE PART OF THE COLON THAT HAD BEEN STAPLED WITH THE FIRST DEVICE THEN THE CLAMP FOR ANASTOMOSIS WAS CHANGED TO COMPLETE THE PROCEDURE. 30-MIN DELAY. COMPETITIVE PROCEDURE USED. ADDITIONAL INVASIVE ACTS. THERE WERE NO PATIENT CONSEQUENCES.
24-11-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | EEA | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A COLORECTAL SURGERY, WHEN OPENING THE DEVICE WITH THE BLACK ROTATION KNOB THAT WAS TURNED WITHOUT PROBLEMS, THE TROCAR FAILED TO PIERCE THE TISSUE. ANOTHER PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2020 | Malfunction | covidien lp llc north haven | 11-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE WOULD NOT FIRE. ANOTHER RELOAD WAS USED AND SAME HANDLE WAS FIRED WITH NO PROBLEMS. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, ON THE DUODENUM, THE INSERTION OPENING WAS CLOSED, BUT FIRING WAS ONLY PERFORMED HALFWAY, THE KNIFE RETURNED BEFORE SOME STAPLES FORMED, THE STAPLE THAT CAME OUT AT THE END SEEMED TO HAVE JUST BEEN PIERCED, THE JAWS OF THE DEVICE LOCKED ON TISSUE. ANOTHER CARTRIDGE WAS TAKEN OUT AND THE INSERTION OPENING WAS CLOSED WITH TWO FIRINGS AND THE SAME HANDLE, ADAPTER AND BATTERY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, ON THE DUODENUM, THE INSERTION OPENING WAS CLOSED, BUT FIRING WAS ONLY PERFORMED HALFWAY, THE KNIFE RETURNED BEFORE SOME STAPLES FORMED, THE STAPLE THAT CAME OUT AT THE END SEEMED TO HAVE JUST BEEN PIERCED, THE JAWS OF THE DEVICE LOCKED ON TISSUE. ANOTHER CARTRIDGE WAS TAKEN OUT AND THE INSERTION OPENING WAS CLOSED WITH TWO FIRINGS AND THE SAME HANDLE, ADAPTER AND BATTERY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, ON THE DUODENUM, THE INSERTION OPENING WAS CLOSED, BUT FIRING WAS ONLY PERFORMED HALFWAY, THE KNIFE RETURNED BEFORE SOME STAPLES FORMED, THE STAPLE THAT CAME OUT AT THE END SEEMED TO HAVE JUST BEEN PIERCED, THE JAWS OF THE DEVICE LOCKED ON TISSUE. ANOTHER CARTRIDGE WAS TAKEN OUT AND THE INSERTION OPENING WAS CLOSED WITH TWO FIRINGS AND THE SAME HANDLE, ADAPTER AND BATTERY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2020 | Injury | us surgical puerto rico | 11-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y, WHILE TRANSECTING THE ILEUM, THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL PART. THE STAPLE LINE WAS FIXED USING ANOTHER RELOAD AND SUTURING.
13-12-2020 | Injury | us surgical puerto rico | 11-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y, WHILE TRANSECTING THE ILEUM, THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL PART. THE STAPLE LINE WAS FIXED USING ANOTHER RELOAD AND SUTURING.
13-12-2020 | Injury | us surgical puerto rico | 11-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y, WHILE TRANSECTING THE ILEUM, THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE AT THE DISTAL PART. THE STAPLE LINE WAS FIXED USING ANOTHER RELOAD AND SUTURING.
14-12-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE HANDLE BROKE OFF AND WAS NOT ABLE TO FIRE. IT WAS NOTED THAT AFTER LOADING THE RELOAD THE COLON WAS KEPT BETWEEN JAWS OF THE STAPLERS AND TRY TO FIRE BY SQUEEZING THE HANDLE AFTER PRESSING THE GREEN BUTTON. IT DID NOT FIRE AND HANDLE BROKEN INSIDE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 11-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE HANDLE BROKE OFF AND WAS NOT ABLE TO FIRE. IT WAS NOTED THAT AFTER LOADING THE RELOAD THE COLON WAS KEPT BETWEEN JAWS OF THE STAPLERS AND TRY TO FIRE BY SQUEEZING THE HANDLE AFTER PRESSING THE GREEN BUTTON. IT DID NOT FIRE AND HANDLE BROKEN INSIDE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON INFORMED US THAT THE KNIFE OF THE STAPLER DIDN'T RETURN, NOT EVEN WHEN USING THE REVERSE BUTTON. THE MANUAL OVERRIDE HAD TO BE USED. THIS ISSUE HAPPENED DURING THE STAPLING OF A STOMACH WHILE REMOVING A GASTRIC BAND. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
01-01-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON INFORMED US THAT THE KNIFE OF THE STAPLER DIDN'T RETURN, NOT EVEN WHEN USING THE REVERSE BUTTON. THE MANUAL OVERRIDE HAD TO BE USED. THIS ISSUE HAPPENED DURING THE STAPLING OF A STOMACH WHILE REMOVING A GASTRIC BAND. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
01-01-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON INFORMED US THAT THE KNIFE OF THE STAPLER DIDN'T RETURN, NOT EVEN WHEN USING THE REVERSE BUTTON. THE MANUAL OVERRIDE HAD TO BE USED. THIS ISSUE HAPPENED DURING THE STAPLING OF A STOMACH WHILE REMOVING A GASTRIC BAND. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
17-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY UPON FIRING, THEY NOTICED A STAPLE MISSING FROM THE OUTER ROW 2/3 DOWN STAPLE LINE. A CLIP WAS APPLIED TO COMPLETE THE CASE. PATIENT CONSEQUENCE WAS NOT REPORTED.
11-11-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY FOR PANCREATIC SPACE-OCCUPYING LESIONS, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE( NOT B SHAPE). CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-11-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SUPER LOW ANTERIOR RESECTION, THE KNIFE STOPPED MOVING FORWARD ON THE WAY OF THE FIRING. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
11-11-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SUPER LOW ANTERIOR RESECTION, THE KNIFE STOPPED MOVING FORWARD ON THE WAY OF THE FIRING. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
17-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, THE DOCTOR EXPERIENCED A DEVICE MISFIRE AND SUBSEQUENT LOCK OUT WHEN FIRING PLEE60A. HE WENT THROUGH ALL BAILOUT PROCEDURES, REVERSE, BATTERY, REVERSE, AND HAD TO RESORT TO MANUAL OVERRIDE. AFTER CRANKING THE MANUAL OVERRIDE, THE DEVICE JAWS STILL WOULD NOT OPEN. FRAGMENTS WERE CREATED BUT WERE EASILY REMOVED. THE PROCEDURE WAS DELAYED BY THIRTY MINUTES HOWEVER THE PROCEDURE WAS SUCCESSFULLY COMPLETED. PATIENT CONSEQUENCE WAS NOT REPORTED.
17-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, THE DOCTOR EXPERIENCED A DEVICE MISFIRE AND SUBSEQUENT LOCK OUT WHEN FIRING PLEE60A. HE WENT THROUGH ALL BAILOUT PROCEDURES, REVERSE, BATTERY, REVERSE, AND HAD TO RESORT TO MANUAL OVERRIDE. AFTER CRANKING THE MANUAL OVERRIDE, THE DEVICE JAWS STILL WOULD NOT OPEN. FRAGMENTS WERE CREATED BUT WERE EASILY REMOVED. THE PROCEDURE WAS DELAYED BY THIRTY MINUTES HOWEVER THE PROCEDURE WAS SUCCESSFULLY COMPLETED. PATIENT CONSEQUENCE WAS NOT REPORTED.
17-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, THE DOCTOR EXPERIENCED A DEVICE MISFIRE AND SUBSEQUENT LOCK OUT WHEN FIRING PLEE60A. HE WENT THROUGH ALL BAILOUT PROCEDURES, REVERSE, BATTERY, REVERSE, AND HAD TO RESORT TO MANUAL OVERRIDE. AFTER CRANKING THE MANUAL OVERRIDE, THE DEVICE JAWS STILL WOULD NOT OPEN. FRAGMENTS WERE CREATED BUT WERE EASILY REMOVED. THE PROCEDURE WAS DELAYED BY THIRTY MINUTES HOWEVER THE PROCEDURE WAS SUCCESSFULLY COMPLETED. PATIENT CONSEQUENCE WAS NOT REPORTED.
17-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, THE DOCTOR EXPERIENCED A DEVICE MISFIRE AND SUBSEQUENT LOCK OUT WHEN FIRING PLEE60A. HE WENT THROUGH ALL BAILOUT PROCEDURES, REVERSE, BATTERY, REVERSE, AND HAD TO RESORT TO MANUAL OVERRIDE. AFTER CRANKING THE MANUAL OVERRIDE, THE DEVICE JAWS STILL WOULD NOT OPEN. FRAGMENTS WERE CREATED BUT WERE EASILY REMOVED. THE PROCEDURE WAS DELAYED BY THIRTY MINUTES HOWEVER THE PROCEDURE WAS SUCCESSFULLY COMPLETED. PATIENT CONSEQUENCE WAS NOT REPORTED.
17-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, THE DOCTOR EXPERIENCED A DEVICE MISFIRE AND SUBSEQUENT LOCK OUT WHEN FIRING PLEE60A. HE WENT THROUGH ALL BAILOUT PROCEDURES, REVERSE, BATTERY, REVERSE, AND HAD TO RESORT TO MANUAL OVERRIDE. AFTER CRANKING THE MANUAL OVERRIDE, THE DEVICE JAWS STILL WOULD NOT OPEN. FRAGMENTS WERE CREATED BUT WERE EASILY REMOVED. THE PROCEDURE WAS DELAYED BY THIRTY MINUTES HOWEVER THE PROCEDURE WAS SUCCESSFULLY COMPLETED. PATIENT CONSEQUENCE WAS NOT REPORTED.
16-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | PROXIMATE*75MM LINEAR CUTTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AFTER USING TLC75 TO CREATE ANASTOMOSIS, OBSERVED CUT LINE FALL APART. CUSTOMER IDENTIFIED THE EVENT AS A MISFIRE. TO ADDRESS THIS ADDITIONAL COLON WAS RESECTED.
14-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE KNIFE STOPPED MOVING FORWARD WITH AN UNEXPECTED SOUND DURING USE. A PART OF THE DEVICE WAS BROKEN OFF. IT WAS RETRIEVED FROM THE PATIENT. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE DEVICE WAS USED DUODENUM. THE TARGET TISSUE WAS COMPRESSED BEFORE THE FIRING AND IT WAS NOT VERY THICK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS NOT CONVERTED TO OPEN. THE PATIENT IS STABLE. WHEN THE SURGEON CHECKED THE VIDEO OF THE SURGERY, IT WAS FOUND THE CARTRIDGE WAS NOT LOADED INTO THE JAW PROPERLY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
14-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE KNIFE STOPPED MOVING FORWARD WITH AN UNEXPECTED SOUND DURING USE. A PART OF THE DEVICE WAS BROKEN OFF. IT WAS RETRIEVED FROM THE PATIENT. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE DEVICE WAS USED DUODENUM. THE TARGET TISSUE WAS COMPRESSED BEFORE THE FIRING AND IT WAS NOT VERY THICK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS NOT CONVERTED TO OPEN. THE PATIENT IS STABLE. WHEN THE SURGEON CHECKED THE VIDEO OF THE SURGERY, IT WAS FOUND THE CARTRIDGE WAS NOT LOADED INTO THE JAW PROPERLY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
14-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE KNIFE STOPPED MOVING FORWARD WITH AN UNEXPECTED SOUND DURING USE. A PART OF THE DEVICE WAS BROKEN OFF. IT WAS RETRIEVED FROM THE PATIENT. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE DEVICE WAS USED DUODENUM. THE TARGET TISSUE WAS COMPRESSED BEFORE THE FIRING AND IT WAS NOT VERY THICK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS NOT CONVERTED TO OPEN. THE PATIENT IS STABLE. WHEN THE SURGEON CHECKED THE VIDEO OF THE SURGERY, IT WAS FOUND THE CARTRIDGE WAS NOT LOADED INTO THE JAW PROPERLY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
14-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE KNIFE STOPPED MOVING FORWARD WITH AN UNEXPECTED SOUND DURING USE. A PART OF THE DEVICE WAS BROKEN OFF. IT WAS RETRIEVED FROM THE PATIENT. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE DEVICE WAS USED DUODENUM. THE TARGET TISSUE WAS COMPRESSED BEFORE THE FIRING AND IT WAS NOT VERY THICK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS NOT CONVERTED TO OPEN. THE PATIENT IS STABLE. WHEN THE SURGEON CHECKED THE VIDEO OF THE SURGERY, IT WAS FOUND THE CARTRIDGE WAS NOT LOADED INTO THE JAW PROPERLY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
17-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, UPON FIRING, THEY NOTICED THREE STAPLES MISSING FROM THE OUTER ROW 2/3 DOWN STAPLE LINE. HE RECLIPPED TO COMPLETE THE CASE. PATIENT CONSEQUENCE WAS NOT REPORTED.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Adhesion(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Anemia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Bacterial Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Calcium Deposits/Calcification | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Diarrhea | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Purulent Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Granuloma | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Low Blood Pressure/ Hypotension | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Septic Shock | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Seroma | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Tachycardia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Vomiting | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Chills | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hernia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Peritonitis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Impaired Healing | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Ventilator Dependent | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Obstruction/Occlusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Diaphoresis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Respiratory Failure | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Cramps | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Ascites | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Bowel Perforation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Constipation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Thrombosis/Thrombus | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Rectal Anastomotic Leakage | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Cramp(s) /Muscle Spasm(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Swelling/ Edema | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
26-12-2018 | Injury | us surgical puerto rico | 11-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR THERAPEUTIC TREATMENT OF LAPARSCOPIC SIGMOID COLECTOMY WITH COLOPROCTOSTOMY ANASTOMOSIS. IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT EXPERIENCED ABDOMINAL PAIN, FLUID AND AIR IN THE PERITONEUM OF THE PATIENTS PELVIS CONSISTENT WITH A PERFORATED BOWEL, TWO LITERS OF STOOL REMOVED FROM ABDOMEN, ANASTOMOTIC DEFECT ON THE PATIENTS RIGHT POSTERIOR LATERAL ASPECT, RESPIRATORY FAILURE, 5 LITERS OF BLOOD AND CLOT WAS REMOVED FROM THE PATIENTS ABDOMEN, BLEEDING AT LEFT GONADAL VESSEL, HEMORRHAGIC SHOCK, ANASTOMOSIS LEAK LOCATED IN THE GASTRIC INTESTINE, DRAINAGE OF 600ML OF HAZY YELLOW COLORED FLUID AND 20 ML OF THICK CREAMY YELLOW FLUID FROM THE PATIENTS LEFT FLANK OF THEIR ABDOMEN, INFECTION, THROMBOSIS, COLITIS, BACTERIAL INFECTION, INFLAMMATION, ABSCESS, PLEURAL EFFUSION, BIBASILAR ATELECTASIS, ILEUS, OPEN WOUND, PROSTATIC CALCIFICATIONS, DIVERTICULITIS IN COLON, HERNIA, OBSTRUCTION, DISTENTION, BLOATING, CONSTIPATION, STOOL PERITONITIS, SEPTIC SHOCK, SIRS [SYSTEMIC INFLAMMATORY RESPONSE SYNDROME], TACHYCARDIA, ANXIETY, CRAMPS, DIAPHORETIC, PALE, HYPOTENSION, NAUSEA, FEVER, NECROTIC TISSUE, GRANULOMA, EDEMA, THROMBOCYTOSIS, ANEMIA, CHILLS, LEUKOCYTOSIS, ADHESIONS, VOMITING, SEROMA AND DIARRHEA. POST-OPERATIVE PATIENT TREATMENT INCLUDED CT SCAN, EXPLORATORY LAPAROTOMY, LOW ANTERIOR RESECTION WAS CREATED WITH AN END COLOSTOMY AND DRAIN PLACEMENT, PLACED ON VENTILATOR, IRRIGATION AND DEBRIDEMENT, X-RAY, MRI, MEDICATIONS, HOSPITALIZATION, IV FLUIDS, WASHOUT OF ABDOMEN WITH EVACUATION OF HEMATOMA CONTROL OF BLEED, TPN [TOTAL PARENTERAL NUTRITION], NG TUBE PLACEMENT, IV FLUIDS, INSERTION OF TWO CHEST TUBES, TRANSFUSION OF 2 UNITS OF FFP [FRESH FROZEN PLASMA], 5 UNITS OF PACKED RED BLOOD CELLS, IV ZOSYN [ANTIBIOTIC] VIA PICC LINE (INSERTED 1/22), ORAL LINEZOLID [ANTIBIOTIC], METRONIDAZOLE, CIPROFLOXACIN, MYCELEX [FOR THRUSH], BRONCHODILATOR THERAPY, WOUND CARE INCLUDING DAKIN¿S SOLUTION AND WET TO DRY DRESSINGS, IV LASIX FOR TISSUE EDEMA, MONITORING OF LAB WORK, ELECTROLYTE REPLACEMENT AS NEEDED, PAIN MANAGEMENT INCLUDING MORPHINE PCA, AND NUTRITION SUPPORT WITH ADVANCEMENT OF DIET AS TOLERATED, HOME HEALTH CARE, EXPLORATORY LAPAROTOMY WITH ENTEROLYSIS (45 MINUTES), SMALL BOWEL RESECTION WITH ENTEROENTEROSTOMY, TAKEDOWN END COLOSTOMY WITH PARTIAL COLECTOMY AND COLOPROCTOSTOMY PELVIC ANASTAMOSIS, MOBILIZATION OF SPLENIC FLEXURE, FLEXIBLE SIGMOIDOSCOPIC EXAMINATION, AND RECONSTRUCTION OF ABDOMINAL WALL WITH VENTRAL HERNIA REPAIR.
17-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, UPON FIRING, THEY NOTICED FOUR STAPLES MISSING FROM THE OUTER ROW, 2/3 DOWN STAPLE LINE. THEY OVERSEWED TO COMPLETE THE CASE. PATIENT CONSEQUENCE WAS NOT REPORTED.
17-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC LEFT HEMIHEPATECTOMY SURGERY, AFTER FIRED, NOTED THAT THE ANVIL WAS BENT. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
09-11-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RADICAL RESECTION OF RECTAL CANCER SURGERY, CS40G COULD NOT FIRE WHEN SEVERING RECTUM TISSUE ON THE FIRST FIRING. CHANGED TO A NEW ONE TO COMPLETE THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES.
11-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED STAPLES MALFORMED. CHANGED TO NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. THIS CASE IS FROM HEALTH AUTHORITY.
11-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED STAPLES MALFORMED. CHANGED TO NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. THIS CASE IS FROM HEALTH AUTHORITY.
18-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE PVS WAS FIRED. BLADE DID NOT RETURN. REVERSE SLIDE BUTTON DID NOT WORK. SURGEON HAD TO USE THE MANUAL BAIL OUT TO GET DEVICE OFF VESSEL. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
18-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE PVS WAS FIRED. BLADE DID NOT RETURN. REVERSE SLIDE BUTTON DID NOT WORK. SURGEON HAD TO USE THE MANUAL BAIL OUT TO GET DEVICE OFF VESSEL. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
18-12-2020 | Malfunction | ethicon endo-surgery | 11-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE PVS WAS FIRED. BLADE DID NOT RETURN. REVERSE SLIDE BUTTON DID NOT WORK. SURGEON HAD TO USE THE MANUAL BAIL OUT TO GET DEVICE OFF VESSEL. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
21-12-2020 | Injury | intuitive surgical | 11-01-2021 | ENDOWRIST | Difficult to Open or Close | A051104 | No Consequences Or Impact To Patient | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
07-12-2020 | Malfunction | intuitive surgical | 11-01-2021 | SUREFORM 60 STAPLER | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: STAPLER APPLIED TO TISSUE AND CLAMPED BUT DID NOT STAPLE. STAPLER REMOVED AND BLEEDING NOTED. NEW STAPLER OBTAINED AND APPLIED TO RESOLVE ISSUE. FDA SAFETY REPORT ID #: (B)(4).
07-12-2020 | Malfunction | intuitive surgical | 11-01-2021 | SUREFORM 60 STAPLER | Failure to Form Staple | A050704 | Tissue Damage | Event Description: STAPLER APPLIED TO TISSUE AND CLAMPED BUT DID NOT STAPLE. STAPLER REMOVED AND BLEEDING NOTED. NEW STAPLER OBTAINED AND APPLIED TO RESOLVE ISSUE. FDA SAFETY REPORT ID #: (B)(4).
03-11-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: SUDDEN LOCK-OUT WHEN STARTED FIRING POWERED ECHELON (PSEE60A), WITH BLACK CARTRIDGE (GST60T). NO PATIENT CONSEQUENCE REPORTED.
03-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LUNG WEDGE RESECTION, WHILE FIRING ACROSS THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LUNG WEDGE RESECTION, WHILE FIRING ACROSS THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LIVER CANCER SURGERY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LIVER CANCER SURGERY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE PROCEDURE, THE RELOAD FAILED TO FIRE. THE GREEN BUTTON WAS ABLE TO BE PRESSED, BUT THE HANDLE COULD NOT BE SQUEEZED. THE SURGEON CHANGED THE RELOAD AND IT WORKED WITH THE SAME HANDLE PROPERLY. THERE WAS NO PATIENT INJURY.
03-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY SURGERY, UPON GASTRIC STAPLING, THE RELOAD STOPPED FIRING. IT WAS REPLACED BY A NEW RELOAD WITH THE SAME BATCH AND WORKED NORMALLY WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTROINTESTINAL PROCEDURE THE SURGEON FIRED THE STAPLER. AFTER FIRING THE SURGEON NOTICED THAT THE STAPLE LINE WAS "MESSED UP". UPON OBSERVATION, THE SURGEON FOUND THE PATIENT HAD A PREVIOUS SURGICAL PROCEDURE WHERE CLIPS WERE USED AND HAD FIRED OVER THOSE CLIPS. A SECOND STAPLER WAS TRIED AND IT WAS USED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCES. HOWEVER, THE SURGEON NOTICED THAT PART OF THE PLASTIC MANIFOLD WAS CHIPPED OFF.
07-10-2020 | Malfunction | covidien lp | 12-01-2021 | GIA STAPLER WITH DST SERIES TECHNOLOGY | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ATTEMPTING TO USE THE STAPLER TO TRANSECT THE PATIENT'S COLON. THERE WERE NO ISSUES WITH THE SCRUB TECH LOADING THE STAPLER OR GETTING THE TWO SIDES TO JOIN ACROSS THE COLON. HOWEVER, DESPITE EFFORTS AND MULTIPLE ATTEMPTS INCLUDING REMOVAL OF THE STAPLER AND REPLACEMENT/REJOINING, NONE OF US WERE ABLE TO CLOSE DOWN THE STAPLER TO FIRE IT. IN OTHER WORDS, THE ANVIL SIDE AND THE STAPLER LOAD SIDE WERE EASILY INTERLOCKED BUT WOULD NOT CLOSE DOWN COMPLETELY OR CLICK IN PLACE SO WE COULD FIRE THE LOAD ACROSS THE COLON. THE STAPLER FROM THE STERILE FIELD, OPENED ANOTHER, AND HAD NO PROBLEM COMPLETING THE PROCEDURE WITH THE REPLACEMENT.
07-10-2020 | Malfunction | covidien lp | 12-01-2021 | GIA STAPLER WITH DST SERIES TECHNOLOGY | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ATTEMPTING TO USE THE STAPLER TO TRANSECT THE PATIENT'S COLON. THERE WERE NO ISSUES WITH THE SCRUB TECH LOADING THE STAPLER OR GETTING THE TWO SIDES TO JOIN ACROSS THE COLON. HOWEVER, DESPITE EFFORTS AND MULTIPLE ATTEMPTS INCLUDING REMOVAL OF THE STAPLER AND REPLACEMENT/REJOINING, NONE OF US WERE ABLE TO CLOSE DOWN THE STAPLER TO FIRE IT. IN OTHER WORDS, THE ANVIL SIDE AND THE STAPLER LOAD SIDE WERE EASILY INTERLOCKED BUT WOULD NOT CLOSE DOWN COMPLETELY OR CLICK IN PLACE SO WE COULD FIRE THE LOAD ACROSS THE COLON. THE STAPLER FROM THE STERILE FIELD, OPENED ANOTHER, AND HAD NO PROBLEM COMPLETING THE PROCEDURE WITH THE REPLACEMENT.
30-10-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN CHARGING, LIQUID LEAKED FROM THE HANDLE AND ADHERED TO THE TERMINAL INSIDE THE CHARGER. THE DEVICE WOULD NOT CHARGE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
30-10-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN CHARGING, LIQUID LEAKED FROM THE HANDLE AND ADHERED TO THE TERMINAL INSIDE THE CHARGER. THE DEVICE WOULD NOT CHARGE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
30-10-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN CHARGING, LIQUID LEAKED FROM THE HANDLE AND ADHERED TO THE TERMINAL INSIDE THE CHARGER. THE DEVICE WOULD NOT CHARGE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
30-10-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN CHARGING, LIQUID LEAKED FROM THE HANDLE AND ADHERED TO THE TERMINAL INSIDE THE CHARGER. THE DEVICE WOULD NOT CHARGE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
14-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY, THE DEVICE CUT BUT DID NOT STAPLE NO STAPLE LINE WAS DEPLOYED. AN ADDITIONAL DEVICE WAS OPENED AND THE OP ENDED WITH NO INFLUENCE ON SURGERY AND PAT. SAFETY.
01-01-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE MALFORMED STAPLES (STAPLING LINE IS NOT STRAIGHT) ON THE THIRD FIRING OF THE BLUE RELOAD WITH CUTTING OF THE STOMACH WHILE THE JAWS ARE CLOSED THERE WAS NO ADVERSE EVENT REPORTED, DR USED OTHER NEW POWERED ECHELON TO COMPLETE THE SURGERY.
01-01-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | WHITE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BILATERAL REDUCTION MASTOPLASTY PROCEDURE, THE STAPLES REMAINED ATTACHED TO THE STAPLER, THEY DO NOT COME OUT CORRECTLY AND THEREFORE THE DEVICE WAS UNUSABLE DURING THE PROCEDURE. DELAY IN SURGERY.
23-11-2020 | Malfunction | celestica de monterrey s.a de c.v | 12-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLES COULD NOT BE FIRED NORMALLY WHEN THE INTESTINE WAS CLOSED. THE SURGEON STOPPED USING IT IMMEDIATELY AND REPLACED THE CUTTING STAPLER AND DISPOSABLE RELOAD FOR THE ENDOSCOPE TO CONTINUE TREATMENT. THERE WAS NO PATIENT INJURY.
23-11-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLES COULD NOT BE FIRED NORMALLY WHEN THE INTESTINE WAS CLOSED. THE SURGEON STOPPED USING IT IMMEDIATELY AND REPLACED THE CUTTING STAPLER AND DISPOSABLE RELOAD FOR THE ENDOSCOPE TO CONTINUE TREATMENT. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY WHEN SEVERING SUPERIOR PULMONARY VEIN TISSUE, AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO A NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2020 | Injury | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL GASTRIC CANCER PROCEDURE, THE DEVICE FIRED, BUT THERE WAS POOR STAPLE FORMATION THAT RESULTED TO INCOMPLETE STAPLE LINE CENTRALLY. TO RESOLVE THE ISSUE, THE SURGEON HAD TO SUTURE THE TISSUE AND COMPLETE THE CASE WITH ANOTHER RELOAD.
15-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE STOMACH DURING LAPAROSCOPIC GASTRIC SLEEVE, WHEN THE SURGEON WAS TRYING TO SQUEEZE THE HANDLE TO FIRE THE SECOND RELOAD, IT JAMMED AND COULD NOT FIRE. THE SURGEON OPENED ANOTHER HANDLE AND HAD THE SAME ISSUE. A COMPETITOR¿S PRODUCT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE STOMACH DURING LAPAROSCOPIC GASTRIC SLEEVE, WHEN THE SURGEON WAS TRYING TO SQUEEZE THE HANDLE TO FIRE THE SECOND RELOAD, IT JAMMED AND COULD NOT FIRE. THE SURGEON OPENED ANOTHER HANDLE AND HAD THE SAME ISSUE. A COMPETITOR¿S PRODUCT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE STOMACH DURING LAPAROSCOPIC GASTRIC SLEEVE, WHEN THE SURGEON WAS TRYING TO SQUEEZE THE HANDLE TO FIRE THE SECOND RELOAD, IT JAMMED AND COULD NOT FIRE. THE SURGEON OPENED ANOTHER HANDLE AND HAD THE SAME ISSUE. A COMPETITOR¿S PRODUCT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN PNEUMONECTOMY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO THE HOME POSITION AT THE 2ND FIRING ON THE LUNG PARENCHYMA. THE TARGET TISSUE WAS A LITTLE THICK, SO IT WAS COMPRESSED FULLY FOR 3 MINUTES BEFORE THE PULSE FIRING. THEN THE KNIFE REVERSE SWITCH WAS USED, BUT A SOUND OF THE BATTERY RUNNING OUT WAS HEARD, AND THE KNIFE DID NOT RETURN. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN PNEUMONECTOMY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO THE HOME POSITION AT THE 2ND FIRING ON THE LUNG PARENCHYMA. THE TARGET TISSUE WAS A LITTLE THICK, SO IT WAS COMPRESSED FULLY FOR 3 MINUTES BEFORE THE PULSE FIRING. THEN THE KNIFE REVERSE SWITCH WAS USED, BUT A SOUND OF THE BATTERY RUNNING OUT WAS HEARD, AND THE KNIFE DID NOT RETURN. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN PNEUMONECTOMY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO THE HOME POSITION AT THE 2ND FIRING ON THE LUNG PARENCHYMA. THE TARGET TISSUE WAS A LITTLE THICK, SO IT WAS COMPRESSED FULLY FOR 3 MINUTES BEFORE THE PULSE FIRING. THEN THE KNIFE REVERSE SWITCH WAS USED, BUT A SOUND OF THE BATTERY RUNNING OUT WAS HEARD, AND THE KNIFE DID NOT RETURN. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
25-04-2020 | Injury | covidien lp llc north haven | 12-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE LITERATURE SOURCE, A STUDY EVALUATED OUTCOME IN PATIENTS WHO UNDERWENT THORACOSCOPIC BILATERAL LUNG WEDGE RESECTIONS BETWEEN 2010 AND 2019. PATIENTS WERE DIVIDED INTO THE SUBXIPHOID APPROACH GROUP, THE ONE-STAGE LATERAL INTERCOSTAL APPROACH GROUP AND THE TWO-STAGE LATERAL INTERCOSTAL APPROACH GROUP. ENDO GIA STAPLER WAS USED FOR THE LUNG RESECTION. POST-OPERATIVE COMPLICATIONS INCLUDED ONE PATIENT WITH A PULMONARY AIR LEAK WHO REQUIRED SURGICAL CLOSURE OF THE PULMONARY LEAK.
13-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, IT WAS NOTED THAT THE STAPLES DID NOT DEPLOY PROPERLY. THE SURGEON, RESIDENT, AND SCRUB TECH NOTED THAT ONE SIDE OF THE STOMACH SEEMED TO CREASE DURING THE STAPLER FIRING AND PULL PROXIMALLY. THE SURGEON STATED IT FELT LIKE SOMETHING WAS DRAGGING WHILE FIRING. NO ABNORMAL SOUNDS WERE NOTED, THE STAPLER MOTOR SOUNDED THE SAME, AND IT DIDN¿T FEEL LIKE THE DEVICE WAS FIRED ACROSS ANYTHING HARD. ON THE REMNANT STOMACH, THE SURGEON OVER SEWED THE STAPLE LINE WITH VICRYL. ON THE SPECIMEN SIDE, AN OPENING WAS VISIBLE AND THE STAPLE LINE WAS INCOMPLETE. THE TECH AND SURGEON NOTED THEY WERE ABLE TO VISUALIZE A SMALL PIECE OF ORANGE MATERIAL LEFT BEHIND ON THE SPECIMEN. IT APPEARED TO BE PART OF THE STAPLE DRIVER. IT WAS REMOVED WITH GRASPERS. STAPLE LINE WAS OVERSEWED WITH VICRYL. REP WAS CALLED INTO THE OR TO DISCUSS AND HELP TROUBLESHOOT. ANOTHER STAPLER WAS OPENED AND UTILIZED ALTHOUGH THE ORIGINAL STAPLER SEEMED TO BE FUNCTIONING PROPERLY. PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A DELAY OF FOUR MINUTES. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
13-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, IT WAS NOTED THAT THE STAPLES DID NOT DEPLOY PROPERLY. THE SURGEON, RESIDENT, AND SCRUB TECH NOTED THAT ONE SIDE OF THE STOMACH SEEMED TO CREASE DURING THE STAPLER FIRING AND PULL PROXIMALLY. THE SURGEON STATED IT FELT LIKE SOMETHING WAS DRAGGING WHILE FIRING. NO ABNORMAL SOUNDS WERE NOTED, THE STAPLER MOTOR SOUNDED THE SAME, AND IT DIDN¿T FEEL LIKE THE DEVICE WAS FIRED ACROSS ANYTHING HARD. ON THE REMNANT STOMACH, THE SURGEON OVER SEWED THE STAPLE LINE WITH VICRYL. ON THE SPECIMEN SIDE, AN OPENING WAS VISIBLE AND THE STAPLE LINE WAS INCOMPLETE. THE TECH AND SURGEON NOTED THEY WERE ABLE TO VISUALIZE A SMALL PIECE OF ORANGE MATERIAL LEFT BEHIND ON THE SPECIMEN. IT APPEARED TO BE PART OF THE STAPLE DRIVER. IT WAS REMOVED WITH GRASPERS. STAPLE LINE WAS OVERSEWED WITH VICRYL. REP WAS CALLED INTO THE OR TO DISCUSS AND HELP TROUBLESHOOT. ANOTHER STAPLER WAS OPENED AND UTILIZED ALTHOUGH THE ORIGINAL STAPLER SEEMED TO BE FUNCTIONING PROPERLY. PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A DELAY OF FOUR MINUTES. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
13-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, IT WAS NOTED THAT THE STAPLES DID NOT DEPLOY PROPERLY. THE SURGEON, RESIDENT, AND SCRUB TECH NOTED THAT ONE SIDE OF THE STOMACH SEEMED TO CREASE DURING THE STAPLER FIRING AND PULL PROXIMALLY. THE SURGEON STATED IT FELT LIKE SOMETHING WAS DRAGGING WHILE FIRING. NO ABNORMAL SOUNDS WERE NOTED, THE STAPLER MOTOR SOUNDED THE SAME, AND IT DIDN¿T FEEL LIKE THE DEVICE WAS FIRED ACROSS ANYTHING HARD. ON THE REMNANT STOMACH, THE SURGEON OVER SEWED THE STAPLE LINE WITH VICRYL. ON THE SPECIMEN SIDE, AN OPENING WAS VISIBLE AND THE STAPLE LINE WAS INCOMPLETE. THE TECH AND SURGEON NOTED THEY WERE ABLE TO VISUALIZE A SMALL PIECE OF ORANGE MATERIAL LEFT BEHIND ON THE SPECIMEN. IT APPEARED TO BE PART OF THE STAPLE DRIVER. IT WAS REMOVED WITH GRASPERS. STAPLE LINE WAS OVERSEWED WITH VICRYL. REP WAS CALLED INTO THE OR TO DISCUSS AND HELP TROUBLESHOOT. ANOTHER STAPLER WAS OPENED AND UTILIZED ALTHOUGH THE ORIGINAL STAPLER SEEMED TO BE FUNCTIONING PROPERLY. PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A DELAY OF FOUR MINUTES. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
21-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | ETS ARTIC LNR CUTR NO KNIFE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PROXIMAL GASTRECTOMY, THE STAPLES WERE NOT DEPLOYED AT THE 1ST FIRING. THE DEVICE WAS REMOVED FROM THE PATIENT AND ATTEMPTED TO OPEN THE JAWS, BUT THE CLOSING LEVER AND THE FIRING TRIGGER COULD NOT RETURN TO THE HOME POSITION. WHEN THE DEVICE WAS CHECKED, IT WAS FOUND ABOUT 2 STAPLES AT THE PROXIMAL END OF THE CARTRIDGE WERE PROTRUDED BUT NOT DEPLOYED. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE DEVICE WAS USED BETWEEN THE ESOPHAGUS AND THE STOMACH. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | ETS ARTIC LNR CUTR NO KNIFE | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PROXIMAL GASTRECTOMY, THE STAPLES WERE NOT DEPLOYED AT THE 1ST FIRING. THE DEVICE WAS REMOVED FROM THE PATIENT AND ATTEMPTED TO OPEN THE JAWS, BUT THE CLOSING LEVER AND THE FIRING TRIGGER COULD NOT RETURN TO THE HOME POSITION. WHEN THE DEVICE WAS CHECKED, IT WAS FOUND ABOUT 2 STAPLES AT THE PROXIMAL END OF THE CARTRIDGE WERE PROTRUDED BUT NOT DEPLOYED. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE DEVICE WAS USED BETWEEN THE ESOPHAGUS AND THE STOMACH. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2020 | Injury | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, THE STAPLER FIRED AND COMPLETE THE FIRING CYCLE HOWEVER THERE POOR STAPLE FORMATION THAT RESULT TO INCOMPLETE STAPLE LINE DISTALLY. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE AND USED ANOTHER RELOAD IN ORDER TO COMPLETE THE CASE.
16-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC WEDGE SEGMENTAL RESECTION PROCEDURE, THE DEVICE FIRED, HOWEVER THERE WAS "BLUE AND YELLOW FRAGMENTS" BREAK OFF FROM THE RELOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | celestica de monterrey s.a de c.v | 12-01-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC WEDGE SEGMENTAL RESECTION PROCEDURE, THE DEVICE FIRED, HOWEVER THERE WAS "BLUE AND Y ELLOW FRAGMENTS" BREAK OFF FROM THE RELOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEVICE FIRING IN THE ANTRUM, DURING AN SLG PROCEDURE, IT WAS NOTED THAT THERE WAS POOR STAPLE FORMATION AND AN INCOMPLETE STAPLE LINE DISTALLY. SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEVICE FIRING IN THE ANTRUM, DURING AN SLG PROCEDURE, IT WAS NOTED THAT THERE WAS POOR STAPLE FORMATION AND AN INCOMPLETE STAPLE LINE DISTALLY. SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ROBOTIC SIGMOID PROCEDURE, WHILE FIRING THE HANDLE GAVE A LOUD CRACKING SOUND, THE FIRING STOPPED HALFWAY, THE SURGEON WAS NOT ABLE TO FIRE THE REST OF THE STAPLE, APPROXIMATELY 15MM OF THE RELOAD WERE JUST FIRED. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ROBOTIC SIGMOID PROCEDURE, WHILE FIRING THE HANDLE GAVE A LOUD CRACKING SOUND, THE FIRING STOPPED HALFWAY, THE SURGEON WAS NOT ABLE TO FIRE THE REST OF THE STAPLE, APPROXIMATELY 15MM OF THE RELOAD WERE JUST FIRED. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ROBOTIC SIGMOID PROCEDURE, WHILE FIRING THE HANDLE GAVE A LOUD CRACKING SOUND, THE FIRING STOPPED HALFWAY, THE SURGEON WAS NOT ABLE TO FIRE THE REST OF THE STAPLE, APPROXIMATELY 15 MM OF THE RELOAD WERE JUST FIRED. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY, THE DEVICE WAS ABLE TO BE FIRED ON THE FIRST FIRING WITHOUT ISSUE, HOWEVER ON THE SECOND FIRING THE DEVICE'S JAWS COULD NOT BE OPENED, WHEN THE SURGEON TRYING TO ADJUST THE POSITION. IN ADDITION THE SURGEON FIND IT HARD TO UNLOAD THE RELOAD SO IT WAS REMOVED BY FORCE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY, THE DEVICE WAS ABLE TO BE FIRED ON THE FIRST FIRING WITHOUT ISSUE, HOWEVER ON THE SECOND FIRING THE DEVICE'S JAWS COULD NOT BE OPENED, WHEN THE SURGEON TRYING TO ADJUST THE POSITION. IN ADDITION THE SURGEON FIND IT HARD TO UNLOAD THE RELOAD SO IT WAS REMOVED BY FORCE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-10-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREVENTIVE MAINTENANCE, THERE WAS NO REPORTED DEVICE ISSUE.
23-10-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREVENTIVE MAINTENANCE, THERE WAS NO REPORTED DEVICE ISSUE.
15-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY, THE DEVICE WAS ABLE TO BE FIRED ON THE FIRST FIRING WITHOUT ISSUE, HOWEVER ON THE SECOND FIRING THE DEVICE'S JAWS COULD NOT BE OPENED, WHEN THE SURGEON TRYING TO ADJUST THE POSITION. IN ADDITION THE SURGEON FIND IT HARD TO UNLOAD THE RELOAD SO IT WAS REMOVED BY FORCE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY, THE DEVICE WAS ABLE TO BE FIRED ON THE FIRST FIRING WITHOUT ISSUE, HOWEVER ON THE SECOND FIRING THE DEVICE'S JAWS COULD NOT BE OPENED, WHEN THE SURGEON TRYING TO ADJUST THE POSITION. IN ADDITION THE SURGEON FIND IT HARD TO UNLOAD THE RELOAD SO IT WAS REMOVED BY FORCE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-11-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED PLUS COMPACT 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC LOBECTOMY, THE KNIFE STOPPED MOVING FORWARD WITH A SOUND OF THE BATTERY RUNNING OUT AT THE 7TH FIRING ON THE LUNG PARENCHYMA. ONLY ABOUT 3 STAPLES AT THE PROXIMAL END OF THE CARTRIDGE WERE DEPLOYED. THE KNIFE REVERSE SWITCH WAS USED TO RETURN THE KNIFE. THE DEVICE WAS ATTEMPTED TO FIRE AGAIN BUT COULD NOT. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
12-11-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | POWERED PLUS COMPACT 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC LOBECTOMY, THE KNIFE STOPPED MOVING FORWARD WITH A SOUND OF THE BATTERY RUNNING OUT AT THE 7TH FIRING ON THE LUNG PARENCHYMA. ONLY ABOUT 3 STAPLES AT THE PROXIMAL END OF THE CARTRIDGE WERE DEPLOYED. THE KNIFE REVERSE SWITCH WAS USED TO RETURN THE KNIFE. THE DEVICE WAS ATTEMPTED TO FIRE AGAIN BUT COULD NOT. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
08-10-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE GREEN BUTTON WAS PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-10-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE GREEN BUTTON WAS PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE'S I-BEAM GOT STUCK AND STOPPED MOVING WITH 45AV. THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
15-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | UNKNOWN EGIA ADAPTER | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE'S I-BEAM GOT STUCK AND STOPPED MOVING WITH 45AVM. THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
12-01-2021 | Malfunction | covidien lp llc north haven | 12-01-2021 | EEA ORVIL | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DISTAL TIP DISCONNECTED FROM THE CIRCULAR PIECE WHEN IT WAS INADVERTENTLY PUT INTO THE ESOPHAGUS IN A DIFFERENT DIRECTION.
15-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | UNKNOWN STAPLING DEVICE | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE'S I-BEAM GOT STUCK AND STOPPED MOVING WITH 45AVM. THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
16-09-2020 | Injury | us surgical puerto rico | 12-01-2021 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A (B)(6) WOMAN WHO UNDERWENT A LAPAROSCOPIC LOW ANTERIOR RESECTION WITH LYMPH NODE DISSECTION DEVELOPED VAGINAL BLEEDING, SUFFERED DISCHARGE OF FECES AND GAS FROM VAGINA AND WAS READMITTED 14 DAYS POST-OPERATIVE THE PROCEDURE. PATIENT WAS DIAGNOSED WITH RVF (RECTOVAGINAL FISTULA) AT THE ANTERIOR RIGHT WALL OF THE ANASTOMOTIC RING AND COMPUTED TOMOGRAPHY (CT) SHOWED AIR IN THE VAGINA. ENEMA USING AMIDOTRIZOIC ACID CONFIRMED THE FORMATION OF RVF. CONSERVATIVE TREATMENTS DID NOT RESOLVE THE FISTULA AND RE-OPERATION WAS REQUIRED.
16-09-2020 | Injury | us surgical puerto rico | 12-01-2021 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A (B)(6) WOMAN WHO UNDERWENT A LAPAROSCOPIC LOW ANTERIOR RESECTION WITH LYMPH NODE DISSECTION DEVELOPED VAGINAL BLEEDING, SUFFERED DISCHARGE OF FECES AND GAS FROM VAGINA AND WAS READMITTED 14 DAYS POST-OPERATIVE THE PROCEDURE. PATIENT WAS DIAGNOSED WITH RVF (RECTOVAGINAL FISTULA) AT THE ANTERIOR RIGHT WALL OF THE ANASTOMOTIC RING AND COMPUTED TOMOGRAPHY (CT) SHOWED AIR IN THE VAGINA. ENEMA USING AMIDOTRIZOIC ACID CONFIRMED THE FORMATION OF RVF. CONSERVATIVE TREATMENTS DID NOT RESOLVE THE FISTULA AND RE-OPERATION WAS REQUIRED.
16-09-2020 | Injury | us surgical puerto rico | 12-01-2021 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A (B)(6) WOMAN WHO UNDERWENT A LAPAROSCOPIC LOW ANTERIOR RESECTION WITH LYMPH NODE DISSECTION DEVELOPED VAGINAL BLEEDING, SUFFERED DISCHARGE OF FECES AND GAS FROM VAGINA AND WAS READMITTED 14 DAYS POST-OPERATIVE THE PROCEDURE. PATIENT WAS DIAGNOSED WITH RVF (RECTOVAGINAL FISTULA) AT THE ANTERIOR RIGHT WALL OF THE ANASTOMOTIC RING AND COMPUTED TOMOGRAPHY (CT) SHOWED AIR IN THE VAGINA. ENEMA USING AMIDOTRIZOIC ACID CONFIRMED THE FORMATION OF RVF. CONSERVATIVE TREATMENTS DID NOT RESOLVE THE FISTULA AND RE-OPERATION WAS REQUIRED.
11-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS), WHILE BEING USED IN THE LUNG, THE HANDLE FIRED THREE TIMES AND THEN SUDDENLY POWERED OFF. THE HANDLE WOULD NOT TURN ON. A MANUAL STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE STAPLER STOPPED AND DID NOT FIRE. THE DEVICE HAD TO BE REPLACED. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS TRIED BUT THE DEVICE STILL DID NOT FIRE. THE RELOAD WAS DAMAGED AND DID NOT WORK ON A NEW HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS TRIED BUT THE DEVICE STILL DID NOT FIRE. THE RELOAD WAS DAMAGED AND DID NOT WORK ON A NEW HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMA PROCEDURE, THE FIRST SET OF THE WAS ABLE TO BE ASSEMBLED WITHOUT ANY PROBLEM, HOWEVER AFTER THE SURGEON PRESSED THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZE AND THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED HOWEVER THE SAME ISSUE OCCURRED DEVICE STILL COULD NOT BE FIRE. IN ADDITION THREE MORE RELOAD HAD A PROBLEM DURING LOADING, WHICH IT COULD NOT BE ATTACHED TO THE HANDLE. THE SURGEON THEN DECIDED TO USED ANOTHER DEVICE FROM OTHER MANUFACTURER TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMA PROCEDURE, THE FIRST SET OF THE WAS ABLE TO BE ASSEMBLED WITHOUT ANY PROBLEM, HOWEVER AFTER THE SURGEON PRESSED THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZE AND THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED HOWEVER THE SAME ISSUE OCCURRED DEVICE STILL COULD NOT BE FIRE. IN ADDITION THREE MORE RELOAD HAD A PROBLEM DURING LOADING, WHICH IT COULD NOT BE ATTACHED TO THE HANDLE. THE SURGEON THEN DECIDED TO USED ANOTHER DEVICE FROM OTHER MANUFACTURER TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-11-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY OF THE LIVER THE STAPLER WAS FIRED ONCE AND STOPPED WORKING. THE SURGEON "THINKS IS WAS DUE TO THE STAPLER ITSELF AND NOT THE BATTERY". THE STAPLER FIRED PROPERLY FIRST TIME. WHEN SECOND ATTEMPT TO FIRE WAS MADE, NOTHING HAPPENED, NO STAPLES, NO CUT. SURGEON OPENED JAWS, REMOVED DEVICE AND PASSED TO TECH.
14-10-2020 | Injury | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMICOLECTOMY, WHEN TRANSECTING THE DESCENDING COLON, IT WAS NOT POSSIBLE SINCE THE HANDLE COULD NOT BE SQUEEZED AFTER PRESSING THE GREEN BUTTON ALTHOUGH THE LOADING UNIT WAS LOADED PROPERLY. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE AND PRESS THE GREEN BUTTON BUT THERE WAS A CRACKING NOISE WHEN SQUEEZING THE HANDLE. AFTER THIS OCCURRED, THE HANDLE WAS NOT ANY LONGER UNDER TENSION TO PULL IT. BECAUSE OF TRYING SEVERAL TIMES, THE COLON WERE SOFTLY DAMAGED AND HAD TO BE TRANSECTED DEEPER WITH A NEW HANDLE AND LOADING UNIT. THE STAPLE ROW WAS VERY WEAK AND THE STAPLE ROW OPENED AT ONE POINT. SINCE THE ANASTOMOSIS WAS LEAKING, IT WAS RESECTED AND REAPPLIED. THE INCIDENT PROLONGED THE ANESTHESIA TIME AND SURGERY TIME FOR 30 MINUTES OR MORE AND THE TISSUE WAS SEVERELY TRAUMATIZED. ON THE SAME DAY THE PATIENT BLED HEAVILY FROM THE ANASTOMOSIS. CONSERVATIVELY, THIS COMPLICATION WAS CONTROLLED.
14-10-2020 | Injury | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMICOLECTOMY, WHEN TRANSECTING THE DESCENDING COLON, IT WAS NOT POSSIBLE SINCE THE HANDLE COULD NOT BE SQUEEZED AFTER PRESSING THE GREEN BUTTON ALTHOUGH THE LOADING UNIT WAS LOADED PROPERLY. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE AND PRESS THE GREEN BUTTON BUT THERE WAS A CRACKING NOISE WHEN SQUEEZING THE HANDLE. AFTER THIS OCCURRED, THE HANDLE WAS NOT ANY LONGER UNDER TENSION TO PULL IT. BECAUSE OF TRYING SEVERAL TIMES, THE COLON WERE SOFTLY DAMAGED AND HAD TO BE TRANSECTED DEEPER WITH A NEW HANDLE AND LOADING UNIT. THE STAPLE ROW WAS VERY WEAK AND THE STAPLE ROW OPENED AT ONE POINT. SINCE THE ANASTOMOSIS WAS LEAKING, IT WAS RESECTED AND REAPPLIED. THE INCIDENT PROLONGED THE ANESTHESIA TIME AND SURGERY TIME FOR 30 MINUTES OR MORE AND THE TISSUE WAS SEVERELY TRAUMATIZED. ON THE SAME DAY THE PATIENT BLED HEAVILY FROM THE ANASTOMOSIS. CONSERVATIVELY, THIS COMPLICATION WAS CONTROLLED.
14-10-2020 | Injury | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMICOLECTOMY, WHEN TRANSECTING THE DESCENDING COLON, IT WAS NOT POSSIBLE SINCE THE HANDLE COULD NOT BE SQUEEZED AFTER PRESSING THE GREEN BUTTON ALTHOUGH THE LOADING UNIT WAS LOADED PROPERLY. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE AND PRESS THE GREEN BUTTON BUT THERE WAS A CRACKING NOISE WHEN SQUEEZING THE HANDLE. AFTER THIS OCCURRED, THE HANDLE WAS NOT ANY LONGER UNDER TENSION TO PULL IT. BECAUSE OF TRYING SEVERAL TIMES, THE COLON WERE SOFTLY DAMAGED AND HAD TO BE TRANSECTED DEEPER WITH A NEW HANDLE AND LOADING UNIT. THE STAPLE ROW WAS VERY WEAK AND THE STAPLE ROW OPENED AT ONE POINT. SINCE THE ANASTOMOSIS WAS LEAKING, IT WAS RESECTED AND REAPPLIED. THE INCIDENT PROLONGED THE ANESTHESIA TIME AND SURGERY TIME FOR 30 MINUTES OR MORE AND THE TISSUE WAS SEVERELY TRAUMATIZED. ON THE SAME DAY THE PATIENT BLED HEAVILY FROM THE ANASTOMOSIS. CONSERVATIVELY, THIS COMPLICATION WAS CONTROLLED.
16-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMA PROCEDURE, THE FIRST SET OF THE WAS ABLE TO BE ASSEMBLED WITHOUT ANY PROBLEM, HOWEVER AFTER THE SURGEON PRESSED THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZE AND THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED HOWEVER THE SAME ISSUE OCCURRED DEVICE STILL COULD NOT BE FIRE. IN ADDITION THREE MORE RELOAD HAD A PROBLEM DURING LOADING, WHICH IT COULD NOT BE ATTACHED TO THE HANDLE. THE SURGEON THEN DECIDED TO USED ANOTHER DEVICE FROM OTHER MANUFACTURER TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMA PROCEDURE, THE FIRST SET OF THE WAS ABLE TO BE ASSEMBLED WITHOUT ANY PROBLEM, HOWEVER AFTER THE SURGEON PRESSED THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZE AND THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED HOWEVER THE SAME ISSUE OCCURRED DEVICE STILL COULD NOT BE FIRE. IN ADDITION THREE MORE RELOAD HAD A PROBLEM DURING LOADING, WHICH IT COULD NOT BE ATTACHED TO THE HANDLE. THE SURGEON THEN DECIDED TO USED ANOTHER DEVICE FROM OTHER MANUFACTURER TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE JAW DOES NOT SPLIT. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMA PROCEDURE, THE FIRST SET OF THE WAS ABLE TO BE ASSEMBLED WITHOUT ANY PROBLEM, HOWEVER AFTER THE SURGEON PRESSED THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZE AND THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED HOWEVER THE SAME ISSUE OCCURRED DEVICE STILL COULD NOT BE FIRE. IN ADDITION THREE MORE RELOAD HAD A PROBLEM DURING LOADING, WHICH IT COULD NOT BE ATTACHED TO THE HANDLE. THE SURGEON THEN DECIDED TO USED ANOTHER DEVICE FROM OTHER MANUFACTURER TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE JAW DOES NOT SPLIT. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE PORTAL WAS DETACHED DURING AN OPEN HEMI-HEPATECTOMY, THE STAPLES WERE NOT WELL FORMED AND COULD NOT BE COMPLETELY CLOSED ON THE SECOND FIRING. THE STAPLE LINE WAS ALSO INCOMPLETE CENTRALLY. THE RELOAD WAS REPLACED WITH ANOTHER MANUFACTURER¿S RELOAD TO COMPLETE THE CUTTING AND CLOSING. THERE WAS NO PATIENT INJURY.
13-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE, THE CDH29A COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | PROXIMATE*PPH PROCEDURE SET | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHEN THE DOCTOR WENT TO FIRE, THE STAPLER MADE THE CUT, BUT NOT THE SUTURE. MANUAL ANASTOMOSIS WAS REQUIRED. THERE WAS NO DAMAGE TO THE PATIENT.
15-12-2020 | Injury | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC PANCREATICODUODENECTOMY, WHILE DETACHING THE DUODENUM, THE SURGEON WAS ABLE TO FULLY SQUEEZE THE HANDLE, BUT SOME STAPLES DID NOT FORM A B SHAPE. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THE STAPLE LINE WAS MANUALLY SUTURED TO FIX THE ISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
15-12-2020 | Injury | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC PANCREATICODUODENECTOMY, WHILE DETACHING THE DUODENUM, THE SURGEON WAS ABLE TO FULLY SQUEEZE THE HANDLE, BUT SOME STAPLES DID NOT FORM A B SHAPE. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THE STAPLE LINE WAS MANUALLY SUTURED TO FIX THE ISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
11-12-2020 | Death | covidien lp llc north haven | 12-01-2021 | EEA ORVIL | Material Twisted/Bent | A040609 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, WHEN INSERTING THE ORAL ANVIL, BEFORE THE HANDLE WAS CONNECTED, ONE OF THE THREE WINGS OF THE ANVIL WAS BENT AND DAMAGED THE TISSUE. THE SURGERY WAS CONVERTED TO OPEN PROCEDURE. A PIECE OF THE ESOPHAGUS WAS RESECTED, AND A NEW ORAL ANVIL WAS USED ON THE SAME HANDLE. THE SURGICAL TIME WAS EXTENDED FOR TWO HOURS. IT WAS ALSO NOTED THAT TWO DAYS AFTER THE SURGERY THERE WAS AN INSUFFICIENCY AT THE ESOPHAGUS ANASTOMOSIS. THE PATIENT HOSPITALIZATION IS EXTENDED.
11-12-2020 | Death | covidien lp llc north haven | 12-01-2021 | EEA ORVIL | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, WHEN INSERTING THE ORAL ANVIL, BEFORE THE HANDLE WAS CONNECTED, ONE OF THE THREE WINGS OF THE ANVIL WAS BENT AND DAMAGED THE TISSUE. THE SURGERY WAS CONVERTED TO OPEN PROCEDURE. A PIECE OF THE ESOPHAGUS WAS RESECTED, AND A NEW ORAL ANVIL WAS USED ON THE SAME HANDLE. THE SURGICAL TIME WAS EXTENDED FOR TWO HOURS. IT WAS ALSO NOTED THAT TWO DAYS AFTER THE SURGERY THERE WAS AN INSUFFICIENCY AT THE ESOPHAGUS ANASTOMOSIS. THE PATIENT HOSPITALIZATION IS EXTENDED.
05-11-2020 | Injury | ethicon endo-surgery | 12-01-2021 | PROXIMATE*CUTTER RELOAD UNIT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT CASE WAS ESOPHAGECTOMY THAT HAPPENED ON (B)(6) 2020, IN THE CASE TO CREATE THE STOMACH TUBE DOCTOR FIRED ONE SR75 AND ONE TCR10. FIRING HAPPENED AT FUNDUS PORTION OF THE STOMACH. AFTER 10 DAYS DUE TO PATIENT CONDITION THEY DONE ENDOSCOPY AND FOUND THAT ALL THE PINS OF STAPLER ARE NOT THERE AND HAS FALLEN TO LUNGS AND TRACHEA, THEY AGAIN DONE THE SURGERY TO OVERCOME THIS AND PATIENT NOW IN VENTILATOR SUPPORT.
14-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE LARGE INTESTINE DURING AN OPEN COLECTOMY, THE STAPLER STUCK AND COULD NOT BE FIRED. A NEW RELOAD WAS USED AND WORKED PROPERLY. THERE WAS NO PATIENT INJURY.
05-11-2020 | Injury | ethicon endo-surgery | 12-01-2021 | 75 MM NTLC SELECTABLE RELOAD | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT CASE WAS ESOPHAGECTOMY THAT HAPPENED ON (B)(6) 2020, IN THE CASE TO CREATE THE STOMACH TUBE DOCTOR FIRED ONE SR75 AND ONE TCR10. FIRING HAPPENED AT FUNDUS PORTION OF THE STOMACH. AFTER 10 DAYS DUE TO PATIENT CONDITION THEY DONE ENDOSCOPY AND FOUND THAT ALL THE PINS OF STAPLER ARE NOT THERE AND HAS FALLEN TO LUNGS AND TRACHEA, THEY AGAIN DONE THE SURGERY TO OVERCOME THIS AND PATIENT NOW IN VENTILATOR SUPPORT.
15-12-2020 | Malfunction | us surgical puerto rico | 12-01-2021 | GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE SURGEON SQUEEZED THE HANDLE TO FIRE, THE STAPLER GOT STUCK AND STOPPED MOVING.
11-12-2020 | Malfunction | covidien lp llc north haven | 12-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE STAPLER GOT STUCK AND DID NOT FIRE. WHEN THE SURGEON CHANGED THE RELOAD IT WORKED NORMALLY. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | ENDO ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT SOME OF THE STAPLES WERE NOT FIRED ON THE TISSUE. ONLY THE BLADE CUT THROUGH THE TARGETED TISSUE, WITH THE TISSUE OPENED WITH A HOLE. SHE HAD TO DO SUTURES TO CLOSE THE HOLE.
04-12-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ESOPHAGEAL OPERATION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-11-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | ILS 29MM, CURVED | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT WORK. THE PROCEDURE CONTINUED WITH A DIFFERENT INSTRUMENT OF THE SAME TYPE FROM ANOTHER DISTRIBUTOR / MANUFACTURER. 20 MIN DELAY. PROCEDURE: LEFT LAPAROSCOPIC HEMICOLECTOMY.
18-01-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING VATS ESOPHAGECTOMY WHEN USING CDH25A TO DO ANASTOMOSIS , STAPLES LINE WERE UNSTABLE AFTER FIRING. TO RESOLVE THE EMERGENCY, SURGEON TRIED TO DO HAND-SEWING TECHNIQUE IN ANASTOMOTIC AREA . THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT SO FAR.
03-11-2020 | Malfunction | ethicon endo-surgery | 12-01-2021 | PROXIMATE*CUTTER RELOAD UNIT | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE 4-5 OF THE STAPLES WAS ALREADY ON THE SURFACE OF THE CARTRIDGE, AND IT WAS EVEN BEFORE IT GOT LOADED ON THE STAPLER. THERE WERE NO PATIENT CONSEQUENCES.
15-12-2020 | Malfunction | covidien | 12-01-2021 | EEA AUTO SUTURE CIRCULAR STAPLER WITH DST SERIES | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: EEA STAPLER MALFUNCTION. FDA SAFETY REPORT ID # (B)(4).
04-01-2021 | Injury | ethicon/ ethicon endo-surgery | 12-01-2021 | ETHICON EEA 29 STAPLER | Activation, Positioning or Separation Problem | A15 | Insufficient Information | Event Description: FROM SURGEON'S OP REPORT: WE THEN PASSED THE EEA 29 EEA STAPLER FROM THE ANAL CANAL AND DEPLOYED THE TINE OF THE STAPLER THROUGH THE MIDPORTION OF OUR CURVED CONTOUR STAPLE LINE. WE SECURED THE ANVIL TO THE TINE OF THE STAPLER AND CLOSED THE STAPLER ENSURING THERE WAS NO EXTRINSIC TISSUE INVOLVED WITH OUR ANASTOMOSIS. UNFORTUNATELY, THERE WAS A MALFUNCTION OF THE STAPLER AND IT DID NOT DEPLOY. THEREFORE, WE HAD TO EXCHANGE STAPLERS AT THIS POINT AND PERFORM THE SAME MANEUVER AS DICTATED EARLIER. THIS WENT SUCCESSFULLY AND WE FORM OUR COLOANAL ANASTOMOSIS WITH THE ETHICON EEA 29 MM STAPLER. WE CHECKED OUR COLOANAL ANASTOMOSIS FOR LEAK. THERE WAS NO EVIDENCE OF AIR LEAK USING A FLEXIBLE SIGMOIDOSCOPE. FDA SAFETY REPORT ID # (B)(4).
04-11-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, KNIFE DID NOT ADVANCE AFTER SEVERAL INITIATIONS.
12-12-2018 | Injury | covidien lp llc north haven | 13-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2018, THE PATIENT UNDERWENT LAP COLON RESECTION. THE PROCEDURE WAS CONVERTED TO OPEN LEFT PARTIAL COLECTOMY. THE PATIENT WAS TAKEN BACK TO THE OR FOR EXPLORATORY LAPAROTOMY AND NOTED TO HAVE A SMALL ANASTOMOTIC LEAK THAT WAS REPAIRED PRIMARILY ON (B)(6) 2018. THE PATIENT WAS AGAIN TAKEN FOR EXPLORATORY SURGERY TO DRAIN AN INTRABDOMINAL ABSCESS AND HAS MULTIPLE CULTURES GROWING PANSENSITIVE PSEUDOMONAS AERUGINOSA. THE PATIENT RETURN TO SURGERY FOR I<(>&<)>D OF ABDOMINAL WALL ABSCESS AND AGAIN FOR ABDOMINAL WASHOUT. THE PATIENT WAS DISCHARGED FROM THE HOSPITAL AT (B)(6) 2019.
12-12-2018 | Injury | covidien lp llc north haven | 13-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Rectal Anastomotic Leakage | Event Description: ACCORDING TO THE REPORTER, ON OR AROUND (B)(6) 2018, THE PATIENT UNDERWENT LAP COLON RESECTION. THE PROCEDURE WAS CONVERTED TO OPEN LEFT PARTIAL COLECTOMY. THE PATIENT WAS TAKEN BACK TO THE OR FOR EXPLORATORY LAPAROTOMY AND NOTED TO HAVE A SMALL ANASTOMOTIC LEAK THAT WAS REPAIRED PRIMARILY ON (B)(6) 2018. THE PATIENT WAS AGAIN TAKEN FOR EXPLORATORY SURGERY TO DRAIN AN INTRABDOMINAL ABSCESS AND HAS MULTIPLE CULTURES GROWING PANSENSITIVE PSEUDOMONAS AERUGINOSA. THE PATIENT RETURN TO SURGERY FOR I<(>&<)>D OF ABDOMINAL WALL ABSCESS AND AGAIN FOR ABDOMINAL WASHOUT. THE PATIENT WAS DISCHARGED FROM THE HOSPITAL AT (B)(6) 2019.
26-03-2019 | Injury | covidien lp llc north haven | 13-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT LAPAROSCOPIC REPAIR OF HIATAL HERNIA AND LAPAROSCOPIC SLEEVE GASTRECTOMY ON OR AROUND (B)(6) 2019. THE PATIENT EXPERIENCED MEDICAL COMPLICATIONS POST OP. INCLUDING PAIN AND NAUSEA.
26-03-2019 | Injury | covidien lp llc north haven | 13-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT LAPAROSCOPIC REPAIR OF HIATAL HERNIA AND LAPAROSCOPIC SLEEVE GASTRECTOMY ON OR AROUND (B)(6) 2019. THE PATIENT EXPERIENCED MEDICAL COMPLICATIONS POST OP. INCLUDING PAIN AND NAUSEA.
26-03-2019 | Injury | covidien lp llc north haven | 13-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT LAPAROSCOPIC REPAIR OF HIATAL HERNIA AND LAPAROSCOPIC SLEEVE GASTRECTOMY ON OR AROUND (B)(6) 2019. THE PATIENT EXPERIENCED MEDICAL COMPLICATIONS POST OP. INCLUDING PAIN AND NAUSEA.
16-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED PLUS COMPACT 45 | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, A CARTRIDGE COULD NOT BE LOADED INTO THE JAW. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. WHEN THE DEVICE WAS CHECKED, IT WAS FOUND THAT THE CARTRIDGE PAN OF THE USED CARTRIDGE WAS LEFT ON THE JAW. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED PLUS COMPACT 45 | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, A CARTRIDGE COULD NOT BE LOADED INTO THE JAW. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. WHEN THE DEVICE WAS CHECKED, IT WAS FOUND THAT THE CARTRIDGE PAN OF THE USED CARTRIDGE WAS LEFT ON THE JAW. NO FURTHER INFORMATION IS AVAILABLE.
27-11-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | PWRD ECH FLEX 45MM 34CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RIGHT LOWER LOBECTOMY SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. THE DEVICE WAS LOCKED ON TISSUE WITH JAWS CLOSED. IT WAS UNKNOWN HOW TO OPEN THE JAW. ALL MEANS HAD BEEN TRIED EXCEPT MANUAL OVERRIDE LEVER. THE INSTRUMENT WAS REMOVED BY FORCE AS THE DEVICE NEVER OPENED. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
18-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE KNIFE MOVED ALL THE WAY BUT STOPPED MOVING BACK TO HOME POSITION AT THE FIRST FIRING. THE KNIFE REVERSE SWITCH WAS USED BUT THE KNIFE DID NOT RETURN. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE DEVICE WAS USED ON THE LUNG. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE KNIFE MOVED ALL THE WAY BUT STOPPED MOVING BACK TO HOME POSITION AT THE FIRST FIRING. THE KNIFE REVERSE SWITCH WAS USED BUT THE KNIFE DID NOT RETURN. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE DEVICE WAS USED ON THE LUNG. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE KNIFE MOVED ALL THE WAY BUT STOPPED MOVING BACK TO HOME POSITION AT THE FIRST FIRING. THE KNIFE REVERSE SWITCH WAS USED BUT THE KNIFE DID NOT RETURN. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE DEVICE WAS USED ON THE LUNG. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING LUNG LOBE TISSUE, AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO ANOTHER TWO RELOADS, COULD NOT FIRE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING LUNG LOBE TISSUE, AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO ANOTHER TWO RELOADS, COULD NOT FIRE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE OPERATION, THE OPERATING ROOM NURSE PICKED UP WHAT THEY EXPECTED TO BE A 60 MM POWER PLUS STAPLER, BUT DESPITE THE MARKING ON BOTH THE KIT AND THE PACKAGING OF THE STAPLER, THEY FOUND A 45 MM POWER PLUS STAPLER. THE NURSE CHANGED THE INSTRUMENT AND THE PROCEDURE COULD CONTINUE WITHOUT COMPLICATIONS FOR THE PATIENT. THERE WAS NO PATIENT CONSEQUENCE.
18-12-2020 | Malfunction | davol inc sub c.r bard inc -1213643 | 13-01-2021 | CAPSURE STRAIGHT | Failure to Form Staple | A050704 | No Consequences Or Impact To Patient | Event Description: AS REPORTED, DURING A LAPAROSCOPIC PROCEDURE ON (B)(6) 2020, WHILE DEPLOYING THE FIRST 10 FASTENERS OF THE BARD/DAVOL CAPSURE STRAIGHT FIXATION DEVICE WITH APPROPRIATE COUNTER-PRESSURE, NO FASTENERS WERE ABLE TO FIXATE INTO THE BARD/DAVOL COMPOSIX L/P MESH CORRECTLY. CRACKING SOUND WAS NOT HEARD; THE FASTENERS PENETRATED THE MESH AND TISSUES PARTLY, HENCE THE SURGEON FOUND IT DIFFICULT TO REMOVE THE LOOSE FASTENERS AS IT WAS DAMAGING THE MESH AND TISSUES. THERE WAS NO REPORTED PATIENT INJURY. THE SURGEON IS AN EXISTING USER OF THE CAPSURE DEVICE.
16-11-2020 | Injury | us surgical puerto rico | 13-01-2021 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER FUNCTIONAL END-TO-END ENTEROSTOMY, THE STAPLE LINE WAS NOT LEAK PROOF. THE LINEAR CUTTER WAS FITTED WITH A NEW STAPLE CARTRIDGE, WHICH WENT TO PLAN. THE DEVICE WAS THEN CLOSED TOGETHER, HANDED OVER, AND USED TO CLOSE THE ANASTOMOSIS AND REMOVE THE RESECTION MATERIAL. WITH THE NEW MAGAZINE, THE LINEAR CUTTER COULD NOT BE ACTIVATED, SO IT WAS REMOVED AND RELOADED. WITH THE DEVICE NOW RELOADED, THE CUTTING/SUTURE PROCESS PROCEEDED WITHOUT INCIDENT. POST OPERATIVELY, 24 HOUR AFTER THE PROCEDURE, THE PATIENT HAD SEPSIS SYMPTOMS. AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE AND PATIENT HAD TO STAY IN HOSPITAL FOR 17 DAYS.
16-11-2020 | Injury | us surgical puerto rico | 13-01-2021 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE REPORTER, AFTER FUNCTIONAL END-TO-END ENTEROSTOMY, THE STAPLE LINE WAS NOT LEAK PROOF. THE LINEAR CUTTER WAS FITTED WITH A NEW STAPLE CARTRIDGE, WHICH WENT TO PLAN. THE DEVICE WAS THEN CLOSED TOGETHER, HANDED OVER, AND USED TO CLOSE THE ANASTOMOSIS AND REMOVE THE RESECTION MATERIAL. WITH THE NEW MAGAZINE, THE LINEAR CUTTER COULD NOT BE ACTIVATED, SO IT WAS REMOVED AND RELOADED. WITH THE DEVICE NOW RELOADED, THE CUTTING/SUTURE PROCESS PROCEEDED WITHOUT INCIDENT. POST OPERATIVELY, 24 HOUR AFTER THE PROCEDURE, THE PATIENT HAD SEPSIS SYMPTOMS. AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE AND PATIENT HAD TO STAY IN HOSPITAL FOR 17 DAYS.
16-11-2020 | Injury | us surgical puerto rico | 13-01-2021 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER FUNCTIONAL END-TO-END ENTEROSTOMY, THE STAPLE LINE WAS NOT LEAK PROOF. THE LINEAR CUTTER WAS FITTED WITH A NEW STAPLE CARTRIDGE, WHICH WENT TO PLAN. THE DEVICE WAS THEN CLOSED TOGETHER, HANDED OVER, AND USED TO CLOSE THE ANASTOMOSIS AND REMOVE THE RESECTION MATERIAL. WITH THE NEW MAGAZINE, THE LINEAR CUTTER COULD NOT BE ACTIVATED, SO IT WAS REMOVED AND RELOADED. WITH THE DEVICE NOW RELOADED, THE CUTTING/SUTURE PROCESS PROCEEDED WITHOUT INCIDENT. POST OPERATIVELY, 24 HOUR AFTER THE PROCEDURE, THE PATIENT HAD SEPSIS SYMPTOMS. AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE AND PATIENT HAD TO STAY IN HOSPITAL FOR 17 DAYS.
16-11-2020 | Injury | us surgical puerto rico | 13-01-2021 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE REPORTER, AFTER FUNCTIONAL END-TO-END ENTEROSTOMY, THE STAPLE LINE WAS NOT LEAK PROOF. THE LINEAR CUTTER WAS FITTED WITH A NEW STAPLE CARTRIDGE, WHICH WENT TO PLAN. THE DEVICE WAS THEN CLOSED TOGETHER, HANDED OVER, AND USED TO CLOSE THE ANASTOMOSIS AND REMOVE THE RESECTION MATERIAL. WITH THE NEW MAGAZINE, THE LINEAR CUTTER COULD NOT BE ACTIVATED, SO IT WAS REMOVED AND RELOADED. WITH THE DEVICE NOW RELOADED, THE CUTTING/SUTURE PROCESS PROCEEDED WITHOUT INCIDENT. POST OPERATIVELY, 24 HOUR AFTER THE PROCEDURE, THE PATIENT HAD SEPSIS SYMPTOMS. AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE AND PATIENT HAD TO STAY IN HOSPITAL FOR 17 DAYS.
13-12-2020 | Injury | us surgical puerto rico | 13-01-2021 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAROSCOPIC LOWER ANTERIOR RESECTION, WHEN ANASTOMOSIS OF SIGMOID COLON TO RECTUM WAS DONE, THE SURGEON REPORTED THAT 50% OF COLORECTAL ANASTOMOSIS HAD FALLEN APART WHEN THE PATIENT WAS BROUGHT BACK TO OPERATING ROOM. ALSO, A DOZEN OF THE STAPLES WERE MALFORMED CAUSING INCOMPLETE STAPLE LINE. OVERSEWING WAS DONE TO FIX THE STAPLE LINE. PATIENT ALSO EXPERIENCED POST OP LEAK AND WAS BROUGHT BACK TO THE HOSPITAL.
13-12-2020 | Injury | us surgical puerto rico | 13-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AFTER LAROSCOPIC LOWER ANTERIOR RESECTION, WHEN ANASTOMOSIS OF SIGMOID COLON TO RECTUM WAS DONE, THE SURGEON REPORTED THAT 50% OF COLORECTAL ANASTOMOSIS HAD FALLEN APART WHEN THE PATIENT WAS BROUGHT BACK TO OPERATING ROOM. ALSO, A DOZEN OF THE STAPLES WERE MALFORMED CAUSING INCOMPLETE STAPLE LINE. OVERSEWING WAS DONE TO FIX THE STAPLE LINE. PATIENT ALSO EXPERIENCED POST OP LEAK AND WAS BROUGHT BACK TO THE HOSPITAL.
17-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-11-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WAS NO LONGER CHARGING AFTER BEING CLEANED WITH THE WRONG SOLUTION. IT WAS STATED THAT THE INCORRECT CLEANING SOLUTION APPEARED TO HAVE LEFT A FILM ON THE CHARGING CONTACTS. TROUBLESHOOTING WAS DONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 - MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE EVIDENCE AVAILABLE THE REPORTED CONDITION THAT THE HANDLE WOULD NOT CHARGE AND THE CONTACTS WERE CONTAMINATED WAS CONFIRMED. THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE. IMPROVEMENTS HAVE BEEN INITIATED TO MITIGATE THIS CONDITION.
12-11-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WAS NO LONGER CHARGING AFTER BEING CLEANED WITH THE WRONG SOLUTION. IT WAS STATED THAT THE INCORRECT CLEANING SOLUTION APPEARED TO HAVE LEFT A FILM ON THE CHARGING CONTACTS. TROUBLESHOOTING WAS DONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 - MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE EVIDENCE AVAILABLE THE REPORTED CONDITION THAT THE HANDLE WOULD NOT CHARGE AND THE CONTACTS WERE CONTAMINATED WAS CONFIRMED. THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE. IMPROVEMENTS HAVE BEEN INITIATED TO MITIGATE THIS CONDITION.
16-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER LOADING THE STAPLES, THE DEVICE DID NOT FIRE. THE GREEN FIRING SAFETY WAS GENTLY TOUCHED, AND THE BLADE WAS OUT BUT THE GRIP DID NOT MOVE. OPENED A NEW HANDLE AND RELOAD TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
23-11-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-11-2020 | Malfunction | covidien lp | 13-01-2021 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: INTRA-OP, STAPLING DEVICE (ENDO GIA) FOR LUNG TISSUE MALFUNCTIONED AND DID NOT SEAL TISSUE AS IT WAS SUPPOSED TO. DURING A THORACOSCOPIC LOBECTOMY, ON RELEASING THE STAPLER THERE WAS A LARGE AMOUNT OF BLEEDING. THE STAPLES APPARENTLY DID NOT HOLD. THE UTILITY INCISION WAS CONVERTED TO THORACOTOMY FOR RAPID CONTROL OF BLEEDING. THORASCOPIC SURGERY CONVERTED TO OPEN LOBECTOMY.
17-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE SURGEON WAS ABLE PRESS THE GREEN BUTTON, AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WAS REMOVED EVENTUALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE SURGEON WAS ABLE PRESS THE GREEN BUTTON, AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WAS REMOVED EVENTUALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE SURGEON WAS ABLE PRESS THE GREEN BUTTON, AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WAS REMOVED EVENTUALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE SURGEON WAS ABLE PRESS THE GREEN BUTTON, AND WAS ABLE TO SQUEEZE THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WAS REMOVED EVENTUALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-08-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC LOBECTOMY PROCEDURE, WHEN OPENING/CLOSING CHECK WAS PERFORMED AFTER CONNECTING THE RELOAD, A CRACKING SOUND WAS HEARD ON THE HANDLE. THE SURGEON THEN STOPPED USING THE DEVICE AND A NEW HANDLE WAS USED TO RESOLVE THE ISSUE WITH THE SAME RELOAD. THE DEVICE WAS NOT USED IN PATIENT. THERE WAS NO PATIENT INJURY.
18-08-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC LOBECTOMY PROCEDURE, WHEN OPENING/CLOSING CHECK WAS PERFORMED AFTER CONNECTING THE RELOAD, A CRACKING SOUND WAS HEARD ON THE HANDLE. THE SURGEON THEN STOPPED USING THE DEVICE AND A NEW HANDLE WAS USED TO RESOLVE THE ISSUE WITH THE SAME RELOAD. THE DEVICE WAS NOT USED IN PATIENT. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE; HOWEVER, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WERE REMOVED EVENTUALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE; HOWEVER, THE DEVICE DID NOT FIRE. IT WAS ALSO NOTED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WERE REMOVED EVENTUALLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-12-2020 | Injury | us surgical puerto rico | 13-01-2021 | GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL PROCEDURE, WHILE RESECTING THE COLON, THE 2 LOADING UNIT WERE ABLE TO FIRE BUT ONLY FEW STAPLE DEPLOYED AT THE BEGINNING, AND AFTER THAT ONLY THE KNIFE RAN AND THE STAPLES WERE NOT PLACED RESULTING AN INCOMPLETE STAPLE LINE. TO RESOLVE THE ISSUE THE SURGEON HAD TO RE-RESECT AND RE-STAPLED THE TISSUE AND USED ANOTHER DEVICE. DUE TO A PROBLEM DURING THE ANASTOMOSIS, THE INTESTINAL TRACT HAD TO BE REMOVED BY ABOUT 10 CM MORE THAN USUAL.
16-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | UNKNOWN ENDO STITCH SULU | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC TOTAL HYSTERECTOMY, WHEN THE NEEDLE WAS BEING PASSED THROUGH TISSUE, THE NEEDLE CAME OUT OF THE DEVICE AND FELL INTO THE CAVITY OF THE PATIENT. A NEEDLE DRIVER WAS USED TO GRAB THE NEEDLE AND PULL THROUGH THE TISSUE. THE SUTURE WAS CUT AND ANOTHER INSTRUMENT AND SUTURE WERE OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien medical products | 13-01-2021 | PREMIUM PLUS CEEA | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE GASTROJEJUNOSTOMY ANASTOMOSIS ON A LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY, THEY HAD DIFFICULTY INSERTING THE ANVIL INTO THE BODY. AFTER COMBINING, TROCAR WAS EXPENDED. ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, UPON DIVISION OF STOMACH, THERE WAS LOADING PROBLEM USING THE FIRST RELOAD. IT WOULD NOT OPEN ALL THE WAY. THE RELOAD WAS REPLACED BEFORE IT CAME INTO CONTACT WITH THE PATIENT. WHEN THE DOCTOR WENT TO FIRE, THERE WAS AN ERROR MESSAGE AND THE RELOAD WOULD NOT FIRE. A NEW RELOAD WAS USED AND IT FIRED WITHOUT ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, UPON DIVISION OF STOMACH, THERE WAS LOADING PROBLEM USING THE FIRST RELOAD. IT WOULD NOT OPEN ALL THE WAY. THE RELOAD WAS REPLACED BEFORE IT CAME INTO CONTACT WITH THE PATIENT. WHEN THE DOCTOR WENT TO FIRE, THERE WAS AN ERROR MESSAGE AND THE RELOAD WOULD NOT FIRE. A NEW RELOAD WAS USED AND IT FIRED WITHOUT ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | ILS 29MM, CURVED | Failure to Cut | A050702 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR COLON RESECTION FOR A RECTAL CANCER PATIENT. THE CIRCULAR STAPLER (CDH29 & CDH33) BOTH MISFIRED. NO AUDIBLE NOISE OR TACTILE FEEDBACK THAT THE ANASTOMOSIS WAS COMPLETE. NO REAL ESTATE LEFT IN PATIENT TO COMPLETE THE ANASTOMOSIS THEREFORE PATIENT WILL HAVE A PERMANENT COLOSTOMY. CASE WAS SIX HOURS IN LENGTH.
18-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | ILS 29MM, CURVED | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR COLON RESECTION FOR A RECTAL CANCER PATIENT. THE CIRCULAR STAPLER (CDH29 & CDH33) BOTH MISFIRED. NO AUDIBLE NOISE OR TACTILE FEEDBACK THAT THE ANASTOMOSIS WAS COMPLETE. NO REAL ESTATE LEFT IN PATIENT TO COMPLETE THE ANASTOMOSIS THEREFORE PATIENT WILL HAVE A PERMANENT COLOSTOMY. CASE WAS SIX HOURS IN LENGTH.
18-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | ILS 29MM, CURVED | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR COLON RESECTION FOR A RECTAL CANCER PATIENT. THE CIRCULAR STAPLER (CDH29 & CDH33) BOTH MISFIRED. NO AUDIBLE NOISE OR TACTILE FEEDBACK THAT THE ANASTOMOSIS WAS COMPLETE. NO REAL ESTATE LEFT IN PATIENT TO COMPLETE THE ANASTOMOSIS THEREFORE PATIENT WILL HAVE A PERMANENT COLOSTOMY. CASE WAS SIX HOURS IN LENGTH.
18-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | ILS 29MM, CURVED | Ejection Problem | A0503 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR COLON RESECTION FOR A RECTAL CANCER PATIENT. THE CIRCULAR STAPLER (CDH29 & CDH33) BOTH MISFIRED. NO AUDIBLE NOISE OR TACTILE FEEDBACK THAT THE ANASTOMOSIS WAS COMPLETE. NO REAL ESTATE LEFT IN PATIENT TO COMPLETE THE ANASTOMOSIS THEREFORE PATIENT WILL HAVE A PERMANENT COLOSTOMY. CASE WAS SIX HOURS IN LENGTH.
18-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | ILS 33MM, CURVED | Failure to Cut | A050702 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR COLON RESECTION FOR A RECTAL CANCER PATIENT. THE CIRCULAR STAPLER (CDH29 & CDH33) BOTH MISFIRED. NO AUDIBLE NOISE OR TACTILE FEEDBACK THAT THE ANASTOMOSIS WAS COMPLETE. NO REAL ESTATE LEFT IN PATIENT TO COMPLETE THE ANASTOMOSIS THEREFORE PATIENT WILL HAVE A PERMANENT COLOSTOMY. CASE WAS SIX HOURS IN LENGTH.
18-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | ILS 33MM, CURVED | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR COLON RESECTION FOR A RECTAL CANCER PATIENT. THE CIRCULAR STAPLER (CDH29 & CDH33) BOTH MISFIRED. NO AUDIBLE NOISE OR TACTILE FEEDBACK THAT THE ANASTOMOSIS WAS COMPLETE. NO REAL ESTATE LEFT IN PATIENT TO COMPLETE THE ANASTOMOSIS THEREFORE PATIENT WILL HAVE A PERMANENT COLOSTOMY. CASE WAS SIX HOURS IN LENGTH.
18-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | ILS 33MM, CURVED | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR COLON RESECTION FOR A RECTAL CANCER PATIENT. THE CIRCULAR STAPLER (CDH29 & CDH33) BOTH MISFIRED. NO AUDIBLE NOISE OR TACTILE FEEDBACK THAT THE ANASTOMOSIS WAS COMPLETE. NO REAL ESTATE LEFT IN PATIENT TO COMPLETE THE ANASTOMOSIS THEREFORE PATIENT WILL HAVE A PERMANENT COLOSTOMY. CASE WAS SIX HOURS IN LENGTH.
18-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | ILS 33MM, CURVED | Ejection Problem | A0503 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR COLON RESECTION FOR A RECTAL CANCER PATIENT. THE CIRCULAR STAPLER (CDH29 & CDH33) BOTH MISFIRED. NO AUDIBLE NOISE OR TACTILE FEEDBACK THAT THE ANASTOMOSIS WAS COMPLETE. NO REAL ESTATE LEFT IN PATIENT TO COMPLETE THE ANASTOMOSIS THEREFORE PATIENT WILL HAVE A PERMANENT COLOSTOMY. CASE WAS SIX HOURS IN LENGTH.
16-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | UNKNOWN EGIA LEGACY ROTIC RELOAD | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A COLECTOMY, THE HANDLE GOT STUCK DURING USE. THE FIRING STOPPED IN THE CENTER. THE HANDLE AND THE RELOAD WERE REPLACED. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED, AFTER STAPLING, THE SURGEON NOTICED THE OPEN ANASTOMOSIS WITHOUT THE STAPLING LINE AND SOME LOOSE STAPLES. IT WAS NECESSARY TO EXCHANGE THE DEVICE FOR ANOTHER ONE TO FINISH THE SURGERY. THE SURGERY WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
16-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A COLECTOMY, THE HANDLE GOT STUCK DURING USE. THE FIRING STOPPED IN THE CENTER. THE HANDLE AND THE RELOAD WERE REPLACED. THERE WAS NO PATIENT INJURY.
20-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN LSG THE SURGEON US THE ECHELON PLEE60A. IN THE LAST FIRING, THE CARTRIDGE - GST60B WASN'T FIRED PROPERLY AND THE STAPLE LINE WAS INVALID/ THERE WERE MISSING STAPLES. THE SURGEON STITCH THE LAST PART OF THE STAPLE LINE AND FINISH THE SURGERY. THE SURGERY WAS DELAYED FOR 5 MINUTES. THERE WERE NO PATIENT CONSEQUENCES.
14-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE OR SEGMENTAL RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. IN ADDITION, THE DEVICE LOCKED ON TISSUE AND WAS NORMALLY REMOVED BY RETURNING THE BLACK KNOB. BOTH HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE OR SEGMENTAL RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. IN ADDITION, THE DEVICE LOCKED ON TISSUE AND WAS NORMALLY REMOVED BY RETURNING THE BLACK KNOB. BOTH HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, WHEN THE SURGEON TRIED TO FIRE THE DEVICE THE SURGEON ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE FIRING COULD NOT BE PERFORMED. IN ADDITION A DURING THE FIRING ATTEMPT, AN ABNORMAL SOUND WAS HEARD, IT RESULTED TO THE DEVICE BEING LOCKED ON THE TISSUE AND WHEN TRYING TO REMOVED, THE HANDLE GEAR CHIPPED AND FELL INTO THE PATIENT CAVITY. THE COMPONENT WAS ABLE TO BE RETRIEVED AND THE SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE BUT SAME ISSUE OCCURRED THE DEVICE STILL COULD NOT BE FIRE SO THE HANDLE WAS ALSO REPLACED USED TOGETHER WITH THE SECOND RELOAD AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, WHEN THE SURGEON TRIED TO FIRE THE DEVICE THE SURGEON ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE FIRING COULD NOT BE PERFORMED. IN ADDITION A DURING THE FIRING ATTEMPT, AN ABNORMAL SOUND WAS HEARD, IT RESULTED TO THE DEVICE BEING LOCKED ON THE TISSUE AND WHEN TRYING TO REMOVED, THE HANDLE GEAR CHIPPED AND FELL INTO THE PATIENT CAVITY. THE COMPONENT WAS ABLE TO BE RETRIEVED AND THE SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE BUT SAME ISSUE OCCURRED THE DEVICE STILL COULD NOT BE FIRE SO THE HANDLE WAS ALSO REPLACED USED TOGETHER WITH THE SECOND RELOAD AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
19-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | PROXIMATE*PPH PROCEDURE SET | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT WHILE FIRING THE STAPLER THE WASHER CUTTING TACTILE FEEDBACK IS NOT THERE AND AFTER FIRED, THERE IS BLEEDING AT CLOCK 9. THE PROCEDURE WAS DELAYED BY 30-MINUTES. SUTURING AT THE BLEEDING END WHERE ANASTOMOSIS IS NOT PERFECTLY DONE. THE PROCEDURE WAS COMPLETED. PROCEDURE: HEMORRHOIDECTOMY
19-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | PROXIMATE*PPH PROCEDURE SET | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT WHILE FIRING THE STAPLER THE WASHER CUTTING TACTILE FEEDBACK IS NOT THERE AND AFTER FIRED, THERE IS BLEEDING AT CLOCK 9. THE PROCEDURE WAS DELAYED BY 30-MINUTES. SUTURING AT THE BLEEDING END WHERE ANASTOMOSIS IS NOT PERFECTLY DONE. THE PROCEDURE WAS COMPLETED. PROCEDURE: HEMORRHOIDECTOMY
19-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | PROXIMATE*PPH PROCEDURE SET | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT WHILE FIRING THE STAPLER THE WASHER CUTTING TACTILE FEEDBACK IS NOT THERE AND AFTER FIRED, THERE IS BLEEDING AT CLOCK 9. THE PROCEDURE WAS DELAYED BY 30-MINUTES. SUTURING AT THE BLEEDING END WHERE ANASTOMOSIS IS NOT PERFECTLY DONE. THE PROCEDURE WAS COMPLETED. PROCEDURE: HEMORRHOIDECTOMY
19-12-2020 | Injury | ethicon endo-surgery | 13-01-2021 | PROXIMATE*PPH PROCEDURE SET | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT WHILE FIRING THE STAPLER THE WASHER CUTTING TACTILE FEEDBACK IS NOT THERE AND AFTER FIRED, THERE IS BLEEDING AT CLOCK 9. THE PROCEDURE WAS DELAYED BY 30-MINUTES. SUTURING AT THE BLEEDING END WHERE ANASTOMOSIS IS NOT PERFECTLY DONE. THE PROCEDURE WAS COMPLETED. PROCEDURE: HEMORRHOIDECTOMY
16-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | UNKNOWN EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC WHIPPLE PROCEDURE, DURING DUODENUM RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. THE SAME ISSUE OCCURRED WITH ANOTHER RELOAD. A MANUAL STAPLER AND A NEW RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-11-2020 | Injury | ethicon endo-surgery | 13-01-2021 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT THE PROF. WAS USING THE CONTOUR DEVICE IN LAR - VERY THICK TISSUE. DEVICE JAMMED AND PIN WAS POSSIBLY BROKEN. THE SURGERY WAS DELAYED BY 15-MINUTES. A NEW DEVICE WAS USED.
13-11-2020 | Injury | ethicon endo-surgery | 13-01-2021 | CONTOUR CURVED CUTTER STAPLER | Material Integrity Problem | A04 | Insufficient Information | Event Description: IT WAS REPORTED THAT THE PROF. WAS USING THE CONTOUR DEVICE IN LAR - VERY THICK TISSUE. DEVICE JAMMED AND PIN WAS POSSIBLY BROKEN. THE SURGERY WAS DELAYED BY 15-MINUTES. A NEW DEVICE WAS USED.
20-02-2020 | Injury | us surgical puerto rico | 13-01-2021 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2018 AND 2019, WHICH EVALUATED THE EFFICACY OF INDOCYANINE GREEN (ICG) FLUORESCENCE ANGIOGRAPHY (FA) IN REDUCING THE INCIDENCE OF ANASTOMOTIC LEAKAGE FOLLOWING STAPLED COLORECTAL ANASTOMOSIS. 377 PATIENTS UNDERWENT SIGMOID AND RECTAL RESECTION AND WERE DIVIDED INTO TWO GROUPS, WITH AND WITHOUT ICG FA. A CIRCULAR STAPLER WAS USED USING A DOUBLE STAPLING TECHNIQUE IN ALL CASES. THE FOLLOWING POST OPERATIVE COMPLICATIONS WERE REPORTED: ANASTOMOTIC LEAKS IN 48 PATIENTS, REOPERATION IN 7 PATIENTS DUE TO ANASTOMOTIC LEAK (6 OF THOSE PATIENTS HAVING ANASTOMOTIC RESECTION AND CREATION OF END COLOSTOMY). THE AVERAGE POSTOPERATIVE LENGTH OF HOSPITAL STAY WAS 8 DAYS. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE WERE; WOUND INFECTION, PARALYTIC ILEUS, URINARY RETENTION, ABDOMINAL BLEEDING AND MECHANICAL ILEUS.
20-02-2020 | Injury | us surgical puerto rico | 13-01-2021 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2018 AND 2019, WHICH EVALUATED THE EFFICACY OF INDOCYANINE GREEN (ICG) FLUORESCENCE ANGIOGRAPHY (FA) IN REDUCING THE INCIDENCE OF ANASTOMOTIC LEAKAGE FOLLOWING STAPLED COLORECTAL ANASTOMOSIS. 377 PATIENTS UNDERWENT SIGMOID AND RECTAL RESECTION AND WERE DIVIDED INTO TWO GROUPS, WITH AND WITHOUT ICG FA. A CIRCULAR STAPLER WAS USED USING A DOUBLE STAPLING TECHNIQUE IN ALL CASES. THE FOLLOWING POST OPERATIVE COMPLICATIONS WERE REPORTED: ANASTOMOTIC LEAKS IN 48 PATIENTS, REOPERATION IN 7 PATIENTS DUE TO ANASTOMOTIC LEAK (6 OF THOSE PATIENTS HAVING ANASTOMOTIC RESECTION AND CREATION OF END COLOSTOMY). THE AVERAGE POSTOPERATIVE LENGTH OF HOSPITAL STAY WAS 8 DAYS. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE WERE; WOUND INFECTION, PARALYTIC ILEUS, URINARY RETENTION, ABDOMINAL BLEEDING AND MECHANICAL ILEUS.
17-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHEN DEVICE WAS BEING APPLIED ON THE LUNG LOBE TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER LOADING THE STAPLES, THE DEVICE DID NOT FIRE. THE GREEN FIRING SAFETY WAS GENTLY TOUCHED, AND THE BLADE WAS OUT BUT THE GRIP DID NOT MOVE. OPENED A NEW HANDLE AND RELOAD TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
13-08-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE GREEN INDICATOR OF THE CHARGER WAS OFF WITHOUT ILLUMINATING OR FLASHING EVEN THOUGH THE HANDLE WAS SET ON THE CHARGER. THE USER RESET THE HANDLE SEVERAL TIMES, BUT THERE WAS NO CHANGE. ALSO, THE CONTACT METAL PART BETWEEN THE HANDLE AND THE CHARGER APPEARED TO BE WET. NOTHING WAS DONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
13-08-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE GREEN INDICATOR OF THE CHARGER WAS OFF WITHOUT ILLUMINATING OR FLASHING EVEN THOUGH THE HANDLE WAS SET ON THE CHARGER. THE USER RESET THE HANDLE SEVERAL TIMES, BUT THERE WAS NO CHANGE. ALSO, THE CONTACT METAL PART BETWEEN THE HANDLE AND THE CHARGER APPEARED TO BE WET. NOTHING WAS DONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
13-08-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE GREEN INDICATOR OF THE CHARGER WAS OFF WITHOUT ILLUMINATING OR FLASHING EVEN THOUGH THE HANDLE WAS SET ON THE CHARGER. THE USER RESET THE HANDLE SEVERAL TIMES, BUT THERE WAS NO CHANGE. ALSO, THE CONTACT METAL PART BETWEEN THE HANDLE AND THE CHARGER APPEARED TO BE WET. NOTHING WAS DONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
05-11-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION PRIOR TO USE ON A PATIENT, THE HANDLE WAS REMOVED FROM THE CHARGER AND STARTED UP. HOWEVER, A HANDLE ERROR DISPLAY WAS SHOWN. THEY HAD TO RESTART THE DEVICE BUT DID NOT IMPROVE. A NEW DEVICE WAS USED. THERE WAS NO PATIENT INVOLVEMENT.
05-11-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION PRIOR TO USE ON A PATIENT, THE HANDLE WAS REMOVED FROM THE CHARGER AND STARTED UP. HOWEVER, A HANDLE ERROR DISPLAY WAS SHOWN. THEY HAD TO RESTART THE DEVICE BUT DID NOT IMPROVE. A NEW DEVICE WAS USED. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2020 | Malfunction | covidien lp llc north haven | 13-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING FIRING OF THE STOMACH ON A LAPAROSCOPIC GASTRIC SLEEVE GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND COULD SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE AT ALL. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE FIRING COULD NOT BE PERFORMED. IN ADDITION, DURING THE FIRING ATTEMPT, AN ABNORMAL SOUND WAS HEARD, IT RESULTED TO THE DEVICE BEING LOCKED ON THE TISSUE AND WHEN TRYING TO REMOVED, THE HANDLE GEAR CHIPPED AND FELL INTO THE PATIENT CAVITY. THE COMPONENT WAS ABLE TO BE RETRIEVED AND THE SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE BUT SAME ISSUE OCCURRED. THE DEVICE STILL COULD NOT BE FIRE SO THE HANDLE WAS ALSO REPLACED WITH THE SECOND RELOAD AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE FIRING COULD NOT BE PERFORMED. IN ADDITION, DURING THE FIRING ATTEMPT, AN ABNORMAL SOUND WAS HEARD, IT RESULTED TO THE DEVICE BEING LOCKED ON THE TISSUE AND WHEN TRYING TO REMOVED, THE HANDLE GEAR CHIPPED AND FELL INTO THE PATIENT CAVITY. THE COMPONENT WAS ABLE TO BE RETRIEVED AND THE SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE BUT SAME ISSUE OCCURRED. THE DEVICE STILL COULD NOT BE FIRE SO THE HANDLE WAS ALSO REPLACED WITH THE SECOND RELOAD AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE FIRING COULD NOT BE PERFORMED. IN ADDITION, DURING THE FIRING ATTEMPT, AN ABNORMAL SOUND WAS HEARD, IT RESULTED TO THE DEVICE BEING LOCKED ON THE TISSUE AND WHEN TRYING TO REMOVED, THE HANDLE GEAR CHIPPED AND FELL INTO THE PATIENT CAVITY. THE COMPONENT WAS ABLE TO BE RETRIEVED AND THE SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE BUT SAME ISSUE OCCURRED. THE DEVICE STILL COULD NOT BE FIRE SO THE HANDLE WAS ALSO REPLACED WITH THE SECOND RELOAD AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, WHEN THE SURGEON TRIED TO FIRE THE DEVICE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE FIRING COULD NOT BE PERFORMED. IN ADDITION, DURING THE FIRING ATTEMPT, AN ABNORMAL SOUND WAS HEARD, IT RESULTED TO THE DEVICE BEING LOCKED ON THE TISSUE AND WHEN TRYING TO REMOVED, THE HANDLE GEAR CHIPPED AND FELL INTO THE PATIENT CAVITY. THE COMPONENT WAS ABLE TO BE RETRIEVED AND THE SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE BUT SAME ISSUE OCCURRED. THE DEVICE STILL COULD NOT BE FIRE SO THE HANDLE WAS ALSO REPLACED WITH THE SECOND RELOAD AND THE FIRING WAS COMPLETED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, DURING STOMACH REMOVAL, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. ONLY THE HANDLE WAS REPLACED AND THE SURGEON COMPLETED THE CASE. THERE WAS NO PATIENT INJURY
14-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE OR SEGMENTAL RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. IN ADDITION, THE DEVICE LOCKED ON TISSUE AND WAS NORMALLY REMOVED BY RETURNING THE BLACK KNOB. BOTH HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE OR SEGMENTAL RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. IN ADDITION, THE DEVICE LOCKED ON TISSUE AND WAS NORMALLY REMOVED BY RETURNING THE BLACK KNOB. BOTH HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 13-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE WAS NOT ABLE TO FIRE. THE SURGEON USED ANOTHER HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-11-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER STOPPED WITH A DEFECT IN THE BULK CARRIER AND WAS IMMEDIATELY REPLACED BY A NEW ONE. PATIENT CONSEQUENCE WAS NOT REPORTED.
21-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | GST WHITE RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING VESSELS, AFTER FIRED WITH GST45W, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | GST WHITE RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING VESSELS, AFTER FIRED WITH GST45W, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | GST WHITE RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING VESSELS, AFTER FIRED WITH GST45W, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | ECHELON FLEX 60 COMPACT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RADICAL RESECTION OF LUNG CANCER SURGERY, WHEN SEVERING PULMONARY TISSUE ON THE FIRST FIRING, AFTER FIRED, NOTED SOME STAPLES WERE MALFORMED WITH IRREGULAR SHAPE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC LOBECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC LOBECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2020 | Malfunction | ethicon endo-surgery | 13-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC LOBECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2020 | Injury | us surgical puerto rico | 14-01-2021 | TA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN LOBECTOMY, WHILE STAPLING THE LOBE, THE DEVICE HAD POOR STAPLE LINE FORMATION. THE SURGEON OVER-SEWED THE STAPLE LINE.
17-12-2020 | Injury | covidien lp llc north haven | 14-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Unspecified Infection | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE STAPLE LINE FORMED CORRECTLY, BUT IT WAS NOT COMPLETE; THE STAPLE LINE DID NOT FORM DISTALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AS WELL. COMPLETED THE RESECTION WITH ANOTHER STAPLER FROM ANOTHER COMPANY AND REMOVED THE SURGICAL PIECE WITH THE OTHER STAPLER STILL BLOCKED WITH THE JAWS CLAMPING THE TISSUE. WHEN THE SURGICAL PIECE WAS REMOVED, THEY FORCED THE INSTRUMENT IN ORDER TO OPEN IT. IT WAS ALSO NECESSARY TO DO ADDITIONAL LUNG RESECTION IN ORDER TO BE ABLE TO REMOVE THE INSTRUMENT BLOCKED ON TISSUE. THERE WAS AN UNANTICIPATED TISSUE LOSS, AND THE SURGICAL TIME WAS EXTENDED FOR ABOUT AN HOUR AS A RESULT OF THIS PROBLEM. THE INITIAL INCISION WAS EXTENDED TO MORE THAN 2.54 CM AND ALSO A NEW INCISION WAS PERFORMED. ON (B)(6) 2021, ON POST SURGERY MEDICAL CHECK, THE INITIAL EXTENDED INCISION THROUGH WHERE OUR INSTRUMENT WAS INSERTED GOT INFECTED.
17-12-2020 | Injury | covidien lp llc north haven | 14-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE STAPLE LINE FORMED CORRECTLY, BUT IT WAS NOT COMPLETE; THE STAPLE LINE DID NOT FORM DISTALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AS WELL. COMPLETED THE RESECTION WITH ANOTHER STAPLER FROM ANOTHER COMPANY AND REMOVED THE SURGICAL PIECE WITH THE OTHER STAPLER STILL BLOCKED WITH THE JAWS CLAMPING THE TISSUE. WHEN THE SURGICAL PIECE WAS REMOVED, THEY FORCED THE INSTRUMENT IN ORDER TO OPEN IT. IT WAS ALSO NECESSARY TO DO ADDITIONAL LUNG RESECTION IN ORDER TO BE ABLE TO REMOVE THE INSTRUMENT BLOCKED ON TISSUE. THERE WAS AN UNANTICIPATED TISSUE LOSS, AND THE SURGICAL TIME WAS EXTENDED FOR ABOUT AN HOUR AS A RESULT OF THIS PROBLEM. THE INITIAL INCISION WAS EXTENDED TO MORE THAN 2.54 CM AND ALSO A NEW INCISION WAS PERFORMED. ON (B)(6) 2021, ON POST SURGERY MEDICAL CHECK, THE INITIAL EXTENDED INCISION THROUGH WHERE OUR INSTRUMENT WAS INSERTED GOT INFECTED.
17-12-2020 | Injury | covidien lp llc north haven | 14-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | Unspecified Infection | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE STAPLE LINE FORMED CORRECTLY, BUT IT WAS NOT COMPLETE; THE STAPLE LINE DID NOT FORM DISTALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AS WELL. COMPLETED THE RESECTION WITH ANOTHER STAPLER FROM ANOTHER COMPANY AND REMOVED THE SURGICAL PIECE WITH THE OTHER STAPLER STILL BLOCKED WITH THE JAWS CLAMPING THE TISSUE. WHEN THE SURGICAL PIECE WAS REMOVED, THEY FORCED THE INSTRUMENT IN ORDER TO OPEN IT. IT WAS ALSO NECESSARY TO DO ADDITIONAL LUNG RESECTION IN ORDER TO BE ABLE TO REMOVE THE INSTRUMENT BLOCKED ON TISSUE. THERE WAS AN UNANTICIPATED TISSUE LOSS, AND THE SURGICAL TIME WAS EXTENDED FOR ABOUT AN HOUR AS A RESULT OF THIS PROBLEM. THE INITIAL INCISION WAS EXTENDED TO MORE THAN 2.54 CM AND ALSO A NEW INCISION WAS PERFORMED. ON (B)(6) 2021, ON POST SURGERY MEDICAL CHECK, THE INITIAL EXTENDED INCISION THROUGH WHERE OUR INSTRUMENT WAS INSERTED GOT INFECTED.
17-12-2020 | Injury | covidien lp llc north haven | 14-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE STAPLE LINE FORMED CORRECTLY, BUT IT WAS NOT COMPLETE; THE STAPLE LINE DID NOT FORM DISTALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AS WELL. COMPLETED THE RESECTION WITH ANOTHER STAPLER FROM ANOTHER COMPANY AND REMOVED THE SURGICAL PIECE WITH THE OTHER STAPLER STILL BLOCKED WITH THE JAWS CLAMPING THE TISSUE. WHEN THE SURGICAL PIECE WAS REMOVED, THEY FORCED THE INSTRUMENT IN ORDER TO OPEN IT. IT WAS ALSO NECESSARY TO DO ADDITIONAL LUNG RESECTION IN ORDER TO BE ABLE TO REMOVE THE INSTRUMENT BLOCKED ON TISSUE. THERE WAS AN UNANTICIPATED TISSUE LOSS, AND THE SURGICAL TIME WAS EXTENDED FOR ABOUT AN HOUR AS A RESULT OF THIS PROBLEM. THE INITIAL INCISION WAS EXTENDED TO MORE THAN 2.54 CM AND ALSO A NEW INCISION WAS PERFORMED. ON (B)(6) 2021, ON POST SURGERY MEDICAL CHECK, THE INITIAL EXTENDED INCISION THROUGH WHERE OUR INSTRUMENT WAS INSERTED GOT INFECTED.
18-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPY, ON STAPLING THE LUNG TISSUE, THE STAPLER WAS USED SUCCESSFULLY ON THE TWO RELOADS. ON THE THIRD RELOAD, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. ANOTHER DEVICE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPY, ON STAPLING THE LUNG TISSUE, THE STAPLER WAS USED SUCCESSFULLY ON THE TWO RELOADS. ON THE THIRD RELOAD, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. ANOTHER DEVICE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-11-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, AND THE STAPLER DID NOT CLOSE THE RELOAD.
05-11-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, AND THE STAPLER DID NOT CLOSE THE RELOAD.
05-11-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, AND THE STAPLER DID NOT CLOSE THE RELOAD.
05-11-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, AND THE STAPLER DID NOT CLOSE THE RELOAD.
18-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | PROXIMATE | Failure to Fire | A050501 | Insufficient Information | Event Description: THE 75MM LINEAR CUTTER FAILED TO FIRE WHILE BEING USED DURING AN ABDOMINAL SURGERY. IT WAS USED THREE TIMES BEFORE AND WORKED PERFECTLY WITHOUT ANY ISSUES HOWEVER, WHEN THE SURGEON TRIED TO USED IT A FOURTH TIME IT WAS UNABLE TO DELIVERED THE LOADED STAPLE. A NEW LINEAR CUTTER WITH THE SAME LOT AND REFERENCE NUMBER WAS OBTAINED AND THIS WITH ONE WORKED WITHOUT MALFUNCTIONING.
16-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, WHEN CHECKING THE OPENING AND CLOSING INTRA-OPERATIVELY, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD, BUT THE JAWS WERE ONLY ABLE TO CLOSE HALFWAY SO THE OPENING AND CLOSING CHECK COULD NOT BE PERFORMED.
16-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN CHECKING THE OPENING AND CLOSING INTRA-OPERATIVELY, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD, BUT THE JAWS WERE ONLY ABLE TO CLOSE HALFWAY SO THE OPENING AND CLOSING CHECK COULD NOT BE PERFORMED.
18-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | RELIATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC UMBILICAL HERNIA REPAIR, WHEN FIXING THE MESH ON THE SECOND FIRING, THE HINGE AT THE ARTICULATING PART CAME OFF. A NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON LAPAROSCOPIC PROCTECTOMY, WHILE APPROACHING THE INTESTINAL TRACT DURING RECTAL RESECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, BUT THE DEVICE STILL DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON LAPAROSCOPIC PROCTECTOMY, WHILE APPROACHING THE INTESTINAL TRACT DURING RECTAL RESECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, BUT THE DEVICE STILL DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY PROCEDURE, THE LOAD FAILED TO FULLY CLOSE WHILE STAPLING. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Injury | us surgical puerto rico | 14-01-2021 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PARTIAL GASTRECTOMY, THE STAPLE LINE WAS BLEEDING AND OVERSEW WAS DONE TO FIX IT. IT WAS REPORTED THAT FOUR LOADING UNITS WERE USED WITH THE OPEN STAPLER.
16-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | SIGNIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, WHEN CHECKING THE OPENING AND CLOSING INTRA-OPERATIVELY, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD, BUT THE JAWS WERE ONLY ABLE TO CLOSE HALFWAY SO THE OPENING AND CLOSING CHECK COULD NOT BE PERFORMED.
16-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | UNKNOWN EGIA ADAPTER | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, WHEN CHECKING THE OPENING AND CLOSING INTRA-OPERATIVELY, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD, BUT THE JAWS WERE ONLY ABLE TO CLOSE HALFWAY SO THE OPENING AND CLOSING CHECK COULD NOT BE PERFORMED.
18-12-2020 | Injury | us surgical puerto rico | 14-01-2021 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PARTIAL GASTRECTOMY, THE STAPLE LINE WAS BLEEDING AND OVERSEW WAS DONE TO FIX IT. IT WAS REPORTED THAT FOUR LOADING UNITS WERE USED WITH THE OPEN STAPLER.
21-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | ILS 21MM, CURVED | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN GASTRECTOMY, THE ANVIL COULD NOT BE ATTACHED TO THE TROCAR BEFORE USE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Injury | us surgical puerto rico | 14-01-2021 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PARTIAL GASTRECTOMY, THE STAPLE LINE WAS BLEEDING AND OVERSEW WAS DONE TO FIX IT. IT WAS REPORTED THAT FOUR LOADING UNITS WERE USED WITH THE OPEN STAPLER.
18-12-2020 | Injury | us surgical puerto rico | 14-01-2021 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PARTIAL GASTRECTOMY, THE STAPLE LINE WAS BLEEDING AND OVERSEW WAS DONE TO FIX IT. IT WAS REPORTED THAT FOUR LOADING UNITS WERE USED WITH THE OPEN STAPLER.
18-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOIDECTOMY, THE SURGEON CLAMPED THE SIGMOID COLON FOR RESECTION WITH THE HANDLE AND 60MM PURPLE RELOAD. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT COULD CLOSE THE JAW COMPLETELY. THE GREEN BUTTON COULD NOT BE PRESSED FOR FIRING AS WELL. ANOTHER SET OF HANDLE AND 60MM PURPLE RELOAD WAS USED BUT HAD THE SAME EXPERIENCE. ANOTHER STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOIDECTOMY, THE SURGEON CLAMPED THE SIGMOID COLON FOR RESECTION WITH THE HANDLE AND 60MM PURPLE RELOAD. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT COULD CLOSE THE JAW COMPLETELY. THE GREEN BUTTON COULD NOT BE PRESSED FOR FIRING AS WELL. ANOTHER SET OF HANDLE AND 60MM PURPLE RELOAD WAS USED BUT HAD THE SAME EXPERIENCE. ANOTHER STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOIDECTOMY, THE SURGEON CLAMPED THE SIGMOID COLON FOR RESECTION WITH THE HANDLE AND 60MM PURPLE RELOAD. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT COULD CLOSE THE JAW COMPLETELY. THE GREEN BUTTON COULD NOT BE PRESSED FOR FIRING AS WELL. ANOTHER SET OF HANDLE AND 60MM PURPLE RELOAD WAS USED BUT HAD THE SAME EXPERIENCE. ANOTHER STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOIDECTOMY, THE SURGEON CLAMPED THE SIGMOID COLON FOR RESECTION WITH THE HANDLE AND 60MM PURPLE RELOAD. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT COULD CLOSE THE JAW COMPLETELY. THE GREEN BUTTON COULD NOT BE PRESSED FOR FIRING AS WELL. ANOTHER SET OF HANDLE AND 60MM PURPLE RELOAD WAS USED BUT HAD THE SAME EXPERIENCE. ANOTHER STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY, STAPLE MALFORMATION. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH NO DELAY. THERE WERE NO PATIENT CONSEQUENCES.
18-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | ENDO ILS 25MM, CURVED | Failure to Fire | A050501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGOJEJUNOSTOMY, LEAKAGE OCCURRED. THE STAPLES AT HALF OF THE ANASTOMOSED SITE WERE NOT DEPLOYED. THEY NEITHER FELL OFF NOR REMAIN STICKING TOGETHER TO THE CARTRIDGE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE PROCEDURE WAS PROLONGED 3 HOURS. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | ENDO ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGOJEJUNOSTOMY, LEAKAGE OCCURRED. THE STAPLES AT HALF OF THE ANASTOMOSED SITE WERE NOT DEPLOYED. THEY NEITHER FELL OFF NOR REMAIN STICKING TOGETHER TO THE CARTRIDGE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE PROCEDURE WAS PROLONGED 3 HOURS. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | ENDO ILS 25MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGOJEJUNOSTOMY, LEAKAGE OCCURRED. THE STAPLES AT HALF OF THE ANASTOMOSED SITE WERE NOT DEPLOYED. THEY NEITHER FELL OFF NOR REMAIN STICKING TOGETHER TO THE CARTRIDGE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE PROCEDURE WAS PROLONGED 3 HOURS. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | ENDO ILS 25MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGOJEJUNOSTOMY, LEAKAGE OCCURRED. THE STAPLES AT HALF OF THE ANASTOMOSED SITE WERE NOT DEPLOYED. THEY NEITHER FELL OFF NOR REMAIN STICKING TOGETHER TO THE CARTRIDGE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE PROCEDURE WAS PROLONGED 3 HOURS. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | ENDO ILS 25MM, CURVED | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGOJEJUNOSTOMY, LEAKAGE OCCURRED. THE STAPLES AT HALF OF THE ANASTOMOSED SITE WERE NOT DEPLOYED. THEY NEITHER FELL OFF NOR REMAIN STICKING TOGETHER TO THE CARTRIDGE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE PROCEDURE WAS PROLONGED 3 HOURS. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | ENDO ILS 25MM, CURVED | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGOJEJUNOSTOMY, LEAKAGE OCCURRED. THE STAPLES AT HALF OF THE ANASTOMOSED SITE WERE NOT DEPLOYED. THEY NEITHER FELL OFF NOR REMAIN STICKING TOGETHER TO THE CARTRIDGE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE PROCEDURE WAS PROLONGED 3 HOURS. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | ENDO ILS 25MM, CURVED | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGOJEJUNOSTOMY, LEAKAGE OCCURRED. THE STAPLES AT HALF OF THE ANASTOMOSED SITE WERE NOT DEPLOYED. THEY NEITHER FELL OFF NOR REMAIN STICKING TOGETHER TO THE CARTRIDGE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE PROCEDURE WAS PROLONGED 3 HOURS. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | ENDO ILS 25MM, CURVED | Patient Device Interaction Problem | A01 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGOJEJUNOSTOMY, LEAKAGE OCCURRED. THE STAPLES AT HALF OF THE ANASTOMOSED SITE WERE NOT DEPLOYED. THEY NEITHER FELL OFF NOR REMAIN STICKING TOGETHER TO THE CARTRIDGE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE PROCEDURE WAS PROLONGED 3 HOURS. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC WEDGE RESECTION OF LUNG, WHEN DEVICE WAS BEING APPLIED ON MARGINAL LUNG TISSUE, THE RELOAD DID NOT FIRE. THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. CHANGED THE RELOAD TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
06-11-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE CANNOT FORMING WELL. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC LEFT SUPERIOR LOBECTOMY, WHEN THE RELOAD WAS AT THE LEFT INTERLOBULAR COLUMN, THE STAPLER WAS ABLE TO FIRE COMPLETELY. THE STAPLES AT THE DISTAL END OF THE RELOAD COULD NOT BE FORMED AND BURST OUT. THE USER REPLACED WITH ANOTHER RELOAD AND USED THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | TA | Failure to Form Staple | A050704 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING CLOSURE OF THE BRONCUS ON AN OPEN PULMONECTOMY, THE STAPLES DID NOT CLOSE IN B SHAPE AND FELL INTO THE PATIENT¿S CAVITY. IT WAS IMMEDIATELY RETRIEVED BY USING A GRASPER AND SUCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | TA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CLOSURE OF THE BRONCUS ON AN OPEN PULMONECTOMY, THE STAPLES DID NOT CLOSE IN B SHAPE AND FELL INTO THE PATIENT¿S CAVITY. IT WAS IMMEDIATELY RETRIEVED BY USING A GRASPER AND SUCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | TA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING CLOSURE OF THE BRONCUS ON AN OPEN PULMONECTOMY, THE STAPLES DID NOT CLOSE IN B SHAPE AND FELL INTO THE PATIENT¿S CAVITY. IT WAS IMMEDIATELY RETRIEVED BY USING A GRASPER AND SUCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | TA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CLOSURE OF THE BRONCUS ON AN OPEN PULMONECTOMY, THE STAPLES DID NOT CLOSE IN B SHAPE AND FELL INTO THE PATIENT¿S CAVITY. IT WAS IMMEDIATELY RETRIEVED BY USING A GRASPER AND SUCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | 55MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, CONFIGURABLE LINEAR STAPLER DIDN¿T FIRE AT ALL. THERE WERE NO PATIENT CONSEQUENCES.
14-01-2021 | Injury | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, THE DEVICE WAS ABLE TO FIRE BUT IT LOCKED ON TISSUE. EXTRA PRESSURE WAS APPLIED TO THE HANDLE WHICH SEEMED TO FREE THE DEVICE SO THE STAPLE LINE COULD BE COMPLETED AND THEN THE JAWS COULD BE OPENED. DUE TO CONCERNS OVER FIRING, THE PATIENT WAS OPENED JUST IN CASE THERE WAS AN UNEXPECTED BLEED.
10-06-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LOWERING PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
19-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, WHEN THE USER DETACHED THE TRANSVERSE COLON, THE HANDLE HAD BEEN FIRED, BUT THE GRIP COULD NOT BE PULLED, AND IT FELT STUCK. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-06-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LOWERING PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
20-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE FIRING KNOB DID NOT MOVE WHEN TRIED TO PUSH THE KNOB AND DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THIS HAPPENED ON THREE RELOADS, EVEN WHEN THE SURGEON REPLACED THE HANDLE. WHEN THE RELOAD WAS REPLACED IT WORKED PROPERLY. THERE WAS NO PATIENT INJURY.
20-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE FIRING KNOB DID NOT MOVE WHEN TRIED TO PUSH THE KNOB AND DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THIS HAPPENED ON THREE RELOADS, EVEN WHEN THE SURGEON REPLACED THE HANDLE. WHEN THE RELOAD WAS REPLACED IT WORKED PROPERLY. THERE WAS NO PATIENT INJURY.
18-05-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | SIGNIA | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING MAINTENANCE, THE SCREEN OF THE HANDLE WENT OUT, WAS BLACK AND HAD NO LIGHT. THE ISSUE OCCURRED ON TWO HANDLES. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
18-05-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING MAINTENANCE, THE SCREEN OF THE HANDLE WENT OUT, WAS BLACK AND HAD NO LIGHT. THE ISSUE OCCURRED ON TWO HANDLES. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
20-12-2020 | Injury | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE PROCEDURE, AFTER A COMPLETE FIRING, IT WAS NOTED THAT THERE WAS MALFORMED STAPLES AND STAPLES DID NOT FORM PROPER B SHAPE. SURGEON OVER SEW THE STAPLE LINE. PATIENT HAD UNANTICIPATED TISSUE LOSS DUE TO THE ISSUE.
27-07-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF LIVER CANCER, WHEN DEVICE WAS BEING APPLIED ON DETACHED THE HEPATIC PORTAL, THE DEVICE CANNOT BE FIRED. THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. RELOAD WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
14-08-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG CANCER SURGERY, WHEN THE USER DETACHED THE BLOOD VESSEL, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED AND USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Injury | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Infection | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE STAPLE LINE FORMED CORRECTLY, BUT IT WAS NOT COM PLETE, THE STAPLE LINE DID NOT FORM DISTALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AS WELL. COMPLETED THE RESECTION WITH ANOTHER STAPLER FROM ANOTHER COMPANY AND REMOVED THE SURGICAL PIECE WITH THE OTHER STAPLER STILL BLOCKED WITH THE JAWS CLAMPING THE TISSUE. WHEN THE SURGICAL PIECE WAS REMOVED, FORCED THE INSTRUMENT IN ORDER TO OPENED IT. IT WAS ALSO NECESSARY TO DO ADDITIONAL LUNG RESECTION IN ORDER TO BE ABLE TO REMOVE THE INSTRUMENT BLOCKED ON TISSUE. THERE WAS AN UNANTICIPATED TISSUE LOSS AND THE SURGICAL TIME EXTENDED FOR ABOUT AN HOUR AS A RESULT OF THIS PROBLEM. THE INITIAL INCISION WAS EXTENDED TO MORE THAN THAN 2.54 CM AND ALSO A NEW INCISION WAS PERFORMED. ON POST SURGERY MEDICAL CHECK, THE INITIAL EXTENDED INCISION THROUGH WHERE THE INSTRUMENT WAS INSERTED GOT INFECTED.
17-12-2020 | Injury | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE STAPLE LINE FORMED CORRECTLY, BUT IT WAS NOT COM PLETE, THE STAPLE LINE DID NOT FORM DISTALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AS WELL. COMPLETED THE RESECTION WITH ANOTHER STAPLER FROM ANOTHER COMPANY AND REMOVED THE SURGICAL PIECE WITH THE OTHER STAPLER STILL BLOCKED WITH THE JAWS CLAMPING THE TISSUE. WHEN THE SURGICAL PIECE WAS REMOVED, FORCED THE INSTRUMENT IN ORDER TO OPENED IT. IT WAS ALSO NECESSARY TO DO ADDITIONAL LUNG RESECTION IN ORDER TO BE ABLE TO REMOVE THE INSTRUMENT BLOCKED ON TISSUE. THERE WAS AN UNANTICIPATED TISSUE LOSS AND THE SURGICAL TIME EXTENDED FOR ABOUT AN HOUR AS A RESULT OF THIS PROBLEM. THE INITIAL INCISION WAS EXTENDED TO MORE THAN THAN 2.54 CM AND ALSO A NEW INCISION WAS PERFORMED. ON POST SURGERY MEDICAL CHECK, THE INITIAL EXTENDED INCISION THROUGH WHERE THE INSTRUMENT WAS INSERTED GOT INFECTED.
17-12-2020 | Injury | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Infection | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE STAPLE LINE FORMED CORRECTLY, BUT IT WAS NOT COM PLETE, THE STAPLE LINE DID NOT FORM DISTALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AS WELL. COMPLETED THE RESECTION WITH ANOTHER STAPLER FROM ANOTHER COMPANY AND REMOVED THE SURGICAL PIECE WITH THE OTHER STAPLER STILL BLOCKED WITH THE JAWS CLAMPING THE TISSUE. WHEN THE SURGICAL PIECE WAS REMOVED, FORCED THE INSTRUMENT IN ORDER TO OPENED IT. IT WAS ALSO NECESSARY TO DO ADDITIONAL LUNG RESECTION IN ORDER TO BE ABLE TO REMOVE THE INSTRUMENT BLOCKED ON TISSUE. THERE WAS AN UNANTICIPATED TISSUE LOSS AND THE SURGICAL TIME EXTENDED FOR ABOUT AN HOUR AS A RESULT OF THIS PROBLEM. THE INITIAL INCISION WAS EXTENDED TO MORE THAN THAN 2.54 CM AND ALSO A NEW INCISION WAS PERFORMED. ON POST SURGERY MEDICAL CHECK, THE INITIAL EXTENDED INCISION THROUGH WHERE THE INSTRUMENT WAS INSERTED GOT INFECTED.
17-12-2020 | Injury | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC SEGMENTECTOMY, THE STAPLE LINE FORMED CORRECTLY, BUT IT WAS NOT COM PLETE, THE STAPLE LINE DID NOT FORM DISTALLY. THE JAWS OF THE DEVICE LOCKED ON TISSUE AS WELL. COMPLETED THE RESECTION WITH ANOTHER STAPLER FROM ANOTHER COMPANY AND REMOVED THE SURGICAL PIECE WITH THE OTHER STAPLER STILL BLOCKED WITH THE JAWS CLAMPING THE TISSUE. WHEN THE SURGICAL PIECE WAS REMOVED, FORCED THE INSTRUMENT IN ORDER TO OPENED IT. IT WAS ALSO NECESSARY TO DO ADDITIONAL LUNG RESECTION IN ORDER TO BE ABLE TO REMOVE THE INSTRUMENT BLOCKED ON TISSUE. THERE WAS AN UNANTICIPATED TISSUE LOSS AND THE SURGICAL TIME EXTENDED FOR ABOUT AN HOUR AS A RESULT OF THIS PROBLEM. THE INITIAL INCISION WAS EXTENDED TO MORE THAN THAN 2.54 CM AND ALSO A NEW INCISION WAS PERFORMED. ON POST SURGERY MEDICAL CHECK, THE INITIAL EXTENDED INCISION THROUGH WHERE THE INSTRUMENT WAS INSERTED GOT INFECTED.
21-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, WHILE LOADING THE RELOAD ONTO THE STAPLER, THE RELOAD WAS ATTACHED TO STAPLER WITHOUT ISSUE. THE STAPLER CYCLED PROPERLY AND RECOGNIZED THE RELOAD. HOWEVER THE RELOAD JAWS DID NOT APPEAR TO CLOSE ALL THE WAY, AND THE RELOAD WOULD NOT FIT INTO A 12MM TROCAR BECAUSE THE JAWS WOULD NOT FULLY CLOSE. THE RELOAD WAS REMOVED AND REINSTALLED BUT AGAIN THE SAME EVENT OCCURRED. ANOTHER RELOAD WAS OPENED AND USED WITH THE SAME HANDLE, ADAPTER AND SHELL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, WHILE LOADING THE RELOAD ONTO THE STAPLER, THE RELOAD WAS ATTACHED TO STAPLER WITHOUT ISSUE. THE STAPLER CYCLED PROPERLY AND RECOGNIZED THE RELOAD. HOWEVER THE RELOAD JAWS DID NOT APPEAR TO CLOSE ALL THE WAY, AND THE RELOAD WOULD NOT FIT INTO A 12MM TROCAR BECAUSE THE JAWS WOULD NOT FULLY CLOSE. THE RELOAD WAS REMOVED AND REINSTALLED BUT AGAIN THE SAME EVENT OCCURRED. ANOTHER RELOAD WAS OPENED AND USED WITH THE SAME HANDLE, ADAPTER AND SHELL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-08-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING THE BRONCHUS IN A LUNG CANCER SURGERY, AFTER CLAMPING THE TISSUE, SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
27-08-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING THE BRONCHUS IN A LUNG CANCER SURGERY, AFTER CLAMPING THE TISSUE, SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
07-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER A GASTRIC SLEEVE OPERATION THAT WAS PERFORMED (B)(6) 2020, THE PATIENT NEEDED A RE-OPERATION DUE TO COAGULATED BLOOD IN THE ABDOMEN FROM THE STAPLE LINE. HOWEVER THIS HAPPENED DESPITE A DRY STAPLE LINE IN THE PRIMARY SURGERY. THE SURGEONS SUCKED OUT THE CLOTTED BLOOD. NO DIRECT SOURCE OF BLEEDING COULD BE FOUND. PUT A DRAINAGE AND THE PATIENT RECOVERED AFTER A COUPLE OF DAYS.
21-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TRANSECTION OF THE ASCENDING COLON ON AN OPEN RIGHT HEMICOLECTOMY, AFTER FIRING, THE SURGEON WAS NOT ABLE TO OPEN THE JAW BY PRESSING THE BLACK RELEASE BUTTON. THE SURGEON WAS ABLE TO RETRIEVE THE KNIFE BY PULLING THE ADJUSTABLE FIRING KNOB BUT FAILED TO OPEN THE JAW. THE SURGEON WAITED FOR SOME TIME AND TRIED AGAIN, AND THE JAW WAS ABLE TO OPEN, AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
16-06-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LUNG CANCER SURGERY, UPON DETACHING BLOOD VESSEL, THE INSTALLATION AND CLOSING STEP WAS NORMAL, BUT IT COULD NOT BE FIRED. IT WAS NOTED THAT THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND COULD NOT SQUEEZE THE HANDLE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LEFT COLECTOMY, UPON RESECTION OF THE SIGMOID, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE RELOAD WAS PREFIRED WHILE USED ON A NEW HANDLE. TWO OTHER SIMILAR RELOADS WITH DIFFERENT LOT NUMBERS WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LEFT COLECTOMY, UPON RESECTION OF THE SIGMOID, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE RELOAD WAS PREFIRED WHILE USED ON A NEW HANDLE. TWO OTHER SIMILAR RELOADS WITH DIFFERENT LOT NUMBERS WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: PROCEDURE: EXTENDED LEFT HEMICOLECTOMY. IT WAS REPORTED THAT THE INSTRUMENT DID NOT TRIGGER. AFTER INSERTING THE BATTERY, THE DISPLAY LIT UP. THE HEAD OF THE INSTRUMENT WAS KNOTTED AND THE CDH31P TRANS ANAL INSERTED. THE STAPLER AND PRESSURE PLATE WERE THEN APPROXIMATED. WHEN THE ORANGE MARK WAS IN THE GREEN AREA, THE SURGEON WERE WAITED TWO MINUTES. THEN UNLOCKED WITH THE RED LEVER AND TRIGGERED WITH THE GRAY LEVER WITHOUT SUCCESS. NO NOISES PERCEPTIBLE. WITH A SECOND STAPLER OF THE SAME SIZE WAS USED TO PRODUCE THE ANASTOMOSIS. NO PATIENT CONSEQUENCES.
07-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER A GASTRIC SLEEVE OPERATION THAT WAS PERFORMED (B)(6) 2020, THE PATIENT NEEDED A RE-OPERATION, SINCE THE PATIENT STARTED TO VOMIT BLOOD. AFTER A GASTROSCOPY THE SURGEONS FOUND A LOT OF BLOOD IN THE STOMACH, PROBABLY FROM THE INNER STAPLE LINE. THE SURGEONS SUCKED OUT THE BLOOD. GAVE BLOOD TRANSFUSION AND FURTHERMORE ADRENALINE TO STOP THE BLEEDING. AFTER TWO DAYS OF MONITORING THE PATIENT COULD RETURN HOME.
07-12-2020 | Injury | ethicon endo-surgery | 14-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Vomiting | Event Description: IT WAS REPORTED THAT AFTER A GASTRIC SLEEVE OPERATION THAT WAS PERFORMED (B)(6) 2020, THE PATIENT NEEDED A RE-OPERATION, SINCE THE PATIENT STARTED TO VOMIT BLOOD. AFTER A GASTROSCOPY THE SURGEONS FOUND A LOT OF BLOOD IN THE STOMACH, PROBABLY FROM THE INNER STAPLE LINE. THE SURGEONS SUCKED OUT THE BLOOD. GAVE BLOOD TRANSFUSION AND FURTHERMORE ADRENALINE TO STOP THE BLEEDING. AFTER TWO DAYS OF MONITORING THE PATIENT COULD RETURN HOME.
17-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. IN ADDITION IT WAS FOUND THAT THE TWO REPORTED RELOAD WAS ALREADY PRE-FIRED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. IN ADDITION IT WAS FOUND THAT THE TWO REPORTED RELOAD WAS ALREADY PRE-FIRED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LUNG BIOPSY PROCEDURE, THE TWO DEVICE'S RELOAD FAILED TO FIRE FULLY, ONLY CUT ON THE FIRST TWO CM OF THE RELOAD. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LUNG BIOPSY PROCEDURE, THE TWO DEVICE'S RELOAD FAILED TO FIRE FULLY, ONLY CUT ON THE FIRST TWO CM OF THE RELOAD. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-07-2020 | Malfunction | covidien lp llc north haven | 14-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREAS SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-07-2020 | Malfunction | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREAS SURGERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL PROCEDURE, LAPAROSCOPIC SIGMOID REPAIR, CDH29P WOULD NOT FIRE. ALL APPEARED TO BE OK BUT WOULD NOT FIRE. NO MOVEMENT AT ALL WHEN TRIGGERED PRESSED. ANOTHER CHD29P WAS OPENED AND THE BATTERY WAS USED IN THE EXISTING STAPLER. SAME RESULT WOULD NOT FIRE. A THIRD CHD29P WAS OPENED AND USED WITH THE ANVIL OF THE OF THE ORIGINAL STAPLER, WHICH STAYED IN PLACE. I ADVISED TO BE SURE THE ANVIL AND TROCAR WERE COMPLETELY CONNECTED AND TO MAKE SURE THE ORANGE BAND WAS NOT VISIBLE. FIRING WAS SUCCESSFUL. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
23-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL PROCEDURE, LAPAROSCOPIC SIGMOID REPAIR, CDH29P WOULD NOT FIRE. ALL APPEARED TO BE OK BUT WOULD NOT FIRE. NO MOVEMENT AT ALL WHEN TRIGGERED PRESSED. ANOTHER CHD29P WAS OPENED AND THE BATTERY WAS USED IN THE EXISTING STAPLER. SAME RESULT WOULD NOT FIRE. A THIRD CHD29P WAS OPENED AND USED WITH THE ANVIL OF THE OF THE ORIGINAL STAPLER, WHICH STAYED IN PLACE. I ADVISED TO BE SURE THE ANVIL AND TROCAR WERE COMPLETELY CONNECTED AND TO MAKE SURE THE ORANGE BAND WAS NOT VISIBLE. FIRING WAS SUCCESSFUL. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
14-01-2021 | Injury | us surgical puerto rico | 14-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, THE DEVICE WAS ABLE TO FIRE BUT IT LOCKED ON TISSUE. EXTRA PRESSURE WAS APPLIED TO THE HANDLE WHICH SEEMED TO FREE THE DEVICE SO THE STAPLE LINE COULD BE COMPLETED AND THEN THE JAWS COULD BE OPENED. DUE TO CONCERNS OVER FIRING, THE PATIENT WAS OPENED JUST IN CASE THERE WAS AN UNEXPECTED BLEED.
22-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLON RESECTION, A CIRCULAR IS USED IN THE PROCEDURE AND AT THE TIME OF OPENING IT, 50% OF THE SUTURE WAS LEFT WITH THE STAPLES OPEN. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH NO DELAY. THERE WERE NO PATIENT CONSEQUENCES.
20-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF RIGHT COLON CANCER SURGERY, WHEN SEVERING THE INTESTINES TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Anemia | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Hypoxia | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Pneumonia | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Septic Shock | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Respiratory Failure | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Rectal Anastomotic Leakage | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
24-04-2019 | Death | covidien lp llc north haven | 14-01-2021 | UNKNOWN STAPLING DEVICE | Adverse Event Without Identified Device or Use Problem | A24 | Renal Impairment | Event Description: ACCORDING TO THE REPORTER, A DEFICIENCY WAS REPORTED AGAINST A GIA 75MM BLUE LOAD STAPLER. THE DEVICE WAS USED IN A PROCEDURE INVOLVING AN EXPLORATORY LAPAROTOMY WITH SMALL BOWEL RESECTION AND REMOVAL OF MESH WITH FISTULA. MEDICAL RECORDS PROVIDED INDICATE THAT BETWEEN FOUR DAYS AND THREE WEEKS AFTER A PROCEDURE WHERE THE DEVICE WAS USED, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAK, ACUTE PAIN, DEVELOPED KIDNEY DISEASE, SEPSIS WITH SEPTIC SHOCK, ACUTE POST-HEMORRHAGIC ANEMIA, CHRONIC HYPOXIA, RESPIRATORY FAILURE, AND PNEUMONIA. POST OPERATIVE TREATMENT DURING THIS TIME INCLUDED INSERTION OF ENDOTRACHEAL AIRWAY, >96 HOURS OF RESPIRATORY VENTILATION, INSERTION OF FEEDING TUBE, DIALYSIS (FROM KIDNEY DISEASE), ABDOMINAL WASHOUT, INTUBATION (FOR HYPOXIA), ADMINISTRATION OF STRESS DOSE STEROIDS, TOTAL PARENTERAL NUTRITION, ADMINISTRATED SUBCUTANEOUS INSULIN FROM TPN/STEROIDS, SURGERY WITH PLACEMENT OF AN END ILEOSTOMY AND GASTROSTOMY TUBE, SURGERY FOR FASCIA CLOSURE/PLACEMENT OF WOUND VAC, ANTIBIOTIC ADMINISTRATION (FOR PNEUMONIA), RIGHT NEPHRECTOMY, AND OPEN APPENDECTOMY. INFORMATION RECEIVED INDICATES THE PATIENT BECAME DECEASED FOUR MONTHS AFTER THE PROCEDURE, DUE TO COMPLICATION OF SMALL BOWEL RESECTION WITH PAIN AND MALNUTRITION. REPORTED UNDERLYING CAUSES INCLUDE CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE.
17-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE FIRING FORCE WAS HIGHER THAN EXPECTED AT THE 4TH FIRING OF FEEA ON THE COLON. A DEVICE WAS FIRED SOMEHOW, BUT THE TARGET TISSUE WAS PUSHED OUT AND COULD NOT BE CUT PROPERLY. THE DEVICE STAPLED AND THE STAPLES WERE UNFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
17-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | 75 MM NTLC SELECTABLE RELOAD | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE FIRING FORCE WAS HIGHER THAN EXPECTED AT THE 4TH FIRING OF FEEA ON THE COLON. A DEVICE WAS FIRED SOMEHOW, BUT THE TARGET TISSUE WAS PUSHED OUT AND COULD NOT BE CUT PROPERLY. THE DEVICE STAPLED AND THE STAPLES WERE UNFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
17-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | 75 MM NTLC SELECTABLE RELOAD | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE FIRING FORCE WAS HIGHER THAN EXPECTED AT THE 4TH FIRING OF FEEA ON THE COLON. A DEVICE WAS FIRED SOMEHOW, BUT THE TARGET TISSUE WAS PUSHED OUT AND COULD NOT BE CUT PROPERLY. THE DEVICE STAPLED AND THE STAPLES WERE UNFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF RECTAL CANCER SURGERY, WHEN SEVERING THE RECTUM TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER SURGERY, AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE( WITH STRAIGHT LEG). CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC HEPATECTOMY SURGERY, WHEN SEVERING LIVER TISSUE, AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | ECHELON ENDOPATH STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, AN OVERCUTTING SITUATION OCCURRED IN THE ENTIRE STAPLE LINE. IT IS UNKNOWN HOW PROCEDURE WAS COMPLETED. PATIENT CONSEQUENCE WAS NOT REPORTED.
17-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | ECHELON ENDOPATH STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, AN OVERCUTTING SITUATION OCCURRED IN THE ENTIRE STAPLE LINE. IT IS UNKNOWN HOW PROCEDURE WAS COMPLETED. PATIENT CONSEQUENCE WAS NOT REPORTED.
21-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE DEPLOYED STAPLES AT THE CUT END WERE UNFORMED AT THE TWELFTH FIRING OF DEVICE. THE DEVICE WAS USED ON THE ASCENDING COLON AND THE ILEUM. THE SURGEON COMMENTED THERE IS A POSSIBILITY THE LYMPH NODE ON THE COLON SIDE WAS SWOLLEN AND THICK, AND THE TARGET TISSUE ON THE ILEUM SIDE WAS NOT VERY THICK. THE TARGET TISSUE WAS NOT COMPRESSED BEFORE THE FIRING, AND THE PULSE FIRING WAS PERFORMED EARLIER. ADDITIONAL SEROSAL MUSCULAR SUTURE WAS PERFORMED TO COMPLETE THE CASE. THE PROCEDURE WAS NOT CONVERTED TO OPEN, THE PATIENT WAS STABLE AFTER PROCEDURE, AND PATIENT STAYED IN HOSPITAL FOR TYPICAL ROUTINE HOSPITAL STAY. NO PATIENT CONSEQUENCE WAS REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE EXTENDED LEFT PARAPAPILLARY NODULE RESECTION SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY, THE KNIFE DID NOT RETURN TO HOME POSITION AFTER IT MOVED ALL THE WAY TO THE DISTAL END OF THE CARTRIDGE. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY, THE KNIFE DID NOT RETURN TO HOME POSITION AFTER IT MOVED ALL THE WAY TO THE DISTAL END OF THE CARTRIDGE. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBECTOMY PROCEDURE, ON THE SECOND FIRING, NOTED THE DEVICE JAW WAS DAMAGED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
15-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2021 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC S1 + 2 RESECTION SURGERY, WHEN SEVERING BRONCHUS TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY.THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS NOT FIRED NORMALLY. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE LOADING UNIT WAS LOADED TO THE GUN AND THE SURGEON WAS READY TO FIRE, SURGEON PRESSED THE GREEN BUTTON AND STARTED FIRING, BUT THE CARTRIDGE DIDN¿T FIRE AND APPEARED PREFIRED. ANOTHER TWO CARTRIDGES OF THE SAME LOT WERE USED AND THE SAME THING HAPPENED, THEN SURGEON USED ANOTHER LOT FOR THE REST OF THE OPERATION WHERE NOTHING WENT WRONG. NO PATIENT INJURY.
17-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE LOADING UNIT WAS LOADED TO THE GUN AND THE SURGEON WAS READY TO FIRE, SURGEON PRESSED THE GREEN BUTTON AND STARTED FIRING, BUT THE CARTRIDGE DIDN¿T FIRE AND APPEARED PREFIRED. ANOTHER TWO CARTRIDGES OF THE SAME LOT WERE USED AND THE SAME THING HAPPENED, THEN SURGEON USED ANOTHER LOT FOR THE REST OF THE OPERATION WHERE NOTHING WENT WRONG. NO PATIENT INJURY.
16-11-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Difficult or Delayed Positioning | A150203 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, WHEN USING THIS RELOAD TO FIRE, THE RESISTANCE WAS EXTREMELY HUGE, BUT THE I-BEAM COULD STILL BE PUSHED FORWARD. THE SURGEON OPTED TO USE ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
18-12-2020 | Injury | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Failure to Cycle | A050703 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE FIRST RELOAD WAS ABLE TO BE FIRED WITH NO ISSUE; HOWEVER, ON THE SECOND FIRING, THE SURGEON WAS ABLE TO FIRE THE RELOAD, BUT IT STARTED PUSHING THE TISSUE OUT OF THE JAWS. THE SURGEON STOPPED AND RETRACTED THE KNIFE BLADE AND OPENED THE JAWS. EXAMINATION OF THE TISSUE REVEALED THAT ONLY ONE SIDE OF STAPLE LINE WAS INTACT AND FINE, AND THE MEDIAL SIDE WAS BLOODY THAT APPEARED TO BE ABSENT OF STAPLES. IN ADDITION, THERE WAS POOR STAPLE FORMATION. TO RESOLVE THE ISSUE, THE SURGEON HAD TO SUTURE THE TISSUE AND USED ANOTHER DEVICE RELOAD IN ORDER TO COMPLETE THE CASE.
18-12-2020 | Injury | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE FIRST RELOAD WAS ABLE TO BE FIRED WITH NO ISSUE; HOWEVER, ON THE SECOND FIRING, THE SURGEON WAS ABLE TO FIRE THE RELOAD, BUT IT STARTED PUSHING THE TISSUE OUT OF THE JAWS. THE SURGEON STOPPED AND RETRACTED THE KNIFE BLADE AND OPENED THE JAWS. EXAMINATION OF THE TISSUE REVEALED THAT ONLY ONE SIDE OF STAPLE LINE WAS INTACT AND FINE, AND THE MEDIAL SIDE WAS BLOODY THAT APPEARED TO BE ABSENT OF STAPLES. IN ADDITION, THERE WAS POOR STAPLE FORMATION. TO RESOLVE THE ISSUE, THE SURGEON HAD TO SUTURE THE TISSUE AND USED ANOTHER DEVICE RELOAD IN ORDER TO COMPLETE THE CASE.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC INTER-LOBAR RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. IN ADDITION, THE DEVICE JAWS LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON REMOVED THE JAWS FROM THE TISSUE AND USED ANOTHER RELOAD IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC INTER-LOBAR RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. IN ADDITION, THE DEVICE JAWS LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON REMOVED THE JAWS FROM THE TISSUE AND USED ANOTHER RELOAD IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ILEOCECUM RESECTION, ON THE FOURTH FIRING, THE SURGEON WAS ABLE TO PRESS THE FIRING BUTTON AND THE FIRING STARTED; HOWEVER, THE CYCLE STOPPED HALFWAY AND COULD NOT CONTINUE. IN ADDITION, THE DEVICE JAWS LOCKED ON THE TISSUE. TO RESOLVE THE ISSUE, THE SURGEON REMOVED THE DEVICE FROM THE PATIENT AND USED ANOTHER RELOAD IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ILEOCECUM RESECTION, ON THE FOURTH FIRING, THE SURGEON WAS ABLE TO PRESS THE FIRING BUTTON AND THE FIRING STARTED; HOWEVER, THE CYCLE STOPPED HALFWAY AND COULD NOT CONTINUE. IN ADDITION, THE DEVICE JAWS LOCKED ON THE TISSUE. TO RESOLVE THE ISSUE, THE SURGEON REMOVED THE DEVICE FROM THE PATIENT AND USED ANOTHER RELOAD IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN PANCREATIC TAIL RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Injury | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Difficult or Delayed Positioning | A150203 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO APPENDECTOMY, WHEN FIRING THE STAPLER, IT STARTED TO BECOME RESISTANT AND AFTER A SNAP THE BLADE LOCKED BEFORE FINISHING THE WHOLE CUT. THE STAPLER LOCKED IN THIS POSITION. WHEN REMOVING THE STAPLER, INADEQUATE STAPLING AND INCOMPLETE CUTTING WERE OBSERVED. IT WAS NECESSARY TO PERFORM SPECIFIC HEMOSTASIS AND STAPLE THE CECUM PART AGAIN. MACERATION OF THE CECUM WAS NOTED. THERE WAS ALSO AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. SURGICAL TIME WAS WAS EXTENDED FOR 30 MINUTES OR MORE. INCISION WAS ALSO EXTENDED.
17-12-2020 | Injury | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO APPENDECTOMY, WHEN FIRING THE STAPLER, IT STARTED TO BECOME RESISTANT AND AFTER A SNAP THE BLADE LOCKED BEFORE FINISHING THE WHOLE CUT. THE STAPLER LOCKED IN THIS POSITION. WHEN REMOVING THE STAPLER, INADEQUATE STAPLING AND INCOMPLETE CUTTING WERE OBSERVED. IT WAS NECESSARY TO PERFORM SPECIFIC HEMOSTASIS AND STAPLE THE CECUM PART AGAIN. MACERATION OF THE CECUM WAS NOTED. THERE WAS ALSO AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. SURGICAL TIME WAS WAS EXTENDED FOR 30 MINUTES OR MORE. INCISION WAS ALSO EXTENDED.
17-12-2020 | Injury | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO APPENDECTOMY, WHEN FIRING THE STAPLER, IT STARTED TO BECOME RESISTANT AND AFTER A SNAP THE BLADE LOCKED BEFORE FINISHING THE WHOLE CUT. THE STAPLER LOCKED IN THIS POSITION. WHEN REMOVING THE STAPLER, INADEQUATE STAPLING AND INCOMPLETE CUTTING WERE OBSERVED. IT WAS NECESSARY TO PERFORM SPECIFIC HEMOSTASIS AND STAPLE THE CECUM PART AGAIN. MACERATION OF THE CECUM WAS NOTED. THERE WAS ALSO AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. SURGICAL TIME WAS WAS EXTENDED FOR 30 MINUTES OR MORE. INCISION WAS ALSO EXTENDED.
17-12-2020 | Injury | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Difficult or Delayed Positioning | A150203 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO APPENDECTOMY, WHEN FIRING THE STAPLER, IT STARTED TO BECOME RESISTANT AND AFTER A SNAP THE BLADE LOCKED BEFORE FINISHING THE WHOLE CUT. THE STAPLER LOCKED IN THIS POSITION. WHEN REMOVING THE STAPLER, INADEQUATE STAPLING AND INCOMPLETE CUTTING WERE OBSERVED. IT WAS NECESSARY TO PERFORM SPECIFIC HEMOSTASIS AND STAPLE THE CECUM PART AGAIN. MACERATION OF THE CECUM WAS NOTED. THERE WAS ALSO AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. INCISION WAS ALSO EXTENDED.
17-12-2020 | Injury | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO APPENDECTOMY, WHEN FIRING THE STAPLER, IT STARTED TO BECOME RESISTANT AND AFTER A SNAP THE BLADE LOCKED BEFORE FINISHING THE WHOLE CUT. THE STAPLER LOCKED IN THIS POSITION. WHEN REMOVING THE STAPLER, INADEQUATE STAPLING AND INCOMPLETE CUTTING WERE OBSERVED. IT WAS NECESSARY TO PERFORM SPECIFIC HEMOSTASIS AND STAPLE THE CECUM PART AGAIN. MACERATION OF THE CECUM WAS NOTED. THERE WAS ALSO AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. INCISION WAS ALSO EXTENDED.
17-12-2020 | Injury | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO APPENDECTOMY, WHEN FIRING THE STAPLER, IT STARTED TO BECOME RESISTANT AND AFTER A SNAP THE BLADE LOCKED BEFORE FINISHING THE WHOLE CUT. THE STAPLER LOCKED IN THIS POSITION. WHEN REMOVING THE STAPLER, INADEQUATE STAPLING AND INCOMPLETE CUTTING WERE OBSERVED. IT WAS NECESSARY TO PERFORM SPECIFIC HEMOSTASIS AND STAPLE THE CECUM PART AGAIN. MACERATION OF THE CECUM WAS NOTED. THERE WAS ALSO AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. INCISION WAS ALSO EXTENDED.
21-12-2020 | Injury | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Difficult or Delayed Positioning | A150203 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BAND TO GASTRIC SLEEVE PROCEDURE, WHEN TRANSECTING/STAPLING THE STOMACH (ON THE 4TH FIRING), THE SURGEON HAD DIFFICULTY FIRING THE RELOAD AND UPON OPENING THE STAPLER, IT WAS NOTICED THAT THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY. IT APPEARED AS THOUGH ALL OF THE ROWS WERE NOT THERE. THERE WAS ALSO INCOMPLETE STAPLE FORMATION. TISSUE DAMAGE WAS NOTED. THE SURGEON IMPLICATED THE STAPLE LINE AND OVER SEWED IT WITH SUTURE. THE NURSE OPENED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE WITHOUT FURTHER INCIDENTS.
21-12-2020 | Injury | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BAND TO GASTRIC SLEEVE PROCEDURE, WHEN TRANSECTING/STAPLING THE STOMACH (ON THE 4TH FIRING), THE SURGEON HAD DIFFICULTY FIRING THE RELOAD AND UPON OPENING THE STAPLER, IT WAS NOTICED THAT THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY. IT APPEARED AS THOUGH ALL OF THE ROWS WERE NOT THERE. THERE WAS ALSO INCOMPLETE STAPLE FORMATION. TISSUE DAMAGE WAS NOTED. THE SURGEON IMPLICATED THE STAPLE LINE AND OVER SEWED IT WITH SUTURE. THE NURSE OPENED ANOTHER HANDLE AND RELOAD TO COMPLETE THE CASE WITHOUT FURTHER INCIDENTS.
21-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, ON A LAPAROSCOPIC GASTROPLASTY BYPASS, ON THE BEGINNING OF THE FIRST CROSS CLAMPING, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND COULD NOT SQUEEZE THE HANDLE. THE LOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY,
17-07-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LUNG CANCER SURGERY, WHEN DETACHING BRONCHUS, POOR STAPLE FORMATION WAS OBSERVED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | PREMIUM PLUS CEEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN END-TO-END ANASTOMOSIS OF ESOPHAGUS AND JEJUNUM IN A TOTAL GASTRECTOMY, THE GREEN BAR WAS VISIBLE IN THE INDICATOR WINDOW. AFTER PRESSING THE FIRING INSURANCE, THE DEVICE COULD NOT BE FIRED. IT WAS LATER REPLACED WITH ANOTHER STAPLER, AND THE SAME FAILED TO FIRE. AFTER REPLACING THE PRODUCT WITH THE SAME MODEL, IT COULD BE SUCCESSFULLY FIRED AND THE OPERATION CONTINUED. NO ADDITIONAL HARM WAS CAUSED TO THE PATIENT.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | PREMIUM PLUS CEEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN END-TO-END ANASTOMOSIS OF ESOPHAGUS AND JEJUNUM IN A TOTAL GASTRECTOMY, THE GREEN BAR WAS VISIBLE IN THE INDICATOR WINDOW. AFTER PRESSING THE FIRING INSURANCE, THE DEVICE COULD NOT BE FIRED. IT WAS LATER REPLACED WITH ANOTHER STAPLER, AND THE SAME FAILED TO FIRE. AFTER REPLACING THE PRODUCT WITH THE SAME MODEL, IT COULD BE SUCCESSFULLY FIRED AND THE OPERATION CONTINUED. NO ADDITIONAL HARM WAS CAUSED TO THE PATIENT.
01-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS NOT FIRED NORMALLY. THERE WAS NO PATIENT INJURY.
09-11-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SECOND FIRING IN A SLEEVE GASTRECTOMY, THE EXCESSIVE TISSUE MESSAGE APPEARED ON THE SCREEN DURING FIRING, BUT POWERED STAPLER FIRING SPEED NEVER SLOWED DOWN OR STOPPED WITH THE MESSAGE. SURGEON CHECKED STAPLE LINE AND IT WAS PERFECT, BUT HAD CONCERNS THAT THE MESSAGE ON SCREEN DID NOT MATCH THE STAPLER STOPPING. NOTHING WAS DONE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
09-11-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SECOND FIRING IN A SLEEVE GASTRECTOMY, THE EXCESSIVE TISSUE MESSAGE APPEARED ON THE SCREEN DURING FIRING, BUT POWERED STAPLER FIRING SPEED NEVER SLOWED DOWN OR STOPPED WITH THE MESSAGE. SURGEON CHECKED STAPLE LINE AND IT WAS PERFECT, BUT HAD CONCERNS THAT THE MESSAGE ON SCREEN DID NOT MATCH THE STAPLER STOPPING. NOTHING WAS DONE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
19-10-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE JAW WOULD NOT CLOSE AND OPEN. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-10-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE DEVICE JAW WOULD NOT CLOSE AND OPEN. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2020 | Malfunction | ethicon endo-surgery | 15-01-2021 | ECHELON FLEX | Entrapment of Device | A150208 | Insufficient Information | Event Description: PLEE60A STAPLER FIRED TWICE THEN ON THIRD FIRE GOT STUCK ON TISSUE AND WOULD NOT FIRE OR OPEN/RELEASE.
30-12-2020 | Malfunction | ethicon endo-surgery | 15-01-2021 | ECHELON FLEX | Failure to Fire | A050501 | Insufficient Information | Event Description: PLEE60A STAPLER FIRED TWICE THEN ON THIRD FIRE GOT STUCK ON TISSUE AND WOULD NOT FIRE OR OPEN/RELEASE.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE EVEN IF THE SURGEON WAS ABLE PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE EVEN IF THE SURGEON WAS ABLE PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-11-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMIHEPATECTOMY, WHILE CUTTING THE LIVER, THERE WAS AN ISSUE IN UNLOADING THE DEVICES. THE RELOADS OF THREE STAPLERS COULD NOT BE PULLED BACK TO DISCONNECT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-11-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMIHEPATECTOMY, WHILE CUTTING THE LIVER, THERE WAS AN ISSUE IN UNLOADING THE DEVICES. THE RELOADS OF THREE STAPLERS COULD NOT BE PULLED BACK TO DISCONNECT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-11-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMIHEPATECTOMY, WHILE CUTTING THE LIVER, THERE WAS AN ISSUE IN UNLOADING THE DEVICES. THE RELOADS OF THREE STAPLERS COULD NOT BE PULLED BACK TO DISCONNECT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-11-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMIHEPATECTOMY, WHILE CUTTING THE LIVER, THERE WAS AN ISSUE IN UNLOADING THE DEVICES. THE RELOADS OF THREE STAPLERS COULD NOT BE PULLED BACK TO DISCONNECT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-11-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMIHEPATECTOMY, WHILE CUTTING THE LIVER, THERE WAS AN ISSUE IN UNLOADING THE DEVICES. THE RELOADS OF THREE STAPLERS COULD NOT BE PULLED BACK TO DISCONNECT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-11-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMIHEPATECTOMY, WHILE CUTTING THE LIVER, THERE WAS AN ISSUE IN UNLOADING THE DEVICES. THE RELOADS OF THREE STAPLERS COULD NOT BE PULLED BACK TO DISCONNECT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-10-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING, THE RED LIGHT WAS ILLUMINATED. LIQUID WAS FOUND TO HAVE LEAKED FROM THE BOTTOM OF THE PO WER HANDLE. THE DEVICE WAS NOT USED AND WAS REPLACED WITH A REPAIR ALTERNATIVE DEMO PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE DEVICE FOUND THAT THE BATTERY CELLS WERE VENTED.
21-10-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING, THE RED LIGHT WAS ILLUMINATED. LIQUID WAS FOUND TO HAVE LEAKED FROM THE BOTTOM OF THE PO WER HANDLE. THE DEVICE WAS NOT USED AND WAS REPLACED WITH A REPAIR ALTERNATIVE DEMO PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE DEVICE FOUND THAT THE BATTERY CELLS WERE VENTED.
21-10-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING, THE RED LIGHT WAS ILLUMINATED. LIQUID WAS FOUND TO HAVE LEAKED FROM THE BOTTOM OF THE PO WER HANDLE. THE DEVICE WAS NOT USED AND WAS REPLACED WITH A REPAIR ALTERNATIVE DEMO PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE DEVICE FOUND THAT THE BATTERY CELLS WERE VENTED.
18-11-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMIHEPATECTOMY, WHILE CUTTING THE LIVER, THERE WAS AN ISSUE IN UNLOADING THE DEVICES. THE RELOADS OF THREE STAPLERS COULD NOT BE PULLED BACK TO DISCONNECT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-11-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMIHEPATECTOMY, WHILE CUTTING THE LIVER, THERE WAS AN ISSUE IN UNLOADING THE DEVICES. THE RELOADS OF THREE STAPLERS COULD NOT BE PULLED BACK TO DISCONNECT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, AS SOON AS THE PRODUCT WAS OPENED, IT BROKE AND THE ACCESSORIES WERE DISCONNECTED. A SECOND STAPLER WAS ALSO REPORTED TO HAVE BEEN BROKEN. THERE WAS NO PATIENT INVOLVEMENT.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, AS SOON AS THE PRODUCT WAS OPENED, IT BROKE AND THE ACCESSORIES WERE DISCONNECTED. A SECOND STAPLER WAS ALSO REPORTED TO HAVE BEEN BROKEN. THERE WAS NO PATIENT INVOLVEMENT.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, AS SOON AS THE PRODUCT WAS OPENED, IT BROKE AND THE ACCESSORIES WERE DISCONNECTED. A SECOND STAPLER WAS ALSO REPORTED TO HAVE BEEN BROKEN. THERE WAS NO PATIENT INVOLVEMENT.
21-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, SOME STAPLES CAME OUT AND SOME DID NOT THEN REPEATED. NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-10-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | SIGNIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE SHELL WAS OPENED FOR THE POWER HANDLE, BUT THE HANDLE DID NOT WORK. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
22-10-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE SHELL WAS OPENED FOR THE POWER HANDLE, BUT THE HANDLE DID NOT WORK. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
22-10-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE SHELL WAS OPENED FOR THE POWER HANDLE, BUT THE HANDLE DID NOT WORK. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
14-12-2020 | Injury | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF A LAPAROSCOPIC RIGHT RADICAL NEPHRECTOMY, THE PATIENT HAD ACUTE BLEEDING OF MORE THAN 500CC REQUIRING BLOOD TRANSFUSION. DUE TO THE BLEEDING, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM TWO DAYS LATER FOR AN EXPLORATORY LAPAROSCOPY WITH ABDOMINAL WASHOUT. ARTERIAL BLEEDING WAS NOTED AT ONE OF THE PRIOR OPERATION¿S STAPLE LINES, WHICH WAS OVER SEWN TO ACHIEVE HEMOSTASIS. THE PATIENT'S HOSPITAL STAY WAS EXTENDED.
01-10-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | EEA ORVIL | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, THE ANVIL WOULD NOT TILT WHEN PAIRED WITH THE HANDLE. THE STAPLER WAS FIRED WITH EASE. THERE WAS NO PATIENT INJURY.
01-10-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | UNKNOWN EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, THE ANVIL WOULD NOT TILT WHEN PAIRED WITH THE HANDLE. THE STAPLER WAS FIRED WITH EASE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACO/WEDGE/SEGMENTAL RESECTION, WHEN RESECTING THE LUNG PARENCHYMA, AFTER THE JAWS WERE OPENED, IT WAS FOUND THAT THE DEVICE STOPPED FIRING HALFWAY. DID NOT LOOK AT THE SCREEN ON THE POWERED STAPLER, SO IT COULD NOT BE CONFIRMED IF EXCESSIVE FORCE WAS DISPLAYED. IT WAS UNKNOWN WHETHER THE STAPLING STOPPAGE WAS DUE TO THE THICKNESS OF THE TISSUE OR WHETHER FOREIGN MATERIAL RESULTED IN RESISTANCE, BUT THERE WAS NO FOREIGN MATERIAL. NEW RELOAD WAS USED RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN ABDOMINAL SURGERY, THE DEVICE DID NOT FIRE. CHANGED THE DEVICE TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN ABDOMINAL SURGERY, THE DEVICE DID NOT FIRE. CHANGED THE DEVICE TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PANCREATIC TAIL RESECTION, THE JAWS OF THE DEVICE LOCKED ON TISSUE, BUT WAS REMOVED EVENTUALLY. IT WAS ALSO NOTED THAT THE PRE-ATTACHED BUTTRESS MATERIAL WAS FOUND TO BE DISPLACED, SO THE USE WAS STOPPED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PANCREATIC TAIL RESECTION, THE JAWS OF THE DEVICE LOCKED ON TISSUE, BUT WAS REMOVED EVENTUALLY. IT WAS ALSO NOTED THAT THE PRE-ATTACHED BUTTRESS MATERIAL WAS FOUND TO BE DISPLACED, SO THE USE WAS STOPPED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CUTTING AND CLOSING THE RECTUM ON A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, THE DEVICE WAS ABLE TO FIRE, HOWEVER, THE STAPLES BURST OUT. AN INCOMPLETE STAPLE LINE PROXIMALLY WAS ALSO NOTED. THE SURGEON OPENED NEW STAPLE TO FIX THE STAPLE LINE AND THE CASE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE USING THE DEVICE ON THE STOMACH IN A LOW ANTERIOR RESECTION, THE RELOADS WERE NOT ABLE TO FIRE EVEN IF THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE. SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE.
11-11-2020 | Malfunction | ethicon endo-surgery | 15-01-2021 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED. USED SUTURE TO OVERSEW. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE USING THE DEVICE ON THE STOMACH IN A LOW ANTERIOR RESECTION, THE RELOADS WERE NOT ABLE TO FIRE EVEN IF THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE. SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE.
21-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE STAPLER WAS NOT ABLE TO FIRE. NO PATIENT INJURY.
11-12-2020 | Malfunction | celestica de monterrey s.a de c.v | 15-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LUNG RESECTION FOR LAPAROSCOPIC WEDGE SURGERY, POOR STAPLE FORMATION WAS OBSERVED ON THE FRONT PART OF FOUR RELOADS USED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, ON THE THIRD FIRING, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE AND THERE WAS A UNUSUAL CRACKING SOUND AS WELL AS PHYSICAL CRACK WHEN ATTEMPTING TO SQUEEZE THE HANDLE. THE SURGEON TRIED TO SQUEEZE THE HANDLE AGAIN FOR TWO TO THREE TIMES, BUT IT DID NOT WORK. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, ON THE THIRD FIRING, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE AND THERE WAS A UNUSUAL CRACKING SOUND AS WELL AS PHYSICAL CRACK WHEN ATTEMPTING TO SQUEEZE THE HANDLE. THE SURGEON TRIED TO SQUEEZE THE HANDLE AGAIN FOR TWO TO THREE TIMES, BUT IT DID NOT WORK. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, ON THE THIRD FIRING, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE AND THERE WAS A UNUSUAL CRACKING SOUND AS WELL AS PHYSICAL CRACK WHEN ATTEMPTING TO SQUEEZE THE HANDLE. THE SURGEON TRIED TO SQUEEZE THE HANDLE AGAIN FOR TWO TO THREE TIMES, BUT IT DID NOT WORK. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, WHEN STAPLING STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD BUT PROBLEM PERSISTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, WHEN STAPLING STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD BUT PROBLEM PERSISTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC BYPASS, WHEN STAPLING STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON NOR SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD BUT PROBLEM PERSISTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPY ASSISTED WEDGE/SEGMENTAL RESECTION, BEFORE FIRING, IT WAS FOUND THAT THERE WAS A FOREIGN MATERIAL LIKE TRANSPARENT PLASTIC ON THE CLAMP COVER, AND WHEN THEY FIRED, IT CAME OFF. FOREIGN MATERIAL GOT ON THE RELOAD THAT WAS PICKED UP WITH FORCEPS AND TOOK IT OUT, SO IT DID NOT FALL INTO THE CAVITY. AN X-RAY WAS PERFORMED JUST FOR SAFETY. THE SAME RELOAD AND HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA | Device Contamination with Chemical or Other Material | A180104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPY ASSISTED WEDGE/SEGMENTAL RESECTION, BEFORE FIRING, IT WAS FOUND THAT THERE WAS A FOREIGN MATERIAL LIKE TRANSPARENT PLASTIC ON THE CLAMP COVER, AND WHEN THEY FIRED, IT CAME OFF. FOREIGN MATERIAL GOT ON THE RELOAD THAT WAS PICKED UP WITH FORCEPS AND TOOK IT OUT, SO IT DID NOT FALL INTO THE CAVITY. AN X-RAY WAS PERFORMED JUST FOR SAFETY. THE SAME RELOAD AND HANDLE WAS USED TO COMPLETE THE CASE.
22-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING STAPLING IN A SLEEVE PROCEDURE, SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE WAS NOT ABLE TO FIRE. SURGEON REPLACED THE STAPLER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN PANCREATICODUODENECTOMY, AT THE TIMING OF PANCREAS RESECTION, THE DEVICE WOULD NOT FIRE - GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. THE DEVICE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED BY PULLING THE RELEASE BLACK RETURN KNOB. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN PANCREATICODUODENECTOMY, AT THE TIMING OF PANCREAS RESECTION, THE DEVICE WOULD NOT FIRE - GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. THE DEVICE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED BY PULLING THE RELEASE BLACK RETURN KNOB. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN PANCREATICODUODENECTOMY, AT THE TIMING OF PANCREAS RESECTION, THE DEVICE WOULD NOT FIRE - GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. THE DEVICE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED BY PULLING THE RELEASE BLACK RETURN KNOB. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 15-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN PANCREATICODUODENECTOMY, AT THE TIMING OF PANCREAS RESECTION, THE DEVICE WOULD NOT FIRE - GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. THE DEVICE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED BY PULLING THE RELEASE BLACK RETURN KNOB. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | us surgical puerto rico | 15-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPY RIGHT UPPER LUNG LOBECTOMY, RESECTION OF LUNG TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. REPLACED THE HANDLE TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | ethicon endo-surgery | 15-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult or Delayed Activation | A150102 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS DIFFICULT TO LOAD WITH FIRST BLACK RELOAD. VERY EXPERIENCED NURSE (>1000 LOADINGS) DESCRIBED THAT IT WAS VERY DIFFICULT TO SNAP THE RELOAD INTO THE CORRECT POSITION WHILST THE RELOAD AND RELOAD ALIGNMENT SLOT WERE ALIGNED. SHE COMMENTED TO THE SURGEON THAT IT WAS DIFFICULT TO LOAD. SURGEON PROCEEDED TO FIRE THE DEVICE AS NORMAL (PULSE FIRING ON FIRST TRANSECTION OF GASTRIC TISSUE) WITH NORMAL FIRING/STAPLE LINE OBSERVED. KNIFE RETURNED TO HOME POSITION AND SURGEON TRIED TO OPEN THE DEVICE AS PER NORMAL (SHE IS VERY EXPERIENCED USER) . SURGEON SAID DEVICE FAILED TO OPEN. SHE RE-TRIED AND AGAIN WOULD NOT OPEN BY SQUEEZING THE CLOSING TRIGGER; SIMULTANEOUSLY PUSHING THE ANVIL RELEASE AND RELEASING THE CLOSING TRIGGER. SHE TRIED MULTIPLE TIMES AND IT FINALLY OPENED. SHE COMMENTED THAT IT FELT LIKE THERE WAS LOWER PRESSURE NEEDED TO FIRE THE DEVICE USING THE GREY TRIGGER AND THAT WHEN TRYING TO OPEN IT FELT LIKE IT JAMMED/STOPPED WHILST OPENING. AFTER SURGERY THEY TRIED IT WITH ANOTHER RELOAD AND FELT IT WAS AGAIN HARD TO LOAD, FIRED AS NORMAL AND WHEN THEY TRIED TO OPEN IT PARTIALLY OPENED ONLY WITH THE CLOSING TRIGGER STOPPING AT ABOUT 50% OPEN AND THE JAW STILL CLOSED. VIDEO IS AVAILABLE. FROM THE VIDEO YOU CAN SEE THAT WHEN THE CLOSING TRIGGER IS RELEASED, IT STOPS OPENING WITH THE RED SAFETY STILL EXPOSED SHOWING IT WONT OPEN FULLY. NO PATIENT CONSEQUENCE REPORTED.
21-12-2020 | Malfunction | ethicon endo-surgery | 15-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS DIFFICULT TO LOAD WITH FIRST BLACK RELOAD. VERY EXPERIENCED NURSE (>1000 LOADINGS) DESCRIBED THAT IT WAS VERY DIFFICULT TO SNAP THE RELOAD INTO THE CORRECT POSITION WHILST THE RELOAD AND RELOAD ALIGNMENT SLOT WERE ALIGNED. SHE COMMENTED TO THE SURGEON THAT IT WAS DIFFICULT TO LOAD. SURGEON PROCEEDED TO FIRE THE DEVICE AS NORMAL (PULSE FIRING ON FIRST TRANSECTION OF GASTRIC TISSUE) WITH NORMAL FIRING/STAPLE LINE OBSERVED. KNIFE RETURNED TO HOME POSITION AND SURGEON TRIED TO OPEN THE DEVICE AS PER NORMAL (SHE IS VERY EXPERIENCED USER) . SURGEON SAID DEVICE FAILED TO OPEN. SHE RE-TRIED AND AGAIN WOULD NOT OPEN BY SQUEEZING THE CLOSING TRIGGER; SIMULTANEOUSLY PUSHING THE ANVIL RELEASE AND RELEASING THE CLOSING TRIGGER. SHE TRIED MULTIPLE TIMES AND IT FINALLY OPENED. SHE COMMENTED THAT IT FELT LIKE THERE WAS LOWER PRESSURE NEEDED TO FIRE THE DEVICE USING THE GREY TRIGGER AND THAT WHEN TRYING TO OPEN IT FELT LIKE IT JAMMED/STOPPED WHILST OPENING. AFTER SURGERY THEY TRIED IT WITH ANOTHER RELOAD AND FELT IT WAS AGAIN HARD TO LOAD, FIRED AS NORMAL AND WHEN THEY TRIED TO OPEN IT PARTIALLY OPENED ONLY WITH THE CLOSING TRIGGER STOPPING AT ABOUT 50% OPEN AND THE JAW STILL CLOSED. VIDEO IS AVAILABLE. FROM THE VIDEO YOU CAN SEE THAT WHEN THE CLOSING TRIGGER IS RELEASED, IT STOPS OPENING WITH THE RED SAFETY STILL EXPOSED SHOWING IT WONT OPEN FULLY. NO PATIENT CONSEQUENCE REPORTED.
21-12-2020 | Malfunction | ethicon endo-surgery | 15-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS DIFFICULT TO LOAD WITH FIRST BLACK RELOAD. VERY EXPERIENCED NURSE (>1000 LOADINGS) DESCRIBED THAT IT WAS VERY DIFFICULT TO SNAP THE RELOAD INTO THE CORRECT POSITION WHILST THE RELOAD AND RELOAD ALIGNMENT SLOT WERE ALIGNED. SHE COMMENTED TO THE SURGEON THAT IT WAS DIFFICULT TO LOAD. SURGEON PROCEEDED TO FIRE THE DEVICE AS NORMAL (PULSE FIRING ON FIRST TRANSECTION OF GASTRIC TISSUE) WITH NORMAL FIRING/STAPLE LINE OBSERVED. KNIFE RETURNED TO HOME POSITION AND SURGEON TRIED TO OPEN THE DEVICE AS PER NORMAL (SHE IS VERY EXPERIENCED USER) . SURGEON SAID DEVICE FAILED TO OPEN. SHE RE-TRIED AND AGAIN WOULD NOT OPEN BY SQUEEZING THE CLOSING TRIGGER; SIMULTANEOUSLY PUSHING THE ANVIL RELEASE AND RELEASING THE CLOSING TRIGGER. SHE TRIED MULTIPLE TIMES AND IT FINALLY OPENED. SHE COMMENTED THAT IT FELT LIKE THERE WAS LOWER PRESSURE NEEDED TO FIRE THE DEVICE USING THE GREY TRIGGER AND THAT WHEN TRYING TO OPEN IT FELT LIKE IT JAMMED/STOPPED WHILST OPENING. AFTER SURGERY THEY TRIED IT WITH ANOTHER RELOAD AND FELT IT WAS AGAIN HARD TO LOAD, FIRED AS NORMAL AND WHEN THEY TRIED TO OPEN IT PARTIALLY OPENED ONLY WITH THE CLOSING TRIGGER STOPPING AT ABOUT 50% OPEN AND THE JAW STILL CLOSED. VIDEO IS AVAILABLE. FROM THE VIDEO YOU CAN SEE THAT WHEN THE CLOSING TRIGGER IS RELEASED, IT STOPS OPENING WITH THE RED SAFETY STILL EXPOSED SHOWING IT WONT OPEN FULLY. NO PATIENT CONSEQUENCE REPORTED.
17-12-2020 | Injury | ethicon endo-surgery | 15-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLEEDING. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS: PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLOOD TRANSFUSION. GENERAL: THORACIC: UROLOGIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, HYPERSENSITIVE REACTION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PLEURAL EFFUSION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PNEUMONIA. (GENERAL: HEPATIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SEPSIS. (GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SURGICAL SITE INFECTION. UROLOGIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, URINARY TRACT INFECTION. (GENERAL: HEPATIC, UROLOGIC, SPLEEN) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, 30-DAY REOPERATION.
17-12-2020 | Injury | ethicon endo-surgery | 15-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLEEDING. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS: PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLOOD TRANSFUSION. GENERAL: THORACIC: UROLOGIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, HYPERSENSITIVE REACTION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PLEURAL EFFUSION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PNEUMONIA. (GENERAL: HEPATIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SEPSIS. (GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SURGICAL SITE INFECTION. UROLOGIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, URINARY TRACT INFECTION. (GENERAL: HEPATIC, UROLOGIC, SPLEEN) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, 30-DAY REOPERATION.
17-12-2020 | Injury | ethicon endo-surgery | 15-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLEEDING. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS: PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLOOD TRANSFUSION. GENERAL: THORACIC: UROLOGIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, HYPERSENSITIVE REACTION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PLEURAL EFFUSION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PNEUMONIA. (GENERAL: HEPATIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SEPSIS. (GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SURGICAL SITE INFECTION. UROLOGIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, URINARY TRACT INFECTION. (GENERAL: HEPATIC, UROLOGIC, SPLEEN) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, 30-DAY REOPERATION.
17-12-2020 | Injury | ethicon endo-surgery | 15-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | A24 | Urinary Tract Infection | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLEEDING. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS: PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLOOD TRANSFUSION. GENERAL: THORACIC: UROLOGIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, HYPERSENSITIVE REACTION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PLEURAL EFFUSION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PNEUMONIA. (GENERAL: HEPATIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SEPSIS. (GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SURGICAL SITE INFECTION. UROLOGIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, URINARY TRACT INFECTION. (GENERAL: HEPATIC, UROLOGIC, SPLEEN) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, 30-DAY REOPERATION.
17-12-2020 | Injury | ethicon endo-surgery | 15-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLEEDING. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS: PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLOOD TRANSFUSION. GENERAL: THORACIC: UROLOGIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, HYPERSENSITIVE REACTION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PLEURAL EFFUSION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PNEUMONIA. (GENERAL: HEPATIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SEPSIS. (GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SURGICAL SITE INFECTION. UROLOGIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, URINARY TRACT INFECTION. (GENERAL: HEPATIC, UROLOGIC, SPLEEN) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, 30-DAY REOPERATION.
17-12-2020 | Injury | ethicon endo-surgery | 15-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Patient Device Interaction Problem | A01 | Pneumonia | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLEEDING. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS: PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLOOD TRANSFUSION. GENERAL: THORACIC: UROLOGIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, HYPERSENSITIVE REACTION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PLEURAL EFFUSION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PNEUMONIA. (GENERAL: HEPATIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SEPSIS. (GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SURGICAL SITE INFECTION. UROLOGIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, URINARY TRACT INFECTION. (GENERAL: HEPATIC, UROLOGIC, SPLEEN) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, 30-DAY REOPERATION.
17-12-2020 | Injury | ethicon endo-surgery | 15-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Patient Device Interaction Problem | A01 | Sepsis | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLEEDING. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS: PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLOOD TRANSFUSION. GENERAL: THORACIC: UROLOGIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, HYPERSENSITIVE REACTION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PLEURAL EFFUSION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PNEUMONIA. (GENERAL: HEPATIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SEPSIS. (GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SURGICAL SITE INFECTION. UROLOGIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, URINARY TRACT INFECTION. (GENERAL: HEPATIC, UROLOGIC, SPLEEN) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, 30-DAY REOPERATION.
17-12-2020 | Injury | ethicon endo-surgery | 15-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Patient Device Interaction Problem | A01 | Urinary Tract Infection | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A RELATED RESEARCH ACTIVITY DATABASE (DRRA) FOR PATIENTS UNDERGOING UNKNOWN PROCEDURE. THE REPORTED COMPLICATION AS PER ICD-9 AND ICD-10 CATEGORIZATION EXPERIENCED BY THE FOLLOWING WITH CORRESPONDING INTERVENTION: THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, AIR LEAK. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLEEDING. GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS: PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, BLOOD TRANSFUSION. GENERAL: THORACIC: UROLOGIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, HYPERSENSITIVE REACTION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PLEURAL EFFUSION. THORACIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, PNEUMONIA. (GENERAL: HEPATIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SEPSIS. (GENERAL: HEPATIC, THORACIC, UROLOGIC, SPLEEN, PANCREAS, PEDIATRIC) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, SURGICAL SITE INFECTION. UROLOGIC PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, URINARY TRACT INFECTION. (GENERAL: HEPATIC, UROLOGIC, SPLEEN) PATIENTS HAD PERI-OPERATIVE COMPLICATION AFTER THE PROCEDURE WITH THE USE OF THE DEVICE AND BEFORE THE DISCHARGE OF THE INDEX HOSPITALIZATION WERE DEFINED WITH THE DIAGNOSIS CODE, 30-DAY REOPERATION.
08-01-2021 | Malfunction | ethicon | 15-01-2021 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2021 AND THE ABSORBABLE STRAPS WERE USED. IT WAS REPORTED THAT THE FIRST STAPLE DID NOT CLOSE PROPERLY, STAPLE SEEMED NOT TO CLOSE AS USUAL, A SECOND DEVICE HAD TO BE UNPACKED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE HANDLE WAS MISFIRING. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE HANDLE WAS MISFIRING. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE HANDLE WAS MISFIRING. THERE WAS NO PATIENT INJURY.
30-03-2021 | Malfunction | us surgical puerto rico | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SEPTATION, WHILE PERFORMING THE ENTERO ENTERO ANASTOMOSIS, THE TWO RELOADS DID NOT FIRE. A SECOND STAPLER WITH A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-01-2021 | Malfunction | covidien lp llc north haven | 18-01-2021 | UNKNOWN STAPLING DEVICE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SEPTATION, WHILE PERFORMING THE ENTERO ENTERO ANASTOMOSIS, THE TWO RELOADS DID NOT FIRE. A SECOND STAPLER WITH A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOIDECTOMY, WHEN COLON WAS CUT AND WHEN ATTEMPTED TO FIRE, IT BECAME STIFF IN THE MIDDLE OF SQUEEZING THE HANDLE AND NO STAPLE WAS ABLE TO FIRE AT ALL. THE TISSUE WAS NOT CUT. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOIDECTOMY, WHEN COLON WAS CUT AND WHEN ATTEMPTED TO FIRE, IT BECAME STIFF IN THE MIDDLE OF SQUEEZING THE HANDLE AND NO STAPLE WAS ABLE TO FIRE AT ALL. THE TISSUE WAS NOT CUT. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-11-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THERE WAS A PROBLEM UNLOADING THE DEVICE, THE STAPLES WERE OUT OF THE LOAD, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD AND HANDLE WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-11-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THERE WAS A PROBLEM UNLOADING THE DEVICE, THE STAPLES WERE OUT OF THE LOAD, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD AND HANDLE WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-11-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THERE WAS A PROBLEM UNLOADING THE DEVICE, THE STAPLES WERE OUT OF THE LOAD, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD AND HANDLE WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | No Display/Image | A090206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE HANDLE TURNED AND MADE THE ADJUSTING NOISE, BUT THE SCREEN DID NOT LIGHT UP; THE INTERNAL SCREEN WAS CRACKED AND THERE WERE DEBRIS INSIDE THE UNIT THAT MADE IT RATTLE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Product Quality Problem | A0201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE HANDLE TURNED AND MADE THE ADJUSTING NOISE, BUT THE SCREEN DID NOT LIGHT UP; THE INTERNAL SCREEN WAS CRACKED AND THERE WERE DEBRIS INSIDE THE UNIT THAT MADE IT RATTLE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE HANDLE TURNED AND MADE THE ADJUSTING NOISE, BUT THE SCREEN DID NOT LIGHT UP; THE INTERNAL SCREEN WAS CRACKED AND THERE WERE DEBRIS INSIDE THE UNIT THAT MADE IT RATTLE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING AND CLOSING THE LUNG PARENCHYMA IN A THORACIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON USED ANOTHER RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
23-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE USING THE DEVICE ON THE BLOOD VESSEL IN A LOBECTOMY, AFTER FIRING, IT WAS NOTED THAT THERE WAS INCOMPLETE STAPLE LINE AT THE PROXIMAL END. THERE WAS ALSO POOR STAPLE FORMATION. THE SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
18-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, THERE WAS PROBLEM UNLOADING THE DEVICE. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, BUT THE LOAD DID NOT FIRE. THE DEVICE WAS EXCHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, THERE WAS PROBLEM UNLOADING THE DEVICE. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, BUT THE LOAD DID NOT FIRE. THE DEVICE WAS EXCHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE SURGICAL ONCOLOGIST PLACED THE POWERED STAPLER DEEP INTO PELVIS FOR THE DISTAL COLON RESECTION DURING A LOW ANTERIOR RESECTION. THE RELOAD WAS FULLY ARTICULATED AND THE TISSUE EXTREMELY THICK. WHEN THE STAPLER WAS CLOSED ON TISSUE THE POWERED STAPLER INDICATED EXTREME TISSUE THICKNESS. THE SURGEON WAITED 30 SECONDS, OPEN AND CLOSED THE STAPLER TO CLAMP BACK DOWN ON THE STAPLER. ONCE THE STAPLER WAS ¿FIRED¿ THE RELOAD ELBOW ¿BLEW OUT¿ FROM THE FRAME. THE STAPLER STOPPED AND DID NOT FIRE. THE SURGEON OPENED THE RELOAD AND NOTICED THE MALFORMED ELBOW ON THE RELOAD. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE SURGICAL ONCOLOGIST PLACED THE POWERED STAPLER DEEP INTO PELVIS FOR THE DISTAL COLON RESECTION DURING A LOW ANTERIOR RESECTION. THE RELOAD WAS FULLY ARTICULATED AND THE TISSUE WAS EXTREMELY THICK. WHEN THE STAPLER WAS CLOSED ON TISSUE, THE POWERED STAPLER INDICATED EXTREME TISSUE THICKNESS. THE SURGEON WAITED 30 SECONDS, OPENED AND CLOSED THE STAPLER TO CLAMP BACK DOWN ON THE STAPLER. ONCE THE STAPLER WAS ¿FIRED¿ THE RELOAD ELBOW ¿BLEW OUT¿ FROM THE FRAME. THE STAPLER STOPPED AND DID NOT FIRE. THE SURGEON OPENED THE RELOAD AND NOTICED THE MALFORMED ELBOW ON THE RELOAD. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-11-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LIVER LOBES AND GALLBLADDER RESECTION PROCEDURE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE SURGEON ENDED UP JUST USING A LIGATURE AROUND THE BASE OF THE LIVER LOBES. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, THERE WAS PROBLEM UNLOADING THE DEVICE. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, BUT THE LOAD DID NOT FIRE. THE DEVICE WAS EXCHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, THERE WAS PROBLEM UNLOADING THE DEVICE. THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, BUT THE LOAD DID NOT FIRE. THE DEVICE WAS EXCHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, WHILE ATTEMPTING TO TRANSECT THE INTERLOBAR FISSURE, THE SURGEON WAS ABLE TO CLAMP THE TISSUE, PRESS THE GREEN BUTTON, AND FIRE DOWN. HOWEVER, THE BLADE STOPPED. THE SURGEON TRIED TO PUSH SEVERAL TIMES MORE, BUT THE DEVICE DID NOT MOVE. THE SURGEON RETRACTED THE BLADE AND ANOTHER RELOAD WAS USED WITH THE HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LUNG RESECTION, WHILE BEING APPLIED AT THE PARENCHYMA, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO ENTER FIRING MODE, BUT THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
03-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE SURGEON ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE COULD NOT BE FIRE, IN ADDITION THE DEVICE WAS DIFFICULT TO UNLOAD. THE SURGEON THEN USED ANOTHER DEVICE RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE SURGEON ABLE TO SQUEEZE THE HANDLE BUT THE DEVICE COULD NOT BE FIRE, IN ADDITION THE DEVICE WAS DIFFICULT TO UNLOAD. THE SURGEON THEN USED ANOTHER DEVICE RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-11-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LIVER LOBES AND GALLBLADDER RESECTION PROCEDURE, THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE SURGEON ENDED UP JUST USING A LIGATURE AROUND THE BASE OF THE LIVER LOBES. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, AFTER UNPACKING THE DEVICE, IT WAS FOUND THAT THE CONNECTION BETWEEN THE RELOAD AND HANDLE BODY WAS DAMAGED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
13-10-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOBECTOMY PROCEDURE, THE SURGEON ACTUATED THE HANDLE THE LOADING DID NOT CLOSED. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-10-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOBECTOMY PROCEDURE, THE SURGEON ACTUATED THE HANDLE THE LOADING DID NOT CLOSED. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LUNG TRANSPLANT PROCEDURE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN CUTTING AND CLOSING THE BRONCHUS ON A LAPAROSCOPIC THORACIC, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. ANOTHER STAPLE WAS USED. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC ENUCLEATION OF HEPATIC METASTASES, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY/SIGMOIDECTOMY, AT THE TIME OF TRANSECTION OF THE SIGMOID COLON IN THE ABDOMINAL CAVITY, WHEN THE RELOAD WAS USED IN THE ABDOMINAL CAVITY WITH POWERED STAPLER, THE FIRING STOPPED WHEN IT ADVANCED ABOUT 15 MM. BLEEDING OCCURRED DUE TO STAPLE NON-FORMATION AT THE TIP WHERE THE FIRING ADVANCED. REPAIR WAS PERFORMED OVER THE ABOVE-MENTIONED STAPLE LINES BY PERFORMING ADDITIONAL FIRING WITH MANUAL STAPLER AND A NEW RELOAD. THERE WAS ALSO TISSUE DAMAGE REPORTED AS A RESULT OF THIS PROBLEM.
23-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY/SIGMOIDECTOMY, AT THE TIME OF TRANSECTION OF THE SIGMOID COLON IN THE ABDOMINAL CAVITY, WHEN THE RELOAD WAS USED IN THE ABDOMINAL CAVITY WITH POWERED STAPLER, THE FIRING STOPPED WHEN IT ADVANCED ABOUT 15 MM. BLEEDING OCCURRED DUE TO STAPLE NON-FORMATION AT THE TIP WHERE THE FIRING ADVANCED. REPAIR WAS PERFORMED OVER THE ABOVE-MENTIONED STAPLE LINES BY PERFORMING ADDITIONAL FIRING WITH MANUAL STAPLER AND A NEW RELOAD. THERE WAS ALSO TISSUE DAMAGE REPORTED AS A RESULT OF THIS PROBLEM.
23-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY/SIGMOIDECTOMY, AT THE TIME OF TRANSECTION OF THE SIGMOID COLON IN THE ABDOMINAL CAVITY, WHEN THE RELOAD WAS USED IN THE ABDOMINAL CAVITY WITH POWERED STAPLER, THE FIRING STOPPED WHEN IT ADVANCED ABOUT 15 MM. BLEEDING OCCURRED DUE TO STAPLE NON-FORMATION AT THE TIP WHERE THE FIRING ADVANCED. REPAIR WAS PERFORMED OVER THE ABOVE-MENTIONED STAPLE LINES BY PERFORMING ADDITIONAL FIRING WITH MANUAL STAPLER AND A NEW RELOAD. THERE WAS ALSO TISSUE DAMAGE REPORTED AS A RESULT OF THIS PROBLEM.
23-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY/SIGMOIDECTOMY, AT THE TIME OF TRANSECTION OF THE SIGMOID COLON IN THE ABDOMINAL CAVITY, WHEN THE RELOAD WAS USED IN THE ABDOMINAL CAVITY WITH POWERED STAPLER, THE FIRING STOPPED WHEN IT ADVANCED ABOUT 15 MM. BLEEDING OCCURRED DUE TO STAPLE NON-FORMATION AT THE TIP WHERE THE FIRING ADVANCED. REPAIR WAS PERFORMED OVER THE ABOVE-MENTIONED STAPLE LINES BY PERFORMING ADDITIONAL FIRING WITH MANUAL STAPLER AND A NEW RELOAD. THERE WAS ALSO TISSUE DAMAGE REPORTED AS A RESULT OF THIS PROBLEM.
23-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY/SIGMOIDECTOMY, AT THE TIME OF TRANSECTION OF THE SIGMOID COLON IN THE ABDOMINAL CAVITY, WHEN THE RELOAD WAS USED IN THE ABDOMINAL CAVITY WITH POWERED STAPLER, THE FIRING STOPPED WHEN IT ADVANCED ABOUT 15 MM. BLEEDING OCCURRED DUE TO STAPLE NON-FORMATION AT THE TIP WHERE THE FIRING ADVANCED. REPAIR WAS PERFORMED OVER THE ABOVE-MENTIONED STAPLE LINES BY PERFORMING ADDITIONAL FIRING WITH MANUAL STAPLER AND A NEW RELOAD. THERE WAS ALSO TISSUE DAMAGE REPORTED AS A RESULT OF THIS PROBLEM.
23-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY/SIGMOIDECTOMY, AT THE TIME OF TRANSECTION OF THE SIGMOID COLON IN THE ABDOMINAL CAVITY, WHEN THE RELOAD WAS USED IN THE ABDOMINAL CAVITY WITH POWERED STAPLER, THE FIRING STOPPED WHEN IT ADVANCED ABOUT 15 MM. BLEEDING OCCURRED DUE TO STAPLE NON-FORMATION AT THE TIP WHERE THE FIRING ADVANCED. REPAIR WAS PERFORMED OVER THE ABOVE-MENTIONED STAPLE LINES BY PERFORMING ADDITIONAL FIRING WITH MANUAL STAPLER AND A NEW RELOAD. THERE WAS ALSO TISSUE DAMAGE REPORTED AS A RESULT OF THIS PROBLEM.
10-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC ENUCLEATION OF HEPATIC METASTASES, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE RELOAD WAS LOADED INTO THE STAPLER, THE RELOAD DID NOT OPEN AND THEREFORE COULD NOT BE USED. TWO NEW RELOADS WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
22-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE RELOAD WAS LOADED INTO THE STAPLER, THE RELOAD DID NOT OPEN AND THEREFORE COULD NOT BE USED. THERE WAS NO PATIENT INVOLVEMENT.
22-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE RELOAD WAS LOADED INTO THE STAPLER, THE RELOAD DID NOT OPEN AND THEREFORE COULD NOT BE USED. THERE WAS NO PATIENT INVOLVEMENT.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY, WHEN THE SURGEON WAS GOING TO SECTION THE PULMONARY VESSELS (MEDIASTINAL TRUNK), THE INSTRUMENT HAD BEEN LOADED PERFECTLY AND THE MESSAGES ON THE OLED SCREEN VERIFIED THAT EVERYTHING WAS FINE. HOWEVER, WHEN STAPLING STARTED, AT APPROXIMATELY 1.5 CM IT CRASHED AND THE OLED SCREEN TURNED OFF. THE TISSUE REMAINED BETWEEN THE JAW OF THE STAPLER AND CLOSED. AFTER A FEW SECONDS, THE MACHINE PRETENDED TO RESET ITSELF AND TURNED ON AND ITS JAWS DROPPED. IT WAS THEN THAT THE DEVICE WAS TAKEN OUT, CLAMMED DOWN THE BLOOD VESSELS (WITH AN INVENTORIED FORCEPS CALLED CLAMP) UNTIL THEY MOUNTED ANOTHER COMPLETE SET OF DEVICE (HANDLE, ADAPTER, SHELL, LOAD) AND PROCEEDED WITH THE SURGERY. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | No Display/Image | A090206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY, WHEN THE SURGEON WAS GOING TO SECTION THE PULMONARY VESSELS (MEDIASTINAL TRUNK), THE INSTRUMENT HAD BEEN LOADED PERFECTLY AND THE MESSAGES ON THE OLED SCREEN VERIFIED THAT EVERYTHING WAS FINE. HOWEVER, WHEN STAPLING STARTED, AT APPROXIMATELY 1.5 CM IT CRASHED AND THE OLED SCREEN TURNED OFF. THE TISSUE REMAINED BETWEEN THE JAW OF THE STAPLER AND CLOSED. AFTER A FEW SECONDS, THE MACHINE PRETENDED TO RESET ITSELF AND TURNED ON AND ITS JAWS DROPPED. IT WAS THEN THAT THE DEVICE WAS TAKEN OUT, CLAMMED DOWN THE BLOOD VESSELS (WITH AN INVENTORIED FORCEPS CALLED CLAMP) UNTIL THEY MOUNTED ANOTHER COMPLETE SET OF DEVICE (HANDLE, ADAPTER, SHELL, LOAD) AND PROCEEDED WITH THE SURGERY. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY, WHEN THE SURGEON WAS GOING TO SECTION THE PULMONARY VESSELS (MEDIASTINAL TRUNK), THE INSTRUMENT HAD BEEN LOADED PERFECTLY AND THE MESSAGES ON THE OLED SCREEN VERIFIED THAT EVERYTHING WAS FINE. HOWEVER, WHEN STAPLING STARTED, AT APPROXIMATELY 1.5 CM IT CRASHED AND THE OLED SCREEN TURNED OFF. THE TISSUE REMAINED BETWEEN THE JAW OF THE STAPLER AND CLOSED. AFTER A FEW SECONDS, THE MACHINE PRETENDED TO RESET ITSELF AND TURNED ON AND ITS JAWS DROPPED. IT WAS THEN THAT THE DEVICE WAS TAKEN OUT, CLAMMED DOWN THE BLOOD VESSELS (WITH AN INVENTORIED FORCEPS CALLED CLAMP) UNTIL THEY MOUNTED ANOTHER COMPLETE SET OF DEVICE (HANDLE, ADAPTER, SHELL, LOAD) AND PROCEEDED WITH THE SURGERY. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY, WHEN THE SURGEON WAS GOING TO SECTION THE PULMONARY VESSELS (MEDIASTINAL TRUNK), THE INSTRUMENT HAD BEEN LOADED PERFECTLY AND THE MESSAGES ON THE OLED SCREEN VERIFIED THAT EVERYTHING WAS FINE. HOWEVER, WHEN STAPLING STARTED, AT APPROXIMATELY 1.5 CM IT CRASHED AND THE OLED SCREEN TURNED OFF. THE TISSUE REMAINED BETWEEN THE JAW OF THE STAPLER AND CLOSED. AFTER A FEW SECONDS, THE MACHINE PRETENDED TO RESET ITSELF AND TURNED ON AND ITS JAWS DROPPED. IT WAS THEN THAT THE DEVICE WAS TAKEN OUT, CLAMMED DOWN THE BLOOD VESSELS (WITH AN INVENTORIED FORCEPS CALLED CLAMP) UNTIL THEY MOUNTED ANOTHER COMPLETE SET OF DEVICE (HANDLE, ADAPTER, SHELL, LOAD) AND PROCEEDED WITH THE SURGERY. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY, WHEN THE SURGEON WAS GOING TO SECTION THE PULMONARY VESSELS (MEDIASTINAL TRUNK), THE INSTRUMENT HAD BEEN LOADED PERFECTLY AND THE MESSAGES ON THE OLED SCREEN VERIFIED THAT EVERYTHING WAS FINE. HOWEVER, WHEN STAPLING STARTED, AT APPROXIMATELY 1.5 CM IT CRASHED AND THE OLED SCREEN TURNED OFF. THE TISSUE REMAINED BETWEEN THE JAW OF THE STAPLER AND CLOSED. AFTER A FEW SECONDS, THE MACHINE PRETENDED TO RESET ITSELF AND TURNED ON AND ITS JAWS DROPPED. IT WAS THEN THAT THE DEVICE WAS TAKEN OUT, CLAMMED DOWN THE BLOOD VESSELS (WITH AN INVENTORIED FORCEPS CALLED CLAMP) UNTIL THEY MOUNTED ANOTHER COMPLETE SET OF DEVICE (HANDLE, ADAPTER, SHELL, LOAD) AND PROCEEDED WITH THE SURGERY. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY, WHEN THE SURGEON WAS GOING TO SECTION THE PULMONARY VESSELS (MEDIASTINAL TRUNK), THE INSTRUMENT HAD BEEN LOADED PERFECTLY AND THE MESSAGES ON THE OLED SCREEN VERIFIED THAT EVERYTHING WAS FINE. HOWEVER, WHEN STAPLING STARTED, AT APPROXIMATELY 1.5 CM IT CRASHED AND THE OLED SCREEN TURNED OFF. THE TISSUE REMAINED BETWEEN THE JAW OF THE STAPLER AND CLOSED. AFTER A FEW SECONDS, THE MACHINE PRETENDED TO RESET ITSELF AND TURNED ON AND ITS JAWS DROPPED. IT WAS THEN THAT THE DEVICE WAS TAKEN OUT, CLAMMED DOWN THE BLOOD VESSELS (WITH AN INVENTORIED FORCEPS CALLED CLAMP) UNTIL THEY MOUNTED ANOTHER COMPLETE SET OF DEVICE (HANDLE, ADAPTER, SHELL, LOAD) AND PROCEEDED WITH THE SURGERY. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY, WHEN THE SURGEON WAS GOING TO SECTION THE PULMONARY VESSELS (MEDIASTINAL TRUNK), THE INSTRUMENT HAD BEEN LOADED PERFECTLY AND THE MESSAGES ON THE OLED SCREEN VERIFIED THAT EVERYTHING WAS FINE. HOWEVER, WHEN STAPLING STARTED, AT APPROXIMATELY 1.5 CM IT CRASHED AND THE OLED SCREEN TURNED OFF. THE TISSUE REMAINED BETWEEN THE JAW OF THE STAPLER AND CLOSED. AFTER A FEW SECONDS, THE MACHINE PRETENDED TO RESET ITSELF AND TURNED ON AND ITS JAWS DROPPED. IT WAS THEN THAT THE DEVICE WAS TAKEN OUT, CLAMMED DOWN THE BLOOD VESSELS (WITH AN INVENTORIED FORCEPS CALLED CLAMP) UNTIL THEY MOUNTED ANOTHER COMPLETE SET OF DEVICE (HANDLE, ADAPTER, SHELL, LOAD) AND PROCEEDED WITH THE SURGERY. THERE WAS NO PATIENT INJURY.
25-02-2018 | Injury | covidien lp llc north haven | 18-01-2021 | UNKNOWN TA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE, A STUDY COMPARED OUTCOMES FOR PATIENTS WHO UNDERWENT LIVING DONOR OPEN LEFT LATERAL SECTIONECTOMY AND LAPAROSCOPIC LEFT LATERAL SECTIONECTOMY BETWEEN JANUARY 2011 AND APRIL 2017. THERE WERE 220 PATIENTS, 141 PATIENTS UNDERWENT OPEN PROCEDURE, AND 79 UNDERWENT LAPAROSCOPIC PROCEDURE. AN OPEN STAPLER WAS USED TO DIVIDE THE LEFT PORTAL VEIN AND A MANUAL STAPLER WAS USED TO DIVIDE THE LEFT HEPATIC VEIN. POST-OPERATIVE COMPLICATIONS INCLUDE: BLOOD LOSS UP TO 800ML.
22-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | EEA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL RESECTION, THE TISSUE GOT OVERLAPPED AT THE TIME OF ANASTOMOSIS SO ONLY HALF OF THE STAPLES WERE PLACED. IN THE END, THE ANASTOMOSIS SITE WAS CLOSE TO THE SIGMOID COLON, SO THE AREA WHERE THE STAPLES WERE NOT PLACED WAS ANASTOMOSED BY HAND SEW. WHEN PURSE-STRING SUTURE WAS PERFORMED TO THE ANVIL, THE TISSUE EXCESSIVELY PILED UP ON ONE SIDE AND FIRING WAS PERFORMED IN A NON-UNIFORM STATE. IT WAS REPORTED THAT THE PROBLEM WAS NOT CAUSED BY THE DEVICE BUT RATHER THE TECHNICALITY SINCE ANASTOMOSIS PART WAS CLOSE TO THE S-SHAPED COLON. THERE WAS NO PATIENT INJURY.
09-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ABSORBATACK | Detachment of Device or Device Component | A0501 | Device Embedded In Tissue or Plaque | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC UNILATERAL INGUINAL HERNIORRAPHY, WHEN INSERTING THE MESH, THE TACKS FELL INTO THE CAVITY OF THE PATIENT AND WAS NOT RETRIEVED. THE SURGICAL TIME EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. ANOTHER DEVICE FROM A DIFFERENT MANUFACTURER WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
22-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT RECTUM RESECTION, THE STAPLER PINS WAS NOT DEPLOYED PROPERLY. STAPLE LINE WAS INCOMPLETE. THE DEVICE WAS LOCKED ON THE TISSUE BUT WAS ABLE TO BE REMOVED NORMALLY. TISSUE DAMAGE WAS OBSERVED. SURGERY WAS CONVERTED TO OPEN AS A RESULT. STAPLE LINE WAS FIXED BY SUTURING. SURGICAL TIME WAS EXTENDED BY AN HOUR AND A HALF. THE PATIENT HAD TO GO WITH TEMPORARY COLOSTOMY. HOSPITAL STAY WAS EXTENDED.
22-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT RECTUM RESECTION, THE STAPLER PINS WAS NOT DEPLOYED PROPERLY. STAPLE LINE WAS INCOMPLETE. THE DEVICE WAS LOCKED ON THE TISSUE BUT WAS ABLE TO BE REMOVED NORMALLY. TISSUE DAMAGE WAS OBSERVED. SURGERY WAS CONVERTED TO OPEN AS A RESULT. STAPLE LINE WAS FIXED BY SUTURING. SURGICAL TIME WAS EXTENDED BY AN HOUR AND A HALF. THE PATIENT HAD TO GO WITH TEMPORARY COLOSTOMY. HOSPITAL STAY WAS EXTENDED.
22-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT RECTUM RESECTION, THE STAPLER PINS WAS NOT DEPLOYED PROPERLY. STAPLE LINE WAS INCOMPLETE. THE DEVICE WAS LOCKED ON THE TISSUE BUT WAS ABLE TO BE REMOVED NORMALLY. TISSUE DAMAGE WAS OBSERVED. SURGERY WAS CONVERTED TO OPEN AS A RESULT. STAPLE LINE WAS FIXED BY SUTURING. SURGICAL TIME WAS EXTENDED BY AN HOUR AND A HALF. THE PATIENT HAD TO GO WITH TEMPORARY COLOSTOMY. HOSPITAL STAY WAS EXTENDED.
10-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC ENUCLEATION OF HEPATIC METASTASES, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LUNG PARENCHYMA RESECTION PROCEDURE, AFTER RESECTING THE TISSUE, THE JAWS WOULD NOT OPEN. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
15-12-2020 | Injury | celestica de monterrey s.a de c.v | 18-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RADICAL RESECTION OF LUNG CANCER, WHILE DETACHING THE TRACHEA, SOME STAPLES DID NOT FORM. THE STAPLE LINE WAS SUTURED BY HAND. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
22-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE BLACK PUSHER THUMB CAN BE MOVED BUT THE STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED.
11-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, 8 DAYS POSTOPERATIVE TO ROBOTIC LAPAROSCOPIC CYSTECTOMY, CLIPS WERE FOUND TO BE OPENED AND INTERNAL BLEEDING OF MORE THAN 500 CC WAS NOTED. THE PATIENT WAS REOPERATED THREE TIMES. MANUAL SUTURING WAS PERFORMED. THE PATIENT'S HOSPITALIZATION WAS EXTENDED.
22-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE BLACK PUSHER THUMB CAN BE MOVED BUT THE STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED.
11-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, 8 DAYS POSTOPERATIVE TO ROBOTIC LAPAROSCOPIC CYSTECTOMY, PATIENT WAS REOPERATED AND FOUND THAT CLIPS WERE OPENED. BLOOD LOSS OF MORE THAN 500 CC WAS NOTED. MANUAL SUTURING WAS PERFORMED. PATIENT HOSPITALIZATION WAS EXTENDED.
23-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED IN THE VESSEL DURING LAPAROSCOPIC LIVER RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. THE SURGEON REPLACED THE PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED IN THE VESSEL DURING LAPAROSCOPIC LIVER RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. THE SURGEON REPLACED THE PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. NO PATIENT INVOLVEMENT.
21-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, THREE DEVICE RELOADS DID NOT FIRE WHEN THE SURGEON TRIED TO CUT THE STOMACH. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Injury | us surgical puerto rico | 18-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT RECTUM RESECTION, THE STAPLER PINS WAS NOT DEPLOYED PROPERLY. STAPLE LINE WAS INCOMPLETE. THE DEVICE WAS LOCKED ON THE TISSUE BUT WAS ABLE TO BE REMOVED NORMALLY. TISSUE DAMAGE WAS OBSERVED. SURGERY WAS CONVERTED TO OPEN AS A RESULT. STAPLE LINE WAS FIXED BY SUTURING. SURGICAL TIME WAS EXTENDED BY AN HOUR AND A HALF. THE PATIENT HAD TO GO WITH TEMPORARY COLOSTOMY. HOSPITAL STAY WAS EXTENDED.
22-12-2020 | Injury | us surgical puerto rico | 18-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW ANTERIOR RESECTION, AT RECTUM RESECTION, THE STAPLER PINS WAS NOT DEPLOYED PROPERLY. STAPLE LINE WAS INCOMPLETE. THE DEVICE WAS LOCKED ON THE TISSUE BUT WAS ABLE TO BE REMOVED NORMALLY. TISSUE DAMAGE WAS OBSERVED. SURGERY WAS CONVERTED TO OPEN AS A RESULT. STAPLE LINE WAS FIXED BY SUTURING. SURGICAL TIME WAS EXTENDED BY AN HOUR AND A HALF. THE PATIENT HAD TO GO WITH TEMPORARY COLOSTOMY. HOSPITAL STAY WAS EXTENDED.
21-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, THREE DEVICE RELOADS DID NOT FIRE WHEN THE SURGEON TRIED TO CUT THE STOMACH. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, THREE DEVICE RELOADS DID NOT FIRE WHEN THE SURGEON TRIED TO CUT THE STOMACH. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-11-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | TA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A RECTOSIGMOIDECTOMY, THE DEVICE WAS DEFECTIVE AND THE CORRECT CLOSURE OF THE STRAIGHT DID NOT OCCUR. AS A RESULT, THE DEVICE DID NOT STAPLING. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
29-09-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS NOT ABLE TO FIRE.
24-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCTECTOMY, WHEN RESECTING RECTUM, THE DEVICE WOULD NOT FIRE. THE GREEN BUTTON COULD NOT BE PRESSED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. NO PATIENT INVOLVEMENT.
22-12-2020 | Injury | us surgical puerto rico | 18-01-2021 | GIA | Fail-Safe Problem | A1602 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE RELOAD HAD BEEN FIRED BUT THE STAPLER WAS HANDED UP AGAIN TO THE SURGEON PRIOR TO PUTTING IN A FRESH LOAD. THE STAPLER WAS FRIED BUT THERE WAS NO STAPLE LINE, ONLY THE CUT BOWEL. THE WHITE BLADE COVER APPARENTLY DID NOT POP UP AS THE SURGEON WAS ABLE TO CLOSE AND FIRE THE STAPLER AGAIN WITH NO STAPLES. THE BOWEL WAS HAND SEWN.
25-02-2018 | Injury | covidien lp llc north haven | 18-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE, A STUDY COMPARED OUTCOMES FOR PATIENTS WHO UNDERWENT LIVING DONOR OPEN LEFT LATERAL SECTIONECTOMY AND LAPAROSCOPIC LEFT LATERAL SECTIONECTOMY BETWEEN JANUARY 2011 AND APRIL 2017. THERE WERE 220 PATIENTS, 141 PATIENTS UNDERWENT OPEN PROCEDURE, AND 79 UNDERWENT LAPAROSCOPIC PROCEDURE. AN OPEN STAPLER WAS USED TO DIVIDE THE LEFT PORTAL VEIN AND A MANUAL STAPLER WAS USED TO DIVIDE THE LEFT HEPATIC VEIN. POST-OPERATIVE COMPLICATIONS INCLUDE: BLOOD LOSS UP TO 800ML.
22-12-2020 | Injury | us surgical puerto rico | 18-01-2021 | GIA | Fail-Safe Problem | A1602 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE RELOAD HAD BEEN FIRED BUT THE STAPLER WAS HANDED UP AGAIN TO THE SURGEON PRIOR TO PUTTING IN A FRESH LOAD. THE STAPLER WAS FRIED BUT THERE WAS NO STAPLE LINE, ONLY THE CUT BOWEL. THE WHITE BLADE COVER APPARENTLY DID NOT POP UP AS THE SURGEON WAS ABLE TO CLOSE AND FIRE THE STAPLER AGAIN WITH NO STAPLES. THE BOWEL WAS HAND SEWN.
20-09-2020 | Injury | covidien lp llc north haven | 18-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A STUDY PERFORMED FROM 2013 TO 2018 ABOUT GASTRIC LEAKS MANAGEMENT AFTER SLEEVE SURGERY, 6 OUT OF 2000 PATIENT WHERE THE ENDOSCOPIC STAPLER WAS APPLIED, THE FOLLOWING POST-OPERATIVE COMPLICATIONS WERE REPORTED: SEPSIS, GASTRIC LEAKS AND FISTULAS REQUIRING REOPERATIONS, AND STAPLE LINE BLEEDING REQUIRING REOPERATION.
20-09-2020 | Injury | covidien lp llc north haven | 18-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A STUDY PERFORMED FROM 2013 TO 2018 ABOUT GASTRIC LEAKS MANAGEMENT AFTER SLEEVE SURGERY, 6 OUT OF 2000 PATIENT WHERE THE ENDOSCOPIC STAPLER WAS APPLIED, THE FOLLOWING POST-OPERATIVE COMPLICATIONS WERE REPORTED: SEPSIS, GASTRIC LEAKS AND FISTULAS REQUIRING REOPERATIONS, AND STAPLE LINE BLEEDING REQUIRING REOPERATION.
20-09-2020 | Injury | covidien lp llc north haven | 18-01-2021 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A STUDY PERFORMED FROM 2013 TO 2018 ABOUT GASTRIC LEAKS MANAGEMENT AFTER SLEEVE SURGERY, 6 OUT OF 2000 PATIENT WHERE THE ENDOSCOPIC STAPLER WAS APPLIED, THE FOLLOWING POST-OPERATIVE COMPLICATIONS WERE REPORTED: SEPSIS, GASTRIC LEAKS AND FISTULAS REQUIRING REOPERATIONS, AND STAPLE LINE BLEEDING REQUIRING REOPERATION.
22-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY, THE NURSE SET THE CARTRIDGE TO THE HANDLE, AFTER THE USER CHECKED THE OPENING AND CLOSING, THE JAWS WOULD NOT OPEN WHEN THE NURSE HANDED IT OVER TO THE SURGEON IN THE CONDITION THAT THE JAWS WERE CLOSED. THERE WAS A PROBLEM UNLOADING THE DEVICE SO THE CARTRIDGE WAS REMOVED BY FORCE AND WAS REPLACED WITH ANOTHER RELOAD AND WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | us surgical puerto rico | 18-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A COLECTOMY, THE NURSE SET THE CARTRIDGE TO THE HANDLE, AFTER THE USER CHECKED THE OPENING AND CLOSING, THE JAWS WOULD NOT OPEN WHEN THE NURSE HANDED IT OVER TO THE SURGEON IN THE CONDITION THAT THE JAWS WERE CLOSED. THERE WAS A PROBLEM UNLOADING THE DEVICE SO THE CARTRIDGE WAS REMOVED BY FORCE AND WAS REPLACED WITH ANOTHER RELOAD AND WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE LOADING THE FINAL STAPLE RELOAD IT WOULD NOT CLOSE. TRIED TO TAKE THE RELOAD OFF AND PLACE BACK ON BUT IT WOULD NOT CAME OFF OF THE ADAPTER. THE RELOAD WAS STUCK ON THE ADAPTER. ANOTHER ADAPTER AND RELOAD WAS USED ON THE SAME HANDLE AND SHELL TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | covidien lp llc north haven | 18-01-2021 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE LOADING THE FINAL STAPLE RELOAD IT WOULD NOT CLOSE. TRIED TO TAKE THE RELOAD OFF AND PLACE BACK ON BUT IT WOULD NOT CAME OFF OF THE ADAPTER. THE RELOAD WAS STUCK ON THE ADAPTER. ANOTHER ADAPTER AND RELOAD WAS USED ON THE SAME HANDLE AND SHELL TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, A DAY POST OPERATIVE A GASTRIC SLEEVE PROCEDURE, PATIENT HAD TO BE BROUGHT BACK TO SURGERY DUE TO BLEEDING ALONG THE STAPLE LINE. SURGEON PERFORMED A DIAG LAP AND HAD TO GIVE TWO UNITS BLOOD, CLIP AND SUTURE THE STAPLE LINE TO CONTROL AS WELL AS EVACUATE THE BLOOD.
21-12-2020 | Injury | covidien lp llc north haven | 18-01-2021 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, A DAY POST OPERATIVE A GASTRIC SLEEVE PROCEDURE, PATIENT HAD TO BE BROUGHT BACK TO SURGERY DUE TO BLEEDING ALONG THE STAPLE LINE. SURGEON PERFORMED A DIAG LAP AND HAD TO GIVE TWO UNITS BLOOD, CLIP AND SUTURE THE STAPLE LINE TO CONTROL AS WELL AS EVACUATE THE BLOOD.
28-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLONECTOMY, ON DETACHING COLON, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. NEW HANDLE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
11-03-2021 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD DID NOT CLOSE FIRMLY FOR STAPLING THE TISSUE.
19-01-2021 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD DID NOT CLOSE FIRMLY FOR STAPLING THE TISSUE.
19-01-2021 | Malfunction | covidien lp llc north haven | 19-01-2021 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD DID NOT CLOSE FIRMLY FOR STAPLING THE TISSUE.
01-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, WHILE CHARGING, THE SCREEN REMAINED BLACK BECAUSE THE PRODUCT CANNOT BE CHARGED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERIES WERE VENTED.
01-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, WHILE CHARGING, THE SCREEN REMAINED BLACK BECAUSE THE PRODUCT CANNOT BE CHARGED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERIES WERE VENTED.
06-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WAS NOT ABLE TO OPERATE. THE FACILITY HAD TWO POWER HANDLES. THE FIRST HANDLE WAS NOT CHARGED AND THE BATTERY CHARGER PINS WERE NOTICED TO BE MELTED AND THERE WAS A LEAKAGE FROM THE POWER HANDLE. THE SECOND POWER HANDLE WAS OPENED, BUT THERE WAS ALSO A PROBLEM WITH IT, SO ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 -MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
06-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WAS NOT ABLE TO OPERATE. THE FACILITY HAD TWO POWER HANDLES. THE FIRST HANDLE WAS NOT CHARGED AND THE BATTERY CHARGER PINS WERE NOTICED TO BE MELTED AND THERE WAS A LEAKAGE FROM THE POWER HANDLE. THE SECOND POWER HANDLE WAS OPENED, BUT THERE WAS ALSO A PROBLEM WITH IT, SO ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 -MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
06-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WAS NOT ABLE TO OPERATE. THE FACILITY HAD TWO POWER HANDLES. THE FIRST HANDLE WAS NOT CHARGED AND THE BATTERY CHARGER PINS WERE NOTICED TO BE MELTED AND THERE WAS A LEAKAGE FROM THE POWER HANDLE. THE SECOND POWER HANDLE WAS OPENED, BUT THERE WAS ALSO A PROBLEM WITH IT, SO ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 -MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
29-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY PROCEDURE, THE DEVICE DID NOT FIRE ACROSS THE TISSUE. A MANUAL HANDLE WAS USED TOGETHER WITH THE SAME RELOAD TO RESOLVE THE ISSUE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
16-09-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | UNKNOWN GIA PRODUCT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL SURGERY FOR ESOPHAGEAL CANCER PROCEDURE, THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
19-08-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LEFT LOWER LUNG RESECTION PROCEDURE, THE DEVICE DID NOT FIRE FOR EXCISION IN THE MIDDLE OF FIRING, RESULTING IN INCOMPLETE EXCISION OF THE PULMONARY SUBSTANCE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHEN TRYING TO ACTIVATE IT, IT DID NOT WORK TO STAPLE THE CECUM. THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE COULD NOT BE SQUEEZED; HENCE, IT DID NOT FIRE. IT SEEMED AS IF IT WAS PREDISPOSED. IT WAS ALSO REPORTED THAT THERE WAS AN ISSUE UNLOADING THE RELOAD. ANOTHER RELOAD OF THE SAME REFERENCE HAD TO BE USED WHICH WORKED NORMALLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHEN TRYING TO ACTIVATE IT, IT DID NOT WORK TO STAPLE THE CECUM. THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE COULD NOT BE SQUEEZED; HENCE, IT DID NOT FIRE. IT SEEMED AS IF IT WAS PREDISPOSED. IT WAS ALSO REPORTED THAT THERE WAS AN ISSUE UNLOADING THE RELOAD. ANOTHER RELOAD OF THE SAME REFERENCE HAD TO BE USED WHICH WORKED NORMALLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHEN TRYING TO ACTIVATE IT, IT DID NOT WORK TO STAPLE THE CECUM. THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE COULD NOT BE SQUEEZED; HENCE, IT DID NOT FIRE. IT SEEMED AS IF IT WAS PREDISPOSED. IT WAS ALSO REPORTED THAT THERE WAS AN ISSUE UNLOADING THE RELOAD. ANOTHER RELOAD OF THE SAME REFERENCE HAD TO BE USED WHICH WORKED NORMALLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A WHIPPLE PROCEDURE, DURING PREPARATION, THE HANDLE DID NOT ACTIVATE AND WAS NOT USED. A MANUAL HANDLE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
29-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A WHIPPLE PROCEDURE, DURING PREPARATION, THE HANDLE DID NOT ACTIVATE AND WAS NOT USED. A MANUAL HANDLE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
29-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A WHIPPLE PROCEDURE, DURING PREPARATION, THE HANDLE DID NOT ACTIVATE AND WAS NOT USED. A MANUAL HANDLE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
14-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHEN TRYING TO ACTIVATE IT, IT DID NOT WORK TO STAPLE THE CECUM. THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE COULD NOT BE SQUEEZED; HENCE, IT DID NOT FIRE. IT SEEMED AS IF IT WAS PREDISPOSED. IT WAS ALSO REPORTED THAT THERE WAS AN ISSUE UNLOADING THE RELOAD. ANOTHER RELOAD OF THE SAME REFERENCE HAD TO BE USED WHICH WORKED NORMALLY TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-09-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT LOWER LOBECTOMY, WHILE BEING APPLIED TO A SEVERED LUNG LOBE, THE GREEN BUTTON WAS PRESSED, BUT THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY SEGMENTECTOMY, WHILE BEING APPLIED TO THE LUNG, THE HANDLE DID NOT TRIGGER THE RELOAD OR CLAMP. THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE WAS NOT SQUEEZED; HENCE, THE DEVICE DID NOT FIRE. THE PRODUCT WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY SEGMENTECTOMY, WHILE BEING APPLIED TO THE LUNG, THE HANDLE DID NOT TRIGGER THE RELOAD OR CLAMP. THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE WAS NOT SQUEEZED; HENCE, THE DEVICE DID NOT FIRE. THE PRODUCT WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, BEFORE STARTING THE SURGERY, HANDLE FAILED TO CHARGE SO THE INSTRUMENT WAS NOT USED IN PATIENT. A DISPOSABLE MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INSTRUMENT NOTED THAT TOWARDS THE END OF THE LOGS, THE HANDLE WAS PLACED ON A CHARGER. THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
26-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, BEFORE STARTING THE SURGERY, HANDLE FAILED TO CHARGE SO THE INSTRUMENT WAS NOT USED IN PATIENT. A DISPOSABLE MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INSTRUMENT NOTED THAT TOWARDS THE END OF THE LOGS, THE HANDLE WAS PLACED ON A CHARGER. THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
26-10-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, BEFORE STARTING THE SURGERY, HANDLE FAILED TO CHARGE SO THE INSTRUMENT WAS NOT USED IN PATIENT. A DISPOSABLE MANUAL STAPLER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INSTRUMENT NOTED THAT TOWARDS THE END OF THE LOGS, THE HANDLE WAS PLACED ON A CHARGER. THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
28-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RADICAL RESECTION OF LEFT LOWER LOBE, WHEN LOBE WAS REMOVED, THE FIRST RELOAD WAS INSTALLED AND READY TO BE FIRED, BUT IT COULD NOT BE FIRED. ANOTHER RELOAD WAS OPENED, THE ACTIVATION PROBLEM DID NOT OCCURRED. AFTER THAT, TWO OTHER RELOADS WERE USED AND THE STAPLES POORLY SHAPED. A PURPLE AND GREEN ROTICULATOR RELOADS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
28-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RADICAL RESECTION OF LEFT LOWER LOBE, WHEN LOBE WAS REMOVED, THE FIRST RELOAD WAS INSTALLED AND READY TO BE FIRED, BUT IT COULD NOT BE FIRED. ANOTHER RELOAD WAS OPENED, THE ACTIVATION PROBLEM DID NOT OCCURRED. AFTER THAT, TWO OTHER RELOADS WERE USED AND THE STAPLES POORLY SHAPED. A PURPLE AND GREEN ROTICULATOR RELOADS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
28-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DETACHING A LUNG LOBE IN A LOBECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. SURGEON REPLACED THE HANDLE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
28-12-2020 | Malfunction | celestica de monterrey s.a de c.v | 19-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC WHIPPLE, UPON DUODENUM, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE BUT BOTH RELOADS WOULD NOT FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2020 | Injury | us surgical puerto rico | 19-01-2021 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, STAPLES WERE MALFORMED AND ANASTOMOSIS LEAKED. SUTURE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
24-12-2020 | Injury | us surgical puerto rico | 19-01-2021 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, STAPLES WERE MALFORMED AND ANASTOMOSIS LEAKED. SUTURE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
24-12-2020 | Injury | us surgical puerto rico | 19-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, STAPLES WERE MALFORMED AND ANASTOMOSIS LEAKED. SUTURE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
24-12-2020 | Injury | us surgical puerto rico | 19-01-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, STAPLES WERE MALFORMED AND ANASTOMOSIS LEAKED. SUTURE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
01-11-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE; THE PRODUCT OPENED, THE ARTICULATING JAW UNABLE TO HOLD A STABLE LOAD. REMOVED FROM THE STERILE FIELD. NO INJURY/NOT USED ON THE PATIENT.
28-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A THORASCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, WHILE CUTTING AND CLOSING THE LUNG PARENCHYMA, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD AND THE SAME HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-10-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | GST BLACK RELOAD, 60MM, 6 ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THREE OF THE BLACK ECHELON RELOADS REF GST60T WERE USED TODAY DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), LEFT UPPER LOBE WEDGE RESECTION, AND BROKE UPON REMOVAL FROM THE STAPLE JAW. THE FIRST RELOAD WAS USED TO TEST FIRE ON A TOWEL AFTER SWITCHING TO A GREEN RELOAD WHICH WAS INTACT AFTER USE. THE RESECTION WAS COMPLETED USING TWO GREEN RELOADS WITH NO ISSUES. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
27-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON LAPAROSCOPIC COLORECTOMY, WHILE BEING APPLIED IN THE RECTAL TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD AND THE SAME HANDLE WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | PROXIMATE*55MM LINEAR CUTTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE THE STAPLER WAS USED A COUPLE OF TIMES SUCCESSFULLY AND THE LAST LOAD THEY TRIED TO FIRE DID NOT EXPEL ANY STAPLES. BUT THE DOCTOR SAID THAT IT DID CUT. THERE WERE NO POOR OUTCOMES DUE TO THIS. THE SURGEON WAS ABLE TO FINISH THE SURGERY. SALES REP WAS NOT IN THE CASE. NO PATIENT CONSEQUENCES WERE REPORTED.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN RIGHT HEMICOLECTOMY, THE FIRING ADVANCED A LITTLE BUT THE HANDLE BECAME UNABLE TO BE SQUEEZED, SO IT WAS STOPPED TO BE USED. IT WAS ALSO STATED THAT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN RIGHT HEMICOLECTOMY, THE FIRING ADVANCED A LITTLE BUT THE HANDLE BECAME UNABLE TO BE SQUEEZED, SO IT WAS STOPPED TO BE USED. IT WAS ALSO STATED THAT THE PROXIMAL STAPLE LINE WAS INCOMPLETE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-01-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | GREEN RELOAD FOR ECHELON 60 | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKOWN PROCEDURE, THE METAL SIDE OF THE CARTRIDGE HAS BEEN DEFORMED AT PROXIMAL. THERE WERE NO PATIENT CONSEQUENCES.
28-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLES DID NOT GET INSERTED AFTER FIRED.
21-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | PWRD ECH FLEX 45MM 34CM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LINGUAL SEGMENT RESECTION OF LEFT UPPER LOBE SURGERY,AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN THE KNIFE AUTOMATICALLY. THE KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN TOTAL GASTRECTOMY, THE JAWS DID NOT OPEN WHEN THE DEVICE WAS CHECKED IF IT WORKS PROPERLY. THE JAWS WERE ATTEMPTED TO OPEN SEVERAL TIMES, THEN THEY OPENED. HOWEVER, AT THE 3RD FIRING, AN UNEXPECTED SOUND WAS HEARD, AND THE JAWS BECAME NOT TO OPEN WHEN THE KNIFE WAS ON THE WAY BACK TO THE HOME POSITION. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
02-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | 55 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, SOME STAPLES FROM THE SR55 CAME OUT (FAIL TO FIRE PROPERLY).
24-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC HEPATECTOMY SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. AS THE TISSUE WAS CUT OFF, REMOVED THE DEVICE FROM THE PATIENT THROUGH A TISSUE GAP. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
26-11-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN UNKNOWN PROCEDURE, THE FIRING KNOB DID NOT MOVE BEFORE USE. ANOTHER DEVICE AND CARTRIDGE WERE USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGUS SURGERY, THE SURGEON FIRED THE STAPLE ON TOP OF THE STAPLE WHICH HE PREVIOUSLY HAS FIRED BEFORE, (WHICH WE CALL AS DOUBLE STAPLING) AND HE CLAIMS THIS AS HIS USUAL SURGICAL METHOD. THIS TIME, THE TISSUE HAS BEEN PUSHED WITH THE PREVIOUS STAPLES AND THE NEXT STAPLE FIRED DID NOT FULLY INSERTED ON THE EXACT TARGETED TISSUE. FORTUNATELY, THE SURGERY HAS BEEN FINISHED SAFELY, BUT THE SURGEON WANTS A CLEAR FOLLOW-UP FROM OUR COMPANY OF THE USED PRODUCT'S ISSUE OF WHY THIS HAS OCCURRED AND POSSIBLY OUR GUIDELINE WITH THE ISSUE OF DOUBLE STAPLING. IT IS UNKNOWN IF THERE WERE ANY PATIENT CONSEQUENCES.
21-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | GST WHITE RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING VESSELS, AFTER FIRED WITH GST45W, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY.THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | ILS 33MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, DOCTOR WAS USING CDH33A IN THE PROCEDURE, LAR. TACTILE FEEDBACK WAS NOT FOUND & THE CIRCULAR STAPLER MISFIRE. SURGEON WAS COMPLETED THE OPERATION BY HAND SEWING METHOD. NO DELAY. PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
23-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY LUNG RESECTION PROCEDURE, THE DEVICE ABLE TO BE FIRED COMPLETELY AND REMOVED FROM THE CAVITY AS USUAL, BUT IT WAS FOUND OUT THAT A DEVICE COMPONENT FELL OFF, IT WAS REPORTED THAT ENCLOSED SMALL DEBRIS FELL INTO THE PATIENT CAVITY. TO RESOLVE THE ISSUE THE DEVICE DEBRIS WAS REMOVED FROM THE CAVITY. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY LUNG RESECTION PROCEDURE, THE DEVICE ABLE TO BE FIRED COMPLETELY AND REMOVED FROM THE CAVITY AS USUAL, BUT IT WAS FOUND OUT THAT A DEVICE COMPONENT FELL OFF, IT WAS REPORTED THAT ENCLOSED SMALL DEBRIS FELL INTO THE PATIENT CAVITY. TO RESOLVE THE ISSUE THE DEVICE DEBRIS WAS REMOVED FROM THE CAVITY. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY LUNG RESECTION PROCEDURE, THE DEVICE WAS ABLE TO BE FIRED COMPLETELY AND REMOVED FROM THE CAVITY AS USUAL, BUT IT WAS FOUND OUT THAT A DEVICE COMPONENT FROM THE RELOAD FELL OFF. IT WAS REPORTED THAT ENCLOSED SMALL DEBRIS FELL INTO THE PATIENT'S CAVITY. TO RESOLVE THE ISSUE, THE DEVICE DEBRIS WAS REMOVED FROM THE CAVITY. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-20 AFC: EGIA HANDLE THICK TISSUE. ALWAYS INSPECT THE TISSUE THICKNESS AND SELECT AN APPROPRIATE STAPLE SIZE PRIOR TO APPLICATION OF THE ENDO GIA¿ ULTRA UNIVERSAL SHORT, ENDO GIA¿ ULTRA UNIVERSAL OR ENDO GIA¿ ULTRA UNIVERSAL XL STAPLER. OVERLY THICK OR THIN TISSUE MAY RESULT IN UNACCEPTABLE STAPLE FORMATION.
23-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY LUNG RESECTION PROCEDURE, THE DEVICE WAS ABLE TO BE FIRED COMPLETELY AND REMOVED FROM THE CAVITY AS USUAL, BUT IT WAS FOUND OUT THAT A DEVICE COMPONENT FROM THE RELOAD FELL OFF. IT WAS REPORTED THAT ENCLOSED SMALL DEBRIS FELL INTO THE PATIENT'S CAVITY. TO RESOLVE THE ISSUE, THE DEVICE DEBRIS WAS REMOVED FROM THE CAVITY. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-20 AFC: EGIA HANDLE THICK TISSUE. ALWAYS INSPECT THE TISSUE THICKNESS AND SELECT AN APPROPRIATE STAPLE SIZE PRIOR TO APPLICATION OF THE ENDO GIA¿ ULTRA UNIVERSAL SHORT, ENDO GIA¿ ULTRA UNIVERSAL OR ENDO GIA¿ ULTRA UNIVERSAL XL STAPLER. OVERLY THICK OR THIN TISSUE MAY RESULT IN UNACCEPTABLE STAPLE FORMATION.
23-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY LUNG RESECTION PROCEDURE, THE DEVICE WAS ABLE TO BE FIRED COMPLETELY AND REMOVED FROM THE CAVITY AS USUAL, BUT IT WAS FOUND OUT THAT A DEVICE COMPONENT FROM THE RELOAD FELL OFF. IT WAS REPORTED THAT ENCLOSED SMALL DEBRIS FELL INTO THE PATIENT'S CAVITY. TO RESOLVE THE ISSUE, THE DEVICE DEBRIS WAS REMOVED FROM THE CAVITY. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-20 AFC: EGIA HANDLE THICK TISSUE. ALWAYS INSPECT THE TISSUE THICKNESS AND SELECT AN APPROPRIATE STAPLE SIZE PRIOR TO APPLICATION OF THE ENDO GIA¿ ULTRA UNIVERSAL SHORT, ENDO GIA¿ ULTRA UNIVERSAL OR ENDO GIA¿ ULTRA UNIVERSAL XL STAPLER. OVERLY THICK OR THIN TISSUE MAY RESULT IN UNACCEPTABLE STAPLE FORMATION.
23-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY LUNG RESECTION PROCEDURE, THE DEVICE WAS ABLE TO BE FIRED COMPLETELY AND REMOVED FROM THE CAVITY AS USUAL, BUT IT WAS FOUND OUT THAT A DEVICE COMPONENT FROM THE RELOAD FELL OFF. IT WAS REPORTED THAT ENCLOSED SMALL DEBRIS FELL INTO THE PATIENT'S CAVITY. TO RESOLVE THE ISSUE, THE DEVICE DEBRIS WAS REMOVED FROM THE CAVITY. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND PLI-20 AFC: EGIA HANDLE THICK TISSUE. ALWAYS INSPECT THE TISSUE THICKNESS AND SELECT AN APPROPRIATE STAPLE SIZE PRIOR TO APPLICATION OF THE ENDO GIA¿ ULTRA UNIVERSAL SHORT, ENDO GIA¿ ULTRA UNIVERSAL OR ENDO GIA¿ ULTRA UNIVERSAL XL STAPLER. OVERLY THICK OR THIN TISSUE MAY RESULT IN UNACCEPTABLE STAPLE FORMATION.
16-11-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE'S FIRING TRIGGER IS NOT POSSIBLE TO CLOSE FOR FIRING. THE INSTRUMENT WAS NOT USABLE. NO PATIENT CONSEQUENCE.
01-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | GST BLACK RELOAD, 60MM, 6 ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE GST60T COULD NOT BE INSERTED INTO THE STAPLER BODY. THE CONNECTING PART WAS CRUSHED. SO, A NEW GST60T WAS REPLACED AND WORKED FINE. THERE WERE NO PATIENT CONSEQUENCES.
01-01-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE, THE STAPLE LINE PRODUCED FROM THIS STAPLER WAS CAUSING EXCESSIVE BLEEDING AND EVERY RELOAD FIRED WITH THIS STAPLER PRODUCED A SPURTING VEIN CAUSING A POOL OF BLOOD TO ACCUMULATE BELOW THE STOMACH. SUTURES WERE TAKEN ALONG THE STAPLE LINE, PRESSURE APPLIED AGAINST IT AND SURGICEL WAS USED TO TONE DOWN BLEEDING. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THE PATIENT HASN'T EXPERIENCED UNEXPECTED OUTCOMES YET,
24-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, WHILE USING THE ECHELON FLEX THROUGHOUT THE SURGERY, WHEN TRIED TO USE IT ON THE SMALL BOWEL, THE STAPLER WOULD NOT FIRE. TRIED TO USE THE MANUAL OVERRIDE FEATURE, WHICH ALSO DID NOT WORK. IT IS UNKNOWN HOW PROCEDURE WAS COMPLETED. PATIENT CONSEQUENCE WAS NOT REPORTED.
01-01-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT IN THREE SEPARATE INSTANCES THE PHYSICIAN STATED THE DEVICE DID NOT FORM A COMPLETE ANASTOMOSIS AND HE HAD TO OVERSEW TO COMPLETE THE CASE. THE PATIENT IS OKAY, FROM THE INFORMATION I COULD GATHER. THE PATIENT DOES NOT HAVE ANY CONSEQUENCES RELATED TO OUR PRODUCT. NO OTHER INFORMATION IS AVAILABLE.
01-04-2021 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPICALLY ASSISTED HIGH ANTERIOR RESECTION OF RECTUM, PRIOR TO PATIENT USE, THE OPENING AND CLOSING OF THE RELOAD WAS CHECKED BEFORE SURGERY, BUT THE RELOAD COULD NOT BE OPENED AND CLOSED AT ALL. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. SURGEON REPLACED THE RELOAD AND SHELL TO RESOLVE THE ISSUE. NO PATIENT INJURY.
01-04-2021 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPICALLY ASSISTED HIGH ANTERIOR RESECTION OF RECTUM, PRIOR TO PATIENT USE, THE OPENING AND CLOSING OF THE RELOAD WAS CHECKED BEFORE SURGERY, BUT THE RELOAD COULD NOT BE OPENED AND CLOSED AT ALL. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY, AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY, AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Separation Failure | A150301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY, AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY, AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER CLAMPING THE JEJUNUM DURING AN OPEN REMOVAL OF JEJUNAL MASS, THE DEVICE COULD NO BE FIRED. A NEW DEVICE WAS USED AND THE OPERATION WAS COMPLETED. THERE WAS NO PATIENT INJURY.
20-12-2020 | Injury | ethicon endo-surgery | 19-01-2021 | ECHELON FLEX 45 STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE PATIENT WAS REOPERATED DUE TO A LEAKING FATTY JEJUNAL ANASTOMOSIS, WHO INITIALLY UNDERWENT BYPASS SURGERY ON (B)(6). AT REOPERATION, THE SURGEON REPAIRED THE LEAK USING MANUAL SUTURE.
25-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY , AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY , AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
15-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING ENDOSCOPIC LOBECTOMY, SOME STAPLES WERE MALFORMED WITH IRREGULAR SHAPE AFTER FIRED. IT IS UNKNOWN HOW PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING USED FOR JEJUNUM RESECTION DURING A LAPAROSCOPIC PANCREATICODUODENECTOMY, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE WAS SQUEEZED. ANOTHER PRODUCT WAS USED. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY, AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY, AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Separation Failure | A150301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY, AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
25-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, USED FOR LUNG PARENCHYMAL RESECTION, THE RELOAD WAS LOADED TO THE HANDLE AND USED THEN THE HANDLE BECAME STIFF HALFWAY, AND IT BECAME UNABLE TO BE SQUEEZED. THE MOST DISTAL SIDE OF THE STAPLES WERE FOUND TO BE U-SHAPED. THE JAWS WERE ALSO LOCKED ON TISSUE. AS THE FIRING BECAME UNABLE TO BE PERFOMED HALFWAY, THE REINFORCED SHEET WAS CUT AND SEPARATED WITH SCISSORS AND REMOVED FROM THE TISSUE. THE SAME EVENT HAPPENED WHEN REPLACE THE RELOAD. AFTER THAT, THE TISSUE WAS SEPARATED BY ANOTHER HANDLE AND RELOAD, FROM THE PART WHICH HAD BEEN STAPLED IN THE HALFWAY. THERE WAS NO PATIENT INJURY.
26-12-2020 | Injury | covidien lp llc north haven | 19-01-2021 | EEA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN HEMORRHOID SURGERY, DURING TISSUE REMOVAL, THE STAPLE LINE WAS INCOMPLETE. MORE THAN HALF OF THE STAPLES WERE NOT FIRED. THERE WERE ALSO POORLY FORMED STAPLES. THE DONUTS WERE INCOMPLETE. THE STAPLE LINE WAS OVER SEWED.
26-12-2020 | Injury | covidien lp llc north haven | 19-01-2021 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN HEMORRHOID SURGERY, DURING TISSUE REMOVAL, THE STAPLE LINE WAS INCOMPLETE. MORE THAN HALF OF THE STAPLES WERE NOT FIRED. THERE WERE ALSO POORLY FORMED STAPLES. THE DONUTS WERE INCOMPLETE. THE STAPLE LINE WAS OVER SEWED.
21-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | GST WHITE RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING VESSELS, AFTER FIRED WITH GST45W, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY.THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING USED FOR JEJUNUM RESECTION DURING A LAPAROSCOPIC PANCREATICODUODENECTOMY, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE WAS SQUEEZED. ANOTHER PRODUCT WAS USED. THERE WAS NO PATIENT INJURY.
24-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, IMMEDIATELY AFTER THE SETUP WAS COMPLETED AND THE CARTRIDGE WAS CONNECTED, THE DISPLAY ON THE HANDLE BLACKED OUT AND STOPPED WORKING. THE ADAPTER WAS UNSTABLE, SO IT WAS PULLED OUT AND REMOVED. THE CARTRIDGE WAS REMOVED AND THE ADAPTER WAS RECONNECTED TO THE HANDLE, BUT IT STILL DID NOT WORK. THE HANDLE WAS HOT AND WAS HOT EVEN WHEN IT WAS TOUCHED OVER THE SHELL. THE PROCEDURE WAS PERFORMED USING ANOTHER MANUAL HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
10-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC EMERGENCY APPENDECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. WHEN IT WAS ATTEMPTED TO BE FIRED, THERE WAS A CRACKING SOUND LIKE THE GEAR WAS CHIPPED AND COULD NOT BE FIRED.THE HANDLE WAS REPLACED AND THE SAME RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE IT WAS ALSO REPORTED THAT THE TISSUE WAS HIGHLY INFLAMED. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC EMERGENCY APPENDECTOMY PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE. WHEN IT WAS ATTEMPTED TO BE FIRED, THERE WAS A CRACKING SOUND LIKE THE GEAR WAS CHIPPED AND COULD NOT BE FIRED.THE HANDLE WAS REPLACED AND THE SAME RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE IT WAS ALSO REPORTED THAT THE TISSUE WAS HIGHLY INFLAMED. THERE WAS NO PATIENT INJURY.
22-10-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | ETS*45 ENDO LIN CUT RELOAD WH | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC DISTAL GASTRECTOMY SURGERY, COULD NOT FIRE WHEN SEVERING JEJUNUM TISSUE . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, IMMEDIATELY AFTER THE SETUP WAS COMPLETED AND THE CARTRIDGE WAS CONNECTED, THE DISPLAY ON THE HANDLE BLACKED OUT AND STOPPED WORKING. THE ADAPTER WAS UNSTABLE, SO IT WAS PULLED OUT AND REMOVED. THE CARTRIDGE WAS REMOVED AND THE ADAPTER WAS RECONNECTED TO THE HANDLE, BUT IT STILL DID NOT WORK. THE HANDLE WAS HOT AND WAS HOT EVEN WHEN IT WAS TOUCHED OVER THE SHELL. THE PROCEDURE WAS PERFORMED USING ANOTHER MANUAL HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
23-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-11-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE FIRST FIRING WAS FINE, SECOND AND THIRD FIRINGS WERE MISFIRES. PHYSICIAN SAID WHEN FIRED IT MADE A HORRIBLE SOUND AND WOULDN'T STAPLE OR CUT. THEY USED A NEW DEVICE TO COMPLETE THE CASE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT CONSEQUENCES.
23-11-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE FIRST FIRING WAS FINE, SECOND AND THIRD FIRINGS WERE MISFIRES. PHYSICIAN SAID WHEN FIRED IT MADE A HORRIBLE SOUND AND WOULDN'T STAPLE OR CUT. THEY USED A NEW DEVICE TO COMPLETE THE CASE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT CONSEQUENCES.
23-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION PROCEDURE, AFTER ARTICULATING THE RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE A LITTLE BIT, BUT THE HANDLE WAS STIFF AND COULD NOT BE SQUEEZED FURTHER. THE DEVICE DID NOT FIRE. THE SURGEON THEN HEARD A CRACK SOUND FROM THE INTERNAL GEAR LIKE IT WAS BROKEN. A COMPETITIVE'S DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2020 | Malfunction | us surgical puerto rico | 19-01-2021 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION PROCEDURE, AFTER ARTICULATING THE RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE A LITTLE BIT, BUT THE HANDLE WAS STIFF AND COULD NOT BE SQUEEZED FURTHER. THE DEVICE DID NOT FIRE. THE SURGEON THEN HEARD A CRACK SOUND FROM THE INTERNAL GEAR LIKE IT WAS BROKEN. A COMPETITIVE'S DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2020 | Malfunction | covidien lp llc north haven | 19-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION PROCEDURE, AFTER ARTICULATING THE RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE A LITTLE BIT, BUT THE HANDLE WAS STIFF AND COULD NOT BE SQUEEZED FURTHER. THE DEVICE DID NOT FIRE. THE SURGEON THEN HEARD A CRACK SOUND FROM THE INTERNAL GEAR LIKE IT WAS BROKEN. A COMPETITIVE'S DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
22-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE DID NOT RETURN TO HOME POSITION COMPLETELY DURING USE. THE KNIFE REVERSE SWITCH WAS USED, BUT THE KNIFE DID NOT RETURN, SO THE MANUAL OVERRIDE LEVER WAS USED TO RETURN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
28-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AFTER THE FIFTH FIRING, THE JAWS WOULD NOT OPEN. THE REVERSE BUTTON WAS TRIED AND THERE WAS NO INDICATION THAT THE BLADE WAS RETURNING. THE MANUAL OVERRIDE WAS TRIED AND AGAIN IT DID NOT APPEAR THE BLADE WAS MOVING. THE SURGEON WORKED WITH THE DEVICE AND EVENTUALLY WAS ABLE TO OPEN THE JAWS AND REMOVE THE DEVICE. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE WITH NO PATIENT CONSEQUENCES.
27-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | GST GOLD RELOAD, 60MM, 6 ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS FOUND THAT A PART OF THE CARTRIDGE HAD BEEN BROKEN OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
29-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AND DURING THE FIRING SEQUENCE, THE KNIFE FINISHED MARCHING, BUT IT DIDN'T COME BACK TO THE START LINE AND DIDN'T MOVE AT THE TIP OF THE JAW. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. PATIENT CONSEQUENCE WAS NOT REPORTED.
27-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | PWRD ECH FLEX 45MM 34CM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RADICAL RESECTION OF LUNG CANCER SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-12-2020 | Malfunction | ethicon endo-surgery | 19-01-2021 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC GASTROPLASTY SURGERY, WHEN SEVERING STOMACH TISSUE ON THE FOURTH FIRING, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH STRAIGHT LEG. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-10-2020 | Malfunction | ethicon | 19-01-2021 | ECHELON CIRCULAR POWERED STAPLER | Failure to Fire | A050501 | Insufficient Information | Event Description: DURING LAP ASSISTED CLOSURE OF COLOSTOMY/MOBILIZATION OF SPLENIC FLEXOR/SMALL BOWEL RESECTION/LYSIS OF ADHESIONS AND PRIMARY REPAIR OF INCISIONAL HERNIA, ECHELON CIRCULAR POWERED STAPLER DID NOT FIRE. THE DEVICE SHOWED AN INDICATOR LIGHT THAT IT WAS READY TO USE BUT WAS APPARENTLY IN A USED TYPE POSITION, SO IT DID NOT WORK AS INTENDED. NO PATIENT INJURY. REF 3005075853-2020-06295.
27-11-2020 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2021 | CAPSURE STRAIGHT | Failure to Fire | A050501 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT A BARD/DAVOL CAPSURE STRAIGHT FIXATION DEVICE WAS USED TO FIXATE AN UNSPECIFIED DUAL MESH DURING AN INTRAPERITONEAL ONLAY MESH (IPOM) TECHNIQUE ON (B)(6) 2020. WITH ADEQUATE COUNTER-PRESSURE, THE SURGEON USED 10 FASTENERS TO FIXATE THE MESH; NO CRACKING SOUND WAS HEARD, AFTER WHICH THE FASTENER WAS NOT COMPLETELY FIRED. THE PROCEDURE WAS COMPLETED USING SUTURES DUE TO WHICH THERE WAS TIME CONSUMPTION AND A DELAY IN PROCEDURE. THERE WERE NO BROKEN OR LOOSE FASTENERS IN THE BODY AND NO REPORTED PATIENT INJURY. AS REPORTED, THE SURGEON WAS NOT DEPLOYING FASTENERS INTO OR NEAR THE COOPER'S LIGAMENT.
27-11-2020 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2021 | CAPSURE STRAIGHT | Mechanical Jam | A0506 | No Consequences Or Impact To Patient | Event Description: IT WAS REPORTED THAT A BARD/DAVOL CAPSURE STRAIGHT FIXATION DEVICE WAS USED TO FIXATE AN UNSPECIFIED DUAL MESH DURING AN INTRAPERITONEAL ONLAY MESH (IPOM) TECHNIQUE ON (B)(6) 2020. WITH ADEQUATE COUNTER-PRESSURE, THE SURGEON USED 10 FASTENERS TO FIXATE THE MESH; NO CRACKING SOUND WAS HEARD, AFTER WHICH THE FASTENER WAS NOT COMPLETELY FIRED. THE PROCEDURE WAS COMPLETED USING SUTURES DUE TO WHICH THERE WAS TIME CONSUMPTION AND A DELAY IN PROCEDURE. THERE WERE NO BROKEN OR LOOSE FASTENERS IN THE BODY AND NO REPORTED PATIENT INJURY. AS REPORTED, THE SURGEON WAS NOT DEPLOYING FASTENERS INTO OR NEAR THE COOPER'S LIGAMENT.
17-11-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | BLUE RELOAD FOR ECHELON 60 | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RIGHT UPPER LOBECTOMY SURGERY, OPEN THE PACKING AND NOTED THAT SOME SLED FELL OFF. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-10-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, THE STAPLES COULDN'T BE LOADED NORMALLY, THE STAPLES WERE NOT STRAIGHT AFTER LOADING, AND THERE WAS AN ANGLE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-11-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE RELOAD DURING A PANCREATIC HEAD RESECTION, ON THE TISSUE, THE SURGEON WAS STRUGGLING WITH THE REMOVAL / OPENING OF THE STAPLER. THE RELOAD COULD NOT BE RELEASED. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND RELEASE THE HANDLE. THERE WAS A DELAY IN THE INTERVENTION. THERE WAS NO PATIENT INJURY.
09-11-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE RELOAD DURING A PANCREATIC HEAD RESECTION, ON THE TISSUE, THE SURGEON WAS STRUGGLING WITH THE REMOVAL / OPENING OF THE STAPLER. THE RELOAD COULD NOT BE RELEASED. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND RELEASE THE HANDLE. THERE WAS A DELAY IN THE INTERVENTION. THERE WAS NO PATIENT INJURY.
09-11-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE RELOAD DURING A PANCREATIC HEAD RESECTION, ON THE TISSUE, THE SURGEON WAS STRUGGLING WITH THE REMOVAL / OPENING OF THE STAPLER. THE RELOAD COULD NOT BE RELEASED. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND RELEASE THE HANDLE. THERE WAS A DELAY IN THE INTERVENTION. THERE WAS NO PATIENT INJURY.
09-11-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER FIRING THE RELOAD DURING A PANCREATIC HEAD RESECTION, ON THE TISSUE, THE SURGEON WAS STRUGGLING WITH THE REMOVAL / OPENING OF THE STAPLER. THE RELOAD COULD NOT BE RELEASED. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND RELEASE THE HANDLE. THERE WAS A DELAY IN THE INTERVENTION. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | covidien lp | 20-01-2021 | EEA | Material Deformation | A0406 | Insufficient Information | Event Description: PROCEDURE - ROBOTIC SIGMOID COLECTOMY. DURING RESECTION OF SIGMOID COLON ANASTOMOTIC, STAPLER MALFUNCTION DURING USE. THE ANVIL SPLAYED AND WOULD NOT RELEASE THE TISSUE AFTER IT WAS DEPLOYED. MULTIPLE ATTEMPTS TO CONNECT THE ANVIL TO THE STAPLER SPIKE WERE UNSUCCESSFUL AS IT APPEARED THE INTERVAL COMPONENT WAS MALFORMED AND NOT ENGAGING THE SPIKE, THIS SUBSEQUENTLY FACILITATED THE SPIKE TO BE WITHDRAWN BACK TO THE RECTAL STUMP AND THE INITIAL SITE COULD NOT BE RENEGOTIATED. DUE TO THE MALFORMATION OF THE ANVIL, SURGEON COMPLETED CASE TO AN OPEN INCISION.
16-12-2020 | Malfunction | covidien lp | 20-01-2021 | EEA | Physical Resistance/Sticking | A0509 | Insufficient Information | Event Description: PROCEDURE - ROBOTIC SIGMOID COLECTOMY. DURING RESECTION OF SIGMOID COLON ANASTOMOTIC, STAPLER MALFUNCTION DURING USE. THE ANVIL SPLAYED AND WOULD NOT RELEASE THE TISSUE AFTER IT WAS DEPLOYED. MULTIPLE ATTEMPTS TO CONNECT THE ANVIL TO THE STAPLER SPIKE WERE UNSUCCESSFUL AS IT APPEARED THE INTERVAL COMPONENT WAS MALFORMED AND NOT ENGAGING THE SPIKE, THIS SUBSEQUENTLY FACILITATED THE SPIKE TO BE WITHDRAWN BACK TO THE RECTAL STUMP AND THE INITIAL SITE COULD NOT BE RENEGOTIATED. DUE TO THE MALFORMATION OF THE ANVIL, SURGEON COMPLETED CASE TO AN OPEN INCISION.
15-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT THE SURGEON WAS FIRING A WHITE RELOAD ON THE JJ ANASTOMOSIS AND THE KNIFE BLADE AND STAPLES DEPLOYED ABOUT THIRTY PERCENT. THE KNIFE BLADE STOPPED ADVANCING. THE SURGEON ATTEMPTED TO OPEN BUT THE STAPLER WOULD NOT OPEN. THE SURGEON ATTEMPTED REVERSE BUT THE DEVICE WOULD NOT OPEN. THE SURGEON THEN ATTEMPTED MANUAL OVERRIDE. THE SURGEON HEARD THE LEVER CLICK THEN CRANKED THE LEVER, BUT THE DEVICE STILL WOULD NOT OPEN. THE SURGEON THEN CUT THE DEVICE FROM THE TISSUE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. ONCE REMOVED THE TECH WAS ABLE TO OPEN THE DEVICE ON THE BACK TABLE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
15-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT THE SURGEON WAS FIRING A WHITE RELOAD ON THE JJ ANASTOMOSIS AND THE KNIFE BLADE AND STAPLES DEPLOYED ABOUT THIRTY PERCENT. THE KNIFE BLADE STOPPED ADVANCING. THE SURGEON ATTEMPTED TO OPEN BUT THE STAPLER WOULD NOT OPEN. THE SURGEON ATTEMPTED REVERSE BUT THE DEVICE WOULD NOT OPEN. THE SURGEON THEN ATTEMPTED MANUAL OVERRIDE. THE SURGEON HEARD THE LEVER CLICK THEN CRANKED THE LEVER, BUT THE DEVICE STILL WOULD NOT OPEN. THE SURGEON THEN CUT THE DEVICE FROM THE TISSUE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. ONCE REMOVED THE TECH WAS ABLE TO OPEN THE DEVICE ON THE BACK TABLE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
23-11-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC-ASSISTED SURGERY FOR RECTAL CANCER, COULD NOT FIRE WHEN SEVERING RECTUM TISSUE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | TA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN ILEAL CONDUIT PROCEDURE, WHILE ON THE SCRUB TABLE, THE JAWS OF THE DEVICE WERE NOT ABLE TO OPEN. THE BLACK RELEASE BUTTON WAS NOT WORKING ON. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | GOLD RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VATS RIGHT UPPER LOBECTOMY SURGERY WHEN SEVERING PULMONARY TISSUE ON THE 1ST FIRING, AFTER FIRED, NOTED SOME STAPLE STAPLES MALFORMED WITH AN IRREGULAR SHAPE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE INSTRUMENT DIDN¿T DO ANYTHING. NOTHING HAPPENED WHEN TRIED TO FIRE THE DEVICE. LIGHTS WERE ON BUT NOTHING ELSE. THERE WERE NO PATIENT CONSEQUENCES.
27-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING ENDOSCOPIC LEFT HEMIHEPATECTOMY SURGERY WHEN SEVERING HEPATIC PEDICLE TISSUE, COULD NOT FIRE FARTHER AFTER FIRED A HALF. NOTED THE STAPLES MALFORMED. CHANGED TO THE NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING ENDOSCOPIC-ASSISTED RADICAL GASTRECTOMY FOR PROXIMAL GASTRIC CANCER SURGERY, WHEN SEVERING ESOPHAGUS AND JEJUNUM TISSUE , AFTER FIRED, NOTED SOME STAPLES MALFORMED . USED SUTURE TO OVERSEW . THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY PROCEDURE, THE DEVICE COULD NOT COMPLETE THE FIRING CYCLE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | UNKNOWN EGIA TRI STAPLE | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY PROCEDURE, THE DEVICE COULD NOT COMPLETE THE FIRING CYCLE. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING THE HEPATIC PEDICLE, IN A LIVER RESECTION, AFTER LOADING THE RELOAD WITH THIS STAPLER TO CLAMP THE TISSUE, THE OPENING OF THE RELOAD COULD NOT BE CLOSED AND THE SURGEON COULD NOT FIRE IT. THE SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE OPEN PANCREATIC SURGERY, AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE( WITH STRAIGHT LEG). CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING THE HEPATIC PEDICLE, IN A LIVER RESECTION, AFTER LOADING THE RELOAD WITH THIS STAPLER TO CLAMP THE TISSUE, THE OPENING OF THE RELOAD COULD NOT BE CLOSED AND THE SURGEON COULD NOT FIRE IT. THE SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING LAPAROSCOPIC RIGHT HEMICOLECTOMY SURGERY, AFTER FIRED, NOTED THAT THE INTESTINAL TISSUE WAS NOT CUT, BUT THERE WERE SOME STAPLES DEPLOYED ON THE TISSUE, THEN CHANGED ANOTHER DEVICE TO COMPLETE THE PROCEDURE. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
30-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE OPEN PANCREATIC SURGERY, AFTER FIRED, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING BARIATRIC SURGERY,AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN THE KNIFE AUTOMATICALLY. THE KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. IT IS UNKNOWN HOW THE SURGEON REMOVED THE DEVICE FROM THE PATIENT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RIGHT UPPER LOBECTOMY SURGERY, AFTER THE FIFTH FIRING, NOTED THAT THE ANVIL WAS DAMAGED (BENT OUTWARD), CHANGED TO ANOTHER DEVICE, AND FIRED, THE DEVICE COULD NOT FIRE AFTER FIRED 1/2, OPEN THE JAW, AND NOTED THE ANVIL WAS BENT UPWARD. CHANGED TO THE NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
27-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC GASTRIC BYPASS, DURING FIRST OR SECOND FIRING DURING GASTRIC BYPASS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. IT WAS NOTICED THAT A FEW STAPLES POPPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC GASTRIC BYPASS, DURING FIRST OR SECOND FIRING DURING GASTRIC BYPASS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. IT WAS NOTICED THAT A FEW STAPLES POPPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RIGHT UPPER LOBECTOMY SURGERY, AFTER THE FIFTH FIRING, NOTED THAT THE ANVIL WAS DAMAGED (BENT OUTWARD), CHANGED TO ANOTHER DEVICE AND FIRED, THE DEVICE COULD NOT FIRE AFTER FIRED 1/2, OPEN THE JAW AND NOTED THE ANVIL WAS BENT UPWARD. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RIGHT UPPER LOBECTOMY SURGERY, AFTER THE FIFTH FIRING, NOTED THAT THE ANVIL WAS DAMAGED (BENT OUTWARD), CHANGED TO ANOTHER DEVICE AND FIRED, THE DEVICE COULD NOT FIRE AFTER FIRED 1/2, OPEN THE JAW AND NOTED THE ANVIL WAS BENT UPWARD. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RIGHT UPPER LOBECTOMY SURGERY, AFTER THE FIFTH FIRING, NOTED THAT THE ANVIL WAS DAMAGED (BENT OUTWARD), CHANGED TO ANOTHER DEVICE AND FIRED, THE DEVICE COULD NOT FIRE AFTER FIRED 1/2, OPEN THE JAW AND NOTED THE ANVIL WAS BENT UPWARD. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
27-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC GASTRIC BYPASS, DURING FIRST OR SECOND FIRING DURING GASTRIC BYPASS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. IT WAS NOTICED THAT A FEW STAPLES POPPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC GASTRIC BYPASS, DURING FIRST OR SECOND FIRING DURING GASTRIC BYPASS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. IT WAS NOTICED THAT A FEW STAPLES POPPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC GASTRIC BYPASS, DURING FIRST OR SECOND FIRING DURING GASTRIC BYPASS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. IT WAS NOTICED THAT A FEW STAPLES POPPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2021 | Injury | covidien lp llc north haven | 20-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE TO CURE FOR LUNG CANCER, WHILE BREAKING THE ESSENCE OF THE LUNG, THERE WAS AN INCOMPLETE STAPLE LINE IN THE DISTAL PART. THE STAPLE LINE WAS FIXED BY HANDSEWING. ANOTHER RELOAD USING THE SAME HANDLE WAS USED TO COMPLETE THE OPERATION.
02-01-2021 | Malfunction | covidien lp llc north haven | 20-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE I-BEAM GOT STUCK AND STOPPED MOVING. IT WAS ALSO NOTED THAT THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE, AND THE DEVICE DID NOT FIRE. THE SAME EVENT OCCURRED TWICE. SURGEON REPLACED NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2021 | Malfunction | covidien lp llc north haven | 20-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE I-BEAM GOT STUCK AND STOPPED MOVING. IT WAS ALSO NOTED THAT THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE, AND THE DEVICE DID NOT FIRE. THE SAME EVENT OCCURRED TWICE. SURGEON REPLACED NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2021 | Malfunction | us surgical puerto rico | 20-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE I-BEAM GOT STUCK AND STOPPED MOVING. IT WAS ALSO NOTED THAT THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE. THE SAME EVENT OCCURRED TWICE. SURGEON REPLACED NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING EXTRATHECAL ANASTOMOSIS GRAYSON SYSTEM ON A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE HANDLE WAS NOT SQUEEZED. THE STAPLER DID NOT FIRE. ANOTHER RELOAD OF THE SAME MODEL WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | covidien lp llc north haven | 20-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER A SUCCESSFUL FIRING IN THE DUODENUM IN A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, WHEN SURGEON TRIED TO TRANSECT THE BODY OF STOMACH, SURGEON ATTEMPTED TO FIRE THE DEVICE HOWEVER, THE BLADE OF THE DEVICE STOPPED. SURGEON TRIED TO FIRE THE DEVICE SEVERAL MORE TIMES BUT THE BLADE DOES NOT MOVE ANYMORE. SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
29-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A BOWEL RESECTION PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE HANDLE, HENCE THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN BOWEL RESECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE SAME ISSUE OCCURRED WITH ANOTHER HANDLE WHILE USING THE SAME RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | UNKNOWN EGIA LEGACY STRAIGHT RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN BOWEL RESECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE SAME ISSUE OCCURRED WITH ANOTHER HANDLE WHILE USING THE SAME RELOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-12-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE ON CUTTING HEMORRHOIDS, THE TWO DEVICES COULD NOT BE FIRED. THE PRODUCT WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Injury | covidien lp llc north haven | 20-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF LUNG CANCER, THE DEVICE ABLE TO FIRE AND COMPLETE THE FIRING CYCLE, HOWEVER THERE WERE POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. TO RESOLVE THE ISSUE THE SURGEON HAD TO SUTURE THE TISSUE AND USED ANOTHER RELOAD IN ORDER TO COMPLETE THE CASE.
04-01-2021 | Injury | us surgical puerto rico | 20-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING DISSECTION OF THE LUNG PARENCHYMA IN A RADICAL RESECTION OF LUNG CANCER, THE DEVICE HAD INCOMPLETE STAPLE LINE AT THE DISTAL END. IT WAS ALSO NOTED THAT THERE WAS POOR STAPLE FORMATION. SURGEON HAND STITCHED THE STAPLE LINE. THE JAW OF THE DEVICE ALSO LOCKED ON TISSUE AND WAS REMOVED BY MANUAL SEPARATION. SURGEON REPLACED THE HANDLE TO COMPLETE THE CASE.
04-01-2021 | Injury | us surgical puerto rico | 20-01-2021 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING DISSECTION OF THE LUNG PARENCHYMA IN A RADICAL RESECTION OF LUNG CANCER, THE DEVICE HAD INCOMPLETE STAPLE LINE AT THE DISTAL END. IT WAS ALSO NOTED THAT THERE WAS POOR STAPLE FORMATION. SURGEON HAND STITCHED THE STAPLE LINE. THE JAW OF THE DEVICE ALSO LOCKED ON TISSUE AND WAS REMOVED BY MANUAL SEPARATION. SURGEON REPLACED THE HANDLE TO COMPLETE THE CASE.
04-01-2021 | Injury | us surgical puerto rico | 20-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING DISSECTION OF THE LUNG PARENCHYMA IN A RADICAL RESECTION OF LUNG CANCER, THE DEVICE HAD INCOMPLETE STAPLE LINE AT THE DISTAL END. IT WAS ALSO NOTED THAT THERE WAS POOR STAPLE FORMATION. SURGEON HAND STITCHED THE STAPLE LINE. THE JAW OF THE DEVICE ALSO LOCKED ON TISSUE AND WAS REMOVED BY MANUAL SEPARATION. SURGEON REPLACED THE HANDLE TO COMPLETE THE CASE.
04-01-2021 | Malfunction | covidien lp llc north haven | 20-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE I-BEAM GOT STUCK AND STOPPED MOVING WITH THE TWO RELOADS, THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE. THE SURGEON REPLACED A NEW ONE TO RESOLVE THE ISSUE. THERE WAS NO DAMAGE TO THE PATIENT AT ALL.
04-01-2021 | Malfunction | covidien lp llc north haven | 20-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE I-BEAM GOT STUCK AND STOPPED MOVING WITH THE TWO RELOADS, THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE. THE SURGEON REPLACED A NEW ONE TO RESOLVE THE ISSUE. THERE WAS NO DAMAGE TO THE PATIENT AT ALL.
04-01-2021 | Malfunction | covidien lp llc north haven | 20-01-2021 | UNKNOWN EGIA ADAPTER | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE I-BEAM GOT STUCK AND STOPPED MOVING WITH THE TWO RELOADS, THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE. THE SURGEON REPLACED A NEW ONE TO RESOLVE THE ISSUE. THERE WAS NO DAMAGE TO THE PATIENT AT ALL.
04-01-2021 | Malfunction | covidien lp llc north haven | 20-01-2021 | UNKNOWN STAPLING DEVICE | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE I-BEAM GOT STUCK AND STOPPED MOVING WITH THE TWO RELOADS, THE SURGEON COULD NOT SQUEEZE THE HANDLE ANYMORE. THE SURGEON REPLACED A NEW ONE TO RESOLVE THE ISSUE. THERE WAS NO DAMAGE TO THE PATIENT AT ALL.
04-01-2021 | Malfunction | us surgical puerto rico | 20-01-2021 | GIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED IN COLON RESECTION DURING HARTMANN PROCEDURE, THE BLADE DID NOT RETRACT WITH THE RETURN OF THE PUSHER, IMPOSSIBLE TO CHANGE THE LOADER. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | us surgical puerto rico | 20-01-2021 | GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED IN COLON RESECTION DURING HARTMANN PROCEDURE, THE BLADE DID NOT RETRACT WITH THE RETURN OF THE PUSHER, IMPOSSIBLE TO CHANGE THE LOADER. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-12-2020 | Injury | us surgical puerto rico | 20-01-2021 | EEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A ROBOTIC LEFT COLECTOMY, DURING ANASTOMOSIS OF DESCENDING COLON TO RECTUM, THERE WAS DIFFICULTY IN RE MOVING THE STAPLER FROM THE CAVITY. AFTER THE STAPLER WAS FIRED, THE KNOB WAS TURNED TWO TIMES UNTIL THE AUDIBLE CLICK WAS HEARD. THERE WAS DIFFICULTY REMOVING THE STAPLER AFTER FIRING, BUT AFTER SEVERAL ATTEMPTS, AND WITH ADDITIONAL TENSION, THE STAPLER CAME OUT WITHOUT THE ANVIL ATTACHED. A SIGMOIDOSCOPY WAS PERFORMED AND THE ANVIL WAS RETRIEVED WITH A LAPAROSCOPIC GRASPER.
28-12-2020 | Injury | us surgical puerto rico | 20-01-2021 | EEA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A ROBOTIC LEFT COLECTOMY, DURING ANASTOMOSIS OF DESCENDING COLON TO RECTUM, THERE WAS DIFFICULTY IN RE MOVING THE STAPLER FROM THE CAVITY. AFTER THE STAPLER WAS FIRED, THE KNOB WAS TURNED TWO TIMES UNTIL THE AUDIBLE CLICK WAS HEARD. THERE WAS DIFFICULTY REMOVING THE STAPLER AFTER FIRING, BUT AFTER SEVERAL ATTEMPTS, AND WITH ADDITIONAL TENSION, THE STAPLER CAME OUT WITHOUT THE ANVIL ATTACHED. A SIGMOIDOSCOPY WAS PERFORMED AND THE ANVIL WAS RETRIEVED WITH A LAPAROSCOPIC GRASPER.
30-12-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC PROCEDURE, THE DEVICE LOAD FIRED ONLY HALFWAY. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-12-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, THE DEVICE COULD NOT BE FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THIRD FIRING OVER STOMACH ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER LOCKED ON TISSUE AND WOULD NOT UNLOCK AFTER HOLDING UP, THEN THE STAPLER POWERED DOWN AND RESTARTED. ONCE RESTARTED THE STAPLER UNLOCKED. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | us surgical puerto rico | 20-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THIRD FIRING OVER STOMACH ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER LOCKED ON TISSUE AND WOULD NOT UNLOCK AFTER HOLDING UP, THEN THE STAPLER POWERED DOWN AND RESTARTED. ONCE RESTARTED THE STAPLER UNLOCKED. THERE WAS NO PATIENT INJURY.
02-12-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THIRD FIRING OVER STOMACH ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER LOCKED ON TISSUE AND WOULD NOT UNLOCK AFTER HOLDING UP, THEN THE STAPLER POWERED DOWN AND RESTARTED. ONCE RESTARTED THE STAPLER UNLOCKED. THERE WAS NO PATIENT INJURY.
28-09-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEMONSTRATION, WHEN THE DEVICE TRIED TO BE CHARGED, THE INDICATOR FLASHED RED IMMEDIATELY. THE HANDLE COULD NOT BE CHARGED. THE HANDLE DID NOT TURN ON AND THE SCREEN FOUND TO HAVE BLACKED OUT. USER TRIED TO CHARGE ANOTHER HANDLE WITH THE SAME CHARGER, IT WAS ABLE TO CHARGE WITHOUT PROBLEM. NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT AN ANALYSIS THE SYSTEM LOGS NOTED THAT THE HANDLE WAS PLACED ON A CHARGER. THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
28-09-2020 | Malfunction | covidien lp llc north haven | 20-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEMONSTRATION, WHEN THE DEVICE TRIED TO BE CHARGED, THE INDICATOR FLASHED RED IMMEDIATELY. THE HANDLE COULD NOT BE CHARGED. THE HANDLE DID NOT TURN ON AND THE SCREEN FOUND TO HAVE BLACKED OUT. USER TRIED TO CHARGE ANOTHER HANDLE WITH THE SAME CHARGER, IT WAS ABLE TO CHARGE WITHOUT PROBLEM. NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT AN ANALYSIS THE SYSTEM LOGS NOTED THAT THE HANDLE WAS PLACED ON A CHARGER. THE NEXT TIME THE HANDLE INITIALIZES IS WITH LOW VOLTAGE DETECTED ON THE CHARGER. THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
21-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE FIRED THREE TIMES AND THE STAPLES WERE MANGLED AND TWISTED. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
31-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE IS NOT MAKING THE PROPER "B" SHAPE WHEN IT STAPLES. THEY STAPLED OVER THE STAPLE LINE AND OVER-SEWED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
17-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING RIGHT UPPER LOBECTOMY SURGERY, AFTER FIRED, NOTED THAT THE JAW COULD NOT OPEN, WAS LOCKED ON TISSUE, AND KNIFE REVERSE BUTTON COULD NOT WORK. UNKNOWN HOW DEVICE WAS REMOVED AND UNKNOWN IF JAWS EVENTUALLY OPENED. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
29-12-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | PWRD ECH FLEX 45MM 28CM SHAFT | Excessive Heating | A1002 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VATS SURGERY, THE SURGEON NOTED THAT THERE WAS "SMOG" COMING FROM THE BATTERY PACKING. THERE WAS NO PATIENT CONSEQUENCE. NO ADDITION INFORMATION COULD BE PROVIDED.
04-01-2020 | Malfunction | ethicon endo-surgery | 20-01-2021 | BLUE RELOAD FOR ECHELON 45 | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, CARTRIDGE PAN WAS DEFORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-01-2021 | Malfunction | intuitive surgical | 20-01-2021 | INTUITIVE SURGICAL DAVINCI XI ENDOWRIST STAPLER 45 | Misfire | A050502 | Blood Loss | Event Description: A (B)(6) YEAR-OLD MALE PATIENT, UNDERGOING ROBOTIC ASSISTED TOTAL ABDOMINAL COLECTOMY. DURING PROCEDURE, THE WHITE VASCULAR LOAD MISFIRED OVER THE COLIC ARTERY TO THE SIGMOID COLON. THE STAPLER LOAD PARTIAL FIRED AND CUT. THERE WAS A SMALL AMOUNT OF BLEEDING FROM THE VESSEL. SURGEON PLACED A CLIP ON VESSEL. INTUITIVE DAVINCI XI SURGICAL SYSTEMS ENDOWRIST STAPLER 44MM RELOAD REFERENCE # 48645W, LOT # M90200831. FDA SAFETY REPORT ID # (B)(4).
03-01-2021 | Malfunction | us surgical puerto rico | 21-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY, THE DEVICE WAS UNABLE TO FIRE -- GREEN BUTTON COULD NOT BE PRESSED NOR HANDLE COULD BE SQUEEZED. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | covidien lp llc north haven | 21-01-2021 | UNKNOWN ENDO STITCH SULU | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y, THE DEVICE JAMMED WHILE SUTURING. IN ADDITION, THE SURGEON HAD DIFFICULTY LOADING AND UNLOADING THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | covidien lp llc north haven | 21-01-2021 | UNKNOWN ENDO STITCH SULU | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y, THE DEVICE JAMMED WHILE SUTURING. IN ADDITION, THE SURGEON HAD DIFFICULTY LOADING AND UNLOADING THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | covidien lp llc north haven | 21-01-2021 | UNKNOWN ENDO STITCH SULU | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y, THE DEVICE JAMMED WHILE SUTURING. IN ADDITION, THE SURGEON HAD DIFFICULTY LOADING AND UNLOADING THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-11-2020 | Malfunction | ethicon endo-surgery | 21-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, LOCKED JAW. RIGHT AFTER CLOSING ON VESSEL, SURGEON WANTED TO ADJUST THE POSITION OF STAPLE BUT COULD NOT OPEN THE JAW. THERE WERE NO PATIENT CONSEQUENCES.
30-10-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE BLUE LAMP ON THE CHARGER ILLUMINATED WHEN THE CHARGER WAS CONNECTED TO THE POWER SOURCE, AND FLASHED GREEN WHEN THE HANDLE WAS SET ON IT, BUT SOMETHING LIKE BURNT WAS FOUND ON THE ELECTRODE PART OF THE CHARGER. THE HANDLE WOULD NOT START UP EVEN THOUGH REMOVED FROM THE CHARGER WHILE THE GREEN LAMP WAS FLASHING.(THERE WAS NO RESPONSE OF EITHER THE SOUND OR THE SCREEN.) THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
30-10-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE BLUE LAMP ON THE CHARGER ILLUMINATED WHEN THE CHARGER WAS CONNECTED TO THE POWER SOURCE, AND FLASHED GREEN WHEN THE HANDLE WAS SET ON IT, BUT SOMETHING LIKE BURNT WAS FOUND ON THE ELECTRODE PART OF THE CHARGER. THE HANDLE WOULD NOT START UP EVEN THOUGH REMOVED FROM THE CHARGER WHILE THE GREEN LAMP WAS FLASHING.(THERE WAS NO RESPONSE OF EITHER THE SOUND OR THE SCREEN.) THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
04-01-2021 | Malfunction | us surgical puerto rico | 21-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC SURGERY, ON LAST STAPLING OF THE STOMACH, IT WAS NOT POSSIBLE TO CLOSE THE HANDLE AND DID NOT FIRE. BOTH HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
28-09-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEMONSTRATION FOR LOW ANTERIOR RESECTION, THE CHARGE STATUS OF THE BATTERY CHARGER FLASHED RED WHEN THE HANDLE WAS BEING CHARGED. WHEN IT WAS REMOVED FROM THE BATTERY CHARGER, NOTHING APPEARED ON THE DISPLAY. THE BATTERY CHARGER WAS REPLACED, BUT THE SAME EVENT OCCURRED. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
28-09-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEMONSTRATION FOR LOW ANTERIOR RESECTION, THE CHARGE STATUS OF THE BATTERY CHARGER FLASHED RED WHEN THE HANDLE WAS BEING CHARGED. WHEN IT WAS REMOVED FROM THE BATTERY CHARGER, NOTHING APPEARED ON THE DISPLAY. THE BATTERY CHARGER WAS REPLACED, BUT THE SAME EVENT OCCURRED. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
03-11-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WOULD BE DEAD WHEN TAKEN OFF THE CHARGER. THE CHARGER WAS ABLE TO WORK ON OTHER HANDLES WITH NO ISSUES. WHEN THE HANDLE WAS TRIED ON ANOTHER CHARGER, BUT THE HANDLE REMAINED DEAD. THE HANDLE WAS TRIED TO BE REBOOTED WITH NO SUCCESS. THERE WAS NO PATIENT INVOLVEMENT. PLI-10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES TO BE VENTED.
03-11-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WOULD BE DEAD WHEN TAKEN OFF THE CHARGER. THE CHARGER WAS ABLE TO WORK ON OTHER HANDLES WITH NO ISSUES. WHEN THE HANDLE WAS TRIED ON ANOTHER CHARGER, BUT THE HANDLE REMAINED DEAD. THE HANDLE WAS TRIED TO BE REBOOTED WITH NO SUCCESS. THERE WAS NO PATIENT INVOLVEMENT. PLI-10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES TO BE VENTED.
28-12-2020 | Malfunction | ethicon endo-surgery | 21-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY ON THE FOURTH FIRING THE BUTTRESSING MATERIAL CAUSED THE TISSUE TO BE PUSHED OUT THE DISTAL END OF THE STAPLER. THIS CAUSED SOME OF THE STAPLES TO BE MALFORMED AND SOME PARTIALLY DEPLOYED. THE SURGEON HAND SEWED THE HOLE IN THE PATIENT¿S STOMACH CLOSED AND COMPLETED THE REST OF THE PROCEDURE WITH THIS SAME DEVICE WITH NO PATIENT CONSEQUENCE. THE PROCEDURE WAS DELAYED BY TEN MINUTES AND THERE WAS SOME MINOR BLEEDING.
19-12-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED IN TISSUE DURING VIDEO RECTOSIGMOIDECTOMY, THE DEVICE WAS SET FOR STAPLING, BUT THE BLADE DID NOT PROGRESS. THE SAME EVENT OCCURRED TWICE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE BEING APPLIED IN TISSUE DURING VIDEO RECTOSIGMOIDECTOMY, THE DEVICE WAS SET FOR STAPLING, BUT THE BLADE DID NOT PROGRESS. THE SAME EVENT OCCURRED TWICE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-01-2020 | Malfunction | ethicon endo-surgery | 21-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BARIATRIC PROCEDURE, THE GST RELOADS ARE VERY DIFFICULT TO INSERT INTO THE DEVICE, REQUIRING THE SURGEON TO INSERT THEM FOR THE SCRUB TECH. THERE WAS NO DELAY AND THE SURGERY WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
17-11-2020 | Malfunction | ethicon endo-surgery | 21-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT CLOSE THE JAW BEFORE INSERTING IT INTO TROCAR. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-12-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY, AFTER CHECKING THE OPENING AND CLOSING, THE PRODUCT WAS INSERTED INTO THE ABDOMINAL CAVITY THROUGH THE PORT AND THE CARTRIDGE WAS OPENED, THE REINFORCEMENT SHEET WAS FOUND TO BE DISENGAGED FROM THE CARTRIDGE. THE JAWS LOCKED ON TISSUE. THE DEVICE WAS ABLE TO BE REMOVED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-12-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | TRI-STAPLE 2.0 | Premature Activation | A150103 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPY-ASSISTED DISTAL GASTRECTOMY, AFTER CHECKING THE OPENING AND CLOSING, THE PRODUCT WAS INSERTED INTO THE ABDOMINAL CAVITY THROUGH THE PORT AND THE CARTRIDGE WAS OPENED, THE REINFORCEMENT SHEET WAS FOUND TO BE DISENGAGED FROM THE CARTRIDGE. THE JAWS LOCKED ON TISSUE. THE DEVICE WAS ABLE TO BE REMOVED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE, THE ISSUE THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE, THE ISSUE THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE, THE ISSUE THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE, THE ISSUE THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Detachment of Device or Device Component | A0501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
04-01-2021 | Injury | covidien lp llc north haven | 21-01-2021 | SIGNIA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO MANUALLY PULLED THE DEVICE FROM THE PATIENT TISSUE USING GRASPER, RESULTING TO A BLEEDING AND TISSUE DAMAGE. IN ADDITION, THE SURGEON FOUND OUT THAT THERE WERE MALFORMED STAPLES IN THE STAPLE LINE. THE USER TRIED HOLDING THE GREEN BUTTONS, ATTACHING AN ALTERNATE POWERED HANDLE AND USING A MANUAL RETRACTION TOO, BUT NONE OF THESE RESOLVE THE ISSUE. THE SURGEON USED NEW HANDLE, ADAPTER AND RELOAD TO COMPLETE THE CASE.
20-04-2021 | Malfunction | covidien lp llc north haven | 21-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-12-2020 | Malfunction | us surgical puerto rico | 21-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AFTER THE FIRST SQUEEZE THE STAPLES DEPLOYED ON THE TISSUE IN A U SHAPE WITHOUT FORMING A PROPER B SHAPE. IT WAS ALSO NOTED THAT THE STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL PART. TWO RELOADS MALFUNCTIONED DURING THIS PROCEDURE. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-01-2021 | Malfunction | covidien lp llc north haven | 21-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE STAPLER WAS ABLE TO FIRE BUT STAPLE LINE AS INCOMPLETE. THE REINFORCEMENT WAS CUT TO COMPLETE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | celestica de monterrey s.a de c.v | 21-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, ON THE RECTUM, THE STAPLES BURST WHEN HALF FIRED, THERE WAS A CLICKING SOUND, AND THE FIRING COULD NOT BE CONTINUED, THE STAPLE LINE WAS INCOMPLETE CENTRALLY. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | celestica de monterrey s.a de c.v | 21-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, ON THE RECTUM, THE STAPLES BURST WHEN HALF FIRED, THERE WAS A CLICKING SOUND, AND THE FIRING COULD NOT BE CONTINUED, THE STAPLE LINE WAS INCOMPLETE CENTRALLY. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 21-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, ON THE RECTUM, THE STAPLES BURST WHEN HALF FIRED, THERE WAS A CLICKING SOUND, AND THE FIRING COULD NOT BE CONTINUED, THE STAPLE LINE WAS INCOMPLETE CENTRALLY. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 21-01-2021 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, ON THE RECTUM, THE STAPLES BURST WHEN HALF FIRED, THERE WAS A CLICKING SOUND, AND THE FIRING COULD NOT BE CONTINUED, THE STAPLE LINE WAS INCOMPLETE CENTRALLY. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-12-2020 | Malfunction | covidien lp llc north haven | 21-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT LOWER LOBECTOMY WITH LYMPH NODE DISSECTION, IT WAS FOUND OUT THAT THE RELOAD COULD NOT BE CLOSED AND COULD NOT BE FIRED NORMALLY. REPLACE WITH A NEW RELOAD, CONTINUE TO OPERATE ON PATIENT. THERE WAS NO PATIENT INJURY.
04-01-2021 | Injury | us surgical puerto rico | 21-01-2021 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING USED ON ESOPHAGUS AND STOMACH ANASTOMOSIS DURING A LAPAROSCOPIC ESOPHAGECTOMY FOR THORACIC ESOPHAGEAL CARCINOMA, THE DONUTS WERE COMPLETE BUT PART OF THE STAPLES COULD NOT FORMED CORRECTLY AND THE ANASTOMOSIS WAS NOT TIGHT. FOR THE PARTS THAT STAPLES WERE NOT WELL SHAPED, IT WAS SUTURED BY HAND.
29-12-2020 | Malfunction | us surgical puerto rico | 21-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE 12TH FIRING, THE DEVICE WAS PLACED ON THE STOMACH AND IT WAS STUCK, WOULD NOT FIRE OR RELEASE WITHOUT WIGGLING. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-12-2020 | Malfunction | us surgical puerto rico | 21-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE 12TH FIRING, THE DEVICE WAS PLACED ON THE STOMACH AND IT WAS STUCK, WOULD NOT FIRE OR RELEASE WITHOUT WIGGLING. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-12-2020 | Malfunction | us surgical puerto rico | 21-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE 12TH FIRING, THE DEVICE WAS PLACED ON THE STOMACH AND IT WAS STUCK, WOULD NOT FIRE OR RELEASE WITHOUT WIGGLING. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-12-2020 | Malfunction | us surgical puerto rico | 21-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE 12TH FIRING, THE DEVICE WAS PLACED ON THE STOMACH AND IT WAS STUCK, WOULD NOT FIRE OR RELEASE WITHOUT WIGGLING. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-08-2016 | Malfunction | us surgical puerto rico | 21-01-2021 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE ENDOSCOPIC STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
15-08-2016 | Malfunction | us surgical puerto rico | 21-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE ENDOSCOPIC STAPLER DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
05-01-2021 | Malfunction | ethicon endo-surgery | 21-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING PULMONARY LOBE TISSUE ON THE THIRD FIRING, AFTER FIRED, NOTED SOME STAPLES WERE MALFORMED WITH IRREGULAR SHAPE.CHANGED TO NEW ONE TO COMPLETE SURGERY . THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2021 | Malfunction | ethicon endo-surgery | 21-01-2021 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN GASTROJEJUNOSTOMY, THE STAPLES AT THE DISTAL END OF THE CARTRIDGE WERE UNFORMED. THIS OCCURRED ON THE SECOND FIRING. THE PROCEDURE WAS NOT CHANGED. THE TISSUE WAS COMPRESSED BEFORE FIRING, THERE WAS NO PROBLEM IN FIRING, AND THE KNIFE WAS NOT STOPPED IN THE FIRING. DOCTOR COMMENTED THE TARGET TISSUE WAS A LITTLE SOLID, THE PATIENT WAS ABOUT 90 YEARS OLD, AND THE PATIENT HAD ADVANCED CANCER. ADDITIONAL HAND SUTURING WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCE TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-07-2020 | Malfunction | ethicon endo-surgery | 21-01-2021 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PULMONARY SPACE OCCUPYING SURGERY, WHEN SEVERING TISSUE, AFTER FIRED, NOTED STAPLES MALFORMED. CHANGED TO NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. THIS CASE WAS REPORTED BY THE HEATH AUTHORITY.
31-12-2020 | Malfunction | ethicon endo-surgery | 21-01-2021 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBECTOMY SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. AS THE TISSUE WAS CUT OFF, REMOVED THE DEVICE FROM THE PATIENT THROUGH TISSUE GAP. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
05-01-2021 | Malfunction | ethicon endo-surgery | 21-01-2021 | POWERED PLUS STANDARD 45 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THE REP DURING A VATS WHEN THE JAWS WERE CLOSED THE CARTRIDGE BROKE APART AND PIECES OF IT FELL INTO THE PATIENT. THE PIECES WERE RECOVERED AND REMOVED. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE WITH NO PATIENT CONSEQUENCES.
05-01-2021 | Malfunction | ethicon endo-surgery | 21-01-2021 | POWERED PLUS STANDARD 45 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THE REP DURING A VATS WHEN THE JAWS WERE CLOSED THE CARTRIDGE BROKE APART AND PIECES OF IT FELL INTO THE PATIENT. THE PIECES WERE RECOVERED AND REMOVED. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE WITH NO PATIENT CONSEQUENCES.
28-12-2020 | Malfunction | ethicon endo-surgery | 21-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WAS LOADED AND FIRED ON VESSEL. THEY NEEDED TO USE DEVICE FOR SECOND VESSEL BUT DEVICE COULD NOT BE OPENED TO RELOAD. IT IS UNKNOWN IF THE DEVICE WAS LOCKED ON TISSUE WITH JAWS CLOSED. IT IS UNKNOWN HOW PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
03-01-2021 | Malfunction | ethicon endo-surgery | 21-01-2021 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LEFT UPPER LOBECTOMY, AFTER INSTALLED RELOAD AND CLOSED THE JAW, PLACED JAW INTO BODY THROUGH TROCAR AS STANDARD PROCEDURE BUT COULD NOT OPEN THE JAW ANYMORE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
05-01-2021 | Malfunction | ethicon endo-surgery | 21-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, THE STAPLES COULD NOT BE DEPLOYED AT THE THIRD FIRING. THE CARTRIDGE COLOR WAS GOLD. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-01-2021 | Malfunction | ethicon endo-surgery | 21-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO HOME POSITION AT THE FIRST FIRING ON THE DUODENUM. THE TARGET TISSUE WAS CUT, AND THE STAPLES WERE DEPLOYED PROPERLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ONE NO FURTHER INFORMATION IS AVAILABLE.
29-12-2020 | Injury | us surgical puerto rico | 22-01-2021 | TA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN RIGHT COLECTOMY, DURING ANASTOMOSIS, THERE WAS POOR STAPLE FORMATION. THE SURGEON HAD TO PICK THE MALFORMED STAPLES OUT OF THE TISSUE. THE STAPLE LINE WAS RESECTED, RESTAPLED, AND THEN OVER SEWN. THERE WAS TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED FOR MORE THAN 10 DAYS, WHICH WAS ALSO DUE TO OTHER VARIABLES.
06-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC SEGMENTECTOMY SURGERY, AFTER FIRED, A SPRING HAD FALLEN OFF FROM THE DEVICE. UNKNOWN HOW THEY OPENED THE JAW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
04-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC SURGERY, ON LAST STAPLING OF THE STOMACH, IT WAS NOT POSSIBLE TO CLOSE THE HANDLE AND DID NOT FIRE. BOTH HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY FOR PANCREATIC CANCER, WHEN RESECTING THE PANCREAS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. ALSO, THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WERE REMOVED BY PULLING BACK THE BLACK RETURN KNOB. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY FOR PANCREATIC CANCER, WHEN RESECTING THE PANCREAS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. ALSO, THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WERE REMOVED BY PULLING BACK THE BLACK RETURN KNOB. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING PULMONARY LOBE TISSUE ON THE THIRD FIRING, AFTER FIRED, NOTED SOME STAPLES WERE MALFORMED WITH IRREGULAR SHAPE.CHANGED TO NEW ONE TO COMPLETE SURGERY . THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY FOR PANCREATIC CANCER, WHEN RESECTING THE PANCREAS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. ALSO, THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WERE REMOVED BY PULLING BACK THE BLACK RETURN KNOB. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY FOR PANCREATIC CANCER, WHEN RESECTING THE PANCREAS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. ALSO, THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WERE REMOVED BY PULLING BACK THE BLACK RETURN KNOB. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2020 | Malfunction | covidien lp llc north haven | 22-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROTECTOMY, THE SURGEON COULD NOT SQUEEZE THE HANDLE; HENCE, THE STAPLER DID NOT FIRE. THE SURGEON STATED THE SURGICAL FIELD WAS NARROW. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBECTOMY. WHEN SERVING THE LUNG TISSUE, THE FIRST FIVE RELOADS GOT MALFORMED STAPLES ISSUE AFTER FIRED, THEN AFTER FIRED ECR60W, NOTED THAT THE TISSUE WAS CUT BUT WITHOUT ANY STAPLES DEPLOYED ON THE TISSUE AND CAUSED BLEEDING, USED SUTURE TO OVERSEW AND STOP BLEEDING. THE PATIENT'S CONDITION IS CURRENTLY STABLE. THE SURGEON NOTED THAT STAPLES STILL IN THE ANVIL POCKET AND TRIED TO FIRE THE DEVICE, FOUND THAT THE JAW OPEN WHEN SQUEEZING DOWN THE FIRING TRIGGER.
06-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBECTOMY. WHEN SERVING THE LUNG TISSUE, THE FIRST FIVE RELOADS GOT MALFORMED STAPLES ISSUE AFTER FIRED, THEN AFTER FIRED ECR60W, NOTED THAT THE TISSUE WAS CUT BUT WITHOUT ANY STAPLES DEPLOYED ON THE TISSUE AND CAUSED BLEEDING, USED SUTURE TO OVERSEW AND STOP BLEEDING. THE PATIENT'S CONDITION IS CURRENTLY STABLE. THE SURGEON NOTED THAT STAPLES STILL IN THE ANVIL POCKET AND TRIED TO FIRE THE DEVICE, FOUND THAT THE JAW OPEN WHEN SQUEEZING DOWN THE FIRING TRIGGER.
06-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBECTOMY. WHEN SERVING THE LUNG TISSUE, THE FIRST FIVE RELOADS GOT MALFORMED STAPLES ISSUE AFTER FIRED, THEN AFTER FIRED ECR60W, NOTED THAT THE TISSUE WAS CUT BUT WITHOUT ANY STAPLES DEPLOYED ON THE TISSUE AND CAUSED BLEEDING, USED SUTURE TO OVERSEW AND STOP BLEEDING. THE PATIENT'S CONDITION IS CURRENTLY STABLE. THE SURGEON NOTED THAT STAPLES STILL IN THE ANVIL POCKET AND TRIED TO FIRE THE DEVICE, FOUND THAT THE JAW OPEN WHEN SQUEEZING DOWN THE FIRING TRIGGER.
06-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOSCOPIC LOBECTOMY. WHEN SERVING THE LUNG TISSUE, THE FIRST FIVE RELOADS GOT MALFORMED STAPLES ISSUE AFTER FIRED, THEN AFTER FIRED ECR60W, NOTED THAT THE TISSUE WAS CUT BUT WITHOUT ANY STAPLES DEPLOYED ON THE TISSUE AND CAUSED BLEEDING, USED SUTURE TO OVERSEW AND STOP BLEEDING. THE PATIENT'S CONDITION IS CURRENTLY STABLE. THE SURGEON NOTED THAT STAPLES STILL IN THE ANVIL POCKET AND TRIED TO FIRE THE DEVICE, FOUND THAT THE JAW OPEN WHEN SQUEEZING DOWN THE FIRING TRIGGER.
21-10-2020 | Injury | us surgical puerto rico | 22-01-2021 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A (B)(6) YEAR OLD MALE UNDERWENT A MINIMALLY INVASIVE IVOR LEWIS OESOPHAGECTOMY FOR OESOPHAGEAL ADENOCARCINOMA. EIGHT DAYS POST-OPERATIVELY, THE PATIENT DEVELOPED AN UPPER GASTROINTESTINAL BLEED, HEMATEMESIS AND REQUIRED INTUBATION FOR AIRWAY PROTECTION. THE PATIENT RECEIVED MULTIPLE BLOOD TRANSFUSIONS, UNDERWENT SEVERAL ENDOSCOPIES AND DEVELOPED ORGAN FAILURE. CHALLENGING DELAYED BLEEDING AFTER AN IVOR LEWIS OESOPHAGECTOMY. 2020 SEAN LIDDLE, ANIRUDH MIRAKHUR AND ESTIFANOS DEBRUL
21-10-2020 | Injury | us surgical puerto rico | 22-01-2021 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Gastrointestinal Hemorrhage | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A (B)(6) YEAR OLD MALE UNDERWENT A MINIMALLY INVASIVE IVOR LEWIS OESOPHAGECTOMY FOR OESOPHAGEAL ADENOCARCINOMA. EIGHT DAYS POST-OPERATIVELY, THE PATIENT DEVELOPED AN UPPER GASTROINTESTINAL BLEED, HEMATEMESIS AND REQUIRED INTUBATION FOR AIRWAY PROTECTION. THE PATIENT RECEIVED MULTIPLE BLOOD TRANSFUSIONS, UNDERWENT SEVERAL ENDOSCOPIES AND DEVELOPED ORGAN FAILURE. CHALLENGING DELAYED BLEEDING AFTER AN IVOR LEWIS OESOPHAGECTOMY. 2020 SEAN LIDDLE, ANIRUDH MIRAKHUR AND ESTIFANOS DEBRUL
21-10-2020 | Injury | us surgical puerto rico | 22-01-2021 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hematemesis | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A (B)(6) YEAR OLD MALE UNDERWENT A MINIMALLY INVASIVE IVOR LEWIS OESOPHAGECTOMY FOR OESOPHAGEAL ADENOCARCINOMA. EIGHT DAYS POST-OPERATIVELY, THE PATIENT DEVELOPED AN UPPER GASTROINTESTINAL BLEED, HEMATEMESIS AND REQUIRED INTUBATION FOR AIRWAY PROTECTION. THE PATIENT RECEIVED MULTIPLE BLOOD TRANSFUSIONS, UNDERWENT SEVERAL ENDOSCOPIES AND DEVELOPED ORGAN FAILURE. CHALLENGING DELAYED BLEEDING AFTER AN IVOR LEWIS OESOPHAGECTOMY. 2020 SEAN LIDDLE, ANIRUDH MIRAKHUR AND ESTIFANOS DEBRUL
19-10-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | ECHELON FLEX | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: SURGICAL TECHNICIAN REPORTED THE ECHELON POWERED STAPLER WOULD NOT "SEAT LOADS." THE TECH WAS UNABLE TO PROPERLY LOAD THE STAPLER AS THE JAW APPEARED TO MALFUNCTION. *THE DEVICE WAS NOT USED ON A PATIENT.
19-10-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | ECHELON FLEX | Defective Component | A0202 | No Clinical Signs, Symptoms or Conditions | Event Description: SURGICAL TECHNICIAN REPORTED THE ECHELON POWERED STAPLER WOULD NOT "SEAT LOADS." THE TECH WAS UNABLE TO PROPERLY LOAD THE STAPLER AS THE JAW APPEARED TO MALFUNCTION. *THE DEVICE WAS NOT USED ON A PATIENT.
07-12-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | PROXIMATE | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE STAFF, THE DEVICE MISFIRED WHILE THE SURGEON AND FIRST ASSIST WAS USING IT.
05-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN USING THE TWO RELOADS, AFTER STAPLING THE STOMACH WHEN THE BLADE WAS RETRACTED FROM THE STAPLER LINE, THERE WAS A WHITE PLASTIC OBJECT THAT BROKE OFF. A NEW STAPLER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-11-2020 | Malfunction | covidien lp llc north haven | 22-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, WHEN FIRING ON THE STOMACH, THE STAPLER WOULD NOT FIRE THE RELOAD WHEN IT WAS ARTICULATED. THE HANDLE WAS NOT SQUEEZED. ANOTHER HANDLE AND RELOAD WERE OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-01-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, THE STAPLER WAS STOPPED AND HE HAD TO OPENED IT MANUALLY; THE DEVICE WAS FULLY FIRED AND STOPPED DURING THE AUTOMATIC KNIFE RETURN. THE JAWS DID NOT EVENTUALLY OPEN. THE SURGEON REPLACED THE PRODUCT WITH OTHER COMPETITORS' PRODUCT. NO DAMAGE CAUSED TO THE PATIENT. NO CHANGE IN THE POST-OPERATIVE CARE OR SIGNIFICANT DELAY.
16-01-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, THE STAPLER WAS STOPPED AND HE HAD TO OPENED IT MANUALLY; THE DEVICE WAS FULLY FIRED AND STOPPED DURING THE AUTOMATIC KNIFE RETURN. THE JAWS DID NOT EVENTUALLY OPEN. THE SURGEON REPLACED THE PRODUCT WITH OTHER COMPETITORS' PRODUCT. NO DAMAGE CAUSED TO THE PATIENT. NO CHANGE IN THE POST-OPERATIVE CARE OR SIGNIFICANT DELAY.
04-01-2021 | Malfunction | celestica de monterrey s.a de c.v | 22-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DETACHMENT OF THE RIGHT COLON INTESTINE IN A LAPAROSCOPIC RADICAL RESECTION OF THE RIGHT COLON, THE SURGEON TRIED TO FIRE THE DEVICE, BUT THE GRIP COULD NOT BE FIRED. LATER, AFTER CHECKING THE RELOAD, SEVERAL STAPLES HAD EMERGED FROM THE INSIDE OF THE RELOAD. THE DEVICE WAS REPLACED WITH ANOTHER MANUFACTURER'S DEVICE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | celestica de monterrey s.a de c.v | 22-01-2021 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DETACHMENT OF THE RIGHT COLON INTESTINE IN A LAPAROSCOPIC RADICAL RESECTION OF THE RIGHT COLON, THE SURGEON TRIED TO FIRE THE DEVICE, BUT THE GRIP COULD NOT BE FIRED. LATER, AFTER CHECKING THE RELOAD, SEVERAL STAPLES HAD EMERGED FROM THE INSIDE OF THE RELOAD. THE DEVICE WAS REPLACED WITH ANOTHER MANUFACTURER'S DEVICE. THERE WAS NO PATIENT INJURY.
01-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | ENDO GIA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAFF AND SURGE ON WERE NOT ABLE TO GET THE STAPLER AND RELOAD TO FIRE. SEVERAL ATTEMPTS WERE MADE TO UNLOAD AND LOAD THE RELOAD WITH NO SUCCESS. THE SURGEON ROTATED THE SHAFT AND CLOSED THE JAWS ON THE TISSUE. IT WAS EXPLAINED THAT THE SURGEON HAD TO HOLD THE HANDLE OF THE STAPLER CLOSED TO KEEP THE JAWS SHUT ON THE TISSUE. THEN THE SURGEON ASKED THE TECH TO PUSH THE GREEN FIRING BUTTON. WHEN THE BUTTON WAS PUSHED, IT RETRACTED INTO THE STAPLER HANDLE SO ONLY HALF OF THE BUTTON WAS VISIBLE. IT WAS ALSO NOTED THAT THE SURGEON WAS STILL HOLDING THE HANDLE IN THE SHUT POSITION. IT WAS MENTIONED THAT IF THE SURGEON RELEASES THE HANDLE THE JAWS WOULD OPEN. THEY WOULD NOT STAY CLOSED ON THE TISSUE. ANOTHER STAPLER WAS OPENED AND CONTINUED TO HAVE ISSUE TRYING TO FIRE. THE SURGEON ENDED UP USING ENDO LOOP IN PLACE OF THE STAPLER. IT WAS ALSO NOTED THAT THE RELOAD WAS NOT ABLE TO BE USED WITH A DIFFERENT STAPLER. THERE WAS NO PATIENT INJURY.
01-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAFF AND SURGE ON WERE NOT ABLE TO GET THE STAPLER AND RELOAD TO FIRE. SEVERAL ATTEMPTS WERE MADE TO UNLOAD AND LOAD THE RELOAD WITH NO SUCCESS. THE SURGEON ROTATED THE SHAFT AND CLOSED THE JAWS ON THE TISSUE. IT WAS EXPLAINED THAT THE SURGEON HAD TO HOLD THE HANDLE OF THE STAPLER CLOSED TO KEEP THE JAWS SHUT ON THE TISSUE. THEN THE SURGEON ASKED THE TECH TO PUSH THE GREEN FIRING BUTTON. WHEN THE BUTTON WAS PUSHED, IT RETRACTED INTO THE STAPLER HANDLE SO ONLY HALF OF THE BUTTON WAS VISIBLE. IT WAS ALSO NOTED THAT THE SURGEON WAS STILL HOLDING THE HANDLE IN THE SHUT POSITION. IT WAS MENTIONED THAT IF THE SURGEON RELEASES THE HANDLE THE JAWS WOULD OPEN. THEY WOULD NOT STAY CLOSED ON THE TISSUE. ANOTHER STAPLER WAS OPENED AND CONTINUED TO HAVE ISSUE TRYING TO FIRE. THE SURGEON ENDED UP USING ENDO LOOP IN PLACE OF THE STAPLER. IT WAS ALSO NOTED THAT THE RELOAD WAS NOT ABLE TO BE USED WITH A DIFFERENT STAPLER. THERE WAS NO PATIENT INJURY.
18-12-2020 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING MAGENRESECTION, THE KNIFE STOPPED AFTER ADVANCING ABOUT 20 MM. THE FIRING WAS UNABLE TO BE PERFORMED AGAIN AFTER WAITING FOR A WHILE, SO A NEW PRODUCT WAS USED AND RESECTION WAS PERFORMED.
04-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING PULMONARY LOBE TISSUE ON THE FIRST FIRING, AFTER FIRED, NOTED ALL THE STAPLES WERE MALFORMED WITH IRREGULAR SHAPE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2020 | Malfunction | covidien lp llc north haven | 22-01-2021 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING MAGEN RESECTION, THE KNIFE STOPPED AFTER ADVANCING ABOUT 20 MM. THE FIRING WAS UNABLE TO BE PERFORMED AGAIN AFTER WAITING FOR A WHILE, SO A NEW PRODUCT WAS USED AND RESECTION WAS PERFORMED.
18-12-2020 | Malfunction | covidien lp llc north haven | 22-01-2021 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING MAGEN RESECTION, THE KNIFE STOPPED AFTER ADVANCING ABOUT 20 MM. THE FIRING WAS UNABLE TO BE PERFORMED AGAIN AFTER WAITING FOR A WHILE, SO A NEW PRODUCT WAS USED AND RESECTION WAS PERFORMED.
31-12-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE THAT THE STAPLER FIRED FINE, BUT THE BLADE WOULD NOT RETRACT. UNKNOWN AS TO HOW THE DEVICE WAS REMOVED FROM THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO DEVICE WILL BE RETURNED.
31-12-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE THAT THE STAPLER FIRED FINE, BUT THE BLADE WOULD NOT RETRACT. UNKNOWN AS TO HOW THE DEVICE WAS REMOVED FROM THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO DEVICE WILL BE RETURNED.
01-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | ENDO GIA ULTRA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAFF AND SURGEON WERE NOT ABLE TO GET THE STAPLER AND RELOAD TO FIRE. SEVERAL ATTEMPTS WERE MADE TO UNLOAD AND LOAD THE RELOAD WITH NO SUCCESS. THE SURGEON ROTATED THE SHAFT AND CLOSED THE JAWS ON THE TISSUE. IT WAS EXPLAINED THAT THE SURGEON HAD TO HOLD THE HANDLE OF THE STAPLER CLOSED TO KEEP THE JAWS SHUT ON THE TISSUE. THEN THE SURGEON ASKED THE TECH TO PUSH THE GREEN FIRING BUTTON. WHEN THE BUTTON WAS PUSHED, IT RETRACTED INTO THE STAPLER HANDLE SO ONLY HALF OF THE BUTTON WAS VISIBLE. IT WAS ALSO NOTED THAT THE SURGEON WAS STILL HOLDING THE HANDLE IN THE SHUT POSITION. IT WAS MENTIONED THAT IF THE SURGEON RELEASES THE HANDLE THE JAWS WOULD OPEN. THEY WOULD NOT STAY CLOSED ON THE TISSUE. ANOTHER STAPLER WAS OPENED AND CONTINUED TO HAVE ISSUE TRYING TO FIRE. THE SURGEON ENDED UP USING ENDO LOOP IN PLACE OF THE STAPLER. IT WAS ALSO NOTED THAT THE RELOAD WAS NOT ABLE TO BE USED WITH A DIFFERENT STAPLER. THERE WAS NO PATIENT INJURY.
01-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAFF AND SURGEON WERE NOT ABLE TO GET THE STAPLER AND RELOAD TO FIRE. SEVERAL ATTEMPTS WERE MADE TO UNLOAD AND LOAD THE RELOAD WITH NO SUCCESS. THE SURGEON ROTATED THE SHAFT AND CLOSED THE JAWS ON THE TISSUE. IT WAS EXPLAINED THAT THE SURGEON HAD TO HOLD THE HANDLE OF THE STAPLER CLOSED TO KEEP THE JAWS SHUT ON THE TISSUE. THEN THE SURGEON ASKED THE TECH TO PUSH THE GREEN FIRING BUTTON. WHEN THE BUTTON WAS PUSHED, IT RETRACTED INTO THE STAPLER HANDLE SO ONLY HALF OF THE BUTTON WAS VISIBLE. IT WAS ALSO NOTED THAT THE SURGEON WAS STILL HOLDING THE HANDLE IN THE SHUT POSITION. IT WAS MENTIONED THAT IF THE SURGEON RELEASES THE HANDLE THE JAWS WOULD OPEN. THEY WOULD NOT STAY CLOSED ON THE TISSUE. ANOTHER STAPLER WAS OPENED AND CONTINUED TO HAVE ISSUE TRYING TO FIRE. THE SURGEON ENDED UP USING ENDO LOOP IN PLACE OF THE STAPLER. IT WAS ALSO NOTED THAT THE RELOAD WAS NOT ABLE TO BE USED WITH A DIFFERENT STAPLER. THERE WAS NO PATIENT INJURY.
25-11-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | 75 MM NTLC SELECTABLE RELOAD | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, A PART OF THE CARTRIDGE WAS BROKEN OFF BEFORE USE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
04-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DETACHMENT OF THE RIGHT COLON INTESTINE IN A LAPAROSCOPIC RADICAL RESECTION OF THE RIGHT COLON, THE SURGEON TRIED TO FIRE THE DEVICE, BUT THE GRIP COULD NOT BE FIRED. LATER, AFTER CHECKING THE RELOAD, SEVERAL STAPLES HAD EMERGED FROM THE INSIDE OF THE RELOAD. THE DEVICE WAS REPLACED WITH ANOTHER MANUFACTURER'S DEVICE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING PULMONARY LOBE TISSUE ON THE FIRST FIRING, AFTER FIRED, NOTED ALL THE STAPLES WERE MALFORMED WITH IRREGULAR SHAPE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, AFTER LOADING THE RELOAD TO THE POWER HANDLE, THE STAPLER WAS ATTEMPTED TO BE FIRED, BUT IT DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | us surgical puerto rico | 22-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING ENTERECTOMY PROCEDURE WHEN THE SURGEON ATTEMPTED TO FIRE, THERE WAS NO PROGRESSION OF THE TRIGGER FOR CLAMP LINE SHOOTING. THERE WAS NO PATIENT INJURY.
03-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING INTESTINAL STRAP ON AN OPEN RE-LAPAROTOMY, THE TRIGGER DID NOT SHOOT. ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, THE DEVICE WAS DIFFICULT TO LOAD, AND THE LATCH PART OF THE DEVICE BROKE OFF WHILE THE SURGEON TRYING TO ARTICULATE THE DEVICE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, THE DEVICE WAS DIFFICULT TO LOAD, AND THE LATCH PART OF THE DEVICE BROKE OFF WHILE THE SURGEON TRYING TO ARTICULATE THE DEVICE. THERE WAS NO PATIENT INJURY.
02-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST FIRING ON LUNG PARENCHYMA IN A THORACOSCOPIC LOBECTOMY, THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON ATTACHED THE RELOAD TO ANOTHER HANDLE AND ADAPTER AND IT FIRED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
23-10-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC LOBECTOMY SURGERY, COULD NOT FIRE WHEN SEVERING LUNG LOBE TISSUE ON THE 1ST FIRING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, BEFORE IT WAS USED ON THE PATIENT IN A PARTIAL GASTRECTOMY, THE CARTRIDGE WAS CONNECTED TO THE HANDLE, BUT IT COULD NOT BE SQUEEZED WHEN CHECKING THE OPENING AND CLOSING. THE HANDLE WAS REPLACED. THERE WAS NO PATIENT INVOLVEMENT.
04-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, BEFORE IT WAS USED ON THE PATIENT IN A PARTIAL GASTRECTOMY, THE CARTRIDGE WAS CONNECTED TO THE HANDLE, BUT IT COULD NOT BE SQUEEZED WHEN CHECKING THE OPENING AND CLOSING. THERE WAS NO PATIENT INVOLVEMENT.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | EEA | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN TOTAL GASTRECTOMY, AFTER MATING THE ORVIL TO THE TROCAR SPIKE, THE ORVIL WOULD CONTINUALLY POP OFF DURING CLOSURE OF DEVICE. THE ORVIL WOULD NOT STAY MATED. A NEW STAPLER AND ORVIL WERE OPENED AND USED TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | EEA ORVIL | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN TOTAL GASTRECTOMY, AFTER MATING THE ORVIL TO THE TROCAR SPIKE, THE ORVIL WOULD CONTINUALLY POP OFF DURING CLOSURE OF DEVICE. THE ORVIL WOULD NOT STAY MATED. A NEW STAPLER AND ORVIL WERE OPENED AND USED TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
01-07-2020 | Malfunction | covidien lp llc north haven | 22-01-2021 | RELIATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING FIXATION ON A LAPAROSCOPIC COELIOSCOPY, THE TACK DISENGAGED INTO THE PATIENT¿S CAVITY AND WAS RETRIEVED WITH A DEVICE. TWO DEVICES HAD THE SAME MALFUNCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VATS, DEVICE JAMMED. OPENED A NEW DEVICE TO CONTINUE SURGERY. THERE WERE NO PATIENT CONSEQUENCES.
02-01-2021 | Malfunction | covidien lp llc north haven | 22-01-2021 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST FIRING ON LUNG PARENCHYMA IN A THORACOSCOPIC LOBECTOMY, THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON ATTACHED THE RELOAD TO ANOTHER HANDLE AND ADAPTER AND IT FIRED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
03-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTROPLASTY SURGERY, USING THE SLEEVE TECHNIQUE, AT THE END OF THE SURGERY, AS USUAL, THE STOMACH IS REMOVED AND A TEST IS PERFORMED TO CHECK THE STAPLE LINE AND THE STOMACH IS INFLATED WITH CO2 UNTIL IT IS COMPLETELY FILLED WITH GAS. IT WAS VERIFIED THAT THE STOMACH DID NOT FILL AND SOON AN ORIFICE (HOLE) WAS FOUND IN THE STAPLE LINE. PART OF THE STAPLE LINE WAS INTACT AT THE BEGINNING, JUST AHEAD OF THE HOLE. DUE TO THE HEIGHT OF THE STAPLING, THE LOAD THAT HAD A DEFECT IS THE BLUE GST60B, BUT IN ADDITION, ONE MORE BLUE LOAD WAS USED, BECAUSE THE STOMACH WAS LARGER THAN NORMAL. IT WAS NOT POSSIBLE TO KNOW WHETHER THE CHARGE THAT SUPPOSEDLY FAILED WAS THE SPARE OR THE LAST CHARGE IN THE KIT. THE RELOADS HAVE THE SAME LOT NUMBER. THERE WAS NO DELAY TO SURGERY, SURGERY WAS SUCCESSFULLY COMPLETED, AND THERE WAS NO PATIENT CONSEQUENCE.
02-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC APPENDECTOMY, WHILE TRANSECTING THE APPENDIX, THE SURGEON PLACED THE RELOAD ON THE TARGETED TISSUE, SQUEEZED THE HANDLE TO CLOSE LOAD AND THEN PUSHED THE GREEN BUTTON, BUT THE HANDLE WOULD NOT ENGAGE INTO FIRING MODE. THE SURGEON OPENED THE JAWS OF THE DEVICE, REMOVED FROM THE PATIENT, AND ATTEMPTED TO REPLICATE THE FIRING ACTIONS OUTSIDE WITHOUT TISSUE IN THE JAWS. THE HANDLE WORKED CORRECTLY SO THE DEVICE WAS THEN REINSERTED INTO THE TROCAR AND PLACED ONTO THE TARGETED TISSUE. THE SURGEON SQUEEZED THE HANDLE TO CLOSE THE LOAD, AND ATTEMPTED TO PUSH THE GREEN BUTTON TO ENGAGE THE FIRING MECHANISM BUT THE SAME ISSUE OCCURRED. ANOTHER HANDLE AND RELOAD WERE OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
03-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTROPLASTY SURGERY, USING THE SLEEVE TECHNIQUE, AT THE END OF THE SURGERY, AS USUAL, THE STOMACH IS REMOVED AND A TEST IS PERFORMED TO CHECK THE STAPLE LINE AND THE STOMACH IS INFLATED WITH CO2 UNTIL IT IS COMPLETELY FILLED WITH GAS. IT WAS VERIFIED THAT THE STOMACH DID NOT FILL AND SOON AN ORIFICE (HOLE) WAS FOUND IN THE STAPLE LINE. PART OF THE STAPLE LINE WAS INTACT AT THE BEGINNING, JUST AHEAD OF THE HOLE. DUE TO THE HEIGHT OF THE STAPLING, THE LOAD THAT HAD A DEFECT IS THE BLUE GST60B, BUT IN ADDITION, ONE MORE BLUE LOAD WAS USED, BECAUSE THE STOMACH WAS LARGER THAN NORMAL. IT WAS NOT POSSIBLE TO KNOW WHETHER THE CHARGE THAT SUPPOSEDLY FAILED WAS THE SPARE OR THE LAST CHARGE IN THE KIT. THE RELOADS HAVE THE SAME LOT NUMBER. THERE WAS NO DELAY TO SURGERY, SURGERY WAS SUCCESSFULLY COMPLETED, AND THERE WAS NO PATIENT CONSEQUENCE.
02-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC APPENDECTOMY, WHILE TRANSECTING THE APPENDIX, THE SURGEON PLACED THE RELOAD ON THE TARGETED TISSUE, SQUEEZED THE HANDLE TO CLOSE LOAD AND THEN PUSHED THE GREEN BUTTON, BUT THE HANDLE WOULD NOT ENGAGE INTO FIRING MODE. THE SURGEON OPENED THE JAWS OF THE DEVICE, REMOVED FROM THE PATIENT, AND ATTEMPTED TO REPLICATE THE FIRING ACTIONS OUTSIDE WITHOUT TISSUE IN THE JAWS. THE HANDLE WORKED CORRECTLY SO THE DEVICE WAS THEN REINSERTED INTO THE TROCAR AND PLACED ONTO THE TARGETED TISSUE. THE SURGEON SQUEEZED THE HANDLE TO CLOSE THE LOAD, AND ATTEMPTED TO PUSH THE GREEN BUTTON TO ENGAGE THE FIRING MECHANISM BUT THE SAME ISSUE OCCURRED. ANOTHER HANDLE AND RELOAD WERE OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
28-12-2020 | Malfunction | covidien lp llc north haven | 22-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC PROCTECTOMY, WHEN APPROACHING TISSUE AND RESECTING THE RECTUM, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD AND THE SAME HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2020 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN REPAIR OF GASTROINTESTINAL ANASTOMOSIS, DURING STOMACH STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE AND THE SAME RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, WHEN SEVERING LUNG TISSUE, THERE WERE THREE CARTRIDGE USED (ONE ECR60G, ECR45D, AND ECR45G), AND AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ANVIL CLAMPING MECHANISM WAS DEFORMED. FUNCTIONAL EVALUATION OF THE LOADING UNIT COULD NOT BE PERFORMED DUE TO THE NOTED DAMAGE.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ANVIL CLAMPING MECHANISM WAS DEFORMED. FUNCTIONAL EVALUATION OF THE LOADING UNIT COULD NOT BE PERFORMED DUE TO THE NOTED DAMAGE.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Component or Accessory Incompatibility | A1701 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ANVIL CLAMPING MECHANISM WAS DEFORMED. FUNCTIONAL EVALUATION OF THE LOADING UNIT COULD NOT BE PERFORMED DUE TO THE NOTED DAMAGE.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ANVIL CLAMPING MECHANISM WAS DEFORMED. FUNCTIONAL EVALUATION OF THE LOADING UNIT COULD NOT BE PERFORMED DUE TO THE NOTED DAMAGE.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ANVIL CLAMPING MECHANISM WAS DEFORMED. FUNCTIONAL EVALUATION OF THE LOADING UNIT COULD NOT BE PERFORMED DUE TO THE NOTED DAMAGE.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ANVIL CLAMPING MECHANISM WAS DEFORMED. FUNCTIONAL EVALUATION OF THE LOADING UNIT COULD NOT BE PERFORMED DUE TO THE NOTED DAMAGE.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ANVIL CLAMPING MECHANISM WAS DEFORMED. FUNCTIONAL EVALUATION OF THE LOADING UNIT COULD NOT BE PERFORMED DUE TO THE NOTED DAMAGE.
05-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS DIFFICULTY LOADING AND THE OPENING AND CLOSING OF THE RELOAD WAS NOT POSSIBLE. THE HANDLE WAS REPLACED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ANVIL CLAMPING MECHANISM WAS DEFORMED. FUNCTIONAL EVALUATION OF THE LOADING UNIT COULD NOT BE PERFORMED DUE TO THE NOTED DAMAGE.
04-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PYLORUS PRESERVING PANCREATODUODECTOMY, ON DUODENUM TRANSECTION, AFTER CLAMPING THE ORGAN, THE DEVICE DID NOT FIRE, THE USER WAS NOT ABLE TO PUSH THE BLACK KNOB. AFTER SURGERY, THE OPERATING ROOM NURSE TRIED TO PUSH THE BLACK KNOB OUT OF OPERATION FIELD, AND COULD NOT BE MOVED TO THE END. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2021 | Malfunction | us surgical puerto rico | 22-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PYLORUS PRESERVING PANCREATODUODECTOMY, ON DUODENUM TRANSECTION, AFTER CLAMPING THE ORGAN, THE DEVICE DID NOT FIRE, THE USER WAS NOT ABLE TO PUSH THE BLACK KNOB. AFTER SURGERY, THE OPERATING ROOM NURSE TRIED TO PUSH THE BLACK KNOB OUT OF OPERATION FIELD, AND COULD NOT BE MOVED TO THE END. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
31-08-2020 | Malfunction | covidien lp llc north haven | 22-01-2021 | UNKNOWN EGIA TRI STAPLE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LUNG PROCEDURE, THE SURGEON WAS ABLE FULLY SQUEEZE THE STAPLER BUT THE STAPLE LINE WAS INCOMPLETE DISTALLY, RESULTING IN AN INCOMPLETE CLOSURE OF TRACHEA. SOME STAPLES WERE ALSO NOT COMPLETELY CLOSED. THE TISSUE WAS CLAMPED WITH AN INSTRUMENT AND ANOTHER RELOAD WAS USED TO FIX THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, THE SPRING FROM THE FIRING TRIGGER CAME AWAY FROM THE POWERED ECHELON FLEX WHEN IT WAS BEING RELOADED. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
27-11-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | CONTOUR CURVED CUTTER STAPLER | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, WHEN THE SURGEON USED THE DEVICE IT DID NOT FIRE. THERE WERE NO PATIENT CONSEQUENCES.
01-01-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE CANNOT FIRE. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
04-11-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | ECHELON ENDOPATH STAPLER | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WAS DEHISCENCE IN THE POUCH STAPLING LINE, IN THE SAME PLACE WHERE AT THE TIME OF SURGERY THERE WAS A SLIGHT BLEEDING.
28-12-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, A LOUD 'BREAKING' CRUNCH WAS HEARD UPON CLOSURE OF STAPLER HANDLE BEFORE FIRING. A YELLOW-LOOKING PLASTIC WAS SEEN BROKEN WITHIN THE STAPLER HANDLE. THE SURGEON DID NOT FEEL A TENSION IN FORCE DURING CLOSURE. EVERYTHING FELT NORMAL. THERE WAS NO DELAY TO PROCEDURE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
28-12-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, A LOUD 'BREAKING' CRUNCH WAS HEARD UPON CLOSURE OF STAPLER HANDLE BEFORE FIRING. A YELLOW-LOOKING PLASTIC WAS SEEN BROKEN WITHIN THE STAPLER HANDLE. THE SURGEON DID NOT FEEL A TENSION IN FORCE DURING CLOSURE. EVERYTHING FELT NORMAL. THERE WAS NO DELAY TO PROCEDURE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE.
07-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC SURGERY FOR LUNG CANCER, WHEN SEVERING VESSELS TISSUE, AFTER FIRED, NOTED STAPLES WERE MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROTOMY WITH PANCREATECTOMY/ADRENALECTOMY PROCEDURE, WHILE FIRING THE PVE35A STAPLER ON A VEIN (THIN TISSUE), THE STAPLER DEPLOYED THE STAPLES AS INTENDED. THE STAPLER SLED THEN FAILED TO RETURN TO THE REGULAR POSITION. THE SURGEON STATED IT DEPLOYED STAPLES PROPERLY BUT CEASED TO RETURN WHEN HE LET GO OF THE FIRING TRIGGER. HE STATED HE HELD THE TRIGGER SEVERAL SECONDS AFTER TRANSECTION TO KEEP THE SLED AT THE DISTAL TIP. AS IT WOULDN'T RETURN, THEY COULDN'T OPEN THE STAPLER. THEY FIRST ATTEMPTED TO USE THE KNIFE RETURN SWITCH BUT IT DID NOT WORK. THEREFORE, THEY USED THE MANUAL OVERRIDE TO RETURN THE KNIFE BLADE TO THE STARTING POSITION. UPON INSPECTION, THEY SAID THE JAWS WERE CLEAR AND NOTED IT WAS A FIRING ON THIN TISSUE. STAPLER WAS OPENED AFTER USING MANUAL OVERRIDE. THEY FINISHED THE CASE WITH THE PCEE60A THAT WAS ALREADY OPENED. THERE WERE NO ADVERSE CONSEQUENCE FOR THE PATIENT.
07-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROTOMY WITH PANCREATECTOMY/ADRENALECTOMY PROCEDURE, WHILE FIRING THE PVE35A STAPLER ON A VEIN (THIN TISSUE), THE STAPLER DEPLOYED THE STAPLES AS INTENDED. THE STAPLER SLED THEN FAILED TO RETURN TO THE REGULAR POSITION. THE SURGEON STATED IT DEPLOYED STAPLES PROPERLY BUT CEASED TO RETURN WHEN HE LET GO OF THE FIRING TRIGGER. HE STATED HE HELD THE TRIGGER SEVERAL SECONDS AFTER TRANSECTION TO KEEP THE SLED AT THE DISTAL TIP. AS IT WOULDN'T RETURN, THEY COULDN'T OPEN THE STAPLER. THEY FIRST ATTEMPTED TO USE THE KNIFE RETURN SWITCH BUT IT DID NOT WORK. THEREFORE, THEY USED THE MANUAL OVERRIDE TO RETURN THE KNIFE BLADE TO THE STARTING POSITION. UPON INSPECTION, THEY SAID THE JAWS WERE CLEAR AND NOTED IT WAS A FIRING ON THIN TISSUE. STAPLER WAS OPENED AFTER USING MANUAL OVERRIDE. THEY FINISHED THE CASE WITH THE PCEE60A THAT WAS ALREADY OPENED. THERE WERE NO ADVERSE CONSEQUENCE FOR THE PATIENT.
07-01-2021 | Malfunction | ethicon endo-surgery | 22-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROTOMY WITH PANCREATECTOMY/ADRENALECTOMY PROCEDURE, WHILE FIRING THE PVE35A STAPLER ON A VEIN (THIN TISSUE), THE STAPLER DEPLOYED THE STAPLES AS INTENDED. THE STAPLER SLED THEN FAILED TO RETURN TO THE REGULAR POSITION. THE SURGEON STATED IT DEPLOYED STAPLES PROPERLY BUT CEASED TO RETURN WHEN HE LET GO OF THE FIRING TRIGGER. HE STATED HE HELD THE TRIGGER SEVERAL SECONDS AFTER TRANSECTION TO KEEP THE SLED AT THE DISTAL TIP. AS IT WOULDN'T RETURN, THEY COULDN'T OPEN THE STAPLER. THEY FIRST ATTEMPTED TO USE THE KNIFE RETURN SWITCH BUT IT DID NOT WORK. THEREFORE, THEY USED THE MANUAL OVERRIDE TO RETURN THE KNIFE BLADE TO THE STARTING POSITION. UPON INSPECTION, THEY SAID THE JAWS WERE CLEAR AND NOTED IT WAS A FIRING ON THIN TISSUE. STAPLER WAS OPENED AFTER USING MANUAL OVERRIDE. THEY FINISHED THE CASE WITH THE PCEE60A THAT WAS ALREADY OPENED. THERE WERE NO ADVERSE CONSEQUENCE FOR THE PATIENT.
02-12-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE TLDG SURGERY, WHEN SEVERING STOMACH TISSUE ON THE SEVENTH FIRING, AFTER FIRED, NOTED SOME STAPLES WERE MALFORMED WITH IRREGULAR SHAPE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-12-2020 | Malfunction | ethicon endo-surgery | 22-01-2021 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, WHEN SEVERING STOMACH TISSUE ON THE THIRD FIRING, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. CHANGED TO NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-12-2020 | Injury | ethicon endo-surgery | 22-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LEFT ATRIAL APPENDECTOMY, MASSIVE BLEEDING OCCURRED WHEN THE PSEE60A LOADING THE GST60G COMPRESSED THE TARGET TISSUE FOR 10 SECONDS AND FIRED AT THE 1ST FIRING ON THE LEFT ATRIAL APPENDAGE. WHEN THE DEVICE WAS CHECKED, IT WAS FOUND THE FORMED STAPLES AT THE THREE ROWS ON A SIDE OF THE PROXIMAL END OF THE CARTRIDGE WERE NOT DEPLOYED BUT REMAINED ON THE CARTRIDGE. THE ARTIFICIAL HEART-LUNG MACHINE WAS USED ALTHOUGH IT WAS NOT SUPPOSED TO BE USED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE BLEEDING WAS STOPPED BY THE COMPRESSION WITH HAND AND THE ADDITIONAL SUTURE. THERE WAS NO PROBLEM FOR THE PATIENT BECAUSE THE BLEEDING WAS STOPPED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
25-12-2020 | Injury | ethicon endo-surgery | 22-01-2021 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LEFT ATRIAL APPENDECTOMY, MASSIVE BLEEDING OCCURRED WHEN THE PSEE60A LOADING THE GST60G COMPRESSED THE TARGET TISSUE FOR 10 SECONDS AND FIRED AT THE 1ST FIRING ON THE LEFT ATRIAL APPENDAGE. WHEN THE DEVICE WAS CHECKED, IT WAS FOUND THE FORMED STAPLES AT THE THREE ROWS ON A SIDE OF THE PROXIMAL END OF THE CARTRIDGE WERE NOT DEPLOYED BUT REMAINED ON THE CARTRIDGE. THE ARTIFICIAL HEART-LUNG MACHINE WAS USED ALTHOUGH IT WAS NOT SUPPOSED TO BE USED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE BLEEDING WAS STOPPED BY THE COMPRESSION WITH HAND AND THE ADDITIONAL SUTURE. THERE WAS NO PROBLEM FOR THE PATIENT BECAUSE THE BLEEDING WAS STOPPED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
14-10-2020 | Malfunction | us surgical puerto rico | 22-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE, ON THE FIRST FIRING, THE GREEN BUTTON OF THE STAPLER WAS ABLE TO BE PRESSED BUT THE HANDLE WAS NOT ABLE TO BE SQUEEZED. THE FEW FIRST PINS WERE POPPED OUT. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2021 | Injury | davol inc sub c.r bard inc -1213643 | 25-01-2021 | CAPSURE STRAIGHT | Patient Device Interaction Problem | A01 | Unspecified Infection | Event Description: IT WAS REPORTED THAT, DURING A LAPAROSCOPIC PROCEDURE ON (B)(6) 2020, THE PATIENT WAS IMPLANTED WITH A BARD/DAVOL 3DMAX MESH FIXATED BY A BARD/DAVOL CAPSURE STRAIGHT FIXATION DEVICE. THE PATIENT WAS FINE AFTER THE SURGERY. ONE WEEK POST-IMPLANT, THE PATIENT HAD COMPLAINTS OF PAIN AND WAS PRESCRIBED ANALGESICS BY THE SURGEON. THE PATIENT RETURNED ON (B)(6) 2021 WITH PAIN AND SWELLING; THE SURGEON SUSPECTED INFECTION WHICH WAS CONFIRMED BY AN ULTRA-SONOGRAM. THE PATIENT WAS PRESCRIBED AN ORAL ANTIBIOTIC AUGMENTIN 650MG TWICE A DAY AND IS RESPONDING WELL. AS REPORTED, MESH REMOVAL MAY OR MAY NOT BE REQUIRED BASED ON THE PATIENT'S RESPONSE.
12-01-2021 | Injury | davol inc sub c.r bard inc -1213643 | 25-01-2021 | CAPSURE STRAIGHT | Patient Device Interaction Problem | A01 | Pain | Event Description: IT WAS REPORTED THAT, DURING A LAPAROSCOPIC PROCEDURE ON (B)(6) 2020, THE PATIENT WAS IMPLANTED WITH A BARD/DAVOL 3DMAX MESH FIXATED BY A BARD/DAVOL CAPSURE STRAIGHT FIXATION DEVICE. THE PATIENT WAS FINE AFTER THE SURGERY. ONE WEEK POST-IMPLANT, THE PATIENT HAD COMPLAINTS OF PAIN AND WAS PRESCRIBED ANALGESICS BY THE SURGEON. THE PATIENT RETURNED ON (B)(6) 2021 WITH PAIN AND SWELLING; THE SURGEON SUSPECTED INFECTION WHICH WAS CONFIRMED BY AN ULTRA-SONOGRAM. THE PATIENT WAS PRESCRIBED AN ORAL ANTIBIOTIC AUGMENTIN 650MG TWICE A DAY AND IS RESPONDING WELL. AS REPORTED, MESH REMOVAL MAY OR MAY NOT BE REQUIRED BASED ON THE PATIENT'S RESPONSE.
12-01-2021 | Injury | davol inc sub c.r bard inc -1213643 | 25-01-2021 | CAPSURE STRAIGHT | Patient Device Interaction Problem | A01 | Swelling | Event Description: IT WAS REPORTED THAT, DURING A LAPAROSCOPIC PROCEDURE ON (B)(6) 2020, THE PATIENT WAS IMPLANTED WITH A BARD/DAVOL 3DMAX MESH FIXATED BY A BARD/DAVOL CAPSURE STRAIGHT FIXATION DEVICE. THE PATIENT WAS FINE AFTER THE SURGERY. ONE WEEK POST-IMPLANT, THE PATIENT HAD COMPLAINTS OF PAIN AND WAS PRESCRIBED ANALGESICS BY THE SURGEON. THE PATIENT RETURNED ON (B)(6) 2021 WITH PAIN AND SWELLING; THE SURGEON SUSPECTED INFECTION WHICH WAS CONFIRMED BY AN ULTRA-SONOGRAM. THE PATIENT WAS PRESCRIBED AN ORAL ANTIBIOTIC AUGMENTIN 650MG TWICE A DAY AND IS RESPONDING WELL. AS REPORTED, MESH REMOVAL MAY OR MAY NOT BE REQUIRED BASED ON THE PATIENT'S RESPONSE.
05-01-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW RECTAL CANCER RESECTION, WHEN THE DISTAL RECTUM WAS DETACHED, THE 45 PURPLE STAPLE WAS USED TO DETACH FOR THE FIRST TIME BUT THE STAPLES COULD NOT BE FIRED. IT WAS THOUGHT THAT IT WAS BECAUSE THE TISSUE WAS TOO THICK SO THE PURPLE RELOAD WAS REPLACED WITH THE GREEN STAPLES, BUT THEY COULD NOT BE FIRED, SO IT WAS REPLACED WITH ANOTHER 45 PURPLE STAPLE AND IT SUCCESSFULLY FIRED. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW RECTAL CANCER RESECTION, WHEN THE DISTAL RECTUM WAS DETACHED, THE 45 PURPLE STAPLE WAS USED TO DETACH FOR THE FIRST TIME BUT THE STAPLES COULD NOT BE FIRED. IT WAS THOUGHT THAT IT WAS BECAUSE THE TISSUE WAS TOO THICK SO THE PURPLE RELOAD WAS REPLACED WITH THE GREEN STAPLES, BUT THEY COULD NOT BE FIRED, SO IT WAS REPLACED WITH ANOTHER 45 PURPLE STAPLE AND IT SUCCESSFULLY FIRED. THERE WAS NO PATIENT INJURY.
05-01-2021 | Injury | us surgical puerto rico | 25-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC TOTAL GASTRECTOMY, THE DEVICE WAS ABLE TO FIRE BUT COULD ONLY PUSH A LITTLE AND UNABLE TO COMPLETE THE CLOSURE, RESULTING TO AN INCOMPLETE STAPLE LINE DISTALLY. IN ADDITION THERE WERE POOR STAPLE FORMATION. TO RESOLVE THE ISSUE THE SURGEON USED SUTURE AND HAND SEW THE TISSUE.
05-01-2021 | Injury | us surgical puerto rico | 25-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC TOTAL GASTRECTOMY, THE DEVICE WAS ABLE TO FIRE BUT COULD ONLY PUSH A LITTLE AND UNABLE TO COMPLETE THE CLOSURE, RESULTING TO AN INCOMPLETE STAPLE LINE DISTALLY. IN ADDITION THERE WERE POOR STAPLE FORMATION. TO RESOLVE THE ISSUE THE SURGEON USED SUTURE AND HAND SEW THE TISSUE.
05-01-2021 | Injury | us surgical puerto rico | 25-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC TOTAL GASTRECTOMY, THE DEVICE WAS ABLE TO FIRE BUT COULD ONLY PUSH A LITTLE AND UNABLE TO COMPLETE THE CLOSURE, RESULTING TO AN INCOMPLETE STAPLE LINE DISTALLY. IN ADDITION THERE WERE POOR STAPLE FORMATION. TO RESOLVE THE ISSUE THE SURGEON USED SUTURE AND HAND SEW THE TISSUE.
06-01-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN DISTAL GASTRECTOMY, PRIOR TO DUODENAL RESECTION, AFTER INSERTING THE HANDLE INTO THE POWER SHELL, THE DISPLAY TURNED OFF AND RESTARTED. THE POWER SHELL DISPLAYED AS A USED SHELL, MAKING IT UNUSABLE. A NEW POWER SHELL WAS OPENED AND USED, BUT THE HANDLE DISPLAY DISAPPEARED AGAIN AND RESTARTED. ANOTHER DEVICE FROM ANOTHER MANUFACTURER USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN DISTAL GASTRECTOMY, PRIOR TO DUODENAL RESECTION, AFTER INSERTING THE HANDLE INTO THE POWER SHELL, THE DISPLAY TURNED OFF AND RESTARTED. THE POWER SHELL DISPLAYED AS A USED SHELL, MAKING IT UNUSABLE. A NEW POWER SHELL WAS OPENED AND USED, BUT THE HANDLE DISPLAY DISAPPEARED AGAIN AND RESTARTED. ANOTHER DEVICE FROM ANOTHER MANUFACTURER USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-03-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD DID NOT CLOSE FIRMLY FOR STAPLING OF THE TISSUE. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
11-03-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD DID NOT CLOSE FIRMLY FOR STAPLING OF THE TISSUE. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien lp llc north haven | 25-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON STAPLING OF THE STOMACH, THE RELOAD DID NOT COMPLETE THE FIRING, THE REMAINING UNFIRED RELOAD HAD UNFORMED STAPLE FORMATION. THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE, ADAPTER AND SHELL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2020 | Malfunction | covidien lp llc north haven | 25-01-2021 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON STAPLING OF THE STOMACH, THE RELOAD DID NOT COMPLETE THE FIRING, THE REMAINING UNFIRED RELOAD HAD UNFORMED STAPLE FORMATION. THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE, ADAPTER AND SHELL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-11-2020 | Malfunction | us surgical puerto rico | 25-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER THE RELOAD WAS CLOSED, SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. SURGEON REPLACED THE HANDLE AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY.
01-11-2020 | Malfunction | ethicon endo-surgery | 25-01-2021 | PROXIMATE*75MM LINEAR CUTTER | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: SEE ATTACHED USER FACILITY MEDWATCH.
01-11-2020 | Malfunction | ethicon endo-surgery | 25-01-2021 | PROXIMATE*75MM LINEAR CUTTER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: SEE ATTACHED USER FACILITY MEDWATCH.
01-11-2020 | Malfunction | ethicon endo-surgery | 25-01-2021 | PROXIMATE*75MM LINEAR CUTTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: SEE ATTACHED USER FACILITY MEDWATCH.
10-12-2020 | Malfunction | ethicon endo-surgery | 25-01-2021 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE FORCEPS GOT STUCK IN THE CLOSED POSITION AFTER SHOWING SIGNS OF WEAKNESS FOR THE BATTERY WHEN STAPLING. THE MANUAL OPENING WAS NOT POSSIBLE. THERE ARE CONSEQUENCES ON THE PATIENT. THE LUNG PARENCHYMA GOT STUCK IN THE FORCEPS INDUCING A TEAR. SPARE DEVICE USED TO FINISH THE PROCEDURE. DELAY REPORTED, NO PRECISION GIVEN TO KNOW HOW MANY MINUTES BUT THIS ISSUE CAUSE AN EXTENSION TIME OF THE PROCEDURE, AND A BREAK OF THE PROCEDURE FOR THE SURGEON AND THE MEDICAL STAFFS. THE EVENT IS CONSIDERED AS ISOLATED AND MINOR CASE. NO MORE INFORMATION AVAILABLE.
05-01-2021 | Injury | covidien lp llc north haven | 25-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY, WHEN THE SURGEON FIRE THE DEVICE, THE KNIFE BLADED ACTIVATED HOWEVER THE FIRING FAILED. IN ADDITION THERE WERE POOR STAPLE FORMATION IT HAPPENED ON TWO RELOADS. THERE WERE TISSUE DAMAGE AS A RESULT OF DEVICE MALFUNCTION. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
05-01-2021 | Injury | covidien lp llc north haven | 25-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY, WHEN THE SURGEON FIRE THE DEVICE, THE KNIFE BLADED ACTIVATED HOWEVER THE FIRING FAILED. IN ADDITION THERE WERE POOR STAPLE FORMATION IT HAPPENED ON TWO RELOADS. THERE WERE TISSUE DAMAGE AS A RESULT OF DEVICE MALFUNCTION. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
05-01-2021 | Injury | us surgical puerto rico | 25-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY, WHEN THE SURGEON FIRE THE DEVICE, THE KNIFE BLADED ACTIVATED HOWEVER THE FIRING FAILED. IN ADDITION THERE WERE POOR STAPLE FORMATION IT HAPPENED ON TWO RELOADS. THERE WERE TISSUE DAMAGE AS A RESULT OF DEVICE MALFUNCTION. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
11-03-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE ABLE TO BE FIE HOWEVER IT STOPPED HALFWAY AND PARTIALLY FIRED.
25-01-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | UNKNOWN STAPLING DEVICE | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE ABLE TO BE "FIE" HOWEVER IT STOPPED HALFWAY AND PARTIALLY FIRED.
25-01-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | UNKNOWN EGIA ADAPTER | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE ABLE TO BE FIRE HOWEVER IT STOPPED HALFWAY AND PARTIALLY FIRED.
24-02-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | UNKNOWN STAPLING DEVICE | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD DID NOT CLOSE FIRMLY FOR STAPLING OF THE TISSUE.
24-02-2021 | Malfunction | covidien lp llc north haven | 25-01-2021 | UNKNOWN EGIA ADAPTER | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD DID NOT CLOSE FIRMLY FOR STAPLING OF THE TISSUE.
05-01-2021 | Injury | covidien lp llc north haven | 25-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, ON CLAMPING THE STOMACH, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, THERE WAS POOR STAPLE FORMATION, THE LOAD CUT THE TISSUE BUT DID NOT STAPLE. THERE WAS TISSUE DAMAGE. THE SURGICAL TIME WAS EXTENDED BY 2 HOURS DUE TO THE PRODUCT PROBLEM. THERE WAS BLOOD LOSS OF 500CC OR MORE. THE DOCTOR REQUESTED A DRAIN TO MONITOR THE BLEEDING AND DRAIN THE FLUIDS. THE SURGEON OPENED OTHER LOADS AND USED THE SAME HANDLE TO CONTAIN THE BLEEDING AND COMPLETE THE CASE.
05-01-2021 | Injury | covidien lp llc north haven | 25-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, ON CLAMPING THE STOMACH, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, THERE WAS POOR STAPLE FORMATION, THE LOAD CUT THE TISSUE BUT DID NOT STAPLE. THERE WAS TISSUE DAMAGE. THE SURGICAL TIME WAS EXTENDED BY 2 HOURS DUE TO THE PRODUCT PROBLEM. THERE WAS BLOOD LOSS OF 500CC OR MORE. THE DOCTOR REQUESTED A DRAIN TO MONITOR THE BLEEDING AND DRAIN THE FLUIDS. THE SURGEON OPENED OTHER LOADS AND USED THE SAME HANDLE TO CONTAIN THE BLEEDING AND COMPLETE THE CASE.
07-01-2021 | Malfunction | ethicon endo-surgery | 25-01-2021 | POWERED PLUS STANDARD 45 | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN PNEUMONECTOMY, THE DEVICE FUNCTIONED SPONTANEOUSLY ALTHOUGH IT WAS NOT TOUCHED AFTER THE CARTRIDGE WAS LOADED INTO THE JAW. THE JAWS COULD BE OPENED. THE CLOSING LEVER WAS STIFF TO PULL. THE DEVICE WAS USED ON THE BRONCHIAL TUB. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-01-2021 | Malfunction | ethicon endo-surgery | 25-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN THE KNIFE AUTOMATICALLY. USED RBRB TO RETURN THE KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-01-2021 | Malfunction | ethicon endo-surgery | 25-01-2021 | ETHICON EEA STAPLER | Therapeutic or Diagnostic Output Failure | A090406 | Insufficient Information | Event Description: DURING PROCEDURE STAPLER WOULD NOT WORK, REPLACED WITH ANOTHER ONE IT WORKED. FDA SAFETY REPORT ID # (B)(4).
19-12-2020 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY PROCEDURE, THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON USED ANOTHER HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | covidien lp llc north haven | 26-01-2021 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE THE SURGEON USED THE DEVICE IN AN IVOR LEWIS ESOPHAGECTOMY, BOTH RELOADS STOPPED FIRING PRIOR TO FULL LENGTH OF STAPLE CARTRIDGE BEING UTILIZED. ANOTHER RELOAD WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-07-2020 | Injury | covidien lp llc north haven | 26-01-2021 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS A POSSIBLE COMPLICATION ON THE DEVICE RELOAD. DUE TO THE DEVICE ISSUE, THE PATIENT HAD TO EXTEND HOSPITAL STAY FOR ABOUT 4 DAYS.
01-06-2020 | Injury | covidien lp llc north haven | 26-01-2021 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE CUSTOMER REPORTED A POTENTIAL COMPLICATION ATTRIBUTED TO THE DEVICE. IT WAS STATED THAT INTRA-OPERATIVELY, THE STAPLER WAS ABLE TO FIRE AND THERE WAS NO INCOMPLETE AND POOR STAPLE LINE. THE EVENT RESULTED TO PATIENT¿S EXTENDED HOSPITALIZATION FOR FOUR DAYS.
06-01-2021 | Malfunction | covidien lp llc north haven | 26-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF COLORECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. BOTH HANDLE AND RELOAD WERE CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | covidien lp llc north haven | 26-01-2021 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF COLORECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. BOTH HANDLE AND RELOAD WERE CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2020 | Malfunction | us surgical puerto rico | 26-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING RETOSIGMOIDECTOMY PROCEDURE, THE STAPLER DID NOT FIRE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, ON REMOVAL OF STOMACH TISSUE, THE STAPLER AND 2 RELOADS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE USER REPLACED WITH ANOTHER MANUFACTURER¿S PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | covidien lp llc north haven | 26-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, ON REMOVAL OF STOMACH TISSUE, THE STAPLER AND 2 RELOADS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE USER REPLACED WITH ANOTHER MANUFACTURER¿S PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-10-2020 | Malfunction | covidien lp llc north haven | 26-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE DID NOT ACTIVATE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10- MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE. IMPROVEMENTS HAVE BEEN INITIATED TO MITIGATE THIS CONDITION.
06-01-2021 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, ON REMOVAL OF THE LUNG TISSUE, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE USER REPLACED WITH ANOTHER MANUFACTURER¿S PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-11-2020 | Malfunction | covidien lp llc north haven | 26-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE DID NOT ACTIVATE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10- MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
06-01-2021 | Malfunction | covidien lp llc north haven | 26-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, ON REMOVAL OF THE LUNG TISSUE, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE USER REPLACED WITH ANOTHER MANUFACTURER¿S PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-01-2020 | Malfunction | ethicon endo-surgery | 26-01-2021 | ECHELON FLEX 45 COMPACT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS CANNOT BE OPENED AFTER FIRING. THERE WERE NO PATIENT CONSEQUENCES.
06-01-2021 | Malfunction | covidien lp llc north haven | 26-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF COLORECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. BOTH HANDLE AND RELOAD WERE CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF COLORECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. BOTH HANDLE AND RELOAD WERE CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF COLORECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. BOTH HANDLE AND RELOAD WERE CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF COLORECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. BOTH HANDLE AND RELOAD WERE CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF COLORECTAL CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. HENCE, THE STAPLER DID NOT FIRE. BOTH HANDLE AND RELOAD WERE CHANGED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-08-2020 | Malfunction | covidien lp llc north haven | 26-01-2021 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE HANDLE WAS MAKING A HIGH PITCH MOTOR SOUND. THERE WAS NO PATIENT INVOLVEMENT.
21-08-2020 | Malfunction | covidien lp llc north haven | 26-01-2021 | SIGNIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE HANDLE WAS MAKING A HIGH PITCH MOTOR SOUND. THERE WAS NO PATIENT INVOLVEMENT.
26-01-2021 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE HANDLE DID NOT WORK. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE WOULD NOT CYCLE DUE TO SHEARED FIRING RACK TEETH DAMAGE.
26-01-2021 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE HANDLE DID NOT WORK. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE WOULD NOT CYCLE DUE TO SHEARED FIRING RACK TEETH DAMAGE.
26-10-2020 | Malfunction | covidien lp llc north haven | 26-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, AFTER REMOVING FROM THE CHARGER, THE DEVICE DID NOT START UP AFTER PRESSING AND HOLDING THE RIGHT AND LEFT GREEN BUTTONS FOR 10 SECONDS. ANOTHER POWER HANDLE IN THE HOSPITAL WAS USED. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 -MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
26-10-2020 | Malfunction | covidien lp llc north haven | 26-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, AFTER REMOVING FROM THE CHARGER, THE DEVICE DID NOT START UP AFTER PRESSING AND HOLDING THE RIGHT AND LEFT GREEN BUTTONS FOR 10 SECONDS. ANOTHER POWER HANDLE IN THE HOSPITAL WAS USED. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 -MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED CONDITION WAS DETERMINED TO BE A RESULT OF DEFICIENT SOFTWARE. THE BATTERY CELLS VENTED, CAUSING THEM TO LEAK ELECTROLYTIC FLUID. THE BATTERY CELLS WOULD BE UNABLE TO CHARGE UNDER THESE CIRCUMSTANCES AND WOULD RESULT IN THE HANDLE ENTERING A CELL UNDER VOLTAGE STATE.
19-10-2020 | Malfunction | covidien lp llc north haven | 26-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE POWER HANDLE DID NOT ACTIVATE AND NOTED A RED FLASH BLINKING ON THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
19-10-2020 | Malfunction | covidien lp llc north haven | 26-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE POWER HANDLE DID NOT ACTIVATE AND NOTED A RED FLASH BLINKING ON THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
19-10-2020 | Malfunction | covidien lp llc north haven | 26-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE POWER HANDLE DID NOT ACTIVATE AND NOTED A RED FLASH BLINKING ON THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES WERE VENTED.
21-09-2020 | Malfunction | us surgical puerto rico | 26-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE STOMACH DURING A LAPAROSCOPIC GASTROPLASTY, THE STAPLER FIRED BUT THE HANDLE BROKE. ANOTHER STAPLER WAS USED TO COMPLETE THE CASE AND ALL THREE RELOADS WORKED FINE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SHEARED TEETH ON THE FIRING RACK MAY OCCUR IN ANY OF THE FOLLOWING CIRCUMSTANCES: FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE. FIRING WITH AN OBSTACLE INCORPORATED IN THE JAWS. ATTEMPTING TO FIRE A RELOAD THAT IS IN INTERLOCK. IN EACH CIRCUMSTANCE, IT WILL BECOME INCREASINGLY DIFFICULT TO ACTUATE THE FIRING HANDLE AND THE INSTRUMENT RETURN KNOBS WILL BE DIFFICULT TO RETRACT. IN ADDITION, STAPLES MAY NOT FORM PROPERLY, AND TISSUE MAY NOT BE FULLY TRANSECTED.
26-01-2021 | Injury | covidien lp llc north haven | 26-01-2021 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE STOMACH, THE STAPLER WAS ABLE TO FIRE, THE PATIENT WAS BROUGHT BACK FOR A SECOND OPERATION TO LOOK FOR ACTIVE BLEEDING ATTRIBUTED TO THE RELOAD SHEET. NONE WAS DISCOVERED. BLOOD TRANSFUSION WAS PERFORMED.
26-01-2021 | Injury | covidien lp llc north haven | 26-01-2021 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE STOMACH, THE STAPLER WAS ABLE TO FIRE, THE PATIENT WAS BROUGHT BACK FOR A SECOND OPERATION TO LOOK FOR ACTIVE BLEEDING ATTRIBUTED TO THE RELOAD SHEET. NONE WAS DISCOVERED. BLOOD TRANSFUSION WAS PERFORMED.
16-03-2020 | Injury | covidien lp llc north haven | 26-01-2021 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON STAPLING OF THE STOMACH, THE PATIENT WAS HOSPITALIZED FOR THREE DAYS ATTRIBUTED TO THE RELOAD SHEET.
23-11-2020 | Injury | covidien lp llc north haven | 26-01-2021 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, FOLLOWING STAPLING OF THE STOMACH DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS A POTENTIAL COMPLICATION ATTRIBUTED TO RELOAD. THIS RESULTED TO EXTENDED PATIENT HOSPITALIZATION FOR FOUR DAYS.
05-01-2021 | Malfunction | davis & geck caribe | 26-01-2021 | UNKNOWN ENDO STITCH SULU | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE WAS DIFFICULT TO TOGGLE, LOAD AND UNLOAD. THE SURGEON WAS STRUGGLING TO TOGGLE AND THE LEFT LOBE OF THE LIVER WAS CAUGHT BY THE DEVICE. THE SURGEON WAS ABLE TO FIX THE INJURY AND CONTINUED WITH THE CASE. THERE WAS BLOOD LOSS OF 10-20CC AND THE OPERATING ROOM (OR) TIME WAS EXTENDED BY 20 MINUTES. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | davis & geck caribe | 26-01-2021 | UNKNOWN ENDO STITCH SULU | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE WAS DIFFICULT TO TOGGLE, LOAD AND UNLOAD. THE SURGEON WAS STRUGGLING TO TOGGLE AND THE LEFT LOBE OF THE LIVER WAS CAUGHT BY THE DEVICE. THE SURGEON WAS ABLE TO FIX THE INJURY AND CONTINUED WITH THE CASE. THERE WAS BLOOD LOSS OF 10-20CC AND THE OPERATING ROOM (OR) TIME WAS EXTENDED BY 20 MINUTES. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | davis & geck caribe | 26-01-2021 | UNKNOWN ENDO STITCH SULU | Component or Accessory Incompatibility | A1701 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE WAS DIFFICULT TO TOGGLE, LOAD AND UNLOAD. THE SURGEON WAS STRUGGLING TO TOGGLE AND THE LEFT LOBE OF THE LIVER WAS CAUGHT BY THE DEVICE. THE SURGEON WAS ABLE TO FIX THE INJURY AND CONTINUED WITH THE CASE. THERE WAS BLOOD LOSS OF 10-20CC AND THE OPERATING ROOM (OR) TIME WAS EXTENDED BY 20 MINUTES. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | davis & geck caribe | 26-01-2021 | UNKNOWN ENDO STITCH SULU | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE WAS DIFFICULT TO TOGGLE, LOAD AND UNLOAD. THE SURGEON WAS STRUGGLING TO TOGGLE AND THE LEFT LOBE OF THE LIVER WAS CAUGHT BY THE DEVICE. THE SURGEON WAS ABLE TO FIX THE INJURY AND CONTINUED WITH THE CASE. THERE WAS BLOOD LOSS OF 10-20CC AND THE OPERATING ROOM (OR) TIME WAS EXTENDED BY 20 MINUTES. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | davis & geck caribe | 26-01-2021 | UNKNOWN ENDO STITCH SULU | Human-Device Interface Problem | A22 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE WAS DIFFICULT TO TOGGLE, LOAD AND UNLOAD. THE SURGEON WAS STRUGGLING TO TOGGLE AND THE LEFT LOBE OF THE LIVER WAS CAUGHT BY THE DEVICE. THE SURGEON WAS ABLE TO FIX THE INJURY AND CONTINUED WITH THE CASE. THERE WAS BLOOD LOSS OF 10-20CC AND THE OPERATING ROOM (OR) TIME WAS EXTENDED BY 20 MINUTES. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | davis & geck caribe | 26-01-2021 | UNKNOWN ENDO STITCH SULU | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE WAS DIFFICULT TO TOGGLE, LOAD AND UNLOAD. THE SURGEON WAS STRUGGLING TO TOGGLE AND THE LEFT LOBE OF THE LIVER WAS CAUGHT BY THE DEVICE. THE SURGEON WAS ABLE TO FIX THE INJURY AND CONTINUED WITH THE CASE. THERE WAS BLOOD LOSS OF 10-20CC AND THE OPERATING ROOM (OR) TIME WAS EXTENDED BY 20 MINUTES. THERE WAS NO PATIENT INJURY.
24-12-2020 | Injury | us surgical puerto rico | 26-01-2021 | EEA | Break | A0401 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING END TO END ANASTOMOSIS ON AN OPEN SUBTOTAL GASTRECTOMY, A CRACKING NOISE WAS HEARD ONCE FIRED. WHEN THE DEVICE WAS RELEASED FROM THE CAVITY, THE ANVIL WAS BROKEN. IT WAS STATED THAT CERTAIN PARTS MAY FELL INTO THE CAVITY. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES FOR MORE SINCE THEY HAD TO FIND THE MISSING PART FROM CAVITY. ASSUMING THAT THERE WAS A MISSING PART FROM THE BROKEN ANVIL. THE OPERATOR TRIED TO LOCATE THE MISSING PART BUT UNABLE TO FIND ANY. AN X-RAY WAS PERFORMED BUT STILL COULD NOT FIND ANY. IT WAS ALSO STATED THAT THE THAT THE PATIENT UNDERGONE A CHEMOTHERAPY OR RADIATION TREATMENT.
24-12-2020 | Injury | us surgical puerto rico | 26-01-2021 | EEA | Detachment of Device or Device Component | A0501 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING END TO END ANASTOMOSIS ON AN OPEN SUBTOTAL GASTRECTOMY, A CRACKING NOISE WAS HEARD ONCE FIRED. WHEN THE DEVICE WAS RELEASED FROM THE CAVITY, THE ANVIL WAS BROKEN. IT WAS STATED THAT CERTAIN PARTS MAY FELL INTO THE CAVITY. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES FOR MORE SINCE THEY HAD TO FIND THE MISSING PART FROM CAVITY. ASSUMING THAT THERE WAS A MISSING PART FROM THE BROKEN ANVIL. THE OPERATOR TRIED TO LOCATE THE MISSING PART BUT UNABLE TO FIND ANY. AN X-RAY WAS PERFORMED BUT STILL COULD NOT FIND ANY. IT WAS ALSO STATED THAT THE THAT THE PATIENT UNDERGONE A CHEMOTHERAPY OR RADIATION TREATMENT.
24-12-2020 | Injury | us surgical puerto rico | 26-01-2021 | EEA | Noise, Audible | A0508 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING END TO END ANASTOMOSIS ON AN OPEN SUBTOTAL GASTRECTOMY, A CRACKING NOISE WAS HEARD ONCE FIRED. WHEN THE DEVICE WAS RELEASED FROM THE CAVITY, THE ANVIL WAS BROKEN. IT WAS STATED THAT CERTAIN PARTS MAY FELL INTO THE CAVITY. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES FOR MORE SINCE THEY HAD TO FIND THE MISSING PART FROM CAVITY. ASSUMING THAT THERE WAS A MISSING PART FROM THE BROKEN ANVIL. THE OPERATOR TRIED TO LOCATE THE MISSING PART BUT UNABLE TO FIND ANY. AN X-RAY WAS PERFORMED BUT STILL COULD NOT FIND ANY. IT WAS ALSO STATED THAT THE THAT THE PATIENT UNDERGONE A CHEMOTHERAPY OR RADIATION TREATMENT.
06-01-2021 | Injury | us surgical puerto rico | 26-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE FIRST FIRING AT THE PYLORUS, THE DEVICE FIRED FULLY BUT THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE PYLORUS WAS CUT WITH POORLY FORMED AND LOOSE STAPLE LINE. THE SAME ISSUE OCCURRED WITH ANOTHER RELOAD. SUTURING WAS DONE THROUGHOUT THE STAPLE LINE. OMENTOPLASTY WAS DONE TO SECURE THE STAPLE LINE. THE KNIFE BAR AND A PIECE OF METAL FELL INTO THE PATIENT¿S CAVITY AND WAS RETRIEVED WITH GRASPERS. THERE WAS BLEEDING AND THE SURGERY WAS EXTENDED BY 1HOUR. THE RELOAD WAS ALSO STUCK TO ADAPTER AND WAS VERY DIFFICULT TO UNLOAD. ANOTHER STAPLING DEVICE WAS USED TO COMPLETE THE CASE.
06-01-2021 | Injury | us surgical puerto rico | 26-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE FIRST FIRING AT THE PYLORUS, THE DEVICE FIRED FULLY BUT THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE PYLORUS WAS CUT WITH POORLY FORMED AND LOOSE STAPLE LINE. THE SAME ISSUE OCCURRED WITH ANOTHER RELOAD. SUTURING WAS DONE THROUGHOUT THE STAPLE LINE. OMENTOPLASTY WAS DONE TO SECURE THE STAPLE LINE. THE KNIFE BAR AND A PIECE OF METAL FELL INTO THE PATIENT¿S CAVITY AND WAS RETRIEVED WITH GRASPERS. THERE WAS BLEEDING AND THE SURGERY WAS EXTENDED BY 1HOUR. THE RELOAD WAS ALSO STUCK TO ADAPTER AND WAS VERY DIFFICULT TO UNLOAD. ANOTHER STAPLING DEVICE WAS USED TO COMPLETE THE CASE.
06-01-2021 | Injury | us surgical puerto rico | 26-01-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE FIRST FIRING AT THE PYLORUS, THE DEVICE FIRED FULLY BUT THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE PYLORUS WAS CUT WITH POORLY FORMED AND LOOSE STAPLE LINE. THE SAME ISSUE OCCURRED WITH ANOTHER RELOAD. SUTURING WAS DONE THROUGHOUT THE STAPLE LINE. OMENTOPLASTY WAS DONE TO SECURE THE STAPLE LINE. THE KNIFE BAR AND A PIECE OF METAL FELL INTO THE PATIENT¿S CAVITY AND WAS RETRIEVED WITH GRASPERS. THERE WAS BLEEDING AND THE SURGERY WAS EXTENDED BY 1HOUR. THE RELOAD WAS ALSO STUCK TO ADAPTER AND WAS VERY DIFFICULT TO UNLOAD. ANOTHER STAPLING DEVICE WAS USED TO COMPLETE THE CASE.
06-01-2021 | Injury | us surgical puerto rico | 26-01-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE FIRST FIRING AT THE PYLORUS, THE DEVICE FIRED FULLY BUT THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE PYLORUS WAS CUT WITH POORLY FORMED AND LOOSE STAPLE LINE. THE SAME ISSUE OCCURRED WITH ANOTHER RELOAD. SUTURING WAS DONE THROUGHOUT THE STAPLE LINE. OMENTOPLASTY WAS DONE TO SECURE THE STAPLE LINE. THE KNIFE BAR AND A PIECE OF METAL FELL INTO THE PATIENT¿S CAVITY AND WAS RETRIEVED WITH GRASPERS. THERE WAS BLEEDING AND THE SURGERY WAS EXTENDED BY 1HOUR. THE RELOAD WAS ALSO STUCK TO ADAPTER AND WAS VERY DIFFICULT TO UNLOAD. ANOTHER STAPLING DEVICE WAS USED TO COMPLETE THE CASE.
06-01-2021 | Injury | us surgical puerto rico | 26-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE FIRST FIRING AT THE PYLORUS, THE DEVICE FIRED FULLY BUT THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE PYLORUS WAS CUT WITH POORLY FORMED AND LOOSE STAPLE LINE. THE SAME ISSUE OCCURRED WITH ANOTHER RELOAD. SUTURING WAS DONE THROUGHOUT THE STAPLE LINE. OMENTOPLASTY WAS DONE TO SECURE THE STAPLE LINE. THE KNIFE BAR AND A PIECE OF METAL FELL INTO THE PATIENT¿S CAVITY AND WAS RETRIEVED WITH GRASPERS. THERE WAS BLEEDING AND THE SURGERY WAS EXTENDED BY 1HOUR. THE RELOAD WAS ALSO STUCK TO ADAPTER AND WAS VERY DIFFICULT TO UNLOAD. ANOTHER STAPLING DEVICE WAS USED TO COMPLETE THE CASE.
06-01-2021 | Injury | us surgical puerto rico | 26-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE FIRST FIRING AT THE PYLORUS, THE DEVICE FIRED FULLY BUT THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE PYLORUS WAS CUT WITH POORLY FORMED AND LOOSE STAPLE LINE. THE SAME ISSUE OCCURRED WITH ANOTHER RELOAD. SUTURING WAS DONE THROUGHOUT THE STAPLE LINE. OMENTOPLASTY WAS DONE TO SECURE THE STAPLE LINE. THE KNIFE BAR AND A PIECE OF METAL FELL INTO THE PATIENT¿S CAVITY AND WAS RETRIEVED WITH GRASPERS. THERE WAS BLEEDING AND THE SURGERY WAS EXTENDED BY 1HOUR. THE RELOAD WAS ALSO STUCK TO ADAPTER AND WAS VERY DIFFICULT TO UNLOAD. ANOTHER STAPLING DEVICE WAS USED TO COMPLETE THE CASE.
09-11-2020 | Injury | covidien lp llc north haven | 26-01-2021 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT WAS BROUGHT BACK TO THE HOSPITAL TO LOOK FOR AN ACTIVE BLEEDING. BLOOD TRANSFUSION WAS PERFORMED AND THE HOSPITALIZATION WAS EXTENDED FOR 3 DAYS. IT WAS ALSO REPORTED THAT THE PATIENT WAS SUBJECTED TO HEART RATE, BLOOD PRESSURE MONITORING AND POSSIBLE INDICATORS OF POST-OPERATIVE BLEEDING WHEN THE ISSUE OCCURRED.
08-01-2021 | Malfunction | ethicon endo-surgery | 26-01-2021 | POWERED PLUS COMPACT 45 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON FELT THE FEELING OF ECHELON STAPLER AT ITS TRIGGER FIRING WAS DIFFERENT FROM REGULAR TIMES. IT SEEMED THAT THE FIRING ITSELF AND STAPLE FORMATION WERE NOT THAT MUCH OF AN ISSUE BUT IT SEEMED TO US THAT IT ALL STARTED WITH THE ISSUE OF THE SURGEON'S ABNORMAL FEELING OF THE TRIGGER. AFTER WE HAVE LOOKED AT THE RETURNED STAPLER, WE FOUND THE METAL SPRING HAS BEEN DISLODGED COMPARE TO THE NEW ONE WE BROUGHT TO CHECK THE DIFFERENCES IN SPRING LOCATION. EVEN THOUGH THE STAPLER WORKED NORMALLY IN THE SURGERY, BUT WE INQUIRE ABOUT ANY POSSIBILITIES OF THE SPRING DISLODGEMENT AND WHAT KIND OF PREVENTION ACTION WE COULD HAVE FOR THIS ISSUE. IT IS UNKNOWN IF THERE WERE PATIENT CONSEQUENCES.
07-01-2021 | Malfunction | ethicon endo-surgery | 26-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC SLEEVE THE DEVICE MADE A FUNNY NOISE AND GOT STUCK ON TISSUE. THE MANUAL OVERRIDE DID NOT WORK. IT IS UNKNOWN HOW THE DEVICE WAS REMOVED. NO PATIENT CONSEQUENCES WERE REPORTED.
07-01-2021 | Malfunction | ethicon endo-surgery | 26-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC SLEEVE THE DEVICE MADE A FUNNY NOISE AND GOT STUCK ON TISSUE. THE MANUAL OVERRIDE DID NOT WORK. IT IS UNKNOWN HOW THE DEVICE WAS REMOVED. NO PATIENT CONSEQUENCES WERE REPORTED.
06-01-2021 | Malfunction | ethicon endo-surgery | 26-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RADICAL GASTRECTOMY FOR GASTRIC CANCER SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT THE KNIFE STOPPED AFTER FORWARD AROUND 2 MILLIMETERS DURING THE RETURN KNIFE PROCESS, THEN THE SURGEON USED THE INSTRUMENT TO PUSH KNIFE TO THE HOME POSITION. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2021 | Malfunction | ethicon endo-surgery | 26-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SEGMENTECTOMY SURGERY WHEN SEVERING VESSELS TISSUE ON THE FIRST AND SECOND FIRING, AFTER FIRED WITH TWO VASECR35A, NOTED STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG). CHANGED TO THE NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-10-2020 | Injury | ethicon endo-surgery | 26-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST-OPERATIVE OF SLEEVE GASTRECTOMY PROCEDURE, A LEAK OCCURRED ON THE LAST FIRING OF A GST60B. THE PATIENT INITIALLY DID WELL, NOW AT HOSPITAL REQUIRING SEVERAL IR AND GI PROCEDURES.
12-12-2020 | Injury | ethicon endo-surgery | 26-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED POST-OPERATIVE TO A SLEEVE GASTRECTOMY PROCEDURE, A LEAK OCCURRED ON THE LAST FIRING OF A GST60B. THE PATIENT HAS HAD 2 EGDS WITH STENTS PLACED WITH ONGOING LEAK. ON TPN.
23-12-2020 | Malfunction | ethicon endo-surgery | 26-01-2021 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS OF THE STAPLER DID NOT CLOSE COMPLETELY. THIS WAS OBVIOUS ON VISUAL INSPECTION. WHEN WE TRIED TO RE-OPEN THE JAWS AND REATTEMPT CLOSURE, THE JAWS WOULD NOT REOPEN. THE STAPLER COULD NOT THEREFORE BE REMOVED FROM THE TISSUE. THIS OCCURRED DURING LIGATION OF A LARGE RENAL VEIN AND LEFT THE PATIENT OPEN TO SIGNIFICANT HARM HAD WE ENCOUNTERED UNCONTROLLABLE BLEEDING. WORKING AROUND THE PARTIALLY CLOSED JAWS AND WITH THE BULK OF THE DEVICE IN THE WAY, WE WERE FORCED TO CONTROL THE RENAL VEIN USING SUTURE LIGATURE, DIVIDE IT, AND THEN ALLOW THE STUMP OF THE RENAL VEIN TO FALL AWAY FROM THE JAWS. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. ONE STAPLER WITH RELOAD AND TEN STERILE RELOADS WILL BE RETURNED.
01-01-2020 | Malfunction | ethicon endo-surgery | 26-01-2021 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THERE WAS A MISFIRE FROM RELOAD. THE SURGERY WAS DELAYED 15 MINUTES. IT IS UNKNOWN IF THERE WERE PATIENT CONSEQUENCES.
29-12-2020 | Malfunction | ethicon endo-surgery | 26-01-2021 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY TOWARDS THE LAST FIRING OF THE DEVICE AT THE TOP OF THE STOMACH THE TISSUE WAS PUSHED OUTWARDS, AND THIS CAUSED THE STAPLES TO BE MALFORMED. THE SURGEON SURMISED THAT THE TISSUE WAS THICK AND MAY HAVE BEEN FOLDED OVER. THE PROCEDURE WAS COMPLETED WITH A SECOND RELOAD WITH NO PATIENT CONSEQUENCES.
06-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC ASSISTED RIGHT HEMICOLECTOMY, THE SURGEON WAS FIRING IT ACROSS THE CAECUM, AND THE DEVI CE DID NOT FIRE. THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED. ANOTHER RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC ASSISTED RIGHT HEMICOLECTOMY, THE SURGEON WAS FIRING IT ACROSS THE CAECUM, AND THE DEVI CE DID NOT FIRE. THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED. ANOTHER RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-12-2020 | Malfunction | covidien lp llc north haven | 27-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN THE DEVICE WAS BEING APPLIED TO THE LUNG DURING A LAPAROSCOPIC WEDGE THORACOSCOPY RESECTION, THE RELOAD WAS ABLE TO BE PLACED NORMALLY BUT WHEN THE BLADE WENT TO SHOOT, IT DID NOT ADVANCE. THE STAPLER WAS ACTIVATED, BUT THE RELOAD DID NOT WORK; IT DID NOT ADVANCE, NEITHER STAPLE NOR CUT. ANOTHER RELOAD WAS USED AND WORKED WELL IN THE SAME STAPLER. THERE WAS NO PATIENT INJURY.
28-10-2015 | Malfunction | covidien lp llc north haven | 27-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT TRIGGER. THERE WAS NO PATIENT INJURY.
28-10-2015 | Malfunction | covidien lp llc north haven | 27-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT TRIGGER. THERE WAS NO PATIENT INJURY.
09-01-2021 | Malfunction | ethicon endo-surgery | 27-01-2021 | ENDOPATH | Misfire | A050502 | Insufficient Information | Event Description: ENDOPATH ETS-FLEX 45, REF (B)(4) WAS USED ON A LAPAROSCOPIC CASE AND IT MISFIRED TWICE WITH 2 DIFFERENT LOADS - BOTH INCLUDED WHEN HANDED INTO MANAGEMENT FOR EVALUATION (RELOAD SAVED BUT NO PACKAGING INFORMATION SAVED).
05-01-2021 | Malfunction | ethicon endo-surgery | 27-01-2021 | ECHELON FLEX | Noise, Audible | A0508 | Insufficient Information | Event Description: ECHELON FLEX STAPLER WAS BEING USED. IT WORKED FINE FOR 2-3 LOADS. THE SCRUBBER HEARD A POP THEN THE STAPLER WOULD NOT LOAD CORRECTLY.
05-01-2021 | Malfunction | ethicon endo-surgery | 27-01-2021 | ECHELON FLEX | Firing Problem | A0505 | Insufficient Information | Event Description: ECHELON FLEX STAPLER WAS BEING USED. IT WORKED FINE FOR 2-3 LOADS. THE SCRUBBER HEARD A POP THEN THE STAPLER WOULD NOT LOAD CORRECTLY.
06-01-2021 | Malfunction | intuitive surgical | 27-01-2021 | SUREFORM | Patient-Device Incompatibility | A0101 | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLER UNABLE TO BE USED. STAPLER WAS CLAMPED ON TISSUE THAT WAS IDENTIFIED AS TOO THICK FOR STAPLER. UPON RELOADING WITH LARGER LOAD, STAPLER WOULD NOT GO INTO TROCAR. ADDITIONAL STAPLER OPENED AND USED. NO HARM OR INJURY TO PATIENT.
06-01-2021 | Malfunction | intuitive surgical | 27-01-2021 | SUREFORM | Device-Device Incompatibility | A1702 | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLER UNABLE TO BE USED. STAPLER WAS CLAMPED ON TISSUE THAT WAS IDENTIFIED AS TOO THICK FOR STAPLER. UPON RELOADING WITH LARGER LOAD, STAPLER WOULD NOT GO INTO TROCAR. ADDITIONAL STAPLER OPENED AND USED. NO HARM OR INJURY TO PATIENT.
05-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE RESECTION, THE JAWS OF THE DEVICE WOULD NOT OPEN. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. THERE WAS A TRACE OF PRE-FIRE. NO PATIENT INJURY.
06-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, DEVICE COULD NOT CLOSE CORRECTLY AND WAS NOT ABLE TO FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
06-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, DEVICE COULD NOT CLOSE CORRECTLY AND WAS NOT ABLE TO FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
06-01-2021 | Malfunction | covidien lp llc north haven | 27-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, DEVICE COULD NOT CLOSE CORRECTLY AND WAS NOT ABLE TO FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
06-01-2021 | Malfunction | covidien lp llc north haven | 27-01-2021 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING SLEEVE GASTRECTOMY, DEVICE COULD NOT CLOSE CORRECTLY AND WAS NOT ABLE TO FIRE. ANOTHER HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
19-11-2015 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT RELEASE LOAD. THERE WAS NO PATIENT INVOLVEMENT.
19-11-2015 | Malfunction | covidien lp llc north haven | 27-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT RELEASE LOAD. THERE WAS NO PATIENT INVOLVEMENT.
19-11-2015 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE STAPLER WOULD NOT RELEASE LOAD.
19-11-2015 | Malfunction | covidien lp llc north haven | 27-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE STAPLER WOULD NOT RELEASE LOAD.
06-01-2020 | Injury | us surgical puerto rico | 27-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTROPLASTY, THE STAPLER WAS FIRED, BUT SOME STAPLES DID NOT CLOSE. THEREFORE, BLEEDING IN THE STAPLE LINE WAS INCREASED. THIS ISSUE OCCURRED ON FOUR UNITS OF RELOADS. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 27-01-2021 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTROPLASTY, THE STAPLER WAS FIRED, BUT SOME STAPLES DID NOT CLOSE. THEREFORE, BLEEDING IN THE STAPLE LINE WAS INCREASED. THIS ISSUE OCCURRED ON FOUR UNITS OF RELOADS. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
06-01-2020 | Injury | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTROPLASTY, THE STAPLER WAS FIRED, BUT SOME STAPLES DID NOT CLOSE. THEREFORE, BLEEDING IN THE STAPLE LINE WAS INCREASED. THIS ISSUE OCCURRED ON FOUR UNITS OF RELOADS. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
05-01-2021 | Malfunction | covidien lp llc north haven | 27-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PARTIAL EXCISION OF THE STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED AND THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | celestica de monterrey s.a de c.v | 27-01-2021 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC LEFT LOWER LOBECTOMY, UPON OPENING THE PACKAGE WHILE PREPARING FOR PULMONARY ARTERY DISSECTION, IT WAS FOUND THAT THE BUCKLE AT THE CONNECTION OF THE RELOAD WAS TWISTED AND BROKEN. THE RELOAD COULD NOT BE LOADED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | covidien lp llc north haven | 27-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE THE SURGEON USED THE DEVICE IN AN IVOR LEWIS ESOPHAGECTOMY, BOTH RELOADS STOPPED FIRING PRIOR TO FULL LENGTH OF STAPLE CARTRIDGE BEING UTILIZED. ANOTHER RELOAD WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-06-2020 | Malfunction | covidien lp llc north haven | 27-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE DEVICE CAME LOCKED AND DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. THE DEVICE ALSO LOCKED ON TISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. THE DEVICE ALSO LOCKED ON TISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | covidien lp llc north haven | 27-01-2021 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. THE DEVICE ALSO LOCKED ON TISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | covidien lp llc north haven | 27-01-2021 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. THE DEVICE ALSO LOCKED ON TISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC WEDGE RESECTION OF LEFT LUNG LOBE, ON PULMONARY LOBE, THE DEVICE DID NOT FIRE. THE SURGEON ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-08-2020 | Malfunction | covidien lp llc north haven | 27-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, WHEN CONNECTING THE LOAD TO THE STAPLER AND ATTEMPTING TO TRIGGER, THE LOAD REMAINED CLOSED AND THE TRIGGER WAS LOCKED, THE STAPLER WAS UNABLE TO FIRE. THE PROCEDURE WAS COMPLETED WITH ANOTHER CHARGE. THERE WAS NO PATIENT INJURY.
21-08-2020 | Malfunction | covidien lp llc north haven | 27-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, WHEN CONNECTING THE LOAD TO THE STAPLER AND ATTEMPTING TO TRIGGER, THE LOAD REMAINED CLOSED AND THE TRIGGER WAS LOCKED, THE STAPLER WAS UNABLE TO FIRE. THE PROCEDURE WAS COMPLETED WITH ANOTHER CHARGE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PANCREATICODUODENECTOMY, UPON RESECTION OF THE DUODENUM, THE STAPLER COULD NOT BE FIRED NORMALLY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. BOTH HANDLE AND RELOAD WERE REPLACED WITH ANOTHER MANUFACTURER¿S PRODUCT TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | covidien lp llc north haven | 27-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PANCREATICODUODENECTOMY, UPON RESECTION OF THE DUODENUM, THE STAPLER COULD NOT BE FIRED NORMALLY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. BOTH HANDLE AND RELOAD WERE REPLACED WITH ANOTHER MANUFACTURER¿S PRODUCT TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
05-01-2021 | Malfunction | covidien lp llc north haven | 27-01-2021 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE RESECTION, THE JAWS OF THE DEVICE WOULD NOT OPEN. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. THERE WAS A TRACE OF PRE-FIRE. NO PATIENT INJURY.
05-01-2021 | Malfunction | covidien lp llc north haven | 27-01-2021 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING WEDGE RESECTION, THE JAWS OF THE DEVICE WOULD NOT OPEN. SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. THERE WAS A TRACE OF PRE-FIRE. NO PATIENT INJURY.
18-05-2020 | Malfunction | covidien lp llc north haven | 27-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING BYPASS GASTROPLASTY, WHEN THE SURGEON ATTEMPTED TO FIRE THE DEVICE, THE DEVICE LOCKED. SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
07-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE ON BOTH RELOADS. SURGEON REPLACED THE HANDLE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
19-11-2015 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THERE WAS DEFECTIVE LOAD FIRING. THE LOAD WAS CHANGED TO RESOLVE THE ISSUE.
07-01-2021 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOW RECTAL RESECTION, WHEN ON RECTUM, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED WITH A NEW ONE AND NEW STAPLE RELOAD TO CONTINUE THE OPERATION. THERE WAS NO PATIENT INJURY.
06-05-2020 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER LOCKED AFTER LOADING. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2020 | Malfunction | us surgical puerto rico | 27-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE, UPON GASTRIC SLEEVE, FIRING ON STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. HANDLE WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | BLUE RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | BLUE RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Dyspnea | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | BLUE RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | BLUE RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | GREEN RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | GREEN RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Dyspnea | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | GREEN RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | GREEN RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | GREEN RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT(TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | GREEN RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Dyspnea | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT(TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | GREEN RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT(TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | GREEN RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT(TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | BLUE RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | BLUE RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Dyspnea | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | BLUE RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
26-11-2020 | Injury | ethicon endo-surgery | 27-01-2021 | BLUE RELOAD FOR ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT ON (B)(6) 2020, A BARIATRIC SURGERY WAS SCHEDULED (GASTRIC BYPASS), AT CLINIC, DURING WHICH THERE WERE PROBLEMS WITH THE SUTURES USED. THERE WAS BLEEDING WHEN USING WHITE RELOADS. LEAK CHECKED WITH METHYLENE BLUE AND BLEEDING WHEN USING ALL BLUE RELOADS. FOR THIS REASON, THE SURGERY WAS PROLONGED AND HAD TO BE COMPLETED WITH OPEN TECHNIQUE. MANAGEMENT IN ICU (INTENSIVE CARE UNIT) AND NPT (TOTAL PARENTERAL NUTRITION) HOSPITALIZATION TIME WAS MORE THAN ONE MONTH.
08-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | TA | Product Quality Problem | A0201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CLOSURE OF ANASTOMOSIS ON AN OPEN ILEOSTOMY REVERSAL, THE RELOAD THAT CAME WITH THE STAPLER DID NOT HAVE ANY STAPLES IN IT. THE STAPLER CAME OUT OF THE BOX WITH ALL THE YELLOW TABS WITHIN THE RELOAD EXTENDED. A NEW STAPLER WAS USED. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | TA | Component Missing | A020602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CLOSURE OF ANASTOMOSIS ON AN OPEN ILEOSTOMY REVERSAL, THE RELOAD THAT CAME WITH THE STAPLER DID NOT HAVE ANY STAPLES IN IT. THE STAPLER CAME OUT OF THE BOX WITH ALL THE YELLOW TABS WITHIN THE RELOAD EXTENDED. A NEW STAPLER WAS USED. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | celestica de monterrey s.a de c.v | 28-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC WEDGE RESECTION, WHEN RESECTING LUNG LOBE, POOR STAPLE FORMATION WAS OBSERVED FROM THE TWO RELOADS. MOST OF THE STAPLES FAILED TO FORM. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-10-2014 | Malfunction | us surgical puerto rico | 28-01-2021 | GIA | Failure to Cut | A050702 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PERFORMANCE OF GASTROENTEROANASTOMOSIS, A LOAD WAS FIRED WITHOUT MAKING THE CUT OF THE TISSUE. IT WAS NECESSARY TO USE ANOTHER RELOAD.
07-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | EEA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE STAPLER WAS DIFFICULT TO BE CLOSED. THE SPIKE WAS DEPLOYED, THE ANVIL WAS JOINED AND CLOSED, BUT THE SAFETY WOULD NOT RELEASE AND THE SURGEON COULD NOT CLOSE THE HANDLE. THE SURGEON OPENED THE STAPLER, RE-DID CONNECTION TO ANVIL AND CLOSURE AND THEN IT WOULD FIRE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | EEA | Positioning Problem | A1502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE STAPLER WAS DIFFICULT TO BE CLOSED. THE SPIKE WAS DEPLOYED, THE ANVIL WAS JOINED AND CLOSED, BUT THE SAFETY WOULD NOT RELEASE AND THE SURGEON COULD NOT CLOSE THE HANDLE. THE SURGEON OPENED THE STAPLER, RE-DID CONNECTION TO ANVIL AND CLOSURE AND THEN IT WOULD FIRE. THERE WAS NO PATIENT INJURY.
28-09-2020 | Malfunction | covidien lp llc north haven | 28-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO THE PROCEDURE, THE POWER HANDLE COULD NOT BE CHARGED WITH THE BATTERY CHARGER. THE CHARGING OF THE REPORTED POWER HANDLE WOULD NOT START EVEN THE REPORTED BATTERY CHARGER WAS REPLACED WITH ANOTHER BATTERY CHARGER WHICH HAD NO PROBLEM. NOTHING WAS INDICATED ON THE DISPLAY OF THE POWER HANDLE. WHEN LOOKING AT THE ELECTRODE SURFACES OF THE REPORTED POWER HANDLE AND THE REPORTED BATTERY CHARGER, THERE WAS A TRACE THAT SEEMED TO BE SHORT CIRCUIT. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES TO BE VENTED.
28-09-2020 | Malfunction | covidien lp llc north haven | 28-01-2021 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO THE PROCEDURE, THE POWER HANDLE COULD NOT BE CHARGED WITH THE BATTERY CHARGER. THE CHARGING OF THE REPORTED POWER HANDLE WOULD NOT START EVEN THE REPORTED BATTERY CHARGER WAS REPLACED WITH ANOTHER BATTERY CHARGER WHICH HAD NO PROBLEM. NOTHING WAS INDICATED ON THE DISPLAY OF THE POWER HANDLE. WHEN LOOKING AT THE ELECTRODE SURFACES OF THE REPORTED POWER HANDLE AND THE REPORTED BATTERY CHARGER, THERE WAS A TRACE THAT SEEMED TO BE SHORT CIRCUIT. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES TO BE VENTED.
28-09-2020 | Malfunction | covidien lp llc north haven | 28-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO THE PROCEDURE, THE POWER HANDLE COULD NOT BE CHARGED WITH THE BATTERY CHARGER. THE CHARGING OF THE REPORTED POWER HANDLE WOULD NOT START EVEN THE REPORTED BATTERY CHARGER WAS REPLACED WITH ANOTHER BATTERY CHARGER WHICH HAD NO PROBLEM. NOTHING WAS INDICATED ON THE DISPLAY OF THE POWER HANDLE. WHEN LOOKING AT THE ELECTRODE SURFACES OF THE REPORTED POWER HANDLE AND THE REPORTED BATTERY CHARGER, THERE WAS A TRACE THAT SEEMED TO BE SHORT CIRCUIT. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES TO BE VENTED.
27-10-2020 | Malfunction | covidien lp llc north haven | 28-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A DISTAL GASTRECTOMY, WHILE CHARGING ON THE CHARGER, THE RED LED ILLUMINATED. THE HANDLE WAS REMOVED AND START-UP WAS ATTEMPTED, BUT IT DID NOT START UP. ALSO, THE GREEN LED ON THE BASE WAS FOUND NOT ILLUMINATING BUT WAS FLASHING, AND IT WAS SEEN THAT THE DISPLAY WAS BLACKOUT. THE OPERATOR USED ANOTHER CHARGER AND HANDLE WHICH WAS MANAGED ON A DIFFERENT CHARGER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. AFTER THE PROCEDURE, WHEN CHARGED A DEMONSTRATION HANDLE BROUGHT BY THE SALES REPRESENTATIVE ON THE CHARGER WHICH HAD THE FAILURE, IT WAS ABLE TO BE CHARGED WITHOUT ANY PROBLEMS. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERIES WERE VENTED.
27-10-2020 | Malfunction | covidien lp llc north haven | 28-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A DISTAL GASTRECTOMY, WHILE CHARGING ON THE CHARGER, THE RED LED ILLUMINATED. THE HANDLE WAS REMOVED AND START-UP WAS ATTEMPTED, BUT IT DID NOT START UP. ALSO, THE GREEN LED ON THE BASE WAS FOUND NOT ILLUMINATING BUT WAS FLASHING, AND IT WAS SEEN THAT THE DISPLAY WAS BLACKOUT. THE OPERATOR USED ANOTHER CHARGER AND HANDLE WHICH WAS MANAGED ON A DIFFERENT CHARGER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. AFTER THE PROCEDURE, WHEN CHARGED A DEMONSTRATION HANDLE BROUGHT BY THE SALES REPRESENTATIVE ON THE CHARGER WHICH HAD THE FAILURE, IT WAS ABLE TO BE CHARGED WITHOUT ANY PROBLEMS. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERIES WERE VENTED.
14-01-2015 | Injury | us surgical puerto rico | 28-01-2021 | GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FULLY FIRED, IT STOPPED AND DID NOT CUT THE LAST PART OF RELOAD. IN ADDITION THERE WERE POOR STAPLE FORMATION. TO RESOLVE THE ISSUE THE SURGEON RE-RESECT AND RE-STAPLE THE TISSUE.
14-01-2015 | Injury | us surgical puerto rico | 28-01-2021 | GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FULLY FIRED, IT STOPPED AND DID NOT CUT THE LAST PART OF RELOAD. IN ADDITION THERE WERE POOR STAPLE FORMATION. TO RESOLVE THE ISSUE THE SURGEON RE-RESECT AND RE-STAPLE THE TISSUE.
08-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOWER ANTERIOR RESECTION, THE STAPLE LINE WAS INCOMPLETE FROM THE CENTER. THE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED BUT THE HANDLE BROKE WHILE USING THE DEVICE. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE.
08-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOWER ANTERIOR RESECTION, THE STAPLE LINE WAS INCOMPLETE FROM THE CENTER. THE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED BUT THE HANDLE BROKE WHILE USING THE DEVICE. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE.
02-10-2020 | Malfunction | covidien lp llc north haven | 28-01-2021 | SIGNIA | Premature End-of-Life Indicator | A160602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, DURING PREPARATION, THERE WAS A HANDLE ERROR MESSAGE, THE DEVICE SHOWED 0 PROCEDURES REMAINING AND THE DEVICE SHOULD HAVE MORE THAN 250. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10- MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE ROOT CAUSE OF THE OBSERVED MOTOR SCREW CONDITION WAS DETERMINED TO BE A RESULT OF A MANUFACTURING ACTIVITY. IT WAS DETERMINED THAT THE THREAD LOCKER APPLIED TO THE SCREWS WAS NOT ACTIVATED PROPERLY.
07-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER CARTRIDGE JAMMED WHILE THE SURGEON WAS FIRING HALFWAY ON THE FUNDUS. THE JAWS LOCKED ON THE TISSUE SO LOAD/UNLOAD BUTTON WAS PRESSED TO REMOVE THE INSTRUMENT FROM THE PATIENT. ANOTHER CARTRIDGE WAS OPENED AND THE SAME HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER CARTRIDGE JAMMED WHILE THE SURGEON WAS FIRING HALFWAY ON THE FUNDUS. THE JAWS LOCKED ON THE TISSUE SO LOAD/UNLOAD BUTTON WAS PRESSED TO REMOVE THE INSTRUMENT FROM THE PATIENT. ANOTHER CARTRIDGE WAS OPENED AND THE SAME HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | covidien lp llc north haven | 28-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC WEDGE RESECTION OF LUNG, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST STEP OF CUTTING STOMACH FOR LAPAROSCOPIC SLEEVE GASTRECTOMY, POOR STAPLE FORMATION WAS OBSERVED. THE STAPLES TWISTED WHEN FIRED. THE STAPLE LINE WAS INCOMPLETE CENTRALLY AND JAWS LOCKED ON TISSUE. BLACK BUTTON WAS RELEASED TO REMOVE THE CARTRIDGE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2021 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST STEP OF CUTTING STOMACH FOR LAPAROSCOPIC SLEEVE GASTRECTOMY, POOR STAPLE FORMATION WAS OBSERVED. THE STAPLES TWISTED WHEN FIRED. THE STAPLE LINE WAS INCOMPLETE CENTRALLY AND JAWS LOCKED ON TISSUE. BLACK BUTTON WAS RELEASED TO REMOVE THE CARTRIDGE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | covidien lp llc north haven | 28-01-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RIGHT COLECTOMY, WHILE DETACHING THE COLON, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | covidien lp llc north haven | 28-01-2021 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOWER ANTERIOR RESECTION, THE STAPLE LINE WAS INCOMPLETE FROM THE CENTER. THE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED BUT THE HANDLE BROKE WHILE USING THE DEVICE. ANOTHER HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE.
08-01-2021 | Injury | covidien lp llc north haven | 28-01-2021 | ENDO GIA | Failure to Cycle | A050703 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE THIRD FIRING, THE STAPLER WAS NOT ABLE TO FORM A B SHAPE. THE BLADE WAS ADVANCED BUT DID NOT CUT THROUGH THE STOMACH, IT PUSHED THE TISSUE OUT AS THE BLADE WAS FIRING. THERE WAS TEAR OF THE SEROSA LAYER OF THE STOMACH AND ANOTHER STAPLER WAS FIRED OVER IT. IT WAS ALSO NOTED THAT THE OTHER RELOAD WAS NOT ABLE TO FIRE. THE SURGICAL TIME WAS ALSO EXTENDED BY MORE THAN 30 MINUTES DUE TO REMOVING THE UNFIRED STAPLER FROM THE TISSUE AND TIME NEED TO TEST THE DEVICE OUT OF THE PATIENT'S BODY. NEW RELOADS WERE USED TO RESOLVE THE ISSUE.
08-01-2021 | Injury | covidien lp llc north haven | 28-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE THIRD FIRING, THE STAPLER WAS NOT ABLE TO FORM A B SHAPE. THE BLADE WAS ADVANCED BUT DID NOT CUT THROUGH THE STOMACH, IT PUSHED THE TISSUE OUT AS THE BLADE WAS FIRING. THERE WAS TEAR OF THE SEROSA LAYER OF THE STOMACH AND ANOTHER STAPLER WAS FIRED OVER IT. IT WAS ALSO NOTED THAT THE OTHER RELOAD WAS NOT ABLE TO FIRE. THE SURGICAL TIME WAS ALSO EXTENDED BY MORE THAN 30 MINUTES DUE TO REMOVING THE UNFIRED STAPLER FROM THE TISSUE AND TIME NEED TO TEST THE DEVICE OUT OF THE PATIENT'S BODY. NEW RELOADS WERE USED TO RESOLVE THE ISSUE.
13-01-2017 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE SURGERY, THE SURGEON PREPARED TO SHOOT AND TRIGGER WAS RELEASED BUT THE CHARGE DID NOT FIRE. THE SURGEON KEPT TRYING TO NO AVAIL.
13-01-2017 | Malfunction | covidien lp llc north haven | 28-01-2021 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE SURGERY, THE SURGEON PREPARED TO SHOOT AND TRIGGER WAS RELEASED BUT THE CHARGE DID NOT FIRE. THE SURGEON KEPT TRYING TO NO AVAIL.
06-05-2015 | Malfunction | us surgical puerto rico | 28-01-2021 | TA PREMIUM | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE.
06-05-2015 | Malfunction | us surgical puerto rico | 28-01-2021 | UNKNOWN TA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE.
06-10-2020 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG SEGMENTECTOMY, THE DEVICE WAS ABLE TO FIRE, HOWEVER, THE STAPLE WAS NOT ABLE TO FORM A PERFECT B SHAPE. IT WAS ALSO NOTED THAT POOR STAPLE FORMATION LED TO BLEEDING. WHILE, THE SECOND RELOAD HAS DEFORMED ANVIL. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENT WAS LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO SHEARED FIRING RACK TEETH DAMAGE.
06-10-2020 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG SEGMENTECTOMY, THE DEVICE WAS ABLE TO FIRE, HOWEVER, THE STAPLE WAS NOT ABLE TO FORM A PERFECT B SHAPE. IT WAS ALSO NOTED THAT POOR STAPLE FORMATION LED TO BLEEDING. WHILE, THE SECOND RELOAD HAS DEFORMED ANVIL. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENT WAS LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO SHEARED FIRING RACK TEETH DAMAGE.
26-01-2015 | Malfunction | us surgical puerto rico | 28-01-2021 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN TRYING TO FIRE THE LOAD, IT STOPPED IN THE MIDDLE OF THE SHOT, MAKING STAPLING IMPOSSIBLE.
29-09-2020 | Malfunction | covidien lp llc north haven | 28-01-2021 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ESOPHAGUS RESECTION PROCEDURE, WHILE BUILDING A GASTRIC SLEEVE, THE DEVICE LOCK ON TISSUE AND WAS REMOVED WITH GENTLE PRESSURE FROM THE JAWS WITH ONE FINGER. IT WAS ALSO REPORTED THAT THE SURGEON NOTICED THAT THERE WERE SEVERAL BLEEDING SPOTS ALONG THE STAPLE LINE AND THE CUTTING LINE WAS FRAYED AFTER THE FIRST AND SECOND STAPLE LINE WERE APPLIED. THE SURGEON MANUALLY SUTURED THE SPOTS TO STOP THE BLEEDING. A THIRD RELOAD WAS USED TO COMPLETE THE PROCEDURE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE DEVICE HAD BEEN USED BEYOND THE RECOMMENDED TISSUE THICKNESS RANGE.
29-09-2020 | Malfunction | covidien lp llc north haven | 28-01-2021 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ESOPHAGUS RESECTION PROCEDURE, WHILE BUILDING A GASTRIC SLEEVE, THE DEVICE LOCK ON TISSUE AND WAS REMOVED WITH GENTLE PRESSURE FROM THE JAWS WITH ONE FINGER. IT WAS ALSO REPORTED THAT THE SURGEON NOTICED THAT THERE WERE SEVERAL BLEEDING SPOTS ALONG THE STAPLE LINE AND THE CUTTING LINE WAS FRAYED AFTER THE FIRST AND SECOND STAPLE LINE WERE APPLIED. THE SURGEON MANUALLY SUTURED THE SPOTS TO STOP THE BLEEDING. A THIRD RELOAD WAS USED TO COMPLETE THE PROCEDURE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE DEVICE HAD BEEN USED BEYOND THE RECOMMENDED TISSUE THICKNESS RANGE.
26-01-2015 | Malfunction | covidien lp llc north haven | 28-01-2021 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN TRYING TO FIRE THE LOAD, IT STOPPED IN THE MIDDLE OF THE SHOT, MAKING STAPLING IMPOSSIBLE.
18-11-2020 | Malfunction | ethicon endo-surgery | 28-01-2021 | PWRD ECH FLEX 60MM 34CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC RIGHT UPPER LOBECTOMY SURGERY, THE THIRD RELOAD CAN NOT BE INSTALLED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-11-2020 | Malfunction | ethicon endo-surgery | 28-01-2021 | ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER | Misfire | A050502 | Insufficient Information | Event Description: STAPLER MISFIRED. ONLY FIRED HALF THE STAPLE. FDA SAFETY REPORT ID#: (B)(4).
08-01-2021 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Difficult or Delayed Positioning | A150203 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, AT THE TIME OF STOMACH RESECTION, THE GREEN LAMP DID NOT FLASH, AND IT WAS RE-CLAMPED SEVERAL TIMES. AFTER A FEW TRIALS, THE GREEN LAMP FLASHED, SO IT FIRED. THEN, THE POWER HANDLE BLACKED OUT AFTER FIRING. NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, AT THE TIME OF STOMACH RESECTION, THE GREEN LAMP DID NOT FLASH, AND IT WAS RE-CLAMPED SEVERAL TIMES. AFTER A FEW TRIALS, THE GREEN LAMP FLASHED, SO IT FIRED. THEN, THE POWER HANDLE BLACKED OUT AFTER FIRING. NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2021 | Injury | us surgical puerto rico | 29-01-2021 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY, FOR THE FOURTH FIRING OF JEJUNAL COLON ANASTOMOSIS FEE (FUNCTIONAL END-TO-END A NASTOMOSIS), THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE. THE PART WHICH WAS NOT FIRED BY STAPLE WAS RESECTED ADDITIONALLY BY USING A NEW RELOAD AND REMOVED.
05-01-2021 | Injury | us surgical puerto rico | 29-01-2021 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY, FOR THE FOURTH FIRING OF JEJUNAL COLON ANASTOMOSIS FEE (FUNCTIONAL END-TO-END A NASTOMOSIS), THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE. THE PART WHICH WAS NOT FIRED BY STAPLE WAS RESECTED ADDITIONALLY BY USING A NEW RELOAD AND REMOVED.
07-08-2020 | Malfunction | us surgical puerto rico | 29-01-2021 | GIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING MEDIAN OPEN LAPAROTOMY BY SEGMENTAL RESECTION OF THE SMALL INTESTINE, THE SURGEON MADE THE CUT BUT THE BLADE DOES NOT RETURN TO ITS PLACE BLOCKING THE DEVICE. ANOTHER STAPLER OF THE SAME REFERENCE WAS PASSED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2021 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AT THE BEGINNING OF A SIGMOIDECTOMY, AFTER TESTING THE STAPLER, RELOAD WAS CLOSED AS IT ENTERS THE TROCAR BUT IT JAWS WOULD NOT RE-OPEN. SURGEON HAD TO FORCE TO REMOVE THE RELOAD. SURGEON USED MANUAL STAPLER TO COMPLETE THE CASE. DEVICE WAS NOT USED ON PATIENT. NO PATIENT INJURY.
08-01-2021 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL HEPATECTOMY PROCEDURE, THE CUSTOMER HAD DIFFICULTY WHEN LOADING THE RELOAD INTO THE HANDLE. THE DEVICE WAS ABLE TO BE FIRED, HOWEVER, AFTER FIRING THERE WAS RESISTANCE WHEN THE BLACK RETURN KNOB WAS RETRACTED. THE RELOAD WAS FORCIBLY REMOVED, SO THE RELOAD JAWS GOT DAMAGED AND THAT WAS WHY THE JAWS COULD NOT CLOSE COMPLETELY. IN ADDITION, THE DEVICE WAS DIFFICULT TO UNLOAD BECAUSE THE RELEASE BUTTON WAS TOO STIFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL HEPATECTOMY PROCEDURE, THE CUSTOMER HAD DIFFICULTY WHEN LOADING THE RELOAD INTO THE HANDLE. THE DEVICE WAS ABLE TO BE FIRED, HOWEVER, AFTER FIRING THERE WAS RESISTANCE WHEN THE BLACK RETURN KNOB WAS RETRACTED. THE RELOAD WAS FORCIBLY REMOVED, SO THE RELOAD JAWS GOT DAMAGED AND THAT WAS WHY THE JAWS COULD NOT CLOSE COMPLETELY. IN ADDITION, THE DEVICE WAS DIFFICULT TO UNLOAD BECAUSE THE RELEASE BUTTON WAS TOO STIFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL HEPATECTOMY PROCEDURE, THE CUSTOMER HAD DIFFICULTY WHEN LOADING THE RELOAD INTO THE HANDLE. THE DEVICE WAS ABLE TO BE FIRED, HOWEVER, AFTER FIRING THERE WAS RESISTANCE WHEN THE BLACK RETURN KNOB WAS RETRACTED. THE RELOAD WAS FORCIBLY REMOVED, SO THE RELOAD JAWS GOT DAMAGED AND THAT WAS WHY THE JAWS COULD NOT CLOSE COMPLETELY. IN ADDITION, THE DEVICE WAS DIFFICULT TO UNLOAD BECAUSE THE RELEASE BUTTON WAS TOO STIFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL HEPATECTOMY PROCEDURE, THE CUSTOMER HAD DIFFICULTY WHEN LOADING THE RELOAD INTO THE HANDLE. THE DEVICE WAS ABLE TO BE FIRED, HOWEVER, AFTER FIRING THERE WAS RESISTANCE WHEN THE BLACK RETURN KNOB WAS RETRACTED. THE RELOAD WAS FORCIBLY REMOVED, SO THE RELOAD JAWS GOT DAMAGED AND THAT WAS WHY THE JAWS COULD NOT CLOSE COMPLETELY. IN ADDITION, THE DEVICE WAS DIFFICULT TO UNLOAD BECAUSE THE RELEASE BUTTON WAS TOO STIFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | us surgical puerto rico | 29-01-2021 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL HEPATECTOMY PROCEDURE, THE CUSTOMER HAD DIFFICULTY WHEN LOADING THE RELOAD INTO THE HANDLE. THE DEVICE WAS ABLE TO BE FIRED, HOWEVER, AFTER FIRING THERE WAS RESISTANCE WHEN THE BLACK RETURN KNOB WAS RETRACTED. THE RELOAD WAS FORCIBLY REMOVED, SO THE RELOAD JAWS GOT DAMAGED AND THAT WAS WHY THE JAWS COULD NOT CLOSE COMPLETELY. IN ADDITION, THE DEVICE WAS DIFFICULT TO UNLOAD BECAUSE THE RELEASE BUTTON WAS TOO STIFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | us surgical puerto rico | 29-01-2021 | ENDO GIA ULTRA | Difficult or Delayed Activation | A150102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL HEPATECTOMY PROCEDURE, THE CUSTOMER HAD DIFFICULTY WHEN LOADING THE RELOAD INTO THE HANDLE. THE DEVICE WAS ABLE TO BE FIRED, HOWEVER, AFTER FIRING THERE WAS RESISTANCE WHEN THE BLACK RETURN KNOB WAS RETRACTED. THE RELOAD WAS FORCIBLY REMOVED, SO THE RELOAD JAWS GOT DAMAGED AND THAT WAS WHY THE JAWS COULD NOT CLOSE COMPLETELY. IN ADDITION, THE DEVICE WAS DIFFICULT TO UNLOAD BECAUSE THE RELEASE BUTTON WAS TOO STIFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | us surgical puerto rico | 29-01-2021 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL HEPATECTOMY PROCEDURE, THE CUSTOMER HAD DIFFICULTY WHEN LOADING THE RELOAD INTO THE HANDLE. THE DEVICE WAS ABLE TO BE FIRED, HOWEVER, AFTER FIRING THERE WAS RESISTANCE WHEN THE BLACK RETURN KNOB WAS RETRACTED. THE RELOAD WAS FORCIBLY REMOVED, SO THE RELOAD JAWS GOT DAMAGED AND THAT WAS WHY THE JAWS COULD NOT CLOSE COMPLETELY. IN ADDITION, THE DEVICE WAS DIFFICULT TO UNLOAD BECAUSE THE RELEASE BUTTON WAS TOO STIFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | us surgical puerto rico | 29-01-2021 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL HEPATECTOMY PROCEDURE, THE CUSTOMER HAD DIFFICULTY WHEN LOADING THE RELOAD INTO THE HANDLE. THE DEVICE WAS ABLE TO BE FIRED, HOWEVER, AFTER FIRING THERE WAS RESISTANCE WHEN THE BLACK RETURN KNOB WAS RETRACTED. THE RELOAD WAS FORCIBLY REMOVED, SO THE RELOAD JAWS GOT DAMAGED AND THAT WAS WHY THE JAWS COULD NOT CLOSE COMPLETELY. IN ADDITION, THE DEVICE WAS DIFFICULT TO UNLOAD BECAUSE THE RELEASE BUTTON WAS TOO STIFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | us surgical puerto rico | 29-01-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ABDOMINAL HEPATECTOMY PROCEDURE, THE CUSTOMER HAD DIFFICULTY WHEN LOADING THE RELOAD INTO THE HANDLE. THE DEVICE WAS ABLE TO BE FIRED, HOWEVER, AFTER FIRING THERE WAS RESISTANCE WHEN THE BLACK RETURN KNOB WAS RETRACTED. THE RELOAD WAS FORCIBLY REMOVED, SO THE RELOAD JAWS GOT DAMAGED AND THAT WAS WHY THE JAWS COULD NOT CLOSE COMPLETELY. IN ADDITION, THE DEVICE WAS DIFFICULT TO UNLOAD BECAUSE THE RELEASE BUTTON WAS TOO STIFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-04-2015 | Injury | covidien lp llc north haven | 29-01-2021 | UNKNOWN EEA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLER CUT THE TISSUE BUT DID NOT DEPLOY STAPLES. THERE WAS TISSUE LOSS AS A RESULT OF THE EVENT.
25-08-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, AT THE MOMENT OF USING THE RELOAD, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER MADE THE TEST BUT THE CLAMP DID NOT ACTIVATE. IT WAS ALSO STATED THAT A COMPONENT FELL INTO THE PATIENT¿S CAVITY. ANOTHER RELOAD WITH THE SAME REFERENCE WAS USED.
25-08-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AT THE MOMENT OF USING THE RELOAD, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER MADE THE TEST BUT THE CLAMP DID NOT ACTIVATE. IT WAS ALSO STATED THAT A COMPONENT FELL INTO THE PATIENT¿S CAVITY. ANOTHER RELOAD WITH THE SAME REFERENCE WAS USED.
25-08-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, AT THE MOMENT OF USING THE RELOAD, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER MADE THE TEST BUT THE CLAMP DID NOT ACTIVATE. IT WAS ALSO STATED THAT A COMPONENT FELL INTO THE PATIENT¿S CAVITY. ANOTHER RELOAD WITH THE SAME REFERENCE WAS USED.
25-08-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AT THE MOMENT OF USING THE RELOAD, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER MADE THE TEST BUT THE CLAMP DID NOT ACTIVATE. IT WAS ALSO STATED THAT A COMPONENT FELL INTO THE PATIENT¿S CAVITY. ANOTHER RELOAD WITH THE SAME REFERENCE WAS USED.
09-01-2021 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AT THE BEGINNING OF A SIGMOIDECTOMY, AFTER TESTING THE STAPLER, RELOAD WAS CLOSED AS IT ENTERS THE TROCAR BUT IT JAWS WOULD NOT RE-OPEN. SURGEON HAD TO FORCE TO REMOVE THE RELOAD. SURGEON USED MANUAL STAPLER TO COMPLETE THE CASE. DEVICE WAS NOT USED ON PATIENT. NO PATIENT INJURY.
10-11-2014 | Malfunction | covidien lp llc north haven | 29-01-2021 | UNKNOWN TRI STAPLE PRODUCT | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. IT WAS LOCKED. ADDITIONALLY, IT WAS STATED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE.
10-11-2014 | Malfunction | covidien lp llc north haven | 29-01-2021 | UNKNOWN TRI STAPLE PRODUCT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. IT WAS LOCKED. ADDITIONALLY, IT WAS STATED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE.
16-11-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA SURGERY, ON MESH FIXATION, UNUSUAL NOISE WAS HEARD WHEN THE SURGEON FIRED AND THE TACKS DID NOT FIXATE THE MESH. USED ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
16-11-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | ABSORBATACK | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA SURGERY, ON MESH FIXATION, UNUSUAL NOISE WAS HEARD WHEN THE SURGEON FIRED AND THE TACKS DID NOT FIXATE THE MESH. USED ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
08-01-2021 | Malfunction | us surgical puerto rico | 29-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING ILEOSTOMY REPAIR, THE PRODUCT HAS NOT BEEN LAUNCHED. ANOTHER PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-10-2020 | Malfunction | us surgical puerto rico | 29-01-2021 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC PROCEDURE, ONE OF THE REPORTED HANDLES HAD AN ISSUE WITH ITS ARTICULATION KNOB. THE ARTICULATION KNOB BROKE OFF AND IT WAS NOTED THAT THE OTHER HANDLE COULD NOT BE LOADED. THE DEVICE WAS NOT USED ON PATIENT. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO SHEARED FIRING RACK TEETH DAMAGE.
10-10-2020 | Malfunction | us surgical puerto rico | 29-01-2021 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC PROCEDURE, ONE OF THE REPORTED HANDLES HAD AN ISSUE WITH ITS ARTICULATION KNOB. THE ARTICULATION KNOB BROKE OFF AND IT WAS NOTED THAT THE OTHER HANDLE COULD NOT BE LOADED. THE DEVICE WAS NOT USED ON PATIENT. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO SHEARED FIRING RACK TEETH DAMAGE.
08-01-2021 | Malfunction | us surgical puerto rico | 29-01-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING ILEOSTOMY REPAIR, THE PRODUCT HAS NOT BEEN LAUNCHED. ANOTHER PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2021 | Malfunction | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, ON THE LUNG TISSUE, THE STAPLER WAS ABLE TO FIRE THE 2 RELOADS AND THERE WAS POOR STAPLE FORMATION. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-11-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, OXIDIZED CONTACTS WERE DISCOVERED. THE DEVICE WAS NOT CHARGING. ANOTHER HANDLE AND CHARGER WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID.
03-11-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, OXIDIZED CONTACTS WERE DISCOVERED. THE DEVICE WAS NOT CHARGING. ANOTHER HANDLE AND CHARGER WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID.
03-11-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, OXIDIZED CONTACTS WERE DISCOVERED. THE DEVICE WAS NOT CHARGING. ANOTHER HANDLE AND CHARGER WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND VENTED BATTERIES, CAUSING THEM TO LEAK ELECTROLYTIC FLUID.
11-01-2021 | Malfunction | celestica de monterrey s.a de c.v | 29-01-2021 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, ON THE LUNG TISSUE, THE STAPLER WAS ABLE TO FIRE THE TWO RELOADS AND THERE WAS POOR STAPLE FORMATION. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-10-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CHARGING INDICATOR ILLUMINATED AND FLASHED, AND OCCASIONALLY RED INDICATOR ILLUMINATED, THE HANDLE WOULD NOT START UP. THE REPORTED HANDLE WAS INSERTED INTO ANOTHER CHARGER AS A TRIAL, CHECK WAS PERFORMED AFTER GREEN INDICATOR ILLUMINATED, BUT THE HANDLE WOULD NOT START UP. ANOTHER HANDLE WAS INSERTED TO THE REPORTED CHARGER AS A TRIAL, IT WAS ABLE TO BE CHARGED AND STARTED UP WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INVOLVEMENT.
26-10-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CHARGING INDICATOR ILLUMINATED AND FLASHED, AND OCCASIONALLY RED INDICATOR ILLUMINATED, THE HANDLE WOULD NOT START UP. THE REPORTED HANDLE WAS INSERTED INTO ANOTHER CHARGER AS A TRIAL, CHECK WAS PERFORMED AFTER GREEN INDICATOR ILLUMINATED, BUT THE HANDLE WOULD NOT START UP. ANOTHER HANDLE WAS INSERTED TO THE REPORTED CHARGER AS A TRIAL, IT WAS ABLE TO BE CHARGED AND STARTED UP WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INVOLVEMENT.
26-10-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CHARGING INDICATOR ILLUMINATED AND FLASHED, AND OCCASIONALLY RED INDICATOR ILLUMINATED, THE HANDLE WOULD NOT START UP. THE REPORTED HANDLE WAS INSERTED INTO ANOTHER CHARGER AS A TRIAL, CHECK WAS PERFORMED AFTER GREEN INDICATOR ILLUMINATED, BUT THE HANDLE WOULD NOT START UP. ANOTHER HANDLE WAS INSERTED TO THE REPORTED CHARGER AS A TRIAL, IT WAS ABLE TO BE CHARGED AND STARTED UP WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INVOLVEMENT.
16-10-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WOULD NOT CHARGE. THE HANDLE DID NOT GIVE ANY IGNITION SIGNAL OR GREEN LIGHT INDICATION WHILE ON THE CHARGER. IT WAS ALSO NOT SHOWING ANY DISPLAY ON THE SCREEN. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHEN THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
16-10-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WOULD NOT CHARGE. THE HANDLE DID NOT GIVE ANY IGNITION SIGNAL OR GREEN LIGHT INDICATION WHILE ON THE CHARGER. IT WAS ALSO NOT SHOWING ANY DISPLAY ON THE SCREEN. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHEN THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
16-10-2020 | Malfunction | covidien lp llc north haven | 29-01-2021 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WOULD NOT CHARGE. THE HANDLE DID NOT GIVE ANY IGNITION SIGNAL OR GREEN LIGHT INDICATION WHILE ON THE CHARGER. IT WAS ALSO NOT SHOWING ANY DISPLAY ON THE SCREEN. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHEN THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
01-01-2020 | Malfunction | ethicon endo-surgery | 29-01-2021 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE ANVIL WOULD NOT ATTACH TO THE DEVICE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A DELAY OF FIVE MINUTES. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
05-01-2021 | Malfunction | ethicon endo-surgery | 29-01-2021 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC TOTAL GASTRECTOMY SURGERY, COULD NOT FIRE WHEN SEVERING JEJUNO-OESOPHAGUS TISSUE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2021 | Malfunction | ethicon endo-surgery | 29-01-2021 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOWER ANTERIOR BOWEL RESECTION PROCEDURE THAT DEVICE WOULD NOT FIRE. THE BATTERY WAS INSERTED INTO DEVICE, GREEN CHECK WAS BACKLIGHT INDICATING FUNCTIONING BATTERY. SURGEON CONNECTED THE ANVIL SPIKE TO TROCAR (NO ORANGE WAS VISIBLE). SURGEON ROTATED THE KNOB TO CLOSE THE DEVICE UNTIL THE ORANGE LINE WAS IN THE GREEN FIRING ZONE. SURGEON PULLED BACK SAFETY AND WENT TO PULL THE FIRING TRIGGER AND THE DEVICE DID NOT FIRE. SURGEON RE-INSERTED BATTERY, STILL DID NOT FIRE. SURGEON OPENED NEW DEVICE (CDH29P FROM SAME LOT #) AND INSERTED JUST THE NEW BATTERY, DEVICE STILL DID NOT FIRE. SURGEON HAD TO DETACH TROCAR & ANVIL, AND REINSERT NEW DEVICE W/ NEW BATTERY TO COMPLETE THE PROCEDURE. THAT DEVICE WORKED AND THERE WAS NO ADVERSE PATIENT CONSEQUENCE INTRA-OPERATIVELY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY BUT WAS DELAYED FOR AN UNKNOWN AMOUNT OF TIME. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
31-12-2020 | Malfunction | ethicon endo-surgery | 29-01-2021 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE OPEN GASTRIC TUMOR RESECTION SURGERY, WHEN SEVERING STOMACH TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE( WITH STRAIGHT LEG). CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-01-2021 | Injury | covidien lp llc north haven | 29-01-2021 | ENDO GIA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE THIRD FIRING, THE STAPLER WAS NOT ABLE TO FORM A B SHAPE. THE BLADE WAS ADVANCED BUT DID NOT CUT THROUGH THE STOMACH, IT PUSHED THE TISSUE OUT AS THE BLADE WAS FIRING. THERE WAS TEAR OF THE SEROSA LAYER OF THE STOMACH AND ANOTHER STAPLER WAS FIRED OVER IT. IT WAS ALSO NOTED THAT THE OTHER RELOAD WAS NOT ABLE TO FIRE. THE SURGICAL TIME WAS ALSO EXTENDED BY MORE THAN 30 MINUTES DUE TO REMOVING THE UNFIRED STAPLER FROM THE TISSUE AND TIME NEED TO TEST THE DEVICE OUT OF THE PATIENT'S BODY. NEW RELOADS WERE USED TO RESOLVE THE ISSUE.
25-12-2020 | Malfunction | us surgical puerto rico | 29-01-2021 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE THREE RELOADS JAWS WOULD NOT CLOSE AND OPEN. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-08-2020 | Injury | ethicon endo-surgery | 29-01-2021 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE SURGERY OF LAPAROSCOPIC RADICAL GASTRECTOMY FOR GASTRIC CANCER, WHEN SEVERING THE INTESTINE TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE THE SURGERY. AFTER THE SURGERY, THE PATIENT'S ANASTOMOTIC SITE BLED. REINFORCEMENT OF FLUID REPLACEMENT AND FASTING TREATMENT WERE GIVEN. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-08-2020 | Injury | ethicon endo-surgery | 29-01-2021 | ILS 25MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE SURGERY OF LAPAROSCOPIC RADICAL GASTRECTOMY FOR GASTRIC CANCER, WHEN SEVERING THE INTESTINE TISSUE, AFTER FIRED, NOTED SOME STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE THE SURGERY. AFTER THE SURGERY, THE PATIENT'S ANASTOMOTIC SITE BLED. REINFORCEMENT OF FLUID REPLACEMENT AND FASTING TREATMENT WERE GIVEN. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2020 | Malfunction | ethicon endo-surgery | 29-01-2021 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE THE DEVICE COULD NOT BE CLOSED THEREFORE COULD NOT BE FIRED. THE DEVICE COULD BE EASILY REMOVED. NO SIGNIFICANT DELAY IN SURGERY; NO HARM TO THE PATIENT.
01-01-2020 | Malfunction | ethicon endo-surgery | 29-01-2021 | CONTOUR CURVED CUTTER STAPLER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE THE DEVICE COULD NOT BE CLOSED THEREFORE COULD NOT BE FIRED. THE DEVICE COULD BE EASILY REMOVED. NO SIGNIFICANT DELAY IN SURGERY; NO HARM TO THE PATIENT.
11-01-2021 | Malfunction | us surgical puerto rico | 29-01-2021 | EEA | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN LOW ANTERIOR RESECTION, BEFORE FIRING, ANVIL CANNOT CONNECT WITH CIRCULAR STAPLER, IT WAS STUCK. THE STAPLER WAS INTACT, CANNOT FIRE. THE SURGEON CHANGED THE DEVICE WITH ANOTHER ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-01-2021 | Malfunction | us surgical puerto rico | 29-01-2021 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN LOW ANTERIOR RESECTION, BEFORE FIRING, ANVIL CANNOT CONNECT WITH CIRCULAR STAPLER, IT WAS STUCK. THE STAPLER WAS INTACT, CANNOT FIRE. THE SURGEON CHANGED THE DEVICE WITH ANOTHER ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | ethicon endo-surgery | 29-01-2021 | CONTOUR CURVED CUTTER STAPLER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE THE DEVICE COULD NOT BE CLOSED THEREFORE COULD NOT BE FIRED. THE DEVICE COULD BE EASILY REMOVED. NO SIGNIFICANT DELAY IN SURGERY; NO HARM TO THE PATIENT.
27-11-2020 | Malfunction | ethicon endo-surgery | 29-01-2021 | CONTOUR TRANSTAR STAPLER SET | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AFTER RELOAD THE THORN COULD NOT BE EXTENDED. A NEW INSTRUMENT HAD TO BE OPENED. AFTER RELOADING, THE RETAINING PIN DID NOT COME OUT. A NEW DEVICE HAD TO BE OPENED. NO PATIENT HARM NOR SIGNIFICANT SURGICAL DELAY REPORTED. PROCEDURE RECTUM PROLAPSE
11-01-2021 | Malfunction | us surgical puerto rico | 01-02-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE TRANSECTING THE GASTRIC TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-04-2021 | Malfunction | celestica de monterrey s.a de c.v | 01-02-2021 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY AND WHILE ASSEMBLING, THE DEVICE WAS DIFFICULT TO LOAD, AND THE DISTAL PART OF THE RELOAD FELL OFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
15-04-2021 | Malfunction | celestica de monterrey s.a de c.v | 01-02-2021 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY AND WHILE ASSEMBLING, THE DEVICE WAS DIFFICULT TO LOAD, AND THE DISTAL PART OF THE RELOAD FELL OFF. THE SURGEON THEN USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
08-01-2021 | Injury | us surgical puerto rico | 01-02-2021 | TA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOWER ANTERIOR RESECTION, DURING THE ANASTOMOSIS OF THE PROXIMAL COLON AND RECTUM, THE DEVICE FIRED BUT THE STAPLES DID NOT DEPLOY. THE STAPLE LINE WAS MANUALLY SUTURED.
11-01-2021 | Malfunction | us surgical puerto rico | 01-02-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE OF THE CLIP APPLIER. ADDITIONALLY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON OF THE STAPLER BUT THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THE STAPLER DID NOT FIRE. IT WAS ALSO STATED THAT TWO RELOADS HAD THE SAME MALFUNCTION. A DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-01-2021 | Malfunction | us surgical puerto rico | 01-02-2021 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE BEING APPLIED TO THE STOMACH, THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE WAS NOT SQUEEZED, HENCE THE STAPLER DID NOT FIRE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-01-2021 | Malfunction | us surgical puerto rico | 01-02-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. THE SURGEON THEN USED ANOTHER DEVICE'S RELOAD TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-01-2021 | Malfunction | us surgical puerto rico | 01-02-2021 | GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE SURGEON SQUEEZED THE HANDLE, BUT THE DEVICE GOT STUCK AND STOPPED MOVING DURING FIRING.
11-01-2021 | Injury | us surgical puerto rico | 01-02-2021 | GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, WHEN DETACHING STOMACH TISSUE, STAPLE WAS FORMED POORLY AT THE MIDDLE PART OF THE TISSUE, AND THERE WAS BLEEDING NOT REACHING 500CC. INCOMPLETE STAPLE LINE WAS ALSO NOTED. RESECTION WAS PERFORMED, AND ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
11-01-2021 | Malfunction | us surgical puerto rico | 01-02-2021 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TRANSECTION OF STOMACH IN AN OPEN GASTRECTOMY, THE RELOAD CANNOT BE FIRED AFTER LOADING ON THE STAPLER. SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
11-01-2021 | Injury | us surgical puerto rico | 01-02-2021 | EEA | Detachment of Device or Device Component | A0501 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PELVIC EXENTERATION PROCEDURE, THE FIRST STAPLER FIRED BUT THE STAPLE LINE WAS TORN CAUSING A HOLE ON THE SIGMOID, AND IT WAS NOTED THAT THE STAPLER WAS ALSO BROKEN. THE SECOND STAPLER FIRED AS WELL, BUT THE ANVIL DISENGAGED FROM THE SPIKE AND EXITED FROM THE LUMEN. THERE WAS TISSUE LOSS AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. THE SURGICAL PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES DUE TO RETRIEVAL OF THE ANVIL, AND TISSUE RESECTION WAS PERFORMED TO RESOLVE THE ISSUE.
11-01-2021 | Injury | us surgical puerto rico | 01-02-2021 | EEA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PELVIC EXENTERATION PROCEDURE, THE FIRST STAPLER FIRED BUT THE STAPLE LINE WAS TORN CAUSING A HOLE ON THE SIGMOID, AND IT WAS NOTED THAT THE STAPLER WAS ALSO BROKEN. THE SECOND STAPLER FIRED AS WELL, BUT THE ANVIL DISENGAGED FROM THE SPIKE AND EXITED FROM THE LUMEN. THERE WAS TISSUE LOSS AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. THE SURGICAL PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES DUE TO RETRIEVAL OF THE ANVIL, AND TISSUE RESECTION WAS PERFORMED TO RESOLVE THE ISSUE.
11-01-2021 | Injury | celestica de monterrey s.a de c.v | 01-02-2021 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC THORACOSCOPIC SEGMENTECTOMY, WHILE DETACHING INTERLOBULAR FISSURE, THE DEVICE WAS ABLE TO FIRE BUT THE STAPLES DID NOT FORM PROPERLY AT THE DISTAL END OF THE CUTTING LINE. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THE SURGEON ADDITIONALLY SUTURED WITH SUTURES.
11-01-2021 | Malfunction | covidien lp llc north haven | 01-02-2021 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF ESOPHAGUS IN AN ESOPHAGECTOMY, THE ANVIL OF THE DEVICE DISENGAGED. SURGEON USED ANOTHER ANVIL TO RESOLVE THE ISSUE. NO PATIENT INJURY.
01-02-2021 | Malfunction | us surgical puerto rico | 01-02-2021 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. IT WAS ALSO NOTED THAT THE RELOAD WAS NOT ABLE TO BE USED WITH THE NEW HANDLE. THERE WAS NO PATIENT INJURY.
11-01-2021 | Malfunction | covidien lp llc north haven | 01-02-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE OF THE CLIP APPLIER. ADDITIONALLY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON OF THE STAPLER BUT THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. THE STAPLER DID NOT FIRE. IT WAS ALSO STATED THAT TWO RELOADS HAD THE SAME MALFUNCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-01-2021 | Malfunction | covidien lp llc north haven | 01-02-2021 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE BEING APPLIED TO THE STOMACH, THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE WAS NOT SQUEEZED HENCE THE STAPLER DID NOT FIRE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2015 | Malfunction | covidien lp llc north haven | 01-02-2021 | UNKNOWN DST TA STAPLER | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER FAILED AFTER THE SECOND LOAD WAS USED, AND THE STAPLES WERE PARTIALLY FIRED. THERE WAS NO PATIENT INJURY.
11-01-2021 | Injury | us surgical puerto rico | 01-02-2021 | EEA | Break | A0401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PELVIC EXENTERATION PROCEDURE, THE FIRST STAPLER FIRED BUT THE STAPLE LINE WAS TORN CAUSING A HOLE ON THE SIGMOID. AND IT WAS NOTED THAT THE STAPLER WAS ALSO BROKEN. THE SECOND STAPLER FIRED AS WELL, BUT THE ANVIL DISENGAGED FROM THE SPIKE AND EXITED FROM THE LUMEN. THERE WAS TISSUE LOSS AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. THE SURGICAL PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES DUE TO RETRIEVAL OF THE ANVIL AND TISSUE RESECTION WAS PERFORMED TO RESOLVE THE ISSUE.
11-01-2021 | Injury | us surgical puerto rico | 01-02-2021 | EEA | Break | A0401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PELVIC EXENTERATION PROCEDURE, THE FIRST STAPLER FIRED BUT THE STAPLE LINE WAS TORN CAUSING A HOLE ON THE SIGMOID. AND IT WAS NOTED THAT THE STAPLER WAS ALSO BROKEN. THE SECOND STAPLER FIRED AS WELL, BUT THE ANVIL DISENGAGED FROM THE SPIKE AND EXITED FROM THE LUMEN. THERE WAS TISSUE LOSS AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. THE SURGICAL PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES DUE TO RETRIEVAL OF THE ANVIL AND TISSUE RESECTION WAS PERFORMED TO RESOLVE THE ISSUE.
11-01-2021 | Injury | us surgical puerto rico | 01-02-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PELVIC EXENTERATION PROCEDURE, THE FIRST STAPLER FIRED BUT THE STAPLE LINE WAS TORN CAUSING A HOLE ON THE SIGMOID. AND IT WAS NOTED THAT THE STAPLER WAS ALSO BROKEN. THE SECOND STAPLER FIRED AS WELL, BUT THE ANVIL DISENGAGED FROM THE SPIKE AND EXITED FROM THE LUMEN. THERE WAS TISSUE LOSS AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. THE SURGICAL PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES DUE TO RETRIEVAL OF THE ANVIL AND TISSUE RESECTION WAS PERFORMED TO RESOLVE THE ISSUE.
11-01-2021 | Injury | us surgical puerto rico | 01-02-2021 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PELVIC EXENTERATION PROCEDURE, THE FIRST STAPLER FIRED BUT THE STAPLE LINE WAS TORN CAUSING A HOLE ON THE SIGMOID. AND IT WAS NOTED THAT THE STAPLER WAS ALSO BROKEN. THE SECOND STAPLER FIRED AS WELL, BUT THE ANVIL DISENGAGED FROM THE SPIKE AND EXITED FROM THE LUMEN. THERE WAS TISSUE LOSS AND TISSUE DAMAGE DUE TO THE REPORTED EVENT. THE SURGICAL PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES DUE TO RETRIEVAL OF THE ANVIL AND TISSUE RESECTION WAS PERFORMED TO RESOLVE THE ISSUE.
02-11-2020 | Malfunction | covidien lp llc north haven | 01-02-2021 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHILE CHARGING, THE BATTERY CHARGER WAS FOUND TO BE FLASHING RED. AS A RESULT, THE POWER HANDLE DID NOT CHARGE. ALTHOUGH, IT WAS INSERTED INTO ANOTHER CHARGER, THE RED INDICATOR ILLUMINATED WITH THE REPORTED HANDLE. ON THE OTHER HAND, ANOTHER POWER HANDLE CHARGED WITHOUT ISSUE WITH THE REPORTED BATTERY CHARGER. THERE WAS NO PATIENT INVOLVEMENT.
02-11-2020 | Malfunction | covidien lp llc north haven | 01-02-2021 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHILE CHARGING, THE BATTERY CHARGER WAS FOUND TO BE FLASHING RED. AS A RESULT, THE POWER HANDLE DID NOT CHARGE. ALTHOUGH, IT WAS INSERTED INTO ANOTHER CHARGER, THE RED INDICATOR ILLUMINATED WITH THE REPORTED HANDLE. ON THE OTHER HAND, ANOTHER POWER HANDLE CHARGED WITHOUT ISSUE WITH THE REPORTED BATTERY CHARGER. THERE WAS NO PATIENT INVOLVEMENT.
14-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREAS PROCEDURE, AT THE TIME OF PANCREATECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS A PROBLEM UNLOADING THE DEVICE, THE LOADER WAS BLOCKED ON THE PANCREAS AND WAS LEFT IN ANALYSIS WITH THE ORGAN BECAUSE IT WAS IMPOSSIBLE TO OPEN. THERE WAS TISSUE DAMAGE NOTED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. THE SURGEON HAD TO MAKE A WIDER RESECTION TO FREE THE OPERATIVE PART WITH THE LOADER TO RESOLVE THE CONCERN.
14-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREAS PROCEDURE, AT THE TIME OF PANCREATECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS A PROBLEM UNLOADING THE DEVICE, THE LOADER WAS BLOCKED ON THE PANCREAS AND WAS LEFT IN ANALYSIS WITH THE ORGAN BECAUSE IT WAS IMPOSSIBLE TO OPEN. THERE WAS TISSUE DAMAGE NOTED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. THE SURGEON HAD TO MAKE A WIDER RESECTION TO FREE THE OPERATIVE PART WITH THE LOADER TO RESOLVE THE CONCERN.
14-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREAS PROCEDURE, AT THE TIME OF PANCREATECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS A PROBLEM UNLOADING THE DEVICE, THE LOADER WAS BLOCKED ON THE PANCREAS AND WAS LEFT IN ANALYSIS WITH THE ORGAN BECAUSE IT WAS IMPOSSIBLE TO OPEN. THERE WAS TISSUE DAMAGE NOTED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. THE SURGEON HAD TO MAKE A WIDER RESECTION TO FREE THE OPERATIVE PART WITH THE LOADER TO RESOLVE THE CONCERN.
14-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREAS PROCEDURE, AT THE TIME OF PANCREATECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS A PROBLEM UNLOADING THE DEVICE, THE LOADER WAS BLOCKED ON THE PANCREAS AND WAS LEFT IN ANALYSIS WITH THE ORGAN BECAUSE IT WAS IMPOSSIBLE TO OPEN. THERE WAS TISSUE DAMAGE NOTED DUE TO THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. THE SURGEON HAD TO MAKE A WIDER RESECTION TO FREE THE OPERATIVE PART WITH THE LOADER TO RESOLVE THE CONCERN.
16-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, AFTER TIME STAPLING IN THE LEFT PULMONARY VEIN DURING A LAPAROSCOPIC LEFT LOWER LOBECTOMY, THE DEVICE HAD A POOR STAPLE FORMATION AND AN INCOMPLETE STAPLE LINE. THE DEVICE DID NOT CLOSE THE TISSUE WAY IT SHOULD HAVE DONE IT. THE PATIENT BLEED A LOT AND REQUIRED TWO UNITS OF BLOOD TRANSFUSION. THE SURGICAL TIME GOT EXTENDED FOR MORE THAN THIRTY MINUTES. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE, SUTURE AND CLIPS WERE ALSO REQUIRED.
21-10-2021 | Injury | covidien | 03-01-2022 | ENDO GIA | Misfire | A050502 | Laceration(s) | Event Description: WHILE PERFORMING A LAPAROSCOPIC APPENDECTOMY, THE ENDO GIA STAPLER 45MM -2.5MM MISFIRED AND DID NOT STAPLE BUT DID CUT. THIS RESULTED IN OPEN BOWEL AND THE SURGERY HAD TO CONVERT TO AN OPEN PROCEDURE. THE STAPLER HANDLE WAS PASSED OFF THE FIELD AND BAGGED FOR INVESTIGATION.
21-10-2021 | Injury | covidien | 03-01-2022 | ENDO GIA | Patient Device Interaction Problem | A01 | Laceration(s) | Event Description: WHILE PERFORMING A LAPAROSCOPIC APPENDECTOMY, THE ENDO GIA STAPLER 45MM -2.5MM MISFIRED AND DID NOT STAPLE BUT DID CUT. THIS RESULTED IN OPEN BOWEL AND THE SURGERY HAD TO CONVERT TO AN OPEN PROCEDURE. THE STAPLER HANDLE WAS PASSED OFF THE FIELD AND BAGGED FOR INVESTIGATION.
11-08-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, PRIOR TO PATIENT CONTACT, AN ERROR MESSAGE SHOWED UP ON THE SCREEN OF THE HANDLE WAS A RED CIRCLE AND LINE THROUGH IT WITH A YELLOW TRIANGLE ABOVE THE RED CIRCLE. IT WAS REPLACED WITH ANOTHER ONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
11-08-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, PRIOR TO PATIENT CONTACT, AN ERROR MESSAGE SHOWED UP ON THE SCREEN OF THE HANDLE WAS A RED CIRCLE AND LINE THROUGH IT WITH A YELLOW TRIANGLE ABOVE THE RED CIRCLE. IT WAS REPLACED WITH ANOTHER ONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE THE DEVICE DID NOT FIRE. THE USER WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE THE DEVICE DID NOT FIRE. THE USER WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE APPLYING ON THE STOMACH TISSUE AFTER THE THIRD FIRING, THE TWO STAPLERS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT A CRACKING SOUND WAS HEARD WHEN SQUEEZING/FIRING AND A COMPONENT BROKE OFF. A 3RD STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE APPLYING ON THE STOMACH TISSUE AFTER THE THIRD FIRING, THE TWO STAPLERS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT A CRACKING SOUND WAS HEARD WHEN SQUEEZING/FIRING AND A COMPONENT BROKE OFF. A 3RD STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE APPLYING ON THE STOMACH TISSUE AFTER THE THIRD FIRING, THE TWO STAPLERS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT A CRACKING SOUND WAS HEARD WHEN SQUEEZING/FIRING AND A COMPONENT BROKE OFF. A 3RD STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Injury | us surgical puerto rico | 03-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE PORT THORACOSCOPIC LOBECTOMY, WHEN DETACHING THE LUNG PARENCHYMA, THE POOR STAPLE FORMATION, IT COULD NOT FORM INTO "B" SHAPE. THE STAPLE LINE WAS ALSO INCOMPLETE CENTRALLY. SUTURING WAS PERFORMED TO FIX THE STAPLE LINE AND A NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE APPLYING ON THE STOMACH TISSUE AFTER THE THIRD FIRING, THE TWO HANDLES AND TWO RELOADS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT A CRACKING SOUND WAS HEARD WHEN SQUEEZING/FIRING AND GEARS IN THE HANDLE BROKE OFF. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE APPLYING ON THE STOMACH TISSUE AFTER THE THIRD FIRING, THE TWO HANDLES AND TWO RELOADS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT A CRACKING SOUND WAS HEARD WHEN SQUEEZING/FIRING AND GEARS IN THE HANDLE BROKE OFF. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHILE APPLYING ON THE STOMACH TISSUE AFTER THE THIRD FIRING, THE TWO HANDLES AND TWO RELOADS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT A CRACKING SOUND WAS HEARD WHEN SQUEEZING/FIRING AND GEARS IN THE HANDLE BROKE OFF. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, AFTER THE HANDLE WAS LOADED WITH A 30 MILLIMETER TAN RELOAD THEN PASSED TO THE STERILE FIELD, IT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED AND THE SAME RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY OF A LAPAROSCOPIC RIGHT COLECTOMY, ON STAPLING THE BOWEL, THERE WAS BLEEDING FROM THE STAPLE LINE. THERE WAS BLOOD LOSS OF 500CC OR MORE DUE TO THE PRODUCT PROBLEM AND THE PATIENT RECEIVED A TRANSFUSION OF 3 UNITS OF BLOOD.
03-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY OF A LAPAROSCOPIC RIGHT COLECTOMY, ON STAPLING THE BOWEL, THERE WAS BLEEDING FROM THE STAPLE LINE. THERE WAS BLOOD LOSS OF 500CC OR MORE DUE TO THE PRODUCT PROBLEM AND THE PATIENT RECEIVED A TRANSFUSION OF 3 UNITS OF BLOOD.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN EVENT, THE DEVICE WAS UNABLE TO FIRE, THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE CANNOT BE SQUEEZED.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN EVENT, THE DEVICE WAS UNABLE TO FIRE, THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE CANNOT BE SQUEEZED.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE.
16-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, THE DEVICE FIRING COULD NOT BE PERFORMED. RELOAD WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
08-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
08-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD STAPLE LINE BLEEDING AFTER PERFORMING LAPAROSCOPIC RIGHT COLECTOMY. SINCE THE BLEEDING HAD STOPPED, TRANSFUSION WAS NOT REQUIRED. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, AFTER A LAPAROSCOPIC COLECTOMY PROCEDURE, THE PATIENT HAD BLEEDING FROM THE STAPLE LINE. THE BLEEDING IS STOPPED TO RESOLVE THE ISSUE. PATIENT HAD SMALL DOSE OF HEPARIN AND A DOSE OF TORADOL.
17-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, AT THE THIRD FIRING OF A SLEEVE GASTRECTOMY AT THE INCISURA ANGULARIS OF THE STOMACH, BOTH TIMES THE STAPLE RELOAD COULD ONLY BE FIRED FRACTIONALLY. THERE WAS INCOMPLETE STAPLE LINE AT THE PROXIMAL END. FAULTY STAPLE LINE HAD TO BE DE-BULKED AND THE DISTANCE FROM THE INCISURA HAD TO BE ESTABLISHED BEFORE A NEW STAPLE LINE/FIRING WAS DEPLOYED. THE STAPLE FAILURE LEAD TO A NEW PATH CLOSER TO THE BOUGIE HAVING TO BE FOLLOWED. SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE NEW STAPLE LINE ALSO HAD TO BE OVERSEWED TO GIVE IT STABILITY AND INTEGRITY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS NOT CUTTING ALL THE WAY, IT WOULD STOP MID WAY AND BECOME INTERMITTENT. THE HANDLE WAS TROUBLESHOOT AND NOTED THAT THE SOFTWARE VERSION 29.3 WAS INSTALLED. AS PER DIRECTION OF THE TECHNICIAN SOFTWARE 29.4 WOULD RECTIFY THIS AND WAS RECOMMENDED TO UPDATE. IT WAS ALSO NOTED THAT THERE WAS A MISHANDLING OF THE ADAPTER WHEN IT CAME TO DETACHING THE LOAD. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS NOT CUTTING ALL THE WAY, IT WOULD STOP MID WAY AND BECOME INTERMITTENT. THE HANDLE WAS TROUBLESHOOT AND NOTED THAT THE SOFTWARE VERSION 29.3 WAS INSTALLED. AS PER DIRECTION OF THE TECHNICIAN SOFTWARE 29.4 WOULD RECTIFY THIS AND WAS RECOMMENDED TO UPDATE. IT WAS ALSO NOTED THAT THERE WAS A MISHANDLING OF THE ADAPTER WHEN IT CAME TO DETACHING THE LOAD. THERE WAS NO PATIENT INJURY.
17-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, AT THE THIRD FIRING OF A SLEEVE GASTRECTOMY AT THE INCISURA ANGULARIS OF THE STOMACH, BOTH TIMES THE STAPLE RELOAD COULD ONLY BE FIRED FRACTIONALLY. THERE WAS INCOMPLETE STAPLE LINE AT THE PROXIMAL END. FAULTY STAPLE LINE HAD TO BE DE-BULKED AND THE DISTANCE FROM THE INCISURA HAD TO BE ESTABLISHED BEFORE A NEW STAPLE LINE/FIRING WAS DEPLOYED. THE STAPLE FAILURE LEAD TO A NEW PATH CLOSER TO THE BOUGIE HAVING TO BE FOLLOWED. SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THE NEW STAPLE LINE ALSO HAD TO BE OVERSEWED TO GIVE IT STABILITY AND INTEGRITY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS NOT CUTTING ALL THE WAY, IT WOULD STOP MID WAY AND BECOME INTERMITTENT. THE HANDLE WAS TROUBLESHOOT AND NOTED THAT THE SOFTWARE VERSION 29.3 WAS INSTALLED. AS PER DIRECTION OF THE TECHNICIAN SOFTWARE 29.4 WOULD RECTIFY THIS AND WAS RECOMMENDED TO UPDATE. IT WAS ALSO NOTED THAT THERE WAS A MISHANDLING OF THE ADAPTER WHEN IT CAME TO DETACHING THE LOAD. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, IT SEEMED LIKE A FIRE COULD NOT BE PERFORMED. NEW PRODUCT WAS TAKEN OUT. NO PATIENT INJURY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, IT SEEMED LIKE A FIRE COULD NOT BE PERFORMED. NEW PRODUCT WAS TAKEN OUT. NO PATIENT INJURY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, IT SEEMED LIKE A FIRE COULD NOT BE PERFORMED. NEW PRODUCT WAS TAKEN OUT. NO PATIENT INJURY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, IT SEEMED LIKE A FIRE COULD NOT BE PERFORMED. NEW PRODUCT WAS TAKEN OUT. NO PATIENT INJURY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING ROBOT-ASSISTED LAPAROSCOPIC TOTAL ESOPHAGECTOMY, WHILE RESECTING THE DISTAL END OF THE ESOPHAGUS USING THE POWERED STAPLER (FIRST FIRING), WHEN AN ATTEMPT WAS MADE TO RELEASE THE DEVICE FROM THE TISSUE, THE KNIFE RETRACTED FOR ABOUT 1.5CM, AND THE MOVEMENT STOPPED. A TISSUE WAS CLAMPED AND THE DEVICE COULD NOT BE REMOVED. BOTH ENDS OF THE TISSUE WERE RESECTED USING AN ENERGY DEVICE. THE STAPLER WAS REMOVED FROM THE PATIENT'S CAVITY WITH A TISSUE PINCHED AT THE TIP. THE NURSE HAD DIFFICULTY UNLOADING THE DEVICE, AND USED MANUAL ADAPTER TO RETRACT THE KNIFE THEN REMOVED THE CARTRIDGE FROM THE ADAPTER. THE DISTAL SIDE OF THE ESOPHAGUS WAS CLOSED USING A SUTURE WITH THE ASSISTANCE OF A ROBOT FOR THE SUBSEQUENT OPERATION.
17-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | SIGNIA | Human-Device Interface Problem | A22 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING ROBOT-ASSISTED LAPAROSCOPIC TOTAL ESOPHAGECTOMY, WHILE RESECTING THE DISTAL END OF THE ESOPHAGUS USING THE POWERED STAPLER (FIRST FIRING), WHEN AN ATTEMPT WAS MADE TO RELEASE THE DEVICE FROM THE TISSUE, THE KNIFE RETRACTED FOR ABOUT 1.5CM, AND THE MOVEMENT STOPPED. A TISSUE WAS CLAMPED AND THE DEVICE COULD NOT BE REMOVED. BOTH ENDS OF THE TISSUE WERE RESECTED USING AN ENERGY DEVICE. THE STAPLER WAS REMOVED FROM THE PATIENT'S CAVITY WITH A TISSUE PINCHED AT THE TIP. THE NURSE HAD DIFFICULTY UNLOADING THE DEVICE, AND USED MANUAL ADAPTER TO RETRACT THE KNIFE THEN REMOVED THE CARTRIDGE FROM THE ADAPTER. THE DISTAL SIDE OF THE ESOPHAGUS WAS CLOSED USING A SUTURE WITH THE ASSISTANCE OF A ROBOT FOR THE SUBSEQUENT OPERATION.
13-09-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATECTOMY, WHEN RESECTING THE PANCREAS, WHEN THE SPECIALIST TRIED TO ARTICULATE THE RELOAD, A SOUND WAS HEARD AND DID NOT NOT ALLOW MAKE THE FIRE. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-09-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | SIGNIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATECTOMY, WHEN RESECTING THE PANCREAS, WHEN THE SPECIALIST TRIED TO ARTICULATE THE RELOAD, A SOUND WAS HEARD AND DID NOT NOT ALLOW MAKE THE FIRE. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-09-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATECTOMY, WHEN RESECTING THE PANCREAS, WHEN THE SPECIALIST TRIED TO ARTICULATE THE RELOAD, A SOUND WAS HEARD AND DID NOT ALLOW MAKE THE FIRE. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING ROBOT-ASSISTED LAPAROSCOPIC TOTAL ESOPHAGECTOMY, WHILE RESECTING THE DISTAL END OF THE ESOPHAGUS USING THE POWERED STAPLER (FIRST FIRING), WHEN AN ATTEMPT WAS MADE TO RELEASE THE DEVICE FROM THE TISSUE, THE KNIFE RETRACTED FOR ABOUT 1.5CM, AND THE MOVEMENT STOPPED. A TISSUE WAS CLAMPED AND THE DEVICE COULD NOT BE REMOVED. BOTH ENDS OF THE TISSUE WERE RESECTED USING AN ENERGY DEVICE. THE STAPLER WAS REMOVED FROM THE PATIENT'S CAVITY WITH A TISSUE PINCHED AT THE TIP. THE NURSE HAD DIFFICULTY UNLOADING THE DEVICE, AND USED MANUAL ADAPTER TO RETRACT THE KNIFE THEN REMOVED THE CARTRIDGE FROM THE ADAPTER. THE DISTAL SIDE OF THE ESOPHAGUS WAS CLOSED USING A SUTURE WITH THE ASSISTANCE OF A ROBOT FOR THE SUBSEQUENT OPERATION.
17-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | SIGNIA | Human-Device Interface Problem | A22 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING ROBOT-ASSISTED LAPAROSCOPIC TOTAL ESOPHAGECTOMY, WHILE RESECTING THE DISTAL END OF THE ESOPHAGUS USING THE POWERED STAPLER (FIRST FIRING), WHEN AN ATTEMPT WAS MADE TO RELEASE THE DEVICE FROM THE TISSUE, THE KNIFE RETRACTED FOR ABOUT 1.5CM, AND THE MOVEMENT STOPPED. A TISSUE WAS CLAMPED AND THE DEVICE COULD NOT BE REMOVED. BOTH ENDS OF THE TISSUE WERE RESECTED USING AN ENERGY DEVICE. THE STAPLER WAS REMOVED FROM THE PATIENT'S CAVITY WITH A TISSUE PINCHED AT THE TIP. THE NURSE HAD DIFFICULTY UNLOADING THE DEVICE, AND USED MANUAL ADAPTER TO RETRACT THE KNIFE THEN REMOVED THE CARTRIDGE FROM THE ADAPTER. THE DISTAL SIDE OF THE ESOPHAGUS WAS CLOSED USING A SUTURE WITH THE ASSISTANCE OF A ROBOT FOR THE SUBSEQUENT OPERATION.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT
17-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING ROBOT-ASSISTED LAPAROSCOPIC TOTAL ESOPHAGECTOMY, WHILE RESECTING THE DISTAL END OF THE ESOPHAGUS USING THE POWERED STAPLER (FIRST FIRING), WHEN AN ATTEMPT WAS MADE TO RELEASE THE DEVICE FROM THE TISSUE, THE KNIFE RETRACTED FOR ABOUT 1.5CM, AND THE MOVEMENT STOPPED. A TISSUE WAS CLAMPED AND THE DEVICE COULD NOT BE REMOVED. BOTH ENDS OF THE TISSUE WERE RESECTED USING AN ENERGY DEVICE. THE STAPLER WAS REMOVED FROM THE PATIENT'S CAVITY WITH A TISSUE PINCHED AT THE TIP. THE NURSE HAD DIFFICULTY UNLOADING THE DEVICE, AND USED MANUAL ADAPTER TO RETRACT THE KNIFE THEN REMOVED THE CARTRIDGE FROM THE ADAPTER. THE DISTAL SIDE OF THE ESOPHAGUS WAS CLOSED USING A SUTURE WITH THE ASSISTANCE OF A ROBOT FOR THE SUBSEQUENT OPERATION.
17-12-2021 | Injury | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Human-Device Interface Problem | A22 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING ROBOT-ASSISTED LAPAROSCOPIC TOTAL ESOPHAGECTOMY, WHILE RESECTING THE DISTAL END OF THE ESOPHAGUS USING THE POWERED STAPLER (FIRST FIRING), WHEN AN ATTEMPT WAS MADE TO RELEASE THE DEVICE FROM THE TISSUE, THE KNIFE RETRACTED FOR ABOUT 1.5CM, AND THE MOVEMENT STOPPED. A TISSUE WAS CLAMPED AND THE DEVICE COULD NOT BE REMOVED. BOTH ENDS OF THE TISSUE WERE RESECTED USING AN ENERGY DEVICE. THE STAPLER WAS REMOVED FROM THE PATIENT'S CAVITY WITH A TISSUE PINCHED AT THE TIP. THE NURSE HAD DIFFICULTY UNLOADING THE DEVICE, AND USED MANUAL ADAPTER TO RETRACT THE KNIFE THEN REMOVED THE CARTRIDGE FROM THE ADAPTER. THE DISTAL SIDE OF THE ESOPHAGUS WAS CLOSED USING A SUTURE WITH THE ASSISTANCE OF A ROBOT FOR THE SUBSEQUENT OPERATION.
11-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, WHILE TRANSECTING THE APPENDIX, 2 HANDLES WERE UNABLE TO FIRE WITH DIFFERENT RELOADS. THE SURGEON USED A POWERED STAPLER WITH A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE USING THE DEVICE TO RESECT THE COMMON BILE DUCT LAPAROSCOPIC CHOLECYSTECTOMY, BECAUSE THE RELOAD WAS 45 MM, THEY TRIED TO SQUEEZE THE HANDLE 3 TIMES, BUT AFTER 2 TIMES, THE HANDLE COULD NOT BE SQUEEZED. THE SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC APPENDECTOMY, WHILE TRANSECTING THE APPENDIX, 2 HANDLES WERE UNABLE TO FIRE WITH DIFFERENT RELOADS. THE SURGEON USED A POWERED STAPLER WITH A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE USING THE DEVICE TO RESECT THE COMMON BILE DUCT LAPAROSCOPIC CHOLECYSTECTOMY, BECAUSE THE RELOAD WAS 45 MM, THEY TRIED TO SQUEEZE THE HANDLE 3 TIMES, BUT AFTER 2 TIMES, THE HANDLE COULD NOT BE SQUEEZED. THE SURGEON REPLACED THE DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-11-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY, THE STAPLER'S GEAR GOT BROKEN OR WAS NOT WORKING. IT WAS NOT POSSIBLE TO FIRE USING THE RELOAD. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-11-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY, THE STAPLER'S GEAR GOT BROKEN OR WAS NOT WORKING. IT WAS NOT POSSIBLE TO FIRE USING THE RELOAD. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-11-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY, THE STAPLER'S GEAR GOT BROKEN OR WAS NOT WORKING. IT WAS NOT POSSIBLE TO FIRE USING THE RELOAD. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2021 | Injury | us surgical puerto rico | 03-01-2022 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY, ALTHOUGH TISSUE WAS INSIDE THE 60 MM CUTLINE, ONLY HALF OF THE STAPLES WERE PLACED. THE JAWS LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THE EVENT LED TO ADDITIONAL COLON RESECTION.
17-12-2021 | Injury | us surgical puerto rico | 03-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLECTOMY, ALTHOUGH TISSUE WAS INSIDE THE 60 MM CUTLINE, ONLY HALF OF THE STAPLES WERE PLACED. THE JAWS LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THE EVENT LED TO ADDITIONAL COLON RESECTION.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE STAPLER WAS UNABLE TO FIRE. THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE CANNOT BE SQUEEZED.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2021 | Malfunction | us surgical puerto rico | 03-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR USE TO A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2021 | Malfunction | covidien lp llc north haven | 03-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBECTOMY, WHILE OPENING THE INTERLOBULAR FISSURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. TO RESOLVE THE ISSUE, THE SURGEON OPENED A NEW RELOAD. THERE WAS NO PATIENT INJURY.
09-12-2021 | Malfunction | intuitive surgical | 03-01-2022 | SUREFORM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE CUSTOMER ENCOUNTERED INCOMPLETE STAPLE LINE WITH THE SUREFORM STAPLER 60 INSTRUMENT. THE CUSTOMER USED A BACKUP INSTRUMENT. THE PROCEDURE WAS COMPLETED WITH NO INJURY TO THE PATIENT. INTUITIVE SURGICAL (IS) MADE MULTIPLE ATTEMPTS TO CONTACT THE REPORTER TO OBTAIN ADDITIONAL INFORMATION ABOUT THE COMPLAINT; HOWEVER, ATTEMPTS WERE NOT SUCCESSFUL.
09-12-2021 | Injury | smith & nephew | 03-01-2022 | TENDON ANCHORS 8 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING PARTIAL THICKNESS ROTATOR CUFF TEAR REPAIR, THREE (3) ANCHORS FROM THE SAME TENDON ANCHORS SET SNAPPED IN HALF AT THE TOP OF THE ¿U¿ DURING INSERTION. THE PIECES WERE REMOVED FROM PATIENT USING HANDHELD GRASPERS. PROCEDURE WAS COMPLETED, AFTER A 10-15MIN DELAY, WITH A BACK-UP DEVICE. NO OTHER COMPLICATIONS WERE REPORTED.
09-12-2021 | Injury | smith & nephew | 03-01-2022 | TENDON ANCHORS 8 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING PARTIAL THICKNESS ROTATOR CUFF TEAR REPAIR, THREE (3) ANCHORS FROM THE SAME TENDON ANCHORS SET SNAPPED IN HALF AT THE TOP OF THE ¿U¿ DURING INSERTION. THE PIECES WERE REMOVED FROM PATIENT USING HANDHELD GRASPERS. PROCEDURE WAS COMPLETED, AFTER A 10-15MIN DELAY, WITH A BACK-UP DEVICE. NO OTHER COMPLICATIONS WERE REPORTED.
08-07-2021 | Injury | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2010 AND AUGUST 2017, A PROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY. A VESSEL SEALER OR OTHER MANUFACTURER DEVICE WERE USED FOR DISSECTION OF GASTROCOLIC LIGAMENT, GASTROEPIPLOIC VESSELS AND GASTRIC GREATER CURVATURE DISSECTION. GASTRIC SECTION WAS PERFORMED USING A 60-MM STAPLER WITH A TRI-STAPLE PURPLE CARTRIDGE OR OTHER COMPANY STAPLER. THERE WERE 1800 PATIENTS IN THE STUDY AND COMPLICATIONS INCLUDED: 27 PATIENTS WITH ABDOMINAL BLEEDING AND 35 PATIENTS WITH STAPLE-LINE LEAKAGE. RELAPAROSCOPY OR ENDOSCOPIC TREATMENT WAS REQUIRED. MEDIAN LENGTH OF HOSPITAL STAY WAS 4 DAYS. ARTICLE: PANAGIOTIS LAINAS, 2021, LEFT HYPOCHONDRIUM OR TRANSUMBILICAL SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY FOR THE TREATMENT OF SEVERE OBESITY: SURGICAL TECHNIQUE AND RESULTS OF A TERTIARY REFERRAL BARIATRIC CENTER, HTTPS://DOI.ORG/10.1007/S11695-021-05563-Y.
08-07-2021 | Injury | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2010 AND AUGUST 2017, A PROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY. A VESSEL SEALER OR OTHER MANUFACTURER DEVICE WERE USED FOR DISSECTION OF GASTROCOLIC LIGAMENT, GASTROEPIPLOIC VESSELS AND GASTRIC GREATER CURVATURE DISSECTION. GASTRIC SECTION WAS PERFORMED USING A 60-MM STAPLER WITH A TRI-STAPLE PURPLE CARTRIDGE OR OTHER COMPANY STAPLER. THERE WERE 1800 PATIENTS IN THE STUDY AND COMPLICATIONS INCLUDED: 27 PATIENTS WITH ABDOMINAL BLEEDING AND 35 PATIENTS WITH STAPLE-LINE LEAKAGE. RELAPAROSCOPY OR ENDOSCOPIC TREATMENT WAS REQUIRED. MEDIAN LENGTH OF HOSPITAL STAY WAS 4 DAYS. ARTICLE: PANAGIOTIS LAINAS, 2021, LEFT HYPOCHONDRIUM OR TRANSUMBILICAL SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY FOR THE TREATMENT OF SEVERE OBESITY: SURGICAL TECHNIQUE AND RESULTS OF A TERTIARY REFERRAL BARIATRIC CENTER, HTTPS://DOI.ORG/10.1007/S11695-021-05563-Y.
17-12-2021 | Injury | us surgical puerto rico | 04-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING SINGLE PORT THORACOSCOPIC LOBECTOMY, WHEN DETACHING LUNG PARENCHYMA, SOME STAPLES COULD NOT FORM "B" SHAPE. ALSO, STAPLE LINE WAS INCOMPLETE CENTRALLY. THE STAPLE LINE WAS SUTURED, AND THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
17-12-2021 | Malfunction | us surgical puerto rico | 04-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE STAPLER WAS UNABLE TO FIRE. THE GREEN BUTTON WAS NOT PRESSED AND THE HANDLE CANNOT BE SQUEEZED.
24-09-2021 | Death | us surgical puerto rico | 04-01-2022 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT IVOR LEWIS MINIMALLY INVASIVE ESOPHAGECTOMY USING THREE DIFFERENT STAPLING TECHNIQUES BETWEEN JANUARY 2012 AND DECEMBER 2019. THE PATIENTS WERE DIVIDED INTO THREE GROUPS: THE REVERSAL PENETRATING TECHNIQUE GROUP, THE TRANSORAL ORAL ANVIL TECHNIQUE GROUP, AND THE PURSE-STRING TECHNIQUE GROUP. A CIRCULAR WERE USED IN ALL THREE GROUPS. AN ORAL ANVIL DEVICE WAS USED ONLY IN THE ORAL ANVIL GROUP. THERE WERE 316 PATIENTS IN THE STUDY AND COMPLICATIONS INCLUDED 5 IN-HOSPITAL/30-DAY MORTALITY. THE CAUSE OF THE REPORTED DEATHS WAS NOT SPECIFIED.
24-09-2021 | Death | covidien lp llc north haven | 04-01-2022 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT IVOR LEWIS MINIMALLY INVASIVE ESOPHAGECTOMY USING THREE DIFFERENT STAPLING TECHNIQUES BETWEEN JANUARY 2012 AND DECEMBER 2019. THE PATIENTS WERE DIVIDED INTO THREE GROUPS: THE REVERSAL PENETRATING TECHNIQUE GROUP, THE TRANSORAL ORAL ANVIL TECHNIQUE GROUP, AND THE PURSE-STRING TECHNIQUE GROUP. AN ENDOSCOPIC LINEAR STAPLER WAS USED IN ALL THREE GROUPS. THERE WERE 316 PATIENTS IN THE STUDY AND COMPLICATIONS INCLUDED 5 IN-HOSPITAL/30-DAY MORTALITY. THE CAUSE OF THE REPORTED DEATHS WAS NOT SPECIFIED.
19-04-2021 | Malfunction | covidien lp llc north haven | 04-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, FIRST HANDLE DID NOT ACTIVATE AND WOULD NOT TURN ON. THE SECOND HANDLE WORKS AND ABLE TO TURN ON THRU A RESET. THERE WAS NO PATIENT INVOLVEMENT.
19-04-2021 | Malfunction | covidien lp llc north haven | 04-01-2022 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, FIRST HANDLE DID NOT ACTIVATE AND WOULD NOT TURN ON. THE SECOND HANDLE WORKS AND ABLE TO TURN ON THRU A RESET. THERE WAS NO PATIENT INVOLVEMENT.
06-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BLEEDING OCCURRED AND B-SHAPE DID NOT WORK WELL. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
06-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BLEEDING OCCURRED AND B-SHAPE DID NOT WORK WELL. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BLEEDING OCCURRED AND B-SHAPE DID NOT WORK WELL. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BLEEDING OCCURRED AND B-SHAPE DID NOT WORK WELL. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
02-11-2021 | Injury | covidien lp llc north haven | 04-01-2022 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Pneumothorax | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY COMPARED THE OUTCOMES OF SINGLE-PORT VERSUS MULTIPORT SURGICAL TECHNIQUE FOR VIDEO ASSISTED THORACOSCOPIC SURGERY IN PEDIATRIC PATIENTS (AGE 12-17) WITH PRIMARY SPONTANEOUS PNEUMOTHORAX BETWEEN AUGUST 2017 AND FEBRUARY 2021. THE STAPLER WITH 45-60 MM PURPLE TRI-STAPLE WAS USED TO RESECT THE REGION OF THE LUNG CONTAINING BLEBS. THERE WERE 7 PATIENTS IN THE STUDY AND ONE PATIENT EXPERIENCED THE COMPLICATION OF RECURRENT PNEUMOTHORAX REQUIRING REPLACEMENT OF CHEST TUBE.
04-01-2022 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
04-01-2022 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
04-01-2022 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
04-01-2022 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
08-12-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 04-01-2022 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A PROCEDURE ON (B)(6) 2021, A BARD/DAVOL OPTIFIX STRAIGHT FIXATION DEVICE WAS USED TO FIXATE AN UNKNOWN MESH. IT WAS REPORTED THAT BROKEN FASTENERS WERE DEPLOYED FROM THE DEVICE WHICH WERE REMOVED FROM THE PATIENT'S BODY. IT WAS ALSO REPORTED THAT THE SURGEON IS AN EXPERIENCED USER OF THE DEVICE. THERE WAS NO REPORTED PATIENT INJURY.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE.THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE.THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
15-12-2021 | Malfunction | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC BYPASS, THE DEVICE DID NOT FIRE AFTER THE GREEN BUTTON IS PRESSED AND THE HANDLE SQUEEZED. THE RELOAD WAS REPLACED AND THE PROCEDURE WENT ON. THERE WAS NO PATIENT INJURY.
18-11-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ILS 33MM, CURVED | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY THE CIRCULAR STAPLER MISFIRED. CASE WAS COMPLETED WITH ANOTHER LIKE DEVICE.
18-11-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ILS 33MM, CURVED | Failure to Eject | A050301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY THE CIRCULAR STAPLER MISFIRED. CASE WAS COMPLETED WITH ANOTHER LIKE DEVICE.
20-12-2021 | Injury | us surgical puerto rico | 04-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DISSECTING THE STOMACH TISSUE, THE STAPLES DID NOT FORM PROPERLY, SOME STAPLES WERE NOT FORMED. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THE SURGEON SUTURED THE STAPLE LINE TO FIX THE STAPLE LINE AND NEW PRODUCT WAS USED TO RESOLVE THE ISSUE.
20-12-2021 | Injury | us surgical puerto rico | 04-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DISSECTING THE STOMACH TISSUE, THE STAPLES DID NOT FORM PROPERLY, SOME STAPLES WERE NOT FORMED. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THE SURGEON SUTURED THE STAPLE LINE TO FIX THE STAPLE LINE AND NEW PRODUCT WAS USED TO RESOLVE THE ISSUE.
10-11-2021 | Malfunction | covidien lp llc north haven | 04-01-2022 | IDRIVE ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, BEFORE USING THE DEVICE ON THE STOMACH IN A CLOSE GASTROPLASTY PROCEDURE, THE COMPONENT OF THE DEVICE DISENGAGED. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SPLENECTOMY, THE DEVICE WAS FIRED ON A VESSEL WHICH LEAD TO SIGNIFICANT BLEEDING IN NUMEROUS LOCATIONS OF THE STAPLE LINE.
10-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SPLENECTOMY, THE DEVICE WAS FIRED ON A VESSEL WHICH LEAD TO SIGNIFICANT BLEEDING IN NUMEROUS LOCATIONS OF THE STAPLE LINE.
10-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SPLENECTOMY, THE DEVICE WAS FIRED ON A VESSEL WHICH LEAD TO SIGNIFICANT BLEEDING IN NUMEROUS LOCATIONS OF THE STAPLE LINE.
10-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SPLENECTOMY, THE DEVICE WAS FIRED ON A VESSEL WHICH LEAD TO SIGNIFICANT BLEEDING IN NUMEROUS LOCATIONS OF THE STAPLE LINE.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, THE STAPLES CANNOT STOP THE BLEEDING. THE ANAL WAS BLEEDING AT 8 O'CLOCK AND 11 O'CLOCK POSITION. TO RESOLVE THE ISSUE, SUTURING WAS PERFORMED. THE PROCEDURE WAS EXTENDED FOR 30 MINUTES OR MORE. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE.
20-12-2021 | Malfunction | us surgical puerto rico | 04-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHEN DETACHING THE BRONCHUS, THE STAPLER WAS ABLE TO FIRE, THERE WAS POOR STAPLE FORMATION THE RELOAD WAS REPLACED AND THE SAME HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHILE USING PVE35A+VASECR35, IT SUDDENLY STOPPED IN THE MIDDLE OF FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2021 | Malfunction | us surgical puerto rico | 04-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC (VAT) LOBECTOMY, WHILE TRANSECTING THE BRONCHUS, THE SURGEON USED THE 1ST HANDLE AND FAILED TO FIRE, THEN CHANGE TO ANOTHER HANDLE WITH THE SAME RELOAD BUT FAILED TO FIRE TOO. THE DOCTOR OPENED ANOTHER RELOAD WITH THE USE OF THE 2ND HANDLE WHICH ALSO DID NOT FIRE. IN THE END, THE USER OPENED A 3RD HANDLE WITH THE 2ND RELOAD TO RESOLVE THE ISSUE. IT WAS NOTED THE OPERATOR WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 04-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC (VAT) LOBECTOMY, WHILE TRANSECTING THE BRONCHUS, THE SURGEON USED THE 1ST HANDLE AND FAILED TO FIRE, THEN CHANGE TO ANOTHER HANDLE WITH THE SAME RELOAD BUT FAILED TO FIRE TOO. THE DOCTOR OPENED ANOTHER RELOAD WITH THE USE OF THE 2ND HANDLE WHICH ALSO DID NOT FIRE. IN THE END, THE USER OPENED A 3RD HANDLE WITH THE 2ND RELOAD TO RESOLVE THE ISSUE. IT WAS NOTED THE OPERATOR WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
30-11-2021 | Injury | intuitive surgical | 04-01-2022 | ENDOWRIST | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS INITIALLY REPORTED THAT AFTER A DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA REPAIR AND SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT EXPERIENCED POST-OPERATIVE COMPLICATIONS AND RETURNED TO THE EMERGENCY ROOM FOR A PROXIMAL STAPLE LINE LEAK ABOUT A WEEK POST-OPERATIVELY (DATE UNSPECIFIED). THIS DA VINCI ASSISTED SURGICAL PROCEDURE WAS THE FIRST OF FOUR ROBOTIC PROCEDURES PERFORMED BY THE SURGEON ON (B)(6)2021. POST-OPERATIVELY, A LEAK TEST WAS PERFORMED AND THE PATIENT UNDERWENT CT IMAGING TO CONFIRM THE LOCATION OF THE LEAK. THE PATIENT UNDERWENT A LAPAROSCOPIC RE-OPERATION, HAD ADDITIONAL SUTURES PLACED TO OVERSEW THE DEFECT, AND A DRAIN WAS PLACED.
20-12-2021 | Malfunction | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC (VAT) LOBECTOMY, WHILE TRANSECTING THE BRONCHUS, THE SURGEON USED THE 1ST HANDLE AND FAILED TO FIRE, THEN CHANGE TO ANOTHER HANDLE WITH THE SAME RELOAD BUT FAILED TO FIRE TOO. THE DOCTOR OPENED ANOTHER RELOAD WITH THE USE OF THE 2ND HANDLE WHICH ALSO DID NOT FIRE. IN THE END, THE USER OPENED A 3RD HANDLE WITH THE 2ND RELOAD TO RESOLVE THE ISSUE. IT WAS NOTED THE OPERATOR WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
30-11-2021 | Injury | intuitive surgical | 04-01-2022 | ENDOWRIST | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS INITIALLY REPORTED THAT AFTER A DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA REPAIR AND SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT EXPERIENCED POST-OPERATIVE COMPLICATIONS AND RETURNED TO THE EMERGENCY ROOM FOR A PROXIMAL STAPLE LINE LEAK ABOUT A WEEK POST-OPERATIVELY (DATE UNSPECIFIED). THIS EVENT IS ASSOCIATED WITH THE SECOND OF FOUR ROBOTIC PROCEDURES PERFORMED BY THE SURGEON ON (B)(6) 2021. POST-OPERATIVELY, A LEAK TEST WAS PERFORMED AND THE PATIENT UNDERWENT CT IMAGING TO CONFIRM THE LOCATION OF THE LEAK. THE PATIENT UNDERWENT A LAPAROSCOPIC RE-OPERATION, HAD ADDITIONAL SUTURES PLACED TO OVERSEW THE STAPLE LINE LEAK, AND A DRAIN WAS PLACED.
30-11-2021 | Injury | intuitive surgical | 04-01-2022 | ENDOWRIST | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS INITIALLY REPORTED THAT AFTER A DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA REPAIR AND SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT EXPERIENCED POST-OPERATIVE COMPLICATIONS AND RETURNED TO THE EMERGENCY ROOM FOR A PROXIMAL STAPLE LINE LEAK ABOUT A WEEK POST-OPERATIVELY (DATE UNSPECIFIED). THIS EVENT IS ASSOCIATED WITH THE SECOND OF FOUR ROBOTIC PROCEDURES PERFORMED BY THE SURGEON ON (B)(6) 2021. POST-OPERATIVELY, A LEAK TEST WAS PERFORMED AND THE PATIENT UNDERWENT CT IMAGING TO CONFIRM THE LOCATION OF THE LEAK. THE PATIENT UNDERWENT A LAPAROSCOPIC RE-OPERATION, HAD ADDITIONAL SUTURES PLACED TO OVERSEW THE STAPLE LINE LEAK, AND A DRAIN WAS PLACED.
20-12-2021 | Injury | us surgical puerto rico | 04-01-2022 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, WHEN THE RELOAD WAS USED WITH THE MANUAL HANDLE IN THE THORACOSCOPY ASSISTED LEFT PULMONARY SEGMENTAL RESECTION, THE BLACK RETURN KNOB OF THE FIRED STAPLER COULD NOT BE RETURNED, NO MATTER WHAT THEY DID, THE CARTRIDGE REMAINED CLAMPED ON THE TISSUE AND COULD NOT BE OPENED. INITIALLY, THE SALES REP ASKED THEM TO PULL THE BLACK RETURN KNOB BACK ON THE TELEPHONE, BUT IT DID NOT MOVE AT ALL. IT DID NOT SEEM TO BE A "TEARFUL PARTING", SO HE/SHE ASKED THEM TO MANUALLY PULL BACK WITH THE CLAMP COVER BUTTON, BUT THIS DID NOT MOVE AT ALL EITHER. PROCEDURE WAS CONVERTED TO LEFT SUPERIOR RESECTION FROM THE SEGMENTAL RESECTION AND THE SPECIMEN WAS IN REMOVED STATE. THE SPECIMEN HAD BEEN ADHERED TO THE CARTRIDGE THAT THE JAWS DID NOT OPEN, ACCORDING TO THE SURGEON, WHEN THE BLACK RETURN KNOB WAS ATTEMPTED TO BE FORCEFULLY PULLED BACK, THE HANDLE GOT BROKE, AND THE CARTRIDGE AND HANDLE CAME OFF. WITH THE SURGEON'S PERMISSION, THE TISSUE WAS CUT OFF AND THE CARTRIDGE WAS COLLECTED.
20-12-2021 | Injury | us surgical puerto rico | 04-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, WHEN THE RELOAD WAS USED WITH THE MANUAL HANDLE IN THE THORACOSCOPY ASSISTED LEFT PULMONARY SEGMENTAL RESECTION, THE BLACK RETURN KNOB OF THE FIRED STAPLER COULD NOT BE RETURNED, NO MATTER WHAT THEY DID, THE CARTRIDGE REMAINED CLAMPED ON THE TISSUE AND COULD NOT BE OPENED. INITIALLY, THE SALES REP ASKED THEM TO PULL THE BLACK RETURN KNOB BACK ON THE TELEPHONE, BUT IT DID NOT MOVE AT ALL. IT DID NOT SEEM TO BE A "TEARFUL PARTING", SO HE/SHE ASKED THEM TO MANUALLY PULL BACK WITH THE CLAMP COVER BUTTON, BUT THIS DID NOT MOVE AT ALL EITHER. PROCEDURE WAS CONVERTED TO LEFT SUPERIOR RESECTION FROM THE SEGMENTAL RESECTION AND THE SPECIMEN WAS IN REMOVED STATE. THE SPECIMEN HAD BEEN ADHERED TO THE CARTRIDGE THAT THE JAWS DID NOT OPEN, ACCORDING TO THE SURGEON, WHEN THE BLACK RETURN KNOB WAS ATTEMPTED TO BE FORCEFULLY PULLED BACK, THE HANDLE GOT BROKE, AND THE CARTRIDGE AND HANDLE CAME OFF. WITH THE SURGEON'S PERMISSION, THE TISSUE WAS CUT OFF AND THE CARTRIDGE WAS COLLECTED.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, WHEN THE RELOAD WAS USED WITH THE MANUAL HANDLE IN THE THORACOSCOPY ASSISTED LEFT PULMONARY SEGMENTAL RESECTION, THE BLACK RETURN KNOB OF THE FIRED STAPLER COULD NOT BE RETURNED, NO MATTER WHAT THEY DID, THE CARTRIDGE REMAINED CLAMPED ON THE TISSUE AND COULD NOT BE OPENED. INITIALLY, THE SALES REP ASKED THEM TO PULL THE BLACK RETURN KNOB BACK ON THE TELEPHONE, BUT IT DID NOT MOVE AT ALL. IT DID NOT SEEM TO BE A "TEARFUL PARTING", SO HE/SHE ASKED THE THEM TO MANUALLY PULL BACK WITH THE CLAMP COVER BUTTON, BUT THIS DID NOT MOVE AT ALL EITHER. PROCEDURE WAS CONVERTED TO LEFT SUPERIOR RESECTION FROM THE SEGMENTAL RESECTION AND THE SPECIMEN WAS IN REMOVED STATE. THE SPECIMEN HAD BEEN ADHERED TO THE CARTRIDGE THAT THE JAWS DID NOT OPEN, ACCORDING TO THE SURGEON, WHEN THE BLACK RETURN KNOB WAS ATTEMPTED TO BE FORCEFULLY PULLED BACK, THE HANDLE GOT BROKE, AND THE CARTRIDGE AND HANDLE CAME OFF. WITH THE SURGEON'S PERMISSION, THE TISSUE WAS CUT OFF AND THE CARTRIDGE WAS COLLECTED.
20-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, WHEN THE RELOAD WAS USED WITH THE MANUAL HANDLE IN THE THORACOSCOPY ASSISTED LEFT PULMONARY SEGMENTAL RESECTION, THE BLACK RETURN KNOB OF THE FIRED STAPLER COULD NOT BE RETURNED, NO MATTER WHAT THEY DID, THE CARTRIDGE REMAINED CLAMPED ON THE TISSUE AND COULD NOT BE OPENED. INITIALLY, THE SALES REP ASKED THEM TO PULL THE BLACK RETURN KNOB BACK ON THE TELEPHONE, BUT IT DID NOT MOVE AT ALL. IT DID NOT SEEM TO BE A "TEARFUL PARTING", SO HE/SHE ASKED THE THEM TO MANUALLY PULL BACK WITH THE CLAMP COVER BUTTON, BUT THIS DID NOT MOVE AT ALL EITHER. PROCEDURE WAS CONVERTED TO LEFT SUPERIOR RESECTION FROM THE SEGMENTAL RESECTION AND THE SPECIMEN WAS IN REMOVED STATE. THE SPECIMEN HAD BEEN ADHERED TO THE CARTRIDGE THAT THE JAWS DID NOT OPEN, ACCORDING TO THE SURGEON, WHEN THE BLACK RETURN KNOB WAS ATTEMPTED TO BE FORCEFULLY PULLED BACK, THE HANDLE GOT BROKE, AND THE CARTRIDGE AND HANDLE CAME OFF. WITH THE SURGEON'S PERMISSION, THE TISSUE WAS CUT OFF AND THE CARTRIDGE WAS COLLECTED.
13-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING A LOBECTOMY THE DOCTOR WAS USING THE PVE35A TO TRANSECT THE PULMONARY ARTERY DURING A LEFT UPPER LOBECTOMY. AFTER COMPLETING THE TRANSECTION THE DEVICE DID NOT FULLY RETRACTED THE BLADE CAUSING A LOCK OUT. DOCTOR WAS UNABLE TO GET THE DEVICE OPEN WITHOUT USING THE MANUAL OVERRIDE AND REPLACING THE DEVICE. NO IMPACT TO PATIENT AND NO ISSUES WITH STAPLE LINE FORMATION OR INTEGRITY. THE SURGERY WAS NOT DELAYED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. NO FRAGMENTS WERE GENERATED.
13-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING A LOBECTOMY THE DOCTOR WAS USING THE PVE35A TO TRANSECT THE PULMONARY ARTERY DURING A LEFT UPPER LOBECTOMY. AFTER COMPLETING THE TRANSECTION THE DEVICE DID NOT FULLY RETRACTED THE BLADE CAUSING A LOCK OUT. DOCTOR WAS UNABLE TO GET THE DEVICE OPEN WITHOUT USING THE MANUAL OVERRIDE AND REPLACING THE DEVICE. NO IMPACT TO PATIENT AND NO ISSUES WITH STAPLE LINE FORMATION OR INTEGRITY. THE SURGERY WAS NOT DELAYED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. NO FRAGMENTS WERE GENERATED.
04-12-2021 | Malfunction | us surgical puerto rico | 04-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY PARTIAL RESECTION OF THE LUNGS, WHILE RESECTING THE LUNG PARENCHYMA, THE TISSUE COULD NOT BE CLAMPED AND THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2021 | Malfunction | us surgical puerto rico | 04-01-2022 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY PARTIAL RESECTION OF THE LUNGS, WHILE RESECTING THE LUNG PARENCHYMA, THE TISSUE COULD NOT BE CLAMPED AND THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2021 | Malfunction | covidien lp llc north haven | 04-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY PARTIAL RESECTION OF THE LUNGS, WHILE RESECTING THE LUNG PARENCHYMA, THE TISSUE COULD NOT BE CLAMPED AND THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2021 | Malfunction | covidien lp llc north haven | 04-01-2022 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY PARTIAL RESECTION OF THE LUNGS, WHILE RESECTING THE LUNG PARENCHYMA, THE TISSUE COULD NOT BE CLAMPED AND THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY, THE DEVICE CLAMPED THE TARGET TISSUE, BUT THE JAWS DID NOT CLOSE WELL AT THE 1ST FIRING. AN UNEXPECTED SOUND WAS HEARD, AND IT WAS FOUND THAT A YELLOW PART WAS VISIBLE FROM THE GAP OF THE TRIGGER. THE DEVICE WAS USED ON THE DUODENUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
13-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY, THE DEVICE CLAMPED THE TARGET TISSUE, BUT THE JAWS DID NOT CLOSE WELL AT THE 1ST FIRING. AN UNEXPECTED SOUND WAS HEARD, AND IT WAS FOUND THAT A YELLOW PART WAS VISIBLE FROM THE GAP OF THE TRIGGER. THE DEVICE WAS USED ON THE DUODENUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
13-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY, THE DEVICE CLAMPED THE TARGET TISSUE, BUT THE JAWS DID NOT CLOSE WELL AT THE 1ST FIRING. AN UNEXPECTED SOUND WAS HEARD, AND IT WAS FOUND THAT A YELLOW PART WAS VISIBLE FROM THE GAP OF THE TRIGGER. THE DEVICE WAS USED ON THE DUODENUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DONOR NEPHRECTOMY PROCEDURE THAT AFTER THE FIRST FIRING THE DEVICE WOULD NOT OPEN. UNKNOWN AS TO HOW THE DEVICE WAS REMOVED. A SECOND STAPLER WAS OPENED AND ON THE FIRST FIRING JAMMED AND WOULD NOT OPEN. TECH WAS ABLE TO GET THE DEVICE OPEN AND WAS ABLE TO RELOAD AND COMPLETE THE PROCEDURE WITH THE SECOND DEVICE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | PROXIMATE*75MM LINEAR CUTTER | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THE PACKING WAS DAMAGED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED. PROCEDURE: GASTROSTOMY.
10-12-2021 | Injury | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, WHILE RESECTING THE COLON, THE DEVICE LOCKED ON TISSUE. THE SURGEON REMOVED BOTH THE ADAPTER AND THE RELOAD AND USED A NEW RELOAD TO RE-STAPLE THE TISSUE. UNEXPECTED TISSUE DAMAGE AND LOSS WERE OBSERVED DUE TO DEVICE ISSUE.
20-12-2021 | Malfunction | covidien lp llc north haven | 04-01-2022 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC INGUINAL HERNIA REPAIR, DURING MESH FIXATION, THE TACK DID NOT STICK EVEN WHEN THE TACKER WAS FIRED. THE TACK COULD NOT BE RELEASED FROM THE TACKER AFTER IT WAS FIRED. WHEN THE TACK WAS ATTACHED AND AN ATTEMPT WAS MADE TO REMOVE THE TACKER, IT WAS FELT LIKE IT WAS BEING PULLED OUT FORCIBLY, AND IT COULD NOT BE USED PROPERLY. A TACK ALSO FELL INTO THE PATIENT¿S CAVITY AND WAS RETRIEVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | covidien lp llc north haven | 04-01-2022 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC INGUINAL HERNIA REPAIR, DURING MESH FIXATION, THE TACK DID NOT STICK EVEN WHEN THE TACKER WAS FIRED. THE TACK COULD NOT BE RELEASED FROM THE TACKER AFTER IT WAS FIRED. WHEN THE TACK WAS ATTACHED AND AN ATTEMPT WAS MADE TO REMOVE THE TACKER, IT WAS FELT LIKE IT WAS BEING PULLED OUT FORCIBLY, AND IT COULD NOT BE USED PROPERLY. A TACK ALSO FELL INTO THE PATIENT¿S CAVITY AND WAS RETRIEVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 04-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SEGMENTECTOMY, DURING THE RESECTION OF THE LUNG SEGMENT, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WITH THE SAME RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-04-2020 | Malfunction | ethicon endo-surgery | 04-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-12-2021 | Injury | ethicon endo-surgery | 04-01-2022 | UNK_ECHELON CIRCULAR POWERED STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THERE WAS AN ANASTOMOTIC LEAK SEVERAL DAYS POST-OP. THE DEVICE WAS FIRED AS NORMAL WITH THE GREEN INDICATOR ILLUMINATED WHICH INDICATED THE FIRING SEQUENCE HAD COMPLETED SUCCESSFULLY. THE SURGEON DID 2 FULL TURNS TO REMOVE THE DEVICE. THE ANASTOMOTIC LEAK OCCURRED WHILE THE PATIENT WAS IN HOSPITAL A FEW DAYS LATER. THE PATIENT IS ON THE ROAD TO RECOVER BUT ANY LONG-TERM CONSEQUENCES ARE UNKNOWN.
08-12-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 04-01-2022 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNSPECIFIED PROCEDURE ON (B)(6) 2021, WHEN THE SURGEON TRIED TO FIXATE AN UNSPECIFIED MESH USING THE BARD/DAVOL OPTIFIX STRAIGHT FIXATION DEVICE, BROKEN FASTENERS WERE DEPLOYED. THE BROKEN FASTENERS WERE REMOVED FROM THE PATIENT'S BODY AND THERE WERE NO LOOSE FASTENERS LEFT IN THE PATIENT. THE SURGEON IS AN EXPERIENCED USER OF THE OPTIFIX DEVICE. THERE WAS NO REPORTED PATIENT INJURY.
01-01-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLE LINE DIDN'T FORM AT THE ANASTOMOSIS WITH THE BLUE RELOAD. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
13-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP COLECTOMY, THEY WENT TO FIRE THE DEVICE AND IT WOULD NOT FIRE. THEY DISENGAGED THE BATTERY AND PUT IT BACK IN. THEY ATTEMPTED A NEW BATTERY FROM A SECOND STAPLER. THE DEVICE WOULD NOT FIRE. THE POWER WAS ON THE LIGHTS WERE ON. THE ANVIL FROM THE FIRST DEVICE WAS STILL USED WITH THE SECOND DEVICE. THE DEVICE WOULD NOT FIRE. CASE COMPLETED WITH A SECOND STAPLER. NO PATIENT HARM.
04-12-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE 60MM ECHELON STAPLER FAILED TO STAPLE PROPERLY AND NEEDED TO BE REPLACED TO FINISH THE PROCEDURE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
27-10-2021 | Malfunction | us surgical puerto rico | 04-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER INTRAOPERATIVELY, WHEN THE INSTRUMENT WAS USED IN THE INTESTINE, IT COULD NOT BE CLAMP AND IT COULD NOT BE USED FOR FIRING. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-12-2021 | Injury | ethicon endo-surgery | 04-01-2022 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT ON (B)(6) 2021 DURING A VATS PROCEDURE THE THORACOTOMY WAS PERFORMED THE DEVICE HAD POOR STAPLE FORMATION. OPERATOR WAS USING THE ECHELON FLEX POWERED STAPLER WITH WHITE VASCULAR RELOADS - VASECR35. IT WAS FOUND OUT BY THE OPERATOR AND THE OR STAFF THAT THE VASECR35 RELOADS THAT DID NOT PERFORM PROPERLY WAS LOT: 459A85. THE OPERATING STAFF STATED THAT THEY HAVE USED THE INSTRUMENTS AS PER THE INSTRUCTIONS, WAITED 15 SECONDS BEFORE FIRING. THEY HAVE ALSO BEEN USING OUR POWERED STAPLERS AND RELOADS FOR OVER YEAR. DID THE PATIENT EXPERIENCE A POST-OP DEVICE MALFUNCTION? YES. IF YES, PLEASE DESCRIBE. STAPLE LINE WAS NOT FORMED PROPERLY. DID THE PATIENT REQUIRE REVISION SURGERY OR HARDWARE REMOVAL? YES.
07-12-2021 | Injury | ethicon endo-surgery | 04-01-2022 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT ON (B)(6) 2021 DURING A VATS PROCEDURE THE THORACOTOMY WAS PERFORMED THE DEVICE HAD POOR STAPLE FORMATION. OPERATOR WAS USING THE ECHELON FLEX POWERED STAPLER WITH WHITE VASCULAR RELOADS - VASECR35. IT WAS FOUND OUT BY THE OPERATOR AND THE OR STAFF THAT THE VASECR35 RELOADS THAT DID NOT PERFORM PROPERLY WAS LOT: 459A85. THE OPERATING STAFF STATED THAT THEY HAVE USED THE INSTRUMENTS AS PER THE INSTRUCTIONS, WAITED 15 SECONDS BEFORE FIRING. THEY HAVE ALSO BEEN USING OUR POWERED STAPLERS AND RELOADS FOR OVER YEAR. DID THE PATIENT EXPERIENCE A POST-OP DEVICE MALFUNCTION? YES. IF YES, PLEASE DESCRIBE. STAPLE LINE WAS NOT FORMED PROPERLY. DID THE PATIENT REQUIRE REVISION SURGERY OR HARDWARE REMOVAL? YES.
08-07-2021 | Death | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2010 AND AUGUST 2017, A PROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY. A VESSEL SEALER OR OTHER MANUFACTURER DEVICE WERE USED FOR DISSECTION OF GASTROCOLIC LIGAMENT, GASTROEPIPLOIC VESSELS AND GASTRIC GREATER CURVATURE DISSECTION. GASTRIC SECTION WAS PERFORMED USING A 60-MM STAPLER WITH A TRI-STAPLE PURPLE CARTRIDGE OR OTHER COMPANY STAPLER. THERE WERE 1800 PATIENTS IN THE STUDY AND COMPLICATIONS INCLUDED: 27 PATIENTS WITH ABDOMINAL BLEEDING AND 35 PATIENTS WITH STAPLE-LINE LEAKAGE. RELAPAROSCOPY OR ENDOSCOPIC TREATMENT WAS REQUIRED. MEDIAN LENGTH OF HOSPITAL STAY WAS 4 DAYS. POSTOPERATIVE MORTALITY WAS 0.05%. CAUSE OF DEATH WAS NOT SPECIFIED. ARTICLE: PANAGIOTIS LAINAS, 2021, LEFT HYPOCHONDRIUM OR TRANSUMBILICAL SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY FOR THE TREATMENT OF SEVERE OBESITY: SURGICAL TECHNIQUE AND RESULTS OF A TERTIARY REFERRAL BARIATRIC CENTER, HTTPS://DOI.ORG/10.1007/S11695-021-05563-Y.
08-07-2021 | Death | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2010 AND AUGUST 2017, A PROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY. A VESSEL SEALER OR OTHER MANUFACTURER DEVICE WERE USED FOR DISSECTION OF GASTROCOLIC LIGAMENT, GASTROEPIPLOIC VESSELS AND GASTRIC GREATER CURVATURE DISSECTION. GASTRIC SECTION WAS PERFORMED USING A 60-MM STAPLER WITH A TRI-STAPLE PURPLE CARTRIDGE OR OTHER COMPANY STAPLER. THERE WERE 1800 PATIENTS IN THE STUDY AND COMPLICATIONS INCLUDED: 27 PATIENTS WITH ABDOMINAL BLEEDING AND 35 PATIENTS WITH STAPLE-LINE LEAKAGE. RELAPAROSCOPY OR ENDOSCOPIC TREATMENT WAS REQUIRED. MEDIAN LENGTH OF HOSPITAL STAY WAS 4 DAYS. POSTOPERATIVE MORTALITY WAS 0.05%. CAUSE OF DEATH WAS NOT SPECIFIED. ARTICLE: PANAGIOTIS LAINAS, 2021, LEFT HYPOCHONDRIUM OR TRANSUMBILICAL SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY FOR THE TREATMENT OF SEVERE OBESITY: SURGICAL TECHNIQUE AND RESULTS OF A TERTIARY REFERRAL BARIATRIC CENTER, HTTPS://DOI.ORG/10.1007/S11695-021-05563-Y.
08-07-2021 | Death | covidien lp llc north haven | 04-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2010 AND AUGUST 2017, A PROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY. A VESSEL SEALER OR OTHER MANUFACTURER DEVICE WERE USED FOR DISSECTION OF GASTROCOLIC LIGAMENT, GASTROEPIPLOIC VESSELS AND GASTRIC GREATER CURVATURE DISSECTION. GASTRIC SECTION WAS PERFORMED USING A 60-MM STAPLER WITH A TRI-STAPLE PURPLE CARTRIDGE OR OTHER COMPANY STAPLER. THERE WERE 1800 PATIENTS IN THE STUDY AND COMPLICATIONS INCLUDED: 27 PATIENTS WITH ABDOMINAL BLEEDING AND 35 PATIENTS WITH STAPLE-LINE LEAKAGE. RELAPAROSCOPY OR ENDOSCOPIC TREATMENT WAS REQUIRED. MEDIAN LENGTH OF HOSPITAL STAY WAS 4 DAYS. POSTOPERATIVE MORTALITY WAS 0.05%. CAUSE OF DEATH WAS NOT SPECIFIED. ARTICLE: PANAGIOTIS LAINAS, 2021, LEFT HYPOCHONDRIUM OR TRANSUMBILICAL SINGLE-INCISION LAPAROSCOPIC SLEEVE GASTRECTOMY FOR THE TREATMENT OF SEVERE OBESITY: SURGICAL TECHNIQUE AND RESULTS OF A TERTIARY REFERRAL BARIATRIC CENTER, HTTPS://DOI.ORG/10.1007/S11695-021-05563-Y.
01-01-2021 | Injury | ethicon endo-surgery | 04-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE SURGEON REPORTED A POST OP LEAK TWO WEEKS AFTER THE BARIATRIC SLEEVE, THERE WAS A SMALL LEAK ON THE PROXIMAL SIDE OF THE STOMACH.
01-12-2021 | Injury | ethicon endo-surgery | 04-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A ROUX EN Y PROCEDURE AND SUBSEQUENTLY DEVELOPED A HEMATOMA. THE PATIENT WAS READMITTED AND THE HEMATOMA REPAIRED. THE HEMATOMA OCCURRED ON THE REMNANT SIDE.
01-11-2021 | Injury | ethicon endo-surgery | 04-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST OP OF A BARIATRIC SLEEVE, THERE WAS A SMALL LEAK ON THE PROXIMAL SIDE OF THE STOMACH. THIS PATIENT IS IN ICU AND SEEMS TO BE DOING A LITTLE BETTER NOW.
01-10-2021 | Malfunction | ethicon endo-surgery | 04-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT ACCORDING TO SURGEON, THE LINEAR CUTTER DID NOT FIRE FROM THE VERY BEGINNING. WILL INQUIRE ADDITIONAL INFO FROM SURGEON. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
07-12-2021 | Injury | intuitive surgical | 04-01-2022 | SUREFORM | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE SURGEON EXPERIENCED A STAPLING ISSUE. INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE SURGEON AND OBTAINED THE FOLLOWING INFORMATION: THE SURGEON HAD SUCCESSFULLY FIRED THE SUREFORM 60 STAPLER WITH A GREEN SUREFORM 60 RELOAD. THEN THE SURGEON ATTEMPTED TO FIRE ON STOMACH TISSUE WITH A BLUE SUREFORM 60 RELOAD. THE SURGEON CLAMPED ON TISSUE AND FIRED. THE SURGEON UNCLAMPED AND FOUND THAT ONLY ABOUT ONE-THIRD TO HALF OF THE STAPLE LINE WAS FORMED. THE STAPLER ALLEGEDLY DID NOT CUT THE TISSUE. THERE WERE STAPLES DEPLOYED THAT FELL ONTO THE TARGET TISSUE, AND THE TISSUE WAS PUSHED OUT. THE SURGEON ATTEMPTED A SECOND FIRE WITH A NEW BLUE RELOAD, MEDIAL TO THE SECOND STAPLE LINE, AND THAT STAPLE FIRE WAS SUCCESSFUL. THE SURGEON HAD TO TAKE IN A LITTLE MORE TISSUE; HOWEVER, THAT DID NOT AFFECT THE SURGERY. THE SURGEON OVERSEWED THE STAPLE LINE AS A PRECAUTION. AS PER ADVICE FROM AN ISI CLINICAL SALES REPRESENTATIVE (CSR), THE SURGEON SWITCHED TO A NEW SUREFORM 60 STAPLER INSTRUMENT FOR THE SUBSEQUENT FIRES. THE SURGEON REMOVED ALL THE DEPLOYED STAPLES THAT FELL ONTO THE TARGET TISSUE. SOME OF THEM WERE B-SHAPED AND SOME WERE OPEN STAPLES. THE SURGEON INDICATED THAT A 40MM FRENCH BOUGIE WAS USED. NO BUTTRESS MATERIAL WAS USED. THERE WAS NO TISSUE BUNCHING OR TENSION. ACCORDING TO THE SURGEON, THERE WERE NO SYSTEM-GENERATED ERRORS. THE SURGEON CONFIRMED THAT A MINOR DEFECT WAS NOTED ON THE EXTRACTED STOMACH TISSUE AND THE SURGEON CLAIMED IT WAS DUE TO THE ISSUE WITH THE SECOND STAPLE FIRE. THERE IS NO VIDEO RECORDING OF THE PROCEDURE AVAILABLE FOR REVIEW. THE SURGEON STATED THAT THEY INSPECTED THE STAPLER RELOAD CARTRIDGE. PER THE SURGEON, THERE WAS AN EXPOSED STAPLE AND HE BELIEVES THAT A STAPLE COULD HAVE BEEN IN FRONT OF THE BLADE. THE PATIENT IS REPORTED TO BE RECOVERING WELL. HOWEVER, AS A PRECAUTION, THE PATIENT'S HOSPITALIZATION WAS EXTENDED BY A DAY. THE SURGEON CLARIFIED THERE WAS NO OOZING FROM THE STAPLE LINE AS INITIALLY REPORTED.
14-12-2021 | Malfunction | celestica de monterrey s.a de c.v | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE SEVERING THE BRONCHUS, THE BLACK HANDLE COULD NOT BE GRASPED AND WAS UNABLE TO BE FIRED. A NEW RELOAD AND THE SAME HANDLE WERE USED TO RESOLVE THE CONCERN. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A THORASCOPIC HYBRID VIDEO-ASSISTED THORACIC SURGERY (VATS), AT THE TIME OF PULMONARY PARENCHYMA RESECTION, WHEN THE DEVICE WAS ATTEMPTED TO BE FIRED, THE TEETH DID NOT ADVANCE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A THORASCOPIC HYBRID VIDEO-ASSISTED THORACIC SURGERY (VATS), AT THE TIME OF PULMONARY PARENCHYMA RESECTION, WHEN THE DEVICE WAS ATTEMPTED TO BE FIRED, THE TEETH DID NOT ADVANCE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC DISTAL PANCREATECTOMY, THE SURGEON APPROACHED THE PANCREAS FOR RESECTION AND ATTEMPTED TO FIRE AFTER SEVERAL CLAMPS. HOWEVER, THE DEVICE DISPLAYED THAT THE FIRING HAS BEEN DONE. ALTHOUGH THE GREEN BUTTON WAS PRESSED, THE FIRING COULD NOT BE PERFORMED. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. IT WAS SEEN THAT THE STAPLES AT THE PROXIMAL SITE OF THE CARTRIDGE HAD BEEN PUSHED UP. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC DISTAL PANCREATECTOMY, THE SURGEON APPROACHED THE PANCREAS FOR RESECTION AND ATTEMPTED TO FIRE AFTER SEVERAL CLAMPS. HOWEVER, THE DEVICE DISPLAYED THAT THE FIRING HAS BEEN DONE. ALTHOUGH THE GREEN BUTTON WAS PRESSED, THE FIRING COULD NOT BE PERFORMED. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. IT WAS SEEN THAT THE STAPLES AT THE PROXIMAL SITE OF THE CARTRIDGE HAD BEEN PUSHED UP. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC BULLOUS RESECTION, WHILE THE LUNG TISSUE WAS BEING DETACHED, THE HANDLE AND 2 RELOADS DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DOCTOR USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC BULLOUS RESECTION, WHILE THE LUNG TISSUE WAS BEING DETACHED, THE HANDLE AND 2 RELOADS DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DOCTOR USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC BULLOUS RESECTION, WHILE THE LUNG TISSUE WAS BEING DETACHED, THE HANDLE AND 2 RELOADS DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DOCTOR USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | GST GREEN RELOAD, 60MM, 6 ROW | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE SURGEON REALIZED THERE WAS A SPECK OF GREEN THAT GOT STAPLED INTO THE PATIENT'S STOMACH. THE SURGEON BELIEVED IT CAME FROM THE CARTRIDGE HOUSING. HE INSPECTED ALL OF THE STAPLE CARTRIDGES AND DID NOT SEE ANY INDICATION THAT THE PIECE HAS BROKEN OFF. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
13-12-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 05-01-2022 | CAPSURE STRAIGHT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A TEP (TOTALLY EXTRAPERITONEAL) RIGHT INGUINAL HERNIA REPAIR PROCEDURE ON (B)(6) 2021, THE BARD/DAVOL CAPSURE STRAIGHT FIXATION DEVICE WAS USED TO FIXATE A BARD/DAVOL 3DMAX MESH. IT WAS REPORTED THAT THE FIRST FASTENER WAS PLACED IN THE COOPER'S LIGAMENT WITHOUT ANY PROBLEM AND WHEN THE SECOND FASTENER WAS FIRED, THE FASTENER HEAD WAS DISPLACED AND THE METAL PART WAS EXPOSED. AS REPORTED, THE SURGEON FELT THAT THE FASTENER HEAD WAS "WOBBLING" WHEN THE THIRD FASTENER WAS FIRED. AS REPORTED, THE DEVICE WAS REMOVED FROM THE PATIENT'S BODY, AND CYCLED IN AIR. THE PROCEDURE WAS COMPLETED USING A NEW CAPSURE DEVICE. THERE WAS NO REPORTED PATIENT INJURY.
13-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | UNK_ECHELON CIRCULAR POWERED STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE DEVICE WAS FIRED AS NORMAL WITH THE GREEN INDICATOR ILLUMINATED WHICH INDICATED THE FIRING SEQUENCE HAD COMPLETED SUCCESSFULLY. THE SURGEON DID 2 FULL TURNS TO REMOVE THE DEVICE. THE ANASTOMOTIC LEAK OCCURRED WHILE THE PATIENT WAS IN HOSPITAL A FEW DAYS LATER. THE PATIENT IS ON THE ROAD TO RECOVERY, BUT ANY LONG-TERM CONSEQUENCES ARE UNKNOWN. PROCEDURE: UNKNOWN.
01-01-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY CASE, THE SURGEON LOADED THE STAPLER AND WAS READY FOR FIRING. DURING FIRING THE FIRING KNOB WENT HALF THE WAY, THEN LOCKED AND NO STAPLES WERE APPLIED, AS IF NO FIRING HAPPENED. AFTER THE SURGEON CHANGED THE RELOAD THE STAPLER WORKED PERFECTLY. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
13-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE DEVICE WAS FIRED AS NORMAL WITH THE GREEN INDICATOR ILLUMINATED WHICH INDICATED THE FIRING SEQUENCE HAD COMPLETED SUCCESSFULLY. THE SURGEON DID 2 FULL TURNS TO REMOVE THE DEVICE. THE ANASTOMOTIC LEAK OCCURRED WHILE THE PATIENT WAS IN HOSPITAL A FEW DAYS LATER. THE PATIENT IS ON THE ROAD TO RECOVERY, BUT ANY LONG-TERM CONSEQUENCES ARE UNKNOWN. PROCEDURE: UNKNOWN.
13-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: IT WAS REPORTED THAT THE DEVICE WAS FIRED AS NORMAL WITH THE GREEN INDICATOR ILLUMINATED WHICH INDICATED THE FIRING SEQUENCE HAD COMPLETED SUCCESSFULLY. THE SURGEON DID 2 FULL TURNS TO REMOVE THE DEVICE. THE ANASTOMOTIC LEAK OCCURRED WHILE THE PATIENT WAS IN HOSPITAL A FEW DAYS LATER. THE PATIENT IS ON THE ROAD TO RECOVERY, BUT ANY LONG-TERM CONSEQUENCES ARE UNKNOWN. PROCEDURE: UNKNOWN.
14-12-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLER IS NOT A FIRED.
15-12-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN IVOR LEWIS ESOPHAGECTOMY WITH LAPAROSCOPIC NEPHRECTOMY, DURING GASTRO-ESOPHAGEAL ANASTOMOSIS, THE WHITE TROCAR PIECE ON THE PROXIMAL END OF THE ORAL ANVIL DID NOT COME OFF WITH THE PLASTIC GUIDE TUBING AND HAD TO BE REMOVED SEPARATELY. AFTER FIRING, THE SUTURE HAD TO BE CUT ON THE ORAL ANVIL IN ORDER TO RELEASE THE STAPLER. IT DID NOT CUT AUTOMATICALLY AFTER THE ANASTOMOSIS. THE SUTURE HAD TO BE CUT BY THE SURGEON IN THE THORACIC CAVITY. THE SUTURE ON THE DISTAL END WAS ALSO BUNCHED UP AS THOUGH IT HAD BEEN TANGLED. THE ORAL ANVIL DISENGAGED, FELL INTO THE CAVITY, AND WAS RETRIEVED. THE ANVIL WAS ALSO DIFFICULT TO OPEN. THE DONUTS WERE COMPLETE AND THE SURGEON WAS SATISFIED WITH THE ANASTOMOSIS DESPITE THE DIFFICULTY REMOVING THE STAPLER FROM THE CAVITY. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN IVOR LEWIS ESOPHAGECTOMY WITH LAPAROSCOPIC NEPHRECTOMY, DURING GASTRO-ESOPHAGEAL ANASTOMOSIS, THE WHITE TROCAR PIECE ON THE PROXIMAL END OF THE ORAL ANVIL DID NOT COME OFF WITH THE PLASTIC GUIDE TUBING AND HAD TO BE REMOVED SEPARATELY. AFTER FIRING, THE SUTURE HAD TO BE CUT ON THE ORAL ANVIL IN ORDER TO RELEASE THE STAPLER. IT DID NOT CUT AUTOMATICALLY AFTER THE ANASTOMOSIS. THE SUTURE HAD TO BE CUT BY THE SURGEON IN THE THORACIC CAVITY. THE SUTURE ON THE DISTAL END WAS ALSO BUNCHED UP AS THOUGH IT HAD BEEN TANGLED. THE ORAL ANVIL DISENGAGED, FELL INTO THE CAVITY, AND WAS RETRIEVED. THE ANVIL WAS ALSO DIFFICULT TO OPEN. THE DONUTS WERE COMPLETE AND THE SURGEON WAS SATISFIED WITH THE ANASTOMOSIS DESPITE THE DIFFICULTY REMOVING THE STAPLER FROM THE CAVITY. THERE WAS NO PATIENT INJURY.
15-12-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN IVOR LEWIS ESOPHAGECTOMY WITH LAPAROSCOPIC NEPHRECTOMY, DURING GASTRO-ESOPHAGEAL ANASTOMOSIS, THE WHITE TROCAR PIECE ON THE PROXIMAL END OF THE ORAL ANVIL DID NOT COME OFF WITH THE PLASTIC GUIDE TUBING AND HAD TO BE REMOVED SEPARATELY. AFTER FIRING, THE SUTURE HAD TO BE CUT ON THE ORAL ANVIL IN ORDER TO RELEASE THE STAPLER. IT DID NOT CUT AUTOMATICALLY AFTER THE ANASTOMOSIS. THE SUTURE HAD TO BE CUT BY THE SURGEON IN THE THORACIC CAVITY. THE SUTURE ON THE DISTAL END WAS ALSO BUNCHED UP AS THOUGH IT HAD BEEN TANGLED. THE ORAL ANVIL DISENGAGED, FELL INTO THE CAVITY, AND WAS RETRIEVED. THE ANVIL WAS ALSO DIFFICULT TO OPEN. THE DONUTS WERE COMPLETE AND THE SURGEON WAS SATISFIED WITH THE ANASTOMOSIS DESPITE THE DIFFICULTY REMOVING THE STAPLER FROM THE CAVITY. THERE WAS NO PATIENT INJURY.
08-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOIDECTOMY, THE SURGEON WAS ABLE TO FIRE THE FIRST RELOAD WITHOUT PROBLEM, BUT THE SECOND RELOAD WAS FAULTY, JAWS WERE JAMMED, AND THE DEVICE DID NOT FIRE. A DEVICE FROM A DIFFERENT MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-02-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SURGERY, COULD NOT FIRE WHEN SEVERING UNKNOWN TISSUE. NOTED THAT CARTRIDGE WAS DAMAGED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-02-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | FLEX 60 ARTICULATING | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SURGERY, COULD NOT FIRE WHEN SEVERING UNKNOWN TISSUE. NOTED THAT CARTRIDGE WAS DAMAGED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PUSH THE GREEN BUTTON AND SQUEEZE THE HANDLE.
08-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOIDECTOMY, THE SURGEON WAS ABLE TO FIRE THE FIRST RELOAD WITHOUT PROBLEM, BUT THE SECOND RELOAD WAS FAULTY, JAWS WERE JAMMED, AND THE DEVICE DID NOT FIRE. A DEVICE FROM A DIFFERENT MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-02-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | BLUE RELOAD FOR ECHELON 60 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: DEVICE DAMAGE & WOULD NOT FIRE. THIS CASE IS FROM HEALTH AUTHORITY. IT WAS REPORTED THAT DURING SURGERY, COULD NOT FIRE WHEN SEVERING UNKNOWN TISSUE. NOTED THAT CARTRIDGE WAS DAMAGED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-02-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | BLUE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: DEVICE DAMAGE & WOULD NOT FIRE. THIS CASE IS FROM HEALTH AUTHORITY. IT WAS REPORTED THAT DURING SURGERY, COULD NOT FIRE WHEN SEVERING UNKNOWN TISSUE. NOTED THAT CARTRIDGE WAS DAMAGED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTRIC CANCER RADICAL RESECTION PROCEDURE, WHILE STOMACH WAS SEVERED, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC DISTAL PANCREATECTOMY, THE SURGEON APPROACHED THE PANCREAS FOR RESECTION AND ATTEMPTED TO FIRE AFTER SEVERAL CLAMPS. HOWEVER, THE DEVICE DISPLAYED THAT THE FIRING HAS BEEN DONE. ALTHOUGH THE GREEN BUTTON WAS PRESSED, THE FIRING COULD NOT BE PERFORMED. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. IT WAS SEEN THAT THE STAPLES AT THE PROXIMAL SITE OF THE CARTRIDGE HAD BEEN PUSHED UP. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC DISTAL PANCREATECTOMY, THE SURGEON APPROACHED THE PANCREAS FOR RESECTION AND ATTEMPTED TO FIRE AFTER SEVERAL CLAMPS. HOWEVER, THE DEVICE DISPLAYED THAT THE FIRING HAS BEEN DONE. ALTHOUGH THE GREEN BUTTON WAS PRESSED, THE FIRING COULD NOT BE PERFORMED. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. IT WAS SEEN THAT THE STAPLES AT THE PROXIMAL SITE OF THE CARTRIDGE HAD BEEN PUSHED UP. THERE WAS NO PATIENT INJURY.
18-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, FOLLOWING LABORATORY EVALUATION, AFTER BEDSIDE-ASSIST STAPLED THROUGH SURGICAL DRAIN TUBING FOR CADAVERIC ESOPHAGECTOMY, STAPLER WAS RETRACTED AND THE TEAM NOTICED DAMAGE TO THE RELOAD. THE PROXIMAL END OF THE CHANNEL ADAPTER SHEARED OFF AT THE TRANSITION TO THE PORTION THAT LEAVES THE COVER TUBE. THERE WAS NO PATIENT INVOLVEMENT.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC GASTRIC BYPASS, WHILE ON THE SECOND FIRING ON THE STOMACH, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. IT STOPPED APPROXIMATELY 10MM FROM THE END OF THE RELOAD AND WOULD NOT CONTINUE AND WOULD NOT RETRACT. THE JAWS OF THE DEVICE WAS LOCKED ON TISSUE AS WELL. SHELL WAS REMOVED AND REATTACHED TO THE ADAPTER BUT STILL NO OPENING OF THE JAWS. A MANUAL RETRACTION TOOL WAS USED TO OPEN THE STAPLE JAWS. THE MIDDLE SLOT IN THE ADAPTOR WAS USED IN CONJUNCTION WITH THE MANUAL RETRACTION TOOL. A MANUAL HANDLE AND NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS SURGICAL TIME EXTENSION OF 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC GASTRIC BYPASS, WHILE ON THE SECOND FIRING ON THE STOMACH, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. IT STOPPED APPROXIMATELY 10MM FROM THE END OF THE RELOAD AND WOULD NOT CONTINUE AND WOULD NOT RETRACT. THE JAWS OF THE DEVICE WAS LOCKED ON TISSUE AS WELL. SHELL WAS REMOVED AND REATTACHED TO THE ADAPTER BUT STILL NO OPENING OF THE JAWS. A MANUAL RETRACTION TOOL WAS USED TO OPEN THE STAPLE JAWS. THE MIDDLE SLOT IN THE ADAPTOR WAS USED IN CONJUNCTION WITH THE MANUAL RETRACTION TOOL. A MANUAL HANDLE AND NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS SURGICAL TIME EXTENSION OF 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
13-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING RECTUM RESECTION IN A LAPAROSCOPIC PROCTECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. A CRACKING SOUND WAS HEARD UPON AN ATTEMPT TO SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON REPLACED THE HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING RECTUM RESECTION IN A LAPAROSCOPIC PROCTECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. A CRACKING SOUND WAS HEARD UPON AN ATTEMPT TO SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON REPLACED THE HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING RECTUM RESECTION IN A LAPAROSCOPIC PROCTECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. A CRACKING SOUND WAS HEARD UPON AN ATTEMPT TO SQUEEZE THE HANDLE. THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON REPLACED THE HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE MALFORMED STAPLES CAUSED BLEEDING UPON FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
10-01-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX 60 LONG | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE MALFORMED STAPLES CAUSED BLEEDING UPON FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PUSH THE GREEN BUTTON AND SQUEEZE THE HANDLE.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC GASTRIC BYPASS, WHILE ON THE SECOND FIRING ON THE STOMACH, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. IT STOPPED APPROXIMATELY 10MM FROM THE END OF THE RELOAD AND WOULD NOT CONTINUE AND WOULD NOT RETRACT. THE JAWS OF THE DEVICE WAS LOCKED ON TISSUE AS WELL. THE SHELL WAS REMOVED AND REATTACHED TO THE ADAPTER BUT STILL NO OPENING OF THE JAWS. A MANUAL RETRACTION TOOL WAS USED TO OPEN THE STAPLE JAWS. THE MIDDLE SLOT IN THE ADAPTOR WAS USED IN CONJUNCTION WITH THE MANUAL RETRACTION TOOL. A MANUAL HANDLE AND NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS SURGICAL TIME EXTENSION OF 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC GASTRIC BYPASS, WHILE ON THE SECOND FIRING ON THE STOMACH, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. IT STOPPED APPROXIMATELY 10MM FROM THE END OF THE RELOAD AND WOULD NOT CONTINUE AND WOULD NOT RETRACT. THE JAWS OF THE DEVICE WAS LOCKED ON TISSUE AS WELL. THE SHELL WAS REMOVED AND REATTACHED TO THE ADAPTER BUT STILL NO OPENING OF THE JAWS. A MANUAL RETRACTION TOOL WAS USED TO OPEN THE STAPLE JAWS. THE MIDDLE SLOT IN THE ADAPTOR WAS USED IN CONJUNCTION WITH THE MANUAL RETRACTION TOOL. A MANUAL HANDLE AND NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS SURGICAL TIME EXTENSION OF 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Injury | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DISSECTING THE STOMACH TISSUE, THE STAPLES DID NOT FORM PROPERLY, SOME STAPLES WERE NOT FORMED. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THE SURGEON SUTURED THE STAPLE LINE TO FIX THE STAPLE LINE AND A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE.
20-12-2021 | Injury | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DISSECTING THE STOMACH TISSUE, THE STAPLES DID NOT FORM PROPERLY, SOME STAPLES WERE NOT FORMED. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THE SURGEON SUTURED THE STAPLE LINE TO FIX THE STAPLE LINE AND A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE.
20-12-2021 | Injury | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DISSECTING THE STOMACH TISSUE, THE STAPLES DID NOT FORM PROPERLY, SOME STAPLES WERE NOT FORMED. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THE SURGEON SUTURED THE STAPLE LINE TO FIX THE STAPLE LINE AND A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE.
20-12-2021 | Injury | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DISSECTING THE STOMACH TISSUE, THE STAPLES DID NOT FORM PROPERLY, SOME STAPLES WERE NOT FORMED. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THE SURGEON SUTURED THE STAPLE LINE TO FIX THE STAPLE LINE AND A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LOBECTOMY, WHILE RESECTING THE LUNG LOBES, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. ANOTHER RELOAD OF THE SAME MODEL WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC GASTRIC BYPASS, WHILE ON THE SECOND FIRING ON THE STOMACH, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. IT STOPPED APPROXIMATELY 10MM FROM THE END OF THE RELOAD AND WOULD NOT CONTINUE AND WOULD NOT RETRACT. THE JAWS OF THE DEVICE WAS LOCKED ON TISSUE AS WELL. THE SHELL WAS REMOVED AND REATTACHED TO THE ADAPTER BUT STILL NO OPENING OF THE JAWS. A MANUAL RETRACTION TOOL WAS USED TO OPEN THE STAPLE JAWS. THE MIDDLE SLOT IN THE ADAPTOR WAS USED IN CONJUNCTION WITH THE MANUAL RETRACTION TOOL. A MANUAL HANDLE AND NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS SURGICAL TIME EXTENSION OF 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC GASTRIC BYPASS, WHILE ON THE SECOND FIRING ON THE STOMACH, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. IT STOPPED APPROXIMATELY 10MM FROM THE END OF THE RELOAD AND WOULD NOT CONTINUE AND WOULD NOT RETRACT. THE JAWS OF THE DEVICE WAS LOCKED ON TISSUE AS WELL. THE SHELL WAS REMOVED AND REATTACHED TO THE ADAPTER BUT STILL NO OPENING OF THE JAWS. A MANUAL RETRACTION TOOL WAS USED TO OPEN THE STAPLE JAWS. THE MIDDLE SLOT IN THE ADAPTOR WAS USED IN CONJUNCTION WITH THE MANUAL RETRACTION TOOL. A MANUAL HANDLE AND NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS SURGICAL TIME EXTENSION OF 30 MINUTES OR MORE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | RELIAMAX | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTROINTESTINAL ANASTOMOSIS PROCEDURE, THE DEVICE DID NOT FIRE. THE BLACK PUSHER THUMB PIECE CANNOT BE MOVED AT ALL. THE LOADING UNIT WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT THE NEW POWERED VASCULAR GUN MISFIRED ON PULMONARY VESSEL. DURING VATS LOBECTOMY. VESSEL WAS BLEEDING IN BETWEEN THE STAPLE LINE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE NEW POWERED VASCULAR GUN MISFIRED ON PULMONARY VESSEL. DURING VATS LOBECTOMY. VESSEL WAS BLEEDING IN BETWEEN THE STAPLE LINE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Cut | A050702 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT THE NEW POWERED VASCULAR GUN MISFIRED ON PULMONARY VESSEL. DURING VATS LOBECTOMY. VESSEL WAS BLEEDING IN BETWEEN THE STAPLE LINE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE NEW POWERED VASCULAR GUN MISFIRED ON PULMONARY VESSEL. DURING VATS LOBECTOMY. VESSEL WAS BLEEDING IN BETWEEN THE STAPLE LINE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT THE NEW POWERED VASCULAR GUN MISFIRED ON PULMONARY VESSEL. DURING VATS LOBECTOMY. VESSEL WAS BLEEDING IN BETWEEN THE STAPLE LINE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE NEW POWERED VASCULAR GUN MISFIRED ON PULMONARY VESSEL. DURING VATS LOBECTOMY. VESSEL WAS BLEEDING IN BETWEEN THE STAPLE LINE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT THE NEW POWERED VASCULAR GUN MISFIRED ON PULMONARY VESSEL. DURING VATS LOBECTOMY. VESSEL WAS BLEEDING IN BETWEEN THE STAPLE LINE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2021 | Injury | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE NEW POWERED VASCULAR GUN MISFIRED ON PULMONARY VESSEL. DURING VATS LOBECTOMY. VESSEL WAS BLEEDING IN BETWEEN THE STAPLE LINE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
02-12-2021 | Malfunction | davis & geck caribe | 05-01-2022 | UNKNOWN ENDO STITCH SULU | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TLH (TOTAL LAPAROSCOPIC HYSTERECTOMY) SURGERY, WHILE SEWING THE VAGINAL CUFF, THE HANDLE DROPPED THE NEEDLE. THE SURGEON USED A NEW HANDLE AND NEEDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-10-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: MW # (B)(4).
16-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AN ISSUE WITH THIS STAPLER FAILING TO OPEN AND SOME DIFFICULTY WITH THE RELOAD. A NEW DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
16-12-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AN ISSUE WITH THIS STAPLER FAILING TO OPEN AND SOME DIFFICULTY WITH THE RELOAD. A NEW DEVICE WAS OPENED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN RIGHT HEMI-HEPATECTOMY, WHEN USED FOR LOBECTOMY, THE ADJUSTABLE TRIGGER BUTTON COULD NOT GO FURTHER AFTER BEING PUSHED HALFWAY (PARTIALLY FIRED). THE RELOAD WAS REPLACED AND TRY TO FIRE AGAIN, BUT DEVICE WAS STILL STUCK AND PARTIALLY FIRED AS WELL. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN RIGHT HEMI-HEPATECTOMY, WHEN USED FOR LOBECTOMY, THE ADJUSTABLE TRIGGER BUTTON COULD NOT GO FURTHER AFTER BEING PUSHED HALFWAY (PARTIALLY FIRED). THE RELOAD WAS REPLACED AND TRY TO FIRE AGAIN, BUT DEVICE WAS STILL STUCK AND PARTIALLY FIRED AS WELL. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE JAW DID NOT OPEN AFTER CLAMPING ON TISSUE. ANOTHER TOOL WAS USED TO OPEN THE JAWS AND AN ALTERNATIVE ADAPTER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE JAW DID NOT OPEN AFTER CLAMPING ON TISSUE. ANOTHER TOOL WAS USED TO OPEN THE JAWS AND AN ALTERNATIVE ADAPTER WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
15-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
15-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Difficult to Remove | A150207 | Fistula | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
15-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
15-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Material Twisted/Bent | A040609 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
15-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Material Twisted/Bent | A040609 | Fistula | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
15-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Material Twisted/Bent | A040609 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
15-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Mechanical Jam | A0506 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
15-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Mechanical Jam | A0506 | Fistula | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
15-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | EEA ORVIL | Mechanical Jam | A0506 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
21-12-2021 | Injury | us surgical puerto rico | 05-01-2022 | UNKNOWN EEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
21-12-2021 | Injury | us surgical puerto rico | 05-01-2022 | UNKNOWN EEA | Difficult to Remove | A150207 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
21-12-2021 | Injury | us surgical puerto rico | 05-01-2022 | UNKNOWN EEA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
21-12-2021 | Injury | us surgical puerto rico | 05-01-2022 | UNKNOWN EEA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ON ANVIL PLACEMENT, IT WAS NOT POSSIBLE TO MAKE THE ANVIL TILT WITH A LAPAROSCOPIC INSTRUMENT TO GET THE ANVIL TO THE UMBRELLA LIKE POSITION. IT WAS NOTED THE ANVIL WAS BENT. THE PROCEDURE WAS CONVERTED TO A LAPAROTOMY RESOLVE THE ISSUE.
01-01-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY, WHENEVER THEY WOULD ARTICULATE THE DEVICE THEY WOULD FIRE ONCE OR TWICE AND IT WOULD STOP WORKING. LIKE IT HAD LOST POWER. LIKE THE BATTERY WAS DEAD. TWO DEVICES WERE USED WITH THE SAME LOT NUMBER AND SAME ISSUE. CASE COMPLETED WITH A THIRD DEVICE. NO PATIENT HARM.
01-01-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY, WHENEVER THEY WOULD ARTICULATE THE DEVICE THEY WOULD FIRE ONCE OR TWICE AND IT WOULD STOP WORKING. LIKE IT HAD LOST POWER. LIKE THE BATTERY WAS DEAD. TWO DEVICES WERE USED WITH THE SAME LOT NUMBER AND SAME ISSUE. CASE COMPLETED WITH A THIRD DEVICE. NO PATIENT HARM.
16-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | EEA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ESOPHAGEAL SURGERY, WHILE IN THE ESOPHAGUS, THE DEVICE PARTIALLY FIRED.
16-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | EEA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ESOPHAGEAL SURGERY, WHILE IN THE ESOPHAGUS, THE DEVICE PARTIALLY FIRED.
12-11-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX 45 STANDARD | Tear, Rip or Hole in Device Packaging | A020504 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC LOW RECTAL RESECTION, OPEN THE STERILE PACKING NOTED THE DEVICE WAS DAMAGED (MIDDLE OF THE DECK UPWRAP). ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-11-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | ECHELON FLEX 45 STANDARD | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC LOW RECTAL RESECTION, OPEN THE STERILE PACKING NOTED THE DEVICE WAS DAMAGED (MIDDLE OF THE DECK UPWRAP). ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
21-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING LAPAROSCOPIC RADICAL RECTAL CANCER SURGERY, WHEN DETACHING BLOOD VESSEL, B-SHAPED STAPLES WOULD NOT FORM PROPERLY. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE RELOAD. THERE WAS NO PATIENT INJURY.
08-10-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING VIDEO ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, AFTER FIRING IN THE BRANCH OF PULMONARY ARTERY ABNORMAL BLEEDING WAS SEEN IN THE STAPLE LINE. THE SURGEON NEEDED TO DO COMPRESSION WITH STERILIZED GAUZE DURING THE EVENT APPROXIMATELY 5 MINS TO CONTROL THE BLEEDING. ADDITIONAL RELOADS WERE USED BUT TO SECTION DIFFERENT BRANCHES OF THE PULMONARY ARTERY AND VEIN WITHOUT ABNORMAL BLEEDING SEEN. THERE WAS NO PATIENT INJURY.