09-12-2025 | Injury | intuitive surgical | 05-01-2026 | SUREFORM | Insufficient Information | E2401 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PANCREATODUODENECTOMY, A STAPLE LINE LEAK AND BLEEDING OCCURRED, PROMPTING CONVERSION TO OPEN SURGERY. FOLLOWING DISSECTION, THE PANCREAS REMAINED ATTACHED TO THE PORTAL VEIN. THE SURGICAL TEAM USED A SUREFORM 30 STAPLER WITH A WHITE RELOAD TO SEPARATE THE TISSUE. ONE FIRING WAS PERFORMED; THE INITIAL STAPLE LINE APPEARED ADEQUATE, AND THE TEAM PROCEEDED WITH RECONSTRUCTION. NO ALARMS OR ISSUES WERE ENCOUNTERED DURING STAPLING. THERE WERE NO OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIALS BETWEEN THE INSTRUMENT JAWS, AND THE STAPLER WAS NOT FIRED OVER AN EXISTING STAPLE LINE. NO BUTTRESS MATERIAL WAS USED. THE TARGET TISSUE WAS NOT CALCIFIED, NOT EXPOSED TO CHEMOTHERAPY OR RADIATION, AND DEMONSTRATED ONLY SLIGHT TENSION WITHOUT BUNCHING. HOWEVER, 15¿20 MINUTES AFTER STAPLING, BLEEDING WAS NOTED AT THE STAPLE LINE NEAR THE SUPERIOR MESENTERIC VEIN AND PORTAL VEIN CONFLUENCE, SUGGESTING THE STAPLE LINE HAD NOT HELD. THERE WERE NO MALFORMED STAPLES OR INCOMPLETE STAPLE LINE IDENTIFIED. EMERGENCY OPEN CONVERSION WAS PERFORMED TO CONTROL THE BLEEDING, AND THE PATIENT WAS STABILIZED FOLLOWING AN ESTIMATED BLOOD LOSS OF ONE LITER. THE BLEEDING SITE WAS OVERSEWN, AND THE PATIENT RECEIVED 3 UNITS OF PACKED RED BLOOD CELLS. POSTOPERATIVELY, THE PATIENT DEVELOPED DELAYED GASTRIC EMPTYING, REQUIRED PROLONGED HOSPITALIZATION, BUT DID NOT REQUIRE ADDITIONAL SURGERY.
01-01-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | nan
12-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON FLEX VASCULAR STAPLER | Misfire | A050502 | Insufficient Information | Event Description: STAPLE MISFIRED AND WOULDN'T OPEN, HAD TO USE MANUAL OPEN.
12-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON FLEX VASCULAR STAPLER | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: STAPLE MISFIRED AND WOULDN'T OPEN, HAD TO USE MANUAL OPEN.
17-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS PARTIAL FIRING OF RELOAD. THERE WAS NO BLEEDING AT THE STAPLE LINE AS IT APPEARED INTACT. NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PROXIMATE*75MM LINEAR CUTTER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THEY CANNOT LOCKED THE HANDLE AND THE STAPLES WERE MALFORMED. NO FURTHER INFORMATION PROVIDED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
17-10-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON LINEAR CUTTER 80MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT HEMICOLECTOMY, THE STAPLER WORKED ON FIRST FIRING, AFTER CLEANING AND RELOADING IT WOULDN'T FIRE AND FELL APART - PIN CAME OUT. SECOND DEVICE WAS OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO DELAY IN THE PROCEDURE AND COMPLETES SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
17-10-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON LINEAR CUTTER 80MM | Packaging Problem | A0205 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT HEMICOLECTOMY, THE STAPLER WORKED ON FIRST FIRING, AFTER CLEANING AND RELOADING IT WOULDN'T FIRE AND FELL APART - PIN CAME OUT. SECOND DEVICE WAS OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO DELAY IN THE PROCEDURE AND COMPLETES SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
20-10-2025 | Injury | ethicon endo-surgery | 05-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT A PATIENT NEEDED TO BE BROUGHT BACK FOR SURGERY DUE TO A LEAK WHILE PERFORMING A LOWER ANTERIOR RESECTION.
15-10-2025 | Injury | ethicon endo-surgery | 05-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE SURGEON HAD A LEAK BUT DID A FULL CIRCUMFERENCE STITCH FOR REINFORCEMENT. THE PROCEDURE WAS DELAYED DUE TO HAVING TO OVER SEW.
10-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ILS 21MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING UNK PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Misfire | A050502 | Abdominal Pain | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Misfire | A050502 | Emotional Changes | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Misfire | A050502 | Pain | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Misfire | A050502 | Distress | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Misfire | A050502 | Unspecified Tissue Injury | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Misfire | A050502 | Infertility | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Adverse Event Without Identified Device or Use Problem | A24 | Emotional Changes | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Adverse Event Without Identified Device or Use Problem | A24 | Distress | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Adverse Event Without Identified Device or Use Problem | A24 | Infertility | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Insufficient Information | E2401 | Abdominal Pain | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Insufficient Information | E2401 | Emotional Changes | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Insufficient Information | E2401 | Pain | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Insufficient Information | E2401 | Distress | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
16-09-2012 | Injury | covidien surgical | 05-01-2026 | MEDTRONIC | Insufficient Information | E2401 | Infertility | Event Description: THE PATIENTS ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. AFTER A LAPAROSCOPIC PROCEDURE AND INCIDENTAL APPENDECTOMY, IN A SUBSEQUENT DIAGNOSTIC LAPAROSCOPY, NUMEROUS LOOSE STAPLES WERE FOUND SCATTERED THROUGHOUT THE PATIENT'S PELVIS, INCLUDING MULTIPLE STAPLES DIRECTLY OVER THE VAGINAL CUFF. THIS INDICATED THAT THE STAPLING SYSTEM MALFUNCTIONED AND/OR MISFIRED RESULTING IN UNCONTROLLED STAPLE ESCAPE INTO THE OPERATIVE FIELD AND/OR FAILURE TO RETAIN STAPLES IN THE INTENDED STAPLE LINE. IT WAS REPORTED THAT AFTER USE AGE OF THE DEVICE, THE PATIENT EXPERIENCED LOOSE STAPLES, STAPLER MALFUNCTION, STAPLER MISFIRE, STAPLER FAILED TO RETAIN STAPLES IN STAPLE LINE, PAIN, LOSS OF FERTILITY, EMOTIONAL DISTRESS, ABDOMINAL/PELVIC PAIN, MENTAL ANGUISH, LOSS OF CAPACITY FOR ENJOYMENT OF LIFE, PERMANENT IMPAIRMENT, AND SUFFERING. POST OPERATIVE PATIENT TREATMENT INCLUDED DIAGNOSTIC LAPAROSCOPY, SURGICAL REMOVAL OF STAPLES, SURGICALLY INDUCED MENOPAUSE, GYNECOLOGIC SURGERIES AND REMOVAL OF REPRODUCTIVE ORGANS.
01-01-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | nan
02-12-2025 | Injury | ethicon endo-surgery | 05-01-2026 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | Insufficient Information | nan
02-12-2025 | Injury | ethicon endo-surgery | 05-01-2026 | ECHELON CONTOUR W/ GREEN RELOAD | Insufficient Information | E2401 | Insufficient Information | nan
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | CIRCULAR MECH STD 29MM | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, UNABLE TO CLICK PIN ON ANVIL. DIFFICULT EVEN EXTERNALLY TO USE. IMPOSSIBLE TO DO ANASTOMOSIS COLO RECTAL COELIO. CHANGE OF DEVICE TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES WERE REPORTED.
10-10-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RIGHT HEMICOLECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, THE SURGEON CLAIMS THAT THE INSTRUMENT DIDN'T CUT AND ALL THE CLIPS NEVER FORMED. THEY ARE NOT ON THE TAPE AS SHAPED. THERE WAS NO DELAY IN THE PROCEDURE AND THE DEVICE WAS REPLACED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
07-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, THE SURGEON CLAIMS THAT THE INSTRUMENT DIDN'T CUT AND ALL THE CLIPS NEVER FORMED. THEY ARE NOT ON THE TAPE AS SHAPED. THERE WAS NO DELAY IN THE PROCEDURE AND THE DEVICE WAS REPLACED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
05-01-2026 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 45 STANDARD | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A CASE OF AN INTERSTITIAL PNEUMONIA, WHEN TRIED TO PERFORM AN INTERSEGMENTAL RESECTION BY THE ECH45S, THE TISSUE WAS INCORPORATED TOO MUCH, AND IT WAS SO HARD THAT THE KNIFE BROKE HALFWAY THROUGH. THE KNIFE THAT HAD FALLEN TO THE PATIENT WAS PICKED UP, THE ECH60S WAS TAKEN OUT, AND THE TISSUE WAS CUT INTO SMALL PIECES, WHICH WAS COMPLETED WITHOUT ANY PROBLEMS. THE DEVICE WAS USED ON THE LUNG PARENCHYMA. THE PATIENT WAS DISCHARGED SAFELY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
05-01-2026 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A CASE OF AN INTERSTITIAL PNEUMONIA, WHEN TRIED TO PERFORM AN INTERSEGMENTAL RESECTION BY THE ECH45S, THE TISSUE WAS INCORPORATED TOO MUCH, AND IT WAS SO HARD THAT THE KNIFE BROKE HALFWAY THROUGH. THE KNIFE THAT HAD FALLEN TO THE PATIENT WAS PICKED UP, THE ECH60S WAS TAKEN OUT, AND THE TISSUE WAS CUT INTO SMALL PIECES, WHICH WAS COMPLETED WITHOUT ANY PROBLEMS. THE DEVICE WAS USED ON THE LUNG PARENCHYMA. THE PATIENT WAS DISCHARGED SAFELY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
05-01-2026 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 45 STANDARD | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A CASE OF AN INTERSTITIAL PNEUMONIA, WHEN TRIED TO PERFORM AN INTERSEGMENTAL RESECTION BY THE ECH45S, THE TISSUE WAS INCORPORATED TOO MUCH, AND IT WAS SO HARD THAT THE KNIFE BROKE HALFWAY THROUGH. THE KNIFE THAT HAD FALLEN TO THE PATIENT WAS PICKED UP, THE ECH60S WAS TAKEN OUT, AND THE TISSUE WAS CUT INTO SMALL PIECES, WHICH WAS COMPLETED WITHOUT ANY PROBLEMS. THE DEVICE WAS USED ON THE LUNG PARENCHYMA. THE PATIENT WAS DISCHARGED SAFELY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
05-01-2026 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 45 STANDARD | No Apparent Adverse Event | A25 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A CASE OF AN INTERSTITIAL PNEUMONIA, WHEN TRIED TO PERFORM AN INTERSEGMENTAL RESECTION BY THE ECH45S, THE TISSUE WAS INCORPORATED TOO MUCH, AND IT WAS SO HARD THAT THE KNIFE BROKE HALFWAY THROUGH. THE KNIFE THAT HAD FALLEN TO THE PATIENT WAS PICKED UP, THE ECH60S WAS TAKEN OUT, AND THE TISSUE WAS CUT INTO SMALL PIECES, WHICH WAS COMPLETED WITHOUT ANY PROBLEMS. THE DEVICE WAS USED ON THE LUNG PARENCHYMA. THE PATIENT WAS DISCHARGED SAFELY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
13-08-2025 | Injury | covidien surgical | 05-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: IT WAS REPORTED THAT, A WEEK AFTER A GASTRIC SLEEVE PROCEDURE, THE PATIENT PRESENTED A ONE-DAY HISTORY OF EXACERBATION OF UPPER ABDOMINAL PAIN RADIATING TO THE CHEST, SHORTNESS OF BREATH, AS WELL AS DRY MOUTH, DIZZINESS, AND WEAKNESS. IT WAS ALSO NOTED THAT THE PATIENT HAD MELENA, VERY ELEVATED RADIOFREQUENCY ABLATION (RFA) LABORATORY VALUES. COMPUTED TOMOGRAPHY (CT) SCAN DESCRIBED COLLECTION AND AIR OUTSIDE THE STOMACH. THE ABDOMEN WAS VERY DISTENDED, SOFT BUT WITH EPIGASTRIC TENDERNESS, IRRITATED, AND THE WOUNDS WERE HEALTHY. AN URGENT REOPERATION WAS FOR MANUAL SUTURING OF THE STAPLE LINE LEAK OF ONE RELOAD.
13-08-2025 | Injury | covidien surgical | 05-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Dyspnea | Event Description: IT WAS REPORTED THAT, A WEEK AFTER A GASTRIC SLEEVE PROCEDURE, THE PATIENT PRESENTED A ONE-DAY HISTORY OF EXACERBATION OF UPPER ABDOMINAL PAIN RADIATING TO THE CHEST, SHORTNESS OF BREATH, AS WELL AS DRY MOUTH, DIZZINESS, AND WEAKNESS. IT WAS ALSO NOTED THAT THE PATIENT HAD MELENA, VERY ELEVATED RADIOFREQUENCY ABLATION (RFA) LABORATORY VALUES. COMPUTED TOMOGRAPHY (CT) SCAN DESCRIBED COLLECTION AND AIR OUTSIDE THE STOMACH. THE ABDOMEN WAS VERY DISTENDED, SOFT BUT WITH EPIGASTRIC TENDERNESS, IRRITATED, AND THE WOUNDS WERE HEALTHY. AN URGENT REOPERATION WAS FOR MANUAL SUTURING OF THE STAPLE LINE LEAK OF ONE RELOAD.
13-08-2025 | Injury | covidien surgical | 05-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fatigue | Event Description: IT WAS REPORTED THAT, A WEEK AFTER A GASTRIC SLEEVE PROCEDURE, THE PATIENT PRESENTED A ONE-DAY HISTORY OF EXACERBATION OF UPPER ABDOMINAL PAIN RADIATING TO THE CHEST, SHORTNESS OF BREATH, AS WELL AS DRY MOUTH, DIZZINESS, AND WEAKNESS. IT WAS ALSO NOTED THAT THE PATIENT HAD MELENA, VERY ELEVATED RADIOFREQUENCY ABLATION (RFA) LABORATORY VALUES. COMPUTED TOMOGRAPHY (CT) SCAN DESCRIBED COLLECTION AND AIR OUTSIDE THE STOMACH. THE ABDOMEN WAS VERY DISTENDED, SOFT BUT WITH EPIGASTRIC TENDERNESS, IRRITATED, AND THE WOUNDS WERE HEALTHY. AN URGENT REOPERATION WAS FOR MANUAL SUTURING OF THE STAPLE LINE LEAK OF ONE RELOAD.
13-08-2025 | Injury | covidien surgical | 05-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Muscle Weakness/Atrophy | Event Description: IT WAS REPORTED THAT, A WEEK AFTER A GASTRIC SLEEVE PROCEDURE, THE PATIENT PRESENTED A ONE-DAY HISTORY OF EXACERBATION OF UPPER ABDOMINAL PAIN RADIATING TO THE CHEST, SHORTNESS OF BREATH, AS WELL AS DRY MOUTH, DIZZINESS, AND WEAKNESS. IT WAS ALSO NOTED THAT THE PATIENT HAD MELENA, VERY ELEVATED RADIOFREQUENCY ABLATION (RFA) LABORATORY VALUES. COMPUTED TOMOGRAPHY (CT) SCAN DESCRIBED COLLECTION AND AIR OUTSIDE THE STOMACH. THE ABDOMEN WAS VERY DISTENDED, SOFT BUT WITH EPIGASTRIC TENDERNESS, IRRITATED, AND THE WOUNDS WERE HEALTHY. AN URGENT REOPERATION WAS FOR MANUAL SUTURING OF THE STAPLE LINE LEAK OF ONE RELOAD.
13-08-2025 | Injury | covidien surgical | 05-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Dizziness | Event Description: IT WAS REPORTED THAT, A WEEK AFTER A GASTRIC SLEEVE PROCEDURE, THE PATIENT PRESENTED A ONE-DAY HISTORY OF EXACERBATION OF UPPER ABDOMINAL PAIN RADIATING TO THE CHEST, SHORTNESS OF BREATH, AS WELL AS DRY MOUTH, DIZZINESS, AND WEAKNESS. IT WAS ALSO NOTED THAT THE PATIENT HAD MELENA, VERY ELEVATED RADIOFREQUENCY ABLATION (RFA) LABORATORY VALUES. COMPUTED TOMOGRAPHY (CT) SCAN DESCRIBED COLLECTION AND AIR OUTSIDE THE STOMACH. THE ABDOMEN WAS VERY DISTENDED, SOFT BUT WITH EPIGASTRIC TENDERNESS, IRRITATED, AND THE WOUNDS WERE HEALTHY. AN URGENT REOPERATION WAS FOR MANUAL SUTURING OF THE STAPLE LINE LEAK OF ONE RELOAD.
13-08-2025 | Injury | covidien surgical | 05-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT, A WEEK AFTER A GASTRIC SLEEVE PROCEDURE, THE PATIENT PRESENTED A ONE-DAY HISTORY OF EXACERBATION OF UPPER ABDOMINAL PAIN RADIATING TO THE CHEST, SHORTNESS OF BREATH, AS WELL AS DRY MOUTH, DIZZINESS, AND WEAKNESS. IT WAS ALSO NOTED THAT THE PATIENT HAD MELENA, VERY ELEVATED RADIOFREQUENCY ABLATION (RFA) LABORATORY VALUES. COMPUTED TOMOGRAPHY (CT) SCAN DESCRIBED COLLECTION AND AIR OUTSIDE THE STOMACH. THE ABDOMEN WAS VERY DISTENDED, SOFT BUT WITH EPIGASTRIC TENDERNESS, IRRITATED, AND THE WOUNDS WERE HEALTHY. AN URGENT REOPERATION WAS FOR MANUAL SUTURING OF THE STAPLE LINE LEAK OF ONE RELOAD.
13-08-2025 | Injury | covidien surgical | 05-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Melena | Event Description: IT WAS REPORTED THAT, A WEEK AFTER A GASTRIC SLEEVE PROCEDURE, THE PATIENT PRESENTED A ONE-DAY HISTORY OF EXACERBATION OF UPPER ABDOMINAL PAIN RADIATING TO THE CHEST, SHORTNESS OF BREATH, AS WELL AS DRY MOUTH, DIZZINESS, AND WEAKNESS. IT WAS ALSO NOTED THAT THE PATIENT HAD MELENA, VERY ELEVATED RADIOFREQUENCY ABLATION (RFA) LABORATORY VALUES. COMPUTED TOMOGRAPHY (CT) SCAN DESCRIBED COLLECTION AND AIR OUTSIDE THE STOMACH. THE ABDOMEN WAS VERY DISTENDED, SOFT BUT WITH EPIGASTRIC TENDERNESS, IRRITATED, AND THE WOUNDS WERE HEALTHY. AN URGENT REOPERATION WAS FOR MANUAL SUTURING OF THE STAPLE LINE LEAK OF ONE RELOAD.
13-08-2025 | Injury | covidien surgical | 05-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: IT WAS REPORTED THAT, A WEEK AFTER A GASTRIC SLEEVE PROCEDURE, THE PATIENT PRESENTED A ONE-DAY HISTORY OF EXACERBATION OF UPPER ABDOMINAL PAIN RADIATING TO THE CHEST, SHORTNESS OF BREATH, AS WELL AS DRY MOUTH, DIZZINESS, AND WEAKNESS. IT WAS ALSO NOTED THAT THE PATIENT HAD MELENA, VERY ELEVATED RADIOFREQUENCY ABLATION (RFA) LABORATORY VALUES. COMPUTED TOMOGRAPHY (CT) SCAN DESCRIBED COLLECTION AND AIR OUTSIDE THE STOMACH. THE ABDOMEN WAS VERY DISTENDED, SOFT BUT WITH EPIGASTRIC TENDERNESS, IRRITATED, AND THE WOUNDS WERE HEALTHY. AN URGENT REOPERATION WAS FOR MANUAL SUTURING OF THE STAPLE LINE LEAK OF ONE RELOAD.
13-08-2025 | Injury | covidien surgical | 05-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT, A WEEK AFTER A GASTRIC SLEEVE PROCEDURE, THE PATIENT PRESENTED A ONE-DAY HISTORY OF EXACERBATION OF UPPER ABDOMINAL PAIN RADIATING TO THE CHEST, SHORTNESS OF BREATH, AS WELL AS DRY MOUTH, DIZZINESS, AND WEAKNESS. IT WAS ALSO NOTED THAT THE PATIENT HAD MELENA, VERY ELEVATED RADIOFREQUENCY ABLATION (RFA) LABORATORY VALUES. COMPUTED TOMOGRAPHY (CT) SCAN DESCRIBED COLLECTION AND AIR OUTSIDE THE STOMACH. THE ABDOMEN WAS VERY DISTENDED, SOFT BUT WITH EPIGASTRIC TENDERNESS, IRRITATED, AND THE WOUNDS WERE HEALTHY. AN URGENT REOPERATION WAS FOR MANUAL SUTURING OF THE STAPLE LINE LEAK OF ONE RELOAD.
12-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Appropriate Term/Code Not Available |  | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE ONLY FIRED HALFWAY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
10-12-2025 | Injury | covidien surgical | 05-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY, BLEEDING OCCURRED AT THE STAPLE LINE. SUTURES WERE APPLIED TO THE STAPLE LINE TO STOP THE BLEEDING. A COLONOSCOPY WAS PERFORMED AS WELL TO CAUTERIZE THE STAPLE LINE, BUT THE PATIENT WAS STILL PASSING BLOOD IN STOOL. IT WAS NOTED THAT THE PATIENT HAD PROLONGED HOSPITALIZATION.
10-12-2025 | Injury | covidien surgical | 05-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY, BLEEDING OCCURRED AT THE STAPLE LINE. SUTURES WERE APPLIED TO THE STAPLE LINE TO STOP THE BLEEDING. A COLONOSCOPY WAS PERFORMED AS WELL TO CAUTERIZE THE STAPLE LINE, BUT THE PATIENT WAS STILL PASSING BLOOD IN STOOL. IT WAS NOTED THAT THE PATIENT HAD PROLONGED HOSPITALIZATION.
18-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BARIATRIC SLEEVE PROCEDURE, IT WAS OBSERVED THAT THE STAPLING PRESENTED DEFORMITY AT THE END OF THE STAPLING AT THE FIRST STAPLE WITH THE GOLD REFILL. AND AFTER THE EXCHANGE WITH THE BLUE RELOAD, IT PRESENTED FAILURE AT THE TIME OF ACTIVATION, AND IT WAS POSSIBLE TO SEE OPEN STAPLES IN THE CAVITY. THERE WAS NO DIRECT IMPACT ON THE PATIENT, I WILL NEED THE SURGEON TO REINFORCE THE STAPLE LINE WITH SUTURES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
10-12-2025 | Injury | covidien surgical | 05-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY, BLEEDING OCCURRED AT THE STAPLE LINE. SUTURES WERE APPLIED TO THE STAPLE LINE TO STOP THE BLEEDING. A COLONOSCOPY WAS PERFORMED AS WELL TO CAUTERIZE THE STAPLE LINE, BUT THE PATIENT WAS STILL PASSING BLOOD IN STOOL. IT WAS NOTED THAT THE PATIENT HAD PROLONGED HOSPITALIZATION.
10-12-2025 | Injury | covidien surgical | 05-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY, BLEEDING OCCURRED AT THE STAPLE LINE. SUTURES WERE APPLIED TO THE STAPLE LINE TO STOP THE BLEEDING. A COLONOSCOPY WAS PERFORMED AS WELL TO CAUTERIZE THE STAPLE LINE, BUT THE PATIENT WAS STILL PASSING BLOOD IN STOOL. IT WAS NOTED THAT THE PATIENT HAD PROLONGED HOSPITALIZATION.
01-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
21-10-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RETROPERITONEAL LYMPHADENECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE STOPPED DURING USE, AND IT WAS NOT POSSIBLE TO PERFORM TRIGGERING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
12-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ENDO ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2025 | Malfunction | davol inc sub c.r bard inc -1213643 | 05-01-2026 | SORBAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, SORBAFIX STAPLES FAIL TO ATTACH AND FALL BACK INTO THE ABDOMEN. THE PROSTHESIS IS THEREFORE NOT FIXED AGAINST THE WALL. DELAY IN OPERATION. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT HARM OR INJURY.
18-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-12-2025 | Malfunction | covidien surgical | 05-01-2026 | SIGPHANDLE | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE DEVICE WAS LOCKED ON TISSUE, WHICH RESULTED IN EXTENDED SURGICAL TIME. THE DEVICE WAS REPLACED, AND AN EMERGENCY TOOL WAS USED TO ADDRESS THE SITUATION.
19-12-2025 | Malfunction | covidien surgical | 05-01-2026 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE DEVICE WAS LOCKED ON TISSUE, WHICH RESULTED IN EXTENDED SURGICAL TIME. THE DEVICE WAS REPLACED, AND AN EMERGENCY TOOL WAS USED TO ADDRESS THE SITUATION.
19-12-2025 | Malfunction | covidien surgical | 05-01-2026 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE DEVICE WAS LOCKED ON TISSUE, WHICH RESULTED IN EXTENDED SURGICAL TIME. THE DEVICE WAS REPLACED, AND AN EMERGENCY TOOL WAS USED TO ADDRESS THE SITUATION.
08-12-2025 | Injury | intuitive surgical | 05-01-2026 | SUREFORM | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.
05-02-2026 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST BLUE RELOAD 45MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
05-02-2026 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
16-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 60MM COMPACT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
16-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 60MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
31-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO ANOTHER THREE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO ANOTHER THREE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO ANOTHER THREE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
31-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO ANOTHER THREE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2025 | Malfunction | covidien surgical | 05-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THERE WAS POOR STAPLE FORMATION. THE DISTAL STAPLE LINE WAS ALSO INCOMPLETE. THE ISSUES HAPPENED WITH THREE DIFFERENT RELOADS. THE STAPLE LINES WERE REINFORCED WITH SUTURE TO RESOLVE THE ISSUE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
22-12-2025 | Malfunction | covidien | 05-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THERE WAS POOR STAPLE FORMATION. THE DISTAL STAPLE LINE WAS ALSO INCOMPLETE. THE ISSUES HAPPENED WITH THREE DIFFERENT RELOADS. THE STAPLE LINES WERE REINFORCED WITH SUTURE TO RESOLVE THE ISSUE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
30-12-2025 | Malfunction | covidien | 05-01-2026 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, IN TISSUE DETACHMENT, THE HANDLE ROD BROKE OFF. NO COMPONENTS FELL INTO THE PATIENT'S CAVITY. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
22-12-2025 | Malfunction | covidien | 05-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THERE WAS POOR STAPLE FORMATION. THE DISTAL STAPLE LINE WAS ALSO INCOMPLETE. THE ISSUES HAPPENED WITH THREE DIFFERENT RELOADS. THE STAPLE LINES WERE REINFORCED WITH SUTURE TO RESOLVE THE ISSUE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
30-12-2025 | Malfunction | covidien | 06-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE SO BOTH, HANDLE AND RELOAD DID NOT FIRE. TO RESOLVE THE ISSUE, BOTH THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE PROCEDURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
30-12-2025 | Malfunction | covidien surgical | 06-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE SO BOTH, HANDLE AND RELOAD DID NOT FIRE. TO RESOLVE THE ISSUE, BOTH THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE PROCEDURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2025 | Malfunction | covidien surgical | 06-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE JAWS DID NOT CLOSE COMPLETELY SO THE STAPLER DID NOT FIRE AS EXPECTED. SEVERAL TIMES, THE SURGEON CLOSED THE RELOADS ON DIFFERENT TYPES OF TISSUE, INCLUDING LUNG PARENCHYMA, VESSEL, AND BRONCHUS, BUT THE EXPECTED FIRING THICKNESS INDICATORS DID NOT APPEAR ON THE SCREEN, AND IT SEEMED THAT THE JAWS DID NOT CLOSE PROPERLY. MULTIPLE ATTEMPTS WERE REQUIRED TO ACHIEVE PROPER DEVICE FUNCTION, AND SOMETIMES AFTER TWO OR THREE TRIES, THE ISSUE WAS RESOLVED AND THE DEVICE FIRED AS INTENDED. HOWEVER, THESE DIFFICULTIES OCCURRED REPEATEDLY IN TWO DIFFERENT PROCEDURES ON THE SAME DAY. TROUBLESHOOTING STEPS INCLUDED CHANGING THE POWER HANDLE, SHELL, AND RELOADS, BUT THE ISSUE PERSISTED. THE POWER HANDLE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2025 | Malfunction | covidien surgical | 06-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE JAWS DID NOT CLOSE COMPLETELY SO THE STAPLER DID NOT FIRE AS EXPECTED. SEVERAL TIMES, THE SURGEON CLOSED THE RELOADS ON DIFFERENT TYPES OF TISSUE, INCLUDING LUNG PARENCHYMA, VESSEL, AND BRONCHUS, BUT THE EXPECTED FIRING THICKNESS INDICATORS DID NOT APPEAR ON THE SCREEN, AND IT SEEMED THAT THE JAWS DID NOT CLOSE PROPERLY. MULTIPLE ATTEMPTS WERE REQUIRED TO ACHIEVE PROPER DEVICE FUNCTION, AND SOMETIMES AFTER TWO OR THREE TRIES, THE ISSUE WAS RESOLVED AND THE DEVICE FIRED AS INTENDED. HOWEVER, THESE DIFFICULTIES OCCURRED REPEATEDLY IN TWO DIFFERENT PROCEDURES ON THE SAME DAY. TROUBLESHOOTING STEPS INCLUDED CHANGING THE POWER HANDLE, SHELL, AND RELOADS, BUT THE ISSUE PERSISTED. THE POWER HANDLE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2025 | Malfunction | covidien surgical | 06-01-2026 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE JAWS DID NOT CLOSE COMPLETELY SO THE STAPLER DID NOT FIRE AS EXPECTED. SEVERAL TIMES, THE SURGEON CLOSED THE RELOADS ON DIFFERENT TYPES OF TISSUE, INCLUDING LUNG PARENCHYMA, VESSEL, AND BRONCHUS, BUT THE EXPECTED FIRING THICKNESS INDICATORS DID NOT APPEAR ON THE SCREEN, AND IT SEEMED THAT THE JAWS DID NOT CLOSE PROPERLY. MULTIPLE ATTEMPTS WERE REQUIRED TO ACHIEVE PROPER DEVICE FUNCTION, AND SOMETIMES AFTER TWO OR THREE TRIES, THE ISSUE WAS RESOLVED AND THE DEVICE FIRED AS INTENDED. HOWEVER, THESE DIFFICULTIES OCCURRED REPEATEDLY IN TWO DIFFERENT PROCEDURES ON THE SAME DAY. TROUBLESHOOTING STEPS INCLUDED CHANGING THE POWER HANDLE, SHELL, AND RELOADS, BUT THE ISSUE PERSISTED. THE POWER HANDLE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2025 | Malfunction | covidien surgical | 06-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) PROCEDURE, THE JAWS DID NOT CLOSE COMPLETELY SO THE STAPLER DID NOT FIRE AS EXPECTED. SEVERAL TIMES, THE SURGEON CLOSED THE RELOADS ON DIFFERENT TYPES OF TISSUE, INCLUDING LUNG PARENCHYMA, VESSEL, AND BRONCHUS, BUT THE EXPECTED FIRING THICKNESS INDICATORS DID NOT APPEAR ON THE SCREEN, AND IT SEEMED THAT THE JAWS DID NOT CLOSE PROPERLY. MULTIPLE ATTEMPTS WERE REQUIRED TO ACHIEVE PROPER DEVICE FUNCTION, AND SOMETIMES AFTER TWO OR THREE TRIES, THE ISSUE WAS RESOLVED AND THE DEVICE FIRED AS INTENDED. TROUBLESHOOTING STEPS INCLUDED CHANGING THE POWER HANDLE, SHELL, AND RELOADS, BUT THE ISSUE PERSISTED. TO RESOLVE THE ISSUE, THE POWER HANDLE WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2025 | Malfunction | covidien surgical | 06-01-2026 | UNK-SIGADAPT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) PROCEDURE, THE JAWS DID NOT CLOSE COMPLETELY SO THE STAPLER DID NOT FIRE AS EXPECTED. SEVERAL TIMES, THE SURGEON CLOSED THE RELOADS ON DIFFERENT TYPES OF TISSUE, INCLUDING LUNG PARENCHYMA, VESSEL, AND BRONCHUS, BUT THE EXPECTED FIRING THICKNESS INDICATORS DID NOT APPEAR ON THE SCREEN, AND IT SEEMED THAT THE JAWS DID NOT CLOSE PROPERLY. MULTIPLE ATTEMPTS WERE REQUIRED TO ACHIEVE PROPER DEVICE FUNCTION, AND SOMETIMES AFTER TWO OR THREE TRIES, THE ISSUE WAS RESOLVED AND THE DEVICE FIRED AS INTENDED. TROUBLESHOOTING STEPS INCLUDED CHANGING THE POWER HANDLE, SHELL, AND RELOADS, BUT THE ISSUE PERSISTED. TO RESOLVE THE ISSUE, THE POWER HANDLE WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2025 | Malfunction | covidien surgical | 06-01-2026 | UNKNOWN EGIA SU | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) PROCEDURE, THE JAWS DID NOT CLOSE COMPLETELY SO THE STAPLER DID NOT FIRE AS EXPECTED. SEVERAL TIMES, THE SURGEON CLOSED THE RELOADS ON DIFFERENT TYPES OF TISSUE, INCLUDING LUNG PARENCHYMA, VESSEL, AND BRONCHUS, BUT THE EXPECTED FIRING THICKNESS INDICATORS DID NOT APPEAR ON THE SCREEN, AND IT SEEMED THAT THE JAWS DID NOT CLOSE PROPERLY. MULTIPLE ATTEMPTS WERE REQUIRED TO ACHIEVE PROPER DEVICE FUNCTION, AND SOMETIMES AFTER TWO OR THREE TRIES, THE ISSUE WAS RESOLVED AND THE DEVICE FIRED AS INTENDED. TROUBLESHOOTING STEPS INCLUDED CHANGING THE POWER HANDLE, SHELL, AND RELOADS, BUT THE ISSUE PERSISTED. TO RESOLVE THE ISSUE, THE POWER HANDLE WAS REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
05-12-2025 | Malfunction | covidien surgical | 06-01-2026 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, THE SHIPPING WEDGE WAS LEFT ON THE DEVICE WHILE THE CARTRIDGE WAS LOADED ONTO THE STAPLER AND THE JAWS WERE OPENED AND CLOSED. THE SHIPPING WEDGE WAS THEN REMOVED WHILE THE JAWS WERE CLOSED. AFTER THE STAPLER WAS USED AND FIRED, TWO SMALL PIECES OF THE SHIPPING WEDGE WERE DISCOVERED IN THE PATIENT CAVITY AND WERE RETRIEVED BY THE SURGEON USING GRASPERS. NO INTERVENTION WAS REQUIRED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
08-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RECTUM PROCEDURE, AFTER FIRED, NOTED THE STAPLES MALFORMED. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ILS 29MM CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RECTAL CANCER RADICAL SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO THE NEW DEVICE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-11-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON CONTOUR W/ BLUE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE WOULD NOT FIRE. SURGICAL DELAY WAS UNKNOWN. NO PATIENT HARM WAS REPORTED.
27-11-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | CIRCULAR MECH STD 29MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING CANCER RESECTION FROM COLON PART, USED A NEW CHD29B, BUT AFTER FIRING AND REMOVING THE DEVICE, HE FOUND NO STAPLES AND NO TISSUE CUT OCCURRED. SO, HE OPENED A NEW CDH29B DEVICE TO CONTINUE THE PROCEDURE.
06-01-2026 | Malfunction | covidien surgical | 06-01-2026 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INGUINAL HERNIA PROCEDURE, WHILE FIXING THE MESH, THE TACKER WOULD NOT LOAD PROPERLY, REQUIRING THE USER TO PRESSED MULTIPLE TIMES FOR THE TACKER TO FIRE. ADDITIONALLY, THE TACKS WERE NOT FIXED ON THE MESH PROPERLY AND FELL INTO THE PATIENT'S CAVITY. THE TACKS THAT FELL WERE RETRIEVED. THIS ISSUE OCCURRED ON TWO DEVICES. TO COMPLETE THE CASE, TWO MORE TACKERS WERE USED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
06-01-2026 | Malfunction | covidien surgical | 06-01-2026 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INGUINAL HERNIA PROCEDURE, WHILE FIXING THE MESH, THE TACKER WOULD NOT LOAD PROPERLY, REQUIRING THE USER TO PRESSED MULTIPLE TIMES FOR THE TACKER TO FIRE. ADDITIONALLY, THE TACKS WERE NOT FIXED ON THE MESH PROPERLY AND FELL INTO THE PATIENT'S CAVITY. THE TACKS THAT FELL WERE RETRIEVED. THIS ISSUE OCCURRED ON TWO DEVICES. TO COMPLETE THE CASE, TWO MORE TACKERS WERE USED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-01-2025 | Injury | ethicon endo-surgery | 06-01-2026 | CIRCULAR MECH STD 29MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, AND A LEAK OCCURRED. THE PATIENT UNDERWENT AN ADDITIONAL PROCEDURE. NO FURTHER INFORMATION AVAILABLE.
12-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
18-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM STANDARD | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | nan
19-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST GOLD RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES WERE FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING AFTER FIRING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
19-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST GOLD RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES WERE FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING AFTER FIRING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
19-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST GOLD RELOAD 45MM 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES WERE FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING AFTER FIRING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
19-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST GOLD RELOAD 45MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES WERE FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING AFTER FIRING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
21-10-2025 | Malfunction | johnson & johnson medical (suzhou | 06-01-2026 | FINAL ASSY 60MM STD DEVICE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THEY COULD NOT OPEN THE JAW AFTER FIRING. THEY USED THE REVERSE BUTTON TO OPEN THE JAW. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT IT WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
21-10-2025 | Malfunction | johnson & johnson medical (suzhou | 06-01-2026 | FINAL ASSY 60MM STD DEVICE | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THEY COULD NOT OPEN THE JAW AFTER FIRING. THEY USED THE REVERSE BUTTON TO OPEN THE JAW. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT IT WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN RESPIRATORY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE STOPPED MIDWAY AT THE 3RD FIRING. THE KNIFE MOVED FORWARD A LITTLE BUT THEN STOPPED WHEN PRESSED AGAIN. THE KNIFE WAS RETURNED, A NEW ONE WAS REPLACED, AND THE SURGERY WAS COMPLETED WITHOUT ANY PROBLEMS. THE CARTRIDGE COLOR WAS BLACK. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
18-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE AN UNKNOWN LAPAROSCOPIC PROCEDURE, IT WAS OBSERVED THAT THERE WAS AN ECH45S IN THE PACKAGE OF AN ECH60S UPON OPENING. THIS WAS FOUND BEFORE USE, SO IT IS UNUSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
03-01-2025 | Malfunction | covidien | 06-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, AFTER FIRING, THE STAPLE LINE DID NOT HOLD OR OPENED UP. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
19-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM COMPACT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WERE 3 UNFORMED STAPLES ALONG THE STAPLE LINE, AFTER FIRING AN ECH60C WITH A BLUE LOAD. THIS WAS THE 3RD FIRING WITH THE DEVICE ON DESCENDING COLON. THE SURGEON DID NOT NOTICE THE MALFORMED STAPLES RIGHT AWAY, BUT NOTICED THEM ON THE SCREEN LATER ON IN THE PROCEDURE. THE STAPLE LEGS WERE STICKING STRAIGHT UP. HE NOTED IT WAS APPX 3 STAPLES, AND IT WAS THE INNERMOST ROW CLOSEST TO THE CUT LINE. UNKNOWN HOW THE STAPLES LOOKED ON THE SPECIMEN SIDE, AS IT WAS ALREADY PACKAGED. THE SURGEON OVERSEWED THE STAPLE LINE AND EVERYTHING WAS FINE. THE DEVICE DID NOT NEED TO BE FIRED AFTER THEN. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
21-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM LONG | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | nan
01-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM LONG | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
22-12-2025 | Malfunction | johnson & johnson medical (suzhou | 06-01-2026 | FINAL ASSY 445MM STD DEVICE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-10-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST BLUE RELOAD 45MM 6 ROW | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY TO AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAD AN UNSPECIFIED MALFUNCTION. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
17-10-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM STANDARD | Structural Problem | A0511 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT SOME OF THE BLADES WERE AUTOMATICALLY FIRED BEFORE USE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST WHITE RELOAD, 45MM, 6 ROW | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE GST45W RELOAD COULD NOT MATCH WITH THE GST GUN. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM STANDARD | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAD AN UNSPECIFIED MALFUNCTION AS IT DID NOT FIRE PROPERLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM STANDARD | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A WHIPPLE PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT INSERT THE SECOND RELOAD AS IF THE JAWS WERE OBSTRUCTED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
10-11-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON FIRED THE STAPLER ON BOWEL AND THE TISSUE WAS STUCK IN THE CROTCH OF THE JAWS AFTER FIRING. THE STAFF HAD TROUBLE REMOVING THE TISSUE FROM THE STAPLER AND OPENED ANOTHER TO FINISH THE CASE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
10-11-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON FIRED THE STAPLER ON BOWEL AND THE TISSUE WAS STUCK IN THE CROTCH OF THE JAWS AFTER FIRING. THE STAFF HAD TROUBLE REMOVING THE TISSUE FROM THE STAPLER AND OPENED ANOTHER TO FINISH THE CASE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT CLOSE WHEN THERE WAS TISSUE INSIDE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-11-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN HARTMANN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT BE FIRED. WHEN CHECKED THE CARTRIDGE, IT WAS LOCKED OUT. THE CARTRIDGE COLOR WAS BLUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
18-11-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM LONG | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE JAWS WOULD NOT CLOSE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
24-11-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
24-11-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM LONG | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP SLEEVE PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS UNABLE TO RELOAD AFTER TWO FIRINGS. THE RELOAD WAS BEING BLOCKED. BECAUSE THEY WERE UNABLE TO RELOAD THE STAPLER THEY NEEDED TO PULL A NEW ONE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
18-11-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | POWERED PLUS COMPACT 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
07-07-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE THE DEVICE DID NOT ACTIVATE NORMALLY. A PHENOMENON WAS ALSO OBSERVED WHERE THE DEVICE WAS FIRED OUTSIDE THE BODY BY THE ARTICULATION CONTROL BUTTON NOT THE FIRING TRIGGER. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
23-12-2025 | Malfunction | covidien surgical | 06-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A CYTOREDUCTIVE SURGERY, WHEN ATTEMPTED, THE INSTRUMENT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED WITH A NEW RELOAD TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
04-01-2026 | Malfunction | covidien | 06-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE FAILED TO FIRE WHEN DETACHING THE RECTUM. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE OF THE STAPLER. REPLACEMENT OF THE HANDLE AND RELOAD WAS REQUIRED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
28-09-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | WHITE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
04-01-2026 | Malfunction | covidien surgical | 06-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE FAILED TO FIRE WHEN DETACHING THE RECTUM. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE OF THE STAPLER. REPLACEMENT OF THE HANDLE AND RELOAD WAS REQUIRED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
24-12-2025 | Malfunction | covidien surgical | 06-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, WHILE PREPARING FOR LUNG PARENCHYMA DETACHMENT, THE RELOAD WAS UNABLE TO FIRE, AND THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON WAS UNABLE TO SQUEEZE THE HANDLE OF THE DEVICE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
26-12-2025 | Malfunction | covidien | 06-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY FOR DETACHMENT OF THE RIGHT COLON, THE RELOAD AND HANDLE WERE UNABLE TO FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE OF BOTH DEVICES. REPLACEMENT OF A NEW RELOAD WAS REQUIRED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
14-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST GREEN RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
24-12-2025 | Malfunction | covidien surgical | 06-01-2026 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN LAPAROSCOPIC ABDOMINAL WALL HERNIA REPAIR, SIGNIFICANT EFFORT WAS REQUIRED TO PULL THE TRIGGER OF THE ABSORBABLE TACKS FIXATOR, AND THE FIRED TACK WAS UNABLE TO PROPERLY SECURE THE MESH. THE DEVICE WAS REPLACED WITH ANOTHER BRAND TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
24-12-2025 | Malfunction | covidien surgical | 06-01-2026 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN LAPAROSCOPIC ABDOMINAL WALL HERNIA REPAIR, SIGNIFICANT EFFORT WAS REQUIRED TO PULL THE TRIGGER OF THE ABSORBABLE TACKS FIXATOR, AND THE FIRED TACK WAS UNABLE TO PROPERLY SECURE THE MESH. THE DEVICE WAS REPLACED WITH ANOTHER BRAND TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
04-01-2026 | Malfunction | covidien | 06-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESECTION OF PANCREATIC BODY AND TAIL LESIONS, WHEN FIRING, THE STAPLES WERE POORLY FORMED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
14-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST GREEN RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
11-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | GST GREEN RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-12-2025 | Malfunction | ethicon endo-surgery | 06-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER PARTIALLY FIRED STAPLES AND PARTIALLY CUT. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | covidien surgical | 06-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, AFTER FIRING, STAPLE LINE BLEEDING OCCURRED. SUTURING WAS PERFORMED TO RESOLVE THE STAPLE LINE BLEEDING.
09-12-2025 | Malfunction | covidien surgical | 06-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, AFTER FIRING, STAPLE LINE BLEEDING OCCURRED. SUTURING WAS PERFORMED TO RESOLVE THE STAPLE LINE BLEEDING.
09-12-2025 | Malfunction | covidien surgical | 06-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, AFTER FIRING, STAPLE LINE BLEEDING OCCURRED. SUTURING WAS PERFORMED TO RESOLVE THE STAPLE LINE BLEEDING.
09-12-2025 | Malfunction | covidien surgical | 06-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, AFTER FIRING, STAPLE LINE BLEEDING OCCURRED. SUTURING WAS PERFORMED TO RESOLVE THE STAPLE LINE BLEEDING.
09-12-2025 | Malfunction | covidien surgical | 06-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, AFTER FIRING, STAPLE LINE BLEEDING OCCURRED. SUTURING WAS PERFORMED TO RESOLVE THE STAPLE LINE BLEEDING.
09-12-2025 | Malfunction | covidien surgical | 06-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, AFTER FIRING, STAPLE LINE BLEEDING OCCURRED. SUTURING WAS PERFORMED TO RESOLVE THE STAPLE LINE BLEEDING.
06-01-2026 | Injury | intuitive surgical | 06-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SIGMOID COLECTOMY PROCEDURE, A STAPLE LINE CREATED AT THE DISTAL PORTION OF THE COLON WITH A SUREFORM 30 CURVED-TIP STAPLER AND SUREFORM 30 BLUE RELOAD DEHISCED. SPECIFIC PROCEDURAL DETAILS WERE NOT PROVIDED. THE DEHISCENCE WAS OBSERVED AT THE SITE WHERE THE SECOND BLUE RELOAD HAD BEEN FIRED. HOW THE ISSUE WAS MANAGED IS UNKNOWN. THE REMAINDER OF THE PROCEDURE WAS COMPLETED ROBOTICALLY WITHOUT FURTHER COMPLICATIONS.
24-12-2025 | Malfunction | covidien | 06-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY, WHEN DETACHING THE INTERLOBULAR FISSURE, THE HANDLE PRODUCED A CLICKING SOUND AND DID NOT RESPOND AFTER THE HANDLE WAS PRESSED. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
24-12-2025 | Malfunction | covidien | 06-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY, WHEN DETACHING THE INTERLOBULAR FISSURE, THE HANDLE PRODUCED A CLICKING SOUND AND DID NOT RESPOND AFTER THE HANDLE WAS PRESSED. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
24-12-2025 | Malfunction | covidien surgical | 06-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY, WHEN DETACHING THE INTERLOBULAR FISSURE, THE HANDLE PRODUCED A CLICKING SOUND AND DID NOT RESPOND AFTER THE HANDLE WAS PRESSED. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
04-12-2025 | Injury | intuitive surgical | 06-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY, THE SUREFORM 60 STAPLER WITH A SUREFORM 60 BLUE RELOAD HAD A PARTIAL FIRE, RESULTING IN TISSUE BUNCHING AND DRAGGING THAT CAUSED A HOLE IN THE PATIENT'S STOMACH. THE INSTRUMENT WAS INSPECTED BEFORE USE, AND NO UNUSUAL FINDINGS WERE NOTED. THE SURGEON SELECTED THIS TYPE OF RELOAD BECAUSE THE STAPLER WAS BEING USED IN THE MIDPORTION OF THE SLEEVE, WHERE THE TISSUE WAS NOT VERY THICK. THE ISSUE OCCURRED DURING THE 3RD OR 4TH STAPLER FIRE. THERE WAS NO EXPOSED BLADE OR CLAMPING ISSUES OBSERVED. NO BUTTRESS MATERIAL WAS USED. ADDITIONALLY, THE STAPLER INSTRUMENT WAS NOT FIRED OVER AN EXISTING STAPLE LINE. THE TARGET TISSUE WAS NOT CALCIFIED, HAD NOT BEEN EXPOSED TO CHEMOTHERAPY OR RADIATION, AND THERE WAS NO TENSION ON THE TISSUE. MALFORMED STAPLES RESEMBLING A "C" AND "X" SHAPE WERE OBSERVED. FURTHERMORE, A SMALL AMOUNT OF BLEEDING OCCURRED. THE AFFECTED SITE WAS VERY CLOSE TO THE VISIGI BOUGIE. THE SURGEON CLOSED THE DEFECT WITH A DOUBLE-LAYER SUTURE AND PLACED A LEAK DRAIN, WHICH IS ROUTINE FOR THIS TYPE OF PROCEDURE. THE EVENT LED TO A PROCEDURAL DELAY OF OVER 30 MINUTES. AN UPPER GASTROINTESTINAL TEST WAS PERFORMED THE NEXT DAY WHICH SHOWED NO LEAK OR OBSTRUCTION. THE PATIENT REMAINED STABLE AND IS SCHEDULED FOR ROUTINE FOLLOW-UP AT THE CLINIC.
24-12-2025 | Malfunction | covidien surgical | 06-01-2026 | ENDO GIA | Misfire | A050502 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY WHILE PREPARING TO DETACH THE LUNG PARENCHYMA, AFTER PRESSING THE SAFETY CATCH OF THE MANUAL HANDLE, THE RELOAD WAS FIRED AND ADVANCED ONLY A SHORT DISTANCE, RESULTING IN POOR FORMATION OF MOST STAPLES. THE DEVICE WAS REPLACED WITH A NEW RELOAD, AND THE PROCEDURE PROCEEDED SMOOTHLY.
24-12-2025 | Malfunction | covidien surgical | 06-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY WHILE PREPARING TO DETACH THE LUNG PARENCHYMA, AFTER PRESSING THE SAFETY CATCH OF THE MANUAL HANDLE, THE RELOAD WAS FIRED AND ADVANCED ONLY A SHORT DISTANCE, RESULTING IN POOR FORMATION OF MOST STAPLES. THE DEVICE WAS REPLACED WITH A NEW RELOAD, AND THE PROCEDURE PROCEEDED SMOOTHLY.
05-01-2026 | Malfunction | covidien | 06-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC DISTAL PANCREATECTOMY, POOR B-SHAPE STAPLE FORMATION WAS OBSERVED WITH THE RELOAD WHILE DETACHING THE PANCREATIC BODY AND TAIL. ADDITIONAL OVERSEWED SUTURING WAS REQUIRED TO ADDRESS THE STAPLE FORMATION ISSUE.
11-12-2025 | Injury | intuitive surgical | 07-01-2026 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED SLEEVE GASTRECTOMY, TWO SUREFORM 60 BLUE RELOADS WERE FIRED AND RESULTED IN STOMACH BLEEDING. THE CUSTOMER REPORTED THAT THIS OCCURRED DUE TO THE STAPLES NOT FORMING THE EXPECTED B-FORMATION. THE PROCEDURE WAS REPORTEDLY COMPLETED AS PLANNED.
11-12-2025 | Injury | intuitive surgical | 07-01-2026 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED SLEEVE GASTRECTOMY, TWO SUREFORM 60 BLUE RELOADS WERE FIRED AND RESULTED IN STOMACH BLEEDING. THE CUSTOMER REPORTED THAT THIS OCCURRED DUE TO THE STAPLES NOT FORMING THE EXPECTED B-FORMATION. THE PROCEDURE WAS REPORTEDLY COMPLETED AS PLANNED.
28-11-2025 | Malfunction | covidien surgical | 07-01-2026 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT REOPEN AFTER FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON 3000 60MM STANDARD | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE JAWS DID NOT OPEN, AND THE KNIFE WOULD NOT RETURN, SO THE PANEL WAS OPENED, AND THE KNIFE WAS RETURNED WITH THE KNIFE REVERSE SWITCH. IT WAS THE 2ND FIRING OF FUNCTIONAL END TO END ANASTOMOSIS, AND SINCE IT HAD BEEN SUTURED, ONLY THE MAIN BODY WAS REPLACED WITH A NEW ONE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON 3000 60MM STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE JAWS DID NOT OPEN, AND THE KNIFE WOULD NOT RETURN, SO THE PANEL WAS OPENED, AND THE KNIFE WAS RETURNED WITH THE KNIFE REVERSE SWITCH. IT WAS THE 2ND FIRING OF FUNCTIONAL END TO END ANASTOMOSIS, AND SINCE IT HAD BEEN SUTURED, ONLY THE MAIN BODY WAS REPLACED WITH A NEW ONE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON 3000 60MM STANDARD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE JAWS DID NOT OPEN, AND THE KNIFE WOULD NOT RETURN, SO THE PANEL WAS OPENED, AND THE KNIFE WAS RETURNED WITH THE KNIFE REVERSE SWITCH. IT WAS THE 2ND FIRING OF FUNCTIONAL END TO END ANASTOMOSIS, AND SINCE IT HAD BEEN SUTURED, ONLY THE MAIN BODY WAS REPLACED WITH A NEW ONE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
23-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST BLACK RELOAD 45MM 6 ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE CARTRIDGE DID NOT FIT INTO THE BODY SMOOTHLY AND SEEMS LIKE THE CARTRIDGE HAD BEEN DESTROYED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
23-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST BLACK RELOAD 45MM 6 ROW | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE CARTRIDGE DID NOT FIT INTO THE BODY SMOOTHLY AND SEEMS LIKE THE CARTRIDGE HAD BEEN DESTROYED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | covidien | 07-01-2026 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RADICAL LUNG CANCER SURGERY, WHILE BEING APPLIED TO SEVERED LUNG TISSUE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
09-12-2025 | Malfunction | covidien | 07-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RADICAL LUNG CANCER SURGERY, WHILE BEING APPLIED TO SEVERED LUNG TISSUE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
09-12-2025 | Malfunction | covidien | 07-01-2026 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RADICAL LUNG CANCER SURGERY, WHILE BEING APPLIED TO SEVERED LUNG TISSUE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
09-12-2025 | Malfunction | covidien | 07-01-2026 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RADICAL LUNG CANCER SURGERY, WHILE BEING APPLIED TO SEVERED LUNG TISSUE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED WITH A NEW ONE TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
15-12-2025 | Malfunction | covidien surgical | 07-01-2026 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, AFTER FIRING, WHEN TRYING TO RETURN THE RELOAD TO A STRAIGHT POSITION FROM BEING ARTICULATED TO REMOVE IT THROUGH THE TROCAR, THE RELOAD WOULD NOT RETURN TO 0°. ULTIMATELY, THE RETRACTION TOOL HAD TO BE USED TO DISARTICULATE THE RELOAD. SUBSEQUENTLY, THE SAME SHAFT AND HANDLE WERE USED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
10-12-2025 | Malfunction | covidien surgical | 07-01-2026 | ABSORBATACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN INGUINAL HERNIA REPAIR, THE DEVICE BROKE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON 3000 60MM STANDARD | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HAND ASSISTED NEPHRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE BLADE DID NOT WANT TO REVERSE, NOT EVEN WITH MANUAL OVERRIDE. AFTER FEW ATTEMPTS, THE OVERRIDE WORKED AND OPENED ANOTHER STAPLER TO COMPLETE THE PROCEDURE SUCCESSFULLY WITH A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
22-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON 3000 60MM STANDARD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HAND ASSISTED NEPHRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE BLADE DID NOT WANT TO REVERSE, NOT EVEN WITH MANUAL OVERRIDE. AFTER FEW ATTEMPTS, THE OVERRIDE WORKED AND OPENED ANOTHER STAPLER TO COMPLETE THE PROCEDURE SUCCESSFULLY WITH A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
22-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON 3000 60MM STANDARD | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HAND ASSISTED NEPHRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE BLADE DID NOT WANT TO REVERSE, NOT EVEN WITH MANUAL OVERRIDE. AFTER FEW ATTEMPTS, THE OVERRIDE WORKED AND OPENED ANOTHER STAPLER TO COMPLETE THE PROCEDURE SUCCESSFULLY WITH A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-12-2025 | Malfunction | covidien surgical | 07-01-2026 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, AFTER FIRING, WHEN TRYING TO RETURN THE RELOAD TO A STRAIGHT POSITION FROM BEING ARTICULATED TO REMOVE IT THROUGH THE TROCAR, THE RELOAD WOULD NOT RETURN TO 0°. ULTIMATELY, THE RETRACTION TOOL HAD TO BE USED TO DISARTICULATE THE RELOAD. SUBSEQUENTLY, THE SAME SHAFT AND HANDLE WERE USED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
12-11-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE CARTRIDGE FELL OFF. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
15-12-2025 | Malfunction | covidien | 07-01-2026 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, AFTER FIRING, WHEN TRYING TO RETURN THE RELOAD TO A STRAIGHT POSITION FROM BEING ARTICULATED TO REMOVE IT THROUGH THE TROCAR, THE RELOAD WOULD NOT RETURN TO 0°. ULTIMATELY, THE RETRACTION TOOL HAD TO BE USED TO DISARTICULATE THE RELOAD. SUBSEQUENTLY, THE SAME SHAFT AND HANDLE WERE USED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
07-11-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
23-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST WHITE RELOAD, 45MM, 6 ROW | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN LAPAROSCOPIC UROLOGICAL PROCEDURE, THE CARTRIDGE WAS USED AT THE 1ST FIRING, AND THE CARTRIDGE WOULD NOT FIT PROPERLY WHEN TRYING TO LOAD THE 2ND FIRING. AFTER TRYING SEVERAL TIMES AND STILL NOT BEING ABLE TO LOAD, IT WAS FOUND THAT THE CARTRIDGE PAN (WHICH HAD ALREADY BEEN SENT AS A DEFECT) WAS STILL LEFT TO THE MAIN BODY WHEN CHECKED THE MAIN BODY. THE PART REMAINING IN THE MAIN BODY WAS REMOVED WITH FORCEPS AND LOADED. THE DOCTOR WANTED TO KNOW WHY THIS WAS POSSIBLE. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
30-11-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RECTUM PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-11-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE GASTRECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-11-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GREEN RELOAD FOR ECHELON 60 | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | nan
12-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMICOLECTOMY, THE CDH29P DID NOT FIRE. A NEW STAPLER PACKAGING WAS OPENED, THE BATTERY OF THE ORIGINAL STAPLER WAS REPLACED AND THEN THE STAPLER DID FIRE. WHEN INSERTING THE ORIGINAL BATTERY, IT WAS ALSO MORE DIFFICULT THAN USUAL. IT SEEMED LIKE THE BATTERY DIDN'T FIT WELL. THE PROCEDURE WAS COMPLETED SUCCESSFULLY AND NO PATIENT CONSEQUENCES WERE REPORTED.
12-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMICOLECTOMY, THE CDH29P DID NOT FIRE. A NEW STAPLER PACKAGING WAS OPENED, THE BATTERY OF THE ORIGINAL STAPLER WAS REPLACED AND THEN THE STAPLER DID FIRE. WHEN INSERTING THE ORIGINAL BATTERY, IT WAS ALSO MORE DIFFICULT THAN USUAL. IT SEEMED LIKE THE BATTERY DIDN'T FIT WELL. THE PROCEDURE WAS COMPLETED SUCCESSFULLY AND NO PATIENT CONSEQUENCES WERE REPORTED.
31-10-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON FLEX | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: THE STAPLER DID NOT COMPLETE ITS FULL CYCLE ON ITS THIRD FIRING. THE BLADE DID NOT RETRACT ONCE THE STAPLES ENGAGED. A MANUAL OVERRIDE WAS PERFORMED, SUCCESSFULLY RELEASING THE DEVICE. A NEW STAPLER WAS USED TO FINISH THE CASE.
11-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLON SURGERY, A * CIRCULAR STAPLER FAILURE WAS REPORTED; A LEAK WAS OBSERVED, AND A LOOSE STAPLE WAS FOUND. A THIRD-PARTY STAPLER SHOULD BE USED TO COMPLETE THE SURGERY. THE PROCEDURE WAS DELAYED BY 30 MINUTES AND NO PATIENT CONSEQUENCES WERE REPORTED.
11-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON SURGERY, A * CIRCULAR STAPLER FAILURE WAS REPORTED; A LEAK WAS OBSERVED, AND A LOOSE STAPLE WAS FOUND. A THIRD-PARTY STAPLER SHOULD BE USED TO COMPLETE THE SURGERY. THE PROCEDURE WAS DELAYED BY 30 MINUTES AND NO PATIENT CONSEQUENCES WERE REPORTED.
11-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Appropriate Term/Code Not Available |  | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLON SURGERY, A * CIRCULAR STAPLER FAILURE WAS REPORTED; A LEAK WAS OBSERVED, AND A LOOSE STAPLE WAS FOUND. A THIRD-PARTY STAPLER SHOULD BE USED TO COMPLETE THE SURGERY. THE PROCEDURE WAS DELAYED BY 30 MINUTES AND NO PATIENT CONSEQUENCES WERE REPORTED.
11-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON SURGERY, A * CIRCULAR STAPLER FAILURE WAS REPORTED; A LEAK WAS OBSERVED, AND A LOOSE STAPLE WAS FOUND. A THIRD-PARTY STAPLER SHOULD BE USED TO COMPLETE THE SURGERY. THE PROCEDURE WAS DELAYED BY 30 MINUTES AND NO PATIENT CONSEQUENCES WERE REPORTED.
02-10-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ETHICON CONTOUR CURVED CUTTER TRIGGER JAMMED. JAWS WOULD NOT CLOSE. NO PATIENT HARM WAS REPORTED.
02-10-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: ETHICON CONTOUR CURVED CUTTER TRIGGER JAMMED. JAWS WOULD NOT CLOSE. NO PATIENT HARM WAS REPORTED.
14-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-12-2025 | Malfunction | covidien surgical | 07-01-2026 | ABSORBATACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HERNIORRHAPHY, THE TACKER BROKE. THE TACKER WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
10-12-2025 | Malfunction | covidien | 07-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT LUNG S2 SEGMENT RESECTION, THE DEVICE COULD NOT BE SQUEEZED ON THE FIRST FIRING ATTEMPT. BOTH THE CARTRIDGE AND THE HANDLE WERE REPLACED. THERE WAS NO PATIENT INJURY.
10-12-2025 | Malfunction | covidien surgical | 07-01-2026 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT LUNG S2 SEGMENT RESECTION, THE DEVICE COULD NOT BE SQUEEZED ON THE FIRST FIRING ATTEMPT. BOTH THE CARTRIDGE AND THE HANDLE WERE REPLACED. THERE WAS NO PATIENT INJURY.
10-12-2025 | Malfunction | surgical | 07-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OBESITY PROCEDURE, THE SIX BLUE CARTRIDGES USED TOGETHER WITH A POWERED STAPLER DID NOT ADVANCE DURING FIRING EVEN THOUGH IT INITIALLY INDICATED LEVEL 1-2 AND SHOWED AN APPROPRIATE LEVEL. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2025 | Malfunction | surgical | 07-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OBESITY PROCEDURE, THE SIX BLUE CARTRIDGES USED TOGETHER WITH A POWERED STAPLER DID NOT ADVANCE DURING FIRING EVEN THOUGH IT INITIALLY INDICATED LEVEL 1-2 AND SHOWED AN APPROPRIATE LEVEL. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2025 | Malfunction | surgical | 07-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OBESITY PROCEDURE, THE SIX BLUE CARTRIDGES USED TOGETHER WITH A POWERED STAPLER DID NOT ADVANCE DURING FIRING EVEN THOUGH IT INITIALLY INDICATED LEVEL 1-2 AND SHOWED AN APPROPRIATE LEVEL. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-10-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT HEMICOLECTOMY (OPEN SURGERY), CANNOT BE FIRED. NO PATIENT CONSEQUENCES WERE REPORTED.
10-12-2025 | Malfunction | surgical | 07-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OBESITY PROCEDURE, THE SIX BLUE CARTRIDGES USED TOGETHER WITH A POWERED STAPLER DID NOT, ADVANCE DURING FIRING EVEN THOUGH IT INITIALLY. INDICATED LEVEL 1-2 AND SHOWED AN APPROPRIATE LEVEL. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2025 | Malfunction | surgical | 07-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OBESITY PROCEDURE, THE SIX BLUE CARTRIDGES USED TOGETHER WITH A POWERED STAPLER DID NOT ADVANCE DURING FIRING EVEN THOUGH IT INITIALLY INDICATED LEVEL 1-2 AND SHOWED AN APPROPRIATE LEVEL. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2025 | Injury | intuitive surgical | 07-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT, DURING A DA VINCI-ASSISTED TRANSTHORACIC ESOPHAGECTOMY WITH CHEST ANASTOMOSIS PROCEDURE, TISSUE SLIPPED OUT OF THE JAWS OF THE SUREFORM 60 STAPLER INSTRUMENT EQUIPPED WITH A SUREFORM 60 GREEN RELOAD DURING THE FOURTH FIRING IN THE ABDOMINAL REGION, RESULTING IN AN INCOMPLETE STAPLE LINE. THE STAPLER INSTRUMENT HAD BEEN THOROUGHLY INSPECTED PRIOR TO USE AND SHOWED NO SIGNS OF DAMAGE OR ABNORMALITIES. SELECTION OF THE GREEN STAPLER RELOAD WAS APPROPRIATE FOR THE TISSUE BEING STAPLED. THE INCIDENT OCCURRED AFTER THE STAPLER RELOAD WAS ENGAGED AND THE TISSUE WAS CLAMPED. DURING THE FIRING PROCESS, THE TISSUE WAS PUSHED FORWARD AND SLIPPED FROM THE JAWS, RESULTING IN UNSPECIFIED FRAGMENTS FALLING INSIDE THE PATIENT. ALL FRAGMENTS WERE SUCCESSFULLY RETRIEVED. IT IS UNCLEAR IF THE FRAGMENTS WERE IN REFERENCE TO LOOSE STAPLES. NO TISSUE BUNCHING OCCURRED, NO ERROR MESSAGES WERE GENERATED, AND THERE WERE NO ISSUES WITH CLAMPING; NOR WAS THE STAPLER INSTRUMENT FIRED OVER AN EXISTING STAPLE LINE. THERE WAS NO BLEEDING OR NEED FOR ADDITIONAL TISSUE REMOVAL. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER RELOAD AND COMPLETING THE FIRING USING THE SAME STAPLER INSTRUMENT. THE PROCEDURE WAS SUCCESSFULLY COMPLETED ROBOTICALLY, AND THERE IS NO CONCERN REGARDING THIS EVENT CAUSING ANY LONG-TERM COMPLICATIONS FOR THE PATIENT.
28-11-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-08-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, CLAIMED DEVICE CANNOT BE FIRED. NO FURTHER INFORMATION. NO PATIENT CONSEQUENCES WERE REPORTED.
10-11-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2025 | Malfunction | johnson & johnson medical (suzhou | 07-01-2026 | FINAL ASSY 60MM STD DEVICE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-11-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GREEN RELOAD FOR ECHELON 60 | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon | 07-01-2026 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2025 AND ABSORBABLE STRAPS WERE USED. DURING A LAPAROSCOPIC HERNIA REPAIR, THE SURGEON OPENED A SECURE STRAP TO FIX HER REPAIR. THERE WAS A PROBLEM WITH THE MECHANISM AND DEPLOYMENT. THE STAPLE WAS COMING OUT OF THE DEVICE BUT DID NOT COMPLETE ITS FIXATION ACTION. THERE WAS NO SIGNIFICANT DELAY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED.
24-02-2026 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST GREEN RELOAD 45MM 6 ROW | Material Integrity Problem | A04 | Foreign Body In Patient | nan
16-12-2025 | Malfunction | covidien | 07-01-2026 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BARIATRIC GASTRIC BYPASS PROCEDURE, THE NURSE LOADED THE RELOAD ONTO THE STAPLER WITHOUT ANY ISSUES, AND THE DEVICE INDICATED THREE GREEN LANES AND CYCLED PROPERLY. WHEN THE SURGEON CLOSED THE JAWS OF THE STAPLER ON TISSUE AND ATTEMPTED TO FIRE, THE DEVICE MADE A NORMAL FIRING SOUND FOR ONE SECOND BEFORE STOPPING, AND A YELLOW TRIANGLE WITH AN EXCLAMATION MARK APPEARED ON THE FAR RIGHT SIDE OF THE STAPLER SCREEN. THE ONLY AVAILABLE ACTION WAS TO OPEN THE JAWS, AND NO STAPLES OR CUTTING OCCURRED TO THE TISSUE. THE RELOAD WAS REPLACED WITH A NEW RELOAD OF THE SAME TYPE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2025 | Malfunction | covidien | 07-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BARIATRIC GASTRIC BYPASS PROCEDURE, THE NURSE LOADED THE RELOAD ONTO THE STAPLER WITHOUT ANY ISSUES, AND THE DEVICE INDICATED THREE GREEN LANES AND CYCLED PROPERLY. WHEN THE SURGEON CLOSED THE JAWS OF THE STAPLER ON TISSUE AND ATTEMPTED TO FIRE, THE DEVICE MADE A NORMAL FIRING SOUND FOR ONE SECOND BEFORE STOPPING, AND A YELLOW TRIANGLE WITH AN EXCLAMATION MARK APPEARED ON THE FAR RIGHT SIDE OF THE STAPLER SCREEN. THE ONLY AVAILABLE ACTION WAS TO OPEN THE JAWS, AND NO STAPLES OR CUTTING OCCURRED TO THE TISSUE. THE RELOAD WAS REPLACED WITH A NEW RELOAD OF THE SAME TYPE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-01-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS NOT CLIPPING WELL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
24-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON 3000 60MM STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ECH45S WAS INCLUDED IN THE PACKAGE OF THE ECH60S, AND WAS REPLACED WITH A NEW ONE. THE ECH45S INSIDE WAS DISCARDED UNUSED. THE SALES REP RECEIVED PACKAGING LOT INFORMATION FROM A CUSTOMER OVER THE PHONE THAT ¿THE ISSUE OCCURRED IN LOT A9874K¿, BUT NO PHOTOGRAPHS WERE PROVIDED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
03-10-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
17-12-2025 | Malfunction | covidien | 07-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, THE DEVICE SHOWED A STAPLING FAILURE AFTER TESTING WITH METHYLENE BLUE. THE DEVICE WAS REPLACED WITH ANOTHER RELOAD FROM THE SAME LOT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-01-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST GOLD RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING WHEN BOTH STAPLE LINE AND CUT LINE WERE COMPLETED. HEMOSTASIS BY APPLICATION OF HEMOSTATIC MATERIAL WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST GOLD RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING WHEN BOTH STAPLE LINE AND CUT LINE WERE COMPLETED. HEMOSTASIS BY APPLICATION OF HEMOSTATIC MATERIAL WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST GOLD RELOAD, 45MM, 6-ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING WHEN BOTH STAPLE LINE AND CUT LINE WERE COMPLETED. HEMOSTASIS BY APPLICATION OF HEMOSTATIC MATERIAL WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST GOLD RELOAD, 45MM, 6-ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING WHEN BOTH STAPLE LINE AND CUT LINE WERE COMPLETED. HEMOSTASIS BY APPLICATION OF HEMOSTATIC MATERIAL WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING THE DEVICE WHEN BOTH STAPLE LINE AND CUT LINE WERE COMPLETE. HEMOSTASIS BY APPLICATION OF HEMOSTATIC MATERIAL WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING THE DEVICE WHEN BOTH STAPLE LINE AND CUT LINE WERE COMPLETE. HEMOSTASIS BY APPLICATION OF HEMOSTATIC MATERIAL WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST GREEN RELOAD, 45MM, 6-ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING THE DEVICE WHEN BOTH STAPLE LINE AND CUT LINE WERE COMPLETE. HEMOSTASIS BY APPLICATION OF HEMOSTATIC MATERIAL WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2025 | Malfunction | ethicon endo-surgery | 07-01-2026 | GST GREEN RELOAD, 45MM, 6-ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING THE DEVICE WHEN BOTH STAPLE LINE AND CUT LINE WERE COMPLETE. HEMOSTASIS BY APPLICATION OF HEMOSTATIC MATERIAL WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
26-12-2025 | Malfunction | johnson & johnson medical (suzhou | 07-01-2026 | CARTRIDGE FINAL ASSY 60MM BLUE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
18-12-2025 | Malfunction | covidien surgical | 07-01-2026 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHILE ON SET-UP, THE ADAPTER WAS AUTOMATICALLY CONNECTING AND DISCONNECTING TO THE POWER SHELL REPEATEDLY, AND THERE WAS A LOOSE CONNECTION FROM THE SHELL TO THE ADAPTER. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
10-12-2025 | Malfunction | surgical | 08-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING AN OBESITY PROCEDURE, THE SIX BLUE CARTRIDGES USED TOGETHER WITH A POWERED STAPLER DID NOT ADVANCE DURING FIRING EVEN THOUGH IT INITIALLY INDICATED LEVEL 1-2 AND SHOWED AN APPROPRIATE LEVEL. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-12-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
04-12-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | nan
04-12-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROXIMAL GASTRECTOMY, DURING CARDIA RESECTION, THE FIRING COULD NOT BE PERFORMED AT ALL, AND A STAPLE APPEARED NEAR THE BASE OF THE JAW OF THE CARTRIDGE. THE RELOAD WAS REPLACED, AND THE SAME ISSUE OCCURRED. A NEW ULTRA STAPLER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-12-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROXIMAL GASTRECTOMY, DURING CARDIA RESECTION, THE FIRING COULD NOT BE PERFORMED AT ALL, AND A STAPLE APPEARED NEAR THE BASE OF THE JAW OF THE CARTRIDGE. THE RELOAD WAS REPLACED, AND THE SAME ISSUE OCCURRED. A NEW ULTRA STAPLER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-12-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROXIMAL GASTRECTOMY, DURING CARDIA RESECTION, THE FIRING COULD NOT BE PERFORMED AT ALL, AND A STAPLE APPEARED NEAR THE BASE OF THE JAW OF THE CARTRIDGE. THE RELOAD WAS REPLACED, AND THE SAME ISSUE OCCURRED. A NEW ULTRA STAPLER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2025 | Malfunction | covidien surgical | 08-01-2026 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC GASTRIC BYPASS, THE POWERED STAPLER WAS ROTATING ON ITS OWN. THE STAFF DISASSEMBLED THE ENTIRE UNIT BUT WAS UNABLE TO PINPOINT WHICH COMPONENT WAS DEFECTIVE. THE HANDLE, ADAPTER AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2025 | Malfunction | covidien surgical | 08-01-2026 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC GASTRIC BYPASS, THE POWERED STAPLER WAS ROTATING ON ITS OWN. THE STAFF DISASSEMBLED THE ENTIRE UNIT BUT WAS UNABLE TO PINPOINT WHICH COMPONENT WAS DEFECTIVE. THE HANDLE, ADAPTER AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
20-10-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR COLORECTAL CANCER ON THE RECTUM, AFTER SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. THE ISSUE OCCURRED ON TWO RELOADS FROM THE SAME BATCH AND ON THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2025 | Malfunction | covidien surgical | 08-01-2026 | ENDO GIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC GASTRIC BYPASS, THE POWERED STAPLER WAS ROTATING ON ITS OWN. THE STAFF DISASSEMBLED THE ENTIRE UNIT BUT WAS UNABLE TO PINPOINT WHICH COMPONENT WAS DEFECTIVE. THE HANDLE, ADAPTER AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
20-10-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR COLORECTAL CANCER ON THE RECTUM, AFTER SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. THE ISSUE OCCURRED ON TWO RELOADS FROM THE SAME BATCH AND ON THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-10-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR COLORECTAL CANCER ON THE RECTUM, AFTER SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. THE ISSUE OCCURRED ON TWO RELOADS FROM THE SAME BATCH AND ON THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2025 | Malfunction | covidien surgical | 08-01-2026 | ENDO GIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | nan
16-12-2025 | Malfunction | covidien surgical | 08-01-2026 | ENDO GIA | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
16-12-2025 | Malfunction | covidien surgical | 08-01-2026 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | nan
16-12-2025 | Malfunction | covidien surgical | 08-01-2026 | SIGNIA | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
16-12-2025 | Malfunction | covidien surgical | 08-01-2026 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | nan
16-12-2025 | Malfunction | covidien surgical | 08-01-2026 | SIGNIA | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
08-01-2026 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Cut | A050702 | Foreign Body In Patient | Event Description: IT WAS REPORTED THAT DURING A PANCREATICODUODENECTOMY, AN UNEXPECTED SOUND WAS HEARD WHEN RESECTING THE GASTRIC BODY, AND THE RESECTION COULD NOT BE DONE PROPERLY. THE SAME THING HAPPENED WHEN TRIED TO FIRE AGAIN WITH ANOTHER CARTRIDGE. A RATTLING SOUND WAS HEARD AT THE 1ST FIRING, AND THE STAPLES WERE DEPLOYED, BUT THE TARGET TISSUE WAS NOT CUT. THE SAME EVENT OCCURRED AT THE 2ND SHOT. ALTHOUGH IT WAS UNCERTAIN WHETHER IT WAS AN ECHELON PART, METAL FRAGMENTS WERE ALSO FOUND, SO IT WAS DETERMINED TO BE A PROBLEM WITH THE MAIN BODY AND REPLACED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
08-01-2026 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON 3000 60MM STANDARD | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: IT WAS REPORTED THAT DURING A PANCREATICODUODENECTOMY, AN UNEXPECTED SOUND WAS HEARD WHEN RESECTING THE GASTRIC BODY, AND THE RESECTION COULD NOT BE DONE PROPERLY. THE SAME THING HAPPENED WHEN TRIED TO FIRE AGAIN WITH ANOTHER CARTRIDGE. A RATTLING SOUND WAS HEARD AT THE 1ST FIRING, AND THE STAPLES WERE DEPLOYED, BUT THE TARGET TISSUE WAS NOT CUT. THE SAME EVENT OCCURRED AT THE 2ND SHOT. ALTHOUGH IT WAS UNCERTAIN WHETHER IT WAS AN ECHELON PART, METAL FRAGMENTS WERE ALSO FOUND, SO IT WAS DETERMINED TO BE A PROBLEM WITH THE MAIN BODY AND REPLACED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
08-01-2026 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON 3000 60MM STANDARD | Noise, Audible | A0508 | Foreign Body In Patient | Event Description: IT WAS REPORTED THAT DURING A PANCREATICODUODENECTOMY, AN UNEXPECTED SOUND WAS HEARD WHEN RESECTING THE GASTRIC BODY, AND THE RESECTION COULD NOT BE DONE PROPERLY. THE SAME THING HAPPENED WHEN TRIED TO FIRE AGAIN WITH ANOTHER CARTRIDGE. A RATTLING SOUND WAS HEARD AT THE 1ST FIRING, AND THE STAPLES WERE DEPLOYED, BUT THE TARGET TISSUE WAS NOT CUT. THE SAME EVENT OCCURRED AT THE 2ND SHOT. ALTHOUGH IT WAS UNCERTAIN WHETHER IT WAS AN ECHELON PART, METAL FRAGMENTS WERE ALSO FOUND, SO IT WAS DETERMINED TO BE A PROBLEM WITH THE MAIN BODY AND REPLACED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
06-12-2025 | Malfunction | johnson & johnson medical (suzhou | 08-01-2026 | FINAL ASSY 60MM STD DEVICE | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER INSTALLED RELOAD CLOSED THE JAW, BUT COULD NOT OPEN THE JAW ANYMORE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
29-10-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE MOVED TO THE DISTAL END BUT DID NOT RETURN AUTOMATICALLY AFTER FIRING. THEY USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
29-10-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE MOVED TO THE DISTAL END BUT DID NOT RETURN AUTOMATICALLY AFTER FIRING. THEY USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
17-11-2025 | Malfunction | johnson & johnson medical (suzhou | 08-01-2026 | FINAL ASSY 60MM STD DEVICE | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER INSTALLED RELOAD CLOSED THE JAW, BUT COULD NOT OPEN THE JAW ANYMORE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
22-12-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS NOT POSSIBLE TO OPEN AFTER BENT. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-11-2025 | Malfunction | johnson & johnson medical (suzhou | 08-01-2026 | FINAL ASSY 60MM STD DEVICE | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER INSTALLED RELOAD CLOSED THE JAW, BUT COULD NOT OPEN THE JAW ANYMORE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
22-12-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | 75MM SELECTABLE NEW TLC | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE NTLC 75 STAPLER HAS BEEN BROKEN WHILE ASSEMBLING. NO PATIENT CONSEQUENCES WERE REPORTED.
06-01-2026 | Malfunction | covidien surgical | 08-01-2026 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR, SOME TACKS WERE HARD TO FIRE AND SOME TACKS DID NOT HOLD WELL WHILE FIRING THE TACKER. A NEW TACKER WAS OPENED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
06-01-2026 | Malfunction | covidien surgical | 08-01-2026 | ABSORBATACK | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR, SOME TACKS WERE HARD TO FIRE AND SOME TACKS DID NOT HOLD WELL WHILE FIRING THE TACKER. A NEW TACKER WAS OPENED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
09-12-2025 | Injury | intuitive surgical | 08-01-2026 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.
01-01-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON 3000 60MM LONG | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON 3000 60MM LONG | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, WHEN THEY FIRED THE CDH29, THE STAPLE LINE WAS UNSTABLE, SO THEY HAD TO RE-STAPLE. NO PATIENT CONSEQUENCES WERE REPORTED.
23-12-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON 3000 60MM STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
27-11-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ENDO ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RIGHT HEMICOLECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-10-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE BATTERY ON THE DEVICE SUDDENLY RAN OUT OF POWER; AND THEREFORE, COULD NOT BE USED. . THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-01-2026 | Malfunction | ethicon endo-surgery | 08-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT BE FIRED WHILE THE MOTOR WAS MOVING; AND THE KNIFE DID NOT RETURN. THERE WAS NO EFFECT ON THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
07-11-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ETHICON 4000 STAPLER STRAIGHT 60 STANDARD | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE MIDDLE BASE OF THE RELOAD STAYED WITH THE JAWS OF THE STAPLER WHILE UNLOADING AND GOT STUCK. THIS MADE US UNABLE TO RELOAD THE STAPLER AND HAVE TO PULL NEW RELOAD AND STAPLER. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-12-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-11-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON 3000 45MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOST BATTERY POWER AFTER 4TH FIRING. THE BATTERY WAS REMOVED AND REPLACED, HOME BUTTON PRESSED BUT STILL NO POWER. THE HANDLE WOULD NOT CLOSE OVER LUNG TISSUE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
05-11-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ECHELON 3000 45MM COMPACT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOST BATTERY POWER AFTER 4TH FIRING. THE BATTERY WAS REMOVED AND REPLACED, HOME BUTTON PRESSED BUT STILL NO POWER. THE HANDLE WOULD NOT CLOSE OVER LUNG TISSUE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
21-11-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | ETHICON 4000 STAPLER STRAIGHT 60 COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A WHIPPLE PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT ACTIVATE WHEN THE SURGEON ATTEMPTED TO FIRE IT. THE RELOAD WAS THEN REPLACED, AND THE STAPLER FUNCTIONED PROPERLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
24-11-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | GOLD RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-11-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THEY FIRED THE FIRST RELOAD AND IT FIRED FINE. THEN THEY WENT TO RELOAD THE DEVICE BUT SUBSEQUENT RELOADS WOULD NOT FIT. THE REP ALSO TRIED TO RELOAD OUTSIDE OF THE STERILE FIELD WITH A NEW RELOAD AND IT WAS VERY DIFFICULT. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
05-11-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN SURGERY, THE CARTRIDGE COULD NOT BE LOADED WHEN REPLACING IT. A REPLACEMENT DEVICE WAS TAKEN OUT AND THE SURGERY CONTINUED. THE PACKAGE WAS OPENED. THE DEVICE WAS NOT USED FOR THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NO FURTHER INFORMATION IS AVAILABLE.
05-11-2025 | Malfunction | ethicon endo-surgery | 08-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN SURGERY, THE CARTRIDGE COULD NOT BE LOADED WHEN REPLACING IT. A REPLACEMENT DEVICE WAS TAKEN OUT AND THE SURGERY CONTINUED. THE PACKAGE WAS OPENED. THE DEVICE WAS NOT USED FOR THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NO FURTHER INFORMATION IS AVAILABLE.
15-12-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING USE OF THE DEVICE IN A THORACIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. ADDITIONALLY, THE STAPLER WAS STUCK IN THE TISSUE. THE DEVICE WAS REPLACED WITH A NEW HANDLE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
15-12-2025 | Malfunction | covidien | 08-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING USE OF THE DEVICE IN A THORACIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. ADDITIONALLY, THE STAPLER WAS STUCK IN THE TISSUE. THE DEVICE WAS REPLACED WITH A NEW HANDLE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
11-11-2025 | Malfunction | covidien surgical | 08-01-2026 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT THE RELOAD FAILED DURING USE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD WAS PARTIALLY FIRED. THE CLAMPING MECHANISM WAS DEFORMED.
08-09-2025 | Malfunction | covidien surgical | 08-01-2026 | ENDO GIA | Retraction Problem | A0510 | Insufficient Information | nan
08-09-2025 | Malfunction | covidien surgical | 08-01-2026 | ENDO GIA | Noise, Audible | A0508 | Insufficient Information | nan
09-01-2026 | Injury | covidien surgical | 09-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC LAPAROSCOPIC SIGMOID PROCEDURE, THE CIRCULAR STAPLER WAS USED, AND AFTER THE SURGERY, A POST-OPERATIVE LEAK OCCURRED. THE LEAK WAS SUSPECTED TO BE A POSSIBLE PANCREATIC LEAK, WHICH WAS IDENTIFIED AS A POTENTIAL CAUSE FOR THE ANASTOMOSIS FAILURE, AS THE ANASTOMOSES FELL APART. THE PATIENT WAS TAKEN BACK TO SURGERY, AND A COLOSTOMY WAS PERFORMED.
09-01-2026 | Injury | covidien surgical | 09-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC LAPAROSCOPIC SIGMOID PROCEDURE, THE CIRCULAR STAPLER WAS USED, AND AFTER THE SURGERY, A POST-OPERATIVE LEAK OCCURRED. THE LEAK WAS SUSPECTED TO BE A POSSIBLE PANCREATIC LEAK, WHICH WAS IDENTIFIED AS A POTENTIAL CAUSE FOR THE ANASTOMOSIS FAILURE, AS THE ANASTOMOSES FELL APART. THE PATIENT WAS TAKEN BACK TO SURGERY, AND A COLOSTOMY WAS PERFORMED.
03-12-2025 | Malfunction | covidien | 09-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | nan
03-12-2025 | Malfunction | covidien | 09-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
20-11-2025 | Injury | covidien surgical | 09-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, EIGHT DAYS AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAKAGE. THE PATIENT WAS READMITTED TO THE HOSPITAL AND UNDERWENT AN ADDITIONAL SURGICAL PROCEDURE, AND THE SURGEON PLACED A STENT. REHOSPITALIZATION LASTED FOR TWO WEEKS.
20-11-2025 | Injury | covidien surgical | 09-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, EIGHT DAYS AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE PATIENT DEVELOPED AN ANASTOMOTIC LEAKAGE. THE PATIENT WAS READMITTED TO THE HOSPITAL AND UNDERWENT AN ADDITIONAL SURGICAL PROCEDURE, AND THE SURGEON PLACED A STENT. REHOSPITALIZATION LASTED FOR TWO WEEKS.
12-12-2025 | Malfunction | covidien surgical | 09-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY SEGMENTECTOMY, THE POWERED STAPLER STOPPED MID-FIRE, FAILING TO FULLY STAPLE THE TISSUE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2025 | Malfunction | covidien surgical | 09-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | nan
12-12-2025 | Malfunction | covidien | 09-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY SEGMENTECTOMY, THE POWERED STAPLER STOPPED MID-FIRE, FAILING TO FULLY STAPLE THE TISSUE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-12-2025 | Malfunction | covidien surgical | 09-01-2026 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE HANDLE FIRING SENSATION WAS STIFF, THERE WAS DISCOMFORT IN THE FIRING SOUND, AND THE FIRED TACK LACKED ADEQUATE FIXATION STRENGTH. AS A RESULT, THE DEVICE COULD NOT BE PROPERLY FIRED AND WAS REPLACED INTRAOPERATIVELY BY OPENING AND USING A NEW TACKER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
03-12-2025 | Malfunction | covidien surgical | 09-01-2026 | ABSORBATACK | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE HANDLE FIRING SENSATION WAS STIFF, THERE WAS DISCOMFORT IN THE FIRING SOUND, AND THE FIRED TACK LACKED ADEQUATE FIXATION STRENGTH. AS A RESULT, THE DEVICE COULD NOT BE PROPERLY FIRED AND WAS REPLACED INTRAOPERATIVELY BY OPENING AND USING A NEW TACKER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
03-12-2025 | Malfunction | covidien surgical | 09-01-2026 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE HANDLE FIRING SENSATION WAS STIFF, THERE WAS DISCOMFORT IN THE FIRING SOUND, AND THE FIRED TACK LACKED ADEQUATE FIXATION STRENGTH. AS A RESULT, THE DEVICE COULD NOT BE PROPERLY FIRED AND WAS REPLACED INTRAOPERATIVELY BY OPENING AND USING A NEW TACKER TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-12-2025 | Malfunction | covidien surgical | 09-01-2026 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | nan
18-12-2025 | Malfunction | covidien surgical | 09-01-2026 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | nan
18-11-2025 | Malfunction | covidien | 09-01-2026 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT INTRAOPERATIVELY, THE INSTRUMENT MADE A STRANGE NOISE, AND THE HANDLE CRACKED DURING FIRING. IT WAS ALSO OBSERVED THAT THE FIRST 5MM OF STAPLES DID NOT FIRE. NO INTERVENTION WAS DONE/REQUIRED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE A SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
18-11-2025 | Malfunction | covidien | 09-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT INTRAOPERATIVELY, THE INSTRUMENT MADE A STRANGE NOISE, AND THE HANDLE CRACKED DURING FIRING. IT WAS ALSO OBSERVED THAT THE FIRST 5MM OF STAPLES DID NOT FIRE. NO INTERVENTION WAS DONE/REQUIRED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE A SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | Unintended Ejection | A050302 | Hemorrhage/Blood Loss/Bleeding | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | Unintended Ejection | A050302 | Post Operative Wound Infection | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | Post Operative Wound Infection | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Post Operative Wound Infection | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | No Apparent Adverse Event | A25 | Post Operative Wound Infection | nan
26-12-2025 | Injury | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 60MM STANDARD | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
13-11-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
29-12-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE CUT BUT DID NOT STAPLE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE WITH A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
02-12-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, WHEN TRYING TO USE THE ECHELON 7 DURING PULMONARY ARTERY RESECTION, A LOCKED-OUT OCCURRED. AFTER REPLACING THE CARTRIDGE, A LOCKED-OUT OCCURRED AGAIN (WHICH WAS RELEASED BY THE MANUAL OVERRIDE LEVER). THE MAIN BODY WAS REPLACED, AND NEW CARTRIDGE WAS ALSO LOADED INTO THE 2ND DEVICE, THEN FIRED OUTSIDE THE BODY CAVITY TO CHECK THE FUNCTION, BUT THE THIRD AND FOURTH CARTRIDGES WERE ALSO LOCKED OUT. A DIFFERENT LOT CARTRIDGE WAS LOADED INTO THE 2ND DEVICE AND USED, IT COULD BE FIRED SUCCESSFULLY, AND THE SURGERY WAS COMPLETED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
14-12-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE FARTHER AFTER FIRED A LITTLE AND THE KNIFE MOVED TO THE DISTAL END BUT DID NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON DID NOT WORK. OPENED THE DEVICE MANUALLY WITH BIG FORCE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE FARTHER AFTER FIRED A LITTLE AND THE KNIFE MOVED TO THE DISTAL END BUT DID NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON DID NOT WORK. OPENED THE DEVICE MANUALLY WITH BIG FORCE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE FARTHER AFTER FIRED A LITTLE AND THE KNIFE MOVED TO THE DISTAL END BUT DID NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON DID NOT WORK. OPENED THE DEVICE MANUALLY WITH BIG FORCE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-12-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE CARTRIDGE WAS DAMAGED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
29-12-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DISTAL GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
24-11-2025 | Malfunction | ethicon | 09-01-2026 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2025 AND ABSORBABLE STRAPS WERE USED. THE ABSORBABLE STRAP DID NOT STAPLE DURING SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS REQUESTED.
02-02-2026 | Malfunction | covidien | 09-01-2026 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | nan
26-11-2025 | Injury | covidien surgical | 09-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: ACCORDING TO THE REPORTER, POSTOPERATIVELY, THE STUMP OF THE DUODENUM WAS LEAKING, AND THERE WAS POSTOPERATIVE INFECTION. A MEDICATION WAS GIVEN TO FIGHT INFECTION.
26-11-2025 | Injury | covidien surgical | 09-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, POSTOPERATIVELY, THE STUMP OF THE DUODENUM WAS LEAKING, AND THERE WAS POSTOPERATIVE INFECTION. A MEDICATION WAS GIVEN TO FIGHT INFECTION.
29-10-2025 | Malfunction | covidien surgical | 09-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE INSTRUMENT MADE STRANGE NOISE. THE DISTAL STAPLE LINE WAS INCOMPLETE. THE STAPLER CUT BUT DID NOT DEPLOY THE STAPLES AND SOME OF THE DEPLOYED STAPLES DID NOT FORM PROPERLY. THE STAPLE LINE BLED. THERE WAS BLOOD LOSS OF 500CC OR MORE. THE STAPLE LINE WAS RESECTED AND RE-STAPLED TO FIX THE ISSUE. THE ISSUE RESULTED TO THE PROCEDURE CHANGING TO A LUNG LOBECTOMY INSTEAD OF SEGMENT RESECTION. THE SURGICAL TIME WAS EXTENDED FOR 20 MINUTES. MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT STAPLE PUSHERS WERE VISIBLE AT THE 2CM CUT LINE.
29-10-2025 | Malfunction | covidien surgical | 09-01-2026 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE INSTRUMENT MADE STRANGE NOISE. THE DISTAL STAPLE LINE WAS INCOMPLETE. THE STAPLER CUT BUT DID NOT DEPLOY THE STAPLES AND SOME OF THE DEPLOYED STAPLES DID NOT FORM PROPERLY. THE STAPLE LINE BLED. THERE WAS BLOOD LOSS OF 500CC OR MORE. THE STAPLE LINE WAS RESECTED AND RE-STAPLED TO FIX THE ISSUE. THE ISSUE RESULTED TO THE PROCEDURE CHANGING TO A LUNG LOBECTOMY INSTEAD OF SEGMENT RESECTION. THE SURGICAL TIME WAS EXTENDED FOR 20 MINUTES. MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT STAPLE PUSHERS WERE VISIBLE AT THE 2CM CUT LINE.
02-01-2025 | Malfunction | covidien | 09-01-2026 | PREMIUM PLUS CEEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE STAPLER DID NOT FIRE. REPEATED ATTEMPTS TO FIRE HAVE FAILED. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2025 | Malfunction | covidien | 09-01-2026 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN A MALFUNCTION OCCURRED IN THE RELEVANT HANDLE, THE SECOND CARTRIDGE STOPPED FIRING AFTER AN ABNORMAL SOUND. HOWEVER, A NEW HANDLE AND A THIRD CARTRIDGE WERE USED TO COMPLETE THE PROCEDURE, AND THE SURGERY WAS ABLE TO CONTINUE WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
22-12-2025 | Malfunction | covidien | 09-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN A MALFUNCTION OCCURRED IN THE RELEVANT HANDLE, THE SECOND CARTRIDGE STOPPED FIRING AFTER AN ABNORMAL SOUND. HOWEVER, A NEW HANDLE AND A THIRD CARTRIDGE WERE USED TO COMPLETE THE PROCEDURE, AND THE SURGERY WAS ABLE TO CONTINUE WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
22-12-2025 | Malfunction | covidien surgical | 09-01-2026 | UNKNOWN EGIA SU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN A MALFUNCTION OCCURRED IN THE RELEVANT HANDLE, THE SECOND CARTRIDGE STOPPED FIRING AFTER AN ABNORMAL SOUND. HOWEVER, A NEW HANDLE AND A THIRD CARTRIDGE WERE USED TO COMPLETE THE PROCEDURE, AND THE SURGERY WAS ABLE TO CONTINUE WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
03-11-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | ECHELON 3000 45MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE AND HAD NO MOTOR SOUND ON THE 1TH FIRING. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL WOULD NOT FIRE AND HAD NO MOTOR SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
29-12-2025 | Malfunction | covidien surgical | 09-01-2026 | EEA | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOID PROCEDURE, THE HANDLE WAS FULLY SQUEEZED UNTIL IT CAME IN CONTACT WITH THE INSTRUMENT BODY, BUT THE INSTRUMENT DID NOT FIRE. THE DEVICE WAS REPLACED WITH ANOTHER MANUFACTURER STAPLER.
04-11-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT FIRE AFTER CLOSING; AND IT WAS STUCK. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-09-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
23-11-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE BECAME LOCKED, FAILED TO OPEN AND NO STAPLES WERE DEPLOYED. THEY HAD TO USE MANUAL RELEASE TO OPEN THE DEVICE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE BECAME LOCKED, FAILED TO OPEN AND NO STAPLES WERE DEPLOYED. THEY HAD TO USE MANUAL RELEASE TO OPEN THE DEVICE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | PWRD ECH FLEX 45MM 28CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT FIRE AFTER REINSERTING THE BATTERY. THE DOCTOR REPLACED THE BATTERY WITH A NEW ONE, AND THE DEVICE FUNCTIONED PROPERLY THEREAFTER. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
03-12-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER JUST LOCK OUT INSTANTLY ON THE ARTERIAL BRANCH. SURGEON ENGAGED THE MANUAL OVERRIDE; AND IT SAFELY CAME OFF THE VESSEL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
03-12-2025 | Malfunction | ethicon endo-surgery | 09-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER JUST LOCK OUT INSTANTLY ON THE ARTERIAL BRANCH. SURGEON ENGAGED THE MANUAL OVERRIDE; AND IT SAFELY CAME OFF THE VESSEL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
12-12-2025 | Malfunction | covidien | 09-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
17-11-2025 | Malfunction | intuitive surgical | 09-01-2026 | SUREFORM | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SURGICAL PROCEDURE, THE SYSTEM WOULD NOT RECOGNIZE THE COLOR OF THE SUREFORM 30 WHITE RELOAD. A SECOND RELOAD HAD A BROKEN TAB. 2 DIFFERENT LOT NUMBERS. NO FRAGMENT FELL INTO THE PATIENT. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY.
17-11-2025 | Malfunction | intuitive surgical | 09-01-2026 | SUREFORM | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SURGICAL PROCEDURE, THE SYSTEM WOULD NOT RECOGNIZE THE COLOR OF THE SUREFORM 30 WHITE RELOAD. A SECOND RELOAD HAD A BROKEN TAB. 2 DIFFERENT LOT NUMBERS. NO FRAGMENT FELL INTO THE PATIENT. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY.
17-11-2025 | Malfunction | intuitive surgical | 09-01-2026 | SUREFORM | Material Fragmentation | A040103 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SURGICAL PROCEDURE, THE SYSTEM WOULD NOT RECOGNIZE THE COLOR OF THE SUREFORM 30 WHITE RELOAD. A SECOND RELOAD HAD A BROKEN TAB. 2 DIFFERENT LOT NUMBERS. NO FRAGMENT FELL INTO THE PATIENT. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY.
01-11-2025 | Malfunction | covidien | 09-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | nan
25-11-2025 | Malfunction | covidien surgical | 09-01-2026 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | nan
25-11-2025 | Malfunction | covidien surgical | 09-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | nan
16-12-2025 | Malfunction | covidien surgical | 09-01-2026 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SACROCOLPOPEXY, WHEN ATTACHING AND PRESSING THE MESH, SEVERAL TACKS CAME OUT AT ONCE, AND OTHERS DID NOT COME OUT. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2025 | Malfunction | covidien surgical | 09-01-2026 | ABSORBATACK | Display or Visual Feedback Problem | A0902 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SACROCOLPOPEXY, WHEN ATTACHING AND PRESSING THE MESH, SEVERAL TACKS CAME OUT AT ONCE, AND OTHERS DID NOT COME OUT. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Bacterial Infection | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Septic Shock | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Organ Dehiscence | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Bacterial Infection | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Septic Shock | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
15-12-2025 | Injury | covidien | 12-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Organ Dehiscence | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD UNDERGONE A ROBOTIC RIGHT COLECTOMY AND WAS DOING WELL. THE SURGEON AND THE SALES REPRESENTATIVE DID NOT OBSERVE ANY COMPLICATIONS. THE ONLY ATYPICAL OBSERVATION WAS THE USE OF THE PURPLE RELOAD IN THE SMALL INTESTINE. AT THE TIME, THE USE OF A BEIGE LOAD WAS SUGGESTED, BUT THE SURGEON DID NOT ACCEPT THE SUGGESTION. THREE DAYS AFTER THE SURGERY, THE PATIENT PRESENTED WITH SEPSIS, BACTERIAL INFECTION, AND SEPTIC SHOCK AND HAD A SEVERE DECREASE IN VITAL SIGNS. INITIALLY, THE PATIENT UNDERWENT MINIMALLY INVASIVE ROBOTIC SURGERY AND PROGRESSED INTO AN EMERGENCY EXPLORATORY LAPAROTOMY. THE SURGEON FOUND A RUPTURE OF THE STAPLE LINE IN THE ILEAL STUMP; THE ILEOCOLIC ANASTOMOSIS WAS INTACT. THE STAPLE LINE STUMP HAD A DEHISCENCE AND LEAKAGE. THE ANASTOMOSES WERE REDONE WITH A MANUAL STAPLER AND PURPLE LOADS, AND AN ILEOSTOMY AND DRAIN WERE ADDED. THE PATIENT REMAINED HOSPITALIZED IN SERIOUS CONDITION, WITH SEPSIS AND ON PROLONGED INTUBATION IN THE INTENSIVE CARE UNIT. THE ILEOSTOMY WAS ONLY TEMPORARY, AND THE PATIENT IS PROGRESSING WELL.
27-12-2025 | Malfunction | covidien surgical | 12-01-2026 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, THE THREE CONSECUTIVE ISSUES WITH A POWERED STAPLER AND STANDARD ADAPTER OCCURRED. IN TWO ATTEMPTS, THE STAPLES WERE DETACHED FROM THE TWO RELOADS, AND IN A THIRD EVENT, THE STAPLES COULD NOT BE INSERTED INTO THE ADAPTER PRIOR TO FIRING. IN EACH CASE, THE AFFECTED RELOAD WAS REPLACED WITH A NEW TRI- STAPLE RELOAD, WHICH WAS SUCCESSFULLY INSERTED AND USED WITHOUT FURTHER ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
27-12-2025 | Malfunction | covidien surgical | 12-01-2026 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, THE THREE CONSECUTIVE ISSUES WITH A POWERED STAPLER AND STANDARD ADAPTER OCCURRED. IN TWO ATTEMPTS, THE STAPLES WERE DETACHED FROM THE TWO RELOADS, AND IN A THIRD EVENT, THE STAPLES COULD NOT BE INSERTED INTO THE ADAPTER PRIOR TO FIRING. IN EACH CASE, THE AFFECTED RELOAD WAS REPLACED WITH A NEW TRI- STAPLE RELOAD, WHICH WAS SUCCESSFULLY INSERTED AND USED WITHOUT FURTHER ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
10-12-2025 | Malfunction | intuitive surgical | 12-01-2026 | SUREFORM | Device Dislodged or Dislocated | A051201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, THE SUREFORM 45 GREEN RELOAD GOT OUT OF THE STAPLER WHILE STAPLING. NO FRAGMENT FELL INTO THE PATIENT. THE USER COMPLETED THE PROCEDURE AND NO KNOWN IMPACT OR PATIENT CONSEQUENCE WAS REPORTED.
18-12-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DISTAL GASTRECTOMY DUODENAL RESECTION, THE FIRST RELOAD PARTIALLY FIRED STOPS MID WAYS AND THE KNIFE BLADE RETRACTED. ON THE SECOND RELOAD THE SHEET WAS DETACHED. A NEW RELOAD WAS USED TO RELOAD THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LIVING TRANSPLANT PROCEDURE¿SPECIFICALLY DURING THE RESECTION OF THE DONOR'S RENAL ARTERY¿THE SUTURE LINE BLED, REQUIRING THE APPLICATION OF ADDITIONAL REFILLS AND CLIPS. PROLONGED SURGICAL TIME WAS NOTED BUT NOT MORE THAN 30 MINUTES. ADDITIONALLY, THE HANDLE MADE A NOISE WHEN ROTATED. TO ADDRESS THE SITUATION, THE SURGICAL TEAM PERFORMED RESECTION AND RE-STAPLING.
17-12-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LIVING TRANSPLANT PROCEDURE¿SPECIFICALLY DURING THE RESECTION OF THE DONOR'S RENAL ARTERY¿THE SUTURE LINE BLED, REQUIRING THE APPLICATION OF ADDITIONAL REFILLS AND CLIPS. PROLONGED SURGICAL TIME WAS NOTED BUT NOT MORE THAN 30 MINUTES. ADDITIONALLY, THE HANDLE MADE A NOISE WHEN ROTATED. TO ADDRESS THE SITUATION, THE SURGICAL TEAM PERFORMED RESECTION AND RE-STAPLING.
18-12-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DISTAL GASTRECTOMY DUODENAL RESECTION, THE FIRST RELOAD PARTIALLY FIRED STOPS MID WAYS AND THE KNIFE BLADE RETRACTED. ON THE SECOND RELOAD, THE SHEET WAS DETACHED. A NEW RELOAD WAS USED TO RELOAD THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2025 | Malfunction | covidien surgical | 12-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DISTAL GASTRECTOMY DUODENAL RESECTION, THE FIRST RELOAD PARTIALLY FIRED STOPS MID WAYS AND THE KNIFE BLADE RETRACTED. ON THE SECOND RELOAD THE SHEET WAS DETACHED. A NEW RELOAD WAS USED TO RELOAD THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | ethicon endo-surgery | 12-01-2026 | ECHELON 3000 60MM STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE NURSE CHECKED A SUTURE LENGTH WHICH IS A 60MM OF TYVEK AND OPENED THE END CUTTER, A SUTURE LENGTH OF 45MM CAME OUT. THE NURSE NOTICED THIS WHEN LOADING THE CARTRIDGE, SO BOTH THE CARTRIDGE AND THE STAPLE LINE REINFORCEMENT HAD ALREADY BEEN OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE 1TH FIRING. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL COULD NOT FIRE AND WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | GST GREEN RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 100MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE GLC STAPLE LINE WAS NOT FULLY CUT THROUGH - UNCLEAR IF THE DISTANCE BETWEEN STAPLE & CUT LINE IS THE INTENDED 9MM OR CAUSED BY AN INCOMPLETE KNIFE BLADE ADVANCEMENT. NO SIGNIFICANT RESISTANCE IN CLOSING AND FIRING OF STAPLER. CASE PROCEEDED WITH A SECOND RELOAD. NO PATIENT CONSEQUENCES WERE REPORTED.
11-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | ENDO ILS 29MM, CURVED | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, STAPLES CANNOT BE FULLY DEPRESSED WHEN INSTALLED, RESULTING IN IMPROPER INSTALLATION. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | nan
17-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | ILS 25MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAP HEMICOLECTOMY SURGERY, AFTER FIRED, NOTED THAT THE TISSUE WAS NOT CUT, CHANGED ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE, IT WAS OBSERVED THAT THE POWERED STAPLER HAD A STICKY BUTTON. THE SURGICAL TEAM HAD TO OPEN A NEW STAPLER YET ITS BLISTER PACK WAS ALREADY TORN, COMPROMISING STERILITY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED USING ALTERNATIVE DEVICES, WITH NO PATIENT CONSEQUENCES REPORTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE, IT WAS OBSERVED THAT THE POWERED STAPLER HAD A STICKY BUTTON. THE SURGICAL TEAM HAD TO OPEN A NEW STAPLER YET ITS BLISTER PACK WAS ALREADY TORN, COMPROMISING STERILITY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED USING ALTERNATIVE DEVICES, WITH NO PATIENT CONSEQUENCES REPORTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE, IT WAS OBSERVED THAT THE POWERED STAPLER HAD A STICKY BUTTON. THE SURGICAL TEAM HAD TO OPEN A NEW STAPLER YET ITS BLISTER PACK WAS ALREADY TORN, COMPROMISING STERILITY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED USING ALTERNATIVE DEVICES, WITH NO PATIENT CONSEQUENCES REPORTED.
16-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | 75MM SELECTABLE NEW TLC | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, WHILST FIRING DEVICE GOT STUCK. SURGEON WAS ABLE TO PUSH THE FIRE BUTTON ALL THE WAY FORWARD, BUT WAS UNABLE TO GET IT BACK. WITH DEVICE IN PATIENT A SECOND NTLC WAS TAKEN TO CUT THROUGH THE COLON AND RELEASE THE STAPLER ON ONE END. THE OTHER END WAS CUT OUT BY A KNIFE. WHEN THE DEVICE WAS OUT OF THE PATIENT THE FIRING BUTTON WAS RETRIEVED WITH A LOT OF FORCE. THE SURGERY WAS DELAYED BY 15 MINUTES AND COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
16-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | 75MM SELECTABLE NEW TLC | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, WHILST FIRING DEVICE GOT STUCK. SURGEON WAS ABLE TO PUSH THE FIRE BUTTON ALL THE WAY FORWARD, BUT WAS UNABLE TO GET IT BACK. WITH DEVICE IN PATIENT A SECOND NTLC WAS TAKEN TO CUT THROUGH THE COLON AND RELEASE THE STAPLER ON ONE END. THE OTHER END WAS CUT OUT BY A KNIFE. WHEN THE DEVICE WAS OUT OF THE PATIENT THE FIRING BUTTON WAS RETRIEVED WITH A LOT OF FORCE. THE SURGERY WAS DELAYED BY 15 MINUTES AND COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
16-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, WHILST FIRING DEVICE GOT STUCK. SURGEON WAS ABLE TO PUSH THE FIRE BUTTON ALL THE WAY FORWARD, BUT WAS UNABLE TO GET IT BACK. WITH DEVICE IN PATIENT A SECOND NTLC WAS TAKEN TO CUT THROUGH THE COLON AND RELEASE THE STAPLER ON ONE END. THE OTHER END WAS CUT OUT BY A KNIFE. WHEN THE DEVICE WAS OUT OF THE PATIENT THE FIRING BUTTON WAS RETRIEVED WITH A LOT OF FORCE. THE SURGERY WAS DELAYED BY 15 MINUTES AND COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
16-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, WHILST FIRING DEVICE GOT STUCK. SURGEON WAS ABLE TO PUSH THE FIRE BUTTON ALL THE WAY FORWARD, BUT WAS UNABLE TO GET IT BACK. WITH DEVICE IN PATIENT A SECOND NTLC WAS TAKEN TO CUT THROUGH THE COLON AND RELEASE THE STAPLER ON ONE END. THE OTHER END WAS CUT OUT BY A KNIFE. WHEN THE DEVICE WAS OUT OF THE PATIENT THE FIRING BUTTON WAS RETRIEVED WITH A LOT OF FORCE. THE SURGERY WAS DELAYED BY 15 MINUTES AND COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
16-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | 75MM SELECTABLE NEW TLC | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, WHILST FIRING DEVICE GOT STUCK. SURGEON WAS ABLE TO PUSH THE FIRE BUTTON ALL THE WAY FORWARD, BUT WAS UNABLE TO GET IT BACK. WITH DEVICE IN PATIENT A SECOND NTLC WAS TAKEN TO CUT THROUGH THE COLON AND RELEASE THE STAPLER ON ONE END. THE OTHER END WAS CUT OUT BY A KNIFE. WHEN THE DEVICE WAS OUT OF THE PATIENT THE FIRING BUTTON WAS RETRIEVED WITH A LOT OF FORCE. THE SURGERY WAS DELAYED BY 15 MINUTES AND COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
16-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | 75MM SELECTABLE NEW TLC | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, WHILST FIRING DEVICE GOT STUCK. SURGEON WAS ABLE TO PUSH THE FIRE BUTTON ALL THE WAY FORWARD, BUT WAS UNABLE TO GET IT BACK. WITH DEVICE IN PATIENT A SECOND NTLC WAS TAKEN TO CUT THROUGH THE COLON AND RELEASE THE STAPLER ON ONE END. THE OTHER END WAS CUT OUT BY A KNIFE. WHEN THE DEVICE WAS OUT OF THE PATIENT THE FIRING BUTTON WAS RETRIEVED WITH A LOT OF FORCE. THE SURGERY WAS DELAYED BY 15 MINUTES AND COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
12-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | CONTOUR CURVED CUTTER STAPLER | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE THE NURSE ADVISED THE DEVICE WOULD NOT HOLD THE CARTRIDGE. NO OTHER DETAILS. NO PATIENT CONSEQUENCES WERE REPORTED.
17-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
19-12-2025 | Malfunction | davol inc sub c.r bard inc -1213643 | 12-01-2026 | SORBAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING THE PROCEDURE THE SURGEON ATTEMPTED TO FIXATE A MESH USING A SORBAFIX ENHANCED FIXATION DEVICE THE DEVICE WAS FIRED SEVERAL TIMES BUT FAILED TO MAKE THE NECESSARY SHOTS FOR STAPLING AND DID NOT SECURE THE MESH WITH THE FASTENERS. IT WAS REPORTED THAT NO LOOSE FASTENER PRESENTED AND THE FIXATION WAS APPLIED TO SOFT STRUCTURE. THE PROCEDURE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
20-12-2025 | Malfunction | davol inc sub c.r bard inc -1213643 | 12-01-2026 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A PROCEDURE USING OPTIFIX FIXATION DEVICE WHILE DEPLOYING, THE FASTENERS WERE FIRING SIDEWAYS, AND THE FASTENERS HAD BROKEN IN HALF. THE LOOSE FASTENERS WERE REMOVED FROM THE PATIENT'S BODY AND A NEW OPTIFIX DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT HARM OR INJURY.
24-11-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | 55MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN SURGERY, COULD NOT FIRE. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-11-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RECTUM PROCEDURE SURGERY, COULD NOT FIRE. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THERE WAS INCOMPLETE STAPLER FIRING, STAPLER PINS DOESN'T DEPLOYED
18-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2025 | Malfunction | ethicon endo-surgery | 12-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, COULD NOT FIRE. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 80MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING COLORECTAL PROCEDURES, OUT OF 20 PATIENTS, FIVE WERE BROUGHT BACK TO THE ICU DUE TO BLEEDING AND ALL REQUIRED BLOOD TRANSFUSIONS.
01-01-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 80MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING COLORECTAL PROCEDURES, OUT OF 20 PATIENTS, FIVE WERE BROUGHT BACK TO THE ICU DUE TO BLEEDING AND ALL REQUIRED BLOOD TRANSFUSIONS.
03-11-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Unintended System Motion | A051206 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
27-11-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 100MM | Failure to Fire | A050501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ANASTOMOSIS THE DEVICE FAILED TO FIRE. HAD TO OPEN A SECOND DEVICE DURING SURGERY. ANASTOMOSIS FAILED, PATIENT HAD TO COME BACK FOR SECOND SURGERY AND WAS VERY UNWELL.
27-11-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 100MM | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A ANASTOMOSIS THE DEVICE FAILED TO FIRE. HAD TO OPEN A SECOND DEVICE DURING SURGERY. ANASTOMOSIS FAILED, PATIENT HAD TO COME BACK FOR SECOND SURGERY AND WAS VERY UNWELL.
27-11-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 100MM | Failure to Fire | A050501 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A ANASTOMOSIS THE DEVICE FAILED TO FIRE. HAD TO OPEN A SECOND DEVICE DURING SURGERY. ANASTOMOSIS FAILED, PATIENT HAD TO COME BACK FOR SECOND SURGERY AND WAS VERY UNWELL.
27-11-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 100MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ANASTOMOSIS THE DEVICE FAILED TO FIRE. HAD TO OPEN A SECOND DEVICE DURING SURGERY. ANASTOMOSIS FAILED, PATIENT HAD TO COME BACK FOR SECOND SURGERY AND WAS VERY UNWELL.
27-11-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 100MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A ANASTOMOSIS THE DEVICE FAILED TO FIRE. HAD TO OPEN A SECOND DEVICE DURING SURGERY. ANASTOMOSIS FAILED, PATIENT HAD TO COME BACK FOR SECOND SURGERY AND WAS VERY UNWELL.
27-11-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 100MM | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A ANASTOMOSIS THE DEVICE FAILED TO FIRE. HAD TO OPEN A SECOND DEVICE DURING SURGERY. ANASTOMOSIS FAILED, PATIENT HAD TO COME BACK FOR SECOND SURGERY AND WAS VERY UNWELL.
01-01-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 80MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING COLORECTAL PROCEDURES, OUT OF 20 PATIENTS, FIVE WERE BROUGHT BACK TO THE ICU DUE TO BLEEDING AND ALL REQUIRED BLOOD TRANSFUSIONS.
01-01-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 80MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED BY THE SALES REPRESENTATIVE THAT DURING COLORECTAL PROCEDURES, OUT OF 20 PATIENTS, FIVE WERE BROUGHT BACK TO THE ICU DUE TO BLEEDING AND ALL REQUIRED BLOOD TRANSFUSIONS. THE DEVICE IS NOT AVAILABLE FOR RETURN.
03-11-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Unintended System Motion | A051206 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | covidien surgical | 12-01-2026 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
01-01-2025 | Injury | ethicon endo-surgery | 12-01-2026 | ECHELON LINEAR CUTTER 80MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING COLORECTAL PROCEDURES, OUT OF 20 PATIENTS, FIVE WERE BROUGHT BACK TO THE ICU DUE TO BLEEDING AND ALL REQUIRED BLOOD TRANSFUSIONS.
03-11-2025 | Malfunction | surgical | 12-01-2026 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | surgical | 12-01-2026 | ENDO GIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | surgical | 12-01-2026 | ENDO GIA | Unintended System Motion | A051206 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | surgical | 12-01-2026 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
03-11-2025 | Malfunction | surgical | 12-01-2026 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN IN THE CAVITY, THE STAPLER WAS STUCK AND WHEN TRYING TO FIRE IT WOULD NOT WORK. SUDDENLY TURNING RIGHT TO LEFT. WHEN THE STAPLER RELOAD WAS ATTEMPTED TO BE REMOVED IT BROKE AT THE JUNCTION WITH THE ADAPTER.
20-10-2025 | Malfunction | covidien surgical | 12-01-2026 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR PROCEDURE, THE TACKS WERE NOT COMING OUT WHILE FIRING AND WERE FALLING DOWN. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
20-10-2025 | Malfunction | covidien surgical | 12-01-2026 | ABSORBATACK | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR PROCEDURE, THE TACKS WERE NOT COMING OUT WHILE FIRING AND WERE FALLING DOWN. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
29-12-2025 | Malfunction | covidien | 12-01-2026 | ENDO GIA ULTRA | No Device Output | A0907 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LUNG LOBECTOMY, THE DEVICE COULD NOT BE GRIPPED PROPERLY AND WAS UNABLE TO FIRE WHILE DETACHING THE LUNG LOBE PARENCHYMA. IT WAS ALSO NOTED THAT THE FEEL OF THE DEVICE WAS DIFFERENT FROM USUAL. A NEW HANDLE AND RELOAD WERE REPLACED WITH A NEW ONE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
29-12-2025 | Malfunction | covidien | 12-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LUNG LOBECTOMY, THE DEVICE COULD NOT BE GRIPPED PROPERLY AND WAS UNABLE TO FIRE WHILE DETACHING THE LUNG LOBE PARENCHYMA. IT WAS ALSO NOTED THAT THE FEEL OF THE DEVICE WAS DIFFERENT FROM USUAL. A NEW HANDLE AND RELOAD WERE REPLACED WITH A NEW ONE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
29-12-2025 | Malfunction | covidien surgical | 12-01-2026 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
10-12-2025 | Malfunction | covidien surgical | 12-01-2026 | nan | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER CANCER RESECTION SURGERY, THE STAPLER WAS UNABLE TO CUT AND THE DEVICE DID NOT FIRE. THE STAPLER WAS REPLACED, AND THE SURGERY CONTINUED. AS A RESULT, THE OPERATION TIME WAS PROLONGED BUT NOT MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
17-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Structural Problem | A0511 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A LAPAROSCOPIC COLECTOMY, WHEN VISUALLY CHECKED IT BEFORE USE, THE STAPLE DRIVERS WERE RAISED TO THE THIRD ROW FROM THE ROOT, AND THE KNIFE HAD SLIGHTLY COME OUT. THE SWING TAB WAS STILL CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
17-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A LAPAROSCOPIC COLECTOMY, WHEN VISUALLY CHECKED IT BEFORE USE, THE STAPLE DRIVERS WERE RAISED TO THE THIRD ROW FROM THE ROOT, AND THE KNIFE HAD SLIGHTLY COME OUT. THE SWING TAB WAS STILL CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2025 | Malfunction | covidien surgical | 13-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY ON THE PULMONARY PARENCHYMA, THE KNIFE STOPPED MIDWAY AND COULD NOT BE FIRED THROUGH THE END. THE KNIFE WAS RETRACTED, AND A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THE PATIENT HAS A C ALCIFICATION, AND THE TISSUE HAS BECOME QUITE HARD DUE TO A COMPLICATION OF PNEUMONIA. THERE WAS NO PATIENT INJURY.
22-12-2025 | Malfunction | covidien surgical | 13-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY ON THE PULMONARY PARENCHYMA, THE KNIFE STOPPED MIDWAY AND COULD NOT BE FIRED THROUGH THE END. THE KNIFE WAS RETRACTED, AND A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THE PATIENT HAS A C ALCIFICATION, AND THE TISSUE HAS BECOME QUITE HARD DUE TO A COMPLICATION OF PNEUMONIA. THERE WAS NO PATIENT INJURY.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON LINEAR CUTTER 80MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE WOULD NOT LOAD THE STAPLES. NO PATIENT HARM HAS BEEN REPORTED.
22-12-2025 | Malfunction | covidien surgical | 13-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY ON THE PULMONARY PARENCHYMA, THE KNIFE STOPPED MIDWAY AND COULD NOT BE FIRED THROUGH THE END. THE KNIFE WAS RETRACTED, AND A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THE PATIENT HAS A C ALCIFICATION, AND THE TISSUE HAS BECOME QUITE HARD DUE TO A COMPLICATION OF PNEUMONIA. THERE WAS NO PATIENT INJURY.
22-12-2025 | Injury | covidien surgical | 13-01-2026 | ENDO GIA | Failure to Read Input Signal | A1301 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY PROCEDURE, THE POWER HANDLE AND THE CARTRIDGE WERE NOT PROPERLY RECOGNIZED WHEN INITIALLY INSTALLED, AND THIS ISSUE OCCURRED MULTIPLE TIMES PRIOR TO PATIENT CONTACT. THE 30MM CARTRIDGE WAS APPLIED TO A VASCULAR STRUCTURE AND FIRED, BUT FIRING DID NOT OCCUR PROPERLY, RESULTING IN VASCULAR LEAKAGE. MALFORMED STAPLES WERE OBSERVED AFTER FIRING, AND THE STAPLE LINE WAS INCOMPLETE. DUE TO THE BLEEDING, THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC (VATS) TO OPEN SURGERY, AND BLEEDING WAS CONTROLLED BY SUTURE. THE OPERATION WAS SUCCESSFULLY COMPLETED.
22-12-2025 | Injury | covidien surgical | 13-01-2026 | ENDO GIA | Failure to Read Input Signal | A1301 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY PROCEDURE, THE POWER HANDLE AND THE CARTRIDGE WERE NOT PROPERLY RECOGNIZED WHEN INITIALLY INSTALLED, AND THIS ISSUE OCCURRED MULTIPLE TIMES PRIOR TO PATIENT CONTACT. THE 30MM CARTRIDGE WAS APPLIED TO A VASCULAR STRUCTURE AND FIRED, BUT FIRING DID NOT OCCUR PROPERLY, RESULTING IN VASCULAR LEAKAGE. MALFORMED STAPLES WERE OBSERVED AFTER FIRING, AND THE STAPLE LINE WAS INCOMPLETE. DUE TO THE BLEEDING, THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC (VATS) TO OPEN SURGERY, AND BLEEDING WAS CONTROLLED BY SUTURE. THE OPERATION WAS SUCCESSFULLY COMPLETED.
22-12-2025 | Injury | covidien surgical | 13-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY PROCEDURE, THE POWER HANDLE AND THE CARTRIDGE WERE NOT PROPERLY RECOGNIZED WHEN INITIALLY INSTALLED, AND THIS ISSUE OCCURRED MULTIPLE TIMES PRIOR TO PATIENT CONTACT. THE 30MM CARTRIDGE WAS APPLIED TO A VASCULAR STRUCTURE AND FIRED, BUT FIRING DID NOT OCCUR PROPERLY, RESULTING IN VASCULAR LEAKAGE. MALFORMED STAPLES WERE OBSERVED AFTER FIRING, AND THE STAPLE LINE WAS INCOMPLETE. DUE TO THE BLEEDING, THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC (VATS) TO OPEN SURGERY, AND BLEEDING WAS CONTROLLED BY SUTURE. THE OPERATION WAS SUCCESSFULLY COMPLETED.
22-12-2025 | Injury | covidien surgical | 13-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY PROCEDURE, THE POWER HANDLE AND THE CARTRIDGE WERE NOT PROPERLY RECOGNIZED WHEN INITIALLY INSTALLED, AND THIS ISSUE OCCURRED MULTIPLE TIMES PRIOR TO PATIENT CONTACT. THE 30MM CARTRIDGE WAS APPLIED TO A VASCULAR STRUCTURE AND FIRED, BUT FIRING DID NOT OCCUR PROPERLY, RESULTING IN VASCULAR LEAKAGE. MALFORMED STAPLES WERE OBSERVED AFTER FIRING, AND THE STAPLE LINE WAS INCOMPLETE. DUE TO THE BLEEDING, THE PROCEDURE WAS CONVERTED FROM LAPAROSCOPIC (VATS) TO OPEN SURGERY, AND BLEEDING WAS CONTROLLED BY SUTURE. THE OPERATION WAS SUCCESSFULLY COMPLETED.
01-01-2025 | Injury | ethicon endo-surgery | 13-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Mechanical Problem | A05 | Insufficient Information | nan
01-01-2025 | Injury | ethicon endo-surgery | 13-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Read Input Signal | A1301 | Insufficient Information | nan
17-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON LINEAR CUTTER 80MM | Break | A0401 | No Clinical Signs, Symptoms or Conditions | nan
17-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON LINEAR CUTTER 80MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
17-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON LINEAR CUTTER 80MM | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | nan
17-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON LINEAR CUTTER 80MM | Packaging Problem | A0205 | No Clinical Signs, Symptoms or Conditions | nan
17-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON LINEAR CUTTER 80MM | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | nan
17-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON LINEAR CUTTER 80MM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | nan
19-12-2025 | Injury | davis & geck caribe | 13-01-2026 | UNKNOWN ENDO STITCH INSTRUMENT | Break | A0401 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC TOTAL HYSTERECTOMY (TLH) AND SALPINGECTOMY OR SLING, WHILE CLOSING THE CUFF, THE NEEDLE BROKE IN TWO. THE TIP HALF OF THE NEEDLE FELL INTO THE PATIENT¿S CAVITY. THE PIECE OF THE NEEDLE WAS FOUND AND RETRIEVED. X-RAY WAS PERFORMED FOR THE EXPRESS PURPOSE OF LOCATING THE NEEDLE OR CONFIRMING THAT ALL PIECES WERE LOCATED. THE LOADING UNIT WAS ALSO DIFFICULT TO UNLOAD. THE SURGERY TOOK EXTRA TIME TO FIND THE BROKEN NEEDLE PIECE AND WENT OVER 85 MINUTES THAN SCHEDULED. ANOTHER LOADING UNIT WAS USED TO COMPLETE THE CASE.
19-12-2025 | Injury | davis & geck caribe | 13-01-2026 | UNKNOWN ENDO STITCH INSTRUMENT | Detachment of Device or Device Component | A0501 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC TOTAL HYSTERECTOMY (TLH) AND SALPINGECTOMY OR SLING, WHILE CLOSING THE CUFF, THE NEEDLE BROKE IN TWO. THE TIP HALF OF THE NEEDLE FELL INTO THE PATIENT¿S CAVITY. THE PIECE OF THE NEEDLE WAS FOUND AND RETRIEVED. X-RAY WAS PERFORMED FOR THE EXPRESS PURPOSE OF LOCATING THE NEEDLE OR CONFIRMING THAT ALL PIECES WERE LOCATED. THE LOADING UNIT WAS ALSO DIFFICULT TO UNLOAD. THE SURGERY TOOK EXTRA TIME TO FIND THE BROKEN NEEDLE PIECE AND WENT OVER 85 MINUTES THAN SCHEDULED. ANOTHER LOADING UNIT WAS USED TO COMPLETE THE CASE.
19-12-2025 | Injury | davis & geck caribe | 13-01-2026 | UNKNOWN ENDO STITCH INSTRUMENT | Human-Device Interface Problem | A22 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC TOTAL HYSTERECTOMY (TLH) AND SALPINGECTOMY OR SLING, WHILE CLOSING THE CUFF, THE NEEDLE BROKE IN TWO. THE TIP HALF OF THE NEEDLE FELL INTO THE PATIENT¿S CAVITY. THE PIECE OF THE NEEDLE WAS FOUND AND RETRIEVED. X-RAY WAS PERFORMED FOR THE EXPRESS PURPOSE OF LOCATING THE NEEDLE OR CONFIRMING THAT ALL PIECES WERE LOCATED. THE LOADING UNIT WAS ALSO DIFFICULT TO UNLOAD. THE SURGERY TOOK EXTRA TIME TO FIND THE BROKEN NEEDLE PIECE AND WENT OVER 85 MINUTES THAN SCHEDULED. ANOTHER LOADING UNIT WAS USED TO COMPLETE THE CASE.
17-12-2025 | Injury | ethicon endo-surgery | 13-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | Insufficient Information | nan
17-12-2025 | Injury | ethicon endo-surgery | 13-01-2026 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | Insufficient Information | nan
17-12-2025 | Injury | ethicon endo-surgery | 13-01-2026 | FLEX 60 ARTICULATING | Appropriate Term/Code Not Available |  | Insufficient Information | nan
18-12-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS PROCEDURE, THE BLADE OF THE RELOAD APPEARED TO CUT BEFORE STAPLES WERE PLACED DURING FIRING. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. IT WAS NOTED THAT THE ADAPTER WAS INSPECTED BY THE REPRESENTATIVE THAT WAS PRESENT ON THE EVENT, THE ADAPTER HAD ONE USE LEFT AND HAD NO ISSUE.
18-12-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS PROCEDURE, THE BLADE OF THE RELOAD APPEARED TO CUT BEFORE STAPLES WERE PLACED DURING FIRING. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. IT WAS NOTED THAT THE ADAPTER WAS INSPECTED BY THE REPRESENTATIVE THAT WAS PRESENT ON THE EVENT, THE ADAPTER HAD ONE USE LEFT AND HAD NO ISSUE.
18-12-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS PROCEDURE, THE BLADE OF THE RELOAD APPEARED TO CUT BEFORE STAPLES WERE PLACED DURING FIRING. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. IT WAS NOTED THAT THE ADAPTER WAS INSPECTED BY THE REPRESENTATIVE THAT WAS PRESENT ON THE EVENT, THE ADAPTER HAD ONE USE LEFT AND HAD NO ISSUE.
18-12-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS PROCEDURE, THE BLADE OF THE RELOAD APPEARED TO CUT BEFORE STAPLES WERE PLACED DURING FIRING. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. IT WAS NOTED THAT THE ADAPTER WAS INSPECTED BY THE REPRESENTATIVE THAT WAS PRESENT ON THE EVENT, THE ADAPTER HAD ONE USE LEFT AND HAD NO ISSUE.
25-12-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARENCHYMA LUNG RESECTION, AFTER FIRING, WHEN OPENING, THE TISSUE GOT CAUGHT (ENTERED) IN THE STAPLE STORAGE PART OF THE CARTRIDGE, AND REMOVAL OF THE TISSUE WAS DIFFICULT. THE JAWS WERE STILL ABLE TO BE OPENED. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2025 | Malfunction | covidien surgical | 13-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN LIVER RESECTION, DURING VASCULAR TREATMENT, THE STAPLER WAS STUCK AFTER A SINGLE FIRING GRIP. THE FIRING WAS STOPPED, AND THE DEVICE WAS WITHDRAWN USING THE BLACK RETURN KNOB. IT WAS REPORTED THAT THE FIRING HAD BEEN PARTIALLY COMPLETED UP TO THE POINT WHERE THE HANDLE WAS GRIPPED. AFTERWARDS, A DIFFERENT CARTRIDGE WAS ATTACHED, THE FIRING WAS PERFORMED, AND THE PROCEDURE WAS COMPLETED WITHOUT FURTHER INCIDENT. THERE WAS NO PATIENT INJURY.
18-12-2025 | Malfunction | covidien surgical | 13-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING JEJUNAL RESECTION, WHILE IN THE JEJUNUM, THE JAW OF THE POWERED STAPLER WOULD NO LONGER OPEN. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2026 | Malfunction | covidien | 13-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GBP (ROUX-EN-Y GASTRIC BYPASS) PROCEDURE, AFTER LOADING, THE HANDLE BECAME BLOCKED AND WAS NOT ABLE TO FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
15-09-2025 | Malfunction | covidien surgical | 13-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE POWERED STAPLER WAS DIFFICULT TO FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
15-09-2025 | Malfunction | covidien surgical | 13-01-2026 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE POWERED STAPLER WAS DIFFICULT TO FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
18-12-2025 | Malfunction | covidien surgical | 13-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING JEJUNAL RESECTION, WHILE IN THE JEJUNUM, THE JAW OF THE POWERED STAPLER WOULD NO LONGER OPEN. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-11-2025 | Malfunction | covidien surgical | 13-01-2026 | ENDO GIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE RELOAD WAS NOT RECOGNIZED. THE DEVICE WAS REPLACED WITH A NEW RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
03-11-2025 | Malfunction | covidien surgical | 13-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE RELOAD WAS NOT RECOGNIZED. THE DEVICE WAS REPLACED WITH A NEW RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
18-12-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING JEJUNAL RESECTION, WHILE IN THE JEJUNUM, THE JAW OF THE POWERED STAPLER WOULD NO LONGER OPEN. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC PARTIAL PULMONARY RESECTION, AT THE PULMONARY PARENCHYMA, THE GREEN BUTTON COULD NOT BE PRESSED, AND THE FIRING COULD NOT BE PERFORMED AT ALL. A NEW RELOAD AND HANDLE WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2025 | Malfunction | covidien surgical | 13-01-2026 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC PARTIAL PULMONARY RESECTION, AT THE PULMONARY PARENCHYMA, THE GREEN BUTTON COULD NOT BE PRESSED, AND THE FIRING COULD NOT BE PERFORMED AT ALL. A NEW RELOAD AND HANDLE WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-12-2025 | Malfunction | covidien surgical | 13-01-2026 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLER KNIFE DID NOT CUT FAR ENOUGH AND THE END OF THE SUTURE THREAD WAS NOT CUT THROUGH, REQUIRING SCISSORS TO COMPLETE THE CUT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
08-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN PROCEDURE, THERE WAS AN UNEXPECTED SOUND DURING FIRING, HIGHER SOUND THAN USUAL. IT COULD BE FIRED TO SOME EXTENT, AND THE LEAK TEST WAS NEGATIVE. THE PROCEDURE WAS COMPLETED AS IT WAS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
08-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN PROCEDURE, THERE WAS AN UNEXPECTED SOUND DURING FIRING, HIGHER SOUND THAN USUAL. IT COULD BE FIRED TO SOME EXTENT, AND THE LEAK TEST WAS NEGATIVE. THE PROCEDURE WAS COMPLETED AS IT WAS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE THE DEVICE MISFIRED. NO PATIENT CONSEQUENCES WERE REPORTED.
04-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE THE DEVICE MISFIRED. NO PATIENT CONSEQUENCES WERE REPORTED.
23-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON 3000 45 STANDARD | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | nan
04-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE MISFIRED DURING SURGERY. NO PATIENT CONSEQUENCES WERE REPORTED.
04-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE MISFIRED DURING SURGERY. NO PATIENT CONSEQUENCES WERE REPORTED.
13-01-2026 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, WHEN TRIED TO FIRE CLAMPING THE TISSUE, THE DEVICE COULD NOT BE FIRED AND THE KNIFE DID NOT MOVE FORWARD. A NEW MAIN BODY AND CARTRIDGE WAS USED TO COMPLETE THE CASE. THE CARTRIDGE COLOR WAS WHITE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
08-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | PWRD ECH FLEX 60MM 28CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAD SUDDEN INABILITY TO USE DURING SURGERY THUS, SERIOUSLY AFFECTING THE OPERATION. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
24-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | GST GOLD RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AFTER FIRING; AND THE ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
24-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | GST GOLD RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AFTER FIRING; AND THE ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
24-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | GST GOLD RELOAD 45MM 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AFTER FIRING; AND THE ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
24-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | GST GOLD RELOAD 45MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AFTER FIRING; AND THE ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
30-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, WHEN THE SURGEON PLANT DOWN ON THE APPENDIX TO FIRE, THE DEVICE FIRED HALFWAY AND THEN STOPPED. THE SURGICAL TEAM REMOVED AND REINSERTED THE BATTERY, AND THE BLADE CAME BACK. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
31-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN EXPLORATORY LAPAROTOMY, THE SURGEON USED THE STAPLER FIRED THE BLUE RELOAD AND WHEN ATTEMPTING TO REMOVE THE RELOAD, THE SILVER SHIELD POPPED OFF. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
31-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN EXPLORATORY LAPAROTOMY, THE SURGEON USED THE STAPLER FIRED THE BLUE RELOAD AND WHEN ATTEMPTING TO REMOVE THE RELOAD, THE SILVER SHIELD POPPED OFF. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-11-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN RECTUM PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
11-12-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | nan
04-01-2026 | Malfunction | ethicon endo-surgery | 13-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2026 | Malfunction | ethicon endo-surgery | 13-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2026 | Malfunction | ethicon endo-surgery | 13-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2026 | Malfunction | ethicon endo-surgery | 13-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN ANTERIOR RESECTION PROCEDURE, WHEN ATTEMPTING TO STAPLE ON TISSUE, THE CS40G STAPLER WAS UNABLE TO FIRE AND STAPLE ON THE TISSUE. THIS STAPLER WAS THEN ISOLATED AND A NEW CS40G WAS OPENED TO CONTINUE THE PROCEDURE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY AND NO PATIENT CONSEQUENCES WERE REPORTED.
04-01-2026 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A IIEPATECTOMY PROCEDURE, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2026 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A IIEPATECTOMY PROCEDURE, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2026 | Malfunction | ethicon endo-surgery | 13-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A IIEPATECTOMY PROCEDURE, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | CONTOUR CURVED CUTTER STAPLER | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, COULD NOT PROPERLY FIRE AS CARTRIDGE ANVIL BROKE. NO PATIENT CONSEQUENCES WERE REPORTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | CONTOUR CURVED CUTTER STAPLER | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, COULD NOT PROPERLY FIRE AS CARTRIDGE ANVIL BROKE. NO PATIENT CONSEQUENCES WERE REPORTED.
13-11-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | 55 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-11-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE CARTRIDGE MISFIRED. YELLOW TAB OPENED, PINS NOT DEPLOYED, NO CUT WAS THERE. STAPLER STUCK. NO PATIENT CONSEQUENCES WERE REPORTED.
20-11-2025 | Malfunction | ethicon endo-surgery | 13-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2025 | Malfunction | covidien surgical | 13-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE, THE STAPLER DID NOT FIRE. THE RELOAD WAS REPLACED AND THE SURGERY WAS COMPLETED SU CCESSFULLY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
04-04-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | nan
04-04-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | nan
26-05-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLORECTAL SURGERY, WHEN DETACHING THE RECTUM, THE USER WAS UNABLE TO INSTALL THE RELOAD FOR THE SECOND TIME. REPLACEMENT OF A NEW RELOAD WAS DONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK DUE TO THINK TISSUE.
26-05-2025 | Malfunction | covidien | 13-01-2026 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLORECTAL SURGERY, WHEN DETACHING THE RECTUM, THE USER WAS UNABLE TO INSTALL THE RELOAD FOR THE SECOND TIME. REPLACEMENT OF A NEW RELOAD WAS DONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK DUE TO THINK TISSUE.
22-12-2025 | Malfunction | intuitive surgical | 13-01-2026 | SUREFORM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SURGICAL PROCEDURE, THE CUSTOMER USED THE SUREFORM 45 STAPLER INSTRUMENT ONCE AND IT WORKED AS INTENDED. THE CUSTOMER WENT TO RELOAD FOR THE NEXT STAPLE, AND THE JAWS OF THE INSTRUMENT WOULD NOT OPEN, NOR RESPOND. THE CUSTOMER COMPLETED THE PROCEDURE, AND NO KNOWN IMPACT OR PATIENT CONSEQUENCE WAS REPORTED.
16-12-2025 | Injury | covidien | 13-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A BYPASS PROCEDURE THE LAST TWO RELOADS DID NOT FORM STAPLES CORRECTLY, RESULTING IN THE STOMACH REMAINING OPEN. AN INCOMPLETE STAPLE LINE WAS NOTED AT THE DISTAL LOCATION. THESE ISSUES EXTENDED THE SURGICAL TIME BY 20 MINUTES TO CUT THE MALFORMED CLIPS AND REQUIRED THE DOCTOR TO CUT MORE CENTIMETERS IN THE STOMACH.
16-12-2025 | Injury | covidien | 13-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A BYPASS PROCEDURE THE LAST TWO RELOADS DID NOT FORM STAPLES CORRECTLY, RESULTING IN THE STOMACH REMAINING OPEN. AN INCOMPLETE STAPLE LINE WAS NOTED AT THE DISTAL LOCATION. THESE ISSUES EXTENDED THE SURGICAL TIME BY 20 MINUTES TO CUT THE MALFORMED CLIPS AND REQUIRED THE DOCTOR TO CUT MORE CENTIMETERS IN THE STOMACH.
16-12-2025 | Injury | covidien | 13-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A BYPASS PROCEDURE THE LAST TWO RELOADS DID NOT FORM STAPLES CORRECTLY, RESULTING IN THE STOMACH REMAINING OPEN. AN INCOMPLETE STAPLE LINE WAS NOTED AT THE DISTAL LOCATION. THESE ISSUES EXTENDED THE SURGICAL TIME BY 20 MINUTES TO CUT THE MALFORMED CLIPS AND REQUIRED THE DOCTOR TO CUT MORE CENTIMETERS IN THE STOMACH.
16-12-2025 | Injury | covidien | 13-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A BYPASS PROCEDURE THE LAST TWO RELOADS DID NOT FORM STAPLES CORRECTLY, RESULTING IN THE STOMACH REMAINING OPEN. AN INCOMPLETE STAPLE LINE WAS NOTED AT THE DISTAL LOCATION. THESE ISSUES EXTENDED THE SURGICAL TIME BY 20 MINUTES TO CUT THE MALFORMED CLIPS AND REQUIRED THE DOCTOR TO CUT MORE CENTIMETERS IN THE STOMACH.
16-12-2025 | Injury | covidien | 13-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A BYPASS PROCEDURE THE LAST TWO RELOADS DID NOT FORM STAPLES CORRECTLY, RESULTING IN THE STOMACH REMAINING OPEN. AN INCOMPLETE STAPLE LINE WAS NOTED AT THE DISTAL LOCATION. THESE ISSUES EXTENDED THE SURGICAL TIME BY 20 MINUTES TO CUT THE MALFORMED CLIPS AND REQUIRED THE DOCTOR TO CUT MORE CENTIMETERS IN THE STOMACH.
16-12-2025 | Injury | covidien | 13-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A BYPASS PROCEDURE THE LAST TWO RELOADS DID NOT FORM STAPLES CORRECTLY, RESULTING IN THE STOMACH REMAINING OPEN. AN INCOMPLETE STAPLE LINE WAS NOTED AT THE DISTAL LOCATION. THESE ISSUES EXTENDED THE SURGICAL TIME BY 20 MINUTES TO CUT THE MALFORMED CLIPS AND REQUIRED THE DOCTOR TO CUT MORE CENTIMETERS IN THE STOMACH.
16-12-2025 | Injury | covidien | 13-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A BYPASS PROCEDURE THE LAST TWO RELOADS DID NOT FORM STAPLES CORRECTLY, RESULTING IN THE STOMACH REMAINING OPEN. AN INCOMPLETE STAPLE LINE WAS NOTED AT THE DISTAL LOCATION. THESE ISSUES EXTENDED THE SURGICAL TIME BY 20 MINUTES TO CUT THE MALFORMED CLIPS AND REQUIRED THE DOCTOR TO CUT MORE CENTIMETERS IN THE STOMACH.
16-12-2025 | Injury | covidien | 13-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A BYPASS PROCEDURE THE LAST TWO RELOADS DID NOT FORM STAPLES CORRECTLY, RESULTING IN THE STOMACH REMAINING OPEN. AN INCOMPLETE STAPLE LINE WAS NOTED AT THE DISTAL LOCATION. THESE ISSUES EXTENDED THE SURGICAL TIME BY 20 MINUTES TO CUT THE MALFORMED CLIPS AND REQUIRED THE DOCTOR TO CUT MORE CENTIMETERS IN THE STOMACH.
23-12-2025 | Malfunction | covidien surgical | 14-01-2026 | ENDO GIA | Mechanics Altered | A0507 | Tissue Breakdown | nan
23-12-2025 | Malfunction | covidien surgical | 14-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | nan
23-12-2025 | Malfunction | covidien surgical | 14-01-2026 | ENDO GIA | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, AFTER FIRING, THE STAPLES DID NOT COME OUT OF THE CARTRIDGE AND ONTO THE TISSUE. IT OCCURRED ON TWO RELOADS. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
23-12-2025 | Malfunction | covidien surgical | 14-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, AFTER FIRING, THE STAPLES DID NOT COME OUT OF THE CARTRIDGE AND ONTO THE TISSUE. IT OCCURRED ON TWO RELOADS. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
10-12-2025 | Malfunction | covidien surgical | 14-01-2026 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, THE DEVICE COULD NOT BE REFILLED DURING THE PLACEMENT OF WALL REINFORCEMENT. FROM THE OUTSET, IT WAS IMPOSSIBLE TO REFILL THE DEVICE WITH THE PROVIDED REFILLS, AND DIFFICULTIES WERE ENCOUNTERED WHEN LOADING THE RELOAD ONTO THE HANDLE. THE RELOAD WAS NOT SECURELY FASTENED ONTO THE SHAFT, AND THERE WAS DIFFICULTY ARTICULATING THE DEVICE. ANOTHER DEVICE WAS USED INSTEAD. NO PATIENT WAS INVOLVED.
23-12-2025 | Malfunction | covidien | 14-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, DURING ILEAL RESECTION, THE STAPLER COULD NOT BE FIRED AT ALL. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien | 14-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY IN A LAPAROSCOPIC APPENDECTOMY, THE SCRUB NURSE LOADED THE RELOAD ONTO THE HANDLE AND CLOSED IT. THE STAPLER WAS HANDED TO SURGEON WHO PLACED IT DOWN THE TROCAR. WHEN THE RELOAD WAS PLACED AROUND THE APPENDIX STUMP, THE JAWS OF THE RELOAD WOULD NOT CLOSE AROUND THE TISSUE AND WERE FAR APART. THE SURGEON OPENED THE JAWS AND CLOSED IT AGAIN, BUT THE SAME ISSUE OCCURRED. THE SURGEON TRIED TO FIRE THE STAPLE BUT COULD NOT SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD AND THE STAPLER FIRED. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien | 14-01-2026 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY IN A LAPAROSCOPIC APPENDECTOMY, THE SCRUB NURSE LOADED THE RELOAD ONTO THE HANDLE AND CLOSED IT. THE STAPLER WAS HANDED TO SURGEON WHO PLACED IT DOWN THE TROCAR. WHEN THE RELOAD WAS PLACED AROUND THE APPENDIX STUMP, THE JAWS OF THE RELOAD WOULD NOT CLOSE AROUND THE TISSUE AND WERE FAR APART. THE SURGEON OPENED THE JAWS AND CLOSED IT AGAIN, BUT THE SAME ISSUE OCCURRED. THE SURGEON TRIED TO FIRE THE STAPLE BUT COULD NOT SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD AND THE STAPLER FIRED. THERE WAS NO PATIENT INJURY.
01-01-2025 | Injury | ethicon endo-surgery | 14-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Appropriate Term/Code Not Available |  | Hemorrhage/Blood Loss/Bleeding | nan
22-12-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE SURGEON SAID IT IS UNABLE TO PRESS THE TRIGGER, BUT THEY DID NOT PRE-PRESS THE TRIGGER BEFORE FIRING. NO PATIENT CONSEQUENCES WERE REPORTED.
24-11-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, INTRAOPERATIVELY THE DEVICE COULD NOT BE CLOSED AND A NEW DEVICE WAS USED TO FINISH THE PROCEDURE WITHOUT COMPLICATION. NO PATIENT CONSEQUENCES WERE REPORTED.
08-01-2026 | Malfunction | covidien | 14-01-2026 | ENDO GIA ULTRA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, AT THE GALLBLADDER ROOT PART, WHEN THE FIRE OCCURRED, A BEEP SOUNDED FROM THE HANDLE AND THE FIRING STOPPED. THE DEVICE WAS CHANGED TO COMPETITOR'S DEVICE TO RESOLVE THE ISSUE.
08-01-2026 | Malfunction | covidien | 14-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, AT THE GALLBLADDER ROOT PART, WHEN THE FIRE OCCURRED, A BEEP SOUNDED FROM THE HANDLE AND THE FIRING STOPPED. THE DEVICE WAS CHANGED TO COMPETITOR'S DEVICE TO RESOLVE THE ISSUE.
04-11-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCTECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-01-2026 | Malfunction | covidien surgical | 14-01-2026 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, AT THE GALLBLADDER ROOT PART, WHEN THE FIRE OCCURRED, A BEEP SOUNDED FROM THE HANDLE AND THE FIRING STOPPED. THE DEVICE WAS CHANGED TO COMPETITOR'S DEVICE TO RESOLVE THE ISSUE.
11-12-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
05-01-2026 | Malfunction | covidien | 14-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, FIRING ISSUES OCCURRED WHEN THE INSTRUMENT DID NOT FIRE AS INTENDED AFTER THE GREEN SAFETY BUTTON WAS PRESSED. THE MANUAL HANDLE DID NOT OPERATE AS EXPECTED, AND SQUEEZING THE HANDLE DID NOT RESULT IN FIRING. THE HANDLE WAS REPLACED WITH ANOTHER MANUAL HANDLE, BUT THE ISSUE PERSISTED UNTIL THE RELOAD WAS REPLACED, AFTER WHICH THE DEVICE FIRED AS INTENDED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
05-01-2026 | Malfunction | covidien surgical | 14-01-2026 | UNKNOWN TRI STAPLE PRODUCT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, FIRING ISSUES OCCURRED WHEN THE INSTRUMENT DID NOT FIRE AS INTENDED AFTER THE GREEN SAFETY BUTTON WAS PRESSED. THE MANUAL HANDLE DID NOT OPERATE AS EXPECTED, AND SQUEEZING THE HANDLE DID NOT RESULT IN FIRING. THE HANDLE WAS REPLACED WITH ANOTHER MANUAL HANDLE, BUT THE ISSUE PERSISTED UNTIL THE RELOAD WAS REPLACED, AFTER WHICH THE DEVICE FIRED AS INTENDED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
05-01-2026 | Malfunction | ethicon endo-surgery | 14-01-2026 | GST WHITE RELOAD, 45MM, 6-ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAD A DAMAGED SLED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
27-11-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Appropriate Term/Code Not Available |  | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A COLOSTOMY REVERSAL PROCEDURE, THE SURGEON HAD DIFFICULTY REMOVING THE DEVICE POST FIRING. THE SURGEON DESCRIBED THE ISSUE AS "HAVING TO RIP THE STAPLER FROM THE PATIENT'S RECTUM". THE STAPLER WAS CLOSED WITH THE ORANGE INDICATOR AT THE BASE OF THE TISSUE COMPRESSION SCALE. THE SURGEON DESCRIBED CLOSING STAPLER AS ROUTINE. THE STAPLER FIRED NORMALLY, THE BREAKAWAY WASHER CRUNCH WAS HEARD, AND THE GREEN CHECK MARK INDICATED A COMPLETED FIRING CYCLE. THE STAPLER WAS OPENED 4 X 1/2 TURNS IN THE ROUTINE FASHION, AND THE DEVICE WAS ROTATED 90° IN BOTH DIRECTIONS. THE SURGEON NOTED THAT WHEN ROTATING THE STAPLER, THE BOWEL ROTATED WITH THE DEVICE. AFTER REMOVAL OF THE STAPLER, THE DONUTS WERE INSPECTED AND WERE OBSERVED TO BE FULLY FORMED AND INTACT. THE BREAKAWAY WASHER WAS OBSERVED WITH THE DONUTS IN THE USUAL FASHION. THE ANASTOMOSIS WAS TESTED WITH A ROUTINE AIR LEAK TEST VIA A RIGID SIGMOIDOSCOPE, AND NO LEAK WAS OBSERVED. THE STAPLE LINE WAS INSPECTED ENDOSCOPICALLY VIA A RIGID SIGMOIDOSCOPE, AND THE STAPLE LINE WAS OBSERVED TO BE FULLY INTACT. THE SURGEON IS EXPERIENCED WITH THE STAPLER AND HAS BEEN USING THE DEVICE FOR 30 MONTHS.
27-11-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A COLOSTOMY REVERSAL PROCEDURE, THE SURGEON HAD DIFFICULTY REMOVING THE DEVICE POST FIRING. THE SURGEON DESCRIBED THE ISSUE AS "HAVING TO RIP THE STAPLER FROM THE PATIENT'S RECTUM". THE STAPLER WAS CLOSED WITH THE ORANGE INDICATOR AT THE BASE OF THE TISSUE COMPRESSION SCALE. THE SURGEON DESCRIBED CLOSING STAPLER AS ROUTINE. THE STAPLER FIRED NORMALLY, THE BREAKAWAY WASHER CRUNCH WAS HEARD, AND THE GREEN CHECK MARK INDICATED A COMPLETED FIRING CYCLE. THE STAPLER WAS OPENED 4 X 1/2 TURNS IN THE ROUTINE FASHION, AND THE DEVICE WAS ROTATED 90° IN BOTH DIRECTIONS. THE SURGEON NOTED THAT WHEN ROTATING THE STAPLER, THE BOWEL ROTATED WITH THE DEVICE. AFTER REMOVAL OF THE STAPLER, THE DONUTS WERE INSPECTED AND WERE OBSERVED TO BE FULLY FORMED AND INTACT. THE BREAKAWAY WASHER WAS OBSERVED WITH THE DONUTS IN THE USUAL FASHION. THE ANASTOMOSIS WAS TESTED WITH A ROUTINE AIR LEAK TEST VIA A RIGID SIGMOIDOSCOPE, AND NO LEAK WAS OBSERVED. THE STAPLE LINE WAS INSPECTED ENDOSCOPICALLY VIA A RIGID SIGMOIDOSCOPE, AND THE STAPLE LINE WAS OBSERVED TO BE FULLY INTACT. THE SURGEON IS EXPERIENCED WITH THE STAPLER AND HAS BEEN USING THE DEVICE FOR 30 MONTHS.
05-01-2026 | Malfunction | ethicon endo-surgery | 14-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
15-12-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN PROCEDURE, THE KNIFE HAD COME OUT WHEN OPENING THE PACKAGE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
05-01-2026 | Malfunction | ethicon endo-surgery | 14-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE BATTERY WAS INSERTED INTO THE STAPLER AND AN ATTEMPT WAS MADE TO THE FIRE THE DEVICE, HOWEVER THE DEVICE DIDN'T OPERATE. AND REMOVED AND REINSERTED THE BATTERY BUT IT STILL FAILED TO FUNCTION. THE PROCEDURE WAS NOT DELAYED AND COMPLETED SUCCESSFULLY. NO PATIENT INVOLVEMENT.
05-01-2026 | Malfunction | ethicon endo-surgery | 14-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | nan
05-01-2026 | Malfunction | ethicon endo-surgery | 14-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
05-01-2026 | Malfunction | ethicon endo-surgery | 14-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | nan
05-01-2026 | Malfunction | ethicon endo-surgery | 14-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | CIRCULAR MECH STD 29MM | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE ANVIL SIDE CAN'T CONNECT TO THE DEVICE. NO PATIENT CONSEQUENCES WERE REPORTED.
22-08-2025 | Malfunction | covidien surgical | 14-01-2026 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE POWER HANDLE WOULD NOT CHARGE DURING PRE-OPERATIVE PREPARATION. THE HANDLE WAS TESTED ON SEVERAL WORKING CHARGERS AND NONE ACTIVATED IT. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. THE HANDLE WAS OPENED UP AND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
22-08-2025 | Malfunction | covidien surgical | 14-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE POWER HANDLE WOULD NOT CHARGE DURING PRE-OPERATIVE PREPARATION. THE HANDLE WAS TESTED ON SEVERAL WORKING CHARGERS AND NONE ACTIVATED IT. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. THE HANDLE WAS OPENED UP AND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
16-12-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | nan
16-12-2025 | Malfunction | ethicon endo-surgery | 14-01-2026 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | nan
05-01-2026 | Malfunction | covidien surgical | 14-01-2026 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PULMONARY PARENCHYMA RESECTION, IN PARTIAL RESECTION OF THE LUNG, THE JAWS OPENING AND CLOSING COULD BE CONFIRMED WITHOUT ANY PROBLEMS, BUT THE JAWS DID NOT CLOSE WHEN APPROACHING THE TISSUE. TO RESOLVE THE ISSUE, A DEVICE FROM ANOTHER LOT NUMBER WAS USED. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien surgical | 14-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PULMONARY PARENCHYMA RESECTION, IN PARTIAL RESECTION OF THE LUNG, THE JAWS OPENING AND CLOSING COULD BE CONFIRMED WITHOUT ANY PROBLEMS, BUT THE JAWS DID NOT CLOSE WHEN APPROACHING THE TISSUE. TO RESOLVE THE ISSUE, A DEVICE FROM ANOTHER LOT NUMBER WAS USED. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien surgical | 14-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PULMONARY PARENCHYMA RESECTION, IN PARTIAL RESECTION OF THE LUNG, THE JAWS OPENING AND CLOSING COULD BE CONFIRMED WITHOUT ANY PROBLEMS, BUT THE JAWS DID NOT CLOSE WHEN APPROACHING THE TISSUE. TO RESOLVE THE ISSUE, A DEVICE FROM ANOTHER LOT NUMBER WAS USED. THERE WAS NO PATIENT INJURY.
22-12-2025 | Malfunction | covidien surgical | 14-01-2026 | MAXTACK | Separation Problem | C070603 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE TACKER MISFIRED WHEN FIRED AT AN INCORRECT ANGLE, RESULTING IN A LOOSE TACK. THE LOOSE TACK WAS SUBSEQUENTLY REMOVED FROM THE PATIENT AND DISCARDED. THERE WAS NO PATIENT INJURY.
23-12-2025 | Malfunction | covidien surgical | 14-01-2026 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A REMOVAL OF THE PANCREAS, THE YELLOW TAB WAS DAMAGED. PART OF THE YELLOW TAB WAS INCLUDED BECAUSE THE DAMAGED PART WAS INCLUDED. THE PART OF THE YELLOW TAB FELL INTO THE PATIENT'S CAVITY BUT WAS RETRIEVED. THERE WAS NO PATIENT INJURY.
20-10-2025 | Malfunction | covidien | 14-01-2026 | ENDO GIA ULTRA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC WEDGE RESECTION OF THE RIGHT LOWER LUNG LOBE, THE CUTTING STAPLER WAS UNABLE TO CUT COMPLETELY, MAKING IT IMPOSSIBLE TO RELEASE THE DISTAL END AND CONTINUE STAPLING. THE INSTRUMENT WAS REPLACED TO COMPLETE THE SURGERY, WHICH PROLONGED THE OPERATIVE TIME.
20-10-2025 | Malfunction | covidien | 14-01-2026 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC WEDGE RESECTION OF THE RIGHT LOWER LUNG LOBE, THE CUTTING STAPLER WAS UNABLE TO CUT COMPLETELY, MAKING IT IMPOSSIBLE TO RELEASE THE DISTAL END AND CONTINUE STAPLING. THE INSTRUMENT WAS REPLACED TO COMPLETE THE SURGERY, WHICH PROLONGED THE OPERATIVE TIME.
26-12-2025 | Malfunction | covidien surgical | 14-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE JAW OF THE DEVICE DID NOT CLOSE COMPLETELY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
26-12-2025 | Malfunction | covidien surgical | 14-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE JAW OF THE DEVICE DID NOT CLOSE COMPLETELY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
26-12-2025 | Malfunction | covidien surgical | 14-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE JAW OF THE DEVICE DID NOT CLOSE COMPLETELY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
12-11-2025 | Malfunction | covidien surgical | 15-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE RELOAD COULD NOT BE CONNECTED TO THE ADAPTER. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE DISTAL CARTRIDGE SUTURE DID NOT RELEASE.
12-11-2025 | Malfunction | covidien surgical | 15-01-2026 | TRI-STAPLE 2.0 | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE RELOAD COULD NOT BE CONNECTED TO THE ADAPTER. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE DISTAL CARTRIDGE SUTURE DID NOT RELEASE.
26-11-2025 | Malfunction | covidien | 15-01-2026 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE, PRIOR TO PATIENT USED, THE HANDLE BECAME BLOCKED, REQUIRING THE USE OF A SECOND DEVICE FROM THE SAME BATCH, WHICH FUNCTIONED CORRECTLY. THIS RESULTED IN AN INCREASE IN THE PROCEDURE DURATION BY APPROXIMATELY 20 MINUTES. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
02-12-2025 | Malfunction | covidien surgical | 15-01-2026 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE FIRING COULD NOT BE PERFORMED AT ALL. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-12-2025 | Malfunction | covidien surgical | 15-01-2026 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE FIRING COULD NOT BE PERFORMED AT ALL. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-12-2025 | Malfunction | covidien | 15-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE FIRING COULD NOT BE PERFORMED AT ALL. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | ECHELON | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLE WOULD NOT CLOSE AT FIRST AND WHEN IT DID, IT WOULD NOT UNCLAMP. THE STAPLE WAS STUCK CLAMPED ON THE PT'S INTESTINES FOR A FEW MINUTES.
12-12-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | ECHELON | Physical Resistance/Sticking | A0509 | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLE WOULD NOT CLOSE AT FIRST AND WHEN IT DID, IT WOULD NOT UNCLAMP. THE STAPLE WAS STUCK CLAMPED ON THE PT'S INTESTINES FOR A FEW MINUTES.
12-12-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | ECHELON | Failure to Seal | A050705 | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLE WOULD NOT CLOSE AT FIRST AND WHEN IT DID, IT WOULD NOT UNCLAMP. THE STAPLE WAS STUCK CLAMPED ON THE PT'S INTESTINES FOR A FEW MINUTES.
14-01-2026 | Malfunction | covidien surgical | 15-01-2026 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC TRANSABDOMINAL PREPERITONEAL (TAPP) HERNIA PROCEDURE, THE TACKER COULD NOT DEPLOY THE TACKS AS SMOOTHLY AND IT DOES NOT SECURE WELL ON THE DEFECT AREA. THE SOUND OF THE TACKER DURING DEPLOYING SOUNDS OFF. THE PHYSICIAN HAD TO SUTURE THE DEFECT AREA TO SECURE THE MESH TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
14-01-2026 | Malfunction | covidien surgical | 15-01-2026 | ABSORBATACK | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC TRANSABDOMINAL PREPERITONEAL (TAPP) HERNIA PROCEDURE, THE TACKER COULD NOT DEPLOY THE TACKS AS SMOOTHLY AND IT DOES NOT SECURE WELL ON THE DEFECT AREA. THE SOUND OF THE TACKER DURING DEPLOYING SOUNDS OFF. THE PHYSICIAN HAD TO SUTURE THE DEFECT AREA TO SECURE THE MESH TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
14-01-2026 | Malfunction | covidien surgical | 15-01-2026 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC TRANSABDOMINAL PREPERITONEAL (TAPP) HERNIA PROCEDURE, THE TACKER COULD NOT DEPLOY THE TACKS AS SMOOTHLY AND IT DOES NOT SECURE WELL ON THE DEFECT AREA. THE SOUND OF THE TACKER DURING DEPLOYING SOUNDS OFF. THE PHYSICIAN HAD TO SUTURE THE DEFECT AREA TO SECURE THE MESH TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-01-2024 | Injury | ethicon endo-surgery | 15-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AND DURING FIRING, WE SAW THE (STOMACH) TISSUE SHIFT ON THE LEFT SIDE OF THE CARTRIDGE. THE STAPLES ON THE RIGHT WERE NORMAL BUT DEFORMED ON THE LEFT. WE SUTURED THE "FAILED" STAPLE WITH STRATAFIX 3-0. AN ADDITIONAL BLUE TEST WAS PERFORMED, WHICH REVEALED LEAKAGE. TISSUECOL SPRAY WAS APPLIED TO THE STAPLE. THE PATIENT WAS GIVEN A NASOGASTRIC TUBE EXTENDING PAST THE STAPLE AND A LARGE DRAIN WAS INSERTED. IT IS NOT YET KNOWN WHETHER THE PATIENT WILL SUFFER PERMANENT DAMAGE. PATIENT STATUS/ OUTCOME / CONSEQUENCES --> YES. PATIENT CONSEQUENCE DESCRIPTION/WAS THERE A CLINICAL OUTCOME EXPERIENCED BY THE PATIENT (INFECTION, INFLAMMATION, ETC.)? --> PROLONGED PROCEDURE WITH INTERVENTIONS TO FIX STAPLE LINE.
01-01-2024 | Injury | ethicon endo-surgery | 15-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Cut | A050702 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AND DURING FIRING, WE SAW THE (STOMACH) TISSUE SHIFT ON THE LEFT SIDE OF THE CARTRIDGE. THE STAPLES ON THE RIGHT WERE NORMAL BUT DEFORMED ON THE LEFT. WE SUTURED THE "FAILED" STAPLE WITH STRATAFIX 3-0. AN ADDITIONAL BLUE TEST WAS PERFORMED, WHICH REVEALED LEAKAGE. TISSUECOL SPRAY WAS APPLIED TO THE STAPLE. THE PATIENT WAS GIVEN A NASOGASTRIC TUBE EXTENDING PAST THE STAPLE AND A LARGE DRAIN WAS INSERTED. IT IS NOT YET KNOWN WHETHER THE PATIENT WILL SUFFER PERMANENT DAMAGE. PATIENT STATUS/ OUTCOME / CONSEQUENCES --> YES. PATIENT CONSEQUENCE DESCRIPTION/WAS THERE A CLINICAL OUTCOME EXPERIENCED BY THE PATIENT (INFECTION, INFLAMMATION, ETC.)? --> PROLONGED PROCEDURE WITH INTERVENTIONS TO FIX STAPLE LINE.
01-01-2024 | Injury | ethicon endo-surgery | 15-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AND DURING FIRING, WE SAW THE (STOMACH) TISSUE SHIFT ON THE LEFT SIDE OF THE CARTRIDGE. THE STAPLES ON THE RIGHT WERE NORMAL BUT DEFORMED ON THE LEFT. WE SUTURED THE "FAILED" STAPLE WITH STRATAFIX 3-0. AN ADDITIONAL BLUE TEST WAS PERFORMED, WHICH REVEALED LEAKAGE. TISSUECOL SPRAY WAS APPLIED TO THE STAPLE. THE PATIENT WAS GIVEN A NASOGASTRIC TUBE EXTENDING PAST THE STAPLE AND A LARGE DRAIN WAS INSERTED. IT IS NOT YET KNOWN WHETHER THE PATIENT WILL SUFFER PERMANENT DAMAGE. PATIENT STATUS/ OUTCOME / CONSEQUENCES --> YES. PATIENT CONSEQUENCE DESCRIPTION/WAS THERE A CLINICAL OUTCOME EXPERIENCED BY THE PATIENT (INFECTION, INFLAMMATION, ETC.)? --> PROLONGED PROCEDURE WITH INTERVENTIONS TO FIX STAPLE LINE.
22-12-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RECTUM PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-01-2025 | Malfunction | covidien | 15-01-2026 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A MINIMALLY INVASIVE SURGERY, THE RETURN KNOB BROKE AND DISENGAGED. THE ISSUE WAS NOTICED WHEN THE SURGEON HAD ALREADY FIRED THE DEVICE AND WAS ABOUT TO RETURN THE KNIFE AND OPEN THE INSTRUMENT. THE SURGEON USED A FORCEPS AND PUSHED THE LEVER INSIDE THE HANDLE BACK SO THE RELOAD OPENED. ONE PIECE OF THE MISSING KNOB WAS FOUND ON THE FLOOR AND THE OTHER ON THE PATIENT. THERE WAS NO PATIENT INJURY.
01-01-2026 | Injury | ethicon endo-surgery | 15-01-2026 | POWERED PLUS LONG 45 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST OP TO A BYPASS PROCEDURE THERE WAS A PATIENT WHO HAD BLEEDING INTO THE STOMACH. I ASKED HIM TO TELL ME MORE ABOUT HIS PERCEPTIONS ABOUT THE CAUSE OF BLEEDING AND HE STARTED BY SAYING THAT HE HAS BEEN DOING BARIATRIC SURGERY FOR MORE THAN 20 YEARS AND STARTED USING OUR ATS, BUT HE BELIEVES THAT THE PROBLEM WAS BECAUSE JOHNSON DID NOT EVOLVE TO THE TRI-STAPLE SYSTEM WHERE HE BELIEVED THIS WAS THE REASON BECAUSE HE DOES NOT HAVE THIS STATISTIC WITH MATERIALS WITH THIS CHARACTERISTIC. FIRST, I ASKED HIM FOR PATIENT DATA AND A REPORT ON WHAT HAPPENED SO THAT WE COULD LOCATE THE MATERIALS, ALTHOUGH IT IS NOT POSSIBLE TO ANALYZE THE MATERIALS INVOLVED DUE TO DISPOSAL, WE WILL MAKE A REPORT TO THE COMPETENT BODIES AND WE COULD TRACK PROBLEMS REPORTED IN OTHER LOCATIONS OF THE SAME LOT. I WENT ON TO INFORM THAT IN THE SURGERY IN WHICH HIS ASSISTANT THAT WAS TAKING PLACE CONCOMITANTLY IN ANOTHER ROOM, I SIGNALED THAT THE DOCTOR WAS NOT RESPECTING THE PROXIMAL LIMIT OF THE TISSUE IN THE STAPLER'S JAW, WHERE TISSUE WAS BEYOND MARKING, WHERE HE WOULD BE AT RISK OF CUTTING WITHOUT A STAPLE, DR. CONFIRMED DURING THIS CONVERSATION THAT THE TEAM'S CONCERN IS REGARDING THE AMOUNT OF PACKAGE LOADS, AS THERE ARE ONLY 6 LOADS, THEY TRY TO SAVE ON STAPLING. I INFORMED DR. THAT THE HOSPITAL DID NOT REFUSE IN ANY PROCEDURE THAT REQUIRED THE SENDING OF MORE LOADS AND THAT EVEN OR TWO OTHER PROCEDURES THAT WOULD COME IN SEQUENCE ALREADY HAD 2 EXTRA LOADS PER PATIENT. DR. SAID THAT THIS COULD BE THE REASON FOR THE BLEEDING AND WHEN I ASKED FOR THE PATIENTS' INFORMATION, HE SAID THAT HE DID NOT WANT TO TOUCH THIS HORNET'S NEST AND THAT THEY WOULD FOLLOW UP ON FUTURE CASES. AS AN ACTION PLAN, WE ARE ALIGNING A NEW TRAINING AGENDA WITH HIS TEAM IN THE CETEB ROOM AND SUGGESTED THE NEED TO INVOLVE DR. IN A CONVERSATION WITH THE SURGEON'S TEAM.
07-01-2026 | Injury | ethicon endo-surgery | 15-01-2026 | POWERED PLUS LONG 45 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST OP TO A BYPASS PROCEDURE THERE WAS A PATIENT WHO HAD BLEEDING INTO THE STOMACH. I ASKED HIM TO TELL ME MORE ABOUT HIS PERCEPTIONS ABOUT THE CAUSE OF BLEEDING AND HE STARTED BY SAYING THAT HE HAS BEEN DOING BARIATRIC SURGERY FOR MORE THAN 20 YEARS AND STARTED USING OUR ATS, BUT HE BELIEVES THAT THE PROBLEM WAS BECAUSE JOHNSON DID NOT EVOLVE TO THE TRI-STAPLE SYSTEM WHERE HE BELIEVED THIS WAS THE REASON BECAUSE HE DOES NOT HAVE THIS STATISTIC WITH MATERIALS WITH THIS CHARACTERISTIC. FIRST, I ASKED HIM FOR PATIENT DATA AND A REPORT ON WHAT HAPPENED SO THAT WE COULD LOCATE THE MATERIALS, ALTHOUGH IT IS NOT POSSIBLE TO ANALYZE THE MATERIALS INVOLVED DUE TO DISPOSAL, WE WILL MAKE A REPORT TO THE COMPETENT BODIES AND WE COULD TRACK PROBLEMS REPORTED IN OTHER LOCATIONS OF THE SAME LOT. I WENT ON TO INFORM THAT IN THE SURGERY IN WHICH HIS ASSISTANT THAT WAS TAKING PLACE CONCOMITANTLY IN ANOTHER ROOM, I SIGNALED THAT THE DOCTOR WAS NOT RESPECTING THE PROXIMAL LIMIT OF THE TISSUE IN THE STAPLER'S JAW, WHERE TISSUE WAS BEYOND MARKING, WHERE HE WOULD BE AT RISK OF CUTTING WITHOUT A STAPLE, DR. CONFIRMED DURING THIS CONVERSATION THAT THE TEAM'S CONCERN IS REGARDING THE AMOUNT OF PACKAGE LOADS, AS THERE ARE ONLY 6 LOADS, THEY TRY TO SAVE ON STAPLING. I INFORMED DR. THAT THE HOSPITAL DID NOT REFUSE IN ANY PROCEDURE THAT REQUIRED THE SENDING OF MORE LOADS AND THAT EVEN OR TWO OTHER PROCEDURES THAT WOULD COME IN SEQUENCE ALREADY HAD 2 EXTRA LOADS PER PATIENT. DR. SAID THAT THIS COULD BE THE REASON FOR THE BLEEDING AND WHEN I ASKED FOR THE PATIENTS' INFORMATION, HE SAID THAT HE DID NOT WANT TO TOUCH THIS HORNET'S NEST AND THAT THEY WOULD FOLLOW UP ON FUTURE CASES. AS AN ACTION PLAN, WE ARE ALIGNING A NEW TRAINING AGENDA WITH HIS TEAM IN THE CETEB ROOM AND SUGGESTED THE NEED TO INVOLVE DR. IN A CONVERSATION WITH THE SURGEON'S TEAM.
01-12-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE GASTRECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-01-2026 | Malfunction | covidien surgical | 15-01-2026 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA MESH IMPLANTATION, THE THREE RELOADS WERE DIFFICULT TO LOAD ONTO THE ARTICUL ATING TIP OF THE HANDLE. THE DEVICES WERE STILL ABLE TO BE USED ON THE PATIENT. THREE NEW RELOADS WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-12-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE COMBINED PANCREATICODUODENECTOMY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2026 | Injury | ethicon endo-surgery | 15-01-2026 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | Restenosis | nan
09-01-2026 | Malfunction | covidien surgical | 15-01-2026 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA MESH IMPLANTATION, THE THREE RELOADS WERE DIFFICULT TO LOAD ONTO THE ARTICUL ATING TIP OF THE HANDLE. THE DEVICES WERE STILL ABLE TO BE USED ON THE PATIENT. THREE NEW RELOADS WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2026 | Malfunction | covidien surgical | 15-01-2026 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA MESH IMPLANTATION, THE THREE RELOADS WERE DIFFICULT TO LOAD ONTO THE ARTICUL ATING TIP OF THE HANDLE. THE DEVICES WERE STILL ABLE TO BE USED ON THE PATIENT. THREE NEW RELOADS WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-12-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SECONDARY REDUCTION SURGERY FOR OVARIAN CANCER, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-12-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ILEOSTOMY RETURN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-12-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RESECTION OF RETROPERITONEAL MASS, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-01-2026 | Malfunction | covidien | 15-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLORECTAL RESECTION, DURING THE TISSUE RESECTION STEP, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
14-01-2026 | Malfunction | covidien | 15-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLORECTAL RESECTION, DURING THE TISSUE RESECTION STEP, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
28-12-2025 | Death | covidien surgical | 15-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Organ Dehiscence | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SURGICAL PROCEDURE FOR INTESTINAL TRANSIT RECONSTRUCTION USING A LINEAR STAPLER AND TWO ADDITIONAL LINEAR RELOADS ON (B)(6) 2025. SURGICAL RE-APPROACH WAS DONE ON (B)(6) 2025 BY FULL OPENING IN THE STAPLE LINE AND PERFORMED SEGMENTAL ENTERECTOMY OF TRANSVERSE ILEUM ANASTOMOSIS USING LINEAR STAPLER IN DISTAL STUMP AND SCALPEL SECTION IN PROXIMAL STUMP DUE TO DEHISCENCE. PATIENT PROGRESSED TO DEATH ON (B)(6) 2025.
28-12-2025 | Death | covidien surgical | 15-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SURGICAL PROCEDURE FOR INTESTINAL TRANSIT RECONSTRUCTION USING A LINEAR STAPLER AND TWO ADDITIONAL LINEAR RELOADS ON (B)(6) 2025. SURGICAL RE-APPROACH WAS DONE ON (B)(6) 2025 BY FULL OPENING IN THE STAPLE LINE AND PERFORMED SEGMENTAL ENTERECTOMY OF TRANSVERSE ILEUM ANASTOMOSIS USING LINEAR STAPLER IN DISTAL STUMP AND SCALPEL SECTION IN PROXIMAL STUMP DUE TO DEHISCENCE. PATIENT PROGRESSED TO DEATH ON (B)(6) 2025.
09-01-2026 | Malfunction | covidien surgical | 15-01-2026 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTECTOMY PROCEDURE TYPICAL OF THE CULMEN OR UPPER LOBE OF THE LEFT UPPER LOBE USING THE SINGLE-PORT TECHNIQUE WITH A WOUND PROTECTION SYSTEM FROM ANOTHER MANUFACTURER, THE SURGEON DETECTED AN ANOMALOUS CONFIGURATION OF THE STAPLES OF THE THREE VASCULAR CARTRIDGES. TWO ON THE ARTERIAL BRANCH AND ONE ON THE VENOUS BRANCH. IT WAS NOTED THAT A POWERED HANDLE WAS USED THEN A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE.
09-01-2026 | Malfunction | covidien surgical | 15-01-2026 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTECTOMY PROCEDURE TYPICAL OF THE CULMEN OR UPPER LOBE OF THE LEFT UPPER LOBE USING THE SINGLE-PORT TECHNIQUE WITH A WOUND PROTECTION SYSTEM FROM ANOTHER MANUFACTURER, THE SURGEON DETECTED AN ANOMALOUS CONFIGURATION OF THE STAPLES OF THE THREE VASCULAR CARTRIDGES. TWO ON THE ARTERIAL BRANCH AND ONE ON THE VENOUS BRANCH. IT WAS NOTED THAT A POWERED HANDLE WAS USED THEN A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE.
09-01-2026 | Malfunction | covidien surgical | 15-01-2026 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTECTOMY PROCEDURE TYPICAL OF THE CULMEN OR UPPER LOBE OF THE LEFT UPPER LOBE USING THE SINGLE-PORT TECHNIQUE WITH A WOUND PROTECTION SYSTEM FROM ANOTHER MANUFACTURER, THE SURGEON DETECTED AN ANOMALOUS CONFIGURATION OF THE STAPLES OF THE THREE VASCULAR CARTRIDGES. TWO ON THE ARTERIAL BRANCH AND ONE ON THE VENOUS BRANCH. IT WAS NOTED THAT A POWERED HANDLE WAS USED THEN A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE.
01-11-2025 | Malfunction | covidien | 15-01-2026 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHEN DETACHING THE LUNG PARENCHYMA, THERE WAS INCOMPLETE SUTURING AND POOR STAPLE FORMATION. ADDITIONALLY, THERE WAS ABNORMAL NOISES FROM THE HANDLE. ADDITIONAL SUTURING WITH A THREAD WAS PERFORMED TO FIX THE STAPLE LINE. A NEW HANDLE WAS THEN USED TO COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
01-11-2025 | Malfunction | covidien | 15-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHEN DETACHING THE LUNG PARENCHYMA, THERE WAS INCOMPLETE SUTURING AND POOR STAPLE FORMATION. ADDITIONALLY, THERE WAS ABNORMAL NOISES FROM THE HANDLE. ADDITIONAL SUTURING WITH A THREAD WAS PERFORMED TO FIX THE STAPLE LINE. A NEW HANDLE WAS THEN USED TO COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
06-01-2026 | Malfunction | johnson & johnson medical (suzhou | 15-01-2026 | FINAL ASSY 445MM STD DEVICE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 15-01-2026 | RELOAD- CURVED CUTTER STAPLER | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE BLADE DOES NOT ADVANCE TO THE END. NO PATIENT CONSEQUENCES WERE REPORTED.
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 15-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
15-01-2026 | Malfunction | covidien surgical | 15-01-2026 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INGUINAL HERNIA PROCEDURE, WHILE FIXING THE MESH, THE TACKER WOULD NOT LOAD PROPERLY, REQUIRING THE USER TO PRESSED MULTIPLE TIMES FOR THE TACKER TO FIRE. ADDITIONALLY, THE TACKS WERE NOT FIXED ON THE MESH PROPERLY AND FELL INTO THE PATIENT'S CAVITY. THE TACKS THAT FELL WERE RETRIEVED. THIS ISSUE OCCURRED ON TWO DEVICES. TO COMPLETE THE CASE, TWO MORE TACKERS WERE USED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
15-01-2026 | Malfunction | covidien surgical | 15-01-2026 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INGUINAL HERNIA PROCEDURE, WHILE FIXING THE MESH, THE TACKER WOULD NOT LOAD PROPERLY, REQUIRING THE USER TO PRESSED MULTIPLE TIMES FOR THE TACKER TO FIRE. ADDITIONALLY, THE TACKS WERE NOT FIXED ON THE MESH PROPERLY AND FELL INTO THE PATIENT'S CAVITY. THE TACKS THAT FELL WERE RETRIEVED. THIS ISSUE OCCURRED ON TWO DEVICES. TO COMPLETE THE CASE, TWO MORE TACKERS WERE USED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
09-12-2025 | Malfunction | covidien | 15-01-2026 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, WHILE ON SET-UP, THE DEVICE'S GREEN BUTTON DID NOT MOVE. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
26-11-2025 | Malfunction | covidien surgical | 15-01-2026 | nan | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE, PRIOR TO PATIENT USED, THE HANDLE BECAME BLOCKED, REQUIRING THE USE OF A SECOND DEVICE FROM THE SAME BATCH, WHICH FUNCTIONED CORRECTLY. THIS RESULTED IN AN INCREASE IN THE PROCEDURE DURATION BY APPROXIMATELY 20 MINUTES. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
09-12-2025 | Malfunction | covidien | 15-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RADICAL LUNG CANCER SURGERY, THE STAPLE RELOAD DID NOT FORM PROPERLY. A NEW RELOAD WAS USED TO RESOLE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2025 | Malfunction | covidien surgical | 15-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT SUPERIOR RESECTION WITH ROBOT SUPPORT PROCEDURE, AT THE LEFT SUPERIOR LOBE BRONCHI, THE STAPLE FORMATION WAS INCOMPLETE AND THE FIRING STOPPED IN THE MIDDLE.
24-12-2025 | Malfunction | covidien surgical | 15-01-2026 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT SUPERIOR RESECTION WITH ROBOT SUPPORT PROCEDURE, AT THE LEFT SUPERIOR LOBE BRONCHI, THE STAPLE FORMATION WAS INCOMPLETE AND THE FIRING STOPPED IN THE MIDDLE.
24-12-2025 | Malfunction | covidien surgical | 15-01-2026 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT SUPERIOR RESECTION WITH ROBOT SUPPORT PROCEDURE, AT THE LEFT SUPERIOR LOBE BRONCHI, THE STAPLE FORMATION WAS INCOMPLETE AND THE FIRING STOPPED IN THE MIDDLE.
22-12-2025 | Injury | intuitive surgical | 16-01-2026 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED APPENDECTOMY PROCEDURE, BLEEDING FROM THE APPENDIX OCCURRED AFTER FIRING A SUREFORM 30 CURVED-TIP STAPLER INSTRUMENT EQUIPPED WITH A SUREFORM 30 WHITE RELOAD. THE INITIAL FIRE WITH A SUREFORM 30 BLUE RELOAD WAS SUCCESSFUL AND UNEVENTFUL. THE WHITE STAPLER RELOAD WAS THEN FIRED ON TARGET ANATOMY AS EXPECTED; HOWEVER, THE RESULTING STAPLE LINE WAS OBSERVED TO BE BLEEDING. ALTHOUGH THE BLEEDING WAS MINIMAL, INTERVENTION WAS REQUIRED. HEMOSTASIS WAS ACHIEVED USING A FENESTRATED BIPOLAR FORCEPS INSTRUMENT TO CAUTERIZE THE TISSUE. THE PROCEDURE WAS COMPLETED ROBOTICALLY WITHOUT FURTHER COMPLICATION. THE PATIENT EXPERIENCED NO POSTOPERATIVE ISSUES, AND THERE ARE NO CONCERNS FOR LONG-TERM COMPLICATIONS.
24-12-2025 | Malfunction | covidien | 16-01-2026 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, AFTER THE FIRING WAS COMPLETED, THE KNIFE BLADE COULD NOT BE RETRACTED. EVEN AFTER RESETTING THE POWER AND REMOVING AND REINSTALLING THE HANDLE, IT STILL DID NOT WORK. THE MANUAL ADAPTER TOOL WAS USED AND ROTATED SEVERAL TIMES, BUT THE BLADE DID NOT APPEAR TO RETRACT. THEREFORE, THE HANDLE WAS REATTACHED, AND THEY WERE ABLE TO REPLACE THE BLADE AND DETACH IT FROM THE TISSUE. ANOTHER HANDLE AND ADAPTER WERE USED FOR THE SUBSEQUENT RESECTION.
24-12-2025 | Malfunction | covidien | 16-01-2026 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, AFTER THE FIRING WAS COMPLETED, THE KNIFE BLADE COULD NOT BE RETRACTED. EVEN AFTER RESETTING THE POWER AND REMOVING AND REINSTALLING THE HANDLE, IT STILL DID NOT WORK. THE MANUAL ADAPTER TOOL WAS USED AND ROTATED SEVERAL TIMES, BUT THE BLADE DID NOT APPEAR TO RETRACT. THEREFORE, THE HANDLE WAS REATTACHED, AND THEY WERE ABLE TO REPLACE THE BLADE AND DETACH IT FROM THE TISSUE. ANOTHER HANDLE AND ADAPTER WERE USED FOR THE SUBSEQUENT RESECTION.
24-12-2025 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, AFTER THE FIRING WAS COMPLETED, THE KNIFE BLADE COULD NOT BE RETRACTED. EVEN AFTER RESETTING THE POWER AND REMOVING AND REINSTALLING THE HANDLE, IT STILL DID NOT WORK. THE MANUAL ADAPTER TOOL WAS USED AND ROTATED SEVERAL TIMES, BUT THE BLADE DID NOT APPEAR TO RETRACT. THEREFORE, THE HANDLE WAS REATTACHED, AND THEY WERE ABLE TO REPLACE THE BLADE AND DETACH IT FROM THE TISSUE. ANOTHER HANDLE AND ADAPTER WERE USED FOR THE SUBSEQUENT RESECTION.
24-12-2025 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, AFTER THE FIRING WAS COMPLETED, THE KNIFE BLADE COULD NOT BE RETRACTED. EVEN AFTER RESETTING THE POWER AND REMOVING AND REINSTALLING THE HANDLE, IT STILL DID NOT WORK. THE MANUAL ADAPTER TOOL WAS USED AND ROTATED SEVERAL TIMES, BUT THE BLADE DID NOT APPEAR TO RETRACT. THEREFORE, THE HANDLE WAS REATTACHED, AND THEY WERE ABLE TO REPLACE THE BLADE AND DETACH IT FROM THE TISSUE. ANOTHER HANDLE AND ADAPTER WERE USED FOR THE SUBSEQUENT RESECTION.
26-12-2025 | Malfunction | covidien surgical | 16-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT INTRAOPERATIVELY, THE PRODUCT DID NOT WORK AS INTENDED. THE RELOAD WAS REPLACED WITH A NEW ONE, AND THE SURGERY PROCEEDED SUCCESSFULLY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
26-12-2025 | Malfunction | covidien surgical | 16-01-2026 | nan | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT INTRAOPERATIVELY, THE PRODUCT DID NOT WORK AS INTENDED. THE RELOAD WAS REPLACED WITH A NEW ONE, AND THE SURGERY PROCEEDED SUCCESSFULLY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
26-12-2025 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT INTRAOPERATIVELY, THE PRODUCT DID NOT WORK AS INTENDED. THE RELOAD WAS REPLACED WITH A NEW ONE, AND THE SURGERY PROCEEDED SUCCESSFULLY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
08-01-2026 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Unraveled Material | A040508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE ISSUE OF BEING UNABLE TO CLOSE POWERSHELL WAS ENCOUNTERED PREOPERATIVELY. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
09-11-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | ECHELON FLEX | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLER WAS FIRED. UPON TRYING TO RELEASE, STAPLER WOULD NOT OPEN. USED EMERGENT RED RELEASE BUTTON. HAD TO REPEAT AND FINALLY IT OPENED.
08-12-2025 | Injury | intuitive surgical | 16-01-2026 | SUREFORM | Firing Problem | A0505 | Hematoma | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, THE PATIENT WAS READMITTED TO THE HOSPITAL DUE TO COMPLICATIONS. THE GASTRIC BYPASS WAS COMPLETED ROBOTICALLY, AND A SUREFORM 60 WHITE RELOAD STAPLE LINE WAS NOTED TO BE BLEEDING AFTER SUCCESSFUL STAPLE LINE. THE SURGEON ENSURED THE STAPLE LINE WAS HEMOSTATIC DURING THE PROCEDURE, INCLUDING OVER-SUTURING THE STAPLE LINE. POSTOPERATIVELY, THE PATIENT PRESENTED TO THE EMERGENCY ROOM, AND A HEMATOMA WAS IDENTIFIED. THE PATIENT'S HEMOGLOBIN LEVEL DROPPED, AND A COMPUTED TOMOGRAPHY ANGIOGRAPHY WAS PERFORMED, WHICH DID NOT REVEAL ANY ACTIVE BLEEDING OR EMBOLIZATION. THE PATIENT WAS HOSPITALIZED FOR A FEW DAYS TO MONITOR THEIR HEMOGLOBIN LEVELS. THE PATIENT IS CURRENTLY DISCHARGED AND RECOVERING WELL.
08-12-2025 | Injury | intuitive surgical | 16-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, THE PATIENT WAS READMITTED TO THE HOSPITAL DUE TO COMPLICATIONS. THE GASTRIC BYPASS WAS COMPLETED ROBOTICALLY, AND A SUREFORM 60 WHITE RELOAD STAPLE LINE WAS NOTED TO BE BLEEDING AFTER SUCCESSFUL STAPLE LINE. THE SURGEON ENSURED THE STAPLE LINE WAS HEMOSTATIC DURING THE PROCEDURE, INCLUDING OVER-SUTURING THE STAPLE LINE. POSTOPERATIVELY, THE PATIENT PRESENTED TO THE EMERGENCY ROOM, AND A HEMATOMA WAS IDENTIFIED. THE PATIENT'S HEMOGLOBIN LEVEL DROPPED, AND A COMPUTED TOMOGRAPHY ANGIOGRAPHY WAS PERFORMED, WHICH DID NOT REVEAL ANY ACTIVE BLEEDING OR EMBOLIZATION. THE PATIENT WAS HOSPITALIZED FOR A FEW DAYS TO MONITOR THEIR HEMOGLOBIN LEVELS. THE PATIENT IS CURRENTLY DISCHARGED AND RECOVERING WELL.
20-10-2025 | Malfunction | covidien surgical | 16-01-2026 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC WEDGE RESECTION OF THE RIGHT LOWER LUNG LOBE, THE CUTTING STAPLER WAS UNABLE TO CUT COMPLETELY, MAKING IT IMPOSSIBLE TO RELEASE THE DISTAL END AND CONTINUE STAPLING. THE INSTRUMENT WAS REPLACED TO COMPLETE THE SURGERY, WHICH PROLONGED THE OPERATIVE TIME.
20-10-2025 | Malfunction | covidien surgical | 16-01-2026 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC WEDGE RESECTION OF THE RIGHT LOWER LUNG LOBE, THE CUTTING STAPLER WAS UNABLE TO CUT COMPLETELY, MAKING IT IMPOSSIBLE TO RELEASE THE DISTAL END AND CONTINUE STAPLING. THE INSTRUMENT WAS REPLACED TO COMPLETE THE SURGERY, WHICH PROLONGED THE OPERATIVE TIME.
18-08-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | ECHELON FLEX 45 STANDARD | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS MISFIRING WHILE INSIDE THE PATIENT; AND NOT WORKING PROPERLY. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE FROZE AND ITS KNIFE DID NOT RETURN BACK AFTER FIRING THEY TRIED THE REVERSE BUTTON BUT IT DID NOT WORK; AND EVENTUALLY, HAD TO USE THE MANUAL OVER RIDE TO RETURN THE KNIFE AND OPEN THE GUN. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE FROZE AND ITS KNIFE DID NOT RETURN BACK AFTER FIRING THEY TRIED THE REVERSE BUTTON BUT IT DID NOT WORK; AND EVENTUALLY, HAD TO USE THE MANUAL OVER RIDE TO RETURN THE KNIFE AND OPEN THE GUN. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE FROZE AND ITS KNIFE DID NOT RETURN BACK AFTER FIRING THEY TRIED THE REVERSE BUTTON BUT IT DID NOT WORK; AND EVENTUALLY, HAD TO USE THE MANUAL OVER RIDE TO RETURN THE KNIFE AND OPEN THE GUN. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
06-01-2026 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PULMONARY LOBECTOMY USING THE POWERED STAPLER, AFTER PRESSING THE GREEN BUTTON, THE DEVICE DID NOT FIRE DESPITE ALL PREPARATIONS BEING COMPLETED. THE RELOAD WAS REPLACED WITH ANOTHER SMALL DIAMETER RELOAD, WHICH WAS FIRED WITHOUT ANY ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2025 | Malfunction | covidien surgical | 16-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE POWERED STAPLER DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2025 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, THE ARTICULATION PART OF THE CARTRIDGE WAS FOUND TO BE UNSTEADY AND WOBBLING MORE THAN USUAL PRIOR TO USE, SHOWING A SIGNIFICANT DIFFERENCE COMPARED TO A NEW CARTRIDGE. HOWEVER, FIRING WAS NOT PERFORMED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2025 | Malfunction | covidien surgical | 16-01-2026 | TRI-STAPLE 2.0 | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, THE ARTICULATION PART OF THE CARTRIDGE WAS FOUND TO BE UNSTEADY AND WOBBLING MORE THAN USUAL PRIOR TO USE, SHOWING A SIGNIFICANT DIFFERENCE COMPARED TO A NEW CARTRIDGE. HOWEVER, FIRING WAS NOT PERFORMED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2025 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, THE ARTICULATION PART OF THE CARTRIDGE WAS FOUND TO BE UNSTEADY AND WOBBLING MORE THAN USUAL PRIOR TO USE, SHOWING A SIGNIFICANT DIFFERENCE COMPARED TO A NEW CARTRIDGE. HOWEVER, FIRING WAS NOT PERFORMED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2025 | Malfunction | covidien | 16-01-2026 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE MANUAL HANDLE DID NOT OPEN DURING SETUP. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien surgical | 16-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, THE RELOAD DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
15-01-2026 | Malfunction | covidien surgical | 16-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC RIGHT LOWER LOBECTOMY, DURING THE TRANSECTION OF THE BRONCHUS, AFTER PARTIAL FIRING, THE HANDLE WAS PRESSED BUT THE STAPLER HAD NO RESPONSE. THE RELOAD HAD A FAILURE TO FIRE TO THE DISTAL END. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-01-2026 | Malfunction | covidien | 16-01-2026 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC RIGHT LOWER LOBECTOMY, DURING THE TRANSECTION OF THE BRONCHUS, AFTER PARTIAL FIRING, THE HANDLE WAS PRESSED BUT THE STAPLER HAD NO RESPONSE. THE RELOAD HAD A FAILURE TO FIRE TO THE DISTAL END. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-11-2025 | Malfunction | covidien surgical | 16-01-2026 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO LITERATURE, A PROSPECTIVE STUDY CONDUCTED BETWEEN JANUARY 2022 AND DECEMBER 2024 EVALUATED 36 PATIENTS UNDERGOING ELECTIVE VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY UNDER SPONTANEOUS VENTILATION WITH INTUBATION, IN WHICH ENDO GIA¿ STAPLERS WERE UTILIZED FOR DISSECTION AND DIVISION OF THE PULMONARY ARTERY, PULMONARY VEIN, FISSURES, AND BRONCHUS. THE STUDY AIMED TO FILL THIS GAP BY CONDUCTING A TWO-COHORT ANALYSIS OF PATIENTS UNDERGOING VATS LOBECTOMY WITH SPONTANEOUS VENTILATION AND INTUBATION (SVI): ONE GROUP WITH CONFIRMED CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND A MATCHED NON-COPD CONTROL GROUP. ALL PROCEDURES WERE COMPLETED SUCCESSFULLY USING A STANDARDIZED UNIPORTAL VATS TECHNIQUE, WITH SYSTEMATIC MEDIASTINAL LYMPH NODE DISSECTION PERFORMED IN ALL CASES. THE STUDY DEMONSTRATED THAT THE ENDO GIA INSTRUMENT PROVIDED RELIABLE STAPLE LINE FORMATION AND EFFECTIVE TISSUE MANAGEMENT DURING CRITICAL PULMONARY RESECTIONS, WITH NO DEVICE-RELATED MALFUNCTIONS OR INTRAOPERATIVE COMPLICATIONS REPORTED. POSTOPERATIVELY, THE MOST COMMONLY OBSERVED COMPLICATION WAS PROLONGED AIR LEAK, WHILE OTHER EVENTS SUCH AS MILD ATELECTASIS AND TRANSIENT ATRIAL FIBRILLATION WERE NOTED AND DEEMED UNRELATED TO THE DEVICE. IMPORTANTLY, NO CONVERSIONS TO THORACOTOMY OR CONTROLLED MECHANICAL VE NTILATION WERE REQUIRED, ALL PATIENTS WERE EXTUBATED IN THE OPERATING ROOM, AND THERE WAS NO 30-DAY MORTALITY. OVERALL OUTCOMES INDICATED STABLE INTRAOPERATIVE PERFORMANCE, ACCEPTABLE COMPLICATION RATES, AND FAVORABLE PATIENT RECOVERY, SUPPORTING THE SAFETY AND EF FECTIVENESS OF ENDO GIA STAPLERS IN MINIMALLY INVASIVE THORACIC SURGERY SETTINGS.
28-12-2025 | Malfunction | covidien surgical | 16-01-2026 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VENTRAL HERNIA (IPOM) PROCEDURE WITH MESH FIXATION, THE FIRST TACKER THE FIRING MECHANISM REPEATEDLY GOT STUCK, AND DURING ATTEMPTED FIXATION, TEARING OF THE MESH OCCURRED. ADDITIONALLY, 4¿5 TACKS CAME OUT FROM THE SHAFT WITHOUT GETTING FIXED INTO THE TISSUE. A SECOND TACKER WAS OPENED, BUT SIMILAR PROBLEMS WERE OBSERVED, INCLUDING THE HANDLE GETTING LOCKED MULTIPLE TIMES AND THE SURGEON BEING ABLE TO FIRE ONLY 7¿8 TACKS OUT OF 15. THESE ISSUES LED TO PROCEDURAL DIFFICULTY AND DISSATISFACTION DURING MESH FIXATION. INTRACORPOREAL SUTURING WAS PERFORMED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
28-12-2025 | Malfunction | covidien surgical | 16-01-2026 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VENTRAL HERNIA (IPOM) PROCEDURE WITH MESH FIXATION, THE FIRST TACKER THE FIRING MECHANISM REPEATEDLY GOT STUCK, AND DURING ATTEMPTED FIXATION, TEARING OF THE MESH OCCURRED. ADDITIONALLY, 4¿5 TACKS CAME OUT FROM THE SHAFT WITHOUT GETTING FIXED INTO THE TISSUE. A SECOND TACKER WAS OPENED, BUT SIMILAR PROBLEMS WERE OBSERVED, INCLUDING THE HANDLE GETTING LOCKED MULTIPLE TIMES AND THE SURGEON BEING ABLE TO FIRE ONLY 7¿8 TACKS OUT OF 15. THESE ISSUES LED TO PROCEDURAL DIFFICULTY AND DISSATISFACTION DURING MESH FIXATION. INTRACORPOREAL SUTURING WAS PERFORMED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-01-2025 | Injury | ethicon endo-surgery | 16-01-2026 | ECHELON CONTOUR GST RELOAD GREEN | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, ECHELON CONTOUR CURVED CUTTER MISFIRED - THIS CAUSED SURGEONS TO HAVE TO CHANGE SURGICAL PLAN FROM RE-JOINING BOWEL TO STOMA FORMATION.
03-11-2025 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LIVER RESECTION, THE DEVICE WOULD ONLY ANGULATE BUT WOULD NOT FIRE. IT WAS NOTED THAT THE DEVICE WAS SUDDENLY TURNING RIGHT TO LEFT. THE JUNCTION BETWEEN THE ADAPTER AND RELOAD BROKE OFF AFTER THE SURGEON RETURNED THE HANDLE TO THE NURSE WHO WAS ATTEMPTING TO REMOVE THE RELOAD. THE DEVICE WAS REPLACED WITH A MANUAL HANDLE AND A NEW RELOAD. IT WAS NOTED THAT THE HANDLE HAD NO ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
03-01-2026 | Malfunction | covidien surgical | 16-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLE LINE DID NOT HOLD, IT OPENED UP AFTER COMPLETE FIRING. THE DEVICE WAS REPLACED WITH A NEW ONE.
17-11-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | ECHELON LINEAR CUTTER 80MM | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL PROCEDURE, THE METAL PIN LOCATED AT THE BACK OF THE STAPLER DISLODGED, RENDERING THE STAPLER INOPERABLE. NO PATIENT CONSEQUENCES WERE REPORTED.
28-12-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | 75MM SELECTABLE NEW TLC | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE; WHILE INSTALLING THE SECOND RELOAD, THE DEVICE BECAME STUCK AND DID NOT ALLOW JAW OPENING. STANDARD TROUBLESHOOTING MANEUVERS WERE ATTEMPTED WITHOUT SUCCESS. NO PATIENT CONSEQUENCES WERE REPORTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | CIRCULAR MECH STD 29MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE CANNOT FIRE. NO PATIENT CONSEQUENCES WERE REPORTED.
30-12-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | CONTOUR CURVED CUTTER STAPLER | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE HEAD WAS UNSTABLE DURING USED AND FALLS OFF WHEN FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT PERFORM THE GRAPPLING WHEN THE CARTRIDGE WAS PLACED. THE CARTRIDGE WAS CHECKED, WHICH SEEMED ADVANCED, BUT WITHOUT A COMPLETE STAPLING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
31-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, AN ISSUE WITH POWER CIRCULAR STAPLER. DURING A PROCEDURE THE OR WENT TO USE IT AND IT APPEARED TO HAVE ALREADY BEEN STARTED. THEY WERE UNABLE TO FIRE IT. THEY DID GET A NEW ONE AND IT WORKED JUST FINE. NO PATIENT CONSEQUENCES WERE REPORTED.
23-12-2025 | Death | intuitive surgical | 16-01-2026 | SUREFORM | Firing Problem | A0505 | Failure to Anastomose | nan
23-12-2025 | Death | intuitive surgical | 16-01-2026 | SUREFORM | Firing Problem | A0505 | Exsanguination | nan
23-12-2025 | Death | intuitive surgical | 16-01-2026 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | nan
23-12-2025 | Death | intuitive surgical | 16-01-2026 | SUREFORM | Firing Problem | A0505 | Great Vessel Perforation | nan
01-10-2025 | Malfunction | johnson & johnson medical (suzhou | 16-01-2026 | CARTRIDGE FINAL ASSY 45MM GOLD | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-10-2025 | Malfunction | johnson & johnson medical (suzhou | 16-01-2026 | CARTRIDGE FINAL ASSY 45MM GOLD | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
18-12-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE THE SURGEON WAS USING THE CDH29P AND HAD DIFFICULTY KEEPING THE ANVIL ATTACHED TO THE TROCAR PORTION OF THE STAPLER. NO PATIENT HARM REPORTED.
13-11-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE BATTERY ON THE DEVICE UNEXPECTEDLY SHUT DOWN BEFORE FIRING SO THE SURGEON COULDN'T USE IT ON THE PATIENT. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-11-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | GST WHITE RELOAD, 45MM, 6 ROW | Manufacturing, Packaging or Shipping Problem | A02 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PULMONARY SEGMENTECTOMY, WHEN THE CIRCULATING NURSE OPENED THE CARTRIDGE PACKAGE AND GAVE IT TO THE SCRUB NURSE, THE CARTRIDGE PAN WAS BROKEN. THE SALES REP ATTENDED THE OPERATION AND DETERMINED THAT THERE WAS NO NEGLIGENCE ON THE PART OF THE NURSE AND THAT THE PROBLEM WAS DUE TO A PRODUCT ASSEMBLY FAILURE. AS IT HAD NOT YET BEEN USED, THERE WAS NO IMPACT ON THE PATIENT. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS A CRACKING SOUND INSIDE THE CASING, AND THE MAGAZINE COULD NOT BE FIRED. THE INSTRUMENT WAS STILL ABLE TO BE OPENED, AND A NEW INSTRUMENT WAS OPENED TO COMPLETE THE OPERATION. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | ETS FLEX ARTICNG LNR CUTR 45MM | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS A CRACKING SOUND INSIDE THE CASING, AND THE MAGAZINE COULD NOT BE FIRED. THE INSTRUMENT WAS STILL ABLE TO BE OPENED, AND A NEW INSTRUMENT WAS OPENED TO COMPLETE THE OPERATION. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | CIRCULAR MECH XL SEAL 29MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE BATTERY WAS DRAINED BEFORE USING IT. NO PATIENT INVOLVEMENT.
13-11-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ESOPHAGUS SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2025 | Malfunction | covidien surgical | 16-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED PREOPERATIVELY, THAT THE DEVICE COULD NOT BE CLOSED UNTIL THE END. THERE WAS NO PATIENT INVOLVED.
22-12-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
03-10-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | PWRD ECH FLEX 45MM 44CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER JAMMED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-12-2025 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED PREOPERATIVELY, THAT THE DEVICE COULD NOT BE CLOSED UNTIL THE END. THERE WAS NO PATIENT INVOLVED.
29-11-2025 | Malfunction | ethicon endo-surgery | 16-01-2026 | GOLD RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-12-2025 | Malfunction | covidien surgical | 16-01-2026 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED PREOPERATIVELY THAT THE DEVICE COULD NOT BE CLOSED UNTIL THE END. THERE WAS NO PATIENT INVOLVED.
09-12-2025 | Malfunction | covidien surgical | 16-01-2026 | ENDO GIA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, WHILE ON SET-UP, THE DEVICE'S GREEN BUTTON DID NOT MOVE. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
11-12-2025 | Malfunction | covidien surgical | 16-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, DURING SETUP, THE JAWS WERE NOT OPENING AT ALL. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
11-12-2025 | Malfunction | covidien surgical | 16-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, DURING SETUP, THE JAWS WERE NOT OPENING AT ALL. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
15-09-2025 | Malfunction | covidien | 16-01-2026 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE DEVICE WAS DIFFICULT TO FIRE. THE PRODUCT DID NOT WORK PROPERLY AND A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
15-09-2025 | Malfunction | covidien | 16-01-2026 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE DEVICE WAS DIFFICULT TO FIRE. THE PRODUCT DID NOT WORK PROPERLY AND A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
11-12-2025 | Malfunction | covidien surgical | 16-01-2026 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, DURING SETUP, THE JAWS WERE NOT OPENING AT ALL. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
28-12-2025 | Malfunction | covidien surgical | 16-01-2026 | MAXTACK | Positioning Failure | A150201 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH (IPOM) REPAIR, AT THE TIME OF FIRING, THE MOTORIZED TACKER HAD AN ISSUE. THE ISSUE WAS RESOLVED BY USING ANOTHER TACKER TO COMPLETE THE CASE.
28-12-2025 | Malfunction | covidien surgical | 16-01-2026 | MAXTACK | Misfire | A050502 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH (IPOM) REPAIR, AT THE TIME OF FIRING, THE MOTORIZED TACKER HAD AN ISSUE. THE ISSUE WAS RESOLVED BY USING ANOTHER TACKER TO COMPLETE THE CASE.
28-12-2025 | Malfunction | covidien surgical | 16-01-2026 | MAXTACK | Device Displays Incorrect Message | A090201 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH (IPOM) REPAIR, AT THE TIME OF FIRING, THE MOTORIZED TACKER HAD AN ISSUE. THE ISSUE WAS RESOLVED BY USING ANOTHER TACKER TO COMPLETE THE CASE.
14-10-2025 | Malfunction | covidien surgical | 16-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE FIRED SUCCESSFULLY. WHEN THE SURGICAL STAFF WANTED TO REMOVE THE EMPTY RELOAD FROM THE ADAPTER, IT GOT STUCK. THE RELOAD AND ADAPTER WERE SUCCESSFULLY REMOVED FROM THE PATIENT. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT RELOAD HAS ALL PUSHERS DEPLOYED AND WAS NOT CLAMPED.
14-10-2025 | Malfunction | covidien surgical | 16-01-2026 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE FIRED SUCCESSFULLY. WHEN THE SURGICAL STAFF WANTED TO REMOVE THE EMPTY RELOAD FROM THE ADAPTER, IT GOT STUCK. THE RELOAD AND ADAPTER WERE SUCCESSFULLY REMOVED FROM THE PATIENT. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT RELOAD HAS ALL PUSHERS DEPLOYED AND WAS NOT CLAMPED.
03-12-2025 | Malfunction | covidien surgical | 17-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE OPERATOR SAID THE RELOAD USED WITH THE POWERED STAPLER HANDLE COULD NOT BE CLOSED UNTIL THE END. THERE WAS NO PATIENT INVOLVED.
03-12-2025 | Malfunction | covidien surgical | 17-01-2026 | SIGPHANDLE | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE OPERATOR SAID THE RELOAD USED WITH THE POWERED STAPLER HANDLE COULD NOT BE CLOSED UNTIL THE END. THERE WAS NO PATIENT INVOLVED.
03-12-2025 | Malfunction | covidien surgical | 17-01-2026 | UNK-SIGADAPT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE OPERATOR SAID THE RELOAD USED WITH THE POWERED STAPLER HANDLE COULD NOT BE CLOSED UNTIL THE END. THERE WAS NO PATIENT INVOLVED.
20-11-2025 | Injury | covidien surgical | 18-01-2026 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING AN AORTIC VALVE REPLACEMENT, THE SURGEON USED THE TRI-STAPLE. AFTER REMOVING THE STAPLER HANDPIECE FROM THE CHEST, THE SURGEON NOTED A PIECE OF PLASTIC INSIDE THE CHEST CAVITY. USER REMOVED THE PLASTIC PIECE AND NOTIFIED THE SURGICAL TEAM. THE SURGEON STATED THAT USER DID NOT SEE ANY OTHER OBJECTS IN THE CHEST CAVITY. THE SURGERY WAS COMPLETED AS SCHEDULED. A POST-OPERATIVE X-RAY WAS TAKEN IN THE HEART CENTER. THE RADIOLOGIST STATED THAT STERNAL WIRING AND A LINEAR CLIP WERE IN PLACE, WITH NO OTHER FOREIGN OBJECTS IDENTIFIED.
24-12-2025 | Malfunction | smith & nephew | 18-01-2026 | TENDON ANCHORS 8 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A SUPRASPINATUS INSUFFICIENCY RUPTURE SURGERY, A TENDON ANCHOR WAS BENT AND BROKE WHEN IMPLANTED INTO THE CUFF. THE PROCEDURE WAS RESUMED, AFTER A 10-MINUTE DELAY, USING AN EQUIVALENT S+N BACK UP. NO FURTHER COMPLICATIONS WERE REPORTED.
24-12-2025 | Malfunction | smith & nephew | 18-01-2026 | TENDON ANCHORS 8 | Material Twisted/Bent | A040609 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A SUPRASPINATUS INSUFFICIENCY RUPTURE SURGERY, A TENDON ANCHOR WAS BENT AND BROKE WHEN IMPLANTED INTO THE CUFF. THE PROCEDURE WAS RESUMED, AFTER A 10-MINUTE DELAY, USING AN EQUIVALENT S+N BACK UP. NO FURTHER COMPLICATIONS WERE REPORTED.
28-11-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE GOT STUCK ON THE 4TH FIRING AND JAMMED. THEY HAD TO DO MANUAL OVERRIDE TO OPEN THE GUN. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
28-11-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE GOT STUCK ON THE 4TH FIRING AND JAMMED. THEY HAD TO DO MANUAL OVERRIDE TO OPEN THE GUN. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-11-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | ECHELON FLEX | Material Deformation | A0406 | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLE LINE DEHISCED SHORTLY AFTER USE.
21-12-2025 | Injury | intuitive surgical | 20-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS (ROUX-EN-Y) PROCEDURE, STOMACH TISSUE WAS INADVERTENTLY INJURED FOLLOWING AN ATTEMPT TO USE A SUREFORM 60 STAPLER WITH A SUREFORM 60 BLUE RELOAD, WHICH GENERATED A "TISSUE TOO THICK TO CONTINUE" ERROR MESSAGE. THE INCIDENT OCCURRED DURING THE GASTRIC POUCH CREATION, SPECIFICALLY ON THE LEFT SIDE ALONG THE GREATER CURVATURE. THE STAPLER INSTRUMENT COULD NOT FULLY CLAMP, AND FIRING WAS NOT INITIATED; HOWEVER, TISSUE WAS DISPLACED FORWARD FROM THE JAWS. THE SURGEON DESCRIBED THE INCIDENT AS A ¿MIGRATORY STAPLE AT THE APEX,¿ AN OCCURRENCE OBSERVED IN SIMILAR CIRCUMSTANCES WITH OTHER STAPLING SYSTEMS. THIS LED TO THE STAPLER RELOAD BEING STUCK TO THE TISSUE, RESULTING IN TISSUE DAMAGE BUT NO BLEEDING OCCURRED. ADDITIONAL TISSUE REMOVAL WAS REQUIRED, AND THE ISSUE WAS RESOLVED USING THE SAME STAPLER INSTRUMENT, AND FIRING A NEW STAPLER RELOAD CLOSER TO THE "GASTRIC POUCH LIMITS". THE SURGEON ATTRIBUTED THE EVENT TO INCOMPLETE CLOSURE OF THE STAPLER RELOAD AND ADHERENCE TO THE TISSUE. DESPITE THIS INCIDENT, THE PROCEDURE WAS COMPLETED ROBOTICALLY WITHOUT ADDITIONAL COMPLICATIONS. THE PATIENT DID NOT EXPERIENCE ANY POSTOPERATIVE COMPLICATIONS AND THERE IS NO CONCERN FOR LONG-TERM COMPLICATIONS RELATED TO THIS EVENT.
02-01-2026 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2026 | SORBAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, SORBAFIX ENHANCED FIXATION DEVICE FAIL TO SECURE THE SCREWS. THE PLATE (MESH) IS THEREFORE NOT FIXED. THE USE OF ANOTHER DEVICE IS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT HARM OR INJURY.
02-10-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | PROXIMATE*RL LINEAR STAPLER | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOTOMY, IT WAS OBSERVED THAT THE LID OF THE PACKAGING OF THE DEVICE WAS TORN DURING ITS UNPACKING, MAKING IT IMPOSSIBLE TO GRASP STERILE. THIS INCIDENT HAD NO CLINICAL CONSEQUENCES FOR THE PATIENT BUT RESULTED IN AN ADDITIONAL COST FOR THE INSTITUTION.
01-01-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | PROXIMATE*RL LINEAR STAPLER | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | nan
04-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, A SURGICAL TECHNOLOGIST LOADED THE CARTRIDGE INTO THE STAPLER THEY USE FREQUENTLY. POSITIONED THE STAPLER ON THE STOMACH AND, WHEN WENT TO FIRE, THE INSTRUMENT WAS JAMMED. REMOVED THE CARTRIDGE, PUT IT BACK IN, AND TRIED TO FIRE AGAIN, BUT IT DIDN¿T WORK. WHEN WE REMOVED THE CARTRIDGE, WE NOTICED THAT THE CUTTING BLADE WAS OUT OF PLACE ¿ THE LITTLE CUTTING BLADE CAME OUT ALONG WITH THE CARTRIDGE. NO PATIENT CONSEQUENCES WERE REPORTED.
04-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, A SURGICAL TECHNOLOGIST LOADED THE CARTRIDGE INTO THE STAPLER THEY USE FREQUENTLY. POSITIONED THE STAPLER ON THE STOMACH AND, WHEN WENT TO FIRE, THE INSTRUMENT WAS JAMMED. REMOVED THE CARTRIDGE, PUT IT BACK IN, AND TRIED TO FIRE AGAIN, BUT IT DIDN¿T WORK. WHEN WE REMOVED THE CARTRIDGE, WE NOTICED THAT THE CUTTING BLADE WAS OUT OF PLACE ¿ THE LITTLE CUTTING BLADE CAME OUT ALONG WITH THE CARTRIDGE. NO PATIENT CONSEQUENCES WERE REPORTED.
22-12-2025 | Injury | intuitive surgical | 20-01-2026 | SUREFORM | Firing Problem | A0505 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A DA VINCI¿ASSISTED APPENDECTOMY PROCEDURE, MALFORMED STAPLES WERE OBSERVED AFTER FIRING A SUREFORM 60 STAPLER INSTRUMENT EQUIPPED WITH A SUREFORM 60 BLUE RELOAD. THE STAPLER INSTRUMENT WAS CLAMPED ONTO TISSUE AND AN ERROR MESSAGE WAS OBSERVED STATING "TISSUE TOO THICK TO CONTINUE." THE CUSTOMER THEN REPLACED THE BLUE STAPLER RELOAD WITH A GREEN 60 RELOAD. THE SUBSEQUENT FIRE WITH THE GREEN STAPLER RELOAD WAS SUCCESSFULLY COMPLETED, WITHOUT COMPLICATIONS. IT REMAINS UNCLEAR WHETHER A NEW BLUE STAPLER RELOAD WAS USED, OR IF THE INITIAL BLUE STAPLER RELOAD, WHICH HAD GENERATED THE ERROR MESSAGE, WAS REINSTALLED FOR A SUBSEQUENT FIRING. AT THE COMPLETION OF THE FIRING, THE STAPLES THAT WERE DEPLOYED WERE OBSERVED TO BE MALFORMED. THERE WAS NO BLEEDING AS A RESULT OF THE EVENT, NO LOOSE STAPLES FELL INSIDE THE PATIENT, AND THE MALFORMED STAPLES WERE NOT REMOVED. THE INTUITIVE SURGICAL, INC. (ISI) CLINICAL SALES MANAGER (CSM) COULD NOT CONFIRM IF ANY HOLES OR GAPS WERE PRESENT IN THE STAPLE LINE DUE TO THE FIRING ISSUE. A SUREFORM 60 GREEN RELOAD WAS SUBSEQUENTLY FIRED SUCCESSFULLY, AND THE PROCEDURE CONTINUED WITHOUT FURTHER COMPLICATIONS. THE PROCEDURE WAS SUCCESSFULLY COMPLETED ROBOTICALLY, NO POSTOPERATIVE IMAGING WAS PERFORMED, AND THE PATIENT HAS NOT EXPERIENCED ANY POSTOPERATIVE COMPLICATIONS OR REQUIRED A RETURN TO THE HOSPITAL.
27-12-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECONSTRUCTION OF BILIARY VIA, WHEN PLACING THE REFILL IN THE LINEAR STAPLER, CUT AND EVIDENCED AN EXPOSED BLADE, THE REFILL DOES NOT ADAPT TO THE STAPLER AND LEAVES IT, FOR THIS REASON A NEW REFILL IS REQUESTED TO PERFORM THE INTESTINAL ANASTOMOSIS. NO PATIENT CONSEQUENCES WERE REPORTED.
26-12-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
29-12-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | FLEX 60 ARTICULATING | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | nan
29-12-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
05-01-2026 | Malfunction | covidien surgical | 20-01-2026 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, THE DEMO POWER HANDLE WOULD NOT CHARGE. IT WAS NOTED THAT THE CHARGER WORKED TO CHARGE THE OTHER HANDLES. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS OPENED UP AND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
05-01-2026 | Malfunction | covidien surgical | 20-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, THE DEMO POWER HANDLE WOULD NOT CHARGE. IT WAS NOTED THAT THE CHARGER WORKED TO CHARGE THE OTHER HANDLES. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS OPENED UP AND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
16-12-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. THEY USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. CHANGED TO A NEW ONE TO COMPLETE THE PROCDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. THEY USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. CHANGED TO A NEW ONE TO COMPLETE THE PROCDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. THEY USED MANUAL OVERRIDE LEVER TO RETURN THE KNIFE. CHANGED TO A NEW ONE TO COMPLETE THE PROCDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
16-01-2026 | Malfunction | covidien surgical | 20-01-2026 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN USING THE RELOAD ON THE UPPER LOBE PULMONARY ARTERY, THE STAPLES FIRED BUT THE DEVICE DID NOT CUT THE TISSUE. THE PROXIMAL STAPLE LINE DID NOT SEAL. TO RESOLVE THE ISSUE, SEVERAL BLUE CLIPS FROM A DIFFERENT MANUFACTURER WERE APPLIED TO REINFORCE, ARTERY DIVIDED USING DEVICE.
16-01-2026 | Malfunction | covidien surgical | 20-01-2026 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN USING THE RELOAD ON THE UPPER LOBE PULMONARY ARTERY, THE STAPLES FIRED BUT THE DEVICE DID NOT CUT THE TISSUE. THE PROXIMAL STAPLE LINE DID NOT SEAL. TO RESOLVE THE ISSUE, SEVERAL BLUE CLIPS FROM A DIFFERENT MANUFACTURER WERE APPLIED TO REINFORCE, ARTERY DIVIDED USING DEVICE.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND OOZING OCCURRED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND OOZING OCCURRED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND OOZING OCCURRED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND OOZING OCCURRED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
24-12-2025 | Malfunction | covidien | 20-01-2026 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE RIGHT LOWER LOBE PARTIAL RESECTION, THERE WAS A CARTRIDGE CONNECTION FAILURE. A NEW CARTRIDGE WAS TAKEN OUT AND THE CONNECTION COULD BE MADE WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS THICK TISSUE. LOCK RING WAS SLIGHTLY OUT OF POSITION. A SCRATCH WAS OBSERVED ON THE CARTRIDGE SIDE JAW, WHICH MIGHT HAVE BEEN MADE WHEN THE I BEAM ADVANCED FORCEFULLY IN CLAMPING THE JAW OR FIRING.
17-12-2025 | Injury | intuitive surgical | 20-01-2026 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC SLEEVE SURGICAL PROCEDURE, THE STAPLE LINE BLED WHERE A SUREFORM 60 GREEN RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. AS A RESULT, THE SURGEON OVER-SEWED THE STAPLE LINE WITH A SUTURE TO CONTROL THE BLEEDING. THE SURGEON EXPLAINED THAT ETHICON SLR BUTTRESS MATERIAL (A 3RD-PARTY MANUFACTURER PRODUCT) WAS INITIALLY USED WITH THE SUREFORM 60 GREEN RELOAD FOR REINFORCEMENT OF THE STAPLE LINE. HOWEVER, DURING THE FIRING SEQUENCE, AN ERROR MESSAGE STATING "TISSUE TOO THICK TO CONTINUE" WAS DISPLAYED, AND FORCE FIRE WAS AN AVAILABLE OPTION. THE BUTTRESS MATERIAL WAS REMOVED, AND A SUREFORM 60 GREEN RELOAD WAS FIRED WITHOUT COMPLICATION. THE STAPLE LINE WAS OBSERVED TO BE OOZING. THE SURGEON THEN OVER-SUTURED THE STAPLE LINE. THE PROCEDURE WAS COMPLETED ROBOTICALLY.
25-11-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | 55 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED PLUS STANDARD 45 | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE AND NOTED THE STAPLES MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED PLUS STANDARD 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE AND NOTED THE STAPLES MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED PLUS STANDARD 45 | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE AND NOTED THE STAPLES MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED PLUS STANDARD 45 | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE AND NOTED THE STAPLES MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
03-12-2025 | Injury | intuitive surgical | 20-01-2026 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, THE PATIENT WAS READMITTED TO THE HOSPITAL DUE TO COMPLICATIONS. DURING THE ROBOTIC PROCEDURE, THERE WAS OOZING FROM THE STAPLE LINE WHERE A SUREFORM 60 WHITE RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. THE LUMENS OF VESSELS COULD BE SEEN THROUGH THE STAPLE LINE. AS A RESULT, THE SURGEON OVER-SUTURED THE STAPLE LINE TO ENSURE HEMOSTASIS. POSTOPERATIVE DAY ONE, THE PATIENT EXPERIENCED EMESIS. AN ENDOSCOPY WAS PERFORMED, WHICH SHOWED A CLOT OBSTRUCTING THE ANASTOMOSIS, AND IT WAS REMOVED. THE PATIENT WAS KEPT HOSPITALIZED FOR A FEW DAYS FOR OBSERVATION.
03-12-2025 | Injury | intuitive surgical | 20-01-2026 | SUREFORM | Firing Problem | A0505 | Vomiting | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, THE PATIENT WAS READMITTED TO THE HOSPITAL DUE TO COMPLICATIONS. DURING THE ROBOTIC PROCEDURE, THERE WAS OOZING FROM THE STAPLE LINE WHERE A SUREFORM 60 WHITE RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. THE LUMENS OF VESSELS COULD BE SEEN THROUGH THE STAPLE LINE. AS A RESULT, THE SURGEON OVER-SUTURED THE STAPLE LINE TO ENSURE HEMOSTASIS. POSTOPERATIVE DAY ONE, THE PATIENT EXPERIENCED EMESIS. AN ENDOSCOPY WAS PERFORMED, WHICH SHOWED A CLOT OBSTRUCTING THE ANASTOMOSIS, AND IT WAS REMOVED. THE PATIENT WAS KEPT HOSPITALIZED FOR A FEW DAYS FOR OBSERVATION.
03-12-2025 | Injury | intuitive surgical | 20-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, THE PATIENT WAS READMITTED TO THE HOSPITAL DUE TO COMPLICATIONS. DURING THE ROBOTIC PROCEDURE, THERE WAS OOZING FROM THE STAPLE LINE WHERE A SUREFORM 60 WHITE RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. THE LUMENS OF VESSELS COULD BE SEEN THROUGH THE STAPLE LINE. AS A RESULT, THE SURGEON OVER-SUTURED THE STAPLE LINE TO ENSURE HEMOSTASIS. POSTOPERATIVE DAY ONE, THE PATIENT EXPERIENCED EMESIS. AN ENDOSCOPY WAS PERFORMED, WHICH SHOWED A CLOT OBSTRUCTING THE ANASTOMOSIS, AND IT WAS REMOVED. THE PATIENT WAS KEPT HOSPITALIZED FOR A FEW DAYS FOR OBSERVATION.
26-11-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | PROXIMATE*75MM LINEAR CUTTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE COLON RESECTION SURGERY, COULD NOT FIRE. CHANGED ANOTHER RELOAD TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-11-2025 | Malfunction | ethicon endo-surgery | 20-01-2026 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE GASTRECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
07-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED PLUS STANDARD 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THEY COULD NOT OPEN THE JAW AFTER FIRING. THEY USED REVERSE BUTTON AND MANUAL OVERRIDE TO OPEN THE JAW AND REMOVED THE DEVICE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THEY COULD NOT OPEN THE JAW AFTER FIRING. THEY USED REVERSE BUTTON AND MANUAL OVERRIDE TO OPEN THE JAW AND REMOVED THE DEVICE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED PLUS STANDARD 45 | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THEY COULD NOT OPEN THE JAW AFTER FIRING. THEY USED REVERSE BUTTON AND MANUAL OVERRIDE TO OPEN THE JAW AND REMOVED THE DEVICE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2026 | Malfunction | covidien surgical | 20-01-2026 | TRI-STAPLE 2.0 | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC-ASSISTED THORACOSCOPIC SURGERY (RATS) LOBECTOMY, ON THE PULMONARY BLOOD VESSEL, THE FIRING WAS ABLE TO BE PERFORMED WITHOUT ANY PROBLEMS, BUT THERE WAS A SIGNIFICANT TWITCHING DURING FIRING, SO IT WAS REGISTERED. THERE WAS NO PATIENT INJURY.
07-01-2026 | Malfunction | covidien surgical | 20-01-2026 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC-ASSISTED THORACOSCOPIC SURGERY (RATS) LOBECTOMY, ON THE PULMONARY BLOOD VESSEL, THE FIRING WAS ABLE TO BE PERFORMED WITHOUT ANY PROBLEMS, BUT THERE WAS A SIGNIFICANT TWITCHING DURING FIRING, SO IT WAS REGISTERED. THERE WAS NO PATIENT INJURY.
20-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN URGENT SURGERY OF COLECTOMY, WHEN CLOSING THE INSERTION SLOTS, THE KNIFE MOVED FORWARD, BUT DID NOT RETURN AT THE 4TH FIRING OF FUNCTIONAL END TO END ANASTOMOSIS. THE KNIFE RETURNED WITH THE MANUAL OVERRIDE LEVER, AND THE SURGERY WAS COMPLETED. THERE WAS NO EFFECT ON THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
20-01-2026 | Malfunction | ethicon endo-surgery | 20-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN URGENT SURGERY OF COLECTOMY, WHEN CLOSING THE INSERTION SLOTS, THE KNIFE MOVED FORWARD, BUT DID NOT RETURN AT THE 4TH FIRING OF FUNCTIONAL END TO END ANASTOMOSIS. THE KNIFE RETURNED WITH THE MANUAL OVERRIDE LEVER, AND THE SURGERY WAS COMPLETED. THERE WAS NO EFFECT ON THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
07-01-2026 | Malfunction | covidien surgical | 20-01-2026 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC-ASSISTED THORACOSCOPIC SURGERY (RATS) LOBECTOMY, ON THE PULMONARY BLOOD VESSEL, THE FIRING WAS ABLE TO BE PERFORMED WITHOUT ANY PROBLEMS, BUT THERE WAS A SIGNIFICANT TWITCHING DURING FIRING, SO IT WAS REGISTERED. THERE WAS NO PATIENT INJURY.
24-08-2025 | Malfunction | covidien surgical | 20-01-2026 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING TESTING AND MAINTENANCE, THE POWER HANDLE WAS NOT CHARGING. ALL PORTS ON THE FOUR-BAY CHARGER WERE CLEANED AND VERIFIED AS WORKING PRIOR TO ATTEMPTING TO CHARGE THE UNIT, BUT THE HANDLE STILL DID NOT CHARGE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHEN THE HANDLE WAS OPENED UP AND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
24-08-2025 | Malfunction | covidien surgical | 20-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING TESTING AND MAINTENANCE, THE POWER HANDLE WAS NOT CHARGING. ALL PORTS ON THE FOUR-BAY CHARGER WERE CLEANED AND VERIFIED AS WORKING PRIOR TO ATTEMPTING TO CHARGE THE UNIT, BUT THE HANDLE STILL DID NOT CHARGE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHEN THE HANDLE WAS OPENED UP AND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
07-01-2026 | Malfunction | covidien surgical | 20-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ESOPHAGECTOMY, PRIOR TO RESECTION OF THE AZYGOS VEIN, THE HANDLE STOPPED OPERATING AND THE DISPLAY WENT BLANK. THE HANDLE, ADAPTER, POWER SHELL, AND THE CARTRIDGE WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien | 20-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN APPENDECTOMY, THE DEVICE DID NOT FIRE AT ALL. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien | 20-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN APPENDECTOMY, THE DEVICE DID NOT FIRE AT ALL. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2026 | Malfunction | covidien | 20-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLE LINE DID NOT HOLD AND OPENED UP AFTER COMPLETE FIRING. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
13-01-2026 | Malfunction | covidien surgical | 21-01-2026 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, THE STAPLER DID NOT AUTOMATICALLY STOP AT THE END OF THE RELOAD. THE RELOAD WAS BENT AT AN ANGLE, AND ONGOING PRESSURE OR FORCE WAS OBSERVED ON THE RELOAD. THE METAL STAPLE COMPONENTS WAS BEING PUSHED OUTWARDS. THE SURGEON EXPERIENCED DIFFICULTY RE-ALIGNING THE STAPLER AND REQUIRED SUFFICIENT PRESSURE AGAINST TISSUES TO REMOVE IT. A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
13-01-2026 | Malfunction | covidien surgical | 21-01-2026 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, THE STAPLER DID NOT AUTOMATICALLY STOP AT THE END OF THE RELOAD. THE RELOAD WAS BENT AT AN ANGLE, AND ONGOING PRESSURE OR FORCE WAS OBSERVED ON THE RELOAD. THE METAL STAPLE COMPONENTS WAS BEING PUSHED OUTWARDS. THE SURGEON EXPERIENCED DIFFICULTY RE-ALIGNING THE STAPLER AND REQUIRED SUFFICIENT PRESSURE AGAINST TISSUES TO REMOVE IT. A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
13-01-2026 | Malfunction | covidien surgical | 21-01-2026 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, THE STAPLER DID NOT AUTOMATICALLY STOP AT THE END OF THE RELOAD. THE RELOAD WAS BENT AT AN ANGLE, AND ONGOING PRESSURE OR FORCE WAS OBSERVED ON THE RELOAD. THE METAL STAPLE COMPONENTS WAS BEING PUSHED OUTWARDS. THE SURGEON EXPERIENCED DIFFICULTY RE-ALIGNING THE STAPLER AND REQUIRED SUFFICIENT PRESSURE AGAINST TISSUES TO REMOVE IT. A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
13-01-2026 | Malfunction | covidien surgical | 21-01-2026 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS, THE STAPLER DID NOT AUTOMATICALLY STOP AT THE END OF THE RELOAD. THE RELOAD WAS BENT AT AN ANGLE, AND ONGOING PRESSURE OR FORCE WAS OBSERVED ON THE RELOAD. THE METAL STAPLE COMPONENTS WAS BEING PUSHED OUTWARDS. THE SURGEON EXPERIENCED DIFFICULTY RE-ALIGNING THE STAPLER AND REQUIRED SUFFICIENT PRESSURE AGAINST TISSUES TO REMOVE IT. A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
06-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND OOZING OCCURRED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND OOZING OCCURRED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND OOZING OCCURRED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND OOZING OCCURRED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2025 | Malfunction | covidien | 21-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: AN ENDO GIA 45MM VASCULAR MEDIUM STAPLE LOAD (REF# EGIA45AVM, LOT#P5J0901), FAILED TO OPEN AND CLOSE ONCE ATTACHED TO THE ENDO GIA STAPLER. IT WAS CONFIRMED TO BE THE STAPLER LOAD AT FAULT ONCE ANOTHER STAPLER LOAD WAS LOADED AND USED WITH NO FURTHER INCIDENT. THIS MALFUNCTION HAD THE POTENTIAL OF CAUSING DIRECT PATIENT HARM IF IT WAS FOUND TO BE CLAMPED BY THE GALLBLADDER AND COULD NOT BE RELEASED.
27-11-2025 | Injury | covidien | 21-01-2026 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A ROBOTIC LAPAROSCOPIC ANTERIOR RESECTION, THERE WAS AN ANASTOMOTIC LEAKAGE.
27-11-2025 | Injury | covidien | 21-01-2026 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A ROBOTIC LAPAROSCOPIC ANTERIOR RESECTION, THERE WAS AN ANASTOMOTIC LEAKAGE.
27-11-2025 | Injury | covidien | 21-01-2026 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abnormal Vaginal Discharge | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A ROBOTIC LAPAROSCOPIC ANTERIOR RESECTION, THERE WAS AN ANASTOMOTIC LEAKAGE.
27-11-2025 | Injury | covidien | 21-01-2026 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A ROBOTIC LAPAROSCOPIC ANTERIOR RESECTION, THERE WAS AN ANASTOMOTIC LEAKAGE.
27-11-2025 | Injury | covidien | 21-01-2026 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | High White Blood Cell Count | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A ROBOTIC LAPAROSCOPIC ANTERIOR RESECTION, THERE WAS AN ANASTOMOTIC LEAKAGE.
20-01-2026 | Malfunction | covidien surgical | 21-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, THE RELOAD WAS USED TO DETACH THE LUNG PARENCHYMA, THE RELOAD WAS NOT FULLY FORMED, AND THE STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL END. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW RELOAD.
20-01-2026 | Malfunction | covidien surgical | 21-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, THE RELOAD WAS USED TO DETACH THE LUNG PARENCHYMA, THE RELOAD WAS NOT FULLY FORMED, AND THE STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL END. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW RELOAD.
20-01-2026 | Malfunction | covidien surgical | 21-01-2026 | ENDO GIA | Insufficient Information | E2401 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, THE RELOAD WAS USED TO DETACH THE LUNG PARENCHYMA, THE RELOAD WAS NOT FULLY FORMED, AND THE STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL END. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW RELOAD.
20-01-2026 | Malfunction | covidien surgical | 21-01-2026 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, THE RELOAD WAS USED TO DETACH THE LUNG PARENCHYMA, THE RELOAD WAS NOT FULLY FORMED, AND THE STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL END. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW RELOAD.
21-01-2026 | Malfunction | covidien | 21-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A BARIATRIC SURGERY, WHEN THE SURGEON USED THE FIRST MANUAL STAPLER, IT WAS USED WITH NO ISSUE. THE SECOND TIME IT WAS FIRED, THE SURGEON HEARD A SUSPICIOUS BROKEN NOISE. ON THE THIRD FIRING, IT DID NOT FIRE. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. WHEN USING THE POWERED HANDLE, THE DEVICE HAS AN END-OF-LIFE ERROR. THERE WAS NO PATIENT INJURY.
21-01-2026 | Malfunction | covidien | 21-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A BARIATRIC SURGERY, WHEN THE SURGEON USED THE FIRST MANUAL STAPLER, IT WAS USED WITH NO ISSUE. THE SECOND TIME IT WAS FIRED, THE SURGEON HEARD A SUSPICIOUS BROKEN NOISE. ON THE THIRD FIRING, IT DID NOT FIRE. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. WHEN USING THE POWERED HANDLE, THE DEVICE HAS AN END-OF-LIFE ERROR. THERE WAS NO PATIENT INJURY.
30-12-2025 | Malfunction | covidien surgical | 21-01-2026 | EEA ORVIL | Incomplete or Inadequate Connection | A1207 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE ORAL ANVIL WAS DIFFICULT TO ATTACH TO THE STAPLER DURING SET-UP. THE DEVICE WAS REPLACED, AND ANOTHER ORAL ANVIL WAS USED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-10-2025 | Malfunction | covidien surgical | 21-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, THERE WAS STAPLE LINE BLEEDING FOLLOWING FIRING OF THE RELOAD. THE ISSUE WAS ADDRESSED BY RESECTING AND RE-STAPLING. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES.
01-10-2025 | Malfunction | covidien surgical | 21-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, THERE WAS STAPLE LINE BLEEDING FOLLOWING FIRING OF THE RELOAD. THE ISSUE WAS ADDRESSED BY RESECTING AND RE-STAPLING. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY LESS THAN 30 MINUTES.
01-01-2025 | Malfunction | covidien surgical | 21-01-2026 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SET-UP PRIOR TO PATIENT CONTACT, THE POWERED HANDLE DID NOT TURN ON DESPITE BEING FULLY CHARGED. THE DEVICE WAS REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE OBSERVED VENTED BATTERY WAS DUE TO BATTERY DEGRADATION (COMPONENT FAILURE).
01-01-2025 | Malfunction | covidien surgical | 21-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SET-UP PRIOR TO PATIENT CONTACT, THE POWERED HANDLE DID NOT TURN ON DESPITE BEING FULLY CHARGED. THE DEVICE WAS REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE OBSERVED VENTED BATTERY WAS DUE TO BATTERY DEGRADATION (COMPONENT FAILURE).
18-12-2025 | Malfunction | covidien surgical | 21-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TISSUE RESECTION, WHILE ON SET-UP, THE DEVICE COULD NOT BE CLAMPED. IT WAS ALSO NOTED THAT THE SQUEEZING OF THE HANDLE WAS VERY STIFF, AND IT TOOK A LOT OF FORCE THAN USUAL DURING RELEASING. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2025 | Malfunction | covidien surgical | 21-01-2026 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TISSUE RESECTION, WHILE ON SET-UP, THE DEVICE COULD NOT BE CLAMPED. IT WAS ALSO NOTED THAT THE SQUEEZING OF THE HANDLE WAS VERY STIFF, AND IT TOOK A LOT OF FORCE THAN USUAL DURING RELEASING. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2025 | Malfunction | covidien surgical | 21-01-2026 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TISSUE RESECTION, WHILE ON SET-UP, THE DEVICE COULD NOT BE CLAMPED. IT WAS ALSO NOTED THAT THE SQUEEZING OF THE HANDLE WAS VERY STIFF, AND IT TOOK A LOT OF FORCE THAN USUAL DURING RELEASING. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A CHOLECYSTECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS NOT WORKING; AND THAT IT JAMMED AND WAS STUCK INSIDE THE CAVITY AS THEY TRIED TO OPEN IT TO PLACE THE TISSUE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
16-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A CHOLECYSTECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS NOT WORKING; AND THAT IT JAMMED AND WAS STUCK INSIDE THE CAVITY AS THEY TRIED TO OPEN IT TO PLACE THE TISSUE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
05-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON 3000 60MM STANDARD | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | nan
08-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GREEN RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GREEN RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | johnson & johnson medical (suzhou | 21-01-2026 | FINAL ASSY 60MM STD DEVICE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
10-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
10-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
17-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
18-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
17-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND THERE WAS SLIGHT OOZING. TO STOP OOZING WITH CLIP. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND THERE WAS SLIGHT OOZING. TO STOP OOZING WITH CLIP. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND THERE WAS SLIGHT OOZING. TO STOP OOZING WITH CLIP. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AFTER FIRING; AND THERE WAS SLIGHT OOZING. TO STOP OOZING WITH CLIP. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-10-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
18-10-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
06-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST GOLD RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON 3000 45 STANDARD | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE PLASTIC SLEEVE WHICH COVERED A TIP OF STAPLER HAD MOISTURE INSIDE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
21-10-2025 | Injury | ethicon endo-surgery | 21-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: PER USER FACILITY MEDWATCH FORM, # MW5179628, DURING AN UNKNOWN PROCEDURE, THE DOCTOR FIRED A STAPLE LOAD. AND THERE WAS A LEAK ALONG THE STAPLE LINE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
11-01-2026 | Injury | ethicon endo-surgery | 21-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A KIDNEY SURGERY DURING INTRAOPERATIVE USE, THE POWERED VASCULAR STAPLER FAILED TO FUNCTION AS INTENDED AT THE TIME OF FIRING, RESULTING IN AN UNSUCCESSFUL STAPLE DEPLOYMENT. THIS LED TO SEVERE BLEEDING, AND IMMEDIATE CORRECTIVE ACTIONS WERE REQUIRED TO CONTROL HEMOSTASIS. DUE TO THE SEVERITY OF BLEEDING AND FAILURE OF THE DEVICE, THE SURGICAL TEAM WAS COMPELLED TO CONVERT THE PROCEDURE FROM LAPAROSCOPIC TO OPEN SURGERY TO ENSURE PATIENT SAFETY. IF OTHER, DESCRIBE -- KIDNEY TRANSPLANT DONOR CASE, WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? -- YES, WAS PROCEDURE SUCCESSFULLY COMPLETED? -- YES, IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? -- NO, OTHER INFORMATION: -- ¿ SEVERE INTRAOPERATIVE BLEEDING ¿ CONVERSION FROM LAPAROSCOPIC TO OPEN SURGERY ¿ PROLONGATION OF OPERATIVE TIME, PATIENT STATUS/ OUTCOME / CONSEQUENCES -- YES, PATIENT CONSEQUENCE DESCRIPTION/WAS THERE A CLINICAL OUTCOME EXPERIENCED BY THE PATIENT (INFECTION, INFLAMMATION, ETC.)? -- SEVERE INTRAOPERATIVE BLEEDING,
14-01-2026 | Injury | ethicon endo-surgery | 21-01-2026 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A WEDGE RESECTION PROCEDURE THE STAPLER WAS FIRED WITH A GREEN RELOAD. STAPLES DID NOT DEPLOY BUT THE KNIFE BLADE FULLY TRANSECTED THE TISSUE. UNKNOWN AMOUNT OF BLOOD LOSS. UNKNOWN ADDITIONAL TIME SPENT IN THE HOSPITAL.
14-01-2026 | Injury | ethicon endo-surgery | 21-01-2026 | ECHELON 3000 45 STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A WEDGE RESECTION PROCEDURE THE STAPLER WAS FIRED WITH A GREEN RELOAD. STAPLES DID NOT DEPLOY BUT THE KNIFE BLADE FULLY TRANSECTED THE TISSUE. UNKNOWN AMOUNT OF BLOOD LOSS. UNKNOWN ADDITIONAL TIME SPENT IN THE HOSPITAL.
15-12-2025 | Injury | intuitive surgical | 21-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED ENDOMETRIOSIS SURGICAL PROCEDURE, TISSUE WAS STAPLED WITH A SUREFORM 45 BLACK RELOAD INSTALLED ON A SUREFORM 45 STAPLER INSTRUMENT. THE PROCEDURE WAS COMPLETED ROBOTICALLY WITH NO INTRAOPERATIVE COMPLICATIONS; SPECIFICALLY, THERE WERE NO COMPLICATIONS WITH THE SUREFORM 45 STAPLER INSTRUMENT. HOWEVER, THE PATIENT DEVELOPED A POSTOPERATIVE LEAK AND RETURNED TO THE HOSPITAL ON POSTOPERATIVE DAY FOUR. THE PATIENT REQUIRED A SUBSEQUENT PROCEDURE THE NEXT DAY; AN EXPLORATORY LAPAROTOMY AND DIVERTING COLOSTOMY WERE PERFORMED.
29-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
05-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | BLUE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS STAPLER MALFUNCTION AND UNEVEN CUTTING EDGES AT THE CREATION OF THE TUBULAR STOMACH. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST BLUE RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | GST BLUE RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | POWERED PLUS STANDARD 45 | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE CARTRIDGE FELL OFF. THE SURGERY CONTINUED WITH A REPLACEMENT WITHOUT ANY PROBLEMS. THE CARTRIDGE COLOR WAS BLACK. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
20-02-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON 3000 45MM COMPACT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | nan
05-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2025 | Injury | intuitive surgical | 21-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT AFTER UNDERGOING A DA VINCI-ASSISTED ENDOMETRIOSIS SURGICAL PROCEDURE, THE PATIENT SUSTAINED A POSTOPERATIVE LEAK AND REQUIRED A SUBSEQUENT PROCEDURE. THE INITIAL ROBOTIC PROCEDURE WAS COMPLETED WITHOUT COMPLICATIONS, INCLUDING NO ISSUES WITH THE SUREFORM 45 STAPLER INSTRUMENT OR THE SUREFORM 45 BLACK RELOAD. THE PATIENT WAS DISCHARGED HOME AND RETURNED TO THEIR RESIDENCE. POSTOPERATIVELY, THE PATIENT DEVELOPED A LEAK AND PRESENTED TO THE EMERGENCY ROOM. THE PATIENT UNDERWENT A SECOND PROCEDURE. HOWEVER, SPECIFIC DETAILS REGARDING THE PROCEDURE ARE UNKNOWN. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO GATHER ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED.
22-12-2025 | Injury | ethicon endo-surgery | 21-01-2026 | PWRD 23MM CURVED CIRCULAR, 18CM SHAFT | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SAITO A, OTSUKA K, GOTO S, ARIYOSHI T, YAMASHITA T, MOTEGI K, KOMOTO M, KISHIMOTO Y, MURAKAMI M, AOKI T. COMPARISON OF ANASTOMOTIC STRICTURE RATES BETWEEN 23- AND 25-MM POWERED CIRCULAR STAPLERS IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS: A PROPENSITY-MATCHED STUDY. ESOPHAGUS. 2025 DEC 22. DOI: 10.1007/S10388-025-01171-2. EPUB AHEAD OF PRINT. THIS STUDY AIMED TO COMPARE STENOSIS RATES BETWEEN NEWLY INTRODUCED 23- AND 25-MM POWERED CIRCULAR STAPLER (CSS) IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS. FROM MAY 2022 TO FEBRUARY 2024, 126 PATIENTS WHO UNDERWENT THORACOSCOPIC MCKEOWN ESOPHAGECTOMY WITH RETROSTERNAL GASTRIC CONDUIT RECONSTRUCTION WERE RETROSPECTIVELY ANALYZED. THEY WERE CATEGORIZED INTO THE 23 MM (N = 52) AND 25-MM (N = 74) CS GROUPS. REPORTED COMPLICATIONS INCLUDE THE INCIDENCE OF ANASTOMOTIC STRICTURE WAS 18.0% (7 CASES) AND 12.8% (5 CASES) IN THE 23- AND 25-MM CS GROUPS, RESPECTIVELY. THE INCIDENCE OF ANASTOMOTIC LEAKAGE AND THE LEAKAGE LOCATION REVEALED NO SIGNIFICANT DIFFERENCES BETWEEN THE TWO GROUPS. IN CONCLUSION, 23- AND 25-MM POWERED CSS ACHIEVED COMPARABLE ANASTOMOTIC STRICTURE RATES IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS DURING MCKEOWN ESOPHAGECTOMY.
22-12-2025 | Injury | ethicon endo-surgery | 21-01-2026 | PWRD 23MM CURVED CIRCULAR, 18CM SHAFT | Patient Device Interaction Problem | A01 | Stenosis | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SAITO A, OTSUKA K, GOTO S, ARIYOSHI T, YAMASHITA T, MOTEGI K, KOMOTO M, KISHIMOTO Y, MURAKAMI M, AOKI T. COMPARISON OF ANASTOMOTIC STRICTURE RATES BETWEEN 23- AND 25-MM POWERED CIRCULAR STAPLERS IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS: A PROPENSITY-MATCHED STUDY. ESOPHAGUS. 2025 DEC 22. DOI: 10.1007/S10388-025-01171-2. EPUB AHEAD OF PRINT. THIS STUDY AIMED TO COMPARE STENOSIS RATES BETWEEN NEWLY INTRODUCED 23- AND 25-MM POWERED CIRCULAR STAPLER (CSS) IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS. FROM MAY 2022 TO FEBRUARY 2024, 126 PATIENTS WHO UNDERWENT THORACOSCOPIC MCKEOWN ESOPHAGECTOMY WITH RETROSTERNAL GASTRIC CONDUIT RECONSTRUCTION WERE RETROSPECTIVELY ANALYZED. THEY WERE CATEGORIZED INTO THE 23 MM (N = 52) AND 25-MM (N = 74) CS GROUPS. REPORTED COMPLICATIONS INCLUDE THE INCIDENCE OF ANASTOMOTIC STRICTURE WAS 18.0% (7 CASES) AND 12.8% (5 CASES) IN THE 23- AND 25-MM CS GROUPS, RESPECTIVELY. THE INCIDENCE OF ANASTOMOTIC LEAKAGE AND THE LEAKAGE LOCATION REVEALED NO SIGNIFICANT DIFFERENCES BETWEEN THE TWO GROUPS. IN CONCLUSION, 23- AND 25-MM POWERED CSS ACHIEVED COMPARABLE ANASTOMOTIC STRICTURE RATES IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS DURING MCKEOWN ESOPHAGECTOMY.
08-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON 3000 60MM COMPACT | Failure to Form Staple | A050704 | Unspecified Tissue Injury | nan
08-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON 3000 60MM COMPACT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
08-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON 3000 60MM COMPACT | No Apparent Adverse Event | A25 | Unspecified Tissue Injury | nan
08-01-2026 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON 3000 60MM COMPACT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
22-12-2025 | Injury | ethicon endo-surgery | 21-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SAITO A, OTSUKA K, GOTO S, ARIYOSHI T, YAMASHITA T, MOTEGI K, KOMOTO M, KISHIMOTO Y, MURAKAMI M, AOKI T. COMPARISON OF ANASTOMOTIC STRICTURE RATES BETWEEN 23- AND 25-MM POWERED CIRCULAR STAPLERS IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS: A PROPENSITY-MATCHED STUDY. ESOPHAGUS. 2025 DEC 22. DOI: 10.1007/S10388-025-01171-2. EPUB AHEAD OF PRINT. THIS STUDY AIMED TO COMPARE STENOSIS RATES BETWEEN NEWLY INTRODUCED 23- AND 25-MM POWERED CIRCULAR STAPLER (CSS) IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS. FROM (B)(6) 2022 TO (B)(6) 2024, 126 PATIENTS WHO UNDERWENT THORACOSCOPIC MCKEOWN ESOPHAGECTOMY WITH RETROSTERNAL GASTRIC CONDUIT RECONSTRUCTION WERE RETROSPECTIVELY ANALYZED. THEY WERE CATEGORIZED INTO THE 23 MM (N = 52) AND 25-MM (N = 74) CS GROUPS. REPORTED COMPLICATIONS INCLUDE THE INCIDENCE OF ANASTOMOTIC STRICTURE WAS 18.0% (7 CASES) AND 12.8% (5 CASES) IN THE 23- AND 25-MM CS GROUPS, RESPECTIVELY. THE INCIDENCE OF ANASTOMOTIC LEAKAGE AND THE LEAKAGE LOCATION REVEALED NO SIGNIFICANT DIFFERENCES BETWEEN THE TWO GROUPS. IN CONCLUSION, 23- AND 25-MM POWERED CSS ACHIEVED COMPARABLE ANASTOMOTIC STRICTURE RATES IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS DURING MCKEOWN ESOPHAGECTOMY.
22-12-2025 | Injury | ethicon endo-surgery | 21-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Patient Device Interaction Problem | A01 | Stenosis | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SAITO A, OTSUKA K, GOTO S, ARIYOSHI T, YAMASHITA T, MOTEGI K, KOMOTO M, KISHIMOTO Y, MURAKAMI M, AOKI T. COMPARISON OF ANASTOMOTIC STRICTURE RATES BETWEEN 23- AND 25-MM POWERED CIRCULAR STAPLERS IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS: A PROPENSITY-MATCHED STUDY. ESOPHAGUS. 2025 DEC 22. DOI: 10.1007/S10388-025-01171-2. EPUB AHEAD OF PRINT. THIS STUDY AIMED TO COMPARE STENOSIS RATES BETWEEN NEWLY INTRODUCED 23- AND 25-MM POWERED CIRCULAR STAPLER (CSS) IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS. FROM (B)(6) 2022 TO (B)(6) 2024, 126 PATIENTS WHO UNDERWENT THORACOSCOPIC MCKEOWN ESOPHAGECTOMY WITH RETROSTERNAL GASTRIC CONDUIT RECONSTRUCTION WERE RETROSPECTIVELY ANALYZED. THEY WERE CATEGORIZED INTO THE 23 MM (N = 52) AND 25-MM (N = 74) CS GROUPS. REPORTED COMPLICATIONS INCLUDE THE INCIDENCE OF ANASTOMOTIC STRICTURE WAS 18.0% (7 CASES) AND 12.8% (5 CASES) IN THE 23- AND 25-MM CS GROUPS, RESPECTIVELY. THE INCIDENCE OF ANASTOMOTIC LEAKAGE AND THE LEAKAGE LOCATION REVEALED NO SIGNIFICANT DIFFERENCES BETWEEN THE TWO GROUPS. IN CONCLUSION, 23- AND 25-MM POWERED CSS ACHIEVED COMPARABLE ANASTOMOTIC STRICTURE RATES IN CERVICAL ESOPHAGOGASTRIC ANASTOMOSIS DURING MCKEOWN ESOPHAGECTOMY.
04-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
04-12-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
03-11-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF THE PANCREAS, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE FARTHER AFTER IT FIRED A LITTLE. IT WAS FURTHER REPORTED THAT THEY COULD NOT OPEN THE JAW. OPENED THE JAW MANUALLY WITH BIG FORCE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
03-11-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF THE PANCREAS, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE FARTHER AFTER IT FIRED A LITTLE. IT WAS FURTHER REPORTED THAT THEY COULD NOT OPEN THE JAW. OPENED THE JAW MANUALLY WITH BIG FORCE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
09-05-2025 | Malfunction | ethicon endo-surgery | 21-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT ACTIVATE WHEN THEY TRIED TO FIRE DURING VASCULAR RESECTION. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
16-01-2026 | Injury | covidien | 21-01-2026 | ENDO GIA | Entrapment of Device | A150208 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN ON THE TRANSECTION OF THE GASTRIC BODY, THE TISSUE WAS P LACED/INSERTED IN THE JAW OF THE CARTRIDGE. THE GREEN BUTTON WAS PRESSED AND FIRED. AFTER FIRING TO THE END, THE DOWN ARROW KEY WAS PRESSED AGAIN. TO OPEN THE JAW, THE UP ARROW KEY WAS PRESSED, BUT THE KNIFE DID NOT RETURN SO THE JAW DID NOT OPEN. IT WAS CONFIRMED THAT EVEN AFTER IT WAS RESTARTED AND WITH THE MANUAL ADAPTER TOOL, THE KNIFE DID NOT RETURN. THEREFORE, THE PROCEDURE WAS SWITCHED TO LAPAROTOMY, AND THE AREA NEAR THE GASTRIC BODY CARTRIDGE WAS TRANSECTED WITH A SCALPEL, AND RELEASED FROM THE TISSUE.
16-01-2026 | Injury | covidien surgical | 21-01-2026 | SIGNIA | Entrapment of Device | A150208 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN ON THE TRANSECTION OF THE GASTRIC BODY, THE TISSUE WAS P LACED/INSERTED IN THE JAW OF THE CARTRIDGE. THE GREEN BUTTON WAS PRESSED AND FIRED. AFTER FIRING TO THE END, THE DOWN ARROW KEY WAS PRESSED AGAIN. TO OPEN THE JAW, THE UP ARROW KEY WAS PRESSED, BUT THE KNIFE DID NOT RETURN SO THE JAW DID NOT OPEN. IT WAS CONFIRMED THAT EVEN AFTER IT WAS RESTARTED AND WITH THE MANUAL ADAPTER TOOL, THE KNIFE DID NOT RETURN. THEREFORE, THE PROCEDURE WAS SWITCHED TO LAPAROTOMY, AND THE AREA NEAR THE GASTRIC BODY CARTRIDGE WAS TRANSECTED WITH A SCALPEL, AND RELEASED FROM THE TISSUE.
16-01-2026 | Injury | covidien surgical | 21-01-2026 | SIGNIA | Entrapment of Device | A150208 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN ON THE TRANSECTION OF THE GASTRIC BODY, THE TISSUE WAS P LACED/INSERTED IN THE JAW OF THE CARTRIDGE. THE GREEN BUTTON WAS PRESSED AND FIRED. AFTER FIRING TO THE END, THE DOWN ARROW KEY WAS PRESSED AGAIN. TO OPEN THE JAW, THE UP ARROW KEY WAS PRESSED, BUT THE KNIFE DID NOT RETURN SO THE JAW DID NOT OPEN. IT WAS CONFIRMED THAT EVEN AFTER IT WAS RESTARTED AND WITH THE MANUAL ADAPTER TOOL, THE KNIFE DID NOT RETURN. THEREFORE, THE PROCEDURE WAS SWITCHED TO LAPAROTOMY, AND THE AREA NEAR THE GASTRIC BODY CARTRIDGE WAS TRANSECTED WITH A SCALPEL, AND RELEASED FROM THE TISSUE.
26-12-2025 | Malfunction | covidien | 22-01-2026 | EEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: 25MM EEA CIRCULAR STAPLER BEING USED IN PATIENT. SURGEON FIRED STAPLER AS DIRECTED AND WHEN IT WAS BEING PULLED OUT, THE ANVIL DETACHED FROM THE STAPLER AND REMAINED INSIDE THE PATIENT. ANVIL WAS SUCCESSFULLY AND SAFELY RETRIEVED.
16-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON FLEX | Therapeutic or Diagnostic Output Failure | A090406 | No Clinical Signs, Symptoms or Conditions | Event Description: POWERED STAPLER DID NOT WORK WHEN SURGEON WAS READY TO STAPLE TISSUE IN APPENDECTOMY CASE.
16-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON FLEX | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: POWERED STAPLER DID NOT WORK WHEN SURGEON WAS READY TO STAPLE TISSUE IN APPENDECTOMY CASE.
29-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT IN THE MIDDLE OF AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS A PROBLEM WITH THE CARTRIDGE WHILE FIRING THE STAPLER. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
29-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT IN THE MIDDLE OF AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS A PROBLEM WITH THE CARTRIDGE WHILE FIRING THE STAPLER. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-12-2025 | Malfunction | covidien surgical | 22-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A VIDEO-ASSISTED THORACIC SURGERY (VATS) PROCEDURE, WHILE IN THE LUNG, THE POWERED STAPLER PARTIALLY FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2025 | Malfunction | covidien surgical | 22-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A VIDEO-ASSISTED THORACIC SURGERY (VATS) PROCEDURE, WHILE IN THE LUNG, THE POWERED STAPLER PARTIALLY FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2025 | Malfunction | covidien surgical | 22-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A VIDEO-ASSISTED THORACIC SURGERY (VATS) PROCEDURE, WHILE IN THE LUNG, THE POWERED STAPLER PARTIALLY FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ENDO ILS 29MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THAT THE TISSUE WAS NOT CUT COMPLETELY. THEN CHANGED THE NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
02-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE STAPLER'S LOAD WAS FIRED BUT DID NOT TRIGGER THE STAPLES. NO FURTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCES REPORTED.
27-10-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON LINEAR CUTTER 80MM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
23-09-2025 | Injury | coopersurgical | 22-01-2026 | INSORB SUBCUTICULAR STAPLER | Use of Device Problem | A23 | Wound Dehiscence | Event Description: PATIENT'S POST-OPERATIVE INCISION DEHISCED WITHIN A FEW HOURS OF USE OF INSORB STAPLER ON SITE. PATIENT REQUIRED TO RETURN TO THE OR FOR ADDITIONAL PROCEDURES TO REPAIR.
23-09-2025 | Injury | coopersurgical | 22-01-2026 | INSORB SUBCUTICULAR STAPLER | Therapeutic or Diagnostic Output Failure | A090406 | Wound Dehiscence | Event Description: PATIENT'S POST-OPERATIVE INCISION DEHISCED WITHIN A FEW HOURS OF USE OF INSORB STAPLER ON SITE. PATIENT REQUIRED TO RETURN TO THE OR FOR ADDITIONAL PROCEDURES TO REPAIR.
09-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER'S JAW COULD NOT BE OPENED AFTER FIRING. THEY USED MANUAL OVERRIDE LEVER BUT JAWS STILL DID NOT OPEN. SURGEON HAD TO CUT BOWEL ON LEFT AND RIGHT OF STAPLER TO REMOVE THE STAPLER TOGETHER WITH THE BOWEL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER'S JAW COULD NOT BE OPENED AFTER FIRING. THEY USED MANUAL OVERRIDE LEVER BUT JAWS STILL DID NOT OPEN. SURGEON HAD TO CUT BOWEL ON LEFT AND RIGHT OF STAPLER TO REMOVE THE STAPLER TOGETHER WITH THE BOWEL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | No Apparent Adverse Event | A25 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER'S JAW COULD NOT BE OPENED AFTER FIRING. THEY USED MANUAL OVERRIDE LEVER BUT JAWS STILL DID NOT OPEN. SURGEON HAD TO CUT BOWEL ON LEFT AND RIGHT OF STAPLER TO REMOVE THE STAPLER TOGETHER WITH THE BOWEL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER'S JAW COULD NOT BE OPENED AFTER FIRING. THEY USED MANUAL OVERRIDE LEVER BUT JAWS STILL DID NOT OPEN. SURGEON HAD TO CUT BOWEL ON LEFT AND RIGHT OF STAPLER TO REMOVE THE STAPLER TOGETHER WITH THE BOWEL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Problem | A1501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER'S JAW COULD NOT BE OPENED AFTER FIRING. THEY USED MANUAL OVERRIDE LEVER BUT JAWS STILL DID NOT OPEN. SURGEON HAD TO CUT BOWEL ON LEFT AND RIGHT OF STAPLER TO REMOVE THE STAPLER TOGETHER WITH THE BOWEL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER'S JAW COULD NOT BE OPENED AFTER FIRING. THEY USED MANUAL OVERRIDE LEVER BUT JAWS STILL DID NOT OPEN. SURGEON HAD TO CUT BOWEL ON LEFT AND RIGHT OF STAPLER TO REMOVE THE STAPLER TOGETHER WITH THE BOWEL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE AND HAD NO MOTOR SOUND. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ETS FLEX ARTICNG LNR CUTR 45MM | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS LOCKED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
21-08-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT WORK. THEY REPLACED THE CARTRIDGE AND IT WORKED PROPERLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
17-04-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | nan
17-10-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLE LINE WAS JAGGED AND DID NOT CUT NICELY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
03-11-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE KNIFE STOPPED UPON FIRING. THEY MANUALLY OPENED THE DEVICE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
03-11-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON 3000 60MM STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE KNIFE STOPPED UPON FIRING. THEY MANUALLY OPENED THE DEVICE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-11-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Manufacturing, Packaging or Shipping Problem | A02 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE PACKAGING OF THE DEVICE WAS PIERCED UPON OPENING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
16-01-2026 | Malfunction | covidien surgical | 22-01-2026 | ABSORBATACK | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC INGUINAL HERNIA REPAIR, DURING MESH FIXATION, NONE OF THE TACKS FIRED WERE ABLE TO HOLD ONTO THE TISSUE CORRECTLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2025 | Malfunction | covidien surgical | 22-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE AFTER FIRING, THE STAPLE LINE DID NOT HOLD AND SUBSEQUENTLY OPENED UP. IT OCCURRED ON TWO RELOADS. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
29-11-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | nan
26-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | PROXIMATE*RL VASCULAR LIN STAP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
22-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON 3000 60MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOT-ASSISTED ESOPHAGECTOMY IT WAS OBSERVED THAT THE DEVICE DID NOT ACTIVATE, SO A NEW ONE WAS TAKEN OUT AND IT COULD BE FIRED A FEW SHOTS. AFTERWARDS, IT DID NOT ACTIVATE EVEN WHEN CLAMPED THE TISSUE, SO IT WAS HANDLED WITH A COMPETITIVE PRODUCT IN THE END. THE CARTRIDGE COLOR WAS BLUE. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ILS 29MM CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN ROBOT-ASSISTED SURGERY, THE ROTATE KNOB WAS ROTATED TWICE AS USUAL WHEN REMOVING AFTER THE FIRING, BUT THE DOCTOR FELT LIKE IT WAS DIFFICULT TO REMOVE, SO HE ROTATED IT ONE MORE TIME AND TRIED TO REMOVE IT. HE PULLED THE DEVICE SLOWLY BECAUSE IT FELT LIKE IT WAS DIFFICULT TO REMOVE, BUT THE ANVIL CAME OFF INSIDE THE INTESTINE WHEN HE APPLIED A LITTLE FORCE. AFTER THAT, WHILE CHECKING WITH THE ENDOSCOPE, IT WAS ABLE TO PUSH THE REMOVED ANVIL WITH FORCEPS AND REMOVE FROM THE ANUS. FORTUNATELY, NOTHING SERIOUS HAPPENED, AND THE PATIENT WAS DISCHARGED SAFELY. THE REPORTED DEVICE WAS USED AS IT WAS TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
03-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON 3000 60MM STANDARD | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY, THE DEVICE CLAMPED THE CLIP, AND IT COULD NOT BE USED PROPERLY AFTERWARDS. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
30-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ILS 29MM CURVED | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ILS 29MM CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ILS 29MM CURVED | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ILS 29MM CURVED | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-01-2026 | Malfunction | covidien surgical | 22-01-2026 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, A STAPLE WAS MISSING FROM THE UPPER STAPLE LINE ROW. ANOTHER RELOAD WAS USED.
01-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN PROCEDURE, THERE WAS NO PROBLEM WITH THE PRODUCT, AND THE HOSPITAL ONLY REQUESTED A COMPLAINT ABOUT THE OUTER BOX AND PACKAGING METHOD. THE STERILIZATION SEAL OF THE CDH MAIN BODY WAS SANDWICHED IN THE OPENING OF THE OUTER BOX, AND THE STERILIZATION SEAL COULD EASILY OPEN WHEN THE OUTER BOX WAS OPENED. SINCE THIS HAPPENED TWICE IN A ROW, THE HOSPITAL ASKED FOR SOME IMPROVEMENT TO PREVENT THIS FROM HAPPENING AGAIN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING WHEN BOTH THE STAPLE AND CUT LINES WERE COMPLETE. CHANGED ANOTHER FOUR TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING WHEN BOTH THE STAPLE AND CUT LINES WERE COMPLETE. CHANGED ANOTHER FOUR TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
03-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED PLUS STANDARD 45 | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. AND THEN DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. AND THEN DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. AND THEN DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED PLUS STANDARD 45 | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. AND THEN DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-10-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | GST GOLD RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
07-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-01-2026 | Malfunction | covidien | 22-01-2026 | EEA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN IVOR LEWIS ESOPHAGECTOMY, THE SUTURE WAS DIFFICULT TO REMOVE FROM THE ORVIL (ORAL ANVIL) PRIOR TO FIRING. EXCESSIVE FORCE WAS REQUIRED TO ATTEMPT REMOVAL, AND THE SUTURE WOULD NOT MOVE DESPITE MULTIPLE ATTEMPTS, INCLUDING MOVING THE ANVIL SIDE TO SIDE, IN AND OUT, AND TWISTING. THE SURGEON DECIDED TO FIRE THE STAPLER WITH THE SUTURE POTENTIALLY IN THE STAPLE LINE. AFTER FIRING, THE SUTURE REMAINED ON THE ANVIL. UPON REMOVAL OF THE STAPLER, SUTURE PIECES WERE RETRIEVED FROM THE CHEST CAVITY, WITH A TOTAL OF THREE PIECES IDENTIFIED: THE SUTURE REEL AND TWO ADDITIONAL PIECES. THE PROCEDURE WAS EXTENDED BY 15 MINUTES DUE TO THESE ISSUES. THE PATIENT WAS ALIVE WITH NO INJURY, AND WAS REPORTEDLY RECOVERING AS EXPECTED ON POSTOPERATIVE DAY TWO.
20-01-2026 | Malfunction | covidien | 22-01-2026 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN IVOR LEWIS ESOPHAGECTOMY, THE SUTURE WAS DIFFICULT TO REMOVE FROM THE ORVIL (ORAL ANVIL) PRIOR TO FIRING. EXCESSIVE FORCE WAS REQUIRED TO ATTEMPT REMOVAL, AND THE SUTURE WOULD NOT MOVE DESPITE MULTIPLE ATTEMPTS, INCLUDING MOVING THE ANVIL SIDE TO SIDE, IN AND OUT, AND TWISTING. THE SURGEON DECIDED TO FIRE THE STAPLER WITH THE SUTURE POTENTIALLY IN THE STAPLE LINE. AFTER FIRING, THE SUTURE REMAINED ON THE ANVIL. UPON REMOVAL OF THE STAPLER, SUTURE PIECES WERE RETRIEVED FROM THE CHEST CAVITY, WITH A TOTAL OF THREE PIECES IDENTIFIED: THE SUTURE REEL AND TWO ADDITIONAL PIECES. THE PROCEDURE WAS EXTENDED BY 15 MINUTES DUE TO THESE ISSUES. THE PATIENT WAS ALIVE WITH NO INJURY, AND WAS REPORTEDLY RECOVERING AS EXPECTED ON POSTOPERATIVE DAY TWO.
06-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
11-02-2026 | Injury | covidien | 22-01-2026 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Stenosis | Event Description: ACCORDING TO THE REPORTER, EIGHT WEEKS POST-OPERATIVE OF A LOW ANTERIOR RESECTION (LAR) FOR ENDOMETRIOSIS PERFORMED ON (B)(6) 2025 WHERE A CIRCULAR STAPLER WAS USED, THE PATIENT HAD STRICTURE. TO TREAT THE PATIENT, THE STRICTURE WAS BALLOONED. THE PATIENT RECOVERED.
11-02-2026 | Injury | covidien | 22-01-2026 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, EIGHT WEEKS POST-OPERATIVE OF A LOW ANTERIOR RESECTION (LAR) FOR ENDOMETRIOSIS PERFORMED ON (B)(6) 2025 WHERE A CIRCULAR STAPLER WAS USED, THE PATIENT HAD STRICTURE. TO TREAT THE PATIENT, THE STRICTURE WAS BALLOONED. THE PATIENT RECOVERED.
11-02-2026 | Injury | covidien | 22-01-2026 | UNKNOWN EEA | Insufficient Information | E2401 | Stenosis | Event Description: ACCORDING TO THE REPORTER, EIGHT WEEKS POST-OPERATIVE OF A LOW ANTERIOR RESECTION (LAR) FOR ENDOMETRIOSIS PERFORMED ON (B)(6) 2025 WHERE A CIRCULAR STAPLER WAS USED, THE PATIENT HAD STRICTURE. TO TREAT THE PATIENT, THE STRICTURE WAS BALLOONED. THE PATIENT RECOVERED.
11-02-2026 | Injury | covidien | 22-01-2026 | UNKNOWN EEA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, EIGHT WEEKS POST-OPERATIVE OF A LOW ANTERIOR RESECTION (LAR) FOR ENDOMETRIOSIS PERFORMED ON (B)(6) 2025 WHERE A CIRCULAR STAPLER WAS USED, THE PATIENT HAD STRICTURE. TO TREAT THE PATIENT, THE STRICTURE WAS BALLOONED. THE PATIENT RECOVERED.
01-01-2026 | Malfunction | ethicon endo-surgery | 22-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN PROCEDURE, THERE WAS NO PROBLEM WITH THE PRODUCT, AND THE HOSPITAL ONLY REQUESTED A COMPLAINT ABOUT THE OUTER BOX AND PACKAGING METHOD. THE STERILIZATION SEAL OF THE CDH MAIN BODY WAS SANDWICHED IN THE OPENING OF THE OUTER BOX, AND THE STERILIZATION SEAL COULD EASILY OPEN WHEN THE OUTER BOX WAS OPENED. SINCE THIS HAPPENED TWICE IN A ROW, THE HOSPITAL ASKED FOR SOME IMPROVEMENT TO PREVENT THIS FROM HAPPENING AGAIN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
20-01-2026 | Malfunction | covidien surgical | 22-01-2026 | SIGNIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y GASTRIC BYPASS, THE FIRST RELOAD FIRED APPROPRIATELY FOR THE POUCH. THE SECOND RELOAD FIRED APPROPRIATELY FOR 2 CM AND STOPPED, THEN SURGED FORWARD IN JERKY MOTION FOR THE NEXT 4 CM. THE SURGEON KEPT THE TRIGGER DEPRESSED THE ENTIRE TIME. AFTER RETRACTING THE BLADE, THE TISSUE ON THE FIRST 2 CM WAS SATISFACTORY, BUT THE REMAINDER OF THE LINE HAD BEEN CUT FOR 4 CM BUT NOT STAPLED. IT APPEARED THAT THIS TISSUE HAD BEEN TORN AND/OR CUT, OR THE ISSUE WAS CAUSED BY A STAPLE DRAGGED BY AN I-BEAM. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE POUCH WAS MADE A LITTLE TIGHTER.
20-01-2026 | Malfunction | covidien surgical | 22-01-2026 | SIGNIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y GASTRIC BYPASS, THE FIRST RELOAD FIRED APPROPRIATELY FOR THE POUCH. THE SECOND RELOAD FIRED APPROPRIATELY FOR 2 CM AND STOPPED, THEN SURGED FORWARD IN JERKY MOTION FOR THE NEXT 4 CM. THE SURGEON KEPT THE TRIGGER DEPRESSED THE ENTIRE TIME. AFTER RETRACTING THE BLADE, THE TISSUE ON THE FIRST 2 CM WAS SATISFACTORY, BUT THE REMAINDER OF THE LINE HAD BEEN CUT FOR 4 CM BUT NOT STAPLED. IT APPEARED THAT THIS TISSUE HAD BEEN TORN AND/OR CUT, OR THE ISSUE WAS CAUSED BY A STAPLE DRAGGED BY AN I-BEAM. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE POUCH WAS MADE A LITTLE TIGHTER.
09-12-2025 | Malfunction | ethicon endo-surgery | 22-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN HEPATECTOMY, IT WAS OBSERVED THAT THE DEVICE COULD NOT BE FIRED WHEN TRIED TO FIRE AS USUAL DURING RESECTING THE HEPATIC VEIN. THE CUSTOMER GAVE UP ON USING THE INSTRUMENT AND FINISHED BY LIGATION WITH THE SUTURE. THE CARTRIDGE COLOR WAS WHITE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
20-01-2026 | Malfunction | covidien surgical | 22-01-2026 | ENDO GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y GASTRIC BYPASS, THE FIRST RELOAD FIRED APPROPRIATELY FOR THE POUCH. THE SECOND RELOAD FIRED APPROPRIATELY FOR 2 CM AND STOPPED, THEN SURGED FORWARD IN JERKY MOTION FOR THE NEXT 4 CM. THE SURGEON KEPT THE TRIGGER DEPRESSED THE ENTIRE TIME. AFTER RETRACTING THE BLADE, THE TISSUE ON THE FIRST 2 CM WAS SATISFACTORY, BUT THE REMAINDER OF THE LINE HAD BEEN CUT FOR 4 CM BUT NOT STAPLED. IT APPEARED THAT THIS TISSUE HAD BEEN TORN AND/OR CUT, OR THE ISSUE WAS CAUSED BY A STAPLE DRAGGED BY AN I-BEAM. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE POUCH WAS MADE A LITTLE TIGHTER.
20-01-2026 | Malfunction | covidien surgical | 22-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y GASTRIC BYPASS, THE FIRST RELOAD FIRED APPROPRIATELY FOR THE POUCH. THE SECOND RELOAD FIRED APPROPRIATELY FOR 2 CM AND STOPPED, THEN SURGED FORWARD IN JERKY MOTION FOR THE NEXT 4 CM. THE SURGEON KEPT THE TRIGGER DEPRESSED THE ENTIRE TIME. AFTER RETRACTING THE BLADE, THE TISSUE ON THE FIRST 2 CM WAS SATISFACTORY, BUT THE REMAINDER OF THE LINE HAD BEEN CUT FOR 4 CM BUT NOT STAPLED. IT APPEARED THAT THIS TISSUE HAD BEEN TORN AND/OR CUT, OR THE ISSUE WAS CAUSED BY A STAPLE DRAGGED BY AN I-BEAM. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE POUCH WAS MADE A LITTLE TIGHTER.
20-01-2026 | Malfunction | covidien surgical | 22-01-2026 | ENDO GIA | Failure to Cut | A050702 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y GASTRIC BYPASS, THE FIRST RELOAD FIRED APPROPRIATELY FOR THE POUCH. THE SECOND RELOAD FIRED APPROPRIATELY FOR 2 CM AND STOPPED, THEN SURGED FORWARD IN JERKY MOTION FOR THE NEXT 4 CM. THE SURGEON KEPT THE TRIGGER DEPRESSED THE ENTIRE TIME. AFTER RETRACTING THE BLADE, THE TISSUE ON THE FIRST 2 CM WAS SATISFACTORY, BUT THE REMAINDER OF THE LINE HAD BEEN CUT FOR 4 CM BUT NOT STAPLED. IT APPEARED THAT THIS TISSUE HAD BEEN TORN AND/OR CUT, OR THE ISSUE WAS CAUSED BY A STAPLE DRAGGED BY AN I-BEAM. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE POUCH WAS MADE A LITTLE TIGHTER.
20-01-2026 | Malfunction | covidien surgical | 22-01-2026 | ENDO GIA | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y GASTRIC BYPASS, THE FIRST RELOAD FIRED APPROPRIATELY FOR THE POUCH. THE SECOND RELOAD FIRED APPROPRIATELY FOR 2 CM AND STOPPED, THEN SURGED FORWARD IN JERKY MOTION FOR THE NEXT 4 CM. THE SURGEON KEPT THE TRIGGER DEPRESSED THE ENTIRE TIME. AFTER RETRACTING THE BLADE, THE TISSUE ON THE FIRST 2 CM WAS SATISFACTORY, BUT THE REMAINDER OF THE LINE HAD BEEN CUT FOR 4 CM BUT NOT STAPLED. IT APPEARED THAT THIS TISSUE HAD BEEN TORN AND/OR CUT, OR THE ISSUE WAS CAUSED BY A STAPLE DRAGGED BY AN I-BEAM. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE POUCH WAS MADE A LITTLE TIGHTER.
02-01-2026 | Malfunction | covidien surgical | 22-01-2026 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRECTOMY, DURING SETUP, THE ADAPTER GO SEPARATED FROM THE SHELL PRIOR TO FIRING. A MANUAL STAPLER WITH THE SAME RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-01-2026 | Malfunction | covidien surgical | 22-01-2026 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRECTOMY, DURING SETUP, THE ADAPTER GO SEPARATED FROM THE SHELL PRIOR TO FIRING. A MANUAL STAPLER WITH THE SAME RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-10-2025 | Injury | covidien surgical | 22-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Stenosis | nan
07-01-2026 | Malfunction | covidien surgical | 22-01-2026 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DISTAL PANCREATECTOMY, WHEN THE CARTRIDGE WAS ATTACHED TO THE HANDLE, THE JAW OPENING AND CLOSING CHECK HAD NO PROBLEMS. AFTER THAT, THE TISSUE WAS APPROACHED, THE GREEN BUTTON WAS PRESSED, AND THE HANDLE WAS NOT SQUEEZED, BUT THE DEVICE DID NOT FIRE AT ALL. A NEW HANDLE AND CARTRIDGE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2026 | Malfunction | covidien | 22-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DISTAL PANCREATECTOMY, WHEN THE CARTRIDGE WAS ATTACHED TO THE HANDLE, THE JAW OPENING AND CLOSING CHECK HAD NO PROBLEMS. AFTER THAT, THE TISSUE WAS APPROACHED, THE GREEN BUTTON WAS PRESSED, AND THE HANDLE WAS NOT SQUEEZED, BUT THE DEVICE DID NOT FIRE AT ALL. A NEW HANDLE AND CARTRIDGE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien | 22-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY, WHILE ASSEMBLING THE STAPLER, THE RELOAD WAS ALREADY CONNECTED TO THE MANUAL HANDLE, AND THE JAWS WERE PARTIALLY OPENED. THE RELOAD WAS NOT USED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-08-2025 | Malfunction | covidien surgical | 22-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE SIX RELOADS USED WITH A POWERED HANDLE AND AN XL ADAPTER PARTIALLY FIRED. TO RESOLVE THE ISSUE, NEW RELOADS WERE USED. NO PATIENT INJURY OCCURRED.
17-12-2025 | Injury | covidien | 22-01-2026 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT A ROBOTIC LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR) INVOLVING AN ANASTOMOSIS RESULTED IN A LEAK EIGHT DAYS POSTOPERATIVELY. A LEAK TEST WAS PERFORMED AND WAS NEGATIVE FOR AIR BUBBLES AND METHYLENE BLUE.
17-12-2025 | Injury | covidien | 22-01-2026 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: IT WAS REPORTED THAT A ROBOTIC LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR) INVOLVING AN ANASTOMOSIS RESULTED IN A LEAK EIGHT DAYS POSTOPERATIVELY. A LEAK TEST WAS PERFORMED AND WAS NEGATIVE FOR AIR BUBBLES AND METHYLENE BLUE.
17-12-2025 | Injury | covidien | 22-01-2026 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Purulent Discharge | Event Description: IT WAS REPORTED THAT A ROBOTIC LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR) INVOLVING AN ANASTOMOSIS RESULTED IN A LEAK EIGHT DAYS POSTOPERATIVELY. A LEAK TEST WAS PERFORMED AND WAS NEGATIVE FOR AIR BUBBLES AND METHYLENE BLUE.
17-12-2025 | Injury | covidien | 22-01-2026 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT A ROBOTIC LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR) INVOLVING AN ANASTOMOSIS RESULTED IN A LEAK EIGHT DAYS POSTOPERATIVELY. A LEAK TEST WAS PERFORMED AND WAS NEGATIVE FOR AIR BUBBLES AND METHYLENE BLUE.
17-12-2025 | Injury | covidien | 22-01-2026 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: IT WAS REPORTED THAT A ROBOTIC LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR) INVOLVING AN ANASTOMOSIS RESULTED IN A LEAK EIGHT DAYS POSTOPERATIVELY. A LEAK TEST WAS PERFORMED AND WAS NEGATIVE FOR AIR BUBBLES AND METHYLENE BLUE.
21-01-2026 | Malfunction | covidien | 23-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RADICAL SURGERY FOR RECTAL CANCER, DURING THE RECTAL SEVERING STEP, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, ALTHOUGH THE GREEN BUTTON COULD BE PRESSED. THE DEVICES WERE REPLACED WITH NEW ONES TO COMPLETE THE PROCEDURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
21-01-2026 | Malfunction | covidien | 23-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RADICAL SURGERY FOR RECTAL CANCER, DURING THE RECTAL SEVERING STEP, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, ALTHOUGH THE GREEN BUTTON COULD BE PRESSED. THE DEVICES WERE REPLACED WITH NEW ONES TO COMPLETE THE PROCEDURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-01-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE FAILED TO CUT THE BOWEL WHEN INSIDE THE PATIENT. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-08-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 60MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE BLADE ON THE DEVICE STOPPED WHILE FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-11-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 60MM LONG | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS LAP ROUX-EN-Y PROCEDURE, THEY USED THE STAPLER 4 TIMES AND IT WAS WORKING/CUTTING PROPERLY. BUT ON THE 5TH LOAD, IT WAS OBSERVED THAT THE KNIFE OF THE STAPLER WAS NOT CUTTING PROPERLY ANYMORE. ON FIRST HORIZONTAL FIRING OF THE STAPLER ON THE GASTRIC POUCH CREATION, IT FIRED BUT DID NOT CUT TISSUE - IT PUSHED TISSUE FORWARD. UPON INSPECTION OF THE JAW, THERE WAS SIGNIFICANT TISSUE BUILDUP IN THE JAW SO NURSE TRIED TO USE TWEEZERS TO REMOVE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE WITH A DELAY OF 15 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-11-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 60MM LONG | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS LAP ROUX-EN-Y PROCEDURE, THEY USED THE STAPLER 4 TIMES AND IT WAS WORKING/CUTTING PROPERLY. BUT ON THE 5TH LOAD, IT WAS OBSERVED THAT THE KNIFE OF THE STAPLER WAS NOT CUTTING PROPERLY ANYMORE. ON FIRST HORIZONTAL FIRING OF THE STAPLER ON THE GASTRIC POUCH CREATION, IT FIRED BUT DID NOT CUT TISSUE - IT PUSHED TISSUE FORWARD. UPON INSPECTION OF THE JAW, THERE WAS SIGNIFICANT TISSUE BUILDUP IN THE JAW SO NURSE TRIED TO USE TWEEZERS TO REMOVE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE WITH A DELAY OF 15 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT TURN ON WHEN THE BATTERY WAS PLUGGED IN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE AND HAD AN ABNORMAL WEAK SOUND (FIRING SPEED WAS ABNORMAL SLOW). ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | POWERED PLUS STANDARD 45 | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE AND HAD AN ABNORMAL WEAK SOUND (FIRING SPEED WAS ABNORMAL SLOW). ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | POWERED PLUS STANDARD 45 | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE AND HAD AN ABNORMAL WEAK SOUND (FIRING SPEED WAS ABNORMAL SLOW). ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 60MM STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THEY COULD NOT CLOSE THE JAW AFTER CLAMPING THE TISSUE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A LOW ANTERIOR RESECTION, ABOUT FOUR HOURS AFTER INSTALLING THE BATTERY, THE DEVICE WAS TRIED TO FIRE, BUT THE BATTERY DID NOT ACTIVATE. THE BATTERY WAS REPLACED, AND IT COULD BE ACTIVATED, SO THERE WAS NO IMPACT ON THE PATIENT. AS IT WAS UNCLEAR WHICH BATTERY DID NOT ACTIVATE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-01-2026 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN LOBECTOMY, WHEN FIRING TO THE PULMONARY ARTERY, THE DEVICE COULD NOT BE FIRED DUE TO IT BEING IMMEDIATELY LOCKED OUT AFTER HEARING THE DRIVING SOUND OF THE BATTERY. AFTER THE KNIFE WAS RETURNED WITH THE KNIFE REVERSE SWITCH, OPENED AND CLOSED, AND IT WAS CONFIRMED THAT THERE WAS NO DAMAGE TO BLOOD VESSEL, BLEEDING, OR OTHER EFFECTS ON THE PATIENT. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
12-01-2026 | Malfunction | covidien | 23-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC TUMOR RESECTION, AFTER ASSEMBLING THE DEVICE AND PRESSING THE GREEN BUTTON, IT WAS FOUND THAT THE HANDLE COULD NOT BE FULLY SQUEEZED. THIS ISSUE OCCURRED ON TWO HANDLES. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
12-01-2026 | Malfunction | covidien | 23-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC TUMOR RESECTION, AFTER ASSEMBLING THE DEVICE AND PRESSING THE GREEN BUTTON, IT WAS FOUND THAT THE HANDLE COULD NOT BE FULLY SQUEEZED. THIS ISSUE OCCURRED ON TWO HANDLES. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | ethicon endo-surgery | 23-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ENDOSCOPIC PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED IN PLACE AND WOULD NOT CUT OR SEAL. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
11-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE KNIFE STOPPED HALFWAY THROUGH. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 23-01-2026 | ENDO ILS 21MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
23-01-2026 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE NOT DEPLOYED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
11-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 45 STANDARD | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE FIRST STAPLER FAILED TO ARTICULATE. A SECOND STAPLER WAS THEN USED, WHICH ARTICULATED ACROSS THE APPENDIX; HOWEVER, AFTER FIRING, IT DID NOT OPEN AND REQUIRED MANUAL OVERRIDE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
11-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE FIRST STAPLER FAILED TO ARTICULATE. A SECOND STAPLER WAS THEN USED, WHICH ARTICULATED ACROSS THE APPENDIX; HOWEVER, AFTER FIRING, IT DID NOT OPEN AND REQUIRED MANUAL OVERRIDE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
10-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOT ASSISTED DISTAL GASTRECTOMY, THE DEVICE COULD BE FIRED AT THE 1ST FIRING, BUT IT COULD NOT BE FIRED WHEN TRYING TO LOAD THE 2ND SHOT. WHEN THE CUSTOMER LOOKED CLOSELY, IT WAS NOTICED THAT THE CARTRIDGE PAN HAD COME OFF AND REMAINED AT THE MAIN BODY SIDE. IT WAS STILL ATTACHED EVEN WHEN LOOKED AT THE RETRIEVED PRODUCT. A NEW PSEE60A AND GST60W WERE TAKEN OUT, AND THEY WERE USED TO COMPLETE THE SURGERY WITHOUT ANY PROBLEMS. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
10-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 60MM STANDARD | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOT ASSISTED DISTAL GASTRECTOMY, THE DEVICE COULD BE FIRED AT THE 1ST FIRING, BUT IT COULD NOT BE FIRED WHEN TRYING TO LOAD THE 2ND SHOT. WHEN THE CUSTOMER LOOKED CLOSELY, IT WAS NOTICED THAT THE CARTRIDGE PAN HAD COME OFF AND REMAINED AT THE MAIN BODY SIDE. IT WAS STILL ATTACHED EVEN WHEN LOOKED AT THE RETRIEVED PRODUCT. A NEW PSEE60A AND GST60W WERE TAKEN OUT, AND THEY WERE USED TO COMPLETE THE SURGERY WITHOUT ANY PROBLEMS. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
15-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON 3000 45 STANDARD | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, IT WAS OBSERVED THAT THE DEVICE MISFIRED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A DELAY OF 5-10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | CARTRIDGE FINAL ASSY 45MM GREEN | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. SUTURE WAS USED TO OVERSEW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
15-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | CARTRIDGE FINAL ASSY 45MM GREEN | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THERE WAS OOZING AFTER FIRING WHEN BOTH STAPLE AND CUT LINES WERE COMPLETE. SUTURE WAS USED TO OVERSEW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
16-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | GST BLACK RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-01-2026 | Malfunction | covidien | 23-01-2026 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHEN THE STOMACH WAS REMOVED, THE STAPLER WAS CLOSED AND COULD NOT BE OPENED. THE RESET BUTTON WAS PULLED BACK TO ITS ORIGINAL POSITION, BUT THE STAPLER DID NOT OPEN. THE USER ATTEMPTED TO RE-CLOSE AND OPEN THE STAPLER, AND PULLED THE RESET BUTTON TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2026 | Malfunction | ethicon endo-surgery | 23-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | nan
08-01-2026 | Malfunction | ethicon endo-surgery | 23-01-2026 | PWRD 23MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-01-2026 | Malfunction | covidien | 23-01-2026 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHEN THE STOMACH WAS REMOVED, THE STAPLER WAS CLOSED AND COULD NOT BE OPENED. THE RESET BUTTON WAS PULLED BACK TO ITS ORIGINAL POSITION, BUT THE STAPLER DID NOT OPEN. THE USER ATTEMPTED TO RE-CLOSE AND OPEN THE STAPLER, AND PULLED THE RESET BUTTON TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2025 | Injury | covidien surgical | 23-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT POST-OPERATIVELY, THERE WAS COMPLETE DEHISCENCE OF THE STAPLED AREA WHICH NECESSITATED REOPERATION.
01-01-2026 | Injury | ethicon endo-surgery | 23-01-2026 | ECHELON LINEAR CUTTER 100MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER AN UNK PROCEDURE, THE SURGEON HAS REPORTED A POST OPERATIVE LEAK FOLLOWING THE PROCEDURE USING THE GLC100 LINEAR STAPLER. PROCEDURE WAS COMPLETED WITH THE DEVICE AS NORMAL, HOWEVER THE SURGEON HAD POST OPERATIVE LEAKING WITH HIS PATIENT. HE ATTRIBUTED THIS TO THE STAPLER AND HAS REVERTED TO OUR TLC STAPLER.
07-01-2026 | Malfunction | covidien | 23-01-2026 | ENDO GIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ESOPHAGECTOMY, AFTER ESOPHAGEAL RESECTION, THE RETURN KNOB DID NOT RETURN, AND THE JAW WOULD NOT OPEN, AND IT WAS DEFINITELY HARDER/STIFFER THAN WHEN USING IT AS USUAL. TO RESOLVE THE ISSUE, THE ARTICULATION RETURNED TO NEUTRAL, THE KNIFE WAS SOMEHOW RETRACTED ALL THE WAY, AND THE JAW OPENED. THERE WAS NO PATIENT INJURY.
07-01-2026 | Malfunction | covidien | 23-01-2026 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ESOPHAGECTOMY, AFTER ESOPHAGEAL RESECTION, THE RETURN KNOB DID NOT RETURN, AND THE JAW WOULD NOT OPEN, AND IT WAS DEFINITELY HARDER/STIFFER THAN WHEN USING IT AS USUAL. TO RESOLVE THE ISSUE, THE ARTICULATION RETURNED TO NEUTRAL, THE KNIFE WAS SOMEHOW RETRACTED ALL THE WAY, AND THE JAW OPENED. THERE WAS NO PATIENT INJURY.
01-01-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS DIFFICULT TO ARTICULATE WHILE INSIDE THE PATIENT. IT WAS FURTHER REPORTED THAT THE STAPLER FIRED INCOMPLETE AS IT WAS ONLY SHOWING ONE ROW OR STAPLES; AND NO CUT WAS ACHIEVED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN LAPAROSCOPIC RESPIRATORY SURGERY, THE DEVICE COULD NOT BE FIRED WHEN RESECTING THE PULMONARY VEIN. THE SLEDS HAD MOVED TOWARD THE TIP WHEN THE SALES REP CHECKED THE CARTRIDGE ON-SITE. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
29-12-2025 | Injury | davis & geck caribe | 23-01-2026 | UNKNOWN ENDO STITCH INSTRUMENT | Break | A0401 | Foreign Body In Patient | nan
17-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE AND HAD NO MOTOR SOUND ON THE FIRING. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL DD NOT FIRE AND HAD NO MOTOR SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
18-12-2025 | Malfunction | ethicon endo-surgery | 23-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO ANOTHER TWO TO CONTINUE THE SURGERY BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-01-2026 | Malfunction | covidien surgical | 24-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SLEEVE PROCEDURE, A LEAK WAS SEEN AROUND THE STAPLE LINE AFTER ALL STAPLES WERE FIRED, WITH NO MALFORMED STAPLES OBSERVED. A LEAK TEST WAS PERFORMED USING 50¿100 ML OF METHYLENE BLUE, WHICH CONFIRMED A POSITIVE LEAK AROUND THE STAPLE LINE DURING THE PROCEDURE. THREE RELOADS WERE USED AND HAD THE SAME ISSUE. SUTURING WAS REQUIRED AS A STAPLE LINE REPLACEMENT TO ADDRESS THE LEAK.
20-01-2026 | Malfunction | covidien surgical | 24-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SLEEVE PROCEDURE, A LEAK WAS SEEN AROUND THE STAPLE LINE AFTER ALL STAPLES WERE FIRED, WITH NO MALFORMED STAPLES OBSERVED. A LEAK TEST WAS PERFORMED USING 50¿100 ML OF METHYLENE BLUE, WHICH CONFIRMED A POSITIVE LEAK AROUND THE STAPLE LINE DURING THE PROCEDURE. THREE RELOADS WERE USED AND HAD THE SAME ISSUE. SUTURING WAS REQUIRED AS A STAPLE LINE REPLACEMENT TO ADDRESS THE LEAK.
20-01-2026 | Malfunction | covidien surgical | 24-01-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SLEEVE PROCEDURE, A LEAK WAS SEEN AROUND THE STAPLE LINE AFTER ALL STAPLES WERE FIRED, WITH NO MALFORMED STAPLES OBSERVED. A LEAK TEST WAS PERFORMED USING 50¿100 ML OF METHYLENE BLUE, WHICH CONFIRMED A POSITIVE LEAK AROUND THE STAPLE LINE DURING THE PROCEDURE. THREE RELOADS WERE USED AND HAD THE SAME ISSUE. SUTURING WAS REQUIRED AS A STAPLE LINE REPLACEMENT TO ADDRESS THE LEAK.
18-01-2026 | Malfunction | covidien | 24-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY, THE DEVICE WAS TOO RIGID AND DIFFICULT TO FIRE. THE MANUAL HANDLE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
06-01-2026 | Malfunction | covidien surgical | 26-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, IN A SINGLE-PORT THORACOSCOPIC LOBECTOMY, WHEN PREPARING TO SEVER THE LUNG PARENCHYMA, THERE WAS POOR STAPLE FORMATION. THE STAPLES WERE UNABLE TO FORM A COMPLETE B-SHAPE. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. SUTURING WAS DONE TO FIX THE STAPLE LINE TO RESOLVE THE ISSUE.
12-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
04-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
04-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
13-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST BLUE RELOAD 45MM 6 ROW | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, NOTED THAT CARTRIDGE BASE LOOSE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST GREEN RELOAD 60MM 6 ROW | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY AND BEFORE USED ON THE PATIENT, NOTED THE PACKING WAS DAMAGED(AS THE PHOTO SHOWS.) ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
13-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | nan
03-11-2025 | Malfunction | covidien | 26-01-2026 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, WHEN USING THE MANUAL STAPLER AND RELOAD, THE SNAP FELT WITHIN THE HANDLE. THE RELOAD CAUSED THE SNAP, AND THE BROKEN IMPRESSION OF THE HANDLE BUT THE SNAP PROBABLY OCCURRED DUE TO THE TISSUE THICKNESS. PLI20 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
03-11-2025 | Malfunction | covidien | 26-01-2026 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, WHEN USING THE MANUAL STAPLER AND RELOAD, THE SNAP FELT WITHIN THE HANDLE. THE RELOAD CAUSED THE SNAP, AND THE BROKEN IMPRESSION OF THE HANDLE BUT THE SNAP PROBABLY OCCURRED DUE TO THE TISSUE THICKNESS. PLI20 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
01-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST GREEN RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2026 | Malfunction | covidien | 26-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RADICAL GASTRECTOMY FOR GASTRIC CANCER, POOR STAPLE FORMATION WAS OBSERVED WITH TWO RELOADS DURING GASTROINTESTINAL ANASTOMOSIS, RESULTING IN AN IMPERFECT B-SHAPE. THE ISSUE WAS RESOLVED BY USING A NEW RELOAD, AND THE CASE WAS COMPLETED SUCCESSFULLY.
07-01-2026 | Malfunction | covidien | 26-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RADICAL GASTRECTOMY FOR GASTRIC CANCER, POOR STAPLE FORMATION WAS OBSERVED WITH TWO RELOADS DURING GASTROINTESTINAL ANASTOMOSIS, RESULTING IN AN IMPERFECT B-SHAPE. THE ISSUE WAS RESOLVED BY USING A NEW RELOAD, AND THE CASE WAS COMPLETED SUCCESSFULLY.
14-01-2026 | Malfunction | johnson & johnson medical (suzhou | 26-01-2026 | CARTRIDGE FINAL ASSY 45MM WHITE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-01-2026 | Malfunction | johnson & johnson medical (suzhou | 26-01-2026 | CARTRIDGE FINAL ASSY 45MM GREEN | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-01-2026 | Malfunction | johnson & johnson medical (suzhou | 26-01-2026 | CARTRIDGE FINAL ASSY 45MM BLUE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2025 | Malfunction | johnson & johnson medical (suzhou | 26-01-2026 | FINAL ASSY 60MM STD DEVICE | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
15-01-2026 | Malfunction | johnson & johnson medical (suzhou | 26-01-2026 | CARTRIDGE FINAL ASSY 60MM GREEN | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AND UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.N O ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2025 | Malfunction | ethicon endo-surgery | 26-01-2026 | GREEN RELOAD FOR ECHELON 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2025 | Malfunction | ethicon endo-surgery | 26-01-2026 | BLUE RELOAD FOR ECHELON 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AND UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THAT THE TISSUE WAS NOT CUT COMPLETELY. THEN CHANGED THE NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
20-01-2026 | Malfunction | covidien | 26-01-2026 | GIA AUTO SUTURE STAPLER DST SERIES 60 MM | Defective Device | A0203 | No Clinical Signs, Symptoms or Conditions | Event Description: GIA STAPLER 60MM AND GIA STAPLER 80MM WERE DEFECTIVE AND WOULD NOT FIRE THROUGH TISSUE. PATIENT CODE: 4582. DEVICE CODE: 2588, 2610. REFERENCE REPORT MW5182742.
20-01-2026 | Malfunction | covidien | 26-01-2026 | GIA AUTO SUTURE STAPLER DST SERIES 60 MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: GIA STAPLER 60MM AND GIA STAPLER 80MM WERE DEFECTIVE AND WOULD NOT FIRE THROUGH TISSUE. PATIENT CODE: 4582. DEVICE CODE: 2588, 2610. REFERENCE REPORT MW5182742.
20-01-2026 | Malfunction | covidien | 26-01-2026 | GIA AUTO SUTURE STAPLER DST SERIES 80MM | Defective Device | A0203 | No Clinical Signs, Symptoms or Conditions | Event Description: GIA STAPLER 60MM AND GIA STAPLER 80MM WERE DEFECTIVE AND WOULD NOT FIRE THROUGH TISSUE. PATIENT CODE: 4582. DEVICE CODE: 2588, 2610. REFERENCE REPORT: MW5182741.
20-01-2026 | Malfunction | covidien | 26-01-2026 | GIA AUTO SUTURE STAPLER DST SERIES 80MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: GIA STAPLER 60MM AND GIA STAPLER 80MM WERE DEFECTIVE AND WOULD NOT FIRE THROUGH TISSUE. PATIENT CODE: 4582. DEVICE CODE: 2588, 2610. REFERENCE REPORT: MW5182741.
01-12-2025 | Malfunction | ethicon endo-surgery | 26-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | ECHELON ENDOPATH STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
09-02-2026 | Injury | covidien surgical | 26-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: IT WAS REPORTED THAT POST-OPERATIVELY, THERE WAS COMPLETE DEHISCENCE OF THE STAPLED AREA WHICH NECESSITATED REOPERATION.
09-02-2026 | Injury | covidien surgical | 26-01-2026 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT POST-OPERATIVELY, THERE WAS COMPLETE DEHISCENCE OF THE STAPLED AREA WHICH NECESSITATED REOPERATION.
16-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST GREEN RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | GST BLUE RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
23-12-2025 | Malfunction | ethicon endo-surgery | 26-01-2026 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2025 | Malfunction | ethicon endo-surgery | 26-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2026 | Malfunction | ethicon endo-surgery | 26-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SIGMOID COLECTOMY SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
15-12-2025 | Malfunction | ethicon endo-surgery | 26-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-11-2025 | Malfunction | covidien surgical | 26-01-2026 | UNKNOWN TACKER | Adverse Event Without Identified Device or Use Problem | A24 | Foreign Body Reaction | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY ANALYZED THE RECURRENCE RATE, COMPLICATIONS, AND CLINICAL OUTCOMES BETWEEN THE USE OF BIOLOGICAL OR SYNTHETIC (PARIETEX COMPOSITE) MESH FOR LAPAROSCOPIC REPAIR OF HIATAL HERNIAS. THERE WERE 174 PATIENTS ENROLLED FROM APRIL 2022 TO FEBRUARY 2024. BOTH MESH TYPES WERE FIXED ON THE DIAPHRAGM WITH ABSORBABLE PROTACKS. POSTOPERATIVE COMPLICATIONS INCLUDED 1 CASE OF MESH-RELATED REACTION, 2 PATIENTS WITH HERNIA RECURRENCE, AND 5 PATIENTS THAT STILL REQUIRED MEDICATION FOR ACID REFLUX.
24-11-2025 | Malfunction | covidien surgical | 26-01-2026 | UNKNOWN TACKER | Adverse Event Without Identified Device or Use Problem | A24 | Hernia | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY ANALYZED THE RECURRENCE RATE, COMPLICATIONS, AND CLINICAL OUTCOMES BETWEEN THE USE OF BIOLOGICAL OR SYNTHETIC (PARIETEX COMPOSITE) MESH FOR LAPAROSCOPIC REPAIR OF HIATAL HERNIAS. THERE WERE 174 PATIENTS ENROLLED FROM APRIL 2022 TO FEBRUARY 2024. BOTH MESH TYPES WERE FIXED ON THE DIAPHRAGM WITH ABSORBABLE PROTACKS. POSTOPERATIVE COMPLICATIONS INCLUDED 1 CASE OF MESH-RELATED REACTION, 2 PATIENTS WITH HERNIA RECURRENCE, AND 5 PATIENTS THAT STILL REQUIRED MEDICATION FOR ACID REFLUX.
03-12-2025 | Injury | intuitive surgical | 26-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS (ROUX-EN Y) PROCEDURE, THE STAPLES AND TISSUE WERE PUSHED, RATHER THAN PROPERLY TRANSECTED, BY THE BLADE OF AN UNSPECIFIED STAPLER RELOAD WHILE USING A SUREFORM 45 STAPLER INSTRUMENT. THE ISSUE WAS RESOLVED BY USING A BACKUP STAPLER INSTRUMENT AND AN UNSPECIFIED STAPLER RELOAD TO RESECT THE DAMAGED SEGMENT OF THE STAPLE LINE AND CREATE A NEW STAPLE LINE. ALTHOUGH OPERATIVE TIME WAS EXTENDED, THE PROCEDURE WAS SUCCESSFULLY COMPLETED ROBOTICALLY WITHOUT FURTHER COMPLICATIONS. ADDITIONAL INFORMATION HAS BEEN REQUESTED; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO GATHER ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED.
29-12-2025 | Injury | intuitive surgical | 26-01-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, TISSUE WAS PUSHED AND NOT PROPERLY STAPLED OR CUT AFTER A SUREFORM 60 BLUE RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. AS A RESULT, A HOLE IN THE INTENDED STAPLE LINE WAS CREATED. THE SURGEON OVER-SUTURED THE HOLE WITHIN THE STAPLE LINE TO REPAIR THE DEFECT. ADDITIONALLY, AN UNPLANNED DRAIN WAS PLACED INTRAOPERATIVELY DUE TO THE EVENT. THE PATIENT REQUIRED AN UPPER GI EVALUATION ON POSTOPERATIVE DAY ONE.
02-01-2026 | Malfunction | covidien surgical | 26-01-2026 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR, THE TACKER DID NOT FIRE. THE DOCTOR USED ANOTHER COMPANY'S PRODUCT AND COMPLETED THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
02-01-2026 | Malfunction | covidien surgical | 26-01-2026 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR, THE TACKER DID NOT FIRE. THE DOCTOR USED ANOTHER COMPANY'S PRODUCT AND COMPLETED THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
03-11-2025 | Malfunction | covidien surgical | 26-01-2026 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DUODENOPANCREATECTOMY FOR ONCOLOGICAL REASONS, ASSOCIATED WITH PARTIAL GASTRECTOMY AND RETROPERI TONEAL LYMPHADENECTOMY, THE OPEN STAPLER MALFUNCTIONED, RESULTING IN INADEQUATE STAPLING OF STRUCTURES. THERE WAS NO PATIENT INJURY.
22-01-2026 | Malfunction | covidien surgical | 27-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-01-2026 | Malfunction | covidien | 27-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-01-2026 | Injury | intuitive surgical | 27-01-2026 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LIVER RESECTION, A SUREFORM 45 STAPLER WITH A WHITE RELOAD ENCOUNTERED CLAMPING AND FIRING DIFFICULTIES, WHICH RESULTED IN A TISSUE TEAR AT THE RIGHT HEPATIC VEIN AND SUBSEQUENT BLEEDING. THERE WERE NO OBSTRUCTIONS NOTED DURING STAPLING, AND THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. ADDITIONALLY, NO BUTTRESS MATERIAL WAS USED; THE TISSUE WAS NOT CALCIFIED, UNDER TENSION, BUNCHED, OR PREVIOUSLY EXPOSED TO CHEMOTHERAPY OR RADIATION. "TISSUE TOO THICK" ERROR MESSAGES WERE OBSERVED DURING THE ATTEMPT. THE ESTIMATED BLOOD LOSS WAS NOT REPORTED. HEMOSTASIS WAS ACHIEVED BY SUTURING THE BLEEDING VESSEL. THE STAPLER AND RELOAD WERE DISPOSED OF AND WILL NOT BE RETURNED FOR EVALUATION.
25-10-2025 | Malfunction | intuitive surgical | 27-01-2026 | SUREFORM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: DURING ROBOT ASSISTED NEPHRECTOMY, SURGEON ATTEMPTED TO STAPLE A VESSEL. IT WAS LOADED WITH A WHITE RELOAD AND POSITIONED TO STAPLE, WOULD NOT ENGAGE OR FIRE. REPOSITIONED AND ATTEMPTED A FEW TIMES WITHOUT SUCCESS. ANOTHER 45 STAPLER WAS OPENED, LOADED, AND WORKED TO STAPLE THE VESSEL. NO HARM TO PATIENT.
25-10-2025 | Malfunction | intuitive surgical | 27-01-2026 | SUREFORM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: DURING ROBOT ASSISTED NEPHRECTOMY, SURGEON ATTEMPTED TO STAPLE A VESSEL. IT WAS LOADED WITH A WHITE RELOAD AND POSITIONED TO STAPLE, WOULD NOT ENGAGE OR FIRE. REPOSITIONED AND ATTEMPTED A FEW TIMES WITHOUT SUCCESS. ANOTHER 45 STAPLER WAS OPENED, LOADED, AND WORKED TO STAPLE THE VESSEL. NO HARM TO PATIENT.
11-11-2025 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON FLEX | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: DURING A LAPAROSCOPIC PANCREATECTOMY/SPLENECTOMY, ONCE THE SPLENIC ARTERY WAS DISSECTED, A CUTTING STAPLER WAS FIRED.STAPLER MISFIRED BUT FORTUNATELY DID NOT CUT. STAPLER MANUALLY RESET AND SURGEON CHANGED TO A DIFFERENT BRAND OF STAPLER TO COMPLETE THE PROCEDURE.
11-11-2025 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON FLEX | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: DURING A LAPAROSCOPIC PANCREATECTOMY/SPLENECTOMY, ONCE THE SPLENIC ARTERY WAS DISSECTED, A CUTTING STAPLER WAS FIRED.STAPLER MISFIRED BUT FORTUNATELY DID NOT CUT. STAPLER MANUALLY RESET AND SURGEON CHANGED TO A DIFFERENT BRAND OF STAPLER TO COMPLETE THE PROCEDURE.
29-12-2025 | Malfunction | covidien surgical | 27-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED. THE RELOAD AND THE POWERED STAPLER SET WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-12-2025 | Malfunction | covidien surgical | 27-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED. THE RELOAD AND THE POWERED STAPLER SET WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-12-2025 | Malfunction | covidien surgical | 27-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED. THE RELOAD AND THE POWERED STAPLER SET WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2025 | Malfunction | covidien surgical | 27-01-2026 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, THE RELOAD INTELLIGENCE CHIP BROKE OFF. IT COULD NOT BE USED EVEN AFTER CONNECTING IT TO THE ADAPTER. TO RESOLVE THE ISSUE, A NEW RELOAD OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INVOLVED.
21-01-2026 | Malfunction | covidien | 27-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A SMALL INTESTINAL RESECTION, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2026 | Malfunction | covidien surgical | 27-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, FOLLOWING THE FIRST FIRING OF THE RELOAD ON THE STOMACH, A SEROSAL TEAR WAS IDENTIFIED ON THE UNDERSIDE OF THE STOMACH AFTER STAPLING, WHICH RESULTED IN OOZING FROM THE STAPLE LINE. THE PROCEDURE TIME WAS EXTENDED BY A FEW MINUTES TO ALLOW THE SURGEON TO OVERSEW THE AFFECTED AREA. STAPLE LINE REINFORCEMENT WAS PERFORMED BY SUTURING TO ACHIEVE HEMOSTASIS, AND FIBRIN SEALANT WAS APPLIED OVER THE ENTIRE STAPLE LINE FOR ADDITIONAL REINFORCEMENT AND HEMOSTASIS.
13-01-2026 | Malfunction | covidien surgical | 27-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, FOLLOWING THE FIRST FIRING OF THE RELOAD ON THE STOMACH, A SEROSAL TEAR WAS IDENTIFIED ON THE UNDERSIDE OF THE STOMACH AFTER STAPLING, WHICH RESULTED IN OOZING FROM THE STAPLE LINE. THE PROCEDURE TIME WAS EXTENDED BY A FEW MINUTES TO ALLOW THE SURGEON TO OVERSEW THE AFFECTED AREA. STAPLE LINE REINFORCEMENT WAS PERFORMED BY SUTURING TO ACHIEVE HEMOSTASIS, AND FIBRIN SEALANT WAS APPLIED OVER THE ENTIRE STAPLE LINE FOR ADDITIONAL REINFORCEMENT AND HEMOSTASIS.
07-10-2024 | Malfunction | covidien | 27-01-2026 | ENDO GIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT INTRAOPERATIVELY, THE RELOAD WAS DIFFICULT TO FIRE WHEN USED WITH THE POWERED STAPLER. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD WAS RECEIVED PARTIALLY FIRED. THE JAWS OF THE RELOAD WERE CLAMPED. THE I BEAM WAS PARTIALLY ADVANCED. FULLY FORMED STAPLES WERE PRESENT IN THE PROXIMAL END OF THE JAWS.
07-10-2024 | Malfunction | covidien | 27-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT INTRAOPERATIVELY, THE RELOAD WAS DIFFICULT TO FIRE WHEN USED WITH THE POWERED STAPLER. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD WAS RECEIVED PARTIALLY FIRED. THE JAWS OF THE RELOAD WERE CLAMPED. THE I BEAM WAS PARTIALLY ADVANCED. FULLY FORMED STAPLES WERE PRESENT IN THE PROXIMAL END OF THE JAWS.
07-10-2024 | Malfunction | covidien | 27-01-2026 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT INTRAOPERATIVELY, THE RELOAD WAS DIFFICULT TO FIRE WHEN USED WITH THE POWERED STAPLER. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD WAS RECEIVED PARTIALLY FIRED. THE JAWS OF THE RELOAD WERE CLAMPED. THE I BEAM WAS PARTIALLY ADVANCED. FULLY FORMED STAPLES WERE PRESENT IN THE PROXIMAL END OF THE JAWS.
26-12-2025 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLE LINE WAS DEFORMED AS THE KNIFE STOPPED IN THE MIDDLE OF THE SEQUENCE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2026 | Malfunction | ethicon endo-surgery | 27-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
05-01-2026 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LIVER TRANSPLANT PROCEDURE, IT WAS OBSERVED THAT THE JAWS FAILED TO RELEASE AND REMAINED CLOSED AFTER FIRING. THE RELEASE BUTTON DID NOT FUNCTION AS INTENDED, AND THE SURGEON EXPERIENCED SIGNIFICANT DIFFICULTY REMOVING THE DEVICE FROM THE TISSUE. MANUAL ATTEMPTS TO DISENGAGE THE JAWS WERE UNSUCCESSFUL INITIALLY, WHICH CAUSED AN INTRAOPERATIVE DELAY AND INCREASED RISK DURING A CRITICAL STAGE OF THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
05-01-2026 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON FLEX 45 STANDARD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LIVER TRANSPLANT PROCEDURE, IT WAS OBSERVED THAT THE JAWS FAILED TO RELEASE AND REMAINED CLOSED AFTER FIRING. THE RELEASE BUTTON DID NOT FUNCTION AS INTENDED, AND THE SURGEON EXPERIENCED SIGNIFICANT DIFFICULTY REMOVING THE DEVICE FROM THE TISSUE. MANUAL ATTEMPTS TO DISENGAGE THE JAWS WERE UNSUCCESSFUL INITIALLY, WHICH CAUSED AN INTRAOPERATIVE DELAY AND INCREASED RISK DURING A CRITICAL STAGE OF THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
26-12-2025 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON 3000 45 STANDARD | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, THE STAPLER SUCCESSFULLY FIRED ONE WHITE LOAD ON THE PEDICLE. A GREEN LOAD WAS THEN ATTEMPTED ACROSS THE SIGMOID; HOWEVER, HALFWAY THROUGH FIRING, THE DEVICE SHUT DOWN AND DID NOT COMPLETE THE FIRING. WHEN PRESSING THE HOME BUTTON, THE JAW DID NOT RETRACT. THE SURGEON IMMEDIATELY USED THE MANUAL OVERRIDE AND WAS ABLE TO OPEN THE JAW SUCCESSFULLY. AN ADDITIONAL STAPLER HAD TO BE OPENED TO COMPLETE THE FIRING. THERE WERE NO PATIENT ADVERSE EVENT. NO ADDITIONAL INFORMATION PROVIDED.
26-12-2025 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, THE STAPLER SUCCESSFULLY FIRED ONE WHITE LOAD ON THE PEDICLE. A GREEN LOAD WAS THEN ATTEMPTED ACROSS THE SIGMOID; HOWEVER, HALFWAY THROUGH FIRING, THE DEVICE SHUT DOWN AND DID NOT COMPLETE THE FIRING. WHEN PRESSING THE HOME BUTTON, THE JAW DID NOT RETRACT. THE SURGEON IMMEDIATELY USED THE MANUAL OVERRIDE AND WAS ABLE TO OPEN THE JAW SUCCESSFULLY. AN ADDITIONAL STAPLER HAD TO BE OPENED TO COMPLETE THE FIRING. THERE WERE NO PATIENT ADVERSE EVENT. NO ADDITIONAL INFORMATION PROVIDED.
26-12-2025 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, THE STAPLER SUCCESSFULLY FIRED ONE WHITE LOAD ON THE PEDICLE. A GREEN LOAD WAS THEN ATTEMPTED ACROSS THE SIGMOID; HOWEVER, HALFWAY THROUGH FIRING, THE DEVICE SHUT DOWN AND DID NOT COMPLETE THE FIRING. WHEN PRESSING THE HOME BUTTON, THE JAW DID NOT RETRACT. THE SURGEON IMMEDIATELY USED THE MANUAL OVERRIDE AND WAS ABLE TO OPEN THE JAW SUCCESSFULLY. AN ADDITIONAL STAPLER HAD TO BE OPENED TO COMPLETE THE FIRING. THERE WERE NO PATIENT ADVERSE EVENT. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 27-01-2026 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
05-12-2025 | Malfunction | covidien | 27-01-2026 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LUNG CANCER RADICAL SURGERY, THE SURGEON HAD DIFFICULTY CLOSING THE RELOAD. AFTER CLOSURE, THE DEVICE COULD NOT BE FIRED OUT PROPERLY. A NEW CUTTING HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED BUT NOT MORE THAN THIRTY MINUTES.
05-12-2025 | Malfunction | covidien | 27-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LUNG CANCER RADICAL SURGERY, THE SURGEON HAD DIFFICULTY CLOSING THE RELOAD. AFTER CLOSURE, THE DEVICE COULD NOT BE FIRED OUT PROPERLY. A NEW CUTTING HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED BUT NOT MORE THAN THIRTY MINUTES.
05-12-2025 | Malfunction | covidien | 27-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LUNG CANCER RADICAL SURGERY, THE SURGEON HAD DIFFICULTY CLOSING THE RELOAD. AFTER CLOSURE, THE DEVICE COULD NOT BE FIRED OUT PROPERLY. A NEW CUTTING HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED BUT NOT MORE THAN THIRTY MINUTES.
05-12-2025 | Malfunction | covidien | 27-01-2026 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LUNG CANCER RADICAL SURGERY, THE SURGEON HAD DIFFICULTY CLOSING THE RELOAD. AFTER CLOSURE, THE DEVICE COULD NOT BE FIRED OUT PROPERLY. A NEW CUTTING HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED BUT NOT MORE THAN THIRTY MINUTES.
26-12-2025 | Malfunction | covidien surgical | 27-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE POWERED STAPLER DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-01-2026 | Malfunction | covidien | 27-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RADICAL COLON CANCER SURGERY, WHEN ON PREPARATION FOR DISSECTING THE COLON, THE RELOAD DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien | 27-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: DUPLICATE OF (B)(4).
05-01-2026 | Malfunction | covidien | 27-01-2026 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: DUPLICATE OF (B)(4).
05-01-2026 | Malfunction | covidien surgical | 27-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | nan
05-01-2026 | Malfunction | covidien surgical | 27-01-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | nan
05-01-2026 | Malfunction | covidien | 27-01-2026 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, THE RETURN BUTTON RETURNED TO ITS ORIGINAL POSITION, THE RELOAD DID NOT OPEN. AFTER REPEATEDLY TRYING TO HOLD THE OPENING AND CLOSING HANDLE, THE RETURN BUTTON RETRACTED AGAIN, AND AFTER REPEATING THE STAPLE WITHDRAWAL OPERATION, THE STAPLE SPREAD APART. THE HANDLE AND STAPLE COULD NOT BE UNLOADED. BOTH THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien | 27-01-2026 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, THE RETURN BUTTON RETURNED TO ITS ORIGINAL POSITION, THE RELOAD DID NOT OPEN. AFTER REPEATEDLY TRYING TO HOLD THE OPENING AND CLOSING HANDLE, THE RETURN BUTTON RETRACTED AGAIN, AND AFTER REPEATING THE STAPLE WITHDRAWAL OPERATION, THE STAPLE SPREAD APART. THE HANDLE AND STAPLE COULD NOT BE UNLOADED. BOTH THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien | 27-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, THE RETURN BUTTON RETURNED TO ITS ORIGINAL POSITION, THE RELOAD DID NOT OPEN. AFTER REPEATEDLY TRYING TO HOLD THE OPENING AND CLOSING HANDLE, THE RETURN BUTTON RETRACTED AGAIN, AND AFTER REPEATING THE STAPLE WITHDRAWAL OPERATION, THE STAPLE SPREAD APART. THE HANDLE AND STAPLE COULD NOT BE UNLOADED. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien | 27-01-2026 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, THE RETURN BUTTON RETURNED TO ITS ORIGINAL POSITION, THE RELOAD DID NOT OPEN. AFTER REPEATEDLY TRYING TO HOLD THE OPENING AND CLOSING HANDLE, THE RETURN BUTTON RETRACTED AGAIN, AND AFTER REPEATING THE STAPLE WITHDRAWAL OPERATION, THE STAPLE SPREAD APART. THE HANDLE AND STAPLE COULD NOT BE UNLOADED. THERE WAS NO PATIENT INJURY.
08-01-2026 | Malfunction | covidien surgical | 28-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL LUNG RESECTION, AN ISSUE OCCURRED WHILE USING THE SECOND CARTRIDGE FOR THE CENTRAL LUNG PARTIAL RESECTION. THE FIRING STOPPED AFTER ADVANCING APPROXIMATELY 3 CM. SUBSEQUENTLY, TWO ADDITIONAL CARTRIDGES WERE USED, AND THE SPECIMEN WAS SUCCESSFULLY OBTAINED. THERE WAS NO PATIENT INJURY.
08-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL LUNG RESECTION, AN ISSUE OCCURRED WHILE USING THE SECOND CARTRIDGE FOR THE CENTRAL LUNG PARTIAL RESECTION. THE FIRING STOPPED AFTER ADVANCING APPROXIMATELY 3 CM. SUBSEQUENTLY, TWO ADDITIONAL CARTRIDGES WERE USED, AND THE SPECIMEN WAS SUCCESSFULLY OBTAINED. THERE WAS NO PATIENT INJURY.
08-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL LUNG RESECTION, AN ISSUE OCCURRED WHILE USING THE SECOND CARTRIDGE FOR THE CENTRAL LUNG PARTIAL RESECTION. THE FIRING STOPPED AFTER ADVANCING APPROXIMATELY 3 CM. SUBSEQUENTLY, TWO ADDITIONAL CARTRIDGES WERE USED, AND THE SPECIMEN WAS SUCCESSFULLY OBTAINED. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
05-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
05-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | ECHELON 3000 45 STANDARD | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | nan
12-01-2026 | Malfunction | johnson & johnson medical (suzhou | 28-01-2026 | CARTRIDGE FINAL ASSY 45MM GOLD | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES WERE FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-01-2026 | Malfunction | johnson & johnson medical (suzhou | 28-01-2026 | CARTRIDGE FINAL ASSY 45MM GOLD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES WERE FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-01-2026 | Malfunction | johnson & johnson medical (suzhou | 28-01-2026 | CARTRIDGE FINAL ASSY 45MM GOLD | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES WERE FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-01-2026 | Malfunction | johnson & johnson medical (suzhou | 28-01-2026 | CARTRIDGE FINAL ASSY 45MM GOLD | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT A FEW STAPLES WERE FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. CHANGED TO ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE BYPASS PROCEDURE, WHILE FIRING ON THE STOMACH, IT WAS OBSERVED THAT THE STAPLER CUT BUT DID NOT STAPLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
24-11-2025 | Malfunction | ethicon endo-surgery | 28-01-2026 | ECHELON FLEX 45 STANDARD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER DID NOT TRIGGER THE REFILLS TWICE, THE MACHINE CRASHED AND LEFT THE STAPLES HALF-OUT. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | GST WHITE RELOAD, 60MM, 6-ROW | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT ONCE THE DEVICE WAS CLOSED, IT COULD NOT BE OPENED AND COULD NOT BE USED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | GST WHITE RELOAD, 60MM, 6-ROW | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT ONCE THE DEVICE WAS CLOSED, IT COULD NOT BE OPENED AND COULD NOT BE USED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | GST WHITE RELOAD, 60MM, 6-ROW | No Apparent Adverse Event | A25 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT ONCE THE DEVICE WAS CLOSED, IT COULD NOT BE OPENED AND COULD NOT BE USED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | GST WHITE RELOAD, 60MM, 6-ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT ONCE THE DEVICE WAS CLOSED, IT COULD NOT BE OPENED AND COULD NOT BE USED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | ETHICON 3D RELOAD 60MM BLUE | Material Integrity Problem | A04 | Foreign Body In Patient | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE, THEY NOTICED SOME BLUE FRAGMENTS ON THE SPECIMEN SIDE ON THE SECOND TO LAST FIRING. THEY LOOKED LIKE BLUE RELOAD PIECES ON THAT SIDE. THEY WERE USING THE ETHICON 4000 STAPLER AND IT WAS A BLUE RELOAD. THERE WERE NO BLUE FRAGMENTS FOUND ON THE PATIENT SIDE AND THE STAPLE LINE LOOKED GOOD. THE RELOAD LOOKED FRAYED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 28-01-2026 | ETHICON 3D RELOAD 60MM BLUE | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE, THEY NOTICED SOME BLUE FRAGMENTS ON THE SPECIMEN SIDE ON THE SECOND TO LAST FIRING. THEY LOOKED LIKE BLUE RELOAD PIECES ON THAT SIDE. THEY WERE USING THE ETHICON 4000 STAPLER AND IT WAS A BLUE RELOAD. THERE WERE NO BLUE FRAGMENTS FOUND ON THE PATIENT SIDE AND THE STAPLE LINE LOOKED GOOD. THE RELOAD LOOKED FRAYED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
27-01-2026 | Malfunction | covidien | 28-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORACOSCOPIC LEFT LOWER LOBECTOMY, WHILE PREPARING TO DETACH THE LOBAR BRONCHUS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE ORIGINAL RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-10-2025 | Injury | covidien | 28-01-2026 | UNKNOWN GIA PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN OCTOBER 2021 TO DECEMBER 2024, A PROSPECTIVE OBSERVATIONAL STUDY COMPARED THE SHORT- AND MID-TERM SURGICAL OUTCOMES OF BILLROTH I AND BILLROTH II/ROUX-EN-Y RECONSTRUCTION FOLLOWING DISTAL GASTRECTOMY IN GASTRIC CANCER PATIENTS. A TOTAL OF 150 PATIENTS WERE ENROLLED IN THE STUDY, WITH 72 UNDERGOING BILLROTH I AND 78 UNDERGOING BILLROTH II/ROUX-EN-Y RECONSTRUCTION. STAPLED END-TO-SIDE ANASTOMOSES WITH (GIA 60 MM LINEAR STAPLER) WERE USED SELECTIVELY WHEN DUODENAL MOBILITY ALLOWED A TENSION-FREE ALIGNMENT. COMPLICATIONS INCLUDED 5 PATIENTS WITH ANASTOMOTIC LEAK, 2 PATIENTS WITH DUODENAL STUMP LEAK, 7 WITH INTRA-ABDOMINAL ABSCESS, 4 WITH POSTOPERATIVE BLEEDING, 11 WITH WOUND INFECTIONS, AND 3 PATIENTS WITH POSTOPERATIVE STRICTURE [OCCLUSION]. HEALTH IMPACTS INCLUDED 17 PATIENT HOSPITAL READMISSIONS, 8 PATIENT REOPERATIONS, AND 11 PATIENTS REQUIRING INTENSIVE CARE UNIT LEVEL CARE. NON PRODUCT RELATED, PROCEDURAL COMPLICATIONS INCLUDED POSTOPERATIVE PULMONARY COMPLICATIONS, CARDIOVASCULAR COMPLICATIONS, PROTON PUMP INHIBITOR MEDICATION USE, ESOPHAGITIS, BILE REFLUX GASTRITIS, AND DUMPING SYNDROME. THE STUDY ALSO EVALUATED 90-DAY PATIENT MORTALITY, NOTING 13 PATIENT DEATHS; HOWEVER, NO PRODUCT ISSUE WAS NOTED, NOR ANY INDICATION THAT THE DEVICE WAS DIRECTLY RELATED TO THE REPORTED DEATHS.
27-10-2025 | Injury | covidien | 28-01-2026 | UNKNOWN GIA PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN OCTOBER 2021 TO DECEMBER 2024, A PROSPECTIVE OBSERVATIONAL STUDY COMPARED THE SHORT- AND MID-TERM SURGICAL OUTCOMES OF BILLROTH I AND BILLROTH II/ROUX-EN-Y RECONSTRUCTION FOLLOWING DISTAL GASTRECTOMY IN GASTRIC CANCER PATIENTS. A TOTAL OF 150 PATIENTS WERE ENROLLED IN THE STUDY, WITH 72 UNDERGOING BILLROTH I AND 78 UNDERGOING BILLROTH II/ROUX-EN-Y RECONSTRUCTION. STAPLED END-TO-SIDE ANASTOMOSES WITH (GIA 60 MM LINEAR STAPLER) WERE USED SELECTIVELY WHEN DUODENAL MOBILITY ALLOWED A TENSION-FREE ALIGNMENT. COMPLICATIONS INCLUDED 5 PATIENTS WITH ANASTOMOTIC LEAK, 2 PATIENTS WITH DUODENAL STUMP LEAK, 7 WITH INTRA-ABDOMINAL ABSCESS, 4 WITH POSTOPERATIVE BLEEDING, 11 WITH WOUND INFECTIONS, AND 3 PATIENTS WITH POSTOPERATIVE STRICTURE [OCCLUSION]. HEALTH IMPACTS INCLUDED 17 PATIENT HOSPITAL READMISSIONS, 8 PATIENT REOPERATIONS, AND 11 PATIENTS REQUIRING INTENSIVE CARE UNIT LEVEL CARE. NON PRODUCT RELATED, PROCEDURAL COMPLICATIONS INCLUDED POSTOPERATIVE PULMONARY COMPLICATIONS, CARDIOVASCULAR COMPLICATIONS, PROTON PUMP INHIBITOR MEDICATION USE, ESOPHAGITIS, BILE REFLUX GASTRITIS, AND DUMPING SYNDROME. THE STUDY ALSO EVALUATED 90-DAY PATIENT MORTALITY, NOTING 13 PATIENT DEATHS; HOWEVER, NO PRODUCT ISSUE WAS NOTED, NOR ANY INDICATION THAT THE DEVICE WAS DIRECTLY RELATED TO THE REPORTED DEATHS.
27-10-2025 | Injury | covidien | 28-01-2026 | UNKNOWN GIA PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN OCTOBER 2021 TO DECEMBER 2024, A PROSPECTIVE OBSERVATIONAL STUDY COMPARED THE SHORT- AND MID-TERM SURGICAL OUTCOMES OF BILLROTH I AND BILLROTH II/ROUX-EN-Y RECONSTRUCTION FOLLOWING DISTAL GASTRECTOMY IN GASTRIC CANCER PATIENTS. A TOTAL OF 150 PATIENTS WERE ENROLLED IN THE STUDY, WITH 72 UNDERGOING BILLROTH I AND 78 UNDERGOING BILLROTH II/ROUX-EN-Y RECONSTRUCTION. STAPLED END-TO-SIDE ANASTOMOSES WITH (GIA 60 MM LINEAR STAPLER) WERE USED SELECTIVELY WHEN DUODENAL MOBILITY ALLOWED A TENSION-FREE ALIGNMENT. COMPLICATIONS INCLUDED 5 PATIENTS WITH ANASTOMOTIC LEAK, 2 PATIENTS WITH DUODENAL STUMP LEAK, 7 WITH INTRA-ABDOMINAL ABSCESS, 4 WITH POSTOPERATIVE BLEEDING, 11 WITH WOUND INFECTIONS, AND 3 PATIENTS WITH POSTOPERATIVE STRICTURE [OCCLUSION]. HEALTH IMPACTS INCLUDED 17 PATIENT HOSPITAL READMISSIONS, 8 PATIENT REOPERATIONS, AND 11 PATIENTS REQUIRING INTENSIVE CARE UNIT LEVEL CARE. NON PRODUCT RELATED, PROCEDURAL COMPLICATIONS INCLUDED POSTOPERATIVE PULMONARY COMPLICATIONS, CARDIOVASCULAR COMPLICATIONS, PROTON PUMP INHIBITOR MEDICATION USE, ESOPHAGITIS, BILE REFLUX GASTRITIS, AND DUMPING SYNDROME. THE STUDY ALSO EVALUATED 90-DAY PATIENT MORTALITY, NOTING 13 PATIENT DEATHS; HOWEVER, NO PRODUCT ISSUE WAS NOTED, NOR ANY INDICATION THAT THE DEVICE WAS DIRECTLY RELATED TO THE REPORTED DEATHS.
27-10-2025 | Injury | covidien | 28-01-2026 | UNKNOWN GIA PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Obstruction/Occlusion | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN OCTOBER 2021 TO DECEMBER 2024, A PROSPECTIVE OBSERVATIONAL STUDY COMPARED THE SHORT- AND MID-TERM SURGICAL OUTCOMES OF BILLROTH I AND BILLROTH II/ROUX-EN-Y RECONSTRUCTION FOLLOWING DISTAL GASTRECTOMY IN GASTRIC CANCER PATIENTS. A TOTAL OF 150 PATIENTS WERE ENROLLED IN THE STUDY, WITH 72 UNDERGOING BILLROTH I AND 78 UNDERGOING BILLROTH II/ROUX-EN-Y RECONSTRUCTION. STAPLED END-TO-SIDE ANASTOMOSES WITH (GIA 60 MM LINEAR STAPLER) WERE USED SELECTIVELY WHEN DUODENAL MOBILITY ALLOWED A TENSION-FREE ALIGNMENT. COMPLICATIONS INCLUDED 5 PATIENTS WITH ANASTOMOTIC LEAK, 2 PATIENTS WITH DUODENAL STUMP LEAK, 7 WITH INTRA-ABDOMINAL ABSCESS, 4 WITH POSTOPERATIVE BLEEDING, 11 WITH WOUND INFECTIONS, AND 3 PATIENTS WITH POSTOPERATIVE STRICTURE [OCCLUSION]. HEALTH IMPACTS INCLUDED 17 PATIENT HOSPITAL READMISSIONS, 8 PATIENT REOPERATIONS, AND 11 PATIENTS REQUIRING INTENSIVE CARE UNIT LEVEL CARE. NON PRODUCT RELATED, PROCEDURAL COMPLICATIONS INCLUDED POSTOPERATIVE PULMONARY COMPLICATIONS, CARDIOVASCULAR COMPLICATIONS, PROTON PUMP INHIBITOR MEDICATION USE, ESOPHAGITIS, BILE REFLUX GASTRITIS, AND DUMPING SYNDROME. THE STUDY ALSO EVALUATED 90-DAY PATIENT MORTALITY, NOTING 13 PATIENT DEATHS; HOWEVER, NO PRODUCT ISSUE WAS NOTED, NOR ANY INDICATION THAT THE DEVICE WAS DIRECTLY RELATED TO THE REPORTED DEATHS.
22-06-2022 | Injury | davol inc sub c.r bard inc -1213643 | 28-01-2026 | OPTIFIX STRAIGHT | Patient-Device Incompatibility | A0101 | Fatigue | Event Description: PER ANSM USER REPORT ((B)(4): IT WAS REPORTED THAT, "SINCE THE INSERTION OF HERNIA REPAIR MESH FOR AN UMBILICAL HERNIA, MY LIFE HAS BEEN HELL, WITH MY HEALTH DETERIORATING, AS CONFIRMED BY A MEDICAL EXAMINATION! MY LIFE IS IN DANGER AND I AM IN INTENSE PAIN, REQUIRING MORPHINE. SINCE THE OPERATION, I HAVE REQUESTED THAT THE MESH BE REMOVED URGENTLY, EVEN IF I HAVE TO GO TO GERMANY TO HAVE IT DONE SINCE THE IMPLANT WAS FITTED, I HAVE BEEN IN AGONY SINCE WAKING UP AFTER THE OPERATION. THE PAIN HAS BEEN DOWNPLAYED BY THE NURSING STAFF, WHO TELL ME I AM EXAGGERATING AND THAT IT IS ALL IN MY HEAD. THE OPERATION SHOULD HAVE BEEN AN OUTPATIENT PROCEDURE, BUT I HAVE BEEN HOSPITALIZED EVER SINCE AND MY LIFE IS NO LONGER THE SAME. COMMENTS: IT IS CRIMINAL TO IMPLANT THIS TYPE OF PROSTHESIS. I AM A CARER AND I WAS NOT SUSPICIOUS ENOUGH. I THOUGHT I COULD TRUST THE SURGEON, BUT THEY STOLE MY LIFE. MY FIVE CHILDREN HAVE LOST THEIR MOTHER, WHO OFTEN HAS TO LIE DOWN BECAUSE OF THE INTOLERABLE PAIN. I AM UNABLE TO SEND YOU THE NECESSARY DOCUMENTS. ERROR DETECTED. PATIENT INFORMATION: CURRENT PATIENT STATUS: SINCE THIS MATERIAL WAS IMPLANTED, I HAVE BEEN SURVIVING. MY GENERAL CONDITION IS DETERIORATING RAPIDLY. I HAVE LOST 25 KG AND AM IN CONSTANT PAIN (EVA 8 ON THE RIGHT SIDE AT THE LEVEL OF THE ABDOMINAL, PELVIC AND INGUINAL PROSTHESIS), PLUS I AM LOSING MOBILITY IN MY RIGHT LEG. I AM EXPERIENCING SEVERE HAIR LOSS AND HAVE LOST MY APPETITE. I AM UNABLE TO EAT AND FEEL OVERWHELMING FATIGUE. I HAVE ALREADY CONTEMPLATED SUICIDE BECAUSE THE PAIN IS SO UNBEARABLE. I AM CURRENTLY SEEING PROFESSOR R IN HAUTEPIERRE TO REQUEST REMOVAL OF THE DEVICE. ACTIONS TAKEN BY THE HEALTHCARE FACILITY TO TREAT THE PATIENT: NR. THE PATIENT WAS IMPLANTED WITH NON-BD MESH (MATERIAL PÉTERS SURGICAL, COUSIN BIOTECH, PROMESH SURG INTRA) FOR UMBILICAL HERNIA REPAIR. THE MESH WAS FIXATED USING OPTIFIX FIXATION DEVICE. THIS REPORT REPRESENTS THE BD OPTIFIX FIXATION DEVICE.
22-06-2022 | Injury | davol inc sub c.r bard inc -1213643 | 28-01-2026 | OPTIFIX STRAIGHT | Patient-Device Incompatibility | A0101 | Pain | Event Description: PER ANSM USER REPORT ((B)(4): IT WAS REPORTED THAT, "SINCE THE INSERTION OF HERNIA REPAIR MESH FOR AN UMBILICAL HERNIA, MY LIFE HAS BEEN HELL, WITH MY HEALTH DETERIORATING, AS CONFIRMED BY A MEDICAL EXAMINATION! MY LIFE IS IN DANGER AND I AM IN INTENSE PAIN, REQUIRING MORPHINE. SINCE THE OPERATION, I HAVE REQUESTED THAT THE MESH BE REMOVED URGENTLY, EVEN IF I HAVE TO GO TO GERMANY TO HAVE IT DONE SINCE THE IMPLANT WAS FITTED, I HAVE BEEN IN AGONY SINCE WAKING UP AFTER THE OPERATION. THE PAIN HAS BEEN DOWNPLAYED BY THE NURSING STAFF, WHO TELL ME I AM EXAGGERATING AND THAT IT IS ALL IN MY HEAD. THE OPERATION SHOULD HAVE BEEN AN OUTPATIENT PROCEDURE, BUT I HAVE BEEN HOSPITALIZED EVER SINCE AND MY LIFE IS NO LONGER THE SAME. COMMENTS: IT IS CRIMINAL TO IMPLANT THIS TYPE OF PROSTHESIS. I AM A CARER AND I WAS NOT SUSPICIOUS ENOUGH. I THOUGHT I COULD TRUST THE SURGEON, BUT THEY STOLE MY LIFE. MY FIVE CHILDREN HAVE LOST THEIR MOTHER, WHO OFTEN HAS TO LIE DOWN BECAUSE OF THE INTOLERABLE PAIN. I AM UNABLE TO SEND YOU THE NECESSARY DOCUMENTS. ERROR DETECTED. PATIENT INFORMATION: CURRENT PATIENT STATUS: SINCE THIS MATERIAL WAS IMPLANTED, I HAVE BEEN SURVIVING. MY GENERAL CONDITION IS DETERIORATING RAPIDLY. I HAVE LOST 25 KG AND AM IN CONSTANT PAIN (EVA 8 ON THE RIGHT SIDE AT THE LEVEL OF THE ABDOMINAL, PELVIC AND INGUINAL PROSTHESIS), PLUS I AM LOSING MOBILITY IN MY RIGHT LEG. I AM EXPERIENCING SEVERE HAIR LOSS AND HAVE LOST MY APPETITE. I AM UNABLE TO EAT AND FEEL OVERWHELMING FATIGUE. I HAVE ALREADY CONTEMPLATED SUICIDE BECAUSE THE PAIN IS SO UNBEARABLE. I AM CURRENTLY SEEING PROFESSOR R IN HAUTEPIERRE TO REQUEST REMOVAL OF THE DEVICE. ACTIONS TAKEN BY THE HEALTHCARE FACILITY TO TREAT THE PATIENT: NR. THE PATIENT WAS IMPLANTED WITH NON-BD MESH (MATERIAL PÉTERS SURGICAL, COUSIN BIOTECH, PROMESH SURG INTRA) FOR UMBILICAL HERNIA REPAIR. THE MESH WAS FIXATED USING OPTIFIX FIXATION DEVICE. THIS REPORT REPRESENTS THE BD OPTIFIX FIXATION DEVICE.
22-06-2022 | Injury | davol inc sub c.r bard inc -1213643 | 28-01-2026 | OPTIFIX STRAIGHT | Patient-Device Incompatibility | A0101 | Depression | Event Description: PER ANSM USER REPORT ((B)(4): IT WAS REPORTED THAT, "SINCE THE INSERTION OF HERNIA REPAIR MESH FOR AN UMBILICAL HERNIA, MY LIFE HAS BEEN HELL, WITH MY HEALTH DETERIORATING, AS CONFIRMED BY A MEDICAL EXAMINATION! MY LIFE IS IN DANGER AND I AM IN INTENSE PAIN, REQUIRING MORPHINE. SINCE THE OPERATION, I HAVE REQUESTED THAT THE MESH BE REMOVED URGENTLY, EVEN IF I HAVE TO GO TO GERMANY TO HAVE IT DONE SINCE THE IMPLANT WAS FITTED, I HAVE BEEN IN AGONY SINCE WAKING UP AFTER THE OPERATION. THE PAIN HAS BEEN DOWNPLAYED BY THE NURSING STAFF, WHO TELL ME I AM EXAGGERATING AND THAT IT IS ALL IN MY HEAD. THE OPERATION SHOULD HAVE BEEN AN OUTPATIENT PROCEDURE, BUT I HAVE BEEN HOSPITALIZED EVER SINCE AND MY LIFE IS NO LONGER THE SAME. COMMENTS: IT IS CRIMINAL TO IMPLANT THIS TYPE OF PROSTHESIS. I AM A CARER AND I WAS NOT SUSPICIOUS ENOUGH. I THOUGHT I COULD TRUST THE SURGEON, BUT THEY STOLE MY LIFE. MY FIVE CHILDREN HAVE LOST THEIR MOTHER, WHO OFTEN HAS TO LIE DOWN BECAUSE OF THE INTOLERABLE PAIN. I AM UNABLE TO SEND YOU THE NECESSARY DOCUMENTS. ERROR DETECTED. PATIENT INFORMATION: CURRENT PATIENT STATUS: SINCE THIS MATERIAL WAS IMPLANTED, I HAVE BEEN SURVIVING. MY GENERAL CONDITION IS DETERIORATING RAPIDLY. I HAVE LOST 25 KG AND AM IN CONSTANT PAIN (EVA 8 ON THE RIGHT SIDE AT THE LEVEL OF THE ABDOMINAL, PELVIC AND INGUINAL PROSTHESIS), PLUS I AM LOSING MOBILITY IN MY RIGHT LEG. I AM EXPERIENCING SEVERE HAIR LOSS AND HAVE LOST MY APPETITE. I AM UNABLE TO EAT AND FEEL OVERWHELMING FATIGUE. I HAVE ALREADY CONTEMPLATED SUICIDE BECAUSE THE PAIN IS SO UNBEARABLE. I AM CURRENTLY SEEING PROFESSOR R IN HAUTEPIERRE TO REQUEST REMOVAL OF THE DEVICE. ACTIONS TAKEN BY THE HEALTHCARE FACILITY TO TREAT THE PATIENT: NR. THE PATIENT WAS IMPLANTED WITH NON-BD MESH (MATERIAL PÉTERS SURGICAL, COUSIN BIOTECH, PROMESH SURG INTRA) FOR UMBILICAL HERNIA REPAIR. THE MESH WAS FIXATED USING OPTIFIX FIXATION DEVICE. THIS REPORT REPRESENTS THE BD OPTIFIX FIXATION DEVICE.
05-01-2026 | Malfunction | davol inc sub c.r bard inc -1213643 | 28-01-2026 | OPTIFIX STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
03-11-2025 | Malfunction | covidien | 28-01-2026 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | nan
07-01-2026 | Malfunction | covidien surgical | 28-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A PROCEDURE, PRIOR TO FIRING, THE JAWS NOT CLOSING COMPLETELY WHEN USED WITH A POWERED STAPLER AND STANDARD ADAPTER. THE ISSUE WAS RESOLVED BY REPLACING THE RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
07-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A PROCEDURE, PRIOR TO FIRING, THE JAWS NOT CLOSING COMPLETELY WHEN USED WITH A POWERED STAPLER AND STANDARD ADAPTER. THE ISSUE WAS RESOLVED BY REPLACING THE RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
07-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A PROCEDURE, PRIOR TO FIRING, THE JAWS NOT CLOSING COMPLETELY WHEN USED WITH A POWERED STAPLER AND STANDARD ADAPTER. THE ISSUE WAS RESOLVED BY REPLACING THE RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
12-01-2026 | Malfunction | covidien surgical | 28-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC CHOLEDOCHOENTEROSTOMY PROCEDURE, WHEN THE DEVICE WAS APPLIED DURING THE ANASTOMOSIS STEP, THE RELOAD WAS ASSOCIATED WITH POOR STAPLE FORMATION, INCOMPLETE DEPLOYMENT OF STAPLES, AND STAPLES THAT COULD NOT BE COMPLETELY CUT. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW STAPLE.
12-01-2026 | Malfunction | covidien surgical | 28-01-2026 | ENDO GIA | Failure to Cut | A050702 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC CHOLEDOCHOENTEROSTOMY PROCEDURE, WHEN THE DEVICE WAS APPLIED DURING THE ANASTOMOSIS STEP, THE RELOAD WAS ASSOCIATED WITH POOR STAPLE FORMATION, INCOMPLETE DEPLOYMENT OF STAPLES, AND STAPLES THAT COULD NOT BE COMPLETELY CUT. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW STAPLE.
12-01-2026 | Malfunction | covidien surgical | 28-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VATS (VIDEO-ASSISTED THORACIC SURGERY) LOBECTOMY, THE JAWS WOULD NOT CLOSE COMPLETELY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
12-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VATS (VIDEO-ASSISTED THORACIC SURGERY) LOBECTOMY, THE JAWS WOULD NOT CLOSE COMPLETELY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
12-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VATS (VIDEO-ASSISTED THORACIC SURGERY) LOBECTOMY, THE JAWS WOULD NOT CLOSE COMPLETELY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
08-01-2026 | Malfunction | covidien | 28-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, THE HANDLE WAS STUCK AND WAS UNABLE TO BE SQUEEZED BY THE SURGEON, AND THE STAPLER CARTRIDGE DID NOT FIRE AT ALL. THE SCRUB NURSE CONFIRMED THAT THE STAPLER CARTRIDGE WAS LOADED CORRECTLY ON THE STAPLER BEFORE PASSING IT OFF TO THE SURGEON FOR USE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-01-2026 | Malfunction | covidien surgical | 28-01-2026 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, THE HANDLE WAS STUCK AND WAS UNABLE TO BE SQUEEZED BY THE SURGEON, AND THE STAPLER CARTRIDGE DID NOT FIRE AT ALL. THE SCRUB NURSE CONFIRMED THAT THE STAPLER CARTRIDGE WAS LOADED CORRECTLY ON THE STAPLER BEFORE PASSING IT OFF TO THE SURGEON FOR USE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-09-2025 | Injury | covidien surgical | 28-01-2026 | UNKNOWN EGIA TRI STAPLE | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO COMPARE THE STAPLER METHOD WITH THE CONVENTIONAL METHOD IN ROBOT-ASSISTED PANCREATICODUODENECTOMY (RPD) AND ITS IMPACT ON THE INCIDENCE OF POSTOPERATIVE PANCREATIC FISTULA. THE STAPLER METHOD INCLUDED USE OF A 60MM PURPLE TRI-STAPLE RELOAD. THERE WERE 134 PATIENTS INCLUDED FROM OCTOBER 2020 TO DECEMBER 2024. THERE WERE EIGHT PATIENTS WITH POSTOPERATIVE PANCREATIC FISTULA ¿ ONE OF WHICH REQUIRED A REOPERATION. TWO PATIENTS HAD POSTOPERATIVE ABDOMINAL BLEEDING AND ONE PATIENT HAD POST-PANCREATECTOMY ACUTE PANCREATITIS. FOR HEALTH IMPACTS, TWO PATIENTS REQUIRED UNSPECIFIED INVASIVE INTERVENTIONS, AND TWO PATIENTS REQUIRED HOSPITAL RE-ADMISSION.
15-09-2025 | Injury | covidien surgical | 28-01-2026 | UNKNOWN EGIA TRI STAPLE | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO COMPARE THE STAPLER METHOD WITH THE CONVENTIONAL METHOD IN ROBOT-ASSISTED PANCREATICODUODENECTOMY (RPD) AND ITS IMPACT ON THE INCIDENCE OF POSTOPERATIVE PANCREATIC FISTULA. THE STAPLER METHOD INCLUDED USE OF A 60MM PURPLE TRI-STAPLE RELOAD. THERE WERE 134 PATIENTS INCLUDED FROM OCTOBER 2020 TO DECEMBER 2024. THERE WERE EIGHT PATIENTS WITH POSTOPERATIVE PANCREATIC FISTULA ¿ ONE OF WHICH REQUIRED A REOPERATION. TWO PATIENTS HAD POSTOPERATIVE ABDOMINAL BLEEDING AND ONE PATIENT HAD POST-PANCREATECTOMY ACUTE PANCREATITIS. FOR HEALTH IMPACTS, TWO PATIENTS REQUIRED UNSPECIFIED INVASIVE INTERVENTIONS, AND TWO PATIENTS REQUIRED HOSPITAL RE-ADMISSION.
15-09-2025 | Injury | covidien surgical | 28-01-2026 | UNKNOWN EGIA TRI STAPLE | Adverse Event Without Identified Device or Use Problem | A24 | Pancreatitis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO COMPARE THE STAPLER METHOD WITH THE CONVENTIONAL METHOD IN ROBOT-ASSISTED PANCREATICODUODENECTOMY (RPD) AND ITS IMPACT ON THE INCIDENCE OF POSTOPERATIVE PANCREATIC FISTULA. THE STAPLER METHOD INCLUDED USE OF A 60MM PURPLE TRI-STAPLE RELOAD. THERE WERE 134 PATIENTS INCLUDED FROM OCTOBER 2020 TO DECEMBER 2024. THERE WERE EIGHT PATIENTS WITH POSTOPERATIVE PANCREATIC FISTULA ¿ ONE OF WHICH REQUIRED A REOPERATION. TWO PATIENTS HAD POSTOPERATIVE ABDOMINAL BLEEDING AND ONE PATIENT HAD POST-PANCREATECTOMY ACUTE PANCREATITIS. FOR HEALTH IMPACTS, TWO PATIENTS REQUIRED UNSPECIFIED INVASIVE INTERVENTIONS, AND TWO PATIENTS REQUIRED HOSPITAL RE-ADMISSION.
15-09-2025 | Injury | covidien surgical | 28-01-2026 | UNKNOWN EGIA TRI STAPLE | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO COMPARE THE STAPLER METHOD WITH THE CONVENTIONAL METHOD IN ROBOT-ASSISTED PANCREATICODUODENECTOMY (RPD) AND ITS IMPACT ON THE INCIDENCE OF POSTOPERATIVE PANCREATIC FISTULA. THE STAPLER METHOD INCLUDED USE OF A 60MM BLACK TRI-STAPLE RELOAD. THERE WERE 134 PATIENTS INCLUDED FROM (B)(6) 2020 TO (B)(6) 2024. THERE WERE EIGHT PATIENTS WITH POSTOPERATIVE PANCREATIC FISTULA ¿ ONE OF WHICH REQUIRED A REOPERATION. TWO PATIENTS HAD POSTOPERATIVE ABDOMINAL BLEEDING AND ONE PATIENT HAD POST-PANCREATECTOMY ACUTE PANCREATITIS. FOR HEALTH IMPACTS, TWO PATIENTS REQUIRED UNSPECIFIED INVASIVE INTERVENTIONS, AND TWO PATIENTS REQUIRED HOSPITAL RE-ADMISSION.
15-09-2025 | Injury | covidien surgical | 28-01-2026 | UNKNOWN EGIA TRI STAPLE | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO COMPARE THE STAPLER METHOD WITH THE CONVENTIONAL METHOD IN ROBOT-ASSISTED PANCREATICODUODENECTOMY (RPD) AND ITS IMPACT ON THE INCIDENCE OF POSTOPERATIVE PANCREATIC FISTULA. THE STAPLER METHOD INCLUDED USE OF A 60MM BLACK TRI-STAPLE RELOAD. THERE WERE 134 PATIENTS INCLUDED FROM (B)(6) 2020 TO (B)(6) 2024. THERE WERE EIGHT PATIENTS WITH POSTOPERATIVE PANCREATIC FISTULA ¿ ONE OF WHICH REQUIRED A REOPERATION. TWO PATIENTS HAD POSTOPERATIVE ABDOMINAL BLEEDING AND ONE PATIENT HAD POST-PANCREATECTOMY ACUTE PANCREATITIS. FOR HEALTH IMPACTS, TWO PATIENTS REQUIRED UNSPECIFIED INVASIVE INTERVENTIONS, AND TWO PATIENTS REQUIRED HOSPITAL RE-ADMISSION.
15-09-2025 | Injury | covidien surgical | 28-01-2026 | UNKNOWN EGIA TRI STAPLE | Adverse Event Without Identified Device or Use Problem | A24 | Pancreatitis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO COMPARE THE STAPLER METHOD WITH THE CONVENTIONAL METHOD IN ROBOT-ASSISTED PANCREATICODUODENECTOMY (RPD) AND ITS IMPACT ON THE INCIDENCE OF POSTOPERATIVE PANCREATIC FISTULA. THE STAPLER METHOD INCLUDED USE OF A 60MM BLACK TRI-STAPLE RELOAD. THERE WERE 134 PATIENTS INCLUDED FROM (B)(6) 2020 TO (B)(6) 2024. THERE WERE EIGHT PATIENTS WITH POSTOPERATIVE PANCREATIC FISTULA ¿ ONE OF WHICH REQUIRED A REOPERATION. TWO PATIENTS HAD POSTOPERATIVE ABDOMINAL BLEEDING AND ONE PATIENT HAD POST-PANCREATECTOMY ACUTE PANCREATITIS. FOR HEALTH IMPACTS, TWO PATIENTS REQUIRED UNSPECIFIED INVASIVE INTERVENTIONS, AND TWO PATIENTS REQUIRED HOSPITAL RE-ADMISSION.
12-11-2025 | Malfunction | covidien | 28-01-2026 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | nan
14-01-2026 | Malfunction | covidien | 28-01-2026 | ENDO GIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOW ANTERIOR RESECTION, THE RECTUM WAS INTENTIONALLY RESECTED WITH TWO FIRINGS, AFTER FIRING THE FIRST 45MM RELOAD, WHEN PRESSING THE CENTER CONTROL AGAIN, THE CARTRIDGE MOVED SLIGHTLY. AFTERWARD, WHEN RETRACTING THE KNIFE WITH THE CENTER CONTROL, THE KNIFE ONLY RETRACTED BY 1 CM. AND DID NOT RETURN TO THE END. THE KNIFE WAS RETRACTED TO THE END USING THE MANUAL ADAPTER TOOL. ANOTHER ADAPTER SHAFT AND RELOAD WERE USED, AND THE RESECTION WAS COMPLETED. THERE WAS NO PATIENT INJURY.
14-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOW ANTERIOR RESECTION, THE RECTUM WAS INTENTIONALLY RESECTED WITH TWO FIRINGS, AFTER FIRING THE FIRST 45MM RELOAD, WHEN PRESSING THE CENTER CONTROL AGAIN, THE CARTRIDGE MOVED SLIGHTLY. AFTERWARD, WHEN RETRACTING THE KNIFE WITH THE CENTER CONTROL, THE KNIFE ONLY RETRACTED BY 1 CM. AND DID NOT RETURN TO THE END. THE KNIFE WAS RETRACTED TO THE END USING THE MANUAL ADAPTER TOOL. ANOTHER ADAPTER SHAFT AND RELOAD WERE USED, AND THE RESECTION WAS COMPLETED. THERE WAS NO PATIENT INJURY.
14-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOW ANTERIOR RESECTION, THE RECTUM WAS INTENTIONALLY RESECTED WITH TWO FIRINGS, AFTER FIRING THE FIRST 45MM RELOAD, WHEN PRESSING THE CENTER CONTROL AGAIN, THE CARTRIDGE MOVED SLIGHTLY. AFTERWARD, WHEN RETRACTING THE KNIFE WITH THE CENTER CONTROL, THE KNIFE ONLY RETRACTED BY 1 CM. AND DID NOT RETURN TO THE END. THE KNIFE WAS RETRACTED TO THE END USING THE MANUAL ADAPTER TOOL. ANOTHER ADAPTER SHAFT AND RELOAD WERE USED, AND THE RESECTION WAS COMPLETED. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A OPEN COLON RESECTION, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A OPEN COLON RESECTION, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A OPEN COLON RESECTION, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A OPEN COLON RESECTION, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A OPEN COLON RESECTION, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A OPEN COLON RESECTION, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A OPEN COLON RESECTION, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A OPEN COLON RESECTION, THE STAPLING WAS NOT COMPLETED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH (IPOM) PROCEDURE, DURING MESH FIXATION, THE DEVICE MALFUNCTIONED, WHEREIN MULTIPLE TACKERS WERE NOT COMPLETELY RELEASED FOLLOWING APPLICATION. DESPITE APPROPRIATE POSITIONING AND STANDARD DEPLOYMENT TECHNIQUE, THE TACKS REMAINED PARTIALLY ATTACHED TO THE DELIVERY MECHANISM INSTEAD OF BEING FULLY RELEASED INTO THE TISSUE. THE SURGEON USED A NEW TACKER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2025 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS SURGERY, THE POWERED HANDLE WAS ARMED WITH THE SHELL, XL ADAPTER, AND RELOAD. THE RELOAD COULD BE MOUNTED BUT THE DEVICE DID NOT FIRE, ANOTHER RELOAD WAS USED AND IT DID NOT WORK EITHER. THE SURGEON ASKED TO CHANGE ALL THE DEVICES, TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2025 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS SURGERY, THE POWERED HANDLE WAS ARMED WITH THE SHELL, XL ADAPTER, AND RELOAD. THE RELOAD COULD BE MOUNTED BUT THE DEVICE DID NOT FIRE, ANOTHER RELOAD WAS USED AND IT DID NOT WORK EITHER. THE SURGEON ASKED TO CHANGE ALL THE DEVICES, TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2025 | Malfunction | covidien | 28-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS SURGERY, THE POWERED HANDLE WAS ARMED WITH THE SHELL, XL ADAPTER, AND RELOAD. THE RELOAD COULD BE MOUNTED BUT THE DEVICE DID NOT FIRE, ANOTHER RELOAD WAS USED AND IT DID NOT WORK EITHER. THE SURGEON ASKED TO CHANGE ALL THE DEVICES, TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2025 | Malfunction | covidien | 28-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS SURGERY, THE POWERED HANDLE WAS ARMED WITH THE SHELL, XL ADAPTER, AND RELOAD. THE RELOAD COULD BE MOUNTED BUT THE DEVICE DID NOT FIRE, ANOTHER RELOAD WAS USED AND IT DID NOT WORK EITHER. THE SURGEON ASKED TO CHANGE ALL THE DEVICES, TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien surgical | 28-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING VIDEO-ASSISTED THORACIC SURGERY (VATS), THE INSTRUMENT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
14-11-2025 | Malfunction | covidien | 28-01-2026 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE MANUAL HANDLE WAS DIFFICULT TO FIRE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
14-11-2025 | Malfunction | covidien | 28-01-2026 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE MANUAL HANDLE WAS DIFFICULT TO FIRE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
26-12-2025 | Malfunction | covidien surgical | 28-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A VIDEO-ASSISTED THORACIC SURGERY (VATS) PROCEDURE, WHILE IN THE LUNG, THE POWERED STAPLER PARTIALLY FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2025 | Malfunction | covidien surgical | 28-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A VIDEO-ASSISTED THORACIC SURGERY (VATS) PROCEDURE, WHILE IN THE LUNG, THE POWERED STAPLER PARTIALLY FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien surgical | 28-01-2026 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, FIRING ISSUE OCCURRED WITH THE RELOAD COMPONENT. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
06-05-2025 | Injury | ethicon endo-surgery | 29-01-2026 | ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SANTULLO F, DE CICCO NARDONE A, EL HALABIEH MA, LODOLI C, ABATINI C, FERRACCI F, CAMPOLO F, BENVENGA G, SCAMBIA G, PACELLI F, IANIERI MM. TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) AFTER SEGMENTAL COLORECTAL RESECTION FOR DEEP ENDOMETRIOSIS: TECHNICAL NOTES AND CASE SERIES. ARCH GYNECOL OBSTET. 2025 AUG;312(2):555-561. DOI: 10.1007/S00404-025-08040-4. EPUB 2025 MAY 6. PMID: 40327066; PMCID: PMC12334526. THE AIM OF THIS STUDY WAS TO EVALUATE THE SAFETY AND FEASIBILITY OF TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) IN PATIENTS UNDERGOING COLORECTAL RESECTION FOR THE TREATMENT OF DEEP ENDOMETRIOSIS (DE) AFFECTING THE BOWEL. BETWEEN JANUARY 2021 AND AUGUST 2024, 33 CONSECUTIVE PATIENTS WITH DE TREATED WITH SEGMENTAL COLORECTAL RESECTION WERE ENROLLED. ECHELON FLEX (EES) AND ENDOPATH STAPLERS (EES) WERE USED TO CARRIED OUT UNTIL REACHING THE RECTAL WALL BELOW THE ENDOMETRIOTIC NODULE, AND THEN THE RECTUM WAS TRANSECTED WITH A LINEAR STAPLER. VICRYL (ETH) WAS USED TO BOUND AT THE HOLE OF THE TIP. REPORTED COMPLICATION: ECHELON FLEX (EES). ENDOPATH STAPLERS (EES). 0 VICRYL (ETH). POSTOPERATIVE BLEEDING OF THE ANASTOMOSIS (N=2). TREATMENT: NOT REPORTED. POSTOPERATIVE ILEUS (N=1). TREATMENT: MEDICALLY TREATED. IN CONCLUSIONS, THE STUDY SHOWS THAT TICA IS A SAFE AND FEASIBLE TECHNIQUE IN PATIENTS WHO UNDERGO LAPAROSCOPIC COLORECTAL RESECTION FOR DE. FURTHER STUDIES ARE STILL REQUIRED TO EVALUATE A POSSIBLE IMPROVEMENT IN FUNCTIONAL OUTCOMES.
06-05-2025 | Injury | ethicon endo-surgery | 29-01-2026 | ECHELON 60 | Adverse Event Without Identified Device or Use Problem | A24 | Obstruction/Occlusion | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SANTULLO F, DE CICCO NARDONE A, EL HALABIEH MA, LODOLI C, ABATINI C, FERRACCI F, CAMPOLO F, BENVENGA G, SCAMBIA G, PACELLI F, IANIERI MM. TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) AFTER SEGMENTAL COLORECTAL RESECTION FOR DEEP ENDOMETRIOSIS: TECHNICAL NOTES AND CASE SERIES. ARCH GYNECOL OBSTET. 2025 AUG;312(2):555-561. DOI: 10.1007/S00404-025-08040-4. EPUB 2025 MAY 6. PMID: 40327066; PMCID: PMC12334526. THE AIM OF THIS STUDY WAS TO EVALUATE THE SAFETY AND FEASIBILITY OF TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) IN PATIENTS UNDERGOING COLORECTAL RESECTION FOR THE TREATMENT OF DEEP ENDOMETRIOSIS (DE) AFFECTING THE BOWEL. BETWEEN JANUARY 2021 AND AUGUST 2024, 33 CONSECUTIVE PATIENTS WITH DE TREATED WITH SEGMENTAL COLORECTAL RESECTION WERE ENROLLED. ECHELON FLEX (EES) AND ENDOPATH STAPLERS (EES) WERE USED TO CARRIED OUT UNTIL REACHING THE RECTAL WALL BELOW THE ENDOMETRIOTIC NODULE, AND THEN THE RECTUM WAS TRANSECTED WITH A LINEAR STAPLER. VICRYL (ETH) WAS USED TO BOUND AT THE HOLE OF THE TIP. REPORTED COMPLICATION: ECHELON FLEX (EES). ENDOPATH STAPLERS (EES). 0 VICRYL (ETH). POSTOPERATIVE BLEEDING OF THE ANASTOMOSIS (N=2). TREATMENT: NOT REPORTED. POSTOPERATIVE ILEUS (N=1). TREATMENT: MEDICALLY TREATED. IN CONCLUSIONS, THE STUDY SHOWS THAT TICA IS A SAFE AND FEASIBLE TECHNIQUE IN PATIENTS WHO UNDERGO LAPAROSCOPIC COLORECTAL RESECTION FOR DE. FURTHER STUDIES ARE STILL REQUIRED TO EVALUATE A POSSIBLE IMPROVEMENT IN FUNCTIONAL OUTCOMES.
06-05-2025 | Injury | ethicon endo-surgery | 29-01-2026 | ECHELON 60 | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SANTULLO F, DE CICCO NARDONE A, EL HALABIEH MA, LODOLI C, ABATINI C, FERRACCI F, CAMPOLO F, BENVENGA G, SCAMBIA G, PACELLI F, IANIERI MM. TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) AFTER SEGMENTAL COLORECTAL RESECTION FOR DEEP ENDOMETRIOSIS: TECHNICAL NOTES AND CASE SERIES. ARCH GYNECOL OBSTET. 2025 AUG;312(2):555-561. DOI: 10.1007/S00404-025-08040-4. EPUB 2025 MAY 6. PMID: 40327066; PMCID: PMC12334526. THE AIM OF THIS STUDY WAS TO EVALUATE THE SAFETY AND FEASIBILITY OF TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) IN PATIENTS UNDERGOING COLORECTAL RESECTION FOR THE TREATMENT OF DEEP ENDOMETRIOSIS (DE) AFFECTING THE BOWEL. BETWEEN JANUARY 2021 AND AUGUST 2024, 33 CONSECUTIVE PATIENTS WITH DE TREATED WITH SEGMENTAL COLORECTAL RESECTION WERE ENROLLED. ECHELON FLEX (EES) AND ENDOPATH STAPLERS (EES) WERE USED TO CARRIED OUT UNTIL REACHING THE RECTAL WALL BELOW THE ENDOMETRIOTIC NODULE, AND THEN THE RECTUM WAS TRANSECTED WITH A LINEAR STAPLER. VICRYL (ETH) WAS USED TO BOUND AT THE HOLE OF THE TIP. REPORTED COMPLICATION: ECHELON FLEX (EES). ENDOPATH STAPLERS (EES). 0 VICRYL (ETH). POSTOPERATIVE BLEEDING OF THE ANASTOMOSIS (N=2). TREATMENT: NOT REPORTED. POSTOPERATIVE ILEUS (N=1). TREATMENT: MEDICALLY TREATED. IN CONCLUSIONS, THE STUDY SHOWS THAT TICA IS A SAFE AND FEASIBLE TECHNIQUE IN PATIENTS WHO UNDERGO LAPAROSCOPIC COLORECTAL RESECTION FOR DE. FURTHER STUDIES ARE STILL REQUIRED TO EVALUATE A POSSIBLE IMPROVEMENT IN FUNCTIONAL OUTCOMES.
06-05-2025 | Injury | ethicon endo-surgery | 29-01-2026 | ECHELON 60 | Patient Device Interaction Problem | A01 | Obstruction/Occlusion | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SANTULLO F, DE CICCO NARDONE A, EL HALABIEH MA, LODOLI C, ABATINI C, FERRACCI F, CAMPOLO F, BENVENGA G, SCAMBIA G, PACELLI F, IANIERI MM. TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) AFTER SEGMENTAL COLORECTAL RESECTION FOR DEEP ENDOMETRIOSIS: TECHNICAL NOTES AND CASE SERIES. ARCH GYNECOL OBSTET. 2025 AUG;312(2):555-561. DOI: 10.1007/S00404-025-08040-4. EPUB 2025 MAY 6. PMID: 40327066; PMCID: PMC12334526. THE AIM OF THIS STUDY WAS TO EVALUATE THE SAFETY AND FEASIBILITY OF TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) IN PATIENTS UNDERGOING COLORECTAL RESECTION FOR THE TREATMENT OF DEEP ENDOMETRIOSIS (DE) AFFECTING THE BOWEL. BETWEEN JANUARY 2021 AND AUGUST 2024, 33 CONSECUTIVE PATIENTS WITH DE TREATED WITH SEGMENTAL COLORECTAL RESECTION WERE ENROLLED. ECHELON FLEX (EES) AND ENDOPATH STAPLERS (EES) WERE USED TO CARRIED OUT UNTIL REACHING THE RECTAL WALL BELOW THE ENDOMETRIOTIC NODULE, AND THEN THE RECTUM WAS TRANSECTED WITH A LINEAR STAPLER. VICRYL (ETH) WAS USED TO BOUND AT THE HOLE OF THE TIP. REPORTED COMPLICATION: ECHELON FLEX (EES). ENDOPATH STAPLERS (EES). 0 VICRYL (ETH). POSTOPERATIVE BLEEDING OF THE ANASTOMOSIS (N=2). TREATMENT: NOT REPORTED. POSTOPERATIVE ILEUS (N=1). TREATMENT: MEDICALLY TREATED. IN CONCLUSIONS, THE STUDY SHOWS THAT TICA IS A SAFE AND FEASIBLE TECHNIQUE IN PATIENTS WHO UNDERGO LAPAROSCOPIC COLORECTAL RESECTION FOR DE. FURTHER STUDIES ARE STILL REQUIRED TO EVALUATE A POSSIBLE IMPROVEMENT IN FUNCTIONAL OUTCOMES.
06-05-2025 | Injury | ethicon endo-surgery | 29-01-2026 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SANTULLO F, DE CICCO NARDONE A, EL HALABIEH MA, LODOLI C, ABATINI C, FERRACCI F, CAMPOLO F, BENVENGA G, SCAMBIA G, PACELLI F, IANIERI MM. TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) AFTER SEGMENTAL COLORECTAL RESECTION FOR DEEP ENDOMETRIOSIS: TECHNICAL NOTES AND CASE SERIES. ARCH GYNECOL OBSTET. 2025 AUG;312(2):555-561. DOI: 10.1007/S00404-025-08040-4. EPUB 2025 MAY 6. PMID: 40327066; PMCID: PMC12334526. THE AIM OF THIS STUDY WAS TO EVALUATE THE SAFETY AND FEASIBILITY OF TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) IN PATIENTS UNDERGOING COLORECTAL RESECTION FOR THE TREATMENT OF DEEP ENDOMETRIOSIS (DE) AFFECTING THE BOWEL. BETWEEN JANUARY 2021 AND AUGUST 2024, 33 CONSECUTIVE PATIENTS WITH DE TREATED WITH SEGMENTAL COLORECTAL RESECTION WERE ENROLLED. ECHELON FLEX (EES) AND ENDOPATH STAPLERS (EES) WERE USED TO CARRIED OUT UNTIL REACHING THE RECTAL WALL BELOW THE ENDOMETRIOTIC NODULE, AND THEN THE RECTUM WAS TRANSECTED WITH A LINEAR STAPLER. VICRYL (ETH) WAS USED TO BOUND AT THE HOLE OF THE TIP. REPORTED COMPLICATION: ECHELON FLEX (EES) ENDOPATH STAPLERS (EES) 0 VICRYL (ETH) POSTOPERATIVE BLEEDING OF THE ANASTOMOSIS (N=2) TREATMENT: NOT REPORTED POSTOPERATIVE ILEUS (N=1) TREATMENT: MEDICALLY TREATED IN CONCLUSIONS, THE STUDY SHOWS THAT TICA IS A SAFE AND FEASIBLE TECHNIQUE IN PATIENTS WHO UNDERGO LAPAROSCOPIC COLORECTAL RESECTION FOR DE. FURTHER STUDIES ARE STILL REQUIRED TO EVALUATE A POSSIBLE IMPROVEMENT IN FUNCTIONAL OUTCOMES.
06-05-2025 | Injury | ethicon endo-surgery | 29-01-2026 | ECHELON ENDOPATH STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Obstruction/Occlusion | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SANTULLO F, DE CICCO NARDONE A, EL HALABIEH MA, LODOLI C, ABATINI C, FERRACCI F, CAMPOLO F, BENVENGA G, SCAMBIA G, PACELLI F, IANIERI MM. TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) AFTER SEGMENTAL COLORECTAL RESECTION FOR DEEP ENDOMETRIOSIS: TECHNICAL NOTES AND CASE SERIES. ARCH GYNECOL OBSTET. 2025 AUG;312(2):555-561. DOI: 10.1007/S00404-025-08040-4. EPUB 2025 MAY 6. PMID: 40327066; PMCID: PMC12334526. THE AIM OF THIS STUDY WAS TO EVALUATE THE SAFETY AND FEASIBILITY OF TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) IN PATIENTS UNDERGOING COLORECTAL RESECTION FOR THE TREATMENT OF DEEP ENDOMETRIOSIS (DE) AFFECTING THE BOWEL. BETWEEN JANUARY 2021 AND AUGUST 2024, 33 CONSECUTIVE PATIENTS WITH DE TREATED WITH SEGMENTAL COLORECTAL RESECTION WERE ENROLLED. ECHELON FLEX (EES) AND ENDOPATH STAPLERS (EES) WERE USED TO CARRIED OUT UNTIL REACHING THE RECTAL WALL BELOW THE ENDOMETRIOTIC NODULE, AND THEN THE RECTUM WAS TRANSECTED WITH A LINEAR STAPLER. VICRYL (ETH) WAS USED TO BOUND AT THE HOLE OF THE TIP. REPORTED COMPLICATION: ECHELON FLEX (EES) ENDOPATH STAPLERS (EES) 0 VICRYL (ETH) POSTOPERATIVE BLEEDING OF THE ANASTOMOSIS (N=2) TREATMENT: NOT REPORTED POSTOPERATIVE ILEUS (N=1) TREATMENT: MEDICALLY TREATED IN CONCLUSIONS, THE STUDY SHOWS THAT TICA IS A SAFE AND FEASIBLE TECHNIQUE IN PATIENTS WHO UNDERGO LAPAROSCOPIC COLORECTAL RESECTION FOR DE. FURTHER STUDIES ARE STILL REQUIRED TO EVALUATE A POSSIBLE IMPROVEMENT IN FUNCTIONAL OUTCOMES.
06-05-2025 | Injury | ethicon endo-surgery | 29-01-2026 | ECHELON ENDOPATH STAPLER | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SANTULLO F, DE CICCO NARDONE A, EL HALABIEH MA, LODOLI C, ABATINI C, FERRACCI F, CAMPOLO F, BENVENGA G, SCAMBIA G, PACELLI F, IANIERI MM. TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) AFTER SEGMENTAL COLORECTAL RESECTION FOR DEEP ENDOMETRIOSIS: TECHNICAL NOTES AND CASE SERIES. ARCH GYNECOL OBSTET. 2025 AUG;312(2):555-561. DOI: 10.1007/S00404-025-08040-4. EPUB 2025 MAY 6. PMID: 40327066; PMCID: PMC12334526. THE AIM OF THIS STUDY WAS TO EVALUATE THE SAFETY AND FEASIBILITY OF TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) IN PATIENTS UNDERGOING COLORECTAL RESECTION FOR THE TREATMENT OF DEEP ENDOMETRIOSIS (DE) AFFECTING THE BOWEL. BETWEEN JANUARY 2021 AND AUGUST 2024, 33 CONSECUTIVE PATIENTS WITH DE TREATED WITH SEGMENTAL COLORECTAL RESECTION WERE ENROLLED. ECHELON FLEX (EES) AND ENDOPATH STAPLERS (EES) WERE USED TO CARRIED OUT UNTIL REACHING THE RECTAL WALL BELOW THE ENDOMETRIOTIC NODULE, AND THEN THE RECTUM WAS TRANSECTED WITH A LINEAR STAPLER. VICRYL (ETH) WAS USED TO BOUND AT THE HOLE OF THE TIP. REPORTED COMPLICATION: ECHELON FLEX (EES) ENDOPATH STAPLERS (EES) 0 VICRYL (ETH) POSTOPERATIVE BLEEDING OF THE ANASTOMOSIS (N=2) TREATMENT: NOT REPORTED POSTOPERATIVE ILEUS (N=1) TREATMENT: MEDICALLY TREATED IN CONCLUSIONS, THE STUDY SHOWS THAT TICA IS A SAFE AND FEASIBLE TECHNIQUE IN PATIENTS WHO UNDERGO LAPAROSCOPIC COLORECTAL RESECTION FOR DE. FURTHER STUDIES ARE STILL REQUIRED TO EVALUATE A POSSIBLE IMPROVEMENT IN FUNCTIONAL OUTCOMES.
06-05-2025 | Injury | ethicon endo-surgery | 29-01-2026 | ECHELON ENDOPATH STAPLER | Patient Device Interaction Problem | A01 | Obstruction/Occlusion | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SANTULLO F, DE CICCO NARDONE A, EL HALABIEH MA, LODOLI C, ABATINI C, FERRACCI F, CAMPOLO F, BENVENGA G, SCAMBIA G, PACELLI F, IANIERI MM. TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) AFTER SEGMENTAL COLORECTAL RESECTION FOR DEEP ENDOMETRIOSIS: TECHNICAL NOTES AND CASE SERIES. ARCH GYNECOL OBSTET. 2025 AUG;312(2):555-561. DOI: 10.1007/S00404-025-08040-4. EPUB 2025 MAY 6. PMID: 40327066; PMCID: PMC12334526. THE AIM OF THIS STUDY WAS TO EVALUATE THE SAFETY AND FEASIBILITY OF TOTALLY INTRACORPOREAL COLORECTAL ANASTOMOSIS (TICA) IN PATIENTS UNDERGOING COLORECTAL RESECTION FOR THE TREATMENT OF DEEP ENDOMETRIOSIS (DE) AFFECTING THE BOWEL. BETWEEN JANUARY 2021 AND AUGUST 2024, 33 CONSECUTIVE PATIENTS WITH DE TREATED WITH SEGMENTAL COLORECTAL RESECTION WERE ENROLLED. ECHELON FLEX (EES) AND ENDOPATH STAPLERS (EES) WERE USED TO CARRIED OUT UNTIL REACHING THE RECTAL WALL BELOW THE ENDOMETRIOTIC NODULE, AND THEN THE RECTUM WAS TRANSECTED WITH A LINEAR STAPLER. VICRYL (ETH) WAS USED TO BOUND AT THE HOLE OF THE TIP. REPORTED COMPLICATION: ECHELON FLEX (EES) ENDOPATH STAPLERS (EES) 0 VICRYL (ETH) POSTOPERATIVE BLEEDING OF THE ANASTOMOSIS (N=2) TREATMENT: NOT REPORTED POSTOPERATIVE ILEUS (N=1) TREATMENT: MEDICALLY TREATED IN CONCLUSIONS, THE STUDY SHOWS THAT TICA IS A SAFE AND FEASIBLE TECHNIQUE IN PATIENTS WHO UNDERGO LAPAROSCOPIC COLORECTAL RESECTION FOR DE. FURTHER STUDIES ARE STILL REQUIRED TO EVALUATE A POSSIBLE IMPROVEMENT IN FUNCTIONAL OUTCOMES.
14-01-2026 | Malfunction | davol inc sub c.r bard inc -1213643 | 29-01-2026 | PERMAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN INCISIONAL HERNIA REPAIR WITHOUT OBSTRUCTION, THE PERMAFIX ENHANCED FIXATION DEVICE FASTENERS FAILED TO FIXATE. IT WAS REPORTED THAT ONLY A PARTIAL TACK WAS DEPLOYED IN THE AREA OF FIXATION, WHICH INVOLVED SOFT TISSUE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY USING AN ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
09-01-2026 | Malfunction | covidien surgical | 29-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, PRIOR TO FIRING, THE JAWS WOULD NOT CLOSE COMPLETELY. THE ISSUE WAS RESOLVED BY REPLACING THE RELOAD, AND THE SURGERY WAS COMPLETED SUCCESSFULLY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
09-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, PRIOR TO FIRING, THE JAWS WOULD NOT CLOSE COMPLETELY. THE ISSUE WAS RESOLVED BY REPLACING THE RELOAD, AND THE SURGERY WAS COMPLETED SUCCESSFULLY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
09-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, PRIOR TO FIRING, THE JAWS WOULD NOT CLOSE COMPLETELY. THE ISSUE WAS RESOLVED BY REPLACING THE RELOAD, AND THE SURGERY WAS COMPLETED SUCCESSFULLY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
27-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID RESECTION, AFTER PRESSING THE GREEN BUTTON, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. A NEW HANDLE WITH THE SAME RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2026 | Malfunction | covidien | 29-01-2026 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING VIDEO-ASSISTED THORACIC SURGERY (VATS), THE INSTRUMENT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
06-01-2026 | Malfunction | covidien | 29-01-2026 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING VIDEO-ASSISTED THORACIC SURGERY (VATS), THE INSTRUMENT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
06-01-2026 | Malfunction | covidien | 29-01-2026 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING VIDEO-ASSISTED THORACIC SURGERY (VATS), THE INSTRUMENT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
06-01-2026 | Malfunction | covidien | 29-01-2026 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING VIDEO-ASSISTED THORACIC SURGERY (VATS), THE INSTRUMENT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED WITH THE SAME PRODUCT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
07-01-2026 | Injury | covidien surgical | 29-01-2026 | EEA ORVIL | Incomplete or Inadequate Connection | A1207 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A TOTAL GASTRECTOMY, DURING THE ANASTOMOSIS OF THE ESOPHAGEAL STUMP AND THE ELEVATED JEJUNUM, THE CIRCULAR STAPLER HANDLE AND THE ORAL ANVIL COULD NOT BE ATTACHED WHICH RESULTED IN THE ANASTOMOSIS BEING UNABLE TO BE PERFORMED. THE PROCEDURE WAS CHANGED FROM LAPAROSCOPY TO LAPAROTOMY AND THORACOTOMY. THE SURGICAL TIME WAS EXTENDED FOR THREE HOURS. IT WAS REPORTED THAT THE PATIENT HAD ACUTE ABDOMINAL PAIN POST-OPERATIVELY.
07-01-2026 | Injury | covidien surgical | 29-01-2026 | EEA ORVIL | Incomplete or Inadequate Connection | A1207 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A TOTAL GASTRECTOMY, DURING THE ANASTOMOSIS OF THE ESOPHAGEAL STUMP AND THE ELEVATED JEJUNUM, THE CIRCULAR STAPLER HANDLE AND THE ORAL ANVIL COULD NOT BE ATTACHED WHICH RESULTED IN THE ANASTOMOSIS BEING UNABLE TO BE PERFORMED. THE PROCEDURE WAS CHANGED FROM LAPAROSCOPY TO LAPAROTOMY AND THORACOTOMY. THE SURGICAL TIME WAS EXTENDED FOR THREE HOURS. IT WAS REPORTED THAT THE PATIENT HAD ACUTE ABDOMINAL PAIN POST-OPERATIVELY.
07-01-2026 | Injury | covidien surgical | 29-01-2026 | EEA ORVIL | Incomplete or Inadequate Connection | A1207 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A TOTAL GASTRECTOMY, DURING THE ANASTOMOSIS OF THE ESOPHAGEAL STUMP AND THE ELEVATED JEJUNUM, THE CIRCULAR STAPLER HANDLE AND THE ORAL ANVIL COULD NOT BE ATTACHED WHICH RESULTED IN THE ANASTOMOSIS BEING UNABLE TO BE PERFORMED. THE PROCEDURE WAS CHANGED FROM LAPAROSCOPY TO LAPAROTOMY AND THORACOTOMY. THE SURGICAL TIME WAS EXTENDED FOR THREE HOURS. IT WAS REPORTED THAT THE PATIENT HAD ACUTE ABDOMINAL PAIN POST-OPERATIVELY.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 60MM LONG | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE TISSUE WAS STUCK ON THE RIDGES OF THE RELOAD. IT WAS THEN THEY SAW THE RELOAD MOVING AT THE PROXIMAL END, EVEN THOUGH IT HAD ALREADY BEEN LOADED PROPERLY EARLIER. THEY OPENED ANOTHER STAPLER AND USED NEW RELOADS AND IT WAS THE SAME. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 60MM LONG | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE TISSUE WAS STUCK ON THE RIDGES OF THE RELOAD. IT WAS THEN THEY SAW THE RELOAD MOVING AT THE PROXIMAL END, EVEN THOUGH IT HAD ALREADY BEEN LOADED PROPERLY EARLIER. THEY OPENED ANOTHER STAPLER AND USED NEW RELOADS AND IT WAS THE SAME. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 60MM LONG | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE TISSUE WAS STUCK ON THE RIDGES OF THE RELOAD. IT WAS THEN THEY SAW THE RELOAD MOVING AT THE PROXIMAL END, EVEN THOUGH IT HAD ALREADY BEEN LOADED PROPERLY EARLIER. THEY OPENED ANOTHER STAPLER AND USED NEW RELOADS AND IT WAS THE SAME. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 60MM LONG | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE TISSUE WAS STUCK ON THE RIDGES OF THE RELOAD. IT WAS THEN THEY SAW THE RELOAD MOVING AT THE PROXIMAL END, EVEN THOUGH IT HAD ALREADY BEEN LOADED PROPERLY EARLIER. THEY OPENED ANOTHER STAPLER AND USED NEW RELOADS AND IT WAS THE SAME. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Break | A0401 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A VERTICAL GASTRECTOMY, THE RELOAD LOCKED ON TISSUE AND COULD NOT BE REMOVED AFTER FIRING. THE SCALPEL FAILED TO RETRACT, THE RELOAD BECAME JAMMED, AND THE RETRACTION KEY BROKE. THE CONTROLS ON THE POWER HANDLE DID NOT WORK TO OPEN THE STAPLE AFTER FIRING THE RELOAD, AND THE MANUAL RETRACT TOOL BROKE. THE SURGEON SLOWLY SLID THE TISSUE BACK INTO PLACE, RELEASING THE RELOAD, AND REINFORCED THE AREA WITH AN ABSORBABLE, MONOFILAMENT SUTURE, WHICH RESULTED IN AN EXTENDED SURGICAL TIME OF 20 MINUTES. THE SURGERY CONTINUED AFTER THE EQUIPMENT WAS REPLACED.
16-12-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 45MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PULMONARY SEGMENTECTOMY, THE KNIFE STOPPED AT THE 3RD FIRING. THE JAWS OPENED, SO THE DEVICE WAS FIRED AGAIN, BUT THE KNIFE STOPPED AGAIN. AT THAT TIME, THE JAWS DID NOT OPEN, SO THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE. THE LUNGS WERE NOT CUT, AND THERE WERE NO PROBLEMS FOR THE PATIENT. THE CARTRIDGE COLOR WAS GREEN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
16-12-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 45MM COMPACT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PULMONARY SEGMENTECTOMY, THE KNIFE STOPPED AT THE 3RD FIRING. THE JAWS OPENED, SO THE DEVICE WAS FIRED AGAIN, BUT THE KNIFE STOPPED AGAIN. AT THAT TIME, THE JAWS DID NOT OPEN, SO THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE. THE LUNGS WERE NOT CUT, AND THERE WERE NO PROBLEMS FOR THE PATIENT. THE CARTRIDGE COLOR WAS GREEN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
18-12-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 60MM STANDARD | Use of Device Problem | A23 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, A NURSE NOT ONLY OPENED/CLOSED THE DEVICE, BUT ALSO PRESSED THE FIRING TRIGGER WHEN FUNCTIONAL TEST BEFORE USE FOR THE PATIENT. THE DEVICE WAS NOT USED FOR THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
27-06-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A PARENCHYMOTOMY ON A HEPATECTOMY, THE STAPLER BECAME STUCK IN THE CLOSED POSITION WITH NO POSSIBILITY OF REOPENING IT. DISSECTION WAS PERFORMED TO REPOSITION THE VASCULAR CLAMPS, AS WELL AS CUTTING THE VEIN ON THE COMMON TRUNK WITH SCISSORS, IN ORDER TO EXTRACT THE STAPLER. THIS RESULTED IN AN APPROXIMATE 90-MINUTE EXTENSION OF THE OPERATIVE TIME, BLEEDING, AND THE SACRIFICE OF THE MIDDLE HEPATIC VEIN.
08-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PARENCHYMOTOMY ON A HEPATECTOMY, THE STAPLER BECAME STUCK IN THE CLOSED POSITION WITH NO POSSIBILITY OF REOPENING IT. DISSECTION WAS PERFORMED TO REPOSITION THE VASCULAR CLAMPS, AS WELL AS CUTTING THE VEIN ON THE COMMON TRUNK WITH SCISSORS, IN ORDER TO EXTRACT THE STAPLER. THIS RESULTED IN AN APPROXIMATE 90-MINUTE EXTENSION OF THE OPERATIVE TIME, BLEEDING, AND THE SACRIFICE OF THE MIDDLE HEPATIC VEIN.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Mechanical Problem | A05 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE FIRST FIRE FAILED. THE SURGEON TRIED TO FIRE AGAIN WITH THE SAME RELOAD BUT FAILED AGAIN, AND THEN NOTED COULD NOT REVERSE THE KNIFE AND COULD NOT OPEN THE JAW. TRIED RBRB TO REVERSE THE KNIFE BUT STILL FAILED. ANOTHER ACCESS WAS BUILT AND ANOTHER DEVICE WAS USED TO CUT OFF THE TISSUE TO REMOVE THE IMPACTED DEVICE. THE DEVICE WAS REMOVED EVENTUALLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE FIRST FIRE FAILED. THE SURGEON TRIED TO FIRE AGAIN WITH THE SAME RELOAD BUT FAILED AGAIN, AND THEN NOTED COULD NOT REVERSE THE KNIFE AND COULD NOT OPEN THE JAW. TRIED RBRB TO REVERSE THE KNIFE BUT STILL FAILED. ANOTHER ACCESS WAS BUILT AND ANOTHER DEVICE WAS USED TO CUT OFF THE TISSUE TO REMOVE THE IMPACTED DEVICE. THE DEVICE WAS REMOVED EVENTUALLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE FIRST FIRE FAILED. THE SURGEON TRIED TO FIRE AGAIN WITH THE SAME RELOAD BUT FAILED AGAIN, AND THEN NOTED COULD NOT REVERSE THE KNIFE AND COULD NOT OPEN THE JAW. TRIED RBRB TO REVERSE THE KNIFE BUT STILL FAILED. ANOTHER ACCESS WAS BUILT AND ANOTHER DEVICE WAS USED TO CUT OFF THE TISSUE TO REMOVE THE IMPACTED DEVICE. THE DEVICE WAS REMOVED EVENTUALLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE FIRST FIRE FAILED. THE SURGEON TRIED TO FIRE AGAIN WITH THE SAME RELOAD BUT FAILED AGAIN, AND THEN NOTED COULD NOT REVERSE THE KNIFE AND COULD NOT OPEN THE JAW. TRIED RBRB TO REVERSE THE KNIFE BUT STILL FAILED. ANOTHER ACCESS WAS BUILT AND ANOTHER DEVICE WAS USED TO CUT OFF THE TISSUE TO REMOVE THE IMPACTED DEVICE. THE DEVICE WAS REMOVED EVENTUALLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE FIRST FIRE FAILED. THE SURGEON TRIED TO FIRE AGAIN WITH THE SAME RELOAD BUT FAILED AGAIN, AND THEN NOTED COULD NOT REVERSE THE KNIFE AND COULD NOT OPEN THE JAW. TRIED RBRB TO REVERSE THE KNIFE BUT STILL FAILED. ANOTHER ACCESS WAS BUILT AND ANOTHER DEVICE WAS USED TO CUT OFF THE TISSUE TO REMOVE THE IMPACTED DEVICE. THE DEVICE WAS REMOVED EVENTUALLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE FIRST FIRE FAILED. THE SURGEON TRIED TO FIRE AGAIN WITH THE SAME RELOAD BUT FAILED AGAIN, AND THEN NOTED COULD NOT REVERSE THE KNIFE AND COULD NOT OPEN THE JAW. TRIED RBRB TO REVERSE THE KNIFE BUT STILL FAILED. ANOTHER ACCESS WAS BUILT AND ANOTHER DEVICE WAS USED TO CUT OFF THE TISSUE TO REMOVE THE IMPACTED DEVICE. THE DEVICE WAS REMOVED EVENTUALLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE FIRST FIRE FAILED. THE SURGEON TRIED TO FIRE AGAIN WITH THE SAME RELOAD BUT FAILED AGAIN, AND THEN NOTED COULD NOT REVERSE THE KNIFE AND COULD NOT OPEN THE JAW. TRIED RBRB TO REVERSE THE KNIFE BUT STILL FAILED. ANOTHER ACCESS WAS BUILT AND ANOTHER DEVICE WAS USED TO CUT OFF THE TISSUE TO REMOVE THE IMPACTED DEVICE. THE DEVICE WAS REMOVED EVENTUALLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE FIRST FIRE FAILED. THE SURGEON TRIED TO FIRE AGAIN WITH THE SAME RELOAD BUT FAILED AGAIN, AND THEN NOTED COULD NOT REVERSE THE KNIFE AND COULD NOT OPEN THE JAW. TRIED RBRB TO REVERSE THE KNIFE BUT STILL FAILED. ANOTHER ACCESS WAS BUILT AND ANOTHER DEVICE WAS USED TO CUT OFF THE TISSUE TO REMOVE THE IMPACTED DEVICE. THE DEVICE WAS REMOVED EVENTUALLY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-01-2026 | Malfunction | covidien surgical | 29-01-2026 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A PARENCHYMOTOMY ON A HEPATECTOMY, THE STAPLER BECAME STUCK IN THE CLOSED POSITION WITH NO POSSIBILITY OF REOPENING IT. DISSECTION WAS PERFORMED TO REPOSITION THE VASCULAR CLAMPS, AS WELL AS CUTTING THE VEIN ON THE COMMON TRUNK WITH SCISSORS, IN ORDER TO EXTRACT THE STAPLER. THIS RESULTED IN AN APPROXIMATE 90-MINUTE EXTENSION OF THE OPERATIVE TIME, BLEEDING, AND THE SACRIFICE OF THE MIDDLE HEPATIC VEIN.
08-01-2026 | Malfunction | covidien surgical | 29-01-2026 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PARENCHYMOTOMY ON A HEPATECTOMY, THE STAPLER BECAME STUCK IN THE CLOSED POSITION WITH NO POSSIBILITY OF REOPENING IT. DISSECTION WAS PERFORMED TO REPOSITION THE VASCULAR CLAMPS, AS WELL AS CUTTING THE VEIN ON THE COMMON TRUNK WITH SCISSORS, IN ORDER TO EXTRACT THE STAPLER. THIS RESULTED IN AN APPROXIMATE 90-MINUTE EXTENSION OF THE OPERATIVE TIME, BLEEDING, AND THE SACRIFICE OF THE MIDDLE HEPATIC VEIN.
13-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE JAW ON THE DEVICE DID NOT OPEN. THERE WAS A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
13-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 45MM COMPACT | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | nan
05-05-2025 | Injury | coopersurgical | 29-01-2026 | INSORB SKIN STAPLER | Failure to Form Staple | A050704 | Wound Dehiscence | Event Description: SUPERFICIAL WOUND DEHISCENCE- INCISION. 3 CM OF THE RIGHT LATERAL INCISION OPENED SUPERFICIALLY ALONG THE SKIN EDGE. THE INSORB STAPLES HAVE FALLEN OUT OF THIS SIDE OF THE INCISION.
05-05-2025 | Injury | coopersurgical | 29-01-2026 | INSORB SKIN STAPLER | Therapeutic or Diagnostic Output Failure | A090406 | Wound Dehiscence | Event Description: SUPERFICIAL WOUND DEHISCENCE- INCISION. 3 CM OF THE RIGHT LATERAL INCISION OPENED SUPERFICIALLY ALONG THE SKIN EDGE. THE INSORB STAPLES HAVE FALLEN OUT OF THIS SIDE OF THE INCISION.
15-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ASCENDING COLECTOMY, IT WAS OBSERVED THAT SOME OF THE STAPLES WERE NOT FULLY FORMED AND A FEW DID NOT HAVE BENT STAPLE LEGS AT ALL AT THE SECOND AND FIFTH FIRING. SURGEON OVERSEWED STAPLE LINE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
17-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE KNIFE ON THE DEVICE DID NOT CUT. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | POWERED PLUS COMPACT 45 | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER WOULD DEPLOY STAPLES BUT WOULD NOT CUT. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. UNKNOWN AS TO HOW THE PROCEDURE WAS COMPLETED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | POWERED PLUS COMPACT 45 | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLER WOULD DEPLOY STAPLES BUT WOULD NOT CUT. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. UNKNOWN AS TO HOW THE PROCEDURE WAS COMPLETED. NO ADDITIONAL INFORMATION PROVIDED.
29-01-2026 | Malfunction | covidien surgical | 29-01-2026 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INGUINAL HERNIA PROCEDURE, WHILE TACKING THE MESH, A TACK FROM THE DEVICE FELL INTO THE CAVITY OF THE PATIENT. THE TACK WAS SUBSEQUENTLY RETRIEVED FROM THE PATIENT. A NEW TACKER WAS USED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
04-04-2025 | Injury | coopersurgical | 29-01-2026 | INSORB SUBCUTICULAR STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: FULL WOUND DEHISCENCE POD 1 S/P PRIMARY C/S AFTER USE OF INSORB STAPLER FOR CLOSURE. RESULTING IN EXTENDED HOSPITALIZATION AND BLOOD TRANSFUSION.
04-04-2025 | Injury | coopersurgical | 29-01-2026 | INSORB SUBCUTICULAR STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: FULL WOUND DEHISCENCE POD 1 S/P PRIMARY C/S AFTER USE OF INSORB STAPLER FOR CLOSURE. RESULTING IN EXTENDED HOSPITALIZATION AND BLOOD TRANSFUSION.
16-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE JAMMED WHEN FIRED. IT MADE INCOMPLETE LINES OF STAPLES; AND THE BLADE RECOIL MECHANISM HAD TO BE USED TO OPEN THE DEVICE. EVEN SO, THE OPENING WAS DIFFICULT. A LOT OF FORCE HAD TO BE APPLIED TO THE FIRING LEVER. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
16-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE JAMMED WHEN FIRED. IT MADE INCOMPLETE LINES OF STAPLES; AND THE BLADE RECOIL MECHANISM HAD TO BE USED TO OPEN THE DEVICE. EVEN SO, THE OPENING WAS DIFFICULT. A LOT OF FORCE HAD TO BE APPLIED TO THE FIRING LEVER. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
16-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE JAMMED WHEN FIRED. IT MADE INCOMPLETE LINES OF STAPLES; AND THE BLADE RECOIL MECHANISM HAD TO BE USED TO OPEN THE DEVICE. EVEN SO, THE OPENING WAS DIFFICULT. A LOT OF FORCE HAD TO BE APPLIED TO THE FIRING LEVER. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS STUCK IN ARTICULATED POSITION IN BETWEEN FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | Hemorrhage/Blood Loss/Bleeding | nan
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | nan
14-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | nan
09-05-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ABDOMINAL ILECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS LOCKED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE WITH A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 60MM STANDARD | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THERE WAS REPETITIVE ERROR OF WAITING (LOADING) BEFORE ACTIVATION. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
28-08-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 60MM STANDARD | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THEY WERE NOT ABLE TO INSERT CARTRIDGE PROPERLY EVEN THOUGH THERE WAS NOT ANY FOREIGN BODIES INSIDE THE JAW OF THE DEVICE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
13-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RIGHT UPPER LOBE LOBECTOMY, AFTER SUCCESSFULLY FIRING THE TWO RELOADS, THE HANDLE BECAME UNUSABLE AND COULD NOT BE SQUEEZED, RESULTING IN THE STAPLER BEING UNABLE TO FIRE. THE ISSUE WAS RESOLVED BY OPENING A NEW HANDLE TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
17-09-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | GST BLUE RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACIC PROCEDURE, THE SURGEON ACTIVATED THE GUN, AND THIS REFILL ONLY ADVANCED THE FIRST TWO STAPLES THEN IT WAS RETURNED. BUT WHEN RE-ACTIVATED, IT REACHED THE SAME PLACE AND DID NOT ADVANCE. THEREFORE, A NEW PISTOL WAS REQUESTED IN ORDER TO CONTINUE WITH THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-09-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS BLOCKED. TO RESOLVE THE ISSUE THEY DID A MANUAL OVERRIDE AND TOOK A NEW PVE. SURGERY WAS PROLONGED WITH 5 MINUTES. WHEN THE STAPLER WAS REOPENED, SURGEON OBSERVED NEITHER CUTTING NOR STAPLING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-09-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS BLOCKED. TO RESOLVE THE ISSUE THEY DID A MANUAL OVERRIDE AND TOOK A NEW PVE. SURGERY WAS PROLONGED WITH 5 MINUTES. WHEN THE STAPLER WAS REOPENED, SURGEON OBSERVED NEITHER CUTTING NOR STAPLING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 60MM STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE JAW DID NOT FULLY OPEN AFTER FIRING; AND WHEN TRYING TO CLOSE THE JAW, THE JAW DID NOT FULLY CLOSE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAD NO FUNCTION. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, AFTER FIRING, A LINE WAS LEFT EXPOSED. SOME STAPLES WERE POORLY FORMED. THE STAPLES WERE STRETCHED OUT. THE STAPLE LINE BLED FOR LESS THAN 3CC. THE STAPLE LINE WAS REINFORCED WITH SUTURE TO RESOLVE THE ISSUE.
14-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, AFTER FIRING, A LINE WAS LEFT EXPOSED. SOME STAPLES WERE POORLY FORMED. THE STAPLES WERE STRETCHED OUT. THE STAPLE LINE BLED FOR LESS THAN 3CC. THE STAPLE LINE WAS REINFORCED WITH SUTURE TO RESOLVE THE ISSUE.
01-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE LOCKED THAT THEY HAD TO USE MANUAL OVERRIDE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
05-11-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 45MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOST BATTERY POWER AFTER 4TH FIRING. THE BATTERY WAS REMOVED AND REPLACED, HOME BUTTON PRESSED BUT STILL NO POWER. THE HANDLE WOULD NOT CLOSE OVER LUNG TISSUE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
05-11-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 60MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOST BATTERY POWER AFTER 4TH FIRING. THE BATTERY WAS REMOVED AND REPLACED, HOME BUTTON PRESSED BUT STILL NO POWER. THE HANDLE WOULD NOT CLOSE OVER LUNG TISSUE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE STAPLER DID NOT FIRE. SUTURING WAS DONE AS STAPLE LINE REPLACEMENT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2026 | Malfunction | covidien surgical | 29-01-2026 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE STAPLER DID NOT FIRE. SUTURING WAS DONE AS STAPLE LINE REPLACEMENT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | GST GOLD RELOAD 45MM 6 ROW | Material Integrity Problem | A04 | Foreign Body In Patient | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, PART OF THE GOLD PLASTIC PART OF THE GST CARTRIDGE WAS CHIPPED AFTER FIRING. THE DEVICE WAS FIRED ON A PREVIOUS STAPLE LINE. THE FRAGMENTS WERE COLLECTED AND DISCARDED TO COMPLETE THE CASE. THE DEVICE WAS USED FOR INTERLOBAR FORMATION. THERE WAS NO EFFECT ON THE PATIENT. THE CARTRIDGE COLOR WAS GREEN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
14-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC WEDGE RESECTION ON LUNG TRANSECTION, THE STAPLER WAS UNABLE TO FIRE, THE GREEN BUTTON WAS PRESSED BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
26-12-2025 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A VIDEO-ASSISTED THORACIC SURGERY (VATS) PROCEDURE, WHILE IN THE LUNG, THE POWERED STAPLER PARTIALLY FIRED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS BLOCKED AFTER SEVERAL USES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | covidien surgical | 29-01-2026 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC WEDGE RESECTION ON LUNG TRANSECTION, THE STAPLER WAS UNABLE TO FIRE, THE GREEN BUTTON WAS PRESSED BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-01-2026 | Malfunction | ethicon endo-surgery | 29-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS BLOCKED AFTER SEVERAL USES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
18-11-2025 | Malfunction | covidien surgical | 29-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LOBECTOMY, WHILE FIRING VASCULAR LOADS, STAPLING PROCEEDED WITHOUT ANY ISSUES AND WAS CONSIDERED SAFE. HOWEVER, WHEN THE SURGEON SWITCHED TO A BLACK RELOAD FOR THE BRONCHUS, THE RELOAD INITIATED THE FIRING SEQUENCE BUT STOPPED ALMOST IMMEDIATELY, RELEASING OPEN STAPLES BUT DID NOT STAPLE THE TISSUE, AND THE KNIFE DID NOT ADVANCE. AFTERWARDS, THE SURGEON OPENED A MANUAL STAPLER AND FIRED ANOTHER BLACK RELOAD FROM THE SAME BATCH WITHOUT ANY ISSUES. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD HAS PUSHERS DEPLOYED UP TO 3CM CUT LINE. THE RELOAD HAD PARTIALLY DEPLOYED STAPLES DISTAL TO THE EXPOSED PUSHERS.
18-11-2025 | Malfunction | covidien surgical | 29-01-2026 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LOBECTOMY, WHILE FIRING VASCULAR LOADS, STAPLING PROCEEDED WITHOUT ANY ISSUES AND WAS CONSIDERED SAFE. HOWEVER, WHEN THE SURGEON SWITCHED TO A BLACK RELOAD FOR THE BRONCHUS, THE RELOAD INITIATED THE FIRING SEQUENCE BUT STOPPED ALMOST IMMEDIATELY, RELEASING OPEN STAPLES BUT DID NOT STAPLE THE TISSUE, AND THE KNIFE DID NOT ADVANCE. AFTERWARDS, THE SURGEON OPENED A MANUAL STAPLER AND FIRED ANOTHER BLACK RELOAD FROM THE SAME BATCH WITHOUT ANY ISSUES. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD HAS PUSHERS DEPLOYED UP TO 3CM CUT LINE. THE RELOAD HAD PARTIALLY DEPLOYED STAPLES DISTAL TO THE EXPOSED PUSHERS.
20-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, A STAPLE FORMATION OCCURRED IN WHICH THE 'B' SHAPE WAS NOT FORMED ON ONE STAPLE IN THE STAPLE LINE. THE ISSUE WAS IDENTIFIED ON THE DISCARDED SIDE OF THE STOMACH, NO INTERVENTION NEEDED TO BE DONE.
01-01-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAD NO FUNCTION. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-01-2026 | Malfunction | covidien surgical | 29-01-2026 | MAXTACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC ABDOMINAL WALL INCISION HERNIA REPAIR, THE HEAD BROKE DURING TACKING, BUT THERE WAS NO PROBLEM WITH THE TACKING. THERE WAS NO PATIENT INJURY.
17-11-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE FIRED ONCE AND THEN STOPPED WORKING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 29-01-2026 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Failure to Read Input Signal | A1301 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AN ISSUE OCCURRED WHILE ATTEMPTING TO TRANSECT THE DUODENUM. THE USER FIRST REPLACED THE POWER HANDLE AND SHELL, THEN TRIED USING THE ORIGINAL LOADING UNIT. HOWEVER, THE STAPLER FAILED TO FIRE THERELOAD, THE LOADING UNIT DID NOT RECOGNIZE. SUBSEQUENTLY, THE STAFF REPLACED THE LOADING UNIT WITH THE ORIGINAL STAPLE CARTRIDGE. THIS EXTENDED THE SURGICAL TIME BY A FEW MINUTES. THE LOADING UNIT AND CARTRIDGE WERE RECOGNIZED BY THE STAPLER, AND THE SURGEON WAS ABLE TO FIRE THE DEVICE SUCCESSFULLY.
08-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AN ISSUE OCCURRED WHILE ATTEMPTING TO TRANSECT THE DUODENUM. THE USER FIRST REPLACED THE POWER HANDLE AND SHELL, THEN TRIED USING THE ORIGINAL LOADING UNIT. HOWEVER, THE STAPLER FAILED TO FIRE THERELOAD, THE LOADING UNIT DID NOT RECOGNIZE. SUBSEQUENTLY, THE STAFF REPLACED THE LOADING UNIT WITH THE ORIGINAL STAPLE CARTRIDGE. THIS EXTENDED THE SURGICAL TIME BY A FEW MINUTES. THE LOADING UNIT AND CARTRIDGE WERE RECOGNIZED BY THE STAPLER, AND THE SURGEON WAS ABLE TO FIRE THE DEVICE SUCCESSFULLY.
08-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AN ISSUE OCCURRED WHILE ATTEMPTING TO TRANSECT THE DUODENUM. THE USER FIRST REPLACED THE POWER HANDLE AND SHELL, THEN TRIED USING THE ORIGINAL LOADING UNIT. HOWEVER, THE STAPLER FAILED TO FIRE THERELOAD, THE LOADING UNIT DID NOT RECOGNIZE. SUBSEQUENTLY, THE STAFF REPLACED THE LOADING UNIT WITH THE ORIGINAL STAPLE CARTRIDGE. THIS EXTENDED THE SURGICAL TIME BY A FEW MINUTES. THE LOADING UNIT AND CARTRIDGE WERE RECOGNIZED BY THE STAPLER, AND THE SURGEON WAS ABLE TO FIRE THE DEVICE SUCCESSFULLY.
08-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AN ISSUE OCCURRED WHILE ATTEMPTING TO TRANSECT THE DUODENUM. THE USER FIRST REPLACED THE POWER HANDLE AND SHELL, THEN TRIED USING THE ORIGINAL LOADING UNIT. HOWEVER, THE STAPLER FAILED TO FIRE THERELOAD, THE LOADING UNIT DID NOT RECOGNIZE. SUBSEQUENTLY, THE STAFF REPLACED THE LOADING UNIT WITH THE ORIGINAL STAPLE CARTRIDGE. THIS EXTENDED THE SURGICAL TIME BY A FEW MINUTES. THE LOADING UNIT AND CARTRIDGE WERE RECOGNIZED BY THE STAPLER, AND THE SURGEON WAS ABLE TO FIRE THE DEVICE SUCCESSFULLY.
07-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A VERTICAL GASTRECTOMY, THE RELOAD LOCKED ON TISSUE AND COULD NOT BE REMOVED AFTER FIRING. THE SCALPEL FAILED TO RETRACT, THE RELOAD BECAME JAMMED, AND THE RETRACTION KEY BROKE. THE CONTROLS ON THE POWER HANDLE DID NOT WORK TO OPEN THE STAPLE AFTER FIRING THE RELOAD, AND THE MANUAL RETRACT TOOL BROKE. THE SURGEON SLOWLY SLID THE TISSUE BACK INTO PLACE, RELEASING THE RELOAD, AND REINFORCED THE AREA WITH AN ABSORBABLE, MONOFILAMENT SUTURE, WHICH RESULTED IN AN EXTENDED SURGICAL TIME OF 20 MINUTES. THE SURGERY CONTINUED AFTER THE EQUIPMENT WAS REPLACED.
07-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A VERTICAL GASTRECTOMY, THE RELOAD LOCKED ON TISSUE AND COULD NOT BE REMOVED AFTER FIRING. THE SCALPEL FAILED TO RETRACT, THE RELOAD BECAME JAMMED, AND THE RETRACTION KEY BROKE. THE CONTROLS ON THE POWER HANDLE DID NOT WORK TO OPEN THE STAPLE AFTER FIRING THE RELOAD, AND THE MANUAL RETRACT TOOL BROKE. THE SURGEON SLOWLY SLID THE TISSUE BACK INTO PLACE, RELEASING THE RELOAD, AND REINFORCED THE AREA WITH AN ABSORBABLE, MONOFILAMENT SUTURE, WHICH RESULTED IN AN EXTENDED SURGICAL TIME OF 20 MINUTES. THE SURGERY CONTINUED AFTER THE EQUIPMENT WAS REPLACED.
07-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A VERTICAL GASTRECTOMY, THE RELOAD LOCKED ON TISSUE AND COULD NOT BE REMOVED AFTER FIRING. THE SCALPEL FAILED TO RETRACT, THE RELOAD BECAME JAMMED, AND THE RETRACTION KEY BROKE. THE CONTROLS ON THE POWER HANDLE DID NOT WORK TO OPEN THE STAPLE AFTER FIRING THE RELOAD, AND THE MANUAL RETRACT TOOL BROKE. THE SURGEON SLOWLY SLID THE TISSUE BACK INTO PLACE, RELEASING THE RELOAD, AND REINFORCED THE AREA WITH AN ABSORBABLE, MONOFILAMENT SUTURE, WHICH RESULTED IN AN EXTENDED SURGICAL TIME OF 20 MINUTES. THE SURGERY CONTINUED AFTER THE EQUIPMENT WAS REPLACED.
07-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A VERTICAL GASTRECTOMY, THE RELOAD LOCKED ON TISSUE AND COULD NOT BE REMOVED AFTER FIRING. THE SCALPEL FAILED TO RETRACT, THE RELOAD BECAME JAMMED, AND THE RETRACTION KEY BROKE. THE CONTROLS ON THE POWER HANDLE DID NOT WORK TO OPEN THE STAPLE AFTER FIRING THE RELOAD, AND THE MANUAL RETRACT TOOL BROKE. THE SURGEON SLOWLY SLID THE TISSUE BACK INTO PLACE, RELEASING THE RELOAD, AND REINFORCED THE AREA WITH AN ABSORBABLE, MONOFILAMENT SUTURE, WHICH RESULTED IN AN EXTENDED SURGICAL TIME OF 20 MINUTES. THE SURGERY CONTINUED AFTER THE EQUIPMENT WAS REPLACED.
07-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A VERTICAL GASTRECTOMY, THE RELOAD LOCKED ON TISSUE AND COULD NOT BE REMOVED AFTER FIRING. THE SCALPEL FAILED TO RETRACT, THE RELOAD BECAME JAMMED, AND THE RETRACTION KEY BROKE. THE CONTROLS ON THE POWER HANDLE DID NOT WORK TO OPEN THE STAPLE AFTER FIRING THE RELOAD, AND THE MANUAL RETRACT TOOL BROKE. THE SURGEON SLOWLY SLID THE TISSUE BACK INTO PLACE, RELEASING THE RELOAD, AND REINFORCED THE AREA WITH AN ABSORBABLE, MONOFILAMENT SUTURE, WHICH RESULTED IN AN EXTENDED SURGICAL TIME OF 20 MINUTES. THE SURGERY CONTINUED AFTER THE EQUIPMENT WAS REPLACED.
07-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A VERTICAL GASTRECTOMY, THE RELOAD LOCKED ON TISSUE AND COULD NOT BE REMOVED AFTER FIRING. THE SCALPEL FAILED TO RETRACT, THE RELOAD BECAME JAMMED, AND THE RETRACTION KEY BROKE. THE CONTROLS ON THE POWER HANDLE DID NOT WORK TO OPEN THE STAPLE AFTER FIRING THE RELOAD, AND THE MANUAL RETRACT TOOL BROKE. THE SURGEON SLOWLY SLID THE TISSUE BACK INTO PLACE, RELEASING THE RELOAD, AND REINFORCED THE AREA WITH AN ABSORBABLE, MONOFILAMENT SUTURE, WHICH RESULTED IN AN EXTENDED SURGICAL TIME OF 20 MINUTES. THE SURGERY CONTINUED AFTER THE EQUIPMENT WAS REPLACED.
08-01-2026 | Malfunction | covidien surgical | 29-01-2026 | TRI-STAPLE 2.0 | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLING WAS NOT PERFORMED WHEN FIRING WAS PERFORMED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE AND THE OPERATION WAS SUCCESSFULLY COMPLETED.
07-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE ADAPTER DOES NOT SEAT PROPERLY INTO THE CLAMSHELL. THE CONNECTION IS LOOSE. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
06-01-2026 | Malfunction | covidien surgical | 29-01-2026 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE FUNCTIONED NORMALLY UP TO THE LCD (LIQUID CRYSTAL DISPLAY) SCREEN GUIDING CARTRIDGE CONNECTION, BUT AFTER THE CARTRIDGE WAS CONNECTED, THE OPENING AND CLOSING OPERATION COULD NOT BE PERFORMED AND UPON RESTARTING, THE CARTRIDGE MOVED LEFT AND RIGHT AUTOMATICALLY AND THEN STOPPED, REMAINING UNABLE TO OPEN OR CLOSE. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
06-01-2026 | Malfunction | covidien surgical | 29-01-2026 | UNKNOWN ENDO GIA SULU | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE FUNCTIONED NORMALLY UP TO THE LCD (LIQUID CRYSTAL DISPLAY) SCREEN GUIDING CARTRIDGE CONNECTION, BUT AFTER THE CARTRIDGE WAS CONNECTED, THE OPENING AND CLOSING OPERATION COULD NOT BE PERFORMED AND UPON RESTARTING, THE CARTRIDGE MOVED LEFT AND RIGHT AUTOMATICALLY AND THEN STOPPED, REMAINING UNABLE TO OPEN OR CLOSE. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
06-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE FUNCTIONED NORMALLY UP TO THE LCD (LIQUID CRYSTAL DISPLAY) SCREEN GUIDING CARTRIDGE CONNECTION, BUT AFTER THE CARTRIDGE WAS CONNECTED, THE OPENING AND CLOSING OPERATION COULD NOT BE PERFORMED AND UPON RESTARTING, THE CARTRIDGE MOVED LEFT AND RIGHT AUTOMATICALLY AND THEN STOPPED, REMAINING UNABLE TO OPEN OR CLOSE. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
06-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE FUNCTIONED NORMALLY UP TO THE LCD (LIQUID CRYSTAL DISPLAY) SCREEN GUIDING CARTRIDGE CONNECTION, BUT AFTER THE CARTRIDGE WAS CONNECTED, THE OPENING AND CLOSING OPERATION COULD NOT BE PERFORMED AND UPON RESTARTING, THE CARTRIDGE MOVED LEFT AND RIGHT AUTOMATICALLY AND THEN STOPPED, REMAINING UNABLE TO OPEN OR CLOSE. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
06-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE FUNCTIONED NORMALLY UP TO THE LCD (LIQUID CRYSTAL DISPLAY) SCREEN GUIDING CARTRIDGE CONNECTION, BUT AFTER THE CARTRIDGE WAS CONNECTED, THE OPENING AND CLOSING OPERATION COULD NOT BE PERFORMED AND UPON RESTARTING, THE CARTRIDGE MOVED LEFT AND RIGHT AUTOMATICALLY AND THEN STOPPED, REMAINING UNABLE TO OPEN OR CLOSE. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
06-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE FUNCTIONED NORMALLY UP TO THE LCD (LIQUID CRYSTAL DISPLAY) SCREEN GUIDING CARTRIDGE CONNECTION, BUT AFTER THE CARTRIDGE WAS CONNECTED, THE OPENING AND CLOSING OPERATION COULD NOT BE PERFORMED AND UPON RESTARTING, THE CARTRIDGE MOVED LEFT AND RIGHT AUTOMATICALLY AND THEN STOPPED, REMAINING UNABLE TO OPEN OR CLOSE. A NEW ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
07-01-2026 | Injury | covidien | 29-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER FIRING, THE STAPLE HAD POOR FORMATION, AND THE STAPLE LINE BLED. SUTURING WAS DONE AS STAPLE LINE REINFORCEMENT. THE PATIENT WAS ADMITTED INTO THE INTENSIVE CARE UNIT, AND THE HOSPITALIZATION WAS EXTENDED.
07-01-2026 | Injury | covidien | 29-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER FIRING, THE STAPLE HAD POOR FORMATION, AND THE STAPLE LINE BLED. SUTURING WAS DONE AS STAPLE LINE REINFORCEMENT. THE PATIENT WAS ADMITTED INTO THE INTENSIVE CARE UNIT, AND THE HOSPITALIZATION WAS EXTENDED.
07-01-2026 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLE LINE WAS INCOMPLETE DISTALLY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
05-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC SURGERY, THE JAWS DID NOT OPEN AFTER STAPLING THE BLOOD VESSEL. TREATMENT INTERVENTION WAS NOT PERFORMED. THERE WAS NO PATIENT INJURY.
05-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC SURGERY, THE JAWS DID NOT OPEN AFTER STAPLING THE BLOOD VESSEL. A DDITIONALLY, AFTER WAITING FOR A SHORT TIME, THE JAW EVENTUALLY OPENED SO NO TREATMENT INTERVENTION WAS PERFORMED. THERE WAS NO PATIENT INJURY.
27-01-2025 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Dull, Blunt | A020101 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, WHILE CREATING THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS, THE OUTER ROW DID NOT FIRE. THE KNIFE BLADE DID NOT CUT SMOOTHLY, RESULTING IN A JAGGED LINE, AND DID NOT APPEAR SHARP. THERE WAS A STAPLE FORMATION ISSUE AND AN INCOMPLETE DISTAL STAPLE LINE. THE PATIENT WAS ON BLOOD THINNERS. SUTURING WITH ABSORBABLE SUTURE WAS REQUIRED AS STAPLE LINE REINFORCEMENT TO RESOLVE THE DEVICE MALFUNCTION.
27-01-2025 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, WHILE CREATING THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS, THE OUTER ROW DID NOT FIRE. THE KNIFE BLADE DID NOT CUT SMOOTHLY, RESULTING IN A JAGGED LINE, AND DID NOT APPEAR SHARP. THERE WAS A STAPLE FORMATION ISSUE AND AN INCOMPLETE DISTAL STAPLE LINE. THE PATIENT WAS ON BLOOD THINNERS. SUTURING WITH ABSORBABLE SUTURE WAS REQUIRED AS STAPLE LINE REINFORCEMENT TO RESOLVE THE DEVICE MALFUNCTION.
27-01-2025 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Failure to Cut | A050702 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, WHILE CREATING THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS, THE OUTER ROW DID NOT FIRE. THE KNIFE BLADE DID NOT CUT SMOOTHLY, RESULTING IN A JAGGED LINE, AND DID NOT APPEAR SHARP. THERE WAS A STAPLE FORMATION ISSUE AND AN INCOMPLETE DISTAL STAPLE LINE. THE PATIENT WAS ON BLOOD THINNERS. SUTURING WITH ABSORBABLE SUTURE WAS REQUIRED AS STAPLE LINE REINFORCEMENT TO RESOLVE THE DEVICE MALFUNCTION.
27-01-2025 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, WHILE CREATING THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS, THE OUTER ROW DID NOT FIRE. THE KNIFE BLADE DID NOT CUT SMOOTHLY, RESULTING IN A JAGGED LINE, AND DID NOT APPEAR SHARP. THERE WAS A STAPLE FORMATION ISSUE AND AN INCOMPLETE DISTAL STAPLE LINE. THE PATIENT WAS ON BLOOD THINNERS. SUTURING WITH ABSORBABLE SUTURE WAS REQUIRED AS STAPLE LINE REINFORCEMENT TO RESOLVE THE DEVICE MALFUNCTION.
27-01-2025 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Dull, Blunt | A020101 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, WHILE CREATING THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS, THE OUTER ROW DID NOT FIRE. THE KNIFE BLADE DID NOT CUT SMOOTHLY, RESULTING IN A JAGGED LINE, AND DID NOT APPEAR SHARP. THERE WAS A STAPLE FORMATION ISSUE AND AN INCOMPLETE DISTAL STAPLE LINE. THE PATIENT WAS ON BLOOD THINNERS. SUTURING WITH ABSORBABLE SUTURE WAS REQUIRED AS STAPLE LINE REINFORCEMENT TO RESOLVE THE DEVICE MALFUNCTION.
27-01-2025 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, WHILE CREATING THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS, THE OUTER ROW DID NOT FIRE. THE KNIFE BLADE DID NOT CUT SMOOTHLY, RESULTING IN A JAGGED LINE, AND DID NOT APPEAR SHARP. THERE WAS A STAPLE FORMATION ISSUE AND AN INCOMPLETE DISTAL STAPLE LINE. THE PATIENT WAS ON BLOOD THINNERS. SUTURING WITH ABSORBABLE SUTURE WAS REQUIRED AS STAPLE LINE REINFORCEMENT TO RESOLVE THE DEVICE MALFUNCTION.
27-01-2025 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Failure to Cut | A050702 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, WHILE CREATING THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS, THE OUTER ROW DID NOT FIRE. THE KNIFE BLADE DID NOT CUT SMOOTHLY, RESULTING IN A JAGGED LINE, AND DID NOT APPEAR SHARP. THERE WAS A STAPLE FORMATION ISSUE AND AN INCOMPLETE DISTAL STAPLE LINE. THE PATIENT WAS ON BLOOD THINNERS. SUTURING WITH ABSORBABLE SUTURE WAS REQUIRED AS STAPLE LINE REINFORCEMENT TO RESOLVE THE DEVICE MALFUNCTION.
27-01-2025 | Malfunction | covidien | 29-01-2026 | ENDO GIA | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS, WHILE CREATING THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS, THE OUTER ROW DID NOT FIRE. THE KNIFE BLADE DID NOT CUT SMOOTHLY, RESULTING IN A JAGGED LINE, AND DID NOT APPEAR SHARP. THERE WAS A STAPLE FORMATION ISSUE AND AN INCOMPLETE DISTAL STAPLE LINE. THE PATIENT WAS ON BLOOD THINNERS. SUTURING WITH ABSORBABLE SUTURE WAS REQUIRED AS STAPLE LINE REINFORCEMENT TO RESOLVE THE DEVICE MALFUNCTION.
05-01-2026 | Malfunction | covidien surgical | 29-01-2026 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC SURGERY, THE JAWS DID NOT OPEN AFTER STAPLING THE BLOOD VESSEL. A DDITIONALLY, AFTER WAITING FOR A SHORT TIME, THE JAW EVENTUALLY OPENED SO NO TREATMENT INTERVENTION WAS PERFORMED. THERE WAS NO PATIENT INJURY.
07-01-2026 | Malfunction | covidien surgical | 29-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | nan
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Calibration Problem | A08 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Unintended Movement | A0512 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Noise, Audible | A0508 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Calibration Problem | A08 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Unintended Movement | A0512 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Noise, Audible | A0508 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Calibration Problem | A08 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Unintended Movement | A0512 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Noise, Audible | A0508 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | TRI-STAPLE 2.0 | Break | A0401 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | TRI-STAPLE 2.0 | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | TRI-STAPLE 2.0 | Unintended Movement | A0512 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
26-01-2026 | Malfunction | covidien surgical | 30-01-2026 | TRI-STAPLE 2.0 | Noise, Audible | A0508 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC ANTERIOR SEGMENTECTOMY WITH GLISSON EN BLOC RESECTION, A POWER HANDLE WITH A STANDARD ADAPTER AND CARTRIDGE WAS IN USE WHEN, DURING NEEDLE INSERTION, AN ABNORMAL MECHANICAL NOISE WAS NOTED AND THE DEVICE ARTICULATED UNINTENTIONALLY, AFTER WHICH THE KNIFE FAILED TO RETURN AND THE JAWS COULD NOT OPEN OR MOVE FORWARD OR BACKWARD, RESULTING IN TISSUE BEING CLAMPED AND UNABLE TO BE RELEASED. MULTIPLE TROUBLESHOOTING STEPS WERE ATTEMPTED INTRAOPERATIVELY, INCLUDING RESTARTING THE HANDLE, RECONNECTING COMPONENTS, CALIBRATION ATTEMPTS, USE OF A MANUAL ADAPTER TOOL, AND REPLACEMENT OF THE HANDLE AND SHELL, ALL WITHOUT RESOLUTION, LEADING TO AN EXTENDED OPERATIVE TIME OF APPROXIMATELY THREE HOURS, AND ULTIMATELY THE ISSUE WAS RESOLVED BY APPLYING A POLYMER LOCKING LIGATION SYSTEM ON THE PATIENT SIDE AND DISSECTING THE TISSUE WITH SCISSORS, WHILE THE SPECIMEN SIDE WAS DIVIDED USING A MONOPOLAR INSTRUMENT, ALLOWING SAFE RELEASE OF THE TISSUE AND COMPLETION OF THE PROCEDURE. THE PATIENT OUTCOME WAS NOTED AS PROLONGED OPERATIVE TIME ONLY, WITH THE CARTRIDGE REMAINING CONNECTED TO THE ADAPTER AND PLANNED FOR CLEANING AND STERILIZATION, AND SUBSEQUENT EVALUATION SUGGESTING A CLIP LODGED IN THE TISSUE AS A CONTRIBUTING FACTOR.
28-01-2026 | Malfunction | covidien | 30-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RADICAL SURGERY FOR HILAR CHOLANGIOCARCINOMA (OPEN PROCEDURE, DETACHMENT OF THE HEPATIC VEIN), THE DEVICE WAS UNABLE TO FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-01-2026 | Injury | ethicon endo-surgery | 30-01-2026 | ECHELON 3000 60MM COMPACT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON INFORMED THAT STAPLE LINE BLED POST-OP AND A SCOPE HAD TO BE DONE. THE ANASTOMOSIS SITE IS A GASTROJEJUNOSTOMY, ECH60C AND A BLUE RELOAD GST60B WAS USED AND HE PREVIOUSLY ALWAYS USED BLUE RELOAD AND THERE WAS NO ISSUE. DID THE PATIENT/USER REQUIRE ANY MEDICAL OR SURGICAL INTERVENTION AS A RESULT OF THE EVENT? YES DESCRIPTION A SCOPE WAS DONE. DID THE PATIENT/USER REQUIRE EXTENDED HOSPITALIZATION AS A RESULT OF THE EVENT? NO. WHAT IS THE PATIENT¿S/USER¿S STATUS/OUTCOME FOLLOWING THE EVENT? THE PATIENT IS FINE NOW. WAS THERE A SIGNIFICANT DELAY? UNKNOWN.
16-01-2025 | Injury | ethicon endo-surgery | 30-01-2026 | ECHELON 3000 60MM COMPACT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | nan
13-01-2026 | Injury | ethicon endo-surgery | 30-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | Obstruction/Occlusion | nan
22-09-2025 | Injury | covidien surgical | 30-01-2026 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO EVALUATE PERIOPERATIVE OUTCOMES OF INFERIOR VENA CAVA SEGMENTAL RESECTION IN RENAL CELL CARCINOMA WITH TUMOR THROMBUS. THE STUDY INCLUDED 271 PATIENTS UNDERGOING RADICAL NEPHRECTOMY FROM (B)(6) 2020 TO (B)(6) 2024. BLOOD TRANSFUSIONS WERE NOTED IN 90 PATIENTS INTRAOPERATIVELY. THERE WERE 63 PATIENTS THAT EXPERIENCED POSTOPERATIVE COMPLICATIONS OF ALL GRADES, INCLUDING 8 PATIENTS WITH CLAVIEN-DINDO GRADE III (SERIOUS INJURY) OR HIGHER; HOWEVER, THERE IS NO FURTHER INFORMATION TO CLASSIFY THE CLINICAL SYMPTOMS/SEVERITY.
22-09-2025 | Injury | covidien surgical | 30-01-2026 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO EVALUATE PERIOPERATIVE OUTCOMES OF INFERIOR VENA CAVA SEGMENTAL RESECTION IN RENAL CELL CARCINOMA WITH TUMOR THROMBUS. THE STUDY INCLUDED 271 PATIENTS UNDERGOING RADICAL NEPHRECTOMY FROM (B)(6) 2020 TO (B)(6) 2024. BLOOD TRANSFUSIONS WERE NOTED IN 90 PATIENTS INTRAOPERATIVELY. THERE WERE 63 PATIENTS THAT EXPERIENCED POSTOPERATIVE COMPLICATIONS OF ALL GRADES, INCLUDING 8 PATIENTS WITH CLAVIEN-DINDO GRADE III (SERIOUS INJURY) OR HIGHER; HOWEVER, THERE IS NO FURTHER INFORMATION TO CLASSIFY THE CLINICAL SYMPTOMS/SEVERITY.
28-01-2026 | Malfunction | covidien surgical | 30-01-2026 | MAXTACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE POWERED TACKER WAS REMOVED AFTER A YELLOW LIGHT APPEARED AND NO TACK DEPLOYED. FOLLOWING THE REMOVAL, MULTIPLE ATTEMPTS WERE MADE TO FIRE THE DEVICE. ON THE THIRD ATTEMPT, MULTIPLE TACKS WERE EXPELLED FROM THE DEVICE ONTO THE FLOOR. AT THAT POINT, THE DECISION WAS MADE TO DISCONTINUE USE AND OPEN A NEW TACKER. THERE WAS NO PATIENT INVOLVEMENT.
28-01-2026 | Malfunction | covidien surgical | 30-01-2026 | MAXTACK | Display or Visual Feedback Problem | A0902 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE POWERED TACKER WAS REMOVED AFTER A YELLOW LIGHT APPEARED AND NO TACK DEPLOYED. FOLLOWING THE REMOVAL, MULTIPLE ATTEMPTS WERE MADE TO FIRE THE DEVICE. ON THE THIRD ATTEMPT, MULTIPLE TACKS WERE EXPELLED FROM THE DEVICE ONTO THE FLOOR. AT THAT POINT, THE DECISION WAS MADE TO DISCONTINUE USE AND OPEN A NEW TACKER. THERE WAS NO PATIENT INVOLVEMENT.
05-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SERVICING, THE HANDLE DID NOT CHARGE WHILE ON THE CHARGER. THE CHARGER WAS CONFIRMED TO WORK WITH OTHER HANDLES. THE DEVICE WAS REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ON THE DATA SAVED TO THE SYSTEM, IT WAS NOTED THAT WHILE THE HANDLE WAS INSERTED INTO A CHARGER, THE BATTERY VOLTAGE DRAINED BELOW OPERABLE LEVELS OVER TIME. THE CAUSE OF THE OBSERVED VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE).
05-01-2026 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SERVICING, THE HANDLE DID NOT CHARGE WHILE ON THE CHARGER. THE CHARGER WAS CONFIRMED TO WORK WITH OTHER HANDLES. THE DEVICE WAS REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ON THE DATA SAVED TO THE SYSTEM, IT WAS NOTED THAT WHILE THE HANDLE WAS INSERTED INTO A CHARGER, THE BATTERY VOLTAGE DRAINED BELOW OPERABLE LEVELS OVER TIME. THE CAUSE OF THE OBSERVED VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE).
25-11-2025 | Malfunction | covidien surgical | 30-01-2026 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY TO EVALUATE WHETHER DELIBERATE SURGEON-CONTROLLED SLOW-FIRING OF THE POWERED STAPLERS REDUCES STUMP OOZING COMPARED WITH DEFAULT AUTOMATIC FAST-FIRING MODE DURING PULMONARY VESSEL TRANSECTION. THE STUDY INCLUDED 139 PATIENTS FROM JANUARY 2023 TO DECEMBER 2024. THE INTRAOPERATIVE COMPLICATION OF PROLONGED OOZING [SLOW BLEEDING] WAS NOTED IN SOME OF THE CASES BUT RESOLVED WITHOUT INTERVENTION OR WITH MINOR INTERVENTIONS INCLUDING MANUAL COMPRESSION, SUTURE LIGATION, AND/OR TOPICAL SEALANT.
29-11-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | nan
08-12-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE MALFUNCTIONED ON THE FIRST LOAD. THERE WAS A SHORT SOUND ON FIRE AND STOP FIRING. COULD NOT RELEASE AND THE EMERGENCY RELEASE WAS USED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-12-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE MALFUNCTIONED ON THE FIRST LOAD. THERE WAS A SHORT SOUND ON FIRE AND STOP FIRING. COULD NOT RELEASE AND THE EMERGENCY RELEASE WAS USED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE MOTOR ON THE DEVICE STOPPED DURING FIRING SEQUENCE, AND THE KNIFE RETUNED TO HOME POSITION. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
27-01-2026 | Malfunction | covidien surgical | 30-01-2026 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE, THE INSTRUMENT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON AND SQUEEZE THE HANDLE, AND THE HANDLE DID NOT RUN IDLE OR FEEL FLOPPY. THE ISSUE WAS RESOLVED BY REPLACING THE RELOAD WITH ANOTHER OF THE SAME LOT NUMBER. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
08-01-2026 | Malfunction | covidien | 30-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A NEPHRECTOMY, THE TWO HANDLES COULD NOT BE SQUEEZED WHEN INITIALLY FIRED DESPITE GIVING SEVERAL RELOADS. THERE WAS A BLOOD LOSS OF APPROXIMATELY 1 LITER OF BLOOD DUE TO THE DEVICE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY NOT MORE THAN 30 MINUTES. THE HANDLE WAS CHANGED, AND THE SAME RELOAD WAS USED TO RESOLVE THE ISSUE.
08-01-2026 | Malfunction | covidien | 30-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A NEPHRECTOMY, THE TWO HANDLES COULD NOT BE SQUEEZED WHEN INITIALLY FIRED DESPITE GIVING SEVERAL RELOADS. THERE WAS A BLOOD LOSS OF APPROXIMATELY 1 LITER OF BLOOD DUE TO THE DEVICE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY NOT MORE THAN 30 MINUTES. THE HANDLE WAS CHANGED, AND THE SAME RELOAD WAS USED TO RESOLVE THE ISSUE.
01-01-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | POWERED PLUS STANDARD 45 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNSPECIFIED PROCEDURE, THE FIRST STAPLER¿S BATTERY DID NOT FUNCTION. A SECOND STAPLER WAS THEN USED, BUT ITS BATTERY ALSO FAILED. A THIRD STAPLER WAS OPENED AND SUCCESSFULLY USED TO COMPLETE THE CASE. NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF THE PANCREAS, IT WAS OBSERVED THAT THE DEVICE COULD BE CLOSED BUT NOT POSSIBLE TO STAPLING AND CUT BECAUSE THE HANDLE GOT JAMMED. THEY REPORTED IT DAMAGED THE PANCREATIC PARENCHYMA WITH BLEEDING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF THE PANCREAS, IT WAS OBSERVED THAT THE DEVICE COULD BE CLOSED BUT NOT POSSIBLE TO STAPLING AND CUT BECAUSE THE HANDLE GOT JAMMED. THEY REPORTED IT DAMAGED THE PANCREATIC PARENCHYMA WITH BLEEDING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF THE PANCREAS, IT WAS OBSERVED THAT THE DEVICE COULD BE CLOSED BUT NOT POSSIBLE TO STAPLING AND CUT BECAUSE THE HANDLE GOT JAMMED. THEY REPORTED IT DAMAGED THE PANCREATIC PARENCHYMA WITH BLEEDING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF THE PANCREAS, IT WAS OBSERVED THAT THE DEVICE COULD BE CLOSED BUT NOT POSSIBLE TO STAPLING AND CUT BECAUSE THE HANDLE GOT JAMMED. THEY REPORTED IT DAMAGED THE PANCREATIC PARENCHYMA WITH BLEEDING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF THE PANCREAS, IT WAS OBSERVED THAT THE DEVICE COULD BE CLOSED BUT NOT POSSIBLE TO STAPLING AND CUT BECAUSE THE HANDLE GOT JAMMED. THEY REPORTED IT DAMAGED THE PANCREATIC PARENCHYMA WITH BLEEDING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF THE PANCREAS, IT WAS OBSERVED THAT THE DEVICE COULD BE CLOSED BUT NOT POSSIBLE TO STAPLING AND CUT BECAUSE THE HANDLE GOT JAMMED. THEY REPORTED IT DAMAGED THE PANCREATIC PARENCHYMA WITH BLEEDING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | ECHELON 3000 60MM LONG | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNSPECIFIED PROCEDURE, IT WAS OBSERVED THAT THE FIRST STAPLER DID NOT FIRE PROPERLY, AND THE STAPLE LINE WAS NOT SECURED. THEY OPENED A NEW STAPLER AND IT WORKED BRIEFLY AND IT BEGAN TO MALFUNCTION. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
12-12-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | ECHELON 3000 60MM STANDARD | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE MISFIRED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 30-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT AFTER FOUR FIRINGS, WHILE INSERTING THE BATTERY GUN, THE FIRING SOUND WAS COMING CONTINUOUSLY. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
12-01-2026 | Malfunction | covidien lp llc north haven | 30-01-2026 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC BYPASS PROCEDURE, THE SHELL WAS UNABLE TO RECOGNIZED AND AT THE END OF THE DISCHARGE, THE CHARGE EMITTED A BURSTING SOUND AND REMAINED LODGED IN THE STOMACH TISSUE, PREVENTING ITS RETRIEVAL AND OPENING. A NEW SHELL AND RELOAD WAS USED TO RESOLVE THE ISSUE.
12-01-2026 | Malfunction | covidien lp llc north haven | 30-01-2026 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC BYPASS PROCEDURE, THE SHELL WAS UNABLE TO RECOGNIZED AND AT THE END OF THE DISCHARGE, THE CHARGE EMITTED A BURSTING SOUND AND REMAINED LODGED IN THE STOMACH TISSUE, PREVENTING ITS RETRIEVAL AND OPENING. A NEW SHELL AND RELOAD WAS USED TO RESOLVE THE ISSUE.
09-01-2026 | Malfunction | covidien surgical | 30-01-2026 | ABSORBATACK | Break | A0401 | Insufficient Information | nan
09-01-2026 | Malfunction | covidien surgical | 30-01-2026 | ABSORBATACK | Positioning Failure | A150201 | Insufficient Information | nan
09-01-2026 | Malfunction | covidien surgical | 30-01-2026 | ABSORBATACK | Detachment of Device or Device Component | A0501 | Insufficient Information | nan
04-09-2025 | Malfunction | covidien | 30-01-2026 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. THERE WERE NO STAPLES WHEN TIGHTENING. SUTURING WAS DONE AS A STAPLE LINE REPLACEMENT TO RESOLVE THE ISSUE. THE PROCEDURE WAS EXTENDED DUE TO THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
04-09-2025 | Malfunction | covidien | 30-01-2026 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. THERE WERE NO STAPLES WHEN TIGHTENING. SUTURING WAS DONE AS A STAPLE LINE REPLACEMENT TO RESOLVE THE ISSUE. THE PROCEDURE WAS EXTENDED DUE TO THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
11-01-2026 | Malfunction | covidien surgical | 30-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC RECTOPEXY, DURING THE RESECTION OF THE SIGMOID COLON, THE RELOAD DID NOT CLOSE PROPERLY AND DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A SECOND RELOAD WAS USED BUT HAD THE SAME ISSUE. A THIRD RELOAD WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE ISSUE.
11-01-2026 | Malfunction | covidien surgical | 30-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC RECTOPEXY, DURING THE RESECTION OF THE SIGMOID COLON, THE RELOAD DID NOT CLOSE PROPERLY AND DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A SECOND RELOAD WAS USED BUT HAD THE SAME ISSUE. A THIRD RELOAD WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE ISSUE.
11-01-2026 | Malfunction | covidien surgical | 30-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC RECTOPEXY, DURING THE RESECTION OF THE SIGMOID COLON, THE RELOAD DID NOT CLOSE PROPERLY AND DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A SECOND RELOAD WAS USED BUT HAD THE SAME ISSUE. A THIRD RELOAD WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE ISSUE.
11-01-2026 | Malfunction | covidien surgical | 30-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC RECTOPEXY, DURING THE RESECTION OF THE SIGMOID COLON, THE RELOAD DID NOT CLOSE PROPERLY AND DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A SECOND RELOAD WAS USED BUT HAD THE SAME ISSUE. A THIRD RELOAD WAS USED TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE ISSUE.
28-10-2025 | Malfunction | intuitive surgical | 02-02-2026 | SUREFORM | Positioning Problem | A1502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE SURGICAL TECH WAS UNABLE TO REMOVE STAPLER RELOAD CARTRIDGE FROM THE SUREFORM 60 STAPLER AFTER USING. SEVERAL TEAM MEMBERS ATTEMPTED TO REMOVE UNSUCCESSFULLY. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER TO OBTAIN ADDITIONAL INFORMATION, HOWEVER, NO FURTHER DETAILS ARE AVAILABLE REGARDING THE REPORTED EVENT.
29-12-2025 | Malfunction | intuitive surgical | 02-02-2026 | SUREFORM | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, WHILE STAPLING (CUTTING) A BLOOD VESSEL, THE MESSAGE ¿TISSUE TOO THICK TO CONTINUE¿ APPEARED. THE INSTRUMENT WAS REMOVED FROM THE PATIENT BODY, AND WHEN THE RELOAD WAS DETACHED, THE RELOAD¿S KNIFE TRACK WAS FOUND TO BE SHAVED/ABRADED. NO FRAGMENT FELL INTO THE PATIENT¿S BODY. THE PROCEDURE WAS COMPLETING AS PLANNED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE CUSTOMER AND OBTAINED THE FOLLOWING INFORMATION: NO FRAGMENT FELL INTO THE PATIENT¿S BODY, AND THE ABSENCE OF FRAGMENTS WAS CONFIRMED DURING THE PROCEDURE. THERE IS NO INSTRUMENT RETURN. NO IMAGES ARE AVAILABLE.
29-12-2025 | Malfunction | intuitive surgical | 02-02-2026 | SUREFORM | Material Fragmentation | A040103 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, WHILE STAPLING (CUTTING) A BLOOD VESSEL, THE MESSAGE ¿TISSUE TOO THICK TO CONTINUE¿ APPEARED. THE INSTRUMENT WAS REMOVED FROM THE PATIENT BODY, AND WHEN THE RELOAD WAS DETACHED, THE RELOAD¿S KNIFE TRACK WAS FOUND TO BE SHAVED/ABRADED. NO FRAGMENT FELL INTO THE PATIENT¿S BODY. THE PROCEDURE WAS COMPLETING AS PLANNED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE CUSTOMER AND OBTAINED THE FOLLOWING INFORMATION: NO FRAGMENT FELL INTO THE PATIENT¿S BODY, AND THE ABSENCE OF FRAGMENTS WAS CONFIRMED DURING THE PROCEDURE. THERE IS NO INSTRUMENT RETURN. NO IMAGES ARE AVAILABLE.
09-01-2026 | Malfunction | covidien | 02-02-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING GASTROPLASTY, WHEN THE SURGEON INSERTED THE RELOAD TO THE STAPLER, THE JAWS DID NOT CLOSE COMPLETELY, REMAINING PARTIALLY OPENED. A NEW RELOAD WAS UTILIZED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
10-10-2025 | Injury | covidien surgical | 02-02-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A LAPAROSCOPIC APPENDECTOMY, DURING WHICH A 45 MM PURPLE LOAD STAPLER RELOAD WAS USED TO TRANSECT THE APPENDICEAL STUMP AND POSTOPERATIVELY, THE PATIENT BECAME INCREASINGLY TACHYCARDIC WITH A SIGNIFICANT DROP IN HEMOGLOBIN ON COMPLETE BLOOD COUNT (CBC), NECESSITATING AN URGENT RETURN TO THE OPERATING ROOM ON THE SAME DATE, WHERE ACTIVE BLEEDING WAS IDENTIFIED AT THE APPENDICEAL STUMP DIRECTLY ALONG THE STAPLE LINE, APPROXIMATELY 1.5 LITERS OF INTRAPERITONEAL BLOOD WERE EVACUATED, AND SURGICAL HEMOSTASIS WAS ACHIEVED, RESULTING IN STABILIZATION OF THE PATIENT FOLLOWING RE-OPERATION.
10-10-2025 | Injury | covidien surgical | 02-02-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tachycardia | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A LAPAROSCOPIC APPENDECTOMY, DURING WHICH A 45 MM PURPLE LOAD STAPLER RELOAD WAS USED TO TRANSECT THE APPENDICEAL STUMP AND POSTOPERATIVELY, THE PATIENT BECAME INCREASINGLY TACHYCARDIC WITH A SIGNIFICANT DROP IN HEMOGLOBIN ON COMPLETE BLOOD COUNT (CBC), NECESSITATING AN URGENT RETURN TO THE OPERATING ROOM ON THE SAME DATE, WHERE ACTIVE BLEEDING WAS IDENTIFIED AT THE APPENDICEAL STUMP DIRECTLY ALONG THE STAPLE LINE, APPROXIMATELY 1.5 LITERS OF INTRAPERITONEAL BLOOD WERE EVACUATED, AND SURGICAL HEMOSTASIS WAS ACHIEVED, RESULTING IN STABILIZATION OF THE PATIENT FOLLOWING RE-OPERATION.
08-12-2025 | Malfunction | covidien | 02-02-2026 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE DIVIDING THE RIGHT UPPER LOBE BRONCHUS, THE STAPLER FAILED, AND THE BRONCHUS WAS NOT SEALED. THE SURGEON SUCCESSFULLY REPAIRED THE BRONCHUS THORACOSCOPICALLY WITHOUT ANY ENDOBRONCHIAL LEAK NOTED IN THE THEATER OR POSTOPERATIVELY BY SUTURING AS A STAPLE LINE REPLACEMENT. SURGICAL TIME WAS EXTENDED BY 20 MINUTES. THE PATIENT WAS DISCHARGED FROM THE HOSPITAL WITHOUT ANY COMPLICATIONS.
19-12-2025 | Malfunction | ethicon endo-surgery | 02-02-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED-UP. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | ECHELON 3000 60MM STANDARD | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT TISSUE CUTTING WAS NOT COMPLETE AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | UNK_ENDO LINEAR CUTTERS ECHELON 3000 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC SLEEVE PROCEDURE, THE FIRING OF ECHELON 3000 DISPLAYED MALFORMED STAPLE LEGS ON THE OUTER MOST ROW OF THE SPECIMEN AND THE PATIENT SIDE. THE OTHER TWO ROWS FORMED CORRECTLY. THEY PERFORMED A LEAK TEST THAT SHOWED NO LEAK. THEY PERFORMED ALL ADDITIONAL FIRINGS WITH THE SAME STAPLER WITH NO FURTHER INCIDENT. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | GST GREEN RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO ANOTHER TWO TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | GST GREEN RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO ANOTHER TWO TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | GST GREEN RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO ANOTHER TWO TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | GST BLUE RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-01-2026 | Malfunction | covidien | 02-02-2026 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) S6 (SEGMENT 6) AREA RESECTION, AT THE TIME OF ARTERY 6 (A6) RESECTION, ALTHOUGH IT WAS DIFFICULT TO FIRE, FIRING COULD BE DONE TO THE END. AFTER COMPLETING THE FIRING, RETRACTING THE KNIFE PROVED EXTREMELY DIFFICULT. IT WAS NOTED THAT THE BENDING WAS 4 STEPS TO THE LEFT, AND IT ONLY RETURNED TO THE FIRST STEP DURING RELEASE. DUE TO IT BEING IMMEDIATELY AFTER FIRING INTO A6, IT WAS RELEASED CAREFULLY, AND IT TOOK ABOUT 10 MINUTES TO COMPLETE THE RELEASE. TREATMENT INTERVENTION WAS NOT PERFORMED/NOT NECESSARY. THERE WAS NO PATIENT INJURY.
16-01-2026 | Malfunction | covidien | 02-02-2026 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) S6 (SEGMENT 6) AREA RESECTION, AT THE TIME OF ARTERY 6 (A6) RESECTION, ALTHOUGH IT WAS DIFFICULT TO FIRE, FIRING COULD BE DONE TO THE END. AFTER COMPLETING THE FIRING, RETRACTING THE KNIFE PROVED EXTREMELY DIFFICULT. IT WAS NOTED THAT THE BENDING WAS 4 STEPS TO THE LEFT, AND IT ONLY RETURNED TO THE FIRST STEP DURING RELEASE. DUE TO IT BEING IMMEDIATELY AFTER FIRING INTO A6, IT WAS RELEASED CAREFULLY, AND IT TOOK ABOUT 10 MINUTES TO COMPLETE THE RELEASE. TREATMENT INTERVENTION WAS NOT PERFORMED/NOT NECESSARY. THERE WAS NO PATIENT INJURY.
15-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | ECHELON 3000 60MM STANDARD | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY FOR COLON CANCER, AFTER FIRING TO THE COLON, THE KNIFE THAT MOVED THROUGH TO THE TIP AND RETURNED WAS OPERATED INTERMITTENTLY EVEN THOUGH THEY WERE NOT TOUCHING THE FIRING TRIGGER. THERE WAS NO PROBLEM WITH THE STAPLE LINE, BUT IT WAS REPLACED WITH A NEW ONE JUST TO BE SAFE, AND AN ADDITIONAL RESECTION WAS PERFORMED. THE CARTRIDGE COLOR WAS GOLD. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED.
16-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, EVERYTHING WENT ACCORDING TO PLAN; THEY USED THREE MAGAZINES AS NORMAL. WITH THE FOURTH MAGAZINE, EVERYTHING SEEMED FINE, BUT UPON OPENING THE SLEEVE, IT WAS DISCOVERED THAT THE STAPLER LINE ON THE STOMACH HAD ALMOST NO STAPLED SUTURE. THEREFORE, THE SURGEON HAD TO OVER-SEW EVERYTHING WITH THREAD TO COMPLETE THE PROCEDURE WITH A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
16-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, EVERYTHING WENT ACCORDING TO PLAN; THEY USED THREE MAGAZINES AS NORMAL. WITH THE FOURTH MAGAZINE, EVERYTHING SEEMED FINE, BUT UPON OPENING THE SLEEVE, IT WAS DISCOVERED THAT THE STAPLER LINE ON THE STOMACH HAD ALMOST NO STAPLED SUTURE. THEREFORE, THE SURGEON HAD TO OVER-SEW EVERYTHING WITH THREAD TO COMPLETE THE PROCEDURE WITH A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
16-01-2026 | Malfunction | covidien surgical | 02-02-2026 | TRI-STAPLE 2.0 | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) S6 (SEGMENT 6) AREA RESECTION, AT THE TIME OF ARTERY 6 (A6) RESECTION, ALTHOUGH IT WAS DIFFICULT TO FIRE, FIRING COULD BE DONE TO THE END. AFTER COMPLETING THE FIRING, RETRACTING THE KNIFE PROVED EXTREMELY DIFFICULT. IT WAS NOTED THAT THE BENDING WAS 4 STEPS TO THE LEFT, AND IT ONLY RETURNED TO THE FIRST STEP DURING RELEASE. DUE TO IT BEING IMMEDIATELY AFTER FIRING INTO A6, IT WAS RELEASED CAREFULLY, AND IT TOOK ABOUT 10 MINUTES TO COMPLETE THE RELEASE. TREATMENT INTERVENTION WAS NOT PERFORMED/NOT NECESSARY. THERE WAS NO PATIENT INJURY.
16-01-2026 | Malfunction | covidien surgical | 02-02-2026 | TRI-STAPLE 2.0 | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) S6 (SEGMENT 6) AREA RESECTION, AT THE TIME OF ARTERY 6 (A6) RESECTION, ALTHOUGH IT WAS DIFFICULT TO FIRE, FIRING COULD BE DONE TO THE END. AFTER COMPLETING THE FIRING, RETRACTING THE KNIFE PROVED EXTREMELY DIFFICULT. IT WAS NOTED THAT THE BENDING WAS 4 STEPS TO THE LEFT, AND IT ONLY RETURNED TO THE FIRST STEP DURING RELEASE. DUE TO IT BEING IMMEDIATELY AFTER FIRING INTO A6, IT WAS RELEASED CAREFULLY, AND IT TOOK ABOUT 10 MINUTES TO COMPLETE THE RELEASE. TREATMENT INTERVENTION WAS NOT PERFORMED/NOT NECESSARY. THERE WAS NO PATIENT INJURY.
01-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | ECHELON 3000 45 STANDARD | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE HOME BUTTON AND ARTICULATING BUTTON DID NOT WORK UNTIL A FEW TRIES. ALSO, THE METAL PART OF A RELOAD CARTRIDGE BECAME STUCK IN THE JAW CHANNEL BETWEEN FIRINGS. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | ECHELON 3000 45 STANDARD | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE HOME BUTTON AND ARTICULATING BUTTON DID NOT WORK UNTIL A FEW TRIES. ALSO, THE METAL PART OF A RELOAD CARTRIDGE BECAME STUCK IN THE JAW CHANNEL BETWEEN FIRINGS. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
12-01-2026 | Malfunction | covidien surgical | 02-02-2026 | EEA ORVIL | Incomplete or Inadequate Connection | A1207 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A MINIMALLY INVASIVE ESOPHAGECTOMY, THE CIRCULAR STAPLER AND FIRST ANVIL WOULD NOT CONNECT AND SEAT PROPERLY. A NEW ANVIL WAS OPENED AND IT WAS VERY DIFFICULT TO CONNECT AND SET THE ANVIL TO THE STAPLER. AFTER MUCH TIME, WHICH EXTENDED 45 MINUTES, AND WORKING, THE ORANGE BAND WAS COVERED, THE SAFETY WAS ABLE TO BE RELEASED AND THE SURGEON FIRED THE STAPLER. AS THE SURGEON OPENED, TURNING THE BLACK WING NUT COUNTER CLOCKWISE TWO FULL TURNS, THE SURGEON FELT A CLICK AND REMOVED THE STAPLER FROM THE PATIENT CAVITY. AT THIS TIME THE STAPLER SEPARATED FROM THE ANVIL AND THE ANVIL REMAINED IN THE PATIENTS TISSUE. THE SURGEON WAS ABLE TO PULL IT THROUGH THE ANASTOMOSIS AND REMOVE IT FROM THE PATIENT. IT APPEARED THAT ALL THE STAPLERS WERE FORMED, HOWEVER, ALL OF THE TISSUE WAS NOT CUT. THERE WAS STILL SOME TISSUE AROUND THE ANVIL AND IT APPEARS THE KNIFE BLADE DID NOT FULLY CUT THE TISSUE. SUTURING WAS PERFORMED AS STAPLE LINE REINFORCEMENT.
12-01-2026 | Malfunction | covidien surgical | 02-02-2026 | EEA ORVIL | Detachment of Device or Device Component | A0501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A MINIMALLY INVASIVE ESOPHAGECTOMY, THE CIRCULAR STAPLER AND FIRST ANVIL WOULD NOT CONNECT AND SEAT PROPERLY. A NEW ANVIL WAS OPENED AND IT WAS VERY DIFFICULT TO CONNECT AND SET THE ANVIL TO THE STAPLER. AFTER MUCH TIME, WHICH EXTENDED 45 MINUTES, AND WORKING, THE ORANGE BAND WAS COVERED, THE SAFETY WAS ABLE TO BE RELEASED AND THE SURGEON FIRED THE STAPLER. AS THE SURGEON OPENED, TURNING THE BLACK WING NUT COUNTER CLOCKWISE TWO FULL TURNS, THE SURGEON FELT A CLICK AND REMOVED THE STAPLER FROM THE PATIENT CAVITY. AT THIS TIME THE STAPLER SEPARATED FROM THE ANVIL AND THE ANVIL REMAINED IN THE PATIENTS TISSUE. THE SURGEON WAS ABLE TO PULL IT THROUGH THE ANASTOMOSIS AND REMOVE IT FROM THE PATIENT. IT APPEARED THAT ALL THE STAPLERS WERE FORMED, HOWEVER, ALL OF THE TISSUE WAS NOT CUT. THERE WAS STILL SOME TISSUE AROUND THE ANVIL AND IT APPEARS THE KNIFE BLADE DID NOT FULLY CUT THE TISSUE. SUTURING WAS PERFORMED AS STAPLE LINE REINFORCEMENT.
12-01-2026 | Malfunction | covidien surgical | 02-02-2026 | EEA ORVIL | Incomplete or Inadequate Connection | A1207 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A MINIMALLY INVASIVE ESOPHAGECTOMY, THE CIRCULAR STAPLER AND FIRST ANVIL WOULD NOT CONNECT AND SEAT PROPERLY. A NEW ANVIL WAS OPENED AND IT WAS VERY DIFFICULT TO CONNECT AND SET THE ANVIL TO THE STAPLER. AFTER MUCH TIME, WHICH EXTENDED 45 MINUTES, AND WORKING, THE ORANGE BAND WAS COVERED, THE SAFETY WAS ABLE TO BE RELEASED AND THE SURGEON FIRED THE STAPLER. AS THE SURGEON OPENED, TURNING THE BLACK WING NUT COUNTER CLOCKWISE TWO FULL TURNS, THE SURGEON FELT A CLICK AND REMOVED THE STAPLER FROM THE PATIENT CAVITY. AT THIS TIME THE STAPLER SEPARATED FROM THE ANVIL AND THE ANVIL REMAINED IN THE PATIENTS TISSUE. THE SURGEON WAS ABLE TO PULL IT THROUGH THE ANASTOMOSIS AND REMOVE IT FROM THE PATIENT. IT APPEARED THAT ALL THE STAPLERS WERE FORMED, HOWEVER, ALL OF THE TISSUE WAS NOT CUT. THERE WAS STILL SOME TISSUE AROUND THE ANVIL AND IT APPEARS THE KNIFE BLADE DID NOT FULLY CUT THE TISSUE. SUTURING WAS PERFORMED AS STAPLE LINE REINFORCEMENT.
19-12-2025 | Malfunction | ethicon endo-surgery | 02-02-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
20-12-2025 | Malfunction | ethicon endo-surgery | 02-02-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED UP. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | PROXIMATE*PPH PROCEDURE SET | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE DIDN'T FIRE DURING HEMORRHOID SURGERY IT WAS GRADE 3. THE PROCEDURE WAS COMPLETED SUCCESSFULLY AND NO PATIENT CONSEQUENCES WERE REPORTED.
06-01-2026 | Injury | intuitive surgical | 02-02-2026 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | nan
08-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC SIGMOID COLECTOMY PROCEDURE, THE DEVICE WOULD NOT CLAMP IN. WHEN THE SURGICAL TEAM ATTEMPTED TO ATTACH THE ANVIL TO THE STAPLER, THE ANVIL WOULD NOT CLICK INTO THE TROCAR ON THE STAPLER, AND THEY COULD NOT GET THE CONNECTION BETWEEN THE ANVIL AND THE TROCAR STAPLER. THE SURGEON THEN REQUESTED A NEW DEVICE, AND A NEW STAPLER WAS OPENED. THEY WERE ABLE TO CONNECT AND CLIP THE ANVIL AND THE TROCAR, AND COMPLETE THE FIRING PROCESS. THERE WERE NO ADVERSE EVENTS OR PATIENT CONSEQUENCES REPORTED.
02-01-2026 | Injury | ethicon endo-surgery | 02-02-2026 | ECHELON LINEAR CUTTER 80MM | Difficult to Open or Close | A051104 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE DID NOT CLOSE AND CAUSED THE MESOCOLON TO BLEED. IT WAS REQUIRED TO REDO THE ANASTOMOSIS AND LENGTHENED THE SURGERY. REPAIR OF THE ANASTOMOSIS WITH RISK OF BLEEDING.
02-01-2026 | Injury | ethicon endo-surgery | 02-02-2026 | ECHELON LINEAR CUTTER 80MM | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE DID NOT CLOSE AND CAUSED THE MESOCOLON TO BLEED. IT WAS REQUIRED TO REDO THE ANASTOMOSIS AND LENGTHENED THE SURGERY. REPAIR OF THE ANASTOMOSIS WITH RISK OF BLEEDING.
02-11-2025 | Malfunction | ethicon endo-surgery | 02-02-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
02-12-2025 | Malfunction | ethicon endo-surgery | 02-02-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
28-12-2025 | Malfunction | ethicon endo-surgery | 02-02-2026 | ILS 29MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
06-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | nan
01-01-2026 | Injury | ethicon endo-surgery | 02-02-2026 | ECHELON LINEAR CUTTER 100MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | nan
24-12-2025 | Malfunction | covidien surgical | 02-02-2026 | SIGNIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT SUPERIOR RESECTION WITH ROBOT SUPPORT PROCEDURE, AT THE LEFT SUPERIOR LOBE BRONCHI, THE STAPLE FORMATION WAS INCOMPLETE AND THE FIRING STOPPED IN THE MIDDLE.
24-12-2025 | Malfunction | covidien surgical | 02-02-2026 | SIGNIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT SUPERIOR RESECTION WITH ROBOT SUPPORT PROCEDURE, AT THE LEFT SUPERIOR LOBE BRONCHI, THE STAPLE FORMATION WAS INCOMPLETE AND THE FIRING STOPPED IN THE MIDDLE.
02-02-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, IT WAS OBSERVED THAT THE DEVICE COULD NOT BE FIRED AT THE 3RD FIRING. THE PREVIOUS TWO SHOTS WERE FIRED WITHOUT ANY PROBLEMS, AND THERE WERE NO PROBLEMS SUCH AS THE STAPLES BEING MALFORMED. THE SURGERY HAS BEEN CONTINUED AND COMPLETED USING A REPLACEMENT. THERE WERE NO PATIENT INJURIES SUCH AS BLEEDING BECAUSE OF THIS EVENT. THE CARTRIDGE COLOR WAS WHITE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
14-01-2026 | Malfunction | covidien surgical | 02-02-2026 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, AFTER FIRING THE RELOAD, THE TISSUE REOPENED, AND SUTURING WAS REQUIRED AS A STAPLE LINE REPLACEMENT. THE SURGICAL TIME WAS EXTENDED BY APPROXIMATELY 30 MINUTES.
01-12-2025 | Malfunction | ethicon endo-surgery | 02-02-2026 | 55MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE BOWEL RESECTION PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2026 | Malfunction | ethicon endo-surgery | 02-02-2026 | ILS 29MM CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.