14-12-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | PWRD 23MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE COULD NOT BE FIRED. THE BATTERY WAS REINSERTED, BUT THE ISSUE DID NOT IMPROVE. THE DEVICE WAS USED FOR THE ANASTOMOSIS BETWEEN THE ESOPHAGUS AND JEJUNUM. THE ANVIL IN QUESTION WAS PLACED AS IT IS, AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS NOT CONVERTED BECAUSE IT WAS OPEN PROCEDURE. THE PATIENT¿S CONDITION IS STABLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. WHEN THE JAWS WERE OPENED, ONLY THE STAPLES AT THE PROXIMAL END SIDE OF THE JAWS WERE FOUND TO BE CAME OUT SLIGHTLY.
14-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. WHEN THE JAWS WERE OPENED, ONLY THE STAPLES AT THE PROXIMAL END SIDE OF THE JAWS WERE FOUND TO BE CAME OUT SLIGHTLY.
14-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA | Expulsion | A010401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. WHEN THE JAWS WERE OPENED, ONLY THE STAPLES AT THE PROXIMAL END SIDE OF THE JAWS WERE FOUND TO BE CAME OUT SLIGHTLY.
14-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. WHEN THE JAWS WERE OPENED, ONLY THE STAPLES AT THE PROXIMAL END SIDE OF THE JAWS WERE FOUND TO BE CAME OUT SLIGHTLY.
14-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. WHEN THE JAWS WERE OPENED, ONLY THE STAPLES AT THE PROXIMAL END SIDE OF THE JAWS WERE FOUND TO BE CAME OUT SLIGHTLY.
13-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | GIA | Difficult or Delayed Positioning | A150203 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE STAPLING THE TISSUE, IT WAS OBSERVED THAT THE FIRING KNOB WAS STIFF AND THE KNIFE DID NOT CUT THE TISSUE SMOOTHLY AND POST-OPERATIVELY THE PATIENT EXPERIENCED A STAPLE LINE LEAK.
13-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | GIA | Failure to Cut | A050702 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE STAPLING THE TISSUE, IT WAS OBSERVED THAT THE FIRING KNOB WAS STIFF AND THE KNIFE DID NOT CUT THE TISSUE SMOOTHLY AND POST-OPERATIVELY THE PATIENT EXPERIENCED A STAPLE LINE LEAK.
13-12-2021 | Injury | covidien lp llc north haven | 05-01-2022 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE STAPLING THE TISSUE, IT WAS OBSERVED THAT THE FIRING KNOB WAS STIFF AND THE KNIFE DID NOT CUT THE TISSUE SMOOTHLY AND POST-OPERATIVELY THE PATIENT EXPERIENCED A STAPLE LINE LEAK.
30-11-2021 | Malfunction | ethicon endo-surgery | 05-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE STAPLER "MISS-FIRED", BUT NO FURTHER INFORMATION WAS PROVIDED.
21-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION FOR COLON CANCER, THE DEVICE WAS UNABLE TO FIRE. THE GREEN BUTTON WAS PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCTECTOMY, WHEN THE RECTUM TISSUE WAS CLAMPED, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND CARTRIDGE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCTECTOMY, WHEN THE RECTUM TISSUE WAS CLAMPED, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND CARTRIDGE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, WHILE THE INTESTINE WAS ANASTOMOSED, THE DEVICE DID NOT FIRE, THE ANVIL CANNOT BE TILTED AFTER FIRING. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, WHILE THE INTESTINE WAS ANASTOMOSED, THE DEVICE DID NOT FIRE, THE ANVIL CANNOT BE TILTED AFTER FIRING. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER RESECTION, WHILE THE BLOOD VESSEL WAS BEING DETACHED, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE SURGEON USED A NEW HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-11-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT LOWER LOBECTOMY, ON LYMPH NODE DISSECTION, THE RELOAD WAS USED TO CLAMP THE TISSUE , BUT IT WAS FOUND TO EXCEED THE APPLICABLE RANGE OF EXCESSIVE FORCE, THEREFORE, IT WAS REPLACED WITH BLACK RELOAD AND FIRING WAS PERFORMED. AFTERWARDS, THE RELOAD THAT HAD NOT BEEN FIRED WAS ATTACHED AGAIN FOR USE AT ANOTHER SITE, BUT THE RELOAD WAS NOT DETECTED. THERE WAS NO PATIENT INJURY.
30-11-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT LOWER LOBECTOMY, ON LYMPH NODE DISSECTION, THE RELOAD WAS USED TO CLAMP THE TISSUE , BUT IT WAS FOUND TO EXCEED THE APPLICABLE RANGE OF EXCESSIVE FORCE, THEREFORE, IT WAS REPLACED WITH BLACK RELOAD AND FIRING WAS PERFORMED. AFTERWARDS, THE RELOAD THAT HAD NOT BEEN FIRED WAS ATTACHED AGAIN FOR USE AT ANOTHER SITE, BUT THE RELOAD WAS NOT DETECTED. THERE WAS NO PATIENT INJURY.
30-11-2021 | Malfunction | celestica de monterrey s.a de c.v | 05-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL LOBECTOMY, WHILE USING THE RELOAD, THE TISSUE RELOAD BURST. THE LOADS WERE RETRIEVED. THE DEVICE WAS REPLACED WITH THE SAME MODEL TO CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-11-2021 | Malfunction | celestica de monterrey s.a de c.v | 05-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL LOBECTOMY, WHILE USING THE RELOAD, THE TISSUE RELOAD BURST. THE LOADS WERE RETRIEVED. THE DEVICE WAS REPLACED WITH THE SAME MODEL TO CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Difficult or Delayed Positioning | A150203 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG LOBECTOMY, WHILE CUTTING THE LUNG TISSUE, SEVERAL STAPLES DID NOT FORM. THE STAPLE LINE WAS ALSO INCOMPLETE CENTRALLY. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG LOBECTOMY, WHILE CUTTING THE LUNG TISSUE, SEVERAL STAPLES DID NOT FORM. THE STAPLE LINE WAS ALSO INCOMPLETE CENTRALLY. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING LAPAROSCOPIC LUNG RESECTION, WHILE CLAMPING THE LUNG LOBES, THE RELOAD COULD NOT BE CLAMPED WHEN PULLING THE GRIP. THE DEVICE DID NOT FIRE - THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE RELOAD. THERE WAS NO PATIENT INJURY.
21-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VERTICAL STAPLING OF THE SECOND BLACK RELOAD ON A LAPAROSCOPIC SLEEVE GASTROPLASTY, THE PLASTIC FILLET COMPONENT FELL INTO PATIENT'S CAVITY. THE CLAMP LINE WAS INTEGRATED BY THE SURGEON TO RESOLVE THE ISSUE. IT WAS NOTED THAT IT WAS NOT NECESSARY TO REPLACE THE RELOAD OR SUTURE CLIP LINE. THERE WAS NO PATIENT INJURY.
10-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LUNG LOBECTOMY, WHILE PERFORMING BRONCHIAL RESECTION, THE GREEN BUTTON WAS PRESSED, BUT THE HANDLE COULD NOT BE SQUEEZED, AND THE DEVICE DID NOT FIRE. ANOTHER HANDLE AND RELOAD WERE OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LUNG LOBECTOMY, WHILE PERFORMING BRONCHIAL RESECTION, THE GREEN BUTTON WAS PRESSED, BUT THE HANDLE COULD NOT BE SQUEEZED, AND THE DEVICE DID NOT FIRE. ANOTHER HANDLE AND RELOAD WERE OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | us surgical puerto rico | 05-01-2022 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRECTOMY, THE JAW OF THE RELOAD DID NOT CLOSE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | covidien lp llc north haven | 05-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, WHILE DETACHING THE BOWEL, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-02-2022 | Malfunction | us surgical puerto rico | 05-01-2022 | EEA | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE ANVIL AND HANDLE WOULD NOT COMBINE WELL. THERE WAS NO PATIENT INVOLVEMENT.
20-12-2021 | Malfunction | ethicon | 05-01-2022 | STAPLER ETHICON | Firing Problem | A0505 | Insufficient Information | Event Description: DURING LAPAROSCOPIC APPENDECTOMY PROCEDURE, THE SURGEON FIRED THE STAPLER AND NOT ALL OF THE STAPLES FIRED. THE STAPLER WAS RELOADED AND FIRED AGAIN. THE STAPLER WORKED CORRECTLY AND THE PROCEDURE CONTINUED. FDA SAFETY REPORT ID# (B)(4).
03-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2022 | ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER | Misfire | A050502 | Insufficient Information | Event Description: ENDOPATH ETS-FLEX 45 FIRED APPROPRIATELY FOR THE FIRST USE, BUT THEN MISFIRED AND THEN UNABLE TO BE RELOADED FOR A SUBSEQUENT USE. FDA SAFETY REPORT ID# (B)(4).
03-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2022 | ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER | Output Problem | A09 | Insufficient Information | Event Description: ENDOPATH ETS-FLEX 45 FIRED APPROPRIATELY FOR THE FIRST USE, BUT THEN MISFIRED AND THEN UNABLE TO BE RELOADED FOR A SUBSEQUENT USE. FDA SAFETY REPORT ID# (B)(4).
30-11-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, OCCURRED IN THE GASTRIC RESECTION DURING LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE CARTRIDGE WAS LOADED, BUT THE JAWS DID NOT CLOSE. WHEN A NEW RELOAD WAS LOADED, JAWS WAS CLOSED WITHOUT ANY PROBLEM AND THE STAPLING COULD BE PERFORMED. NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A THORASCOPIC HYBRID VIDEO-ASSISTED THORACIC SURGERY (VATS), AT THE TIME OF PULMONARY PARENCHYMA RESECTION, WHEN THE DEVICE WAS ATTEMPTED TO BE FIRED, THE TEETH DID NOT ADVANCE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HIGH ANTERIOR RESECTION, THE DEVICE COULD NOT BE FIRED. THE BATTERY FROM ANOTHER DEVICE WAS INSERTED INTO THE DEVICE IN QUESTION, BUT THE ISSUE DID NOT IMPROVE. ANOTHER THIRD DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
08-11-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE DISTAL GASTRECTOMY SURGERY, COULD NOT FIRE WHEN SEVERING GASTRIC TISSUE ON THE 1ST FIRING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT UPON RECEIPT OF THE DEVICE IT WAS NOTICED THAT THE INNER PEEL POUCH WAS SEPARATED ON THE PLASTIC. NO PATIENT INVOLVEMENT.
01-01-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT UPON RECEIPT OF THE DEVICE IT WAS NOTICED THAT THE INNER PEEL POUCH WAS SEPARATED ON THE PLASTIC. NO PATIENT INVOLVEMENT.
01-01-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT UPON RECEIPT OF DEVICE IT WAS NOTICED THAT THE INNER PEEL POUCH WAS SEPARATED ON THE PLASTIC. NO PATIENT INVOLVEMENT.
01-01-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT UPON RECEIPT OF THE DEVICE IT WAS NOTICED THAT THE INNER PEEL POUCH WAS SEPARATED ON THE PLASTIC. NO PATIENT INVOLVEMENT.
19-01-2021 | Injury | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT COMPLICATED WITH DEHISCENCE OF A TERMINALIZED LATERO-LATERAL ILEOCOLIC ANASTOMOSIS PERFORMED WITH NTLC75 PERFORMED ON 19TH NOVEMBER. REMAINING INTACT ANASTOMOSIS. THE DISASSEMBLY AND PREPARATION OF TERMINAL ILEOSTOMY IN (B)(6) WAS CARRIED OUT. SMI LEVEL II ADMISSION.
19-01-2021 | Injury | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | Sepsis | Event Description: IT WAS REPORTED THAT COMPLICATED WITH DEHISCENCE OF A TERMINALIZED LATERO-LATERAL ILEOCOLIC ANASTOMOSIS PERFORMED WITH NTLC75 PERFORMED ON 19TH NOVEMBER. REMAINING INTACT ANASTOMOSIS. THE DISASSEMBLY AND PREPARATION OF TERMINAL ILEOSTOMY IN (B)(6) WAS CARRIED OUT. SMI LEVEL II ADMISSION.
19-01-2021 | Injury | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | Hernia | Event Description: IT WAS REPORTED THAT COMPLICATED WITH DEHISCENCE OF A TERMINALIZED LATERO-LATERAL ILEOCOLIC ANASTOMOSIS PERFORMED WITH NTLC75 PERFORMED ON 19TH NOVEMBER. REMAINING INTACT ANASTOMOSIS. THE DISASSEMBLY AND PREPARATION OF TERMINAL ILEOSTOMY IN (B)(6) WAS CARRIED OUT. SMI LEVEL II ADMISSION.
19-01-2021 | Injury | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | Anxiety | Event Description: IT WAS REPORTED THAT COMPLICATED WITH DEHISCENCE OF A TERMINALIZED LATERO-LATERAL ILEOCOLIC ANASTOMOSIS PERFORMED WITH NTLC75 PERFORMED ON 19TH NOVEMBER. REMAINING INTACT ANASTOMOSIS. THE DISASSEMBLY AND PREPARATION OF TERMINAL ILEOSTOMY IN (B)(6) WAS CARRIED OUT. SMI LEVEL II ADMISSION.
19-01-2021 | Injury | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT COMPLICATED WITH DEHISCENCE OF A TERMINALIZED LATERO-LATERAL ILEOCOLIC ANASTOMOSIS PERFORMED WITH NTLC75 PERFORMED ON 19TH NOVEMBER. REMAINING INTACT ANASTOMOSIS. THE DISASSEMBLY AND PREPARATION OF TERMINAL ILEOSTOMY IN (B)(6) WAS CARRIED OUT. SMI LEVEL II ADMISSION.
19-01-2021 | Injury | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: IT WAS REPORTED THAT COMPLICATED WITH DEHISCENCE OF A TERMINALIZED LATERO-LATERAL ILEOCOLIC ANASTOMOSIS PERFORMED WITH NTLC75 PERFORMED ON 19TH NOVEMBER. REMAINING INTACT ANASTOMOSIS. THE DISASSEMBLY AND PREPARATION OF TERMINAL ILEOSTOMY IN (B)(6) WAS CARRIED OUT. SMI LEVEL II ADMISSION.
19-01-2021 | Injury | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Hernia | Event Description: IT WAS REPORTED THAT COMPLICATED WITH DEHISCENCE OF A TERMINALIZED LATERO-LATERAL ILEOCOLIC ANASTOMOSIS PERFORMED WITH NTLC75 PERFORMED ON 19TH NOVEMBER. REMAINING INTACT ANASTOMOSIS. THE DISASSEMBLY AND PREPARATION OF TERMINAL ILEOSTOMY IN (B)(6) WAS CARRIED OUT. SMI LEVEL II ADMISSION.
19-01-2021 | Injury | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Anxiety | Event Description: IT WAS REPORTED THAT COMPLICATED WITH DEHISCENCE OF A TERMINALIZED LATERO-LATERAL ILEOCOLIC ANASTOMOSIS PERFORMED WITH NTLC75 PERFORMED ON 19TH NOVEMBER. REMAINING INTACT ANASTOMOSIS. THE DISASSEMBLY AND PREPARATION OF TERMINAL ILEOSTOMY IN (B)(6) WAS CARRIED OUT. SMI LEVEL II ADMISSION.
08-12-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ABDOMINAL SURGERY THERE WAS AN ISSUE WITH LOADING THE CARTRIDGE IN ADDITION TO A FIRING ISSUE. THE FIRING ISSUE WAS THAT THE KNIFE GOT STUCK IN THE RELOAD, COULDN¿T BE RETURNED. THERE WERE NO PATIENT CONSEQUENCES OR CHANGE IN POST-OPERATIVE CARE MENTIONED.
08-12-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ABDOMINAL SURGERY THERE WAS AN ISSUE WITH LOADING THE CARTRIDGE IN ADDITION TO A FIRING ISSUE. THE FIRING ISSUE WAS THAT THE KNIFE GOT STUCK IN THE RELOAD, COULDN¿T BE RETURNED. THERE WERE NO PATIENT CONSEQUENCES OR CHANGE IN POST-OPERATIVE CARE MENTIONED.
16-12-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | FLEX 60 ARTICULATING | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A LAPAROSCOPIC HYSTERECTOMY CASE THAT AFTER A FIRING OF THE STAPLER AND IT WAS PREPARED TO BE RELOADED IT FELT ¿STICKY¿. CARTRIDGE WAS RELOADED. ONCE STAPLER WAS INTRODUCED INTO PATIENT, THE CARTRIDGE POPPED OFF THE STAPLER AND WAS RECOVERED. THE SURGEON THEN COMPLETED THE PROCEDURE WITH A NEW STAPLER. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A LAPAROSCOPIC HYSTERECTOMY CASE THAT AFTER A FIRING OF THE STAPLER AND IT WAS PREPARED TO BE RELOADED IT FELT ¿STICKY¿. CARTRIDGE WAS RELOADED. ONCE STAPLER WAS INTRODUCED INTO PATIENT, THE CARTRIDGE POPPED OFF THE STAPLER AND WAS RECOVERED. THE SURGEON THEN COMPLETED THE PROCEDURE WITH A NEW STAPLER. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
28-10-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN INGUINAL HERNIOGRAPHY, WHILE ACTIVATING THE DEVICE, THE TACK DID NOT HOLD AND FELL INTO SURGICAL FIELD. THE SURGEON APPLIED MANUAL SUTURING TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-10-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN INGUINAL HERNIOGRAPHY, WHILE ACTIVATING THE DEVICE, THE TACK DID NOT HOLD AND FELL INTO SURGICAL FIELD. THE SURGEON APPLIED MANUAL SUTURING TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE CRACKING SOUND WAS HEARD THRICE WHEN THE STOMACH IS BEING STAPLED. A NEW DEVICE WAS USED TO CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
15-11-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE CRACKING SOUND WAS HEARD THRICE WHEN THE STOMACH IS BEING STAPLED. A NEW DEVICE WAS USED TO CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY, MALFUNCTION OF STAPLER FORMATION.
15-11-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE STAPLING THE STOMACH, THE STAPLER WAS ABLE TO FIRE, A CRACKING SOUND WAS HEARD THREE TIMES. A COMPETITIVE POWERED STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE STAPLING THE STOMACH, THE STAPLER WAS ABLE TO FIRE, A CRACKING SOUND WAS HEARD THREE TIMES. A COMPETITIVE POWERED STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-01-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT FAIL TO FIRE DURING AN ANTERIOR RESECTION. THERE WERE NO PATIENT CONSEQUENCES.
22-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING LAPAROSCOPIC SLEEVE GASTRECTOMY WHILE TRANSECTING THE GASTRIC, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING LAPAROSCOPIC SLEEVE GASTRECTOMY WHILE TRANSECTING THE GASTRIC, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE.
29-11-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE ECHELON FLEX STAPLER WITH BLUE (1) AND GOLD (1) CHARGE FAILED TO STAPLE, INTERRUPTING THE STAPLE LINE AFTER THE FIRST SHOT. IT WAS NECESSARY TO ACTIVATE THE FAULT BUTTON AND CHANGE THE LOAD ON TWO OCCASIONS. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-11-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RADICAL RESECTION OF PANCREATIC CANCER SURGERY, COULD NOT FIRE WHEN SEVERING PANCREATIC TISSUE ON THE "2TH" FIRING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-11-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BILLROTH 2 PROCEDURE, WHILE CUTTING THE STOMACH BODY, THE RELOADS HAD AN UNKNOWN ISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-11-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BILLROTH 2 PROCEDURE, WHILE CUTTING THE STOMACH BODY, THE RELOADS HAD AN UNKNOWN ISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | PREMIUM PLUS CEEA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN TOTAL GASTRECTOMY, WHILE PERFORMING THE END-TO-SIDE ESOPHAGOJEJUNOSTOMY, THE DEVICE STAPLED COMPLETELY BUT THE BLADE DID NOT CUT. TO RESOLVE THE ISSUE, THE SURGEON DID A MANUAL RESECTION. THERE WAS A NO PATIENT INJURY.
22-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TOT HE REPORTER DURING LAPAROSCOPIC BARIATRIC SURGERY, WHILE IN THE STOMACH, THE DEVICE DID NOT FIRE. ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED. TO RESOLVE THE ISSUE, ANOTHER CARTRIDGE WAS USED AND IT WORKED WELL. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ENDO GIA | Component or Accessory Incompatibility | A1701 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION FOR A LAPAROSCOPIC PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE. THE SURGEON WAS NOT ABLE TO PRESSED THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE, HENCE THE DEVICE DID NOT FIRE. TO COMPLETE THE CASE, ANOTHER RELOAD WAS USED.
22-12-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ENDO GIA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION FOR A LAPAROSCOPIC PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE. THE SURGEON WAS NOT ABLE TO PRESSED THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE, HENCE THE DEVICE DID NOT FIRE. TO COMPLETE THE CASE, ANOTHER RELOAD WAS USED.
22-12-2021 | Injury | us surgical puerto rico | 06-01-2022 | EEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY (SIGMOID COLON), ON ANASTOMOSIS, THE ANVIL COULD NOT BE REMOVED WITH THE HANDLE AND REMAINED IN THE BODY. AN ADDITIONAL RESECTION AND ANASTOMOSIS WAS PERFORMED TO REMOVE THE ANVIL. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. SURGICAL INCISION WAS ALSO EXTENDED BY 1 INCH.
21-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | TA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RADICAL SURGERY FOR ESOPHAGEAL CANCER, WHILE IN THE SEVERED BLOOD VESSEL, THE DEVICE PARTIALLY FIRED. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | TA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN RADICAL SURGERY FOR ESOPHAGEAL CANCER, WHILE DETACHING BLOOD VESSEL, THE DEVICE PARTIALLY FIRED. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A LAPAROSCOPY, THE JAWS DID NOT OPEN BEFORE CLAMPING ON THE TISSUE. ANOTHER CARTRIDGE WAS USED TO RESOLVE THE CONCERN.
30-11-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | EEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, THE CARTRIDGE FACE WAS SEPARATED. THERE WAS NO PATIENT INVOLVEMENT.
22-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING RECONSTRUCTION OF THE LARGE INTESTINE, THE FIRST FIRING WITH THE HANDLE AND RELOAD WAS USED WITHOUT ANY PROBLEM HOWEVER ON THE SECOND FIRING, THE HANDLE DID NOT ADVANCED WHEN THE RELOAD WAS REPLACED WITH A NEW ONE. IN THE THIRD FIRING IT WAS ALSO USED WITHOUT ANY PROBLEM. THE HANDLE DID NOT ADVANCE AGAIN IN THE 4TH FIRING OF THE LAST FIRING. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE.
22-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING RECONSTRUCTION OF THE LARGE INTESTINE, THE FIRST FIRING WITH THE HANDLE AND RELOAD WAS USED WITHOUT ANY PROBLEM HOWEVER ON THE SECOND FIRING, THE HANDLE DID NOT ADVANCED WHEN THE RELOAD WAS REPLACED WITH A NEW ONE. IN THE THIRD FIRING IT WAS ALSO USED WITHOUT ANY PROBLEM. THE HANDLE DID NOT ADVANCE AGAIN IN THE 4TH FIRING OF THE LAST FIRING. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE.
22-12-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING RECONSTRUCTION OF THE LARGE INTESTINE, THE FIRST FIRING WITH THE HANDLE AND RELOAD WAS USED WITHOUT ANY PROBLEM HOWEVER ON THE SECOND FIRING, THE HANDLE DID NOT ADVANCED WHEN THE RELOAD WAS REPLACED WITH A NEW ONE. IN THE THIRD FIRING IT WAS ALSO USED WITHOUT ANY PROBLEM. THE HANDLE DID NOT ADVANCE AGAIN IN THE 4TH FIRING OF THE LAST FIRING. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE.
22-12-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING RECONSTRUCTION OF THE LARGE INTESTINE, THE FIRST FIRING WITH THE HANDLE AND RELOAD WAS USED WITHOUT ANY PROBLEM HOWEVER ON THE SECOND FIRING, THE HANDLE DID NOT ADVANCED WHEN THE RELOAD WAS REPLACED WITH A NEW ONE. IN THE THIRD FIRING IT WAS ALSO USED WITHOUT ANY PROBLEM. THE HANDLE DID NOT ADVANCE AGAIN IN THE 4TH FIRING OF THE LAST FIRING. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE.
17-03-2020 | Malfunction | ethicon endo-surgery | 06-01-2022 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-09-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC BODY AND TAIL RESECTION PROCEDURE, WHILE DETACHING THE TISSUE AND BLOOD VESSELS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-07-2021 | Injury | covidien lp llc north haven | 06-01-2022 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT LIVER RESECTION FOR COLORECTAL LIVER METASTASES (CRLM) BETWEEN 2012 AND 2019. THE AIM OF THE STUDY WAS TO COMPARE THE POSTOPERATIVE OUTCOMES AND LONG-TERM SURVIVALS OF PATIENTS UNDERGOING PARENCHYMAL-SPARING RESECTION (PSR) WITH THOSE PATIENTS UNDERGOING NON-PSR IN THE ENTIRE COHORT OF CONSECUTIVE PATIENTS TREATED IN THE CENTER CRLM AFTER STRATIFYING THE PATIENTS ACCORDING TO THE BURDEN THE TUMOR BURDEN SCORE. PARENCHYMAL TRANSECTIONS WERE PERFORMED WITH THE STAPLER COMBINED WITH OTHER DEVICES. THERE WERE 153 PATIENTS IN THE STUDY AND COMPLICATIONS INCLUDED: POSTOPERATIVE BLEEDING. BLOOD TRANSFUSIONS WERE REQUIRED. NO OTHER INTERVENTIONS WERE REPORTED.
10-10-2020 | Malfunction | ethicon endo-surgery | 06-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Excessive Heating | A1002 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, OPENED THE PACKING, BEFORE USED ON THE PATIENT, NOTED THE BATTERY WAS WITH ABNORMAL HIGH TEMPERATURE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
02-12-2021 | Injury | ethicon endo-surgery | 06-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Cut | A050702 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE PLEE60A & GST60D WERE TAKEN OUT FROM STERILE PACK AND 6 RELOADS WERE FIRED IN MGB. DURING FIRING DOC EXPERIENCE THE RUGGED FIRING AND FIRED SURFACES ARE IN ZIG ZAG FASHION. DR. FELT IT'S DUE TO PLEE60A AND REQUESTED FOR REPLACEMENT WITH NEW ONE. DR. IS KEEN TO KNOW THE REASON FOR LEAK FROM STAPLED EDGES AND ANALYSIS OF COMPLAINT SAMPLE.
02-12-2021 | Injury | ethicon endo-surgery | 06-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE PLEE60A & GST60D WERE TAKEN OUT FROM STERILE PACK AND 6 RELOADS WERE FIRED IN MGB. DURING FIRING DOC EXPERIENCE THE RUGGED FIRING AND FIRED SURFACES ARE IN ZIG ZAG FASHION. DR. FELT IT'S DUE TO PLEE60A AND REQUESTED FOR REPLACEMENT WITH NEW ONE. DR. IS KEEN TO KNOW THE REASON FOR LEAK FROM STAPLED EDGES AND ANALYSIS OF COMPLAINT SAMPLE.
02-12-2021 | Injury | ethicon endo-surgery | 06-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Patient Device Interaction Problem | A01 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE PLEE60A & GST60D WERE TAKEN OUT FROM STERILE PACK AND 6 RELOADS WERE FIRED IN MGB. DURING FIRING DOC EXPERIENCE THE RUGGED FIRING AND FIRED SURFACES ARE IN ZIG ZAG FASHION. DR. FELT IT'S DUE TO PLEE60A AND REQUESTED FOR REPLACEMENT WITH NEW ONE. DR. IS KEEN TO KNOW THE REASON FOR LEAK FROM STAPLED EDGES AND ANALYSIS OF COMPLAINT SAMPLE.
09-07-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-09-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, THE STAPLES FORMED POORLY AFTER DETACHING THE LUNG PARENCHYMA, TISSUES, ARTERIES, AND VEINS OF THE PULMONARY LOBE. THIS HAD CAUSED BLEEDING, BUT ANOTHER RELOAD OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF COLON CANCER, WHILE DETACHING THE BLOOD VESSEL, THE DEVICE HAD POOR STAPLE FORMATION. SOME STAPLES DID NOT FORM A B SHAPE. THE SURGEON REPLACED THE RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-07-2021 | Malfunction | ethicon endo-surgery | 06-01-2022 | GST BLACK RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANTERIOR RESECTION PROCEDURE, WHILE CUTTING COLON, THE DEVICE DID NOT FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS CHANGED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANTERIOR RESECTION PROCEDURE, WHILE CUTTING COLON, THE DEVICE DID NOT FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS CHANGED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2021 | Malfunction | covidien lp llc north haven | 06-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE.
17-12-2021 | Malfunction | us surgical puerto rico | 06-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE.
08-12-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC PULMONARY RESECTION, SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HANDLE COULD NOT BE SQUEEZED. DEVICE WAS NOT ABLE TO FIRE. SURGEON USED ANOTHER RELOAD FOR THE SAME HANDLE FUNCTIONED WITH OTHER RELOADS. NO PATIENT INJURY. PLI20: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK AND WAS CONFIRMED IN THE 1ST AND 2ND ON THE FIRING RACK.
14-12-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, AT PULMONARY PARENCHYMA RESECTION, THE DEVICE WAS UNABLE TO FIRE - GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. THE JAWS LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, AT PULMONARY PARENCHYMA RESECTION, THE DEVICE WAS UNABLE TO FIRE - GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. THE JAWS LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 07-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING THE INSTRUMENT FOR A LAPAROSCOPIC WEDGE RESECTION, THE FIRST RELOAD COULD BE USED NORMALLY, BUT THIS RESULTED IN ONE "CREAKY" SOUND. WHEN TRYING TO FIRE THE SECOND RELOAD, THERE WAS A LOUD BANG AND FROM THAT MOMENT THE HANDLE WAS NO LONGER FUNCTIONAL. THE MAGAZINES COULD OPENED AND CLOSED, BUT NO LONGER FIRED EVEN IF THE GREEN BUTTON AND HANDLE COULD BE SQUEEZED BY THE SURGEON. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 07-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING THE INSTRUMENT FOR A LAPAROSCOPIC WEDGE RESECTION, THE FIRST RELOAD COULD BE USED NORMALLY, BUT THIS RESULTED IN ONE "CREAKY" SOUND. WHEN TRYING TO FIRE THE SECOND RELOAD, THERE WAS A LOUD BANG AND FROM THAT MOMENT THE HANDLE WAS NO LONGER FUNCTIONAL. THE MAGAZINES COULD OPENED AND CLOSED, BUT NO LONGER FIRED EVEN IF THE GREEN BUTTON AND HANDLE COULD BE SQUEEZED BY THE SURGEON. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
20-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
20-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | WHITE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
20-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
20-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | WHITE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
20-03-2020 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, SURGEON HAS TAKEN THE STERILE PIECE OF TCR55 AND FOLLOWED ALL NECESSARY FIRING SEQUENCE, BUT STAPLER DID NOT GET FIRED PROPERLY AND HE TOOK ANOTHER NEW STERILE UNIT TO COMPLETE THE OPERATION. THERE WAS NO DELAY TO THE SURGERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY.
01-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, SURGEON HAS TAKEN THE STERILE PIECE OF TCR55 AND FOLLOWED ALL NECESSARY FIRING SEQUENCE, BUT STAPLER DID NOT GET FIRED PROPERLY AND HE TOOK ANOTHER NEW STERILE UNIT TO COMPLETE THE OPERATION. THERE WAS NO DELAY TO THE SURGERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY.
02-07-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-09-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | PROXIMATE*75MM LINEAR CUTTER | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLES WERE EXPELLED PRIOR TO FIRING. THE DEVICE WAS RELOADED AND IT OCCURRED AGAIN. THERE WERE NO PATIENT CONSEQUENCES.
19-01-2020 | Injury | us surgical puerto rico | 07-01-2022 | TA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LUNG RESECTION, ON THE VESSEL, THE SURGEON NOTED A CLICK WAS HEARD BUT THE DEVICE NEVER FIRED. THE DOCTOR SUTURED THE VESSEL TO RESOLVE THE ISSUE. IT WAS NOTED THAT WHEN THE PRODUCT WAS TAKEN BACK TO THE SCRUB TABLE THE USER WAS ABLE TO FIRE.
13-07-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2021 | Injury | covidien lp llc north haven | 07-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOIDECTOMY, WHILE ON RESECTION OF SIGMOID. THE FIRING STOPPED AT MIDWAY, BUT THE KNIFE BAR AUTOMATICALLY RETURNED. TO FIX THE ISSUE, ONE ADDITIONAL CARTRIDGE WAS USED FOR RESECTING THE COLON.
06-07-2021 | Death | covidien lp llc north haven | 07-01-2022 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT LIVER RESECTION FOR COLORECTAL LIVER METASTASES (CRLM) BETWEEN 2012 AND 2019. THE AIM OF THE STUDY WAS TO COMPARE THE POSTOPERATIVE OUTCOMES AND LONG-TERM SURVIVALS OF PATIENTS UNDERGOING PARENCHYMAL-SPARING RESECTION (PSR) WITH THOSE PATIENTS UNDERGOING NON-PSR IN THE ENTIRE COHORT OF CONSECUTIVE PATIENTS TREATED IN THE CENTER CRLM AFTER STRATIFYING THE PATIENTS ACCORDING TO THE BURDEN THE TUMOR BURDEN SCORE. PARENCHYMAL TRANSECTIONS WERE PERFORMED WITH THE STAPLER COMBINED WITH OTHER DEVICES. THERE WERE 153 PATIENTS IN THE STUDY AND COMPLICATIONS INCLUDED: THERE WERE 2 DEATHS WITHIN 90 DAYS, ONE DUE TO AN UNSPECIFIED CAUSE.
13-07-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-03-2020 | Malfunction | ethicon endo-surgery | 07-01-2022 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. PROCEDURE: LAP COLON PROCEDURE.
27-10-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. PROCEDURE: ESOPHAGEAL
28-09-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
27-09-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE DEVICE WAS CONNECTED TO A COMPUTER WHEN AN ATTEMPT WAS MADE TO UPGRADE THE POWER HANDLE, WHICH HAD BEEN CHARGED. THE SCREEN BLACKED OUT DUE TO DISCONNECTION DURING UPDATE OPERATION AND UPDATE WAS NOT COMPLETED. WHEN IT WAS PLUGGED TO THE CHARGER AGAIN, THE BATTERY CHARGER FLASHED RED, MAKING CHARGING IMPOSSIBLE. A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
27-09-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | SIGNIA | Application Program Version or Upgrade Problem | A110206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE DEVICE WAS CONNECTED TO A COMPUTER WHEN AN ATTEMPT WAS MADE TO UPGRADE THE POWER HANDLE, WHICH HAD BEEN CHARGED. THE SCREEN BLACKED OUT DUE TO DISCONNECTION DURING UPDATE OPERATION AND UPDATE WAS NOT COMPLETED. WHEN IT WAS PLUGGED TO THE CHARGER AGAIN, THE BATTERY CHARGER FLASHED RED, MAKING CHARGING IMPOSSIBLE. A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
27-09-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE DEVICE WAS CONNECTED TO A COMPUTER WHEN AN ATTEMPT WAS MADE TO UPGRADE THE POWER HANDLE, WHICH HAD BEEN CHARGED. THE SCREEN BLACKED OUT DUE TO DISCONNECTION DURING UPDATE OPERATION AND UPDATE WAS NOT COMPLETED. WHEN IT WAS PLUGGED TO THE CHARGER AGAIN, THE BATTERY CHARGER FLASHED RED, MAKING CHARGING IMPOSSIBLE. A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
16-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE REVERSAL OR A HARTMANN PROCEDURE (OPEN) THE SURGEON ATTACHED THE ANVIL TO THE TROCAR RIGHT ON THE DEVICE AS NORMAL. THE ASSISTING CONSULTANT HELPED AT THE BOTTOM END TO DIAL DOWN AND START CLOSING THE GUN. AS THE CONSULTANT CLOSED THE GUN AND ATTEMPTED TO FIRE IT, THE GUN DID NOT FIRE. A WHICH POINT WE REALIZED THAT THE ANVIL WAS NOT ATTACHED. WE TRIED AGAIN TO ATTACH THE ANVIL AND LISTENED FOR THE CLICK AND ENSURED THE ORANGE MARKER WAS NOT LONGER THERE AND TRIED AGAIN. UNFORTUNATELY IT CAME OFF AGAIN AS THE SURGEON CLOSED THE GONE. FINALLY THE CONSULTANT HELD THE ANVIL N PLACE AND FIRED THE GUN AND THIS TIME IT WORKED. CONSULTANT EXPERIENCE: USED CDH FOR 30 YEARS.
16-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE REVERSAL OR A HARTMANN PROCEDURE (OPEN) THE SURGEON ATTACHED THE ANVIL TO THE TROCAR RIGHT ON THE DEVICE AS NORMAL. THE ASSISTING CONSULTANT HELPED AT THE BOTTOM END TO DIAL DOWN AND START CLOSING THE GUN. AS THE CONSULTANT CLOSED THE GUN AND ATTEMPTED TO FIRE IT, THE GUN DID NOT FIRE. A WHICH POINT WE REALIZED THAT THE ANVIL WAS NOT ATTACHED. WE TRIED AGAIN TO ATTACH THE ANVIL AND LISTENED FOR THE CLICK AND ENSURED THE ORANGE MARKER WAS NOT LONGER THERE AND TRIED AGAIN. UNFORTUNATELY IT CAME OFF AGAIN AS THE SURGEON CLOSED THE GONE. FINALLY THE CONSULTANT HELD THE ANVIL N PLACE AND FIRED THE GUN AND THIS TIME IT WORKED. CONSULTANT EXPERIENCE: USED CDH FOR 30 YEARS.
02-06-2020 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-11-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. PROCEDURE: ESOPHAGEAL.
15-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, WHEN SEVERING SHORT GASTRIC VESSELS TISSUE ON THE SECOND FIRING, AFTER FIRED, NOTED SOME STAPLES MALFORMED . CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
02-07-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Eject | A050301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED FAILURE TO OPEN THE CDH29P WHEN PERFORMING LAPAROSCOPIC COLORECTAL ANASTOMOSIS. THEY COULD NOT REMOVE THE MACHINE AND THE ANASTOMOSIS OPENED. THERE WERE NO PATIENT CONSEQUENCES.
08-09-2020 | Malfunction | ethicon endo-surgery | 07-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | ILS 33MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE CDH STAPLER CUT AND NOT STAPLED. THERE WERE NO PATIENT CONSEQUENCES.
23-09-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, DURING CHARGING, HANDLE WAS NOT LOADING WITH CHARGER AND CHARGER WAS IN RED LIGHT BAY. THERE WAS NO PATIENT INVOLVEMENT.
23-09-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, DURING CHARGING, HANDLE WAS NOT LOADING WITH CHARGER AND CHARGER WAS IN RED LIGHT BAY. THERE WAS NO PATIENT INVOLVEMENT.
17-01-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESECTION LEFT COLON, THE POWERED CIRCULAR WOULD NOT STAPLE WHEN THE SAFETY WAS RELEASED. IT LIT UP BUT DID NOT FIRE. DEVICE DID NOT FIRE OR STAPLE. A SECOND BATTERY WAS ATTEMPTED WITH NO CHANGE. CASE COMPLETED WITH ANOTHER STAPLER. NO PATIENT HARM.
22-12-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBECTOMY, AFTER THE THORAX WAS BROKEN, AFTER DETACHING INTERLOBULAR FISSURE UNDER THE THORACOSCOPE, THE BLACK STAPLE COULD NOT BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. REPLACED THE HANDLE, BUT STILL COULD NOT FIRE. FINALLY, A NEW BLACK STAPLE WAS REPLACED, AND THE FIRING WAS SUCCESSFUL WITH THE REPLACED HANDLE. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | us surgical puerto rico | 07-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBECTOMY, AFTER THE THORAX WAS BROKEN, AFTER DETACHING INTERLOBULAR FISSURE UNDER THE THORACOSCOPE, THE BLACK STAPLE COULD NOT BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. REPLACED THE HANDLE, BUT STILL COULD NOT FIRE. FINALLY, A NEW BLACK STAPLE WAS REPLACED, AND THE FIRING WAS SUCCESSFUL WITH THE REPLACED HANDLE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-09-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-08-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST WHITE RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, WHEN SEVERING LUNG TISSUE ON THE FIRST FIRING, AFTER FIRED, NOTED SOME STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GREEN RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, WHEN SEVERING LUNG TISSUE ON THE FIRST FIRING, AFTER FIRED, NOTED SOME STAPLES MALFORMED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2022 | Malfunction | covidien lp llc north haven | 07-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY SURGERY, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE BUT THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | us surgical puerto rico | 07-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY SURGERY, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE BUT THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPY, WHEN ATTEMPTING TO FIRE TO CREATE UPPER RIGHT MIDDLE LOBE, GREEN BUTTON COULD NOT BE PRESSED. A NEW RELOAD FROM ANOTHER LOT WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-09-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-09-2022 | Injury | us surgical puerto rico | 07-01-2022 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF A HEMICOLECTOMY, THE PATIENT HAD A BOWEL LEAK.
16-09-2022 | Injury | us surgical puerto rico | 07-01-2022 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF A HEMICOLECTOMY, THE PATIENT HAD A BOWEL LEAK.
05-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2021 | Malfunction | us surgical puerto rico | 07-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-09-2022 | Injury | us surgical puerto rico | 07-01-2022 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE WAS USED ON THE PATIENT HOWEVER THE SURGEON FELT THAT THE STAPLE LINE WAS NOT ADEQUATE. THE SURGEON CUT THE STAPLES OUT DURING THE SURGERY, AND PROVIDED THE PATIENT WITH AN OSTOMY INSTEAD OF AN ANASTOMOSIS.
23-12-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-09-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | SIGNIA | Device Contamination with Chemical or Other Material | A180104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE PRODUCT WAS OPENED, A FOREIGN MATERIAL WAS FOUND TO BE INSIDE. IT WAS REPLACED WITH A NEW ONE. THERE WAS NO PATIENT INVOLVEMENT.
16-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PSEE60A DEVICE WAS INACTIVE, THE BATTERY SEEMED TO BE DEAD DURING THE 9TH CARTRIDGE FIRING. HCP A. R. WAS DOING MINI-BYPASS SURGERY. THE KNIFE SEEMED TO BE STILL IN PLACE, NOT FIRED. THE FIRING BUTTON WAS INACTIVE. 8 CARTRIDGES WERE FIRED SUCCESSFULLY WITH THE SAME DEVICE. THE BATTERY WAS VERY HOT AFTER 8 FIRINGS. SURGERY DELAYED 10 MINUTES DUE TO THE EVENT. A NEW DEVICE WAS USED TO SUCCESSFULLY COMPLETELY THE CASE. NO PATIENT CONSEQUENCES NOTED.
16-12-2021 | Malfunction | ethicon endo-surgery | 07-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Excessive Heating | A1002 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PSEE60A DEVICE WAS INACTIVE, THE BATTERY SEEMED TO BE DEAD DURING THE 9TH CARTRIDGE FIRING. HCP A. R. WAS DOING MINI-BYPASS SURGERY. THE KNIFE SEEMED TO BE STILL IN PLACE, NOT FIRED. THE FIRING BUTTON WAS INACTIVE. 8 CARTRIDGES WERE FIRED SUCCESSFULLY WITH THE SAME DEVICE. THE BATTERY WAS VERY HOT AFTER 8 FIRINGS. SURGERY DELAYED 10 MINUTES DUE TO THE EVENT. A NEW DEVICE WAS USED TO SUCCESSFULLY COMPLETELY THE CASE. NO PATIENT CONSEQUENCES NOTED.
14-12-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC DISTAL GASTRECTOMY, WHILE PERFORMING GASTROSTOMY, A STRONG RESISTANCE WAS FELT WHEN CLOSING AND FIRING THE STAPLER. A STRANGE NOISE WAS ALSO HEARD AFTER TRYING TO FIRE, AND THE DEVICE DID NOT MOVE FORWARD AFTER FIRING HALFWAY. THE GEAR INSIDE THE HANDLE WAS FOUND MISSING. ANOTHER HANDLE OF THE SAME TYPE AND NEW RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | us surgical puerto rico | 07-01-2022 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING STOMACH BODY DURING A LAPAROSCOPIC STOMACH PROCEDURE, WHEN THE NEW RELOAD WAS LOADED ON THE HANDLE THE JAWS WOULD NOT CLOSE COMPLETELY. IT WAS IN AN OPEN STATE. THE HANDLE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | us surgical puerto rico | 07-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE LIVER RESECTION, WHILE CLOSING THE LEFT LIVER PEDICLE. THE DEVICE DID NOT FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. TO COMPLETE THE CASE, THE SURGEON REPLACED THE DEVICE WITH ANOTHER MANUFACTURER'S PRODUCT. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | covidien lp llc north haven | 07-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE LIVER RESECTION, WHILE CLOSING THE LEFT LIVER PEDICLE. THE DEVICE DID NOT FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. TO COMPLETE THE CASE, THE SURGEON REPLACED THE DEVICE WITH ANOTHER MANUFACTURER'S PRODUCT. THERE WAS NO PATIENT INJURY.
22-12-2021 | Injury | covidien lp llc north haven | 07-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Abscess | Event Description: ACCORDING TO THE REPORTER, SEVEN DAYS AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT DEVELOPED NECROSIS THAT RESULTED TO LEAK. THE STAPLE FORMATION DURING THE SURGERY WAS POOR. AS A RESULT, THERE WAS TISSUE DAMAGE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED. TO RESOLVE THE ISSUE, THE PATIENT HAD TO BE RE-OPERATED, THE HOLE IN STOMACH WAS CLOSED VIA SUTURE AND STENT WAS IMPLANTED.
22-12-2021 | Injury | covidien lp llc north haven | 07-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Necrosis | Event Description: ACCORDING TO THE REPORTER, SEVEN DAYS AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT DEVELOPED NECROSIS THAT RESULTED TO LEAK. THE STAPLE FORMATION DURING THE SURGERY WAS POOR. AS A RESULT, THERE WAS TISSUE DAMAGE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED. TO RESOLVE THE ISSUE, THE PATIENT HAD TO BE RE-OPERATED, THE HOLE IN STOMACH WAS CLOSED VIA SUTURE AND STENT WAS IMPLANTED.
22-12-2021 | Injury | covidien lp llc north haven | 07-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, SEVEN DAYS AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT DEVELOPED NECROSIS THAT RESULTED TO LEAK. THE STAPLE FORMATION DURING THE SURGERY WAS POOR. AS A RESULT, THERE WAS TISSUE DAMAGE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED. TO RESOLVE THE ISSUE, THE PATIENT HAD TO BE RE-OPERATED, THE HOLE IN STOMACH WAS CLOSED VIA SUTURE AND STENT WAS IMPLANTED.
22-12-2021 | Injury | covidien lp llc north haven | 07-01-2022 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE REPORTER, SEVEN DAYS AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT DEVELOPED NECROSIS THAT RESULTED TO LEAK. THE STAPLE FORMATION DURING THE SURGERY WAS POOR. AS A RESULT, THERE WAS TISSUE DAMAGE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED. TO RESOLVE THE ISSUE, THE PATIENT HAD TO BE RE-OPERATED, THE HOLE IN STOMACH WAS CLOSED VIA SUTURE AND STENT WAS IMPLANTED.
22-12-2021 | Injury | covidien lp llc north haven | 07-01-2022 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: ACCORDING TO THE REPORTER, SEVEN DAYS AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT DEVELOPED NECROSIS THAT RESULTED TO LEAK. THE STAPLE FORMATION DURING THE SURGERY WAS POOR. AS A RESULT, THERE WAS TISSUE DAMAGE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED. TO RESOLVE THE ISSUE, THE PATIENT HAD TO BE RE-OPERATED, THE HOLE IN STOMACH WAS CLOSED VIA SUTURE AND STENT WAS IMPLANTED.
22-12-2021 | Injury | covidien lp llc north haven | 07-01-2022 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, SEVEN DAYS AFTER A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT DEVELOPED NECROSIS THAT RESULTED TO LEAK. THE STAPLE FORMATION DURING THE SURGERY WAS POOR. AS A RESULT, THERE WAS TISSUE DAMAGE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED. TO RESOLVE THE ISSUE, THE PATIENT HAD TO BE RE-OPERATED, THE HOLE IN STOMACH WAS CLOSED VIA SUTURE AND STENT WAS IMPLANTED.
17-11-2021 | Malfunction | ethicon | 09-01-2022 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Loss of or Failure to Bond | A040102 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN INGUINAL HERNIA REPAIR PROCEDURE ON (B)(6) 2021 AND THE ABSORBABLE STRAPS WERE USED. IT WAS REPORTED THAT THE DURING THE PROCEDURE, AFTER PLACING THE MESH AND STARTING THE FIXATION WITH THE STRAPS, THE EQUIPMENT DID NOT FIX THE MESH, CAUSED BLEEDING IN THE CAVITY AND LOCKED AND. IT WAS NECESSARY TO SUTURE THE MESH WITH A SUTURE THREAD. IT WAS REPORTED THAT THE PRODUCT WAS STORED IN ACCORDANCE WITH THE MANUFACTURER'S INSTRUCTIONS. IT WAS ALSO REPORTED THAT THE PATIENT HAD NO PERMANENT DAMAGE AND DID NOT NEED TO BE HOSPITALIZED. ALSO, THE PATIENT DID NOT NEED TO BE RE-OPERATED TO PREVENT OR CORRECT ANY DAMAGE, AND THERE WAS NO SERIOUS DAMAGE. ADDITIONAL INFORMATION WAS REQUESTED.
11-11-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE INSTRUMENT WAS CLOSED FOR INSERTION. AFTER INSERTION IT COULD ONLY BE OPENED AGAIN AFTER MULTIPLE ATTEMPTS. SURGERY WAS COMPLETED WITH ANOTHER INSTRUMENT JUST IN CASE. NO ADVERSE PATIENT HARM.
29-09-2020 | Malfunction | ethicon endo-surgery | 10-01-2022 | PROXIMATE*PPH PROCEDURE SET | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE HEMORRHOIDECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. THE GAP BETWEEN THE NAILS, ANASTOMY SLIGHTLY BLEEDING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-09-2020 | Malfunction | ethicon endo-surgery | 10-01-2022 | PROXIMATE*PPH PROCEDURE SET | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE HEMORRHOIDECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. THE GAP BETWEEN THE NAILS, ANASTOMY SLIGHTLY BLEEDING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-06-2020 | Malfunction | ethicon endo-surgery | 10-01-2022 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCTRECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-11-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE PROCEDURE AND WHEN IT WAS TIME TO FIRE THE ECHELON GST, NOTHING HAPPENED, LIKE IF THE BATTERY HAD NO LIFE. THEY TRIED REMOVING IT, REINTRODUCE IT, TEST IT, NOTHING. THEY OPENED ANOTHER PSEE60A AND WORKED FINE, FINISHED THE CASE.
23-11-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING OPEN PANCREATICODUODENAL SURGERY, COULD NOT FIRE WHEN SEVERING SMALL INTESTINE TISSUE ON THE FIRST FIRING. CHECKED THE SLED¿S POSITION, IT IS LOCKOUT ISSUE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE DEVICE COULD NOT BE FIRED. THE INDICATOR WAS WITHIN THE GREEN AREA AND THE TISSUE COMPRESSION SCALE BACKLIGHT WAS ILLUMINATED. THE ANVIL WAS REATTACHED TO THE TROCAR, BUT THE ISSUE DID NOT IMPROVE. THE GREEN CHECK MARK WAS ILLUMINATED. THE DEVICE WAS USED FOR THE ANASTOMOSIS BETWEEN THE COLON AND THE RECTUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-09-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER ,DURING CHARGING, THE HANDLE LOOKED COMPLETELY NORMAL IN THE CHARGING STATION, THE GREEN LED WAS FLASHING. BUT WOULD NOT WAKE UP EVEN WITH A RESET. WHEN CHARGED TO A DIFFERENT CHARGER AT THE SALES REPRESENTATIVE HOME, THE GREEN LED DOES NOT START FLASHING, ONLY THE BLUE ONE IS LIT. TO RESOLVE THE ISSUE, A NEW HANDLE AND CHARGER WAS USED. THERE WAS NO PATIENT INVOLVED.
22-09-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER ,DURING CHARGING, THE HANDLE LOOKED COMPLETELY NORMAL IN THE CHARGING STATION, THE GREEN LED WAS FLASHING. BUT WOULD NOT WAKE UP EVEN WITH A RESET. WHEN CHARGED TO A DIFFERENT CHARGER AT THE SALES REPRESENTATIVE HOME, THE GREEN LED DOES NOT START FLASHING, ONLY THE BLUE ONE IS LIT. TO RESOLVE THE ISSUE, A NEW HANDLE AND CHARGER WAS USED. THERE WAS NO PATIENT INVOLVED.
16-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED: PROXIMAL RECTUM STUMP TRANSECTED WITH COMPETITOR DEVICE DISTAL RECTAL STUMP WAS TRANSECTED WITH CS40G WITH THE GREATER CURVATURE OF THE JAW FACING DOWNWARDS AFTER CLOSING THE CONTOUR STAPLER ONTO THE COLON, SURGEON WAITED A FEW SECOND BEFORE FIRING ONCE TRANSECTED, SURGEON A WANTED TO OVERSEW THE STAPLE LINE WITH PDS 3-0 JUST AS SURGEON A WAS ABOUT TO SEW OVER THE STAPLE LINE, HE REALIZED THAT THE STAPLE LINE WAS INCOMPLETE AND THE MUCOSA WAS POUTING FROM THE STAPLE LINE SURGEON A USED A FORCEP TO PROBE ON THE POUTED MUCOSA AND THE STAPLE LINE SLOWLY OPENED UP BY ITSELF A SECOND RELOAD WAS LOADED ONTO THE CONTOUR STAPLER, AND FIRED 2CM DISTAL FROM THE FIRST TRANSECTION THIS TIME, SURGEON B AND SURGEON A REALIZED THAT THERE WAS ONLY STAPLE LINE ON THE ANTERIOR SIDE AND NOT ON THE POSTERIOR STAPLE LINE WAS NOT COMPLETE AND DID NOT FULLY CAPTURE THE COLON A THIRD RELOAD WAS LOADED AND FIRED AGAIN 3-4CM DISTAL FROM THE SECOND TRANSECTION. STAPLE LINE WAS FULLY FORMED AND COLON TRANSECTED PDS 3-0 WAS SEWED OVER THE STAPLE LINE EVERYTHING WAS OK SURGERY WAS DELAYED BY 15 MINUTES THE SURGERY WAS COMPLETED SUCCESSFULLY.
16-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | CONTOUR CURVED CUTTER STAPLER | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED: PROXIMAL RECTUM STUMP TRANSECTED WITH COMPETITOR DEVICE DISTAL RECTAL STUMP WAS TRANSECTED WITH CS40G WITH THE GREATER CURVATURE OF THE JAW FACING DOWNWARDS AFTER CLOSING THE CONTOUR STAPLER ONTO THE COLON, SURGEON WAITED A FEW SECOND BEFORE FIRING ONCE TRANSECTED, SURGEON A WANTED TO OVERSEW THE STAPLE LINE WITH PDS 3-0 JUST AS SURGEON A WAS ABOUT TO SEW OVER THE STAPLE LINE, HE REALIZED THAT THE STAPLE LINE WAS INCOMPLETE AND THE MUCOSA WAS POUTING FROM THE STAPLE LINE SURGEON A USED A FORCEP TO PROBE ON THE POUTED MUCOSA AND THE STAPLE LINE SLOWLY OPENED UP BY ITSELF A SECOND RELOAD WAS LOADED ONTO THE CONTOUR STAPLER, AND FIRED 2CM DISTAL FROM THE FIRST TRANSECTION THIS TIME, SURGEON B AND SURGEON A REALIZED THAT THERE WAS ONLY STAPLE LINE ON THE ANTERIOR SIDE AND NOT ON THE POSTERIOR STAPLE LINE WAS NOT COMPLETE AND DID NOT FULLY CAPTURE THE COLON A THIRD RELOAD WAS LOADED AND FIRED AGAIN 3-4CM DISTAL FROM THE SECOND TRANSECTION. STAPLE LINE WAS FULLY FORMED AND COLON TRANSECTED PDS 3-0 WAS SEWED OVER THE STAPLE LINE EVERYTHING WAS OK SURGERY WAS DELAYED BY 15 MINUTES THE SURGERY WAS COMPLETED SUCCESSFULLY.
31-08-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | POWERED PLUS STANDARD 45 | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER WORKED FINE FOR THE FIRST 5 RELOADS. THE 6TH RELOAD COULD NOT BE INSERTED BECAUSE SOMETHING WAS BLOCKING THE RELOAD AND KEPT PREVENTING IT FROM INSERTION. NEW STAPLER WAS UNPACKED TO BE ABLE TO FINISH OPERATION. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TOTAL PROCTOCOLECTOMY, THE DEVICE COULD NOT BE FIRED ALTHOUGH THE TRIGGER WAS PULLED. THE DEVICE WAS OPENED/CLOSED, AND THE BATTERY WAS REINSERTED BUT THE ISSUE DID NOT IMPROVE. THE TISSUE COMPRESSION SCALE BACKLIGHT WAS ILLUMINATED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE. THE ANVIL WAS NOT REPLACED. THE PROCEDURE WAS NOT CONVERTED.
30-04-2020 | Malfunction | ethicon endo-surgery | 10-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-04-2020 | Malfunction | ethicon endo-surgery | 10-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE PORT LOBECTOMY, WHILE RESECTIONING THE LUNG TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE EVEN AFTER PRESSING THE GREEN BUTTON. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE COLECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREATODUODENECTOMY SURGERY, WHILE ON GASTRIC RESECTION THE DEVICE DID NOT FIRE. THE FIRING KNOB COULD NOT BE PUSHED FORWARD AT ALL. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED IN THE COURSE OF A ANTERIOR RESECTION PROCEDURE USING CDH25A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE DID NOT PROCEED, AND STAPLES ALSO WEREN'T FIRED. AS AN EMERGENCY MEASURE, THE SURGEON DETACHED IT FROM THE TISSUE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES FOR THE PATIENT SO FAR.
14-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT COLECTOMY, THE PHYSICIAN TRIED TO RESECT THE LEFT COLON, CLOSED THE RELOAD ON THE TISSUE AND TRIED TO START STAPLING, THE DEVICE DID NOT FIRE. A SNAP WAS HEARD AS IF THE TISSUE WITHIN THE RELOAD WAS TOO THICK. IT DID NOT STAPLE. THE SURGEON USED A NEW HANDLE AND RELOAD THE RESOLVE THE ISSUE. IT WAS REPORTED THAT THE PATIENT DIED BUT ACCORDING TO THE PHYSICIAN THERE WAS NO LINK BETWEEN THE DEATH AND THE DEVICE ISSUE. IT WAS A PROCEDURE THAT WAS NOT PLANNED.
14-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT COLECTOMY, THE PHYSICIAN TRIED TO RESECT THE LEFT COLON, CLOSED THE RELOAD ON THE TISSUE AND TRIED TO START STAPLING, THE DEVICE DID NOT FIRE. A SNAP WAS HEARD AS IF THE TISSUE WITHIN THE RELOAD WAS TOO THICK. IT DID NOT STAPLE. THE SURGEON USED A NEW HANDLE AND RELOAD THE RESOLVE THE ISSUE. IT WAS REPORTED THAT THE PATIENT DIED BUT ACCORDING TO THE PHYSICIAN THERE WAS NO LINK BETWEEN THE DEATH AND THE DEVICE ISSUE. IT WAS A PROCEDURE THAT WAS NOT PLANNED.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ESOPHAGEAL SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT COLECTOMY, THE PHYSICIAN TRIED TO RESECT THE LEFT COLON, CLOSED THE RELOAD ON THE TISSUE AND TRIED TO START STAPLING, THE DEVICE DID NOT FIRE. A SNAP WAS HEARD AS IF THE TISSUE WITHIN THE RELOAD WAS TOO THICK. IT DID NOT STAPLE. THE SURGEON USED A NEW HANDLE AND RELOAD THE RESOLVE THE ISSUE. IT WAS REPORTED THAT THE PATIENT DIED BUT ACCORDING TO THE PHYSICIAN THERE WAS NO LINK BETWEEN THE DEATH AND THE DEVICE ISSUE. IT WAS A PROCEDURE THAT WAS NOT PLANNED.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLE DID NOT DEPLOY. THERE WERE NO PATIENT CONSEQUENCES.
23-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, ON ILEOTOMY AND INTESTINAL RESECTION, THE STAPLER WAS ABLE TO FIRE, THERE WAS POOR STAPLE FORMATION, A PORTION OF THE STAPLE LINE WAS NOT PROPERLY FORMED AND WAS INCOMPLETE. AN ADDITIONAL RESECTION WAS MADE BY USING A NEW RELOAD TO REMOVE THE AREA NEAR THE CORRESPONDING STAPLE LINE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC GASTRECTOMY, DURING JEJUNAL TRANSECTION, AFTER THE CLAMP TEST AND PRESSING THE GREEN BUTTON, WHEN ATTEMPTING TO FIRE, THE DISPLAY BECAME WHITE AND DISPLAYED 0. THE INDICATION OF THE HANDLE AND ADAPTER WERE IN GREEN. WHEN NEW EGIA45AVM WAS REPLACED, THE DEVICE COULD BE USED WITHOUT ANY PROBLEM, THE OPERATION WAS COMPLETED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC GASTRECTOMY, DURING JEJUNAL TRANSECTION, AFTER THE CLAMP TEST AND PRESSING THE GREEN BUTTON, WHEN ATTEMPTING TO FIRE, THE DISPLAY BECAME WHITE AND DISPLAYED 0. THE INDICATION OF THE HANDLE AND ADAPTER WERE IN GREEN. WHEN NEW EGIA45AVM WAS REPLACED, THE DEVICE COULD BE USED WITHOUT ANY PROBLEM, THE OPERATION WAS COMPLETED SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATECTOMY, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATECTOMY, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | BLUE RELOAD FOR ECHELON 60 | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: DAMAGED PACKAGING. THIS CASE IS FROM THE HEALTH AUTHORITY. IT WAS REPORTED THAT DURING THE SURGERY, BEFORE USED ON THE PATIENT AND OPENED THE PACKAGE, NOTED THE CORNER OF PACKAGE WAS NOT CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
24-12-2021 | Malfunction | celestica de monterrey s.a de c.v | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE SINGLE PORT LOBECTOMY, THE DEVICE DID NOT FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. TO RESOLVE THE ISSUE, THE SURGEON REPLACED WITH A NEW DEVICE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-08-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE PORT LOBECTOMY PROCEDURE, WHILE RESECTIONING THE LUNG TISSUE, THE DEVICE DID NOT FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE PORT LOBECTOMY PROCEDURE, WHILE RESECTIONING THE LUNG TISSUE, THE DEVICE DID NOT FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, THE DEVICE HAD AN UNKNOWN ISSUE. THE PROCEDURE TIME WAS EXTENDED BY 30 MINUTES. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, THE DEVICE HAD AN UNKNOWN ISSUE. THE PROCEDURE TIME WAS EXTENDED BY 30 MINUTES. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, THE DEVICE HAD AN UNKNOWN ISSUE. THE PROCEDURE TIME WAS EXTENDED BY 30 MINUTES. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | UNKNOWN ENDO GIA SULU | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, THE DEVICE HAD AN UNKNOWN ISSUE. THE PROCEDURE TIME WAS EXTENDED BY 30 MINUTES. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, THE DEVICE HAD AN UNKNOWN ISSUE. THE PROCEDURE TIME WAS EXTENDED BY 30 MINUTES. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-08-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE PORT THORACOSCOPIC LOBECTOMY PROCEDURE, WHILE LOADING THE STAPLES TO BE FIRED, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER PRIOR USE TO A PROCEDURE THE DEVICE DID NOT FIRE, ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. THERE WAS NO PATIENT INVOLVED.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEPLOYED STAPLES AT 7 O¿CLOCK FROM THE ANUS SIDE WERE UNFORMED. AN UNEXPECTED SOUND WAS HEARD AT THE FIRING. THERE IS A POSSIBILITY THAT THE RED SAFETY WAS NOT RELEASED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | ILS 29MM, CURVED | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEPLOYED STAPLES AT 7 O¿CLOCK FROM THE ANUS SIDE WERE UNFORMED. AN UNEXPECTED SOUND WAS HEARD AT THE FIRING. THERE IS A POSSIBILITY THAT THE RED SAFETY WAS NOT RELEASED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | ILS 29MM, CURVED | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEPLOYED STAPLES AT 7 O¿CLOCK FROM THE ANUS SIDE WERE UNFORMED. AN UNEXPECTED SOUND WAS HEARD AT THE FIRING. THERE IS A POSSIBILITY THAT THE RED SAFETY WAS NOT RELEASED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER PRIOR USE TO A PROCEDURE THE DEVICE DID NOT FIRE, ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. THERE WAS NO PATIENT INVOLVED.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER PRIOR USE TO A PROCEDURE THE DEVICE DID NOT FIRE, ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. THERE WAS NO PATIENT INVOLVED.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS NOT ABLE TO FIRE.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS NOT ABLE TO FIRE.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | PROXIMATE*RL VASCULAR LIN STAP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, SURGEON COULD NOT OPEN DEVICE AFTER FIRING COMPLETE. THERE WERE NO PATIENT CONSEQUENCES.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, WHILE TRANSECTING THE LIVER PARENCHYMA, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE EVEN AFTER PRESSING THE GREEN BUTTON. THE SURGEON USED A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS NOT ABLE TO FIRE.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS NOT ABLE TO FIRE.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
24-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY PROCEDURE, WHILE TRANSECTING GASTRIC FUNDUS, THE DEVICE DID NOT FIRE. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-12-2021 | Injury | ethicon endo-surgery | 10-01-2022 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED, ANASTOMOTIC LEAK. THE RADIOLOGIST HIGHLIGHTED BADLY FORMED POINTS. POST-OP TO LOW ANTERIOR RESECTION.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SINGLE PORT THORACOSCOPIC LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, WHILE PERFORMING SEGMENT RESECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN AFTER PRESSING THE GREEN BUTTON. THE SURGEON USED A NEW HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE PORT LOBECTOMY, WHILE RESECTIONING THE LUNG TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE EVEN AFTER PRESSING THE GREEN BUTTON. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLORECTAL SURGERY, THE HANDLE STILL SQUEEZED BUT COULD NOT FIRE AND BROKE. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE HANDLE WITH A NEW ONE. THERE WAS NO PATIENT INJURY.
25-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLORECTAL SURGERY, THE HANDLE STILL SQUEEZED BUT COULD NOT FIRE AND BROKE. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE HANDLE WITH A NEW ONE. THERE WAS NO PATIENT INJURY.
25-12-2021 | Malfunction | covidien lp llc north haven | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE STOMACH DURING A LAPAROSCOPIC PANCREATICDUODENUM SURGERY, THE DEVICE COULD NOT FIRE AFTER PRESSING THE GREEN BUTTON. THE RELOAD WAS REPLACED TO CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
26-08-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY THE DEVICE DID NOT STAPLE ALL THE WAY. NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | POWERED PLUS COMPACT 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, AN UNEXPECTED SOUND WAS HEARD AT THE FIRING. THE STAPLES AT THE END OF THE BRONCHIAL TUBE WERE DEPLOYED BUT OTHER STAPLES FELL OFF THE TISSUE. ONLY THE PROXIMAL END OF THE TISSUE WAS CUT SLIGHTLY. BEFORE THE BRONCHIAL TUBE WAS CUT, THE INTERLOBAR WAS CUT AND NO PROBLEM WAS OBSERVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | POWERED PLUS COMPACT 45 | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, AN UNEXPECTED SOUND WAS HEARD AT THE FIRING. THE STAPLES AT THE END OF THE BRONCHIAL TUBE WERE DEPLOYED BUT OTHER STAPLES FELL OFF THE TISSUE. ONLY THE PROXIMAL END OF THE TISSUE WAS CUT SLIGHTLY. BEFORE THE BRONCHIAL TUBE WAS CUT, THE INTERLOBAR WAS CUT AND NO PROBLEM WAS OBSERVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | POWERED PLUS COMPACT 45 | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, AN UNEXPECTED SOUND WAS HEARD AT THE FIRING. THE STAPLES AT THE END OF THE BRONCHIAL TUBE WERE DEPLOYED BUT OTHER STAPLES FELL OFF THE TISSUE. ONLY THE PROXIMAL END OF THE TISSUE WAS CUT SLIGHTLY. BEFORE THE BRONCHIAL TUBE WAS CUT, THE INTERLOBAR WAS CUT AND NO PROBLEM WAS OBSERVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
24-12-2021 | Malfunction | celestica de monterrey s.a de c.v | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE PORT LOBECTOMY PROCEDURE, WHILE RESECTIONING THE LUNG TISSUE, THE DEVICE DID NOT FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | celestica de monterrey s.a de c.v | 10-01-2022 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE PORT LOBECTOMY PROCEDURE, WHILE RESECTIONING THE LUNG TISSUE, THE DEVICE DID NOT FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, THE DEVICE WAS LOCKED OUT. THE KNIFE REVERSE SWITCH WAS USED, BUT THE KNIFE DID NOT RETURN TO HOME POSITION. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, THE DEVICE WAS LOCKED OUT. THE KNIFE REVERSE SWITCH WAS USED, BUT THE KNIFE DID NOT RETURN TO HOME POSITION. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2021 | Malfunction | ethicon endo-surgery | 10-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, THE DEVICE WAS LOCKED OUT. THE KNIFE REVERSE SWITCH WAS USED, BUT THE KNIFE DID NOT RETURN TO HOME POSITION. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
24-12-2021 | Malfunction | us surgical puerto rico | 10-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TOT HE REPORTER DURING THE SINGLE PORT LOBECTOMY, THE DEVICE DID NOT FIRE ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. TO RESOLVE THE ISSUE, THE SURGEON REPLACED WITH NEW DEVICE. THERE WAS NO PATIENT INJURY.
01-01-2021 | Death | ethicon endo-surgery | 10-01-2022 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE A MISFIRE IN A WHITE RELOAD GST60W TOOK PLACE, THE LOT# WAS U4050R.
01-01-2021 | Death | ethicon endo-surgery | 10-01-2022 | GST WHITE RELOAD, 60MM, 6 ROW | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE A MISFIRE IN A WHITE RELOAD GST60W TOOK PLACE, THE LOT# WAS U4050R.
01-01-2021 | Death | ethicon endo-surgery | 10-01-2022 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE A MISFIRE IN A WHITE RELOAD GST60W TOOK PLACE, THE LOT# WAS U40268.
01-01-2021 | Death | ethicon endo-surgery | 10-01-2022 | GST WHITE RELOAD, 60MM, 6 ROW | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE A MISFIRE IN A WHITE RELOAD GST60W TOOK PLACE, THE LOT# WAS U40268.
15-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THERE WERE NO PROBLEMS WITH THE LOADING PERFORMED DURING THE PARTIAL PULMONARY RESECTION, AND AFTER APPROACHING THE LUNG, AN ATTEMPT WAS MADE TO FIRE, BUT FAILED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE.
01-01-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE ANVIL CANNOT OPEN AFTER FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THERE WERE NO PROBLEMS WITH THE LOADING PERFORMED DURING THE PARTIAL PULMONARY RESECTION, AND AFTER APPROACHING THE LUNG, AN ATTEMPT WAS MADE TO FIRE, BUT FAILED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE.
23-01-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTROPLASTY PROCEDURE, IN THE SECOND STAPLE, THE LOAD OPENED THE STAPLE LINE, AND IT WAS NECESSARY TO USE A NEW LOAD. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PARTIAL LUNG RESECTION, THE DEVICE WAS UNABLE TO FIRE. THE GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | GOLD RELOAD FOR ECHELON 60 | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: LATERAL PART OF THE STEEL MATERIAL IS NOT PROPERLY ATTACHED TO RELOAD, DID THE EVENT OCCUR DURING STERILE PROCESSING? UNKNOWN, DID THE EVENT OCCUR DURING FIELD INSPECTION? --> UNKNOWN, DID THE EVENT OCCUR DURING INTERNAL SERVIC ACTIVITIES SUCH AS CALIBRATION? UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES NO, WAS OTHER MEDICAL IERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN, IS THE PATIENT PART OF A CLINICAL STUDY UNKNOWN, IP-01331835. DEVICE PROPERTY OF: NONE, DEVICE IN POSSESSION OF -NONE. BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST. TRUE.
21-11-2021 | Injury | covidien lp llc north haven | 11-01-2022 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER AFTER THE PROCEDURE THE DEVICE ENCOUNTERED AN UNKNOWN ISSUE.
04-11-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON LAPAROSCOPIC BARIATRIC SLEEVE, WHEN THE LOAD WAS FIRED, THE HANDLE MADE A NOISE. FOR THE SECOND FIRING DEVICE DID NOT FIRE. THE DEVICE WAS DIFFICULT TO UNLOAD. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-11-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON LAPAROSCOPIC BARIATRIC SLEEVE, WHEN THE LOAD WAS FIRED, THE HANDLE MADE A NOISE. FOR THE SECOND FIRING DEVICE DID NOT FIRE. THE DEVICE WAS DIFFICULT TO UNLOAD. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-11-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON LAPAROSCOPIC BARIATRIC SLEEVE, WHEN THE LOAD WAS FIRED, THE HANDLE MADE A NOISE. FOR THE SECOND FIRING DEVICE DID NOT FIRE. THE DEVICE WAS DIFFICULT TO UNLOAD. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS WOBBLED AFTER A CARTRIDGE WAS LOADED INTO. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE. ANOTHER DEVICE WAS USED TO CONTINUE THE CASE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, DURING INSTALLATION, THE JAWS WOULD NOT OPEN. THERE WAS NO PATIENT INVOLVEMENT.
07-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | EEA | Failure to Align | A050701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE CREATING ANASTOMOSIS IN A COLOSTOMY TAKE DOWN, THE ANVIL TILTED PRIOR TO FIRING. TO RESOLVE THE ISSUE, A NEW DEVICE WAS OPENED. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY FOR ENDOMETRIOSIS, AT DISCOIDAL COLON RESECTION, GREEN MARK WAS NOTED BUT THE DEVICE DID NOT FIRE NOR CUT THE TISSUE. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, DURING INSTALLATION, THE JAWS WOULD NOT OPEN. THERE WAS NO PATIENT INVOLVEMENT.
27-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. NEW PRODUCT WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. NEW PRODUCT WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN RELOAD WAS CONNECTED TO THE HANDLE, THE JAWS WOULD NOT OPEN. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | EEA ORVIL | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN TOTAL GASTRECTOMY AND WHIPPLE, WHEN THE ESOPHAGUS WAS BEING ANASTOMOSED, THE DELIVERY TUBE WAS DIFFICULT TO REMOVE AFTER ONE OF THE TWO SUTURES WAS CUT. THE SURGEON CUT THE SECOND SUTURE TO ASSIST IN REMOVING THE TUBE AND TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Injury | celestica de monterrey s.a de c.v | 11-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Injury | celestica de monterrey s.a de c.v | 11-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE WAS DONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | celestica de monterrey s.a de c.v | 11-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT FORM PROPERLY, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO FIX THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | celestica de monterrey s.a de c.v | 11-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT FORM PROPERLY, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO FIX THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
06-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC RADICAL GASTRECTOMY FOR GASTRIC CANCER SURGERY, COULD NOT FIRE WHEN SEVERING GASTRIC TISSUE ON THE 3TH FIRING. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC SEGMENTECTOMY, WHEN DETACHING LUNG TISSUE, THE DEVICE DID NOT FIRE. THE GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN RELOAD WAS CONNECTED TO THE HANDLE, THE JAWS WOULD NOT OPEN. THERE WAS NO PATIENT INVOLVEMENT.
27-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED.
01-06-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, PRIOR TO PATIENT CONTACT, HANDLE SHOWED AN ERROR MESSAGE AND IT DOES NOT WORK. THIS POWER HANDLE HAD BEEN UPDATED ONE WEEK AGO TO THE LATEST VERSION 29.3.2, THEN IT WAS USED IN 3-4 CASES WITHOUT ANY ERROR. DURING ASSEMBLY TO THE POWER SHELL IT SHOWED A GRAPHIC WITH A HANDLE AND A RED CIRCLE WITH A LINE GOING THROUGH IT WAS DISPLAYED ON THE SCREEN THEN IT STOPPED WORKING. IT WAS TRIED TO RECONNECT TO THE CHARGER AND REBOOTING THE HANDLE BUT UNFORTUNATELY ALL TRIALS FAILED. THERE WAS NO PATIENT INVOLVEMENT.
01-06-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, PRIOR TO PATIENT CONTACT, HANDLE SHOWED AN ERROR MESSAGE AND IT DOES NOT WORK. THIS POWER HANDLE HAD BEEN UPDATED ONE WEEK AGO TO THE LATEST VERSION 29.3.2, THEN IT WAS USED IN 3-4 CASES WITHOUT ANY ERROR. DURING ASSEMBLY TO THE POWER SHELL IT SHOWED A GRAPHIC WITH A HANDLE AND A RED CIRCLE WITH A LINE GOING THROUGH IT WAS DISPLAYED ON THE SCREEN THEN IT STOPPED WORKING. IT WAS TRIED TO RECONNECT TO THE CHARGER AND REBOOTING THE HANDLE BUT UNFORTUNATELY ALL TRIALS FAILED. THERE WAS NO PATIENT INVOLVEMENT.
01-06-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, PRIOR TO PATIENT CONTACT, HANDLE SHOWED AN ERROR MESSAGE AND IT DOES NOT WORK. THIS POWER HANDLE HAD BEEN UPDATED ONE WEEK AGO TO THE LATEST VERSION 29.3.2, THEN IT WAS USED IN 3-4 CASES WITHOUT ANY ERROR. DURING ASSEMBLY TO THE POWER SHELL IT SHOWED A GRAPHIC WITH A HANDLE AND A RED CIRCLE WITH A LINE GOING THROUGH IT WAS DISPLAYED ON THE SCREEN THEN IT STOPPED WORKING. IT WAS TRIED TO RECONNECT TO THE CHARGER AND REBOOTING THE HANDLE BUT UNFORTUNATELY ALL TRIALS FAILED. THERE WAS NO PATIENT INVOLVEMENT.
01-06-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, PRIOR TO PATIENT CONTACT, HANDLE SHOWED AN ERROR MESSAGE AND IT DOES NOT WORK. THIS POWER HANDLE HAD BEEN UPDATED ONE WEEK AGO TO THE LATEST VERSION 29.3.2, THEN IT WAS USED IN 3-4 CASES WITHOUT ANY ERROR. DURING ASSEMBLY TO THE POWER SHELL IT SHOWED A GRAPHIC WITH A HANDLE AND A RED CIRCLE WITH A LINE GOING THROUGH IT WAS DISPLAYED ON THE SCREEN THEN IT STOPPED WORKING. IT WAS TRIED TO RECONNECT TO THE CHARGER AND REBOOTING THE HANDLE BUT UNFORTUNATELY ALL TRIALS FAILED. THERE WAS NO PATIENT INVOLVEMENT.
04-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | UNKNOWN STAPLING DEVICE | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY, THE RELOAD WAS PARTIALLY FIRED. THERE WERE REMAINING STAPLES IN THE RELOAD. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN REMOVING THE AFFECTED TISSUE DURING A LAPAROSCOPIC PNEUMOTHORAX, DEVICE OPENED AND CLOSED OUTSIDE THE PATIENT WHEN IT WAS INSERTED THROUGH THE TROCAR, IT NO LONGER OPENED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY, THE RELOAD WAS PARTIALLY FIRED. THERE WERE REMAINING STAPLES IN THE RELOAD. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-09-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING, GREEN LIGHT WAS SHOWING ON THE CHARGER BUT HANDLE WILL NOT POWER AND LIGHT EMITTING DIODE (LED) SCREEN WAS BLACK. THERE WAS NO PATIENT INVOLVEMENT.
21-09-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING, GREEN LIGHT WAS SHOWING ON THE CHARGER BUT HANDLE WILL NOT POWER AND LIGHT EMITTING DIODE (LED) SCREEN WAS BLACK. THERE WAS NO PATIENT INVOLVEMENT.
23-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED , NOTED THE STAPLES MALFORMED . USED THE NEW ONE TO CONTINUE SURGERY. COULD NOT INSTALL THE DEVICE , CHANGED ANOTHER NEW DEVICE , COULD INSTALL INTO THE GUN. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | PWRD ECH FLEX 45MM 34CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED , NOTED THE STAPLES MALFORMED . USED THE NEW ONE TO CONTINUE SURGERY. COULD NOT INSTALL THE DEVICE , CHANGED ANOTHER NEW DEVICE , COULD INSTALL INTO THE GUN. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER MALFUNCTIONED. KNIFE DID NOT RETURN HOME FULLY. TRIED KNIFE-REVERSE SWITCH 2-3 TIMES BUT NOTHING HAPPENED. GUN UNRESPONSIVE. EVENTUALLY AFTER TRYING EVERYTHING AND ANYTHING THE GUN STARTED WORKING AND THE KNIFE RETURNED HOME. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER MALFUNCTIONED. KNIFE DID NOT RETURN HOME FULLY. TRIED KNIFE-REVERSE SWITCH 2-3 TIMES BUT NOTHING HAPPENED. GUN UNRESPONSIVE. EVENTUALLY AFTER TRYING EVERYTHING AND ANYTHING THE GUN STARTED WORKING AND THE KNIFE RETURNED HOME. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER MALFUNCTIONED. KNIFE DID NOT RETURN HOME FULLY. TRIED KNIFE-REVERSE SWITCH 2-3 TIMES BUT NOTHING HAPPENED. GUN UNRESPONSIVE. EVENTUALLY AFTER TRYING EVERYTHING AND ANYTHING THE GUN STARTED WORKING AND THE KNIFE RETURNED HOME. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
13-10-2021 | Malfunction | us surgical puerto rico | 11-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, WHILE THE TISSUE WAS BEING CUT, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN AFTER PRESSING THE GREEN BUTTON. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-09-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, A HANDLE ERROR SHOWING RED LINE AND WARNING SIGN ON TOP WAS DISPLAYED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
21-09-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, A HANDLE ERROR SHOWING RED LINE AND WARNING SIGN ON TOP WAS DISPLAYED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
19-11-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN CHECKING THE OPENING/CLOSING OF THE JAWS, AN ABNORMAL SOUND WAS HEARD. THE SURGEON TESTED THE DEVICE OUT OF FIELD AND FOUND THAT DEVICE WAS UNABLE TO FIRE. ANOTHER PRODUCT WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
19-11-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHEN CHECKING THE OPENING/CLOSING OF THE JAWS, AN ABNORMAL SOUND WAS HEARD. THE SURGEON TESTED THE DEVICE OUT OF FIELD AND FOUND THAT DEVICE WAS UNABLE TO FIRE. ANOTHER PRODUCT WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
01-01-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER WAS STUCK DURING SURGERY. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER WAS STUCK DURING SURGERY. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS LOCKED ON THE SPLENIC ARTERY. THE SURGEON CHANGED THE ANGULATION OF THE DEVICE BUT AT THE STAPLING, THE DEVICE GOT JAMMED. IT WAS IMPOSSIBLE TO OPEN THE JAWS, INCLUDING BY MANUAL METHOD. THE MECHANISM DID NOT INTERLOCK. THE DEVICE WAS OPENED WITH ANOTHER INSTRUMENT. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. HEMORRHAGE RISK. A SUTURE OF THE ARTERY WAS PERFORMED WITH THREADS.
22-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS LOCKED ON THE SPLENIC ARTERY. THE SURGEON CHANGED THE ANGULATION OF THE DEVICE BUT AT THE STAPLING, THE DEVICE GOT JAMMED. IT WAS IMPOSSIBLE TO OPEN THE JAWS, INCLUDING BY MANUAL METHOD. THE MECHANISM DID NOT INTERLOCK. THE DEVICE WAS OPENED WITH ANOTHER INSTRUMENT. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. HEMORRHAGE RISK. A SUTURE OF THE ARTERY WAS PERFORMED WITH THREADS.
22-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS LOCKED ON THE SPLENIC ARTERY. THE SURGEON CHANGED THE ANGULATION OF THE DEVICE BUT AT THE STAPLING, THE DEVICE GOT JAMMED. IT WAS IMPOSSIBLE TO OPEN THE JAWS, INCLUDING BY MANUAL METHOD. THE MECHANISM DID NOT INTERLOCK. THE DEVICE WAS OPENED WITH ANOTHER INSTRUMENT. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. HEMORRHAGE RISK. A SUTURE OF THE ARTERY WAS PERFORMED WITH THREADS.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION AND GASTRIC TUBE ANASTOMOSIS, WHEN CREATING GASTRIC TUBE, THE FIRST FIRING WAS PERFORMED BUT KNIFE WOULD NOT RETRACT - LOCKED ON TISSUE. THE KNIFE WAS RETRACTED AND DEVICE WAS ABLE TO BE REMOVED AFTER RESTARTING THE HANDLE. AFTERWARDS, A HANDLE ERROR MARK WAS DISPLAYED ON THE SCREEN. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION AND GASTRIC TUBE ANASTOMOSIS, WHEN CREATING GASTRIC TUBE, THE FIRST FIRING WAS PERFORMED BUT KNIFE WOULD NOT RETRACT - LOCKED ON TISSUE. THE KNIFE WAS RETRACTED AND DEVICE WAS ABLE TO BE REMOVED AFTER RESTARTING THE HANDLE. AFTERWARDS, A HANDLE ERROR MARK WAS DISPLAYED ON THE SCREEN. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION AND GASTRIC TUBE ANASTOMOSIS, WHEN CREATING GASTRIC TUBE, THE FIRST FIRING WAS PERFORMED BUT KNIFE WOULD NOT RETRACT - LOCKED ON TISSUE. THE KNIFE WAS RETRACTED AND DEVICE WAS ABLE TO BE REMOVED AFTER RESTARTING THE HANDLE. AFTERWARDS, A HANDLE ERROR MARK WAS DISPLAYED ON THE SCREEN. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION AND GASTRIC TUBE ANASTOMOSIS, WHEN CREATING GASTRIC TUBE, THE FIRST FIRING WAS PERFORMED BUT KNIFE WOULD NOT RETRACT - LOCKED ON TISSUE. THE KNIFE WAS RETRACTED AND DEVICE WAS ABLE TO BE REMOVED AFTER RESTARTING THE HANDLE. AFTERWARDS, A HANDLE ERROR MARK WAS DISPLAYED ON THE SCREEN. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | SIGNIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION AND GASTRIC TUBE ANASTOMOSIS, WHEN CREATING GASTRIC TUBE, THE FIRST FIRING WAS PERFORMED BUT KNIFE WOULD NOT RETRACT - LOCKED ON TISSUE. THE KNIFE WAS RETRACTED AND DEVICE WAS ABLE TO BE REMOVED AFTER RESTARTING THE HANDLE. AFTERWARDS, A HANDLE ERROR MARK WAS DISPLAYED ON THE SCREEN. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP APPENDECTOMY, THE SURGEON FIRED THE STAPLER AND WHEN THEY OPENED THE STAPLER ABOUT ONE CENTIMETER OF THE PORTION CLOSEST TO THE APEX OF THE STAPLER DEPLOYED. CASE COMPLETED WITH A NEW STAPLER BELOW THE OTHER STAPLE LINE WHICH REQUIRED TWO STAPLE LOADS. NO PATIENT HARM. DEVICE IS AVAILABLE FOR RETURN.
27-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THAT THE FIRST TWO FIRINGS WITH GOLD CARTRIDGES AND STAPLE LINE REINFORCEMENT THE DEVICE FIRED AS EXPECTED. ON THE THIRD FIRING WITH A BLUE CARTRIDGE AND NO STAPLE LINE REINFORCEMENT, THE SURGEON WAS PULSE FIRING AND IT APPEARED THAT THE DEVICE WAS PUSHING THE TISSUE WHILE FIRING. ONCE REMOVED IT WAS NOTICED THAT THERE WERE SEVERAL FORMED STAPLES BUNCHED UP AT THE END OF THE STAPLE LINE. WHILE ATTEMPTING THE FOURTH FIRING THE DEVICE SOUNDED AS IF SOMETHING WAS WRONG AND IT APPEARED TO BE PUSHING TISSUE. DEVICE WAS REVERSED AND REMOVED FROM THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. PRIOR TO REMOVING THE SPECIMEN, A LEAK TEST WAS PERFORMED, AND NO LEAKS WERE FOUND. WHILE REMOVING THE SPECIMEN IT WAS NOTICED THAT THERE WAS A LARGE HOLE WHERE STAPLES SHOULD HAVE BEEN. SURGEON THEN OVERSEWED THE CORRESPONDING AREA ON THE PATIENT SIDE TO STRENGTHEN THE STAPLE LINE. A SECOND LEAK TEST WAS THEN PERFORMED WITH NO LEAKS DETECTED. SURGEON THEN USED VISTASEAL ON THE STAPLE LINE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
27-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THAT THE FIRST TWO FIRINGS WITH GOLD CARTRIDGES AND STAPLE LINE REINFORCEMENT THE DEVICE FIRED AS EXPECTED. ON THE THIRD FIRING WITH A BLUE CARTRIDGE AND NO STAPLE LINE REINFORCEMENT, THE SURGEON WAS PULSE FIRING AND IT APPEARED THAT THE DEVICE WAS PUSHING THE TISSUE WHILE FIRING. ONCE REMOVED IT WAS NOTICED THAT THERE WERE SEVERAL FORMED STAPLES BUNCHED UP AT THE END OF THE STAPLE LINE. WHILE ATTEMPTING THE FOURTH FIRING THE DEVICE SOUNDED AS IF SOMETHING WAS WRONG AND IT APPEARED TO BE PUSHING TISSUE. DEVICE WAS REVERSED AND REMOVED FROM THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. PRIOR TO REMOVING THE SPECIMEN, A LEAK TEST WAS PERFORMED, AND NO LEAKS WERE FOUND. WHILE REMOVING THE SPECIMEN IT WAS NOTICED THAT THERE WAS A LARGE HOLE WHERE STAPLES SHOULD HAVE BEEN. SURGEON THEN OVERSEWED THE CORRESPONDING AREA ON THE PATIENT SIDE TO STRENGTHEN THE STAPLE LINE. A SECOND LEAK TEST WAS THEN PERFORMED WITH NO LEAKS DETECTED. SURGEON THEN USED VISTASEAL ON THE STAPLE LINE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION AND GASTRIC TUBE ANASTOMOSIS, WHEN CREATING GASTRIC TUBE, THE FIRST FIRING WAS PERFORMED BUT KNIFE WOULD NOT RETRACT - LOCKED ON TISSUE. THE KNIFE WAS RETRACTED AND DEVICE WAS ABLE TO BE REMOVED AFTER RESTARTING THE HANDLE. AFTERWARDS, A HANDLE ERROR MARK WAS DISPLAYED ON THE SCREEN. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | TRI-STAPLE 2.0 | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ESOPHAGUS RESECTION AND GASTRIC TUBE ANASTOMOSIS, WHEN CREATING GASTRIC TUBE, THE FIRST FIRING WAS PERFORMED BUT KNIFE WOULD NOT RETRACT - LOCKED ON TISSUE. THE KNIFE WAS RETRACTED AND DEVICE WAS ABLE TO BE REMOVED AFTER RESTARTING THE HANDLE. AFTERWARDS, A HANDLE ERROR MARK WAS DISPLAYED ON THE SCREEN. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY, THE DEVICE DID NOT FIRE - THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE. ANOTHER DEVICE WAS USED TO CONTINUE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 11-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PARTIAL LUNG RESECTION, THE DEVICE WAS UNABLE TO FIRE - GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE CARTRIDGE LOADED ON A PCEE60A FELL OFF IN THE PATIENT WHEN THE DEVICE WAS INSERTED INTO THE PATIENT AT THE 1ST FIRING. THE CARTRIDGE WAS RETRIEVED BY A FORCEPS AND NOT USED FOR THE PATIENT. THE 2ND AND 3RD FIRINGS WERE COMPLETED WITH NO PROBLEM. AFTER THAT, A 2 TO 3 MM WHITE PART WAS FOUND INSIDE THE PATIENT. THE DEVICE WAS USED ON THE LUNG. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE CARTRIDGE LOADED ON A PCEE60A FELL OFF IN THE PATIENT WHEN THE DEVICE WAS INSERTED INTO THE PATIENT AT THE 1ST FIRING. THE CARTRIDGE WAS RETRIEVED BY A FORCEPS AND NOT USED FOR THE PATIENT. THE 2ND AND 3RD FIRINGS WERE COMPLETED WITH NO PROBLEM. AFTER THAT, A 2 TO 3 MM WHITE PART WAS FOUND INSIDE THE PATIENT. THE DEVICE WAS USED ON THE LUNG. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE CARTRIDGE LOADED ON A PCEE60A FELL OFF IN THE PATIENT WHEN THE DEVICE WAS INSERTED INTO THE PATIENT AT THE 1ST FIRING. THE CARTRIDGE WAS RETRIEVED BY A FORCEPS AND NOT USED FOR THE PATIENT. THE 2ND AND 3RD FIRINGS WERE COMPLETED WITH NO PROBLEM. AFTER THAT, A 2 TO 3 MM WHITE PART WAS FOUND INSIDE THE PATIENT. THE DEVICE WAS USED ON THE LUNG. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | FLEX 60 ARTICULATING | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID RESECTION THE 3RD FIRING OF THE MANUAL STAPLER, THE DEVICE KNIFE RETURNED BUT DEVICE DID NOT OPEN. RETURN BUTTON DID NOT WORK, KNIFE WAS RETURNED BUT DEVICE DID NOT WANT TO OPEN WITH SIGMOID BETWEEN THE STAPLER. SURGEON NEEDED TO CUT OUT THE STAPLER FROM THE SIGMOID AND MANUALLY CLOSED THE ANASTOMOSIS.
23-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID RESECTION THE 3RD FIRING OF THE MANUAL STAPLER, THE DEVICE KNIFE RETURNED BUT DEVICE DID NOT OPEN. RETURN BUTTON DID NOT WORK, KNIFE WAS RETURNED BUT DEVICE DID NOT WANT TO OPEN WITH SIGMOID BETWEEN THE STAPLER. SURGEON NEEDED TO CUT OUT THE STAPLER FROM THE SIGMOID AND MANUALLY CLOSED THE ANASTOMOSIS.
20-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING USE IN THE OPERATING FIELD, THE DEVICE ¿STICKS¿ IN ¿CLOSED BUNDLE¿ MODE AND COMPONENT (SPRING) WAS FOUND IN THE OPERATING FIELD. ILEUM RESECTION INTERVENTION FOR PERFORATION, DEVICE USED ON ILEUM. DURATION OF THE OPERATION TWO HOURS. NO CONSEQUENCE FOR THE PATIENT. THE PROBLEM WAS FOUND AT THE END OF THE INTERVENTION.
20-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, DEVICE COULD NOT FIRE ON THE ESOPHAGEAL TISSUE ON THE FIRST FIRING . CHANGED TO ANOTHER DEVICE. AFTER FIRED, REMOVED THE DEVICE AND NOTED SOME STAPLES MALFORMED . CHANGED TO ANOTHER DEVICE TO USE, AND THEY HAD THE SAME PROBLEM. CHANGED TO ANOTHER NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
20-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, DEVICE COULD NOT FIRE ON THE ESOPHAGEAL TISSUE ON THE FIRST FIRING . CHANGED TO ANOTHER DEVICE. AFTER FIRED, REMOVED THE DEVICE AND NOTED SOME STAPLES MALFORMED . CHANGED TO ANOTHER DEVICE TO USE, AND THEY HAD THE SAME PROBLEM. CHANGED TO ANOTHER NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
08-11-2021 | Injury | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE THAT THE JAWS OF THE STAPLER CLOSED COMPLETELY BUT THE STAPLER MISS-FIRED AND STAPLES DID NOT FORM. UNKNOWN AS TO HOW THE DEVICE WAS REMOVED FROM TISSUE. PROCEDURE WAS CONVERTED TO AN OPEN ILEOCECECTOMY. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-11-2021 | Injury | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE THAT THE JAWS OF THE STAPLER CLOSED COMPLETELY BUT THE STAPLER MISS-FIRED AND STAPLES DID NOT FORM. UNKNOWN AS TO HOW THE DEVICE WAS REMOVED FROM TISSUE. PROCEDURE WAS CONVERTED TO AN OPEN ILEOCECECTOMY. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-11-2021 | Injury | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE THAT THE JAWS OF THE STAPLER CLOSED COMPLETELY BUT THE STAPLER MISS-FIRED AND STAPLES DID NOT FORM. UNKNOWN AS TO HOW THE DEVICE WAS REMOVED FROM TISSUE. PROCEDURE WAS CONVERTED TO AN OPEN ILEOCECECTOMY. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-11-2021 | Injury | ethicon endo-surgery | 11-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE THAT THE JAWS OF THE STAPLER CLOSED COMPLETELY BUT THE STAPLER MISS-FIRED AND STAPLES DID NOT FORM. UNKNOWN AS TO HOW THE DEVICE WAS REMOVED FROM TISSUE. PROCEDURE WAS CONVERTED TO AN OPEN ILEOCECECTOMY. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
17-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE ECHELON FLEX STAPLER DID NOT STAPLE AND LOCKED IN THE OPENING, IN ADDITION TO NOT PRESSING THE TISSUE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
17-12-2021 | Malfunction | ethicon endo-surgery | 11-01-2022 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE ECHELON FLEX STAPLER DID NOT STAPLE AND LOCKED IN THE OPENING, IN ADDITION TO NOT PRESSING THE TISSUE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
17-11-2021 | Injury | intuitive surgical | 11-01-2022 | ENDOWRIST | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LOW ANTERIOR RESECTION (LAR) SURGICAL PROCEDURE, THE STAPLE LINE FORMED AS EXPECTED WITH AN ENDOWRIST STAPLER 45 INSTRUMENT AND GREEN RELOAD, THEN SUBSEQUENTLY THE STAPLE LINE OPENED BACK UP. IT WAS REPORTED THAT A SECOND GREEN RELOAD WAS USED AND THE SAME BEHAVIOR WAS OBSERVED AGAIN (CAPTURED IN THIS REPORT). MORE SPECIFICALLY, IT WAS REPORTED THAT THE TWO STAPLE LINES WERE FORMED AND B-SHAPE FORMATION (INTENDED FORMATION) WAS OBSERVED, BUT SLIGHTLY AFTER STAPLING WAS PERFORMED, THE STAPLE LINE OPENED BACK UP, WHICH RESULTED IN A NEED TO EXCISE ADDITIONAL COLON TISSUE THAT WAS NOT PLANNED TO BE EXCISED. THE THIRD RELOAD STAPLED AS EXPECTED/INTENDED AND THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE DA VINCI SURGICAL SYSTEM. THE INITIAL REPORTER (NURSE) INDICATED THAT THE PATIENT DID NOT EXPERIENCE ANY INTRA-OPERATIVE/POST-OPERATIVE COMPLICATIONS AND THAT THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE DA VINCI SURGICAL SYSTEM, INSTRUMENTS, AND ACCESSORIES ONCE THE THIRD RELOAD WAS SUCCESSFUL. PER FOLLOW-UP WITH THE INITIAL REPORTER, THE SURGEON INDICATED THAT THIS EVENT DID NOT IMPACT THE PROCEDURE AS A WHOLE AND THAT THERE WAS NO EFFECT TO THE PATIENT¿S HEALTH. THE SURGEON INDICATED THAT THERE WAS NO CONCERN REGARDING THIS EVENT POTENTIALLY CAUSING LONG-TERM COMPLICATIONS FOR THE PATIENT. TRANSECTION WAS REPORTED TO HAVE BEEN NORMAL AND THE STAPLES WERE NOTED TO HAVE FORMED NORMALLY INTO A B-SHAPE FORMATION BEFORE THE STAPLE LINE OPENED. THERE WAS NO REPORT OF TISSUE TEARING, BUT ALSO NO CONFIRMATION THAT THE STAPLES BECAME UNFORMED AFTER INITIALLY FORMING. AT THIS TIME, THE EXACT WAY IN WHICH THE STAPLE LINE ¿OPENED-UP¿ REMAINS UNKNOWN. THE SAME ENDOWRIST STAPLER 45 INSTRUMENT WAS USED FOR ALL RELOAD FIRES AND PERFORMED AS EXPECTED. THERE WERE NO ERROR MESSAGES. IT WAS REPORTED THAT THIS REPORTED BEHAVIOR WAS OBSERVED IN BACK-TO-BACK FIRES AND ONLY SHORTLY AFTER BEGINNING THE PROCEDURE. THERE WAS NO REPORT OF TISSUE BUNCHING AND NO REPORTS OF STAPLES DEPLOYING AT AN INCORRECT TIME. IT WAS CONFIRMED THAT ALL STAPLES DEPLOYED AND NO STAPLES WERE PRESENT IN THE RELOAD CARTRIDGE FOLLOWING STAPLING. THERE WERE NO ISSUES WITH THE INSTRUMENT/RELOAD STICKING TO TISSUE. THE SURGEON DID NOT ENCOUNTER OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. NO BUTTRESS MATERIAL WAS USED. THERE WERE NO MALFORMED STAPLES, THE STAPLES APPEARED TO HAVE BEEN DEPLOYED NORMALLY/AS INTENDED BEFORE THEY ¿OPENED-UP.¿ THERE WERE NO CLAMPING ISSUES NOTED. THE TISSUE WAS NOT CALCIFIED AND WAS NOT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY PRIOR TO THE PROCEDURE. THERE WAS NO TISSUE TENSION AND NO TISSUE BUNCHING. THIS MEDICAL DEVICE REPORT (MDR) CAPTURES THE REPORT OF THE SECOND RELOAD THAT EXPERIENCED THE SAME STAPLING ISSUE DURING THE REPORTED EVENT. THE FIRST RELOAD IS REPORTED UNDER MDR WITH PATIENT IDENTIFIER #(B)(6).
17-11-2021 | Injury | intuitive surgical | 11-01-2022 | ENDOWRIST | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LOW ANTERIOR RESECTION (LAR) SURGICAL PROCEDURE, THE STAPLE LINE FORMED AS EXPECTED WITH AN ENDOWRIST STAPLER 45 INSTRUMENT AND GREEN RELOAD THEN SUBSEQUENTLY THE STAPLE LINE OPENED BACK UP. IT WAS REPORTED THAT A SECOND GREEN RELOAD WAS USED AND THE SAME BEHAVIOR WAS OBSERVED AGAIN. MORE SPECIFICALLY, IT WAS REPORTED THAT THE TWO STAPLE LINES WERE FORMED AND B-SHAPE FORMATION (INTENDED FORMATION) WAS OBSERVED, BUT SLIGHTLY AFTER STAPLING WAS PERFORMED, THE STAPLE LINE OPENED BACK UP, WHICH RESULTED IN A NEED TO EXCISE ADDITIONAL COLON TISSUE THAT WAS NOT PLANNED TO BE EXCISED. THE THIRD RELOAD STAPLED AS EXPECTED/INTENDED AND THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE DA VINCI SURGICAL SYSTEM. THE INITIAL REPORTER INDICATED THAT THE PATIENT DID NOT EXPERIENCE ANY INTRA-OPERATIVE/POST-OPERATIVE COMPLICATIONS AND THAT THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE DA VINCI SURGICAL SYSTEM, INSTRUMENTS, AND ACCESSORIES ONCE THE THIRD RELOAD FIRE WAS SUCCESSFUL. PER FOLLOW-UP WITH THE INITIAL REPORTER, THE SURGEON INDICATED THAT THIS EVENT DID NOT IMPACT THE PROCEDURE AS A WHOLE AND THAT THERE WAS NO EFFECT TO THE PATIENT¿S HEALTH. THE SURGEON INDICATED THAT THERE WAS NO CONCERN REGARDING THIS EVENT POTENTIALLY CAUSING LONG-TERM COMPLICATIONS FOR THE PATIENT. TRANSECTION WAS REPORTED TO HAVE BEEN NORMAL AND THE STAPLES WERE NOTED TO HAVE FORMED NORMALLY INTO A B-SHAPE FORMATION BEFORE THE STAPLE LINE OPENED. THERE WAS NO REPORT OF TISSUE TEARING, BUT ALSO NO CONFIRMATION THAT THE STAPLES BECAME UNFORMED AFTER INITIALLY FORMING. AT THIS TIME, THE EXACT WAY IN WHICH THE STAPLE LINE ¿OPENED-UP¿ REMAINS UNKNOWN. THE SAME ENDOWRIST STAPLER 45 INSTRUMENT WAS USED FOR ALL RELOAD FIRES AND PERFORMED AS EXPECTED. THERE WERE NO ERROR MESSAGES. IT WAS REPORTED THAT THIS BEHAVIOR OF THE STAPLE LINE OPENING UP WAS OBSERVED AFTER BACK-TO-BACK FIRES AND ONLY SHORTLY AFTER BEGINNING THE PROCEDURE. THERE WAS NO REPORT OF TISSUE BUNCHING AND NO REPORTS OF STAPLES DEPLOYING AT AN INCORRECT TIME. IT WAS CONFIRMED THAT ALL STAPLES DEPLOYED AND NO STAPLES WERE PRESENT IN THE RELOAD CARTRIDGE FOLLOWING STAPLING. THERE WERE NO ISSUES WITH THE INSTRUMENT/RELOAD STICKING TO TISSUE. THE SURGEON DID NOT ENCOUNTER OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. NO BUTTRESS MATERIAL WAS USED. THERE WERE NO MALFORMED STAPLES, THE STAPLES APPEARED TO HAVE BEEN DEPLOYED NORMALLY AND AS INTENDED BEFORE THEY ALLEGEDLY ¿OPENED-UP.¿ THERE WERE NO CLAMPING ISSUES NOTED. THE TISSUE WAS NOT CALCIFIED AND WAS NOT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY PRIOR TO THE PROCEDURE. THERE WAS NO TISSUE TENSION AND NO TISSUE BUNCHING. THIS MEDICAL DEVICE REPORT (MDR) CAPTURES THE REPORT OF THE FIRST RELOAD THAT EXPERIENCED THE ALLEGED STAPLING ISSUE DURING THE REPORTED EVENT. THE SECOND RELOAD IS REPORTED UNDER MDR WITH PATIENT IDENTIFIER #(B)(4).
05-01-2022 | Malfunction | ethicon | 11-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER | Mechanical Jam | A0506 | Insufficient Information | Event Description: STAPLER USED IN LAPAROSCOPIC APPENDECTOMY FAILED, JAMMED WHEN THE TRIGGER WAS PULLED. FDA SAFETY REPORT ID# (B)(4).
05-01-2022 | Malfunction | ethicon | 11-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER | Therapeutic or Diagnostic Output Failure | A090406 | Insufficient Information | Event Description: STAPLER USED IN LAPAROSCOPIC APPENDECTOMY FAILED, JAMMED WHEN THE TRIGGER WAS PULLED. FDA SAFETY REPORT ID# (B)(4).
17-12-2021 | Injury | covidien lp llc north haven | 12-01-2022 | EEA ORVIL | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A TOTAL GASTRECTOMY VIA OPEN PROCEDURE, DURING THE ESOPHAGUS TO BOWEL ANASTOMOSIS. THE TUBING WAS DIFFICULT TO REMOVE FROM THE ANVIL AFTER ONE OF THE TWO SUTURES WAS CUT. THE SURGEON THEN CUT THE SECOND SUTURE TO REMOVE THE DELIVERY TUBE. SOME SUTURES WERE LEFT IN THE TISSUE AND WERE REMOVED WITH FORCEPS. THE PATIENT DEVELOPED LEAKAGE AT THE ANASTOMOTIC SITE AS EVIDENCED BY FLUID IN THE DRAIN. THE CAUSE OF LEAK WAS UNKNOWN. A STENT WAS PLACED TO MITIGATE THE LEAK. IT WAS STATED THAT A FEW DAYS AFTER THE PROCEDURE, THE SCOPE SHOWED THAT A SUTURE WAS STILL VISIBLE.
17-12-2021 | Injury | covidien lp llc north haven | 12-01-2022 | EEA ORVIL | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A TOTAL GASTRECTOMY VIA OPEN PROCEDURE, DURING THE ESOPHAGUS TO BOWEL ANASTOMOSIS. THE TUBING WAS DIFFICULT TO REMOVE FROM THE ANVIL AFTER ONE OF THE TWO SUTURES WAS CUT. THE SURGEON THEN CUT THE SECOND SUTURE TO REMOVE THE DELIVERY TUBE. SOME SUTURES WERE LEFT IN THE TISSUE AND WERE REMOVED WITH FORCEPS. THE PATIENT DEVELOPED LEAKAGE AT THE ANASTOMOTIC SITE AS EVIDENCED BY FLUID IN THE DRAIN. THE CAUSE OF LEAK WAS UNKNOWN. A STENT WAS PLACED TO MITIGATE THE LEAK. IT WAS STATED THAT A FEW DAYS AFTER THE PROCEDURE, THE SCOPE SHOWED THAT A SUTURE WAS STILL VISIBLE.
29-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. STOPPED OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
18-06-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | PROXIMATE | Misfire | A050502 | Insufficient Information | Event Description: MD WAS USING THE TLC75 75MM LINEAR CUTTER, DURING THE 3RD USE, IT MISFIRED, BLADE WOULD NOT ADVANCE DURING MID-DEPLOYMENT. THE STAPLES WERE NOT FIRED, THE STAPLES CAME OUT IN AN "U" SHAPE NOT THE "B" SHAPE. THIS IS A DISPOSABLE, STAPLER, THEREFORE ANOTHER STAPLER WAS OPENED AND NOT CHARGED TO PATIENT. THIS PARTICULAR STAPLER: LOT#U40T59, #TLC75. TYPE OF SURGERY/PROCEDURE SCHEDULED: ILEOSTOMY TAKEDOWN.
18-06-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | PROXIMATE | Material Deformation | A0406 | Insufficient Information | Event Description: MD WAS USING THE TLC75 75MM LINEAR CUTTER, DURING THE 3RD USE, IT MISFIRED, BLADE WOULD NOT ADVANCE DURING MID-DEPLOYMENT. THE STAPLES WERE NOT FIRED, THE STAPLES CAME OUT IN AN "U" SHAPE NOT THE "B" SHAPE. THIS IS A DISPOSABLE, STAPLER, THEREFORE ANOTHER STAPLER WAS OPENED AND NOT CHARGED TO PATIENT. THIS PARTICULAR STAPLER: LOT#U40T59, #TLC75. TYPE OF SURGERY/PROCEDURE SCHEDULED: ILEOSTOMY TAKEDOWN.
08-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | 75MM SELECTABLE NEW TLC | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ABDOMINAL SURGERY THERE WAS AN ISSUE WITH LOADING THE CARTRIDGE IN ADDITION TO A FIRING ISSUE. NEITHER LEFT NOR RIGHT COULD BE PROPERLY MANIPULATED. THE FIRING ISSUE WAS THAT THE KNIFE GOT STUCK IN THE RELOAD, COULDN¿T BE RETURNED. THERE WERE NO PATIENT CONSEQUENCES OR CHANGE IN POST-OPERATIVE CARE MENTIONED.
08-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ABDOMINAL SURGERY THERE WAS AN ISSUE WITH LOADING THE CARTRIDGE IN ADDITION TO A FIRING ISSUE. NEITHER LEFT NOR RIGHT COULD BE PROPERLY MANIPULATED. THE FIRING ISSUE WAS THAT THE KNIFE GOT STUCK IN THE RELOAD, COULDN¿T BE RETURNED. THERE WERE NO PATIENT CONSEQUENCES OR CHANGE IN POST-OPERATIVE CARE MENTIONED.
12-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2022 | UNKNOWN STAPLING DEVICE | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE.
12-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2022 | UNKNOWN EGIA ADAPTER | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE.
06-04-2022 | Malfunction | us surgical puerto rico | 12-01-2022 | TA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE CARTRIDGE DID NOT FIRMLY CLOSE FOR STAPLING THE TISSUE.
27-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG WEDGE RESECTION SURGERY, WHILE REMOVING THE LUNG TISSUE, THE DEVICE DID NOT FIRE FOR BOTH RELOADS. THE SURGEON WAS ABLE TO PRESS THE GREEN BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE SURGEON USED A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG WEDGE RESECTION SURGERY, WHILE REMOVING THE LUNG TISSUE, THE DEVICE DID NOT FIRE FOR BOTH RELOADS. THE SURGEON WAS ABLE TO PRESS THE GREEN BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. THE SURGEON USED A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC, THE STAPLE LINE DID NOT HOLD AT INITIAL FIRING. WHEN THE SURGEON FIRED AGAIN, IT MADE A STRANGE NOISE AND CANNOT BE FIRED .THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC, THE STAPLE LINE DID NOT HOLD AT INITIAL FIRING. WHEN THE SURGEON FIRED AGAIN, IT MADE A STRANGE NOISE AND CANNOT BE FIRED .THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC, THE STAPLE LINE DID NOT HOLD AT INITIAL FIRING. WHEN THE SURGEON FIRED AGAIN, IT MADE A STRANGE NOISE AND CANNOT BE FIRED .THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | GOLD RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY CASE WHILE USING GOLD RELOAD WITH POWERED ECHELON STAPLER THE LATERAL SIDE OF THE RELOAD DIDN'T APPLY THE STAPLES AND THE RELOAD ITSELF WAS DAMAGED FROM THE LATERAL SURFACE, WITH NO EFFECT ON THE MEDIAL SIDE OF THE RELOAD, AND THE DOCTOR CONFIRMED THERE ARE NO PATIENT CONSEQUENCES.
01-01-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | GOLD RELOAD FOR ECHELON 60 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY CASE WHILE USING GOLD RELOAD WITH POWERED ECHELON STAPLER THE LATERAL SIDE OF THE RELOAD DIDN'T APPLY THE STAPLES AND THE RELOAD ITSELF WAS DAMAGED FROM THE LATERAL SURFACE, WITH NO EFFECT ON THE MEDIAL SIDE OF THE RELOAD, AND THE DOCTOR CONFIRMED THERE ARE NO PATIENT CONSEQUENCES.
22-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SIGMOID RESECTION PROCEDURE, WHILE THE BOWEL WAS BEING ANASTOMOSED, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS OPENED AND FIRED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PNEUMONECTOMY, ON PULMONARY PARENCHYMA DISSECTION, AFTER PRESSING THE GREEN BUTTON, AN ATTEMPT WAS MADE TO FIRE, BUT FAILED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | UNKNOWN ENDO GIA SULU | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN LAPAROTOMY DISTAL GASTRECTOMY, AN UNUSUAL GEAR SOUND WAS HEARD DURING THE FIRST AND SECOND FIRING, BUT THE DEVICE WAS FIRED WITHOUT ANY PROBLEM. AFTER THE SECOND FIRING (THE RELOAD WAS AT ABOUT 30 DEGREES ANGLE), THE RELOAD COULD NOT BE RETURNED TO NEUTRAL POSITION AND THE RELOAD COULD NOT BE REMOVED. POWER APPROACH WAS PERFORMED, AND THE DEVICE WAS RESTARTED., THE RELOAD WAS ABLE TO RETURN TO NEUTRAL POSITION AND WAS REMOVED. THERE WAS NO PATIENT INJURY.
15-09-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, DURING DEVICE MAINTENANCE, THE HANDLE WOULD NOT TURN ON AFTER BEING CHARGED. OTHER HANDLES WERE TESTED ON THE CHARGER WITH NO ISSUES. THERE WAS NO PATIENT INVOLVEMENT.
15-09-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, DURING DEVICE MAINTENANCE, THE HANDLE WOULD NOT TURN ON AFTER BEING CHARGED. OTHER HANDLES WERE TESTED ON THE CHARGER WITH NO ISSUES. THERE WAS NO PATIENT INVOLVEMENT.
01-01-2021 | Injury | ethicon endo-surgery | 12-01-2022 | CIRCULAR MECH STD 29MM | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT THERE WAS 2 BACK-TO- BACK MISFIRES OF CDH25B AND CDH29B CIRCULAR STAPLERS. RESIDENT ATTEMPTED FIRST FIRING. UPON REMOVAL 2 FULL ROTATIONS, TORE ANASTOMOSIS. ENOUGH TISSUE REMAINED TO ATTEMPT A SECOND ANASTOMOSIS. ATTENDING FIRED THE SECOND STAPLER - USED A DIFFERENT SIZE. CLOSED TILL FINGERTIP TIGHT IN THE GREEN ZONE. FIRED, OPENED 2 FULL ROTATIONS AND UPON 90- DEGREE ROTATION TO BEGIN REMOVAL TORE TISSUE. WILL ATTEMPT TO HAND SEW THE TORN TISSUE OF THE ANASTOMOSIS. PATIENT REQUIRED OSTOMY. SURGERY WAS DELAYED BY 20-MINUTES. PROCEDURE: UNKNOWN
01-01-2021 | Injury | ethicon endo-surgery | 12-01-2022 | CIRCULAR MECH STD 29MM | Ejection Problem | A0503 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT THERE WAS 2 BACK-TO- BACK MISFIRES OF CDH25B AND CDH29B CIRCULAR STAPLERS. RESIDENT ATTEMPTED FIRST FIRING. UPON REMOVAL 2 FULL ROTATIONS, TORE ANASTOMOSIS. ENOUGH TISSUE REMAINED TO ATTEMPT A SECOND ANASTOMOSIS. ATTENDING FIRED THE SECOND STAPLER - USED A DIFFERENT SIZE. CLOSED TILL FINGERTIP TIGHT IN THE GREEN ZONE. FIRED, OPENED 2 FULL ROTATIONS AND UPON 90- DEGREE ROTATION TO BEGIN REMOVAL TORE TISSUE. WILL ATTEMPT TO HAND SEW THE TORN TISSUE OF THE ANASTOMOSIS. PATIENT REQUIRED OSTOMY. SURGERY WAS DELAYED BY 20-MINUTES. PROCEDURE: UNKNOWN
09-01-2022 | Malfunction | us surgical puerto rico | 12-01-2022 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
27-09-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN CHARGING, THE BATTERY CHARGER FLASHED RED AND THE HANDLE COULD NOT BE CHARGED. A SIMILAR EVENT OCCURRED WHEN ANOTHER CHARGER WAS USED TO TRY CHARGING THE HANDLE AGAIN. THE CHARGERS WORKED FINE WITH OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT.
27-09-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN CHARGING, THE BATTERY CHARGER FLASHED RED AND THE HANDLE COULD NOT BE CHARGED. A SIMILAR EVENT OCCURRED WHEN ANOTHER CHARGER WAS USED TO TRY CHARGING THE HANDLE AGAIN. THE CHARGERS WORKED FINE WITH OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT.
31-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. THEN CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. THEN CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2021 | Injury | ethicon endo-surgery | 12-01-2022 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED, ANASTOMOTIC LEAK. THE RADIOLOGIST HIGHLIGHTED BADLY FORMED POINTS. PROCEDURE: LOW ANTERIOR RESECTION.
24-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
24-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE.
23-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE CDH25P DID NOT FUNCTION AT THE 1ST FIRING. THE DEVICE WAS REPLACED TO THE CDH29P AND THE BATTERY WAS REPLACED BUT IT DID NOT FUNCTION AS WELL. THE DEVICE WAS USED BETWEEN THE COLON AND THE RECTUM. ANOTHER SECOND DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE CDH25P DID NOT FUNCTION AT THE 1ST FIRING. THE DEVICE WAS REPLACED TO THE CDH29P AND THE BATTERY WAS REPLACED BUT IT DID NOT FUNCTION AS WELL. THE DEVICE WAS USED BETWEEN THE COLON AND THE RECTUM. ANOTHER SECOND DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
03-03-2022 | Malfunction | covidien lp llc north haven | 12-01-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE.
27-12-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN THE USER ATTACHED THE CARTRIDGE TO THE HANDLE, WHILE THE DEVICE WAS PLACED ON THE TISSUE IN THE CARTRIDGE JAWS AND THE OPERATOR SQUEEZED THE HANDLE TO CLOSE THE JAWS, THE SURGEON PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE, BUT FOUND THE SITUATION WHERE IT COULD NOT SQUEEZE FIRMLY AS IF IT WAS LOCKED AND DID NOT FIRE. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW PRODUCT WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-12-2021 | Injury | ethicon endo-surgery | 12-01-2022 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED, ANASTOMOTIC LEAK. THE RADIOLOGIST HIGHLIGHTED BADLY FORMED POINTS. PROCEDURE: LOW ANTERIOR RESECTION.
01-12-2021 | Injury | ethicon endo-surgery | 12-01-2022 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: IT WAS REPORTED, ANASTOMOTIC LEAK. THE RADIOLOGIST HIGHLIGHTED BADLY FORMED POINTS. PROCEDURE: LOW ANTERIOR RESECTION.
27-12-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE PROCEDURE THE DEVICE DID NOT FIRE. ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZED THE HANDLE BUT THE IT DID NOT DEPLOY. THERE WAS NO PATIENT INVOLVED IN THE EVENT.
26-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | GOLD RELOAD FOR ECHELON 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: THIS CASE IS FROM THE HEALTH AUTHORITY. IT WAS REPORTED THAT DURING THE SURGERY, WHEN TO USE, COULD NOT OPEN THE JAW, WHICH COULD NOT BE USED. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
27-12-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN REMOVING THE AFFECTED TISSUE DURING A LAPAROSCOPIC PNEUMOTHORAX, DEVICE OPENED AND CLOSED OUTSIDE THE PATIENT WHEN IT WAS INSERTED THROUGH THE TROCAR, IT NO LONGER OPENED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ESOPHAGEAL SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-10-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, WHILE THE TISSUE WAS BEING CUT, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN AFTER PRESSING THE GREEN BUTTON. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2021 | Injury | ethicon endo-surgery | 12-01-2022 | CIRCULAR MECH STD 25MM | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT THERE WAS 2 BACK-TO- BACK MISFIRES OF CDH25B AND CDH29B CIRCULAR STAPLERS. RESIDENT ATTEMPTED FIRST FIRING. UPON REMOVAL 2 FULL ROTATIONS, TORE ANASTOMOSIS. ENOUGH TISSUE REMAINED TO ATTEMPT A SECOND ANASTOMOSIS. ATTENDING FIRED THE SECOND STAPLER - USED A DIFFERENT SIZE. CLOSED TILL FINGERTIP TIGHT IN THE GREEN ZONE. FIRED, OPENED 2 FULL ROTATIONS AND UPON 90- DEGREE ROTATION TO BEGIN REMOVAL TORE TISSUE. WILL ATTEMPT TO HAND SEW THE TORN TISSUE OF THE ANASTOMOSIS. PATIENT REQUIRED OSTOMY. SURGERY WAS DELAYED BY 20-MINUTES. PROCEDURE: UNKNOWN.
01-01-2021 | Injury | ethicon endo-surgery | 12-01-2022 | CIRCULAR MECH STD 25MM | Ejection Problem | A0503 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT THERE WAS 2 BACK-TO- BACK MISFIRES OF CDH25B AND CDH29B CIRCULAR STAPLERS. RESIDENT ATTEMPTED FIRST FIRING. UPON REMOVAL 2 FULL ROTATIONS, TORE ANASTOMOSIS. ENOUGH TISSUE REMAINED TO ATTEMPT A SECOND ANASTOMOSIS. ATTENDING FIRED THE SECOND STAPLER - USED A DIFFERENT SIZE. CLOSED TILL FINGERTIP TIGHT IN THE GREEN ZONE. FIRED, OPENED 2 FULL ROTATIONS AND UPON 90- DEGREE ROTATION TO BEGIN REMOVAL TORE TISSUE. WILL ATTEMPT TO HAND SEW THE TORN TISSUE OF THE ANASTOMOSIS. PATIENT REQUIRED OSTOMY. SURGERY WAS DELAYED BY 20-MINUTES. PROCEDURE: UNKNOWN.
17-12-2021 | Injury | us surgical puerto rico | 12-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A TOTAL GASTRECTOMY VIA OPEN PROCEDURE, DURING THE ESOPHAGUS TO BOWEL ANASTOMOSIS. THE TUBING WAS DIFFICULT TO REMOVE FROM THE ANVIL AFTER ONE OF THE TWO SUTURES WAS CUT. THE SURGEON THEN CUT THE SECOND SUTURE TO REMOVE THE DELIVERY TUBE. SOME SUTURES WERE LEFT IN THE TISSUE AND WERE REMOVED WITH FORCEPS. THE PATIENT DEVELOPED LEAKAGE AT THE ANASTOMOTIC SITE AS EVIDENCED BY FLUID IN THE DRAIN. THE CAUSE OF LEAK IS UNKNOWN. A STENT WAS PLACED TO MITIGATE THE LEAK. IT WAS STATED THAT A FEW DAYS AFTER THE PROCEDURE, THE SCOPE SHOWED THAT A SUTURE WAS STILL VISIBLE.
27-12-2021 | Malfunction | covidien lp llc north haven | 12-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE DEVICE DID NOT FIRE. ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZED THE HANDLE BUT IT DID NOT DEPLOY. THERE WAS NO PATIENT INVOLVED IN THE EVENT.
27-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Injury | celestica de monterrey s.a de c.v | 12-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT FORM PROPERLY, AND THE STAPLE LINE WAS INCOMPLETE IN THE DISTAL. SUTURE REINFORCEMENT WAS DONE TO FIX THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Injury | celestica de monterrey s.a de c.v | 12-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT FORM PROPERLY, AND THE STAPLE LINE WAS INCOMPLETE IN THE DISTAL. SUTURE REINFORCEMENT WAS DONE TO FIX THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Injury | celestica de monterrey s.a de c.v | 12-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT FORM PROPERLY, AND THE STAPLE LINE WAS INCOMPLETE IN THE DISTAL. SUTURE REINFORCEMENT WAS DONE TO FIX THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Injury | celestica de monterrey s.a de c.v | 12-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT FORM PROPERLY, AND THE STAPLE LINE WAS INCOMPLETE IN THE DISTAL. SUTURE REINFORCEMENT WAS DONE TO FIX THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | us surgical puerto rico | 12-01-2022 | EEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE WHITE RING ABOVE THE STAPLER HEAD WAS DETACHED. THERE WAS NO PATIENT INVOLVEMENT.
08-02-2021 | Injury | us surgical puerto rico | 12-01-2022 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN JANUARY 2012 AND JANUARY 2020. A RETROSPECTIVE STUDY EVALUATED OUTCOMES OF 84 PATIENTS WHO UNDERWENT CYTOREDUCTIVE SURGERY (CRS) WITH HYPERTHERMIC INTRA-PERITONEAL CHEMOTHERAPY (HIPEC) AND RECTAL ANASTOMOSIS, WITH AND WITHOUT A DIVERTING LOOP ILEOSTOMY. THE OBJECTIVE OF THIS STUDY WAS TO COMPARE THE POSTOPERATIVE MORTALITY AND MORBIDITY RATES OF PATIENTS UNDERGOING CRS WITH HIPEC AND RECTAL ANASTOMOSIS, WITH AND WITHOUT A DIVERTING LOOP ILEOSTOMY, IN ORDER TO IDENTIFY WHICH PATIENTS MAY BENEFIT FROM SUCH A PREVENTIVE STRATEGY. THERE WERE 29 PATIENTS UNDERGOING THE ELECTIVE DIVERTING LOOP ILEOSTOMY. THE STAPLER WAS USED FOR RECTAL ANASTOMOSIS WITH THE FOLLOWING REPORTED COMPLICATIONS: 7 RECTAL ANASTOMOTIC LEAKS: 3 REQUIRING INTRAVENOUS ANTIBIOTICS FOR PELVIC ABSCESS, 1 REQUIRING PERCUTANEOUS DRAINAGE AND 3 REQUIRING SURGICAL CREATION OF PERMANENT COLOSTOMY. HOSPITAL READMISSIONS WERE ALSO REPORTED. COMPLICATIONS UNRELATED TO THE STAPLER INCLUDED ANASTOMOTIC LEAKS ASSOCIATED WITH THE REVERSAL OF THE DIVERTING LOOP ILEOSTOMY.
08-02-2021 | Injury | us surgical puerto rico | 12-01-2022 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN JANUARY 2012 AND JANUARY 2020. A RETROSPECTIVE STUDY EVALUATED OUTCOMES OF 84 PATIENTS WHO UNDERWENT CYTOREDUCTIVE SURGERY (CRS) WITH HYPERTHERMIC INTRA-PERITONEAL CHEMOTHERAPY (HIPEC) AND RECTAL ANASTOMOSIS, WITH AND WITHOUT A DIVERTING LOOP ILEOSTOMY. THE OBJECTIVE OF THIS STUDY WAS TO COMPARE THE POSTOPERATIVE MORTALITY AND MORBIDITY RATES OF PATIENTS UNDERGOING CRS WITH HIPEC AND RECTAL ANASTOMOSIS, WITH AND WITHOUT A DIVERTING LOOP ILEOSTOMY, IN ORDER TO IDENTIFY WHICH PATIENTS MAY BENEFIT FROM SUCH A PREVENTIVE STRATEGY. THERE WERE 29 PATIENTS UNDERGOING THE ELECTIVE DIVERTING LOOP ILEOSTOMY. THE STAPLER WAS USED FOR RECTAL ANASTOMOSIS WITH THE FOLLOWING REPORTED COMPLICATIONS: 7 RECTAL ANASTOMOTIC LEAKS: 3 REQUIRING INTRAVENOUS ANTIBIOTICS FOR PELVIC ABSCESS, 1 REQUIRING PERCUTANEOUS DRAINAGE AND 3 REQUIRING SURGICAL CREATION OF PERMANENT COLOSTOMY. HOSPITAL READMISSIONS WERE ALSO REPORTED. COMPLICATIONS UNRELATED TO THE STAPLER INCLUDED ANASTOMOTIC LEAKS ASSOCIATED WITH THE REVERSAL OF THE DIVERTING LOOP ILEOSTOMY.
04-11-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. CHANGED ANOTHER ONE TO FIRE AGAIN, THE SAME PROBLEM HAPPENED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE DEVICE COULD NOT BE FIRED ALTHOUGH THE TISSUE COMPRESSION SCALE BACKLIGHT WAS ILLUMINATED. THE DEVICE WAS USED ON THE RECTUM. THE ANVIL IN QUESTION WAS PLACED AS IT IS AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
04-11-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: THIS CASE IS FROM THE HEALTH AUTHORITY. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. CHANGED ANOTHER ONE TO FIRE AGAIN, THE SAME PROBLEM HAPPENED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-12-2021 | Injury | intuitive surgical | 12-01-2022 | SUREFORM | Patient Device Interaction Problem | A01 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY PROCEDURE, A PARTIAL STAPLER FIRE ALLEGEDLY OCCURRED INVOLVING A SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT. THE EVENT OCCURRED ON THE FIFTH STAPLER FIRE OF A GREEN STAPLER RELOAD WHILE THE SURGEON WAS STAPLING THE BRONCHUS STUMP. ACCORDING TO THE SURGEON, THE INSTRUMENT WAS VISUALLY INSPECTED PRIOR TO USE WITH NO ABNORMALITY FOUND. THE SURGEON SELECTED THE GREEN RELOAD BECAUSE IT WAS SUITABLE. THERE WERE NO CLAMPING ISSUES. THE MESSAGE, ¿PAUSING FOR COMPRESSIONS,¿ APPEARED 3 TIMES. IN ADDITION, A ¿TISSUE TOO THICK TO CONTINUE¿ MESSAGE AND AN UNCLAMPING RECOMMENDATION WERE DISPLAYED DURING THE FOURTH PAUSE. THERE WERE ALSO MALFORMED STAPLES SEEN. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. NO BUTTRESS MATERIAL FOR REINFORCING WAS USED. THE TARGET TISSUE WAS NON-CALCIFIED, AND NOT EXPOSED TO CHEMOTHERAPY OR RADIATION. THERE WAS NO TISSUE TENSION AND TISSUE BUNCHING. THE SURGEON RESOLVED THE ISSUE BY SUTURING THE BRONCHIAL STUMP.
13-12-2021 | Injury | intuitive surgical | 12-01-2022 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY PROCEDURE, A PARTIAL STAPLER FIRE ALLEGEDLY OCCURRED INVOLVING A SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT. THE EVENT OCCURRED ON THE FIFTH STAPLER FIRE OF A GREEN STAPLER RELOAD WHILE THE SURGEON WAS STAPLING THE BRONCHUS STUMP. ACCORDING TO THE SURGEON, THE INSTRUMENT WAS VISUALLY INSPECTED PRIOR TO USE WITH NO ABNORMALITY FOUND. THE SURGEON SELECTED THE GREEN RELOAD BECAUSE IT WAS SUITABLE. THERE WERE NO CLAMPING ISSUES. THE MESSAGE, ¿PAUSING FOR COMPRESSIONS,¿ APPEARED 3 TIMES. IN ADDITION, A ¿TISSUE TOO THICK TO CONTINUE¿ MESSAGE AND AN UNCLAMPING RECOMMENDATION WERE DISPLAYED DURING THE FOURTH PAUSE. THERE WERE ALSO MALFORMED STAPLES SEEN. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. NO BUTTRESS MATERIAL FOR REINFORCING WAS USED. THE TARGET TISSUE WAS NON-CALCIFIED, AND NOT EXPOSED TO CHEMOTHERAPY OR RADIATION. THERE WAS NO TISSUE TENSION AND TISSUE BUNCHING. THE SURGEON RESOLVED THE ISSUE BY SUTURING THE BRONCHIAL STUMP.
08-04-2021 | Malfunction | ethicon endo-surgery | 12-01-2022 | GOLD RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-11-2021 | Injury | ethicon endo-surgery | 13-01-2022 | CIRCULAR MECH XL SEAL 29MM | Failure to Cut | A050702 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A HARTMAN'S PROCEDURE, LEAK DURING THE COLON ANASTOMOSIS. THE DONUT WAS NOT COMPLETE. IT WAS NOTICED ON THE DISTAL END. CASE COMPLETED BY OVER SEWING.
30-11-2021 | Injury | ethicon endo-surgery | 13-01-2022 | CIRCULAR MECH XL SEAL 29MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A HARTMAN'S PROCEDURE, LEAK DURING THE COLON ANASTOMOSIS. THE DONUT WAS NOT COMPLETE. IT WAS NOTICED ON THE DISTAL END. CASE COMPLETED BY OVER SEWING.
08-04-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | PWRD ECH FLEX 45MM 34CM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT ABNORMAL SOUND WAS DETECTED WHEN USING THE DEVICE DURING PROCEDURE. AFTER FIRED, THE DEVICE FELL APART. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
08-04-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | PWRD ECH FLEX 45MM 34CM SHAFT | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT ABNORMAL SOUND WAS DETECTED WHEN USING THE DEVICE DURING PROCEDURE. AFTER FIRED, THE DEVICE FELL APART. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
30-12-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE JAW DID NOT OPEN AFTER FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
28-12-2021 | Malfunction | covidien lp llc north haven | 13-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE LOAD OF THE STAPLES DID NOT FIRED WHEN IT WAS TRIGGERED. A NEW RELOAD OF THE SAME CODE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY,AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-01-2022 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY,AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-01-2022 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY,AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-12-2021 | Malfunction | us surgical puerto rico | 13-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATECTOMY, WHEN THE TISSUE OF THE PANCREAS WAS REMOVED, THE STAPLER AND 2 CONSECUTIVE RELOADS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-11-2021 | Malfunction | covidien lp llc north haven | 13-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC BYPASS PROCEDURE, THE FIRST RELOAD WAS CONNECTED AND TESTED, BUT THE JAWS DID NOT CLOSE COMPLETELY. ANOTHER RELOAD WAS TAKEN AND USED, BUT SAME ISSUE OCCURRED. A THIRD RELOAD WAS TAKEN OUT AND WAS USED WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE, ON THE INNER ROW OF THE LEFT CUT LINE ON THE APPROXIMAL THIRD, A MALFORMED. THIS HAPPENED ON ALL SIX FIRINGS. CASE COMPLETED BY OVER SEWING THE ENTIRE STAPLE LINE. A FIBRIN SEALANT WAS ALSO USED. NO PATIENT HARM.
29-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, WHILE TRANSECTING THE SMALL INTESTINE, BLEEDING OCCURRED ON THE RIGHT SIDE OF THE STAPLE LINE. BLEEDING CONTROL WAS PERFORMED USING A SUTURE.
29-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, WHILE TRANSECTING THE SMALL INTESTINE, BLEEDING OCCURRED ON THE RIGHT SIDE OF THE STAPLE LINE. BLEEDING CONTROL WAS PERFORMED USING A SUTURE.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | TRI-STAPLE 2.0 | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE TO A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE TO A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE TO A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
28-12-2021 | Malfunction | us surgical puerto rico | 13-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATECTOMY, WHEN THE TISSUE OF THE PANCREAS WAS REMOVED, THE STAPLER AND 2 CONSECUTIVE RELOADS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE POWER HANDLE AND 2 RELOADS WERE ABLE TO FIRE, THE PATIENT HAD BLOOD LOSS OF MORE THAN 500CC. BLOOD TRANSFUSION WAS DONE TO RESOLVE THE ISSUE.
18-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Syncope/Fainting | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE POWER HANDLE AND 2 RELOADS WERE ABLE TO FIRE, THE PATIENT HAD BLOOD LOSS OF MORE THAN 500CC. BLOOD TRANSFUSION WAS DONE TO RESOLVE THE ISSUE.
18-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE POWER HANDLE AND 2 RELOADS WERE ABLE TO FIRE, THE PATIENT HAD BLOOD LOSS OF MORE THAN 500CC. BLOOD TRANSFUSION WAS DONE TO RESOLVE THE ISSUE.
18-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE POWER HANDLE AND 2 RELOADS WERE ABLE TO FIRE, THE PATIENT HAD BLOOD LOSS OF MORE THAN 500CC. BLOOD TRANSFUSION WAS DONE TO RESOLVE THE ISSUE.
18-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Syncope/Fainting | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE POWER HANDLE AND 2 RELOADS WERE ABLE TO FIRE, THE PATIENT HAD BLOOD LOSS OF MORE THAN 500CC. BLOOD TRANSFUSION WAS DONE TO RESOLVE THE ISSUE.
18-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE POWER HANDLE AND 2 RELOADS WERE ABLE TO FIRE, THE PATIENT HAD BLOOD LOSS OF MORE THAN 500CC. BLOOD TRANSFUSION WAS DONE TO RESOLVE THE ISSUE.
04-01-2022 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS THAT DURING A LAPAROSCOPIC APPENDECTOMY ONCE PLACE THROUGH THE TROCAR THE DEVICE WOULD NOT OPEN. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE TO A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE TO A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE TO A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE TO A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | SIGNIA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE TO A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE TO A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
19-04-2022 | Malfunction | us surgical puerto rico | 13-01-2022 | GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE CARTRIDGE DID NOT CLOSE FIRMLY AND FIRED FOR STAPLING THE TISSUE.
25-12-2021 | Malfunction | covidien lp llc north haven | 13-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN EMERGENCY APPENDECTOMY PROCEDURE, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
28-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | UNKNOWN SIGNIA EGIA ADAPTER | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF LAPAROSCOPIC SLEEVE PROCEDURE, FOR THE FIRST TWO FIRINGS, THE STAPLES WERE MALFORMED AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY AS THE DEVICE STOPPED FIRING HALF WAY THROUGH THE STAPLE LOAD. ANOTHER DEVICE WAS FIRED TO COMPLETE THE CASE. THERE WAS TISSUE LOSS AND PERMANENT TISSUE DAMAGE. AFTER 4 DAYS, THE PATIENT WAS RETURNED TO THE HOSPITAL FOR REOPERATION DUE A LEAK IN THE SAME STAPLE LINE. THE PATIENT¿S HOSPITAL STAY WAS EXTENDED.
10-12-2021 | Injury | covidien lp llc north haven | 13-01-2022 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE SURGERY, THE PATIENT HAD STAPLE LINE BLEEDING. THIS RESULTED TO 500CC BLOOD LOSS. THE PATIENT WAS BROUGHT BACK AND THE SURGEON FIXED THE BLEEDING.
29-12-2021 | Malfunction | us surgical puerto rico | 13-01-2022 | GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RIGHT COLORECTAL SURGERY, WHILE PERFORMING SIDE-TO-SIDE ANASTOMOSIS, TWO CLAMPS COULD BE COMBINED AT THE HINGE PIN, BUT THE HANDLE COULD NOT BE CLOSED, THE JAWS WOULD NOT CLOSE, AND THE DEVICE COULD NOT BE FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-12-2021 | Malfunction | covidien lp llc north haven | 13-01-2022 | EEA | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RECTAL PROLAPSE PROCEDURE, THE ANVIL WAS REPEATEDLY DISENGAGING ITSELF FROM THE HEAD EVEN BEFORE FIRING, AND THE DEVICE WAS DIFFICULT OR UNABLE TO LOAD. THE DOCTOR USED ANOTHER DEVICE FROM A DIFFERENT MANUFACTURER TO COMPLETE THE CASE AND RESOLVE THE ISSUE. IT WAS NOTED THAT WHEN ATTACHING THE ANVIL TO THE INSTRUMENT, THE BLACK TWIST KNOB WAS NOT HELD FIRMLY TO PREVENT THE TROCAR FROM MOVING BACK INTO THE HEAD OF THE DEVICE. THERE WAS NO PATIENT INJURY.
28-12-2021 | Malfunction | covidien lp llc north haven | 13-01-2022 | EEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RECTAL PROLAPSE PROCEDURE, THE ANVIL WAS REPEATEDLY DISENGAGING ITSELF FROM THE HEAD EVEN BEFORE FIRING, AND THE DEVICE WAS DIFFICULT OR UNABLE TO LOAD. THE DOCTOR USED ANOTHER DEVICE FROM A DIFFERENT MANUFACTURER TO COMPLETE THE CASE AND RESOLVE THE ISSUE. IT WAS NOTED THAT WHEN ATTACHING THE ANVIL TO THE INSTRUMENT, THE BLACK TWIST KNOB WAS NOT HELD FIRMLY TO PREVENT THE TROCAR FROM MOVING BACK INTO THE HEAD OF THE DEVICE. THERE WAS NO PATIENT INJURY.
16-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE BLEEDING OCCURRED AT THE STAPLE LINE ON THE STOMACH. THIS RESULTED TO MORE THAN 500CC. THE PATIENT WAS BOUGHT BACK FOR RE-OPERATION AND BLOOD TRANSFUSION WAS ADMINISTERED.
16-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE BLEEDING OCCURRED AT THE STAPLE LINE ON THE STOMACH. THIS RESULTED TO MORE THAN 500CC. THE PATIENT WAS BOUGHT BACK FOR RE-OPERATION AND BLOOD TRANSFUSION WAS ADMINISTERED.
14-09-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO A NEW DEVICE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
16-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE BLEEDING OCCURRED AT THE STAPLE LINE ON THE STOMACH. THIS RESULTED TO MORE THAN 500CC. THE PATIENT WAS BOUGHT BACK FOR RE-OPERATION AND BLOOD TRANSFUSION WAS ADMINISTERED.
16-12-2021 | Injury | us surgical puerto rico | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE BLEEDING OCCURRED AT THE STAPLE LINE ON THE STOMACH. THIS RESULTED TO MORE THAN 500CC. THE PATIENT WAS BOUGHT BACK FOR RE-OPERATION AND BLOOD TRANSFUSION WAS ADMINISTERED.
01-09-2022 | Malfunction | covidien lp llc north haven | 13-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE OPERATIVELY, ON A LAPAROSCOPY, WHEN THE RELOAD WAS CONNECTED TO THE STAPLER, THE USER TRIED TO OPEN AND CLOSE THE UNIT USING THE HANDLE BUT NOTICED THAT IT WAS LOCKED. THE RELOAD DID NOT WORK AND WAS NEVER USED ON THE PATIENT. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-11-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE POWERED STAPLER MISFIRED WHILE WORKING DURING A SLEEVE GASTRECTOMY CASE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
13-01-2022 | Injury | covidien lp llc north haven | 13-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A PROCEDURE, THE PATIENT HAD BLEEDING. THE PATIENT HOSPITALIZATION WAS EXTENDED.
04-01-2022 | Malfunction | ethicon endo-surgery | 13-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SURGERY, NOTED THAT A COMPONENT OF DEVICE FELL OFF FROM THE DEVICE DURING THE 3TH FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | GST GOLD RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLE FORMATION WAS NOT FULLY COMPLETED. THERE WAS ONE SINGLE STAPLE THAT WAS MALFORMED. AFTER CHECKING, THE WOUND DID NOT BLEED.
01-01-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | GST GOLD RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLE FORMATION WAS NOT FULLY COMPLETED. THERE WAS ONE SINGLE STAPLE THAT WAS MALFORMED. AFTER CHECKING, THE WOUND DID NOT BLEED.
01-01-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | GST GOLD RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLE FORMATION WAS NOT FULLY COMPLETED. THERE WAS ONE SINGLE STAPLE THAT WAS MALFORMED. AFTER CHECKING, THE WOUND DID NOT BLEED.
01-01-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | GST GOLD RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT STAPLE FORMATION WAS NOT FULLY COMPLETED. THERE WAS ONE SINGLE STAPLE THAT WAS MALFORMED. AFTER CHECKING, THE WOUND DID NOT BLEED.
20-12-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | FLEX 60 ARTICULATING | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, A NEW STAPLER IS USED, DURING THE FIRST SHOT THE STAPLER IS ACTIVATED NORMALLY, THE SURGEON TRIES TO OPEN THE STAPLER AND THE RELEASE BUTTON DOES NOT WORK. THE JAW MUST BE FORCED TO RELEASE THE TISSUE, FINDING THAT THE STAPLE LINE IS NOT COMPLETE. THE DOCTOR MUST USE A NEW STAPLER AND REFILL TO FINISH THE PROCEDURE. THE PATIENT RECOVERS WITHOUT ANY DIFFICULTY. THE SURGERY WAS NOT DELAYED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITHOUT ANY PATIENT CONSEQUENCE REPORTED.
20-12-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, A NEW STAPLER IS USED, DURING THE FIRST SHOT THE STAPLER IS ACTIVATED NORMALLY, THE SURGEON TRIES TO OPEN THE STAPLER AND THE RELEASE BUTTON DOES NOT WORK. THE JAW MUST BE FORCED TO RELEASE THE TISSUE, FINDING THAT THE STAPLE LINE IS NOT COMPLETE. THE DOCTOR MUST USE A NEW STAPLER AND REFILL TO FINISH THE PROCEDURE. THE PATIENT RECOVERS WITHOUT ANY DIFFICULTY. THE SURGERY WAS NOT DELAYED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITHOUT ANY PATIENT CONSEQUENCE REPORTED.
20-12-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, A NEW STAPLER IS USED, DURING THE FIRST SHOT THE STAPLER IS ACTIVATED NORMALLY, THE SURGEON TRIES TO OPEN THE STAPLER AND THE RELEASE BUTTON DOES NOT WORK. THE JAW MUST BE FORCED TO RELEASE THE TISSUE, FINDING THAT THE STAPLE LINE IS NOT COMPLETE. THE DOCTOR MUST USE A NEW STAPLER AND REFILL TO FINISH THE PROCEDURE. THE PATIENT RECOVERS WITHOUT ANY DIFFICULTY. THE SURGERY WAS NOT DELAYED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITHOUT ANY PATIENT CONSEQUENCE REPORTED.
20-12-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LIVER RESECTION THERE WAS AN UNKNOWN PATIENT CONSEQUENCE.
22-11-2021 | Injury | ethicon endo-surgery | 13-01-2022 | GST GREEN RELOAD, 60MM, 6 ROW | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED VIA CLINICAL TRIAL THE PATIENT EXPERIENCED, POST-OPERATIVE PAIN, RIGHT APICAL PNEUMOTHORAX, AND POST-OPERATIVE DELIRIUM. THE EVENT WAS NOT RELATED TO THE STUDY DEVICE.
22-11-2021 | Injury | ethicon endo-surgery | 13-01-2022 | GST GREEN RELOAD, 60MM, 6 ROW | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED VIA CLINICAL TRIAL THE PATIENT EXPERIENCED, POST-OPERATIVE PAIN, RIGHT APICAL PNEUMOTHORAX, AND POST-OPERATIVE DELIRIUM. THE EVENT WAS NOT RELATED TO THE STUDY DEVICE.
22-11-2021 | Injury | ethicon endo-surgery | 13-01-2022 | GST GREEN RELOAD, 60MM, 6 ROW | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: IT WAS REPORTED VIA CLINICAL TRIAL THE PATIENT EXPERIENCED, POST-OPERATIVE PAIN, RIGHT APICAL PNEUMOTHORAX, AND POST-OPERATIVE DELIRIUM. THE EVENT WAS NOT RELATED TO THE STUDY DEVICE.
21-12-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY, THE KNIFE MOVED SLIGHTLY BUT STOPPED MOVING FORWARD. THE KNIFE REVERSE SWITCH WAS USED TO RETURN THE KNIFE TO HOME POSITION. THE DEVICE WAS USED ON THE BLOOD VESSEL. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-01-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE MISFIRED. NO OTHER INFORMATION IS AVAILABLE. NO PATIENT CONSEQUENCE REPORTED.
01-01-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE MISFIRED. NO OTHER INFORMATION IS AVAILABLE. NO PATIENT CONSEQUENCE REPORTED.
01-01-2021 | Malfunction | ethicon endo-surgery | 13-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE MISFIRED. NO OTHER INFORMATION IS AVAILABLE. NO PATIENT CONSEQUENCE REPORTED.
02-08-2021 | Injury | ethicon endo-surgery | 14-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THREE DAYS POST OP OF A SIGMOID COLON RESECTION PROCEDURE THE PATIENT RETURNED WITH AN ANASTOMOSIS LEAK. DURING THE PROCEDURE SURGEON THOUGHT HE SAW A BUBBLE ON THE FIRST LEAK TEST THEN TESTED AGAIN AND SAW NONE BUT TO BE SAFE VISTASEAL WAS APPLIED TO THE ANASTOMOSIS. POST-OP LEAK WAS REPAIRED DOING A COLOSTOMY.
05-11-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: SURGEON CLOSED STAPLER WITH WHITE LOAD LOADED; FIRED STAPLER; BRIEF BUZZING NOISE THEN STOPPED. STAPLER WOULD NOT OPEN. THE REP WAS AVAILABLE BY PHONE AND INSTRUCTED THE SURGEON TO USE THE REVERSE BUTTON. THE REVERSE BUTTON WOULD NOT MOVE. THE REP INSTRUCTED THE SURGEON TO USE THE MANUAL OVERRIDE AND THEN THE SURGEON WAS ABLE TO OPEN THE STAPLER. THE KNIFE HAD NOT DEPLOYED.
05-11-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX | Mechanical Jam | A0506 | Insufficient Information | Event Description: SURGEON CLOSED STAPLER WITH WHITE LOAD LOADED; FIRED STAPLER; BRIEF BUZZING NOISE THEN STOPPED. STAPLER WOULD NOT OPEN. THE REP WAS AVAILABLE BY PHONE AND INSTRUCTED THE SURGEON TO USE THE REVERSE BUTTON. THE REVERSE BUTTON WOULD NOT MOVE. THE REP INSTRUCTED THE SURGEON TO USE THE MANUAL OVERRIDE AND THEN THE SURGEON WAS ABLE TO OPEN THE STAPLER. THE KNIFE HAD NOT DEPLOYED.
05-11-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX | Activation Failure | A150101 | Insufficient Information | Event Description: SURGEON CLOSED STAPLER WITH WHITE LOAD LOADED; FIRED STAPLER; BRIEF BUZZING NOISE THEN STOPPED. STAPLER WOULD NOT OPEN. THE REP WAS AVAILABLE BY PHONE AND INSTRUCTED THE SURGEON TO USE THE REVERSE BUTTON. THE REVERSE BUTTON WOULD NOT MOVE. THE REP INSTRUCTED THE SURGEON TO USE THE MANUAL OVERRIDE AND THEN THE SURGEON WAS ABLE TO OPEN THE STAPLER. THE KNIFE HAD NOT DEPLOYED.
05-11-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX | Noise, Audible | A0508 | Insufficient Information | Event Description: SURGEON CLOSED STAPLER WITH WHITE LOAD LOADED; FIRED STAPLER; BRIEF BUZZING NOISE THEN STOPPED. STAPLER WOULD NOT OPEN. THE REP WAS AVAILABLE BY PHONE AND INSTRUCTED THE SURGEON TO USE THE REVERSE BUTTON. THE REVERSE BUTTON WOULD NOT MOVE. THE REP INSTRUCTED THE SURGEON TO USE THE MANUAL OVERRIDE AND THEN THE SURGEON WAS ABLE TO OPEN THE STAPLER. THE KNIFE HAD NOT DEPLOYED.
08-10-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX | Battery Problem | A0705 | No Clinical Signs, Symptoms or Conditions | Event Description: DEAD BATTERY.
08-10-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: TRIGGER RELEASE AND EMERGENCY RELEASE DID NOT WORK.
08-10-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX | Fail-Safe Did Not Operate | A160201 | No Clinical Signs, Symptoms or Conditions | Event Description: TRIGGER RELEASE AND EMERGENCY RELEASE DID NOT WORK.
08-11-2021 | Injury | ethicon endo-surgery | 14-01-2022 | ENDO ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST-OP OF A SIGMOID COLON RESECTION PATIENT RETURNED WITH AN ANASTOMOSIS LEAK. TREATMENT OF THE LEAK IS CURRENTLY UNKNOWN.
07-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2022 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOWER ANTERIOR RESECTION, THERE WAS A MISFIRE/FAILURE OF THE STAPLER, IT STAPLED BUT DID NOT CUT THE TISSUE, THEY HAD TO CUT THE COLON AND CUT THE STAPLED PART TO REMOVE THE STAPLER FROM THE RECTUM. THERE WERE NO PATIENT CONSEQUENCES.
07-12-2020 | Malfunction | ethicon endo-surgery | 14-01-2022 | ILS 29MM, CURVED | Failure to Eject | A050301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOWER ANTERIOR RESECTION, THERE WAS A MISFIRE/FAILURE OF THE STAPLER, IT STAPLED BUT DID NOT CUT THE TISSUE, THEY HAD TO CUT THE COLON AND CUT THE STAPLED PART TO REMOVE THE STAPLER FROM THE RECTUM. THERE WERE NO PATIENT CONSEQUENCES.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Failure to Anastomose | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Wound Dehiscence | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Adhesion(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Bacterial Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Cellulitis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Chest Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Cyst(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Purulent Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Emotional Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Fever | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Hematoma | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Muscle Weakness/Atrophy | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Nausea | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Necrosis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Pleural Effusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Pneumonia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Sepsis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Tachycardia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Vomiting | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Hernia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Impaired Healing | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Post Operative Wound Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Weight Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Unspecified Tissue Injury | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Swelling/ Edema | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Leak/Splash | A0504 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Wound Dehiscence | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Adhesion(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Bacterial Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Cellulitis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Chest Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Cyst(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Purulent Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Emotional Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Fever | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Hematoma | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Muscle Weakness/Atrophy | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Nausea | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Necrosis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Pleural Effusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Pneumonia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Sepsis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Tachycardia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Vomiting | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Hernia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Impaired Healing | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Post Operative Wound Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Weight Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Swelling/ Edema | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Difficult to Remove | A150207 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Failure to Anastomose | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Wound Dehiscence | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Adhesion(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Bacterial Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Cellulitis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Chest Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Cyst(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Purulent Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Emotional Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Fever | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Hematoma | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Muscle Weakness/Atrophy | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Nausea | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Necrosis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Pleural Effusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Pneumonia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Sepsis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Tachycardia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Vomiting | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Hernia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Impaired Healing | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Post Operative Wound Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Weight Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Swelling/ Edema | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
19-12-2019 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. DURING A PROCEDURE, THE PATIENT UNDERWENT A SIGMOIDECTOMY COLORECTAL ANASTOMOSIS, THE SURGEON NOTED A PIN POPPED OUT OF PLACE WHEN THE ANVIL WAS PULLED THROUGH THE COLONIC WALL. THE SURGEON WAS UNABLE TO MOBILIZE THE ANVIL AND HAD TO INSTEAD ENTER THROUGH ANOTHER PORT TO REMOVE IT. A SECOND STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. THE SURGEON NOTED A SMALL LEAK PROXIMAL TO THE ANASTOMOSIS WHICH APPEARED DUE TO THE REMOVAL OF THE FIRST STAPLER. THE LEAK WAS REPAIRED WITH SUTURES. THREE DAYS LATER, POST-OPERATIVE, THE PATIENT EXPERIENCED VOMITING, TACHYCARDIC, STOMACH BECAME TIGHT AND DISTENDED, INFECTION OF THE INCISION LINE, CONTAMINATION OF THE DEHISCED ANASTOMOSIS, PHYSICALLY WEAKENED, WEIGHT LOSS, FEVER, PNEUMONIA. THE PATIENT TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, ADDITIONAL MEDICAL PROCEDURES, FIT WITH COLOSTOMY, SURGICAL INTERVENTION, INTUBATED AT VARIOUS POINTS, DRAINAGE, A WOUND VAC.
24-12-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOT-ASSISTED LOW ANTERIOR RESECTION, THE DEVICE DID NOT FUNCTION AT THE FIRING. THE TISSUE COMPRESSION SCALE BACKLIGHT WAS NOT ILLUMINATED. THE DEVICE WAS USED ON THE RECTUM. THE ANVIL IN QUESTION WAS PLACED AS IT IS AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. AFTER THE SURGERY ANOTHER BATTERY WAS INSERTED INTO THE DEVICE IN QUESTION, BUT THE TISSUE COMPRESSION SCALE BACKLIGHT WAS NOT ILLUMINATED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-12-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | PROXIMATE*75MM LINEAR CUTTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE COULD NOT BE OPENED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-01-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | CIRCULAR MECH XL SEAL 25MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE HEAD SURGEON PERFORMED SURGERY ON THE ESOPHAGUS. IN HIS OPINION HE CONNECTED THE CAP OF THE STAPLER CORRECTLY WITH THE STAPLER ITSELF, BUT SURPRISINGLY THE STAPLE LINE WAS INCOMPLETE. JUST THE VENTRAL PART WAS CORRECTLY STAPLED. THE REST OF THE CIRCLE WAS UNTOUCHED, EVEN UNFORMED STAPLES WERE LYING IN THE SITUS. THERE WAS NO DELAY TO THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES.
23-01-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | CIRCULAR MECH XL SEAL 25MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE HEAD SURGEON PERFORMED SURGERY ON THE ESOPHAGUS. IN HIS OPINION HE CONNECTED THE CAP OF THE STAPLER CORRECTLY WITH THE STAPLER ITSELF, BUT SURPRISINGLY THE STAPLE LINE WAS INCOMPLETE. JUST THE VENTRAL PART WAS CORRECTLY STAPLED. THE REST OF THE CIRCLE WAS UNTOUCHED, EVEN UNFORMED STAPLES WERE LYING IN THE SITUS. THERE WAS NO DELAY TO THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES.
23-01-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | CIRCULAR MECH XL SEAL 25MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE HEAD SURGEON PERFORMED SURGERY ON THE ESOPHAGUS. IN HIS OPINION HE CONNECTED THE CAP OF THE STAPLER CORRECTLY WITH THE STAPLER ITSELF, BUT SURPRISINGLY THE STAPLE LINE WAS INCOMPLETE. JUST THE VENTRAL PART WAS CORRECTLY STAPLED. THE REST OF THE CIRCLE WAS UNTOUCHED, EVEN UNFORMED STAPLES WERE LYING IN THE SITUS. THERE WAS NO DELAY TO THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES.
04-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO HOME POSITION AND THE JAWS DID NOT OPEN WHEN THE DEVICE WAS USED ON THE LUNG. THE MANUAL OVERRIDE LEVER WAS USED TO OPEN THE JAWS. THE STAPLES WERE DEPLOYED. ANOTHER CARTRIDGE WAS LOADED INTO THE DEVICE IN QUESTION, BUT IT DID NOT FUNCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
04-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RESPIRATORY SURGERY, THE KNIFE MOVED ALL THE WAY BUT DID NOT RETURN TO HOME POSITION AND THE JAWS DID NOT OPEN WHEN THE DEVICE WAS USED ON THE LUNG. THE MANUAL OVERRIDE LEVER WAS USED TO OPEN THE JAWS. THE STAPLES WERE DEPLOYED. ANOTHER CARTRIDGE WAS LOADED INTO THE DEVICE IN QUESTION, BUT IT DID NOT FUNCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-07-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE "TRIGGER RELEASE AND EMERGENCY RELEASE DID NOT WORK." THERE WERE NO PATIENT CONSEQUENCES REPORTED.
26-07-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE "TRIGGER RELEASE AND EMERGENCY RELEASE DID NOT WORK." THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-12-2021 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT ANASTOMOSE.
16-12-2021 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT ANASTOMOSE.
16-12-2021 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT ANASTOMOSE.
16-12-2021 | Injury | us surgical puerto rico | 14-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT ANASTOMOSE.
01-01-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT ¿THE JAWS WOULD NOT OPEN WHEN PLACED IN THE PATIENT FOR USE. ANOTHER STAPLER WAS THEN USED AND WORKED FINE.¿ NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE TIP COULD NOT BE OPENED. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2021 | Injury | ethicon endo-surgery | 14-01-2022 | CONTOUR CURVED CUTTER STAPLER | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT AFTER RELOADING THE CS40G CARTRIDGE, THE SURGEON FIRED THE DEVICE, BUT THE JAW DID NOT OPEN WITH THE RELEASE BUTTON. THUS, THE SURGEON HAD TO FORCE THE JAW TO OPEN. DUE TO THE FORCE WHILE OPENING, THE RECTUM OF PATIENT WAS MINIMALLY INJURED. PROCEDURE THAT WAS A LAR OPERATION BUT CHANGED TO AN ILEOSTOMY PROCEDURE.
04-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-12-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE DEVICE COULD NOT BE FIRED ALTHOUGH THE TRIGGER WAS PULLED. THE DEVICE WAS USED BETWEEN THE COLON AND THE RECTUM. ANOTHER BATTERY WAS INSERTED INTO THE DEVICE IN QUESTION, BUT THE ISSUE DID NOT IMPROVE. THE TISSUE COMPRESSION SCALE BACKLIGHT WAS ILLUMINATED. THE ANVIL WAS REMOVED FROM THE TROCAR, A TUBE WAS ATTACHED TO THE TROCAR, AND THE DEVICE WAS REMOVED FROM THE PATIENT. THE ANVIL IN QUESTION WAS NOT REPLACED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS NOT CONVERTED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THE PATIENT'S CONDITION IS STABLE. NO FURTHER INFORMATION IS AVAILABLE.
30-11-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE NTLC75 COULDN'T BE FIRED DURING AN EOT PROCEDURE (EMERGENCY ROOM PROCEDURE). WHEN ASKED ABOUT THE INCIDENT DETAILS, THE NURSE SAID THAT WHEN LOADING THE STAPLE, THE HOME BUTTON WAS POSITIONED ARROW-ARROW, THOUGH SHE COULDN'T RECALL WHETHER SHE HAD OPENED THE LOCKING LEVER OR NOT. EVENTUALLY THEY USED A NEW NTLC75 WITH NEW SR75. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
03-11-2021 | Malfunction | covidien lp llc north haven | 14-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC BYPASS PROCEDURE, THE FIRST RELOAD WAS CONNECTED AND TESTED, BUT THE JAWS DID NOT CLOSE COMPLETELY. ANOTHER RELOAD WAS TAKEN AND USED, BUT SAME ISSUE OCCURRED. A THIRD RELOAD WAS TAKEN OUT AND WAS USED WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
01-10-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DEVICE FIRED ONE RELOAD AS PER NORMAL. SECOND RELOAD WOULD NOT FIRE. PLASTIC COMPONENTS OF RELOAD PRESENT AS BROKEN. THERE WAS NO DELAY TO THE PROCEDURE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
01-10-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | 75MM SELECTABLE NEW TLC | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DEVICE FIRED ONE RELOAD AS PER NORMAL. SECOND RELOAD WOULD NOT FIRE. PLASTIC COMPONENTS OF RELOAD PRESENT AS BROKEN. THERE WAS NO DELAY TO THE PROCEDURE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
17-12-2021 | Injury | intuitive surgical | 14-01-2022 | ENDOWRIST | Patient Device Interaction Problem | A01 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED PULMONARY LOBECTOMY, THE CURVED-TIP 30 STAPLER MADE A LOUD GRINDING NOISE AND STOPPED 1/3 OF THE WAY DURING THE STAPLING PROCESS ON THE PULMONARY ARTERY. A NEW STAPLER AND RELOAD WERE USED AND THE SURGEON PROCEEDED WITH THE CASE. THE CASE WAS COMPLETED ROBOTICALLY WITH NO PATIENT INJURY.
17-12-2021 | Injury | intuitive surgical | 14-01-2022 | ENDOWRIST | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED PULMONARY LOBECTOMY, THE CURVED-TIP 30 STAPLER MADE A LOUD GRINDING NOISE AND STOPPED 1/3 OF THE WAY DURING THE STAPLING PROCESS ON THE PULMONARY ARTERY. A NEW STAPLER AND RELOAD WERE USED AND THE SURGEON PROCEEDED WITH THE CASE. THE CASE WAS COMPLETED ROBOTICALLY WITH NO PATIENT INJURY.
15-12-2021 | Malfunction | us surgical puerto rico | 14-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY PROCEDURE, WHILE ASSEMBLING THE DEVICE AND PERFORMING A CALIBRATION TEST OUTSIDE THE PATIENT'S CAVITY, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. THE TRIGGER WAS SOFT. THE HANDLE CANNOT BE SQUEEZED AND WAS UNABLE TO FIRE. A NEW HANDLE WAS OPENED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-11-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WAS NOT FIRED. THERE WAS NO DELAY TO THE SURGERY. NO FRAGMENTS GENERATED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
29-12-2021 | Malfunction | us surgical puerto rico | 14-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESECTION OF PANCREATIC CANCER, WHILE DETACHING THE BLOOD VESSEL, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON USED A NEW HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2021 | Injury | celestica de monterrey s.a de c.v | 14-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING LUNG DURING A LAPAROSCOPIC SEGMENTAL RESECTION OF LUNG, THE DEVICE DID NOT FORM INTO A GOOD B SHAPE. SURGEON SUTURED THE CUT PART.
30-12-2021 | Malfunction | us surgical puerto rico | 14-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING BLOOD VESSEL DURING A LAPAROSCOPIC LIVER RESECTION, THE DEVICE DID NOT FIRE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2021 | Malfunction | covidien lp llc north haven | 14-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING BLOOD VESSEL DURING A LAPAROSCOPIC LIVER RESECTION, THE DEVICE DID NOT FIRE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2021 | Malfunction | covidien lp llc north haven | 14-01-2022 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING BLOOD VESSEL DURING A LAPAROSCOPIC LIVER RESECTION, THE DEVICE DID NOT FIRE. SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | FLEX 60 ARTICULATING | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED SURGERY ON THE LIVER, GRABBED A VESSEL WITH THE STAPLER AND DECIDED TO GRAB IT FROM ANOTHER ANGLE. HE WASN´T ABLE TO LOOSE THE INSTRUMENT NEITHER BY USING THE HANDLES NOR USING THE BUTTONS ON THE INSTRUMENT. THE HEAD SURGEON OPENED A NEW ECHELON WITH SAME RESULT. THE THIRD INSTRUMENT PERFORMED ACCORDINGLY. NO HARM OCCURRED TO THE PATIENT.
06-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED SURGERY ON THE LIVER, GRABBED A VESSEL WITH THE STAPLER AND DECIDED TO GRAB IT FROM ANOTHER ANGLE. HE WASN´T ABLE TO LOOSE THE INSTRUMENT NEITHER BY USING THE HANDLES NOR USING THE BUTTONS ON THE INSTRUMENT. THE HEAD SURGEON OPENED A NEW ECHELON WITH SAME RESULT. THE THIRD INSTRUMENT PERFORMED ACCORDINGLY. NO HARM OCCURRED TO THE PATIENT.
27-12-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AFTER THE RED INTERLOCK WAS UNLOCKED, EVEN THOUGH THE ORANGE INDICATOR WAS INSIDE THE GREEN SCALE AND THE DISPLAY WAS ILLUMINATED, PRESSING THE TRIGGER DID NOT ACTIVATE THE SUTURE. ALSO TRIED TO REMOVE AND INSERT THE BATTERY SEVERAL TIMES. THE INTERVENTION WAS COMPLETED USING A SECOND SUTURE LEAVING THE ANVIL ALREADY INSERTED. THERE WERE NO PATIENT CONSEQUENCES.
30-12-2021 | Malfunction | us surgical puerto rico | 14-01-2022 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, AFTER THE PRODUCT PACKAGE WAS OPENED, THE PRODUCT WAS FOUND TO BE IN A SCATTERED STATE (DISENGAGED) AND COULD NOT BE USED NORMALLY. THE DEVICE WAS REPLACED TO PROCEED WITH THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
06-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | FLEX 60 ARTICULATING | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED SURGERY ON THE LIVER, GRABBED A VESSEL WITH THE STAPLER AND DECIDED TO GRAB IT FROM ANOTHER ANGLE. HE WASN´T ABLE TO LOOSE THE INSTRUMENT NEITHER BY USING THE HANDLES NOR USING THE BUTTONS ON THE INSTRUMENT. THE HEAD SURGEON OPENED A NEW ECHELON WITH SAME RESULT. THE THIRD INSTRUMENT PERFORMED ACCORDINGLY. NO HARM OCCURRED TO THE PATIENT.
06-01-2022 | Malfunction | ethicon endo-surgery | 14-01-2022 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED SURGERY ON THE LIVER, GRABBED A VESSEL WITH THE STAPLER AND DECIDED TO GRAB IT FROM ANOTHER ANGLE. HE WASN´T ABLE TO LOOSE THE INSTRUMENT NEITHER BY USING THE HANDLES NOR USING THE BUTTONS ON THE INSTRUMENT. THE HEAD SURGEON OPENED A NEW ECHELON WITH SAME RESULT. THE THIRD INSTRUMENT PERFORMED ACCORDINGLY. NO HARM OCCURRED TO THE PATIENT.
03-12-2021 | Malfunction | us surgical puerto rico | 14-01-2022 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE ATTEMPTING THE GASTROINTESTINAL ANASTOMOSIS, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-11-2021 | Malfunction | ethicon endo-surgery | 14-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC CHOLECYSTECTOMY THE ATS STAPLER DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
29-12-2021 | Malfunction | us surgical puerto rico | 14-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE DETACHING THE RECTUM, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, THE HANDLE OF THE DEVICE CANNOT BE SQUEEZED AND WAS UNABLE TO FIRE. A NEW HANDLE WAS USED TO PROCEED WITH THE PROCEDURE. THERE WAS NO PATIENT INJURY.
30-12-2021 | Malfunction | us surgical puerto rico | 14-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHILE REMOVING THE PULMONARY LOBE THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE EVEN AFTER PRESSING THE GREEN BUTTON. THE SURGEON USED A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2021 | Malfunction | covidien lp llc north haven | 14-01-2022 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING STOMACH TRANSECTION IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLES WERE FIRED LESS THAN HALFWAY THROUGH THE LOAD WHEN IT RETRACTED. THIS RESULTS TO A DISTALLY INCOMPLETE STAPLE LINE. TO RESOLVE THE ISSUE, A NEW CARTRIDGE WAS USED. THERE WAS NO PATIENT INJURY.
06-12-2021 | Injury | intuitive surgical | 14-01-2022 | ENDOWRIST | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED SLEEVE GASTRECTOMY PROCEDURE COMPLETED ON (B)(6) 2021, THE PATIENT EXPERIENCED POST-OPERATIVE COMPLICATIONS AND RETURNED TO THE HOSPITAL EIGHT DAYS POST-OPERATIVELY FOR A STAPLE LINE LEAK. THE SURGEON BELIEVES THE SUREFORM 60 STAPLES WERE INVOLVED WITH THE STAPLE LINE LEAK. POST-OPERATIVELY, A LEAK TEST WAS PERFORMED TO CONFIRM THE LOCATION OF THE LEAK BUT THE RESULTS WERE NOT PROVIDED ARE UNKNOWN. THE PATIENT WAS READMITTED TO THE HOSPITAL AND IS ON ANTIBIOTIC THERAPY TO RESOLVE THE POST-OPERATIVE COMPLICATION. THE PATIENT¿S CURRENT STATUS IS STABLE AND IMPROVING. ON (B)(6) 2021, INTUITIVE SURGICAL, INC (ISI) OBTAINED ADDITIONAL INFORMATION ABOUT THE EVENT: THE INTUITIVE SURGICAL, INC (ISI) CLINICAL SALES REPRESENTATIVE (CSR) CONFIRMED THAT THIS EVENT IS ASSOCIATED WITH THE THIRD AND LAST OF THREE ROBOTIC PROCEDURES PERFORMED BY THE SURGEON ON (B)(6) 2021. THE CSR REPORTED THAT THE CASE WAS COMPLETED AS PLANNED WITH NO SYSTEM ERRORS, STAPLING ISSUES, OR INTRAOPERATIVE COMPLICATIONS. THE SURGEON REPORTEDLY FIRED 5 RELOADS. THE TARGET STOMACH TISSUE WAS NOT PREVIOUSLY EXPOSED TO RADIATION OR CHEMOTHERAPY. THERE WERE NO PAUSES FOR COMPRESSION WHILE STAPLING WITH THE SUREFORM STAPLER 60 INSTRUMENT. BUTTRESS MATERIAL WAS USED. HOWEVER, THE SURGEON DID NO ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL IN BETWEEN THE JAWS OF THE STAPLER INSTRUMENT. THERE WERE NO MALFORMED STAPLES OBSERVED. THERE WAS NO TISSUE TENSION. THE DA VINCI-ASSISTED PROCEDURE WAS COMPLETED ROBOTICALLY AND IT WAS REPORTED THAT ALL INSTRUMENTS WORKED AS EXPECTED AND AS INTENDED.
14-01-2022 | Malfunction | us surgical puerto rico | 14-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2021 | Malfunction | covidien lp llc north haven | 14-01-2022 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC DUODENAL SWITCH PROCEDURE, WHILE IN THE STOMACH, THE FIRST TWO RELOADS WERE STOPPING MIDFIRE AND NOT MAKING THROUGH THE STANDARD TISSUE THICKNESS, CREATING EXCESS TRS (TISSUE REINFORCEMENT SYSTEM) BUNCHING. A THIRD RELOAD WAS USED, BUT IT ALSO FIRED PARTIALLY. A NON-TRS RELOAD WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2021 | Malfunction | us surgical puerto rico | 14-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON HOWEVER, THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON REPLACED THE DEVICE WITH THE SAME MODEL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2021 | Malfunction | covidien lp llc north haven | 14-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON HOWEVER, THE DEVICE WAS NOT ABLE TO FIRE. THE SURGEON REPLACED THE DEVICE WITH THE SAME MODEL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | ethicon | 14-01-2022 | SECURESTRAP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: SECURE STRAP TACKER WAS WORKING THROUGH FIRST PART OF CASE. SURGEON REPORTED TACKER JAMMED AND WAS UNABLE TO CONTINUE USING. FDA SAFETY REPORT ID# (B)(4).
30-12-2021 | Malfunction | us surgical puerto rico | 16-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE LUNGS DURING A LAPAROSCOPIC SEGMENTECTOMY PROCEDURE, THE DEVICE DID NOT FIRE AND THE BLACK CIRCLE OF THE ARROW INDICATING THE HANDLE BROKE OFF. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2021 | Malfunction | us surgical puerto rico | 16-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE LUNGS DURING A LAPAROSCOPIC SEGMENTECTOMY PROCEDURE, THE DEVICE DID NOT FIRE AND THE BLACK CIRCLE OF THE ARROW INDICATING THE HANDLE BROKE OFF. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
31-12-2021 | Malfunction | us surgical puerto rico | 16-01-2022 | ENDO GIA ULTRA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY, TWO MANUAL HANDLES MISFIRED.
31-12-2021 | Malfunction | covidien lp llc north haven | 16-01-2022 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY, TWO MANUAL HANDLES MISFIRED.
31-12-2021 | Malfunction | us surgical puerto rico | 16-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING OPEN RADICAL TRANSVERSE COLORECTAL SURGERY, WHEN CUTTING THE TRANSVERSE COLON TISSUE, THE OPEN STAPLER DID NOT FIRE. IT WAS ALSO REPORTED THAT THE RELOAD THAT CAME PRELOADED ALSO DID NOT FIRE. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE RELOAD. THERE WAS NO PATIENT INJURY.
12-05-2021 | Malfunction | covidien lp llc north haven | 16-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN SITTING ON CHARGER, HANDLE WILL NOT CHARGE-FLASHING RED WHILE ON CHARGER. THERE WAS NO PATIENT INVOLVEMENT.
12-05-2021 | Malfunction | covidien lp llc north haven | 16-01-2022 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN SITTING ON CHARGER, HANDLE WILL NOT CHARGE-FLASHING RED WHILE ON CHARGER. THERE WAS NO PATIENT INVOLVEMENT.
10-06-2021 | Malfunction | covidien lp llc north haven | 16-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE HANDLE DID NOT TURN ON AND WHEN PUT ON THE CHARGER, THE GREEN LIGHT BLINKS FOR NINETY SECONDS AND THEN BLINKS RED. ANOTHER HANDLE WAS PLACED ON THE SAME CHARGER AND IT WORK PROPERLY. THERE WAS NO PATIENT INVOLVED.
10-06-2021 | Malfunction | covidien lp llc north haven | 16-01-2022 | SIGNIA | Electrical /Electronic Property Problem | A07 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE HANDLE DID NOT TURN ON AND WHEN PUT ON THE CHARGER, THE GREEN LIGHT BLINKS FOR NINETY SECONDS AND THEN BLINKS RED. ANOTHER HANDLE WAS PLACED ON THE SAME CHARGER AND IT WORK PROPERLY. THERE WAS NO PATIENT INVOLVED.
20-07-2021 | Malfunction | covidien lp llc north haven | 16-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE HANDLE DID NOT TURN ON. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
20-07-2021 | Malfunction | covidien lp llc north haven | 16-01-2022 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE HANDLE DID NOT TURN ON. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
07-12-2021 | Malfunction | covidien lp llc north haven | 17-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, WHEN THE RELOAD WAS ATTACHED TO THE HANDLE AND AN ATTEMPT WAS MADE TO PERFORM THE FIRST FIRING IN THE INTESTINAL COLON RESECTION. AN ERROR SOUND WAS GENERATED FROM THE HANDLE AFTER THE TISSUE WAS CLAMPED. THE MONITOR'S RELOAD FIELD WAS FOUND TO BE INDICATED ZERO WHICH SHOULD BE INDICATE AFTER FIRING. THE FIRING COULD NOT EVEN BE STARTED. IT WAS NOTED THAT THE GREEN BUTTON COULD BE PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED. TO RESOLVE THE ISSUE, A NEW RELOAD WAS TAKEN OUT AND ATTACHED TO THE SAME HANDLE, THE DEVICE COULD BE USED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
07-12-2021 | Malfunction | covidien lp llc north haven | 17-01-2022 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, WHEN THE RELOAD WAS ATTACHED TO THE HANDLE AND AN ATTEMPT WAS MADE TO PERFORM THE FIRST FIRING IN THE INTESTINAL COLON RESECTION. AN ERROR SOUND WAS GENERATED FROM THE HANDLE AFTER THE TISSUE WAS CLAMPED. THE MONITOR'S RELOAD FIELD WAS FOUND TO BE INDICATED ZERO WHICH SHOULD BE INDICATE AFTER FIRING. THE FIRING COULD NOT EVEN BE STARTED. IT WAS NOTED THAT THE GREEN BUTTON COULD BE PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED. TO RESOLVE THE ISSUE, A NEW RELOAD WAS TAKEN OUT AND ATTACHED TO THE SAME HANDLE, THE DEVICE COULD BE USED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
09-12-2021 | Malfunction | us surgical puerto rico | 17-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE LUNG RESECTION IN A LAPAROSCOPIC PNEUMONECTOMY PROCEDURE, THE HANDLE HAD AN ABNORMAL NOISE DURING FIRING. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE SHEARED TEETH DAMAGE WAS CONFIRMED IN THE 1ST AND 2ND ON THE FIRING RACK.
09-12-2021 | Malfunction | us surgical puerto rico | 17-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE LUNG RESECTION IN A LAPAROSCOPIC PNEUMONECTOMY PROCEDURE, THE HANDLE HAD AN ABNORMAL NOISE DURING FIRING. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE SHEARED TEETH DAMAGE WAS CONFIRMED IN THE 1ST AND 2ND ON THE FIRING RACK.
13-10-2021 | Injury | us surgical puerto rico | 17-01-2022 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING RECTUM SURGERY, THE CIRCULAR STAPLER WAS ABLE TO CUT THE ANASTOMOSIS, BUT DID NOT PUT STAPLES ON ONE SIDE. THE STAPLE ROW WAS OVERSEWN. A RELIEVING STOMA WAS PLACED TO PREVENT AN ANASTOMOTIC LEAK.
01-03-2021 | Malfunction | us surgical puerto rico | 17-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATIC RESECTION, ON LIGATION OF LEFT PORTAL VEIN AND LEFT HEPATOBILIARY, STAPLE FORMATION FAILED ON TWO STAPLERS. SOME STAPLE WAS FALLING OUT WITH INCOMPLETE FORMATION. BOTH STAPLERS FIRED COMPLETELY BUT THERE WAS AN INCOMPLETE STAPLE LINE ON DISTAL AREA ON THE FIRST RELOAD AND CENTRAL AREA ON THE SECOND RELOAD. STAPLE CARTRIDGES GOT BROKEN WHILE FIRING. THE VESSEL WAS TEMPORARILY BURST BECAUSE STAPLE DID NOT HAVE SUCCESSFUL FORMATION. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. AFTER FIRING, THE RELOAD FAILED TO BE REMOVED FROM THE HANDLE. TO COMPLETE THE CASE ANOTHER STAPLER FROM THE COMPETITOR WAS USED.
01-03-2021 | Malfunction | us surgical puerto rico | 17-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATIC RESECTION, ON LIGATION OF LEFT PORTAL VEIN AND LEFT HEPATOBILIARY, STAPLE FORMATION FAILED ON TWO STAPLERS. SOME STAPLE WAS FALLING OUT WITH INCOMPLETE FORMATION. BOTH STAPLERS FIRED COMPLETELY BUT THERE WAS AN INCOMPLETE STAPLE LINE ON DISTAL AREA ON THE FIRST RELOAD AND CENTRAL AREA ON THE SECOND RELOAD. STAPLE CARTRIDGES GOT BROKEN WHILE FIRING. THE VESSEL WAS TEMPORARILY BURST BECAUSE STAPLE DID NOT HAVE SUCCESSFUL FORMATION. SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. AFTER FIRING, THE RELOAD FAILED TO BE REMOVED FROM THE HANDLE. TO COMPLETE THE CASE ANOTHER STAPLER FROM THE COMPETITOR WAS USED.
22-12-2021 | Malfunction | ethicon | 17-01-2022 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC INGUINAL HERNIA REPAIR ON (B)(6) 2021 AND THE ABSORBABLE STRAP WAS USED. IT WAS REPORTED THAT IT BOUNCED RIGHT OFF THE TISSUE AND DID NOT PIERCE THE TISSUE, ONLY THE MESH. IT WAS REPORTED THAT THE SURGEON HAD APPROPRIATE PRESSURE AND APPROXIMATION. IT WAS REPORTED THAT THE SURGERY WAS DELAYED FOUR MINUTES. IT WAS REPORTED THAT FRAGMENTS WERE GENERATED AND REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION. IT WAS REPORTED THAT THE REPLACEMENT HAD A DIFFERENT LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED.
10-12-2021 | Malfunction | ethicon | 17-01-2022 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2021 AND THE ABSORBABLE STRAP WAS USED. IT WAS REPORTED THAT THE STRAP COULD NOT BE DEPLOYED WELL. IT WAS REPORTED THAT WHEN THE TRIGGER WAS GRASPED, THE STRAP FELL OFF THE JAWS. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED.
22-09-2021 | Malfunction | us surgical puerto rico | 17-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, WHILE BEING APPLIED IN THE PANCREAS HEAD, BY FIRING THE INSTRUMENT, THERE WAS A CRACK SOUND AND LARGE BANG WITH PROPER STAPLE LINE, BUT THE HANDLE COULD NOT BE USED ANYMORE AND WAS BROKEN. THE SURGEON TOOK ANOTHER HANDLE AND NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-09-2021 | Malfunction | us surgical puerto rico | 17-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, WHILE BEING APPLIED IN THE PANCREAS HEAD, BY FIRING THE INSTRUMENT, THERE WAS A CRACK SOUND AND LARGE BANG WITH PROPER STAPLE LINE, BUT THE HANDLE COULD NOT BE USED ANYMORE AND WAS BROKEN. THE SURGEON TOOK ANOTHER HANDLE AND NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2021 | Malfunction | covidien lp llc north haven | 17-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE CUTTING THE LUNGS DURING A LAPAROSCOPIC SEGMENTECTOMY PROCEDURE, THE DEVICE DID NOT FIRE AND THE BLACK CIRCLE OF THE ARROW INDICATING THE HANDLE BROKE OFF. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-09-2021 | Malfunction | us surgical puerto rico | 18-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, WHILE SUTURING THE HILAR OF THE LUNGS, THE STAPLE BURST REPEATEDLY FROM THE FIVE RELOADS. SOME STAPLES COULD NOT FORM INTO "B" SHAPE. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. TO COMPLETE THE CASE, THE SURGEON REPLACED THE RELOAD. THERE WAS NO PATIENT INJURY.
27-09-2021 | Malfunction | us surgical puerto rico | 18-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, WHILE SUTURING THE HILAR OF THE LUNGS, THE STAPLE BURST REPEATEDLY FROM THE FIVE RELOADS. SOME STAPLES COULD NOT FORM INTO "B" SHAPE. THE STAPLE LINE WAS INCOMPLETE CENTRALLY. TO COMPLETE THE CASE, THE SURGEON REPLACED THE RELOAD. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | covidien lp llc north haven | 18-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, WHILE TRANSECTING THE LIVER PARENCHYMA, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE EVEN AFTER PRESSING THE GREEN BUTTON. THE SURGEON USED A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 18-01-2022 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC PROCEDURE ON (B)(6) 2021, THE BARD/DAVOL OPTIFIX STRAIGHT DEVICE WAS USED TO FIXATE AN UNKNOWN MESH. IT WAS REPORTED THAT BROKEN FASTENERS WERE DEPLOYED FROM THE DEVICE. THE SURGEON IS AN EXPERIENCED USER OF THE DEVICE. THERE WAS NO REPORTED PATIENT INJURY.
31-12-2021 | Malfunction | us surgical puerto rico | 18-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING OPEN RADICAL TRANSVERSE COLORECTAL SURGERY, WHEN CUTTING THE TRANSVERSE COLON TISSUE, THE OPEN STAPLER DID NOT FIRE. IT WAS ALSO REPORTED THAT THE RELOAD THAT CAME PRELOADED ALSO DID NOT FIRE. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE RELOAD. THERE WAS NO PATIENT INJURY.
30-12-2021 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF LOW RECTAL CANCER, WHILE DETACHING THE RECTUM, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2021 | Injury | celestica de monterrey s.a de c.v | 18-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, AFTER ANASTOMOSIS AND DETACHING THE DISTAL COLON, THE STAPLES WERE POORLY FORMED AND THE TISSUE WAS OOZING BLOOD. THE ANASTOMOSIS SITE WAS SUTURED TO STOP THE BLEEDING. THERE WAS NO PATIENT INJURY.
30-12-2021 | Injury | celestica de monterrey s.a de c.v | 18-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLECTOMY, AFTER ANASTOMOSIS AND DETACHING THE DISTAL COLON, THE STAPLES WERE POORLY FORMED AND THE TISSUE WAS OOZING BLOOD. THE ANASTOMOSIS SITE WAS SUTURED TO STOP THE BLEEDING. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | covidien lp llc north haven | 18-01-2022 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN HEMORRHOID SURGERY, THE DEVICE DID NOT FIRE. THE STAPLER DID NOT WORK AT ALL. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-12-2021 | Malfunction | us surgical puerto rico | 18-01-2022 | EEA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ANTERIOR RESECTION OF THE RECTUM, ANASTOMOSIS WAS CARRIED OUT BUT THE ANVIL BECAME STUCK IN THE COLON AFTER ACTIVATION OF THE CLAMP. THE SURGEON HAD TO FORCE THE SYSTEM TO EXTRACT IT IN ORDER TO COMPLETE HIS PROCEDURE. ONCE THE FORCEPS WERE OUT, THE ANVIL WAS NO LONGER VERTICAL AND A SMALL PIECE OF METAL WAS FOUND WITH THE ANVIL/DONUTS OF THE ANASTOMOSIS. THERE WAS NO PATIENT INJURY.
28-12-2021 | Malfunction | us surgical puerto rico | 18-01-2022 | EEA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ANTERIOR RESECTION OF THE RECTUM, ANASTOMOSIS WAS CARRIED OUT BUT THE ANVIL BECAME STUCK IN THE COLON AFTER ACTIVATION OF THE CLAMP. THE SURGEON HAD TO FORCE THE SYSTEM TO EXTRACT IT IN ORDER TO COMPLETE HIS PROCEDURE. ONCE THE FORCEPS WERE OUT, THE ANVIL WAS NO LONGER VERTICAL AND A SMALL PIECE OF METAL WAS FOUND WITH THE ANVIL/DONUTS OF THE ANASTOMOSIS. THERE WAS NO PATIENT INJURY.
23-12-2021 | Injury | us surgical puerto rico | 18-01-2022 | EEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A SIGMOIDECTOMY, AFTER TERMINAL COLORECTAL ANASTOMOSIS, THERE WAS DIFFICULTY IN REMOVING THE STAPLER FROM THE CAVITY. WHEN REMOVING THE CLAMP, THE ANVIL TRIPPED FROM THE CLAMP. THERE WAS A SERIOUS TEAR ON THE ANTERIOR SURFACE OF THE ANASTOMOSIS. THE ANVIL WAS TRIED TO BE INSERTED TO RESOLVE THE ISSUE AND FINISH THE ANASTOMOSIS.
23-12-2021 | Injury | us surgical puerto rico | 18-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A SIGMOIDECTOMY, AFTER TERMINAL COLORECTAL ANASTOMOSIS, THERE WAS DIFFICULTY IN REMOVING THE STAPLER FROM THE CAVITY. WHEN REMOVING THE CLAMP, THE ANVIL TRIPPED FROM THE CLAMP. THERE WAS A SERIOUS TEAR ON THE ANTERIOR SURFACE OF THE ANASTOMOSIS. THE ANVIL WAS TRIED TO BE INSERTED TO RESOLVE THE ISSUE AND FINISH THE ANASTOMOSIS.
21-12-2021 | Malfunction | us surgical puerto rico | 18-01-2022 | GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SUBTOTAL COLECTOMY, WHILE FIRING THE OPEN STAPLER, STAPLES SLACKED, THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE. THE SURGEON USED ANOTHER DEVICE TO RESOLVE THE ISSUE.
03-01-2022 | Malfunction | us surgical puerto rico | 18-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BULLECTOMY, WHILE RESECTING THE BULLA, THE SECOND RELOAD MISFIRED AND COULD NOT BE ABLE TO UNLOAD IT AND THE FOURTH RELOAD GOT STUCK WHILE FIRING. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON USED ANOTHER INSTRUMENT TO OPEN THE JAW FROM THE TISSUE AND USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | us surgical puerto rico | 18-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BULLECTOMY, WHILE RESECTING THE BULLA, THE SECOND RELOAD MISFIRED AND COULD NOT BE ABLE TO UNLOAD IT AND THE FOURTH RELOAD GOT STUCK WHILE FIRING. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON USED ANOTHER INSTRUMENT TO OPEN THE JAW FROM THE TISSUE AND USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | us surgical puerto rico | 18-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BULLECTOMY, WHILE RESECTING THE BULLA, THE SECOND RELOAD MISFIRED AND COULD NOT BE ABLE TO UNLOAD IT AND THE FOURTH RELOAD GOT STUCK WHILE FIRING. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON USED ANOTHER INSTRUMENT TO OPEN THE JAW FROM THE TISSUE AND USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | us surgical puerto rico | 18-01-2022 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BULLECTOMY, WHILE RESECTING THE BULLA, THE SECOND RELOAD MISFIRED AND COULD NOT BE ABLE TO UNLOAD IT AND THE FOURTH RELOAD GOT STUCK WHILE FIRING. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON USED ANOTHER INSTRUMENT TO OPEN THE JAW FROM THE TISSUE AND USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | us surgical puerto rico | 18-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BULLECTOMY, WHILE RESECTING THE BULLA, THE SECOND RELOAD MISFIRED AND COULD NOT BE ABLE TO UNLOAD IT AND THE FOURTH RELOAD GOT STUCK WHILE FIRING. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON USED ANOTHER INSTRUMENT TO OPEN THE JAW FROM THE TISSUE AND USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | us surgical puerto rico | 18-01-2022 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BULLECTOMY, WHILE RESECTING THE BULLA, THE SECOND RELOAD MISFIRED AND COULD NOT BE ABLE TO UNLOAD IT AND THE FOURTH RELOAD GOT STUCK WHILE FIRING. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON USED ANOTHER INSTRUMENT TO OPEN THE JAW FROM THE TISSUE AND USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2021 | Malfunction | covidien lp llc north haven | 18-01-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, WHILE IN THE MIDDLE OF THE DISSECTION DURING A BARIATRIC SURGERY, THE MANUAL STAPLER KNIFE STOPPED AND THE DISTAL END HAS INCOMPLETE STAPLE LINE. THE SURGEON USED ANOTHER STAPLER TO RESOLVE THE ISSUE.
01-01-2021 | Malfunction | ethicon endo-surgery | 18-01-2022 | POWERED PLUS STANDARD 45 | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS DAMAGED WITHIN BOX, SO NOT STERILE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
04-01-2022 | Malfunction | us surgical puerto rico | 18-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY ON CARDIAN CARCINOMA, WHEN USING THE DEVICE ON THE LESSER CURVATURE OF THE STOMACH, IT CANNOT BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | covidien lp llc north haven | 18-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY ON CARDIAN CARCINOMA, WHEN USING THE DEVICE ON THE LESSER CURVATURE OF THE STOMACH, IT CANNOT BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 18-01-2022 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNSPECIFIED PROCEDURE ON (B)(6) 2021, A BARD/DAVOL OPTIFIX STRAIGHT FIXATION DEVICE WAS USED TO FIXATE AN UNSPECIFIED MESH. IT WAS REPORTED THAT THE FIRST TWO FASTENERS DEPLOYED WERE BROKEN AND THOSE WERE REMOVED FROM THE PATIENT¿S BODY AND THE DEVICE WAS CYCLED IN AIR. AS REPORTED, THERE WAS NO FURTHER PROBLEM AND THE SAME DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY.
21-12-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, WHILE IN THE MIDDLE OF THE DISSECTION DURING A BARIATRIC SURGERY, MANUAL STAPLER KNIFE STOPPED AND THE DISTAL END HAS INCOMPLETE STAPLE LINE. SURGEON USED ANOTHER STAPLER TO RESOLVE THE ISSUE.
03-01-2022 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER. DURING A LAPAROSCOPIC DISTAL GASTRECTOMY, WHILE DETACHING A GASTRIC BODY, THERE WAS POOR STAPLE FORMATION. THE STAPLES WERE NOT WELL AND NOT FIXED. TO RESOLVE THE ISSUE, THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
04-10-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE ALTHOUGH THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. AT THE TIME OF ACTIVATING THE HANDLE, IT SPLITS AND CAUSED THE RELOAD TO BE DAMAGED AS WELL (IT DID NOT LET IT ACTIVATE).
04-10-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE ALTHOUGH THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. AT THE TIME OF ACTIVATING THE HANDLE, IT SPLITS AND CAUSED THE RELOAD TO BE DAMAGED AS WELL (IT DID NOT LET IT ACTIVATE).
04-10-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE ALTHOUGH THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. AT THE TIME OF ACTIVATING THE HANDLE, IT SPLITS AND CAUSED THE RELOAD TO BE DAMAGED AS WELL (IT DID NOT LET IT ACTIVATE).
04-10-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE ALTHOUGH THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. AT THE TIME OF ACTIVATING THE HANDLE, IT SPLITS AND CAUSED THE RELOAD TO BE DAMAGED AS WELL (IT DID NOT LET IT ACTIVATE).
11-01-2022 | Malfunction | ethicon endo-surgery | 19-01-2022 | ECHELON FLEX ENDOPATH | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLER WOULD JAM WHEN TRYING TO FIRE -- TRIED 2 DIFFERENT BLACK LOADS WITH PERISTRIPS AND STILL WOULD NOT FIRE. STAPLER REPLACED WITH NEW ONE AND IT WORKED.
04-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE ALTHOUGH THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. AT THE TIME OF ACTIVATING THE HANDLE, IT SPLITS AND CAUSED THE RELOAD TO BE DAMAGED AS WELL (IT DID NOT LET IT ACTIVATE).
04-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE ALTHOUGH THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. AT THE TIME OF ACTIVATING THE HANDLE, IT SPLITS AND CAUSED THE RELOAD TO BE DAMAGED AS WELL (IT DID NOT LET IT ACTIVATE).
04-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE ALTHOUGH THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. AT THE TIME OF ACTIVATING THE HANDLE, IT SPLITS AND CAUSED THE RELOAD TO BE DAMAGED AS WELL (IT DID NOT LET IT ACTIVATE).
04-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE ALTHOUGH THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZED THE HANDLE. AT THE TIME OF ACTIVATING THE HANDLE, IT SPLITS AND CAUSED THE RELOAD TO BE DAMAGED AS WELL (IT DID NOT LET IT ACTIVATE).
22-12-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOID, THE STAPLES WERE LOCKED AND THE STAPLER DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Injury | covidien lp llc north haven | 19-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, DURING THE TRANSECTION OF THE LUNG, THE KNIFE STOPPED AND THE STAPLE FORMATION WAS POOR. THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. THE STAPLE LINE WAS FIXED BY FIRING A NEW ONE BESIDE THE STAPLE LINE. THE PUSHER ALSO BROKE OFF.
03-01-2022 | Injury | covidien lp llc north haven | 19-01-2022 | ENDO GIA | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, DURING THE TRANSECTION OF THE LUNG, THE KNIFE STOPPED AND THE STAPLE FORMATION WAS POOR. THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. THE STAPLE LINE WAS FIXED BY FIRING A NEW ONE BESIDE THE STAPLE LINE. THE PUSHER ALSO BROKE OFF.
30-12-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | TA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A BOWEL RESECTION SURGERY, THE DEVICE WAS DIFFICULT TO REOPEN FOR ADJUSTMENT. THE SURGEON WAS ABLE TO WORK THE DEVICE TO BE REMOVED ON THE TISSUE. IT WAS OPENED BY PRESSING THE BLACK KNOB WITH DIFFICULTY. SURGEON USED A NEW DEVICE TO CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Injury | us surgical puerto rico | 19-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, DURING THE TRANSECTION OF THE LUNG, THE KNIFE STOPPED AND THE STAPLE FORMATION WAS POOR. THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. THE STAPLE LINE WAS FIXED BY FIRING A NEW ONE BESIDE THE STAPLE LINE. THE PUSHER ALSO BROKE OFF.
03-01-2022 | Injury | us surgical puerto rico | 19-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOBECTOMY, DURING THE TRANSECTION OF THE LUNG, THE KNIFE STOPPED AND THE STAPLE FORMATION WAS POOR. THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. THE STAPLE LINE WAS FIXED BY FIRING A NEW ONE BESIDE THE STAPLE LINE. THE PUSHER ALSO BROKE OFF.
29-12-2021 | Malfunction | celestica de monterrey s.a de c.v | 19-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LOBECTOMY, WHILE DIVIDING THE LOBE AND THE BRONCHUS, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN WHEN THE GREEN BUTTON WAS PRESSED. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-12-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LOBECTOMY, WHILE DIVIDING THE LOBE AND THE BRONCHUS, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN WHEN THE GREEN BUTTON WAS PRESSED. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE INSPECTION, THE USER STATED THAT AFTER SOFTWARE THE HANDLE WOULD NOT TURN ON. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED IN THE EVENT.
07-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE INSPECTION, THE USER STATED THAT AFTER SOFTWARE THE HANDLE WOULD NOT TURN ON. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED IN THE EVENT.
04-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PERIODICAL CHECK UP, BUT THE HANDLE DID NOT START UP. THE CHARGER LAMP WAS ILLUMINATING GREEN. THERE WAS NO PATIENT INVOLVEMENT.
04-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PERIODICAL CHECK UP, BUT THE HANDLE DID NOT START UP. THE CHARGER LAMP WAS ILLUMINATING GREEN. THERE WAS NO PATIENT INVOLVEMENT.
27-09-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE HANDLE WAS UNABLE TO BE CHARGED. IT WAS SET TO THE CHARGER, AND IT FLASHED GREEN, BUT IT DID NOT ILLUMINATE GREEN AFTER SEVERAL HOURS, AND THE HANDLE DID NOT START UP. THE DEVICE COULD NOT BE CHARGED EVEN AFTER IT WAS RESTARTED. NO PATIENT INVOLVEMENT.
27-09-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE HANDLE WAS UNABLE TO BE CHARGED. IT WAS SET TO THE CHARGER, AND IT FLASHED GREEN, BUT IT DID NOT ILLUMINATE GREEN AFTER SEVERAL HOURS, AND THE HANDLE DID NOT START UP. THE DEVICE COULD NOT BE CHARGED EVEN AFTER IT WAS RESTARTED. NO PATIENT INVOLVEMENT.
04-01-2022 | Injury | covidien lp llc north haven | 19-01-2022 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPY SURGERY LOBECTOMY, WHILE STAPLING THE LUNG, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE PATIENT EXPERIENCED UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. THE SURGEON NEEDED TO CUT THE TISSUE FROM THE JAWS TO REMOVE THE DEVICE.
04-01-2022 | Injury | covidien lp llc north haven | 19-01-2022 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPY SURGERY LOBECTOMY, WHILE STAPLING THE LUNG, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE PATIENT EXPERIENCED UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. THE SURGEON NEEDED TO CUT THE TISSUE FROM THE JAWS TO REMOVE THE DEVICE.
07-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO SURGERY, THE SCREEN OF THE POWERED HANDLE WAS BLACKOUT AND DID NOT START UP DESPITE BEING CHARGED. THE ISSUE OCCURRED AFTER VERSION UPGRADE. THE CHARGER WORKED FINE WITH A DIFFERENT HANDLE. THERE WAS NO PATIENT INVOLVED.
07-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO SURGERY, THE SCREEN OF THE POWERED HANDLE WAS BLACKOUT AND DID NOT START UP DESPITE BEING CHARGED. THE ISSUE OCCURRED AFTER VERSION UPGRADE. THE CHARGER WORKED FINE WITH A DIFFERENT HANDLE. THERE WAS NO PATIENT INVOLVED.
22-12-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING REMOVAL OF THE APPENDIX IN A LAPAROSCOPIC APPENDECTOMY, THE DEVICE HAD POOR STAPLE FORMATION. STAPLES THAT WERE FORMED BEYOND THE SPECIMEN FELL INTO THE ABDOMINAL CAVITY. USED SUCTION/IRRIGATOR OR LAP INSTRUMENTS TO COLLECT THE LOOSE STAPLES AND REMOVED THEM FROM THE PATIENT. NO PATIENT INJURY.
22-12-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING REMOVAL OF THE APPENDIX IN A LAPAROSCOPIC APPENDECTOMY, THE DEVICE HAD POOR STAPLE FORMATION. STAPLES THAT WERE FORMED BEYOND THE SPECIMEN FELL INTO THE ABDOMINAL CAVITY. USED SUCTION/IRRIGATOR OR LAP INSTRUMENTS TO COLLECT THE LOOSE STAPLES AND REMOVED THEM FROM THE PATIENT. NO PATIENT INJURY.
01-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CHARGER ILLUMINATED RED. THE HANDLE ALSO DID NOT ACTIVATE. THEY BROUGHT A CHARGER FOR THE DEMONSTRATION AND INSERTED THE REPORTED HANDLE, BUT THE HANDLE STILL DID NOT ACTIVATE. THERE WAS A LIQUID-LIKE SUBSTANCE ATTACHED TO BOTH THE CHARGER AND THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
01-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CHARGER ILLUMINATED RED. THE HANDLE ALSO DID NOT ACTIVATE. THEY BROUGHT A CHARGER FOR THE DEMONSTRATION AND INSERTED THE REPORTED HANDLE, BUT THE HANDLE STILL DID NOT ACTIVATE. THERE WAS A LIQUID-LIKE SUBSTANCE ATTACHED TO BOTH THE CHARGER AND THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
01-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CHARGER ILLUMINATED RED. THE HANDLE ALSO DID NOT ACTIVATE. THEY BROUGHT A CHARGER FOR THE DEMONSTRATION AND INSERTED THE REPORTED HANDLE, BUT THE HANDLE STILL DID NOT ACTIVATE. THERE WAS A LIQUID-LIKE SUBSTANCE ATTACHED TO BOTH THE CHARGER AND THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
06-12-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHILE STAPLING THE LUNG PARENCHYMA, THE TWO STAPLERS WERE UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN BOWEL RESECTION PROCEDURE, WHILE THE SURGEON WAS ATTEMPTING TO FIRE THE STAPLER ACROSS THE BOWEL, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ANVIL DISCONNECTED FROM THE SPIKE WHILE CLOSING THE STAPLER. NO PATIENT HARM. THERE WERE NO PATIENT CONSEQUENCES.
06-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE FAILED TO CHARGE. A MANUAL STAPLER WAS USED INSTEAD. THE CHARGER WORKED WITH OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT.
06-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE FAILED TO CHARGE. A MANUAL STAPLER WAS USED INSTEAD. THE CHARGER WORKED WITH OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT.
11-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, A RED LIGHT WAS NOTED DURING CHARGING. BATTERY LEAK WAS NOTED ON THE HANDLE. A REPLACEMENT WAS PROVIDED. THERE WAS NO PATIENT INVOLVEMENT.
11-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, A RED LIGHT WAS NOTED DURING CHARGING. BATTERY LEAK WAS NOTED ON THE HANDLE. A REPLACEMENT WAS PROVIDED. THERE WAS NO PATIENT INVOLVEMENT.
11-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, A RED LIGHT WAS NOTED DURING CHARGING. BATTERY LEAK WAS NOTED ON THE HANDLE. A REPLACEMENT WAS PROVIDED. THERE WAS NO PATIENT INVOLVEMENT.
30-12-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY PROCEDURE, WHEN DIVIDING THE COLON ASCENDENS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | us surgical puerto rico | 19-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE USING THE DEVICE ON GASTRIC PARENCHYMA DURING LAPAROTOMY GASTRIC DISSECTION, THE SURGEON COULD NOT PUSH THE FIRING KNOB WHEN THEY TRIED TO FIRE THE DEVICE. IT WAS SAID THAT IT WAS ABLE TO FIRE SUCCESSFULLY WHEN USING A NEW PRODUCT. NO PATIENT INJURY.
07-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICE MAINTENANCE, THE POWERED STAPLER FAILED TO CHARGE. THERE WAS A BATTERY FLUID LEAKAGE FOUND ON THE DEVICE.
07-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICE MAINTENANCE, THE POWERED STAPLER FAILED TO CHARGE. THERE WAS A BATTERY FLUID LEAKAGE FOUND ON THE DEVICE.
07-10-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICE MAINTENANCE, THE POWERED STAPLER FAILED TO CHARGE. THERE WAS A BATTERY FLUID LEAKAGE FOUND ON THE DEVICE.
30-12-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY PROCEDURE, WHEN DIVIDING THE COLON ASCENDENS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY PROCEDURE, WHEN DIVIDING THE COLON ASCENDENS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | PROXIMATE*PPH PROCEDURE SET | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON WAS USING OUR PPH03 STAPLER FOR THE TREATMENT OF HEMORRHOID. STAPLER GOT MISFIRED, IT CUT WITHOUT STAPLING.
28-12-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | PROXIMATE*PPH PROCEDURE SET | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON WAS USING OUR PPH03 STAPLER FOR THE TREATMENT OF HEMORRHOID. STAPLER GOT MISFIRED, IT CUT WITHOUT STAPLING.
04-01-2022 | Malfunction | us surgical puerto rico | 19-01-2022 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SURGERY FOR RECTAL CANCER, WHILE DETACHING THE INTESTINES, THE RELOAD WAS PARTIALLY DISENGAGED FROM THE DEVICE. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN WHEN THE GREEN BUTTON WAS PRESSED. THE SURGEON USED A NEW HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Injury | covidien lp llc north haven | 19-01-2022 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPY SURGERY LOBECTOMY, WHILE STAPLING THE LUNG, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE PATIENT EXPERIENCED UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. THE SURGEON NEEDED TO CUT THE TISSUE FROM THE JAWS TO REMOVE THE DEVICE.
17-11-2021 | Injury | us surgical puerto rico | 19-01-2022 | GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A DIAGNOSTIC LAPAROTOMY, IT WAS CONVERTED TO AN EXPLORATORY LAPAROTOMY AND WHILE ON BOWEL RESECTION THE STAPLER GOT STUCK. THE DEVICE WAS CLAMPED DOWN AND WOULD NOT OPEN DESPITE MULTIPLE ATTEMPTS TO UNLOCK THE INSTRUMENT. THIS CAUSED THE BOWEL THAT WAS TRAPPED IN THE STAPLER TO BE SURGICALLY CUT TO REMOVE THE PRODUCT FROM THE PATIENT.
17-11-2021 | Injury | us surgical puerto rico | 19-01-2022 | GIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A DIAGNOSTIC LAPAROTOMY, IT WAS CONVERTED TO AN EXPLORATORY LAPAROTOMY AND WHILE ON BOWEL RESECTION THE STAPLER GOT STUCK. THE DEVICE WAS CLAMPED DOWN AND WOULD NOT OPEN DESPITE MULTIPLE ATTEMPTS TO UNLOCK THE INSTRUMENT. THIS CAUSED THE BOWEL THAT WAS TRAPPED IN THE STAPLER TO BE SURGICALLY CUT TO REMOVE THE PRODUCT FROM THE PATIENT.
17-11-2021 | Injury | us surgical puerto rico | 19-01-2022 | GIA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A DIAGNOSTIC LAPAROTOMY, IT WAS CONVERTED TO AN EXPLORATORY LAPAROTOMY AND WHILE ON BOWEL RESECTION THE STAPLER GOT STUCK. THE DEVICE WAS CLAMPED DOWN AND WOULD NOT OPEN DESPITE MULTIPLE ATTEMPTS TO UNLOCK THE INSTRUMENT. THIS CAUSED THE BOWEL THAT WAS TRAPPED IN THE STAPLER TO BE SURGICALLY CUT TO REMOVE THE PRODUCT FROM THE PATIENT.
17-11-2021 | Injury | us surgical puerto rico | 19-01-2022 | GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A DIAGNOSTIC LAPAROTOMY, IT WAS CONVERTED TO AN EXPLORATORY LAPAROTOMY AND WHILE ON BOWEL RESECTION THE STAPLER GOT STUCK. THE DEVICE WAS CLAMPED DOWN AND WOULD NOT OPEN DESPITE MULTIPLE ATTEMPTS TO UNLOCK THE INSTRUMENT. THIS CAUSED THE BOWEL THAT WAS TRAPPED IN THE STAPLER TO BE SURGICALLY CUT TO REMOVE THE PRODUCT FROM THE PATIENT.
04-01-2022 | Injury | covidien lp llc north haven | 19-01-2022 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPY SURGERY LOBECTOMY, WHILE STAPLING THE LUNG, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE PATIENT EXPERIENCED UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. THE SURGEON NEEDED TO CUT THE TISSUE FROM THE JAWS TO REMOVE THE DEVICE.
04-01-2022 | Injury | covidien lp llc north haven | 19-01-2022 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPY SURGERY LOBECTOMY, WHILE STAPLING THE LUNG, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE PATIENT EXPERIENCED UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE. THE SURGEON NEEDED TO CUT THE TISSUE FROM THE JAWS TO REMOVE THE DEVICE.
31-12-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | GST60W SINGLE PRODUCT INSTRUMENT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
31-12-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | GST60W SINGLE PRODUCT INSTRUMENT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
21-12-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING UNKNOWN PROCEDURE, THE ANVIL DISCONNECTS FROM SPIKE WHILE CLOSING THE STAPLER. WE EXPERIENCED BOTH THE AUDIBLE CLICK AND THE VISUAL CUE OF THE ANVIL OF THE ANVIL BEING FULLY MATED WITH THE SPIKE. AND THEN NOTICED THEM DISCONNECT UPON CLOSING. THEY NOTICED IT EARLY ON, SO THEY WERE ABLE TO REOPEN AND REDEPLOY TO THE SPIKE TO RECONNECT AND FIRE SUCCESSFULLY. NO PATIENT HARM. THERE WERE NO PATIENT CONSEQUENCES.
12-08-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE DEVICE COULD NOT BE FIRED AFTER CLOSING. NO PATIENT INJURY.
31-12-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | GST60W SINGLE PRODUCT INSTRUMENT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
31-12-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | GST60W SINGLE PRODUCT INSTRUMENT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-08-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE THE DEVICE COULD NOT BE FIRED AFTER CLOSING. NO PATIENT INJURY.
12-08-2021 | Malfunction | covidien lp llc north haven | 19-01-2022 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE THE DEVICE COULD NOT BE FIRED AFTER CLOSING. NO PATIENT INJURY.
11-10-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCTRECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-11-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE MANUAL LINER CUTTING STAPLER BROKE UPON INSERTION OF THE RELOAD.
18-11-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE MANUAL LINER CUTTING STAPLER BROKE UPON INSERTION OF THE RELOAD.
18-11-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE MANUAL LINER CUTTING STAPLER BROKE UPON INSERTION OF THE RELOAD.
18-11-2021 | Malfunction | us surgical puerto rico | 19-01-2022 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE MANUAL LINER CUTTING STAPLER BROKE UPON INSERTION OF THE RELOAD.
07-10-2021 | Malfunction | ethicon endo-surgery | 19-01-2022 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCTRECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-11-2021 | Malfunction | davis & geck caribe | 19-01-2022 | UNKNOWN ENDO STITCH INSTRUMENT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HYSTERECTOMY, THE NEEDLE OF THE DEVICE BROKE AND FELL OFF THE DEVICE WHILE SUTURING THE VAGINAL CUFF.
04-01-2022 | Injury | us surgical puerto rico | 19-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE 2 HANDLES AND 2 RELOADS WERE ABLE TO FIRE, THE STAPLE LINE WAS INCOMPLETE DISTALLY, THERE WAS POOR STAPLE FORMATION. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
04-01-2022 | Injury | celestica de monterrey s.a de c.v | 20-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE 2 HANDLES AND 2 RELOADS WERE ABLE TO FIRE, THE STAPLE LINE WAS INCOMPLETE DISTALLY, THERE WAS POOR STAPLE FORMATION. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
04-01-2022 | Injury | celestica de monterrey s.a de c.v | 20-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE 2 HANDLES AND 2 RELOADS WERE ABLE TO FIRE, THE STAPLE LINE WAS INCOMPLETE DISTALLY, THERE WAS POOR STAPLE FORMATION. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
04-01-2022 | Injury | celestica de monterrey s.a de c.v | 20-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE 2 HANDLES AND 2 RELOADS WERE ABLE TO FIRE, THE STAPLE LINE WAS INCOMPLETE DISTALLY, THERE WAS POOR STAPLE FORMATION. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
04-01-2022 | Injury | celestica de monterrey s.a de c.v | 20-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE 2 HANDLES AND 2 RELOADS WERE ABLE TO FIRE, THE STAPLE LINE WAS INCOMPLETE DISTALLY, THERE WAS POOR STAPLE FORMATION. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
04-01-2022 | Injury | covidien lp llc north haven | 20-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG SEGMENTECTOMY PROCEDURE, THE STAPLE FORMATION WAS POOR, AND THE STAPLE LINE WAS INCOMPLETE IN THE CENTER. THE STAPLE LINE WAS FIXED WITH SUTURE. THE STAPLER WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Injury | covidien lp llc north haven | 20-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG SEGMENTECTOMY PROCEDURE, THE STAPLE FORMATION WAS POOR, AND THE STAPLE LINE WAS INCOMPLETE IN THE CENTER. THE STAPLE LINE WAS FIXED WITH SUTURE. THE STAPLER WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT HEMICOLECTOMY, A PART OF THE TWO DEVICES DISENGAGED WHILE RELOADING. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING RIGHT HEMICOLECTOMY, A PART OF THE TWO DEVICES DISENGAGED WHILE RELOADING. A NEW STAPLER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Injury | us surgical puerto rico | 20-01-2022 | EEA | Product Quality Problem | A0201 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLON RESECTION, TOWARDS THE END OF PROCEDURE, THE PROXIMAL RING WAS THIN. THE SURGEON HAD TO OVERSEW THE THIN PROXIMAL RING TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
04-01-2022 | Injury | us surgical puerto rico | 20-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLON RESECTION, TOWARDS THE END OF PROCEDURE, THE PROXIMAL RING WAS THIN. THE SURGEON HAD TO OVERSEW THE THIN PROXIMAL RING TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
20-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, A COMPONENT OF THE DEVICE DISENGAGED. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE IN ORDER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
14-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | EEA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LOWER ANTERIOR RESECTION, DURING THE CREATION OF THE ANASTOMOSIS OF THE COLON TISSUE, THE DEVICE DID NOT FIRE. THE ANASTOMOSIS WAS HAND-SEWN TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE ISSUE. THE SURGEON DIVERTED THE COLON TO LET THE HAND SEWED ANASTOMOSIS HEAL BEFORE ALLOWING HE SUBSTANCES TO FLOW THROUGH IT.
04-01-2022 | Injury | us surgical puerto rico | 20-01-2022 | EEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ASSISTED PROCTECTOMY WITH ILEAL ANAL ANASTOMOSIS, WHILE ANASTOMOSING THE COLON, THE STAPLE DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE. THE DEVICE WAS DIFFICULT TO REMOVE AFTER FIRING. THE ANVIL ALSO DID NOT TILT. WHEN THE DEVICE WAS REMOVED, THE TISSUE TORE. THE SURGEON ENDED UP RE-DOING THE ANASTOMOSIS WITH HAND-SEWN TECHNIQUE. THIS RESULTED TO ONE HOUR EXTENDED SURGERY TIME.
04-01-2022 | Injury | us surgical puerto rico | 20-01-2022 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ASSISTED PROCTECTOMY WITH ILEAL ANAL ANASTOMOSIS, WHILE ANASTOMOSING THE COLON, THE STAPLE DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE. THE DEVICE WAS DIFFICULT TO REMOVE AFTER FIRING. THE ANVIL ALSO DID NOT TILT. WHEN THE DEVICE WAS REMOVED, THE TISSUE TORE. THE SURGEON ENDED UP RE-DOING THE ANASTOMOSIS WITH HAND-SEWN TECHNIQUE. THIS RESULTED TO ONE HOUR EXTENDED SURGERY TIME.
15-11-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON ATTEMPTED TO MAKE A CUT USING THE DEVICE BUT THE STOPPED MID-WAY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-11-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON ATTEMPTED TO MAKE A CUT USING THE DEVICE BUT THE STOPPED MID-WAY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-11-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON ATTEMPTED TO MAKE A CUT USING THE DEVICE BUT THE STOPPED MID-WAY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-11-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SIGMOID COLECTOMY, THE SURGEON ATTEMPTED TO MAKE A CUT USING THE DEVICE BUT THE STOPPED MID-WAY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
31-12-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, A COMPONENT DISENGAGED FROM THE DEVICE AND THE DEVICE PARTIALLY FIRED.
31-12-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, A COMPONENT DISENGAGED FROM THE DEVICE AND THE DEVICE PARTIALLY FIRED.
31-12-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, A COMPONENT DISENGAGED FROM THE DEVICE AND THE DEVICE PARTIALLY FIRED.
31-12-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, A COMPONENT DISENGAGED FROM THE DEVICE AND THE DEVICE PARTIALLY FIRED.
03-01-2022 | Injury | covidien lp llc north haven | 20-01-2022 | GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, WHILE TRANSECTING THE DUODENUM DURING A WHIPPLE SURGERY, THE SECOND RELOAD DID NOT FORM PROPERLY RESULTING TO STAPLE LINE BLEEDING. THE SURGEON OVERSEWED THE STAPLE LINE TO RESOLVE THE ISSUE.
03-01-2022 | Injury | covidien lp llc north haven | 20-01-2022 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, WHILE TRANSECTING THE DUODENUM DURING A WHIPPLE SURGERY, THE SECOND RELOAD DID NOT FORM PROPERLY RESULTING TO STAPLE LINE BLEEDING. THE SURGEON OVERSEWED THE STAPLE LINE TO RESOLVE THE ISSUE.
03-04-2021 | Injury | covidien lp llc north haven | 20-01-2022 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN JANUARY 2015 TO MAY 2020, A RETROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT ENDOSCOPIC ULTRASOUND-GUIDED GASTROENTEROSTOMY (EUS-GE) AND LAPAROSCOPIC GASTROENTEROSTOMY (L-GE) FOR GASTRIC OUTLET OBSTRUCTION. 77 PATIENTS WERE TREATED WITH EUS-GE AND 48 PATIENTS WITH L-GE. A 60-MM STAPLER OR NON-COMPANY STAPLER WAS USED IN PATIENTS UNDERGOING L-GE. COMPANY PRODUCTS WERE NOT USED IN EUS-GE. POSTOPERATIVE COMPLICATIONS INCLUDED: ANASTOMOTIC LEAKS ON 3 PATIENTS, ANASTOMOTIC BLEEDING ON 2 INDIVIDUALS AND DYSFUNCTIONAL SURGICAL ANASTOMOSIS. SURGICAL REINTERVENTION ON 3 PEOPLE AND ENDOSCOPIC REINTERVENTIONS ON 4 PATIENTS WERE REQUIRED TO TREAT POSTOPERATIVE BLEEDING AND ANASTOMOTIC LEAK. PLACEMENT OF A TRANS ANASTOMOTIC STENT TO TREAT A DYSFUNCTIONAL SURGICAL ANASTOMOSIS WAS NOTED. EXTENDED HOSPITAL STAY OF A MEAN OF 8 DAYS WAS REPORTED.
03-04-2021 | Injury | covidien lp llc north haven | 20-01-2022 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN JANUARY 2015 TO MAY 2020, A RETROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT ENDOSCOPIC ULTRASOUND-GUIDED GASTROENTEROSTOMY (EUS-GE) AND LAPAROSCOPIC GASTROENTEROSTOMY (L-GE) FOR GASTRIC OUTLET OBSTRUCTION. 77 PATIENTS WERE TREATED WITH EUS-GE AND 48 PATIENTS WITH L-GE. A 60-MM STAPLER OR NON-COMPANY STAPLER WAS USED IN PATIENTS UNDERGOING L-GE. COMPANY PRODUCTS WERE NOT USED IN EUS-GE. POSTOPERATIVE COMPLICATIONS INCLUDED: ANASTOMOTIC LEAKS ON 3 PATIENTS, ANASTOMOTIC BLEEDING ON 2 INDIVIDUALS AND DYSFUNCTIONAL SURGICAL ANASTOMOSIS. SURGICAL REINTERVENTION ON 3 PEOPLE AND ENDOSCOPIC REINTERVENTIONS ON 4 PATIENTS WERE REQUIRED TO TREAT POSTOPERATIVE BLEEDING AND ANASTOMOTIC LEAK. PLACEMENT OF A TRANS ANASTOMOTIC STENT TO TREAT A DYSFUNCTIONAL SURGICAL ANASTOMOSIS WAS NOTED. EXTENDED HOSPITAL STAY OF A MEAN OF 8 DAYS WAS REPORTED.
03-04-2021 | Injury | covidien lp llc north haven | 20-01-2022 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN JANUARY 2015 TO MAY 2020, A RETROSPECTIVE STUDY ANALYZED OUTCOMES OF PATIENTS WHO UNDERWENT ENDOSCOPIC ULTRASOUND-GUIDED GASTROENTEROSTOMY (EUS-GE) AND LAPAROSCOPIC GASTROENTEROSTOMY (L-GE) FOR GASTRIC OUTLET OBSTRUCTION. 77 PATIENTS WERE TREATED WITH EUS-GE AND 48 PATIENTS WITH L-GE. A 60-MM STAPLER OR NON-COMPANY STAPLER WAS USED IN PATIENTS UNDERGOING L-GE. COMPANY PRODUCTS WERE NOT USED IN EUS-GE. POSTOPERATIVE COMPLICATIONS INCLUDED: ANASTOMOTIC LEAKS ON 3 PATIENTS, ANASTOMOTIC BLEEDING ON 2 INDIVIDUALS AND DYSFUNCTIONAL SURGICAL ANASTOMOSIS. SURGICAL REINTERVENTION ON 3 PEOPLE AND ENDOSCOPIC REINTERVENTIONS ON 4 PATIENTS WERE REQUIRED TO TREAT POSTOPERATIVE BLEEDING AND ANASTOMOTIC LEAK. PLACEMENT OF A TRANS ANASTOMOTIC STENT TO TREAT A DYSFUNCTIONAL SURGICAL ANASTOMOSIS WAS NOTED. EXTENDED HOSPITAL STAY OF A MEAN OF 8 DAYS WAS REPORTED.
12-08-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A SURGICAL PROCEDURE FOR LUNG CANCER, THE STAPLER COULD NOT BE FIRED DURING CLOSURE. A NEW RELOAD WAS USED AS A REPLACEMENT. THERE WAS NO PATIENT INJURY. FROM (B)(4)- ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPY LUNG CANCER SURGERY, WHILE DISSECTING THE LUNG TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-09-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER COULD NOT BE CLOSED AFTER RELOADING THE HANDLE TO FIRE THE LUNG TISSUE. REPLACE THE HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC CARDIA CANCER RESECTION, WHILE RESECTING THE CARDIA FROM THE LESSER CURVATURE OF THE STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | covidien lp llc north haven | 20-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE RESECTING THE COLON, THE DEVICE WAS UNABLE TO FIRE THE SECOND FIRING AS THE RELOAD WAS DIFFICULT TO INSTALL. THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE. THE DEVICE WAS NOT USED IN PATIENT. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | covidien lp llc north haven | 20-01-2022 | GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE RESECTING THE COLON, THE DEVICE WAS UNABLE TO FIRE THE SECOND FIRING AS THE RELOAD WAS DIFFICULT TO INSTALL. THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE. THE DEVICE WAS NOT USED IN PATIENT. THERE WAS NO PATIENT INJURY.
16-09-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHILE CHARGING, THE CHARGER FLASHED RED AND WHEN THE HANDLE WAS REMOVED IT DID NOT START-UP. ANOTHER CHARGER WAS USED BUT THE SAME ISSUE OCCURRED. IT WAS NOTED THAT THE CHARGER FUNCTIONED NORMALLY WITH OTHER HANDLE. THE SURGEON THEN USED A NEW HANDLE TO RESOLVE THE ISSUE.
16-09-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHILE CHARGING, THE CHARGER FLASHED RED AND WHEN THE HANDLE WAS REMOVED IT DID NOT START-UP. ANOTHER CHARGER WAS USED BUT THE SAME ISSUE OCCURRED. IT WAS NOTED THAT THE CHARGER FUNCTIONED NORMALLY WITH OTHER HANDLE. THE SURGEON THEN USED A NEW HANDLE TO RESOLVE THE ISSUE.
21-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING THE COLON RESECTION, AFTER STAPLE FIRE, STAPLE LINE LOOKED FINE BUT WHEN INSPECTING INSIDE THE BOWEL, THE SURGEON DISCOVERED THAT THERE WAS A HOLE. THE INCISION WAS EXTENDED BUT NOT BY MORE THAN ONE INCH. THE SURGEON SUTURED THE HOLE AND DECIDED TO CREATE A DIVERTING COLOSTOMY FOR BOWEL HEALING AND PLANS TO BRING PATIENT BACK FOR ANOTHER PROCEDURE TAKEDOWN COLOSTOMY.
21-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE COLON RESECTION, AFTER STAPLE FIRE, STAPLE LINE LOOKED FINE BUT WHEN INSPECTING INSIDE THE BOWEL, THE SURGEON DISCOVERED THAT THERE WAS A HOLE. THE INCISION WAS EXTENDED BUT NOT BY MORE THAN ONE INCH. THE SURGEON SUTURED THE HOLE AND DECIDED TO CREATE A DIVERTING COLOSTOMY FOR BOWEL HEALING AND PLANS TO BRING PATIENT BACK FOR ANOTHER PROCEDURE TAKEDOWN COLOSTOMY.
27-12-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2022 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR (TAPP) PROCEDURE ON (B)(6) 2021, THE BARD/DAVOL OPTIFIX STRAIGHT FIXATION DEVICE WAS USED TO FIXATE 3DMAX LIGHT MESH, THE FASTENER WAS BROKEN AND THE BARREL INSERT WAS DETACHED. IT WAS REPORTED THAT THE BROKEN FASTENER WAS REMOVED FROM THE PATIENT'S BODY AND ANOTHER OPTIFIX DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY.
27-12-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2022 | OPTIFIX STRAIGHT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR (TAPP) PROCEDURE ON (B)(6) 2021, THE BARD/DAVOL OPTIFIX STRAIGHT FIXATION DEVICE WAS USED TO FIXATE 3DMAX LIGHT MESH, THE FASTENER WAS BROKEN AND THE BARREL INSERT WAS DETACHED. IT WAS REPORTED THAT THE BROKEN FASTENER WAS REMOVED FROM THE PATIENT'S BODY AND ANOTHER OPTIFIX DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY.
15-12-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | EEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE THE, THE BAYONET WAS NOT FIXED, THUS BEING DETACHED, IT ACTIVATED BEFORE THE OCCURRENCE. THERE WAS NO PATIENT INJURY.
05-01-2022 | Injury | us surgical puerto rico | 20-01-2022 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC COLORECTAL ANASTOMOSIS, THE STAPLE LINE WAS INCOMPLETE. THE DEVICE PARTIALLY FIRED. THE DEVICE WAS CHANGED TO RESOLVE THE ISSUE ON ORDER TO COMPLETE THE CASE.
04-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE BACK PUSHER DID NOT MOVE AT ALL. THERE WAS NO PATIENT INVOLVEMENT.
05-01-2022 | Malfunction | covidien lp llc north haven | 20-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MINI GASTRIC BYPASS PROCEDURE, WHILE STAPLING ON THE STOMACH, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA | Break | A0401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, THE TWO HANDLES WERE ABLE TO FIRE. THE STAPLES DID NOT FORM PROPERLY ON THE THREE RELOADS AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
27-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, THE TWO HANDLES WERE ABLE TO FIRE. THE STAPLES DID NOT FORM PROPERLY ON THE THREE RELOADS AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
27-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, THE TWO HANDLES WERE ABLE TO FIRE. THE STAPLES DID NOT FORM PROPERLY ON THE THREE RELOADS AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
27-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, THE TWO HANDLES WERE ABLE TO FIRE. THE STAPLES DID NOT FORM PROPERLY ON THE THREE RELOADS AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
27-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, THE TWO HANDLES WERE ABLE TO FIRE. THE STAPLES DID NOT FORM PROPERLY ON THE THREE RELOADS AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
27-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, THE TWO HANDLES WERE ABLE TO FIRE. THE STAPLES DID NOT FORM PROPERLY ON THE THREE RELOADS AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
27-12-2021 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT FORM PROPERLY, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO FIX THE STAPLE LINE.
06-04-2022 | Malfunction | covidien lp llc north haven | 20-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROPLASTY, THE HANDLE WAS ABLE TO RECOGNIZE THE RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING DIGITAL GASTRECTOMY, THE JAWS OF THE DEVICE WOULD NOT OPEN AFTER STOMACH RESECTION.
04-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DIGITAL GASTRECTOMY, THE JAWS OF THE DEVICE WOULD NOT OPEN AFTER STOMACH RESECTION.
29-12-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE JAWS WERE OPENED INSIDE THE ABDOMINAL CAVITY, THE SHEET ON THE TIP SIDE OF THE RELOAD CAME OFF AND COULD NOT BE USED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
05-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL PANCREATECTOMY WHILE REMOVING THE PANCREAS BODY AND TAIL, THE DEVICE CANNOT BE FIRED NORMALLY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. TO RESOLVE THE ISSUE, BOTH HANDLE AND RELOAD WERE REPLACED WITH ANOTHER MANUFACTURER'S PRODUCT. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DISTAL PANCREATECTOMY WHILE REMOVING THE PANCREAS BODY AND TAIL, THE DEVICE CANNOT BE FIRED NORMALLY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. TO RESOLVE THE ISSUE, BOTH HANDLE AND RELOAD WERE REPLACED WITH ANOTHER MANUFACTURER'S PRODUCT. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | covidien lp llc north haven | 20-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLES DID NOT FORM PROPERLY ON TWO DEVICES.
05-01-2022 | Malfunction | covidien lp llc north haven | 20-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLES DID NOT FORM PROPERLY ON TWO DEVICES.
05-01-2022 | Injury | celestica de monterrey s.a de c.v | 20-01-2022 | ENDO GIA | Misfire | A050502 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL SURGERY FOR RECTAL CANCER, WHEN DISSECTING RECTAL TISSUE, PART OF STAPLES DID NOT FORM AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE SURGEON USED A NEW DEVICE AND SUTURE WAS USED ENDOSCOPICALLY FOR REPAIR.
05-01-2022 | Injury | celestica de monterrey s.a de c.v | 20-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL SURGERY FOR RECTAL CANCER, WHEN DISSECTING RECTAL TISSUE, PART OF STAPLES DID NOT FORM AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE SURGEON USED A NEW DEVICE AND SUTURE WAS USED ENDOSCOPICALLY FOR REPAIR.
05-01-2022 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA ULTRA | Misfire | A050502 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL SURGERY FOR RECTAL CANCER, WHEN DISSECTING RECTAL TISSUE, PART OF STAPLES DID NOT FORM AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE SURGEON USED A NEW DEVICE AND SUTURE WAS USED ENDOSCOPICALLY FOR REPAIR.
22-11-2021 | Malfunction | davis & geck caribe | 20-01-2022 | UNKNOWN ENDO STITCH SULU | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE BROKE WHILE GRABBING THE POSTERIOR LEFT VAGINAL CUFF. THE SUTURE WAS REMOVED AND REPLACED, BUT SUTURE SNAPPED. A NEW HANDLE AND NEW RELOAD WERE OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE GUN WAS STUCK. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Excessive Heating | A1002 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, NOTED THE BATTERY WAS WITH ABNORMAL HIGH TEMPERATURE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
27-12-2021 | Injury | ethicon endo-surgery | 20-01-2022 | ILS 33MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST OP OF A LOW ANTERIOR RESECTION, PATIENT WITH POST-OP BLEED. NO FURTHER INFORMATION WAS PROVIDED.
16-11-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT FIRE. NO HARM TO PATIENT.
29-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | CONTOUR CURVED CUTTER STAPLER | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DID NOT CUT THE TISSUE. THERE WERE NO PATIENT CONSEQUENCES.
29-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | CONTOUR CURVED CUTTER STAPLER | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, DID NOT CUT THE TISSUE. THERE WERE NO PATIENT CONSEQUENCES.
27-12-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2022 | SORBAFIX ENHANCED | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR (TAPP) PROCEDURE ON (B)(6) 2021, THE BARD/DAVOL SORBAFIX ENHANCED FIXATION DEVICE WAS USED TO FIXATE A BARD/DAVOL 3DMAX LIGHT MESH. IT WAS REPORTED THAT THE DEVICE DEPLOYED A BROKEN FASTENER AND DID NOT FIXATE THE MESH. IT WAS ALSO REPORTED THAT THE BROKEN FASTENER WAS REMOVED FROM THE PATIENT'S BODY AND ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY.
16-11-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | GST BLUE RELOAD, 45MM, 6 ROW | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THE CARTRIDGE PAN SEPARATION AFTER OPENED THE PACKING. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
29-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | CIRCULAR MECH XL SEAL 29MM | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL PROCEDURE THAT AFTER ATTACHING TO THE ANVIL THE DEVICE WAS FIRED BUT THEY WERE NOT ABLE TO REMOVE THE DEVICE FROM THE PATIENT. REP GAVE SURGEON INSTRUCTIONS OVER THE PHONE, BUT SURGEON STATED HE WAS PROBABLY GOING TO HAVE TO CUT THE DEVICE OUT TO REMOVE IT. SURGEON STATED THAT THEY DID NOT HEAR THE WASHER BREAK. THE CASE WAS COMPLETED WITH A THIRD DEVICE, A 31 POWERED CIRCULAR. TO COMPLETE A DIFFERENT ANASTOMOSIS. THERE WERE CHECKING FOR A LEAK AND WERE ABLE TO GET THE DEVICE OUT. UNKNOWN THE RESULTS OF THE LEAK. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
29-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | CIRCULAR MECH XL SEAL 29MM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL PROCEDURE THAT AFTER ATTACHING TO THE ANVIL THE DEVICE WAS FIRED BUT THEY WERE NOT ABLE TO REMOVE THE DEVICE FROM THE PATIENT. REP GAVE SURGEON INSTRUCTIONS OVER THE PHONE, BUT SURGEON STATED HE WAS PROBABLY GOING TO HAVE TO CUT THE DEVICE OUT TO REMOVE IT. SURGEON STATED THAT THEY DID NOT HEAR THE WASHER BREAK. THE CASE WAS COMPLETED WITH A THIRD DEVICE, A 31 POWERED CIRCULAR. TO COMPLETE A DIFFERENT ANASTOMOSIS. THERE WERE CHECKING FOR A LEAK AND WERE ABLE TO GET THE DEVICE OUT. UNKNOWN THE RESULTS OF THE LEAK. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
29-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | CIRCULAR MECH XL SEAL 29MM | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL PROCEDURE THAT AFTER ATTACHING TO THE ANVIL THE DEVICE WAS FIRED BUT THEY WERE NOT ABLE TO REMOVE THE DEVICE FROM THE PATIENT. REP GAVE SURGEON INSTRUCTIONS OVER THE PHONE, BUT SURGEON STATED HE WAS PROBABLY GOING TO HAVE TO CUT THE DEVICE OUT TO REMOVE IT. SURGEON STATED THAT THEY DID NOT HEAR THE WASHER BREAK. THE CASE WAS COMPLETED WITH A THIRD DEVICE, A 31 POWERED CIRCULAR. TO COMPLETE A DIFFERENT ANASTOMOSIS. THERE WERE CHECKING FOR A LEAK AND WERE ABLE TO GET THE DEVICE OUT. UNKNOWN THE RESULTS OF THE LEAK. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
29-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | CIRCULAR MECH XL SEAL 29MM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL PROCEDURE THAT AFTER ATTACHING TO THE ANVIL THE DEVICE WAS FIRED BUT THEY WERE NOT ABLE TO REMOVE THE DEVICE FROM THE PATIENT. REP GAVE SURGEON INSTRUCTIONS OVER THE PHONE, BUT SURGEON STATED HE WAS PROBABLY GOING TO HAVE TO CUT THE DEVICE OUT TO REMOVE IT. SURGEON STATED THAT THEY DID NOT HEAR THE WASHER BREAK. THE CASE WAS COMPLETED WITH A THIRD DEVICE, A 31 POWERED CIRCULAR. TO COMPLETE A DIFFERENT ANASTOMOSIS. THERE WERE CHECKING FOR A LEAK AND WERE ABLE TO GET THE DEVICE OUT. UNKNOWN THE RESULTS OF THE LEAK. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
16-12-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2022 | OPTIFIX STRAIGHT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNKNOWN PROCEDURE ON (B)(6) 2021, A BARD/DAVOL OPTIFIX STRAIGHT FIXATION DEVICE WAS USED TO FIXATE THE UNKNOWN MESH TO THE AREA OF HARD BONE PART. IT WAS REPORTED THAT THE DEVICE WAS DAMAGED AND FAILED TO DEPLOY A FASTENER. THERE WAS NO REPORTED PATIENT INJURY. EVALUATION OF THE RETURNED SAMPLE FOUND THE BARREL INSERT IS DETACHED FROM THE CANNULA TIP, AS SUCH WE ARE SUBMITTING THIS MALFUNCTION MDR TO DOCUMENT OUR FINDINGS.
22-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY, A MOTOR SOUND WAS HEARD A SECOND BUT THEN THE KNIFE STOPPED MOVING FORWARD. THE KNIFE REVERSE SWITCH WAS USED TO RELEASE THE DEVICE. THE DEVICE WAS USED ON THE INFERIOR PULMONARY ARTERY. THE STAPLES WERE NOT DEPLOYED OR THE KNIFE DID NOT CUT THE TARGET TISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | BLUE RELOAD FOR ECHELON 60 | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WAS A MISFIRED RELOAD WHILE STAPLING IN SLEEVE GASTRECTOMY CASE. THERE WAS A 20 MINUTE DELAY IN SURGERY AND THERE WERE NO PATIENT CONSEQUENCES.
01-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WAS A MISFIRED RELOAD WHILE STAPLING IN SLEEVE GASTRECTOMY CASE. THERE WAS A 20 MINUTE DELAY IN SURGERY AND THERE WERE NO PATIENT CONSEQUENCES.
11-11-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | RELOADS FOR LINEAR CUTTERS | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE COULD NOT BE FIRED AT THE 1ST FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
28-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | ILS 25MM, STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC PULMONARY RESECTION, SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HANDLE COULD NOT BE SQUEEZED. DEVICE WAS NOT ABLE TO FIRE. SURGEON USED ANOTHER RELOAD FOR THE SAME HANDLE FUNCTIONED WITH OTHER RELOADS. NO PATIENT INJURY.
08-12-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC PULMONARY RESECTION, SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT HANDLE COULD NOT BE SQUEEZED. DEVICE WAS NOT ABLE TO FIRE. SURGEON USED ANOTHER RELOAD FOR THE SAME HANDLE FUNCTIONED WITH OTHER RELOADS. NO PATIENT INJURY.
29-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCTRECTOMY SURGERY, AFTER FIRED, NOTED THAT THE TISSUE WAS NOT CUT, BUT THERE WERE SOME STAPLES DEPLOYED ON THE TISSUE, THEN CHANGED ANOTHER DEVICE TO COMPLETE THE PROCEDURE.
11-10-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE DEVICE WOULD NOT CHARGE COMPLETELY. EVEN WHEN HANDLE WAS RE-INSERTED, THE CHARGER FLASHED GREEN OR RED. THE CHARGER WORKED FINE WITH ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT AN ANALYSIS OF THE SYSTEM LOGS NOTED THAT THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
11-10-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE DEVICE WOULD NOT CHARGE COMPLETELY. EVEN WHEN HANDLE WAS RE-INSERTED, THE CHARGER FLASHED GREEN OR RED. THE CHARGER WORKED FINE WITH ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT AN ANALYSIS OF THE SYSTEM LOGS NOTED THAT THE HANDLE WAS OPENED UP AND BOTH BATTERIES WERE FOUND TO BE VENTED.
10-01-2022 | Malfunction | ethicon endo-surgery | 20-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-11-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | WHITE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC SEGMENTAL LUNG SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-01-2021 | Injury | ethicon endo-surgery | 20-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | Failure to Anastomose | Event Description: IT WAS REPORTED DURING A LOWER ANTERIOR BOWEL RESECTION PROCEDURE, STAPLER DIDN¿T FIRE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
29-01-2021 | Injury | ethicon endo-surgery | 20-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED DURING A LOWER ANTERIOR BOWEL RESECTION PROCEDURE, STAPLER DIDN¿T FIRE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
29-01-2021 | Injury | ethicon endo-surgery | 20-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Eject | A050301 | Failure to Anastomose | Event Description: IT WAS REPORTED DURING A LOWER ANTERIOR BOWEL RESECTION PROCEDURE, STAPLER DIDN¿T FIRE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
01-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | PROXIMATE*RL LINEAR STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE HAD MADE A STRANGE NOISE WHEN STAPLING AND CONSEQUENTLY A BAD STAPLING (DISCONTINUITY OF THE STAPLING LINE) -NOT TIGHT. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES.
01-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE HAD MADE A STRANGE NOISE WHEN STAPLING AND CONSEQUENTLY A BAD STAPLING (DISCONTINUITY OF THE STAPLING LINE) -NOT TIGHT. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES.
01-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | PROXIMATE*RL LINEAR STAPLER | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE HAD MADE A STRANGE NOISE WHEN STAPLING AND CONSEQUENTLY A BAD STAPLING (DISCONTINUITY OF THE STAPLING LINE) -NOT TIGHT. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES.
14-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE DID NOT RETURN TO HOME POSITION. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
14-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE DID NOT RETURN TO HOME POSITION. THE MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
22-12-2021 | Injury | ethicon endo-surgery | 20-01-2022 | GST WHITE RELOAD, 45MM, 6 ROW | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST OP TO A GASTRIC BYPASS SURGERY, THE DOCTOR INDICATES THAT A PREVIOUSLY OPERATED PATIENT IS ADMITTED TO THE INTERMEDIATE SERVICE DUE TO DIGESTIVE BLEEDING. IT REFERS TO "THE BLEEDING COMES FROM ONE OF THE ANASTOMOSES". THE PATIENT IS BEING TREATED WITH WELL-TOLERATED TRANEXAMIC ACID, TO DATE WITHOUT THE NEED FOR REOPERATION.
22-12-2021 | Injury | ethicon endo-surgery | 20-01-2022 | GST WHITE RELOAD, 45MM, 6 ROW | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST OP TO A GASTRIC BYPASS SURGERY, THE DOCTOR INDICATES THAT A PREVIOUSLY OPERATED PATIENT IS ADMITTED TO THE INTERMEDIATE SERVICE DUE TO DIGESTIVE BLEEDING. IT REFERS TO "THE BLEEDING COMES FROM ONE OF THE ANASTOMOSES". THE PATIENT IS BEING TREATED WITH WELL-TOLERATED TRANEXAMIC ACID, TO DATE WITHOUT THE NEED FOR REOPERATION.
26-12-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE IN TISSUE DURING A LAPAROSCOPIC APPENDECTOMY, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZE THE HANDLE. TO RESOLVE THE ISSUE, THE SURGEON REPLACED IT WITH A NEW HANDLE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Perforation | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD A LAPAROSCOPIC BAND REMOVAL ON (B)(6) 2021 THEN HAD A LAPAROSCOPIC SLEEVE ON (B)(6) 2021. ON (B)(6) 2022, THE PATIENT WAS BROUGHT DUE TO A LEAK. WHEN WENT IN, THE STAPLE LINE WAS PERFORATED AT THE FUNDUS OF THE STOMACH. THE SURGEON WAS ABLE TO FIX THE LEAK USING A PURPLE RELOAD. THE PATIENT WAS STILL IN THE HOSPITAL AS OF (B)(6) 2022.
03-01-2022 | Injury | us surgical puerto rico | 20-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Perforation | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD A LAPAROSCOPIC BAND REMOVAL ON (B)(6) 2021 THEN HAD A LAPAROSCOPIC SLEEVE ON (B)(6) 2021. ON (B)(6) 2022, THE PATIENT WAS BROUGHT DUE TO A LEAK. WHEN WENT IN, THE STAPLE LINE WAS PERFORATED AT THE FUNDUS OF THE STOMACH. THE SURGEON WAS ABLE TO FIX THE LEAK USING A PURPLE RELOAD. THE PATIENT WAS STILL IN THE HOSPITAL AS OF (B)(6) 2022.
12-12-2021 | Malfunction | covidien lp llc north haven | 20-01-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
12-12-2021 | Malfunction | us surgical puerto rico | 20-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
04-01-2022 | Malfunction | covidien lp llc north haven | 20-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, THE RELOAD WAS LOCKED AND DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, THE RELOAD WAS LOCKED AND DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, THE RELOAD WAS LOCKED AND DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, WHILE REMOVING THE RECTUM, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, WHILE REMOVING THE RECTUM, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | covidien lp llc north haven | 20-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC PERCUTANEOUS FISTULA CLOSURE, WHILE FIRING ACROSS THE STOMACH, THE TISSUE WAS VERY THICK AND THE STAPLE FIRE NEVER ENGAGED FULLY TO FIRE AND THE HANDLE MADE A CRACKING SOUND. THE SURGEON THEN USED A BIGGER RELOAD AND A NEW HANDLE WHERE WHEN FIRED THE HANDLE CRACKED AGAIN BUT WERE ABLE TO CONTINUE THE FIRING THROUGH TO THE END. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN TOTAL MESORECTAL EXCISION, ON RECTAL CLOSURE AND DETACHMENT, THE DEVICE COULD NOT CUT AND CLOSE. A NEW RELOAD WAS USED TO RESOLVE THE CONCERN. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN TOTAL MESORECTAL EXCISION, ON RECTAL CLOSURE AND DETACHMENT, THE DEVICE COULD NOT CUT AND CLOSE. A NEW RELOAD WAS USED TO RESOLVE THE CONCERN. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | us surgical puerto rico | 20-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN TOTAL MESORECTAL EXCISION, ON RECTAL CLOSURE AND DETACHMENT, THE DEVICE COULD NOT CUT AND CLOSE. A NEW RELOAD WAS USED TO RESOLVE THE CONCERN. THERE WAS NO PATIENT INJURY.
01-12-2021 | Malfunction | ethicon endo-surgery | 20-01-2022 | ENDOPATH ETS FLEX 45 | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: DURING SURGICAL PROCEDURE FOR NEPHRECTOMY, A SINGLE STAPLE LOAD WAS FIRED OVER THE RENAL HILUM. THE VEINS WERE DIVIDED AND WELL STAPLED BUT THE ARTERY ESCAPED IN A CEPHALAD DIRECTION. A 2ND STAPLE LOAD WAS INTRODUCED AND FIRED BUT THIS ONE MISFIRED AND THE STAPLE LOAD COULD NOT BE OPENED TO BE RELEASED. MANIPULATION DEMONSTRATED BRISK ARTERIAL BLEEDING CONSISTENT WITH AN INJURY TO THE RENAL ARTERY AT THE LEVEL OF THE MISFIRED STAPLE LOAD. FDA SAFETY REPORT ID# (B)(4).
06-12-2021 | Malfunction | us surgical puerto rico | 21-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHILE STAPLING THE LUNG PARENCHYMA, THE 2 STAPLERS WERE UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2021 | Malfunction | covidien lp llc north haven | 21-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING REMOVAL OF THE APPENDIX IN A LAPAROSCOPIC APPENDECTOMY, THE DEVICE HAD POOR STAPLE FORMATION. STAPLES THAT WERE FORMED BEYOND THE SPECIMEN FELL INTO THE ABDOMINAL CAVITY. USED SUCTION/IRRIGATOR OR LAP INSTRUMENTS TO COLLECT THE LOOSE STAPLES AND REMOVE THEM FROM THE PATIENT. NO PATIENT INJURY.
22-12-2021 | Malfunction | covidien lp llc north haven | 21-01-2022 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING REMOVAL OF THE APPENDIX IN A LAPAROSCOPIC APPENDECTOMY, THE DEVICE HAD POOR STAPLE FORMATION. STAPLES THAT WERE FORMED BEYOND THE SPECIMEN FELL INTO THE ABDOMINAL CAVITY. USED SUCTION/IRRIGATOR OR LAP INSTRUMENTS TO COLLECT THE LOOSE STAPLES AND REMOVE THEM FROM THE PATIENT. NO PATIENT INJURY.
11-11-2021 | Malfunction | covidien lp llc north haven | 21-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE.
11-11-2021 | Malfunction | us surgical puerto rico | 21-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE.
05-01-2022 | Malfunction | us surgical puerto rico | 21-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF THE LEFT COLON, ON REMOVAL, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | us surgical puerto rico | 21-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF THE LEFT COLON, ON REMOVAL, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | us surgical puerto rico | 21-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHEN DIVIDING THE TISSUE, THE DEVICES DID NOT FIRE. THE SURGEON WAS BALE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | celestica de monterrey s.a de c.v | 21-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHEN DIVIDING THE TISSUE, THE DEVICES DID NOT FIRE. THE SURGEON WAS BALE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Injury | celestica de monterrey s.a de c.v | 21-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR LUNG CANCER, WHILE DETACHING THE LUNG TISSUE, THE STAPLES DID NOT FORM PROPERLY AND THE DEVICE PARTIALLY FIRED. IN WHICH RESULTED TO A INCOMPLETE STAPLE LINE DISTALLY. THE SURGEON USED A NEW RELOAD AND SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE.
06-01-2022 | Injury | celestica de monterrey s.a de c.v | 21-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR LUNG CANCER, WHILE DETACHING THE LUNG TISSUE, THE STAPLES DID NOT FORM PROPERLY AND THE DEVICE PARTIALLY FIRED. IN WHICH RESULTED TO A INCOMPLETE STAPLE LINE DISTALLY. THE SURGEON USED A NEW RELOAD AND SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE.
06-01-2022 | Malfunction | covidien lp llc north haven | 21-01-2022 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOIDS, AFTER PURSTRING STAGE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2021 | Malfunction | us surgical puerto rico | 21-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE PROCEDURE THE DEVICE DID NOT FIRE. ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZED THE HANDLE BUT THE IT DID NOT DEPLOY. THERE WAS NO PATIENT INVOLVED IN THE EVENT.
27-12-2021 | Malfunction | us surgical puerto rico | 21-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE PROCEDURE THE DEVICE DID NOT FIRE. ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZED THE HANDLE BUT THE IT DID NOT DEPLOY. THERE WAS NO PATIENT INVOLVED IN THE EVENT.
25-11-2021 | Malfunction | us surgical puerto rico | 21-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS ABLE TO FIRE AND THE STAPLES WERE FINE UPON INSPECTION; HOWEVER, A LOUD SOUND WAS HEARD DURING THE SECOND FIRING. THE STAPLE LINE WAS PERFECT AND THE SURGICAL TEAM DID NOT THINK THAT THE RELOAD WAS AT FAULT, RATHER THE TISSUE WAS VERY THICK. A NEW STAPLER WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-11-2021 | Malfunction | us surgical puerto rico | 21-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, THE STAPLER WAS ABLE TO FIRE AND THE STAPLES WERE FINE UPON INSPECTION; HOWEVER, A LOUD SOUND WAS HEARD DURING THE SECOND FIRING. THE STAPLE LINE WAS PERFECT AND THE SURGICAL TEAM DID NOT THINK THAT THE RELOAD WAS AT FAULT, RATHER THE TISSUE WAS VERY THICK. A NEW STAPLER WAS OPENED AND USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Injury | us surgical puerto rico | 21-01-2022 | ENDO GIA | Break | A0401 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, AFTER FIRING THE LOADING UNIT, THE DEVICE WAS BROKEN, AND THE COMPONENT DISENGAGED INTO CAVITY. AN X-RAY WAS PERFORMED FOR THE EXPRESS PURPOSE OF LOCATING THE COMPONENT. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. ALL THE BROKEN COMPONENTS WERE RETRIEVED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
04-01-2022 | Injury | us surgical puerto rico | 21-01-2022 | ENDO GIA | Detachment of Device or Device Component | A0501 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, AFTER FIRING THE LOADING UNIT, THE DEVICE WAS BROKEN, AND THE COMPONENT DISENGAGED INTO CAVITY. AN X-RAY WAS PERFORMED FOR THE EXPRESS PURPOSE OF LOCATING THE COMPONENT. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE. ALL THE BROKEN COMPONENTS WERE RETRIEVED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
06-01-2022 | Malfunction | covidien lp llc north haven | 21-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCTECTOMY, WHILE TRANSECTING THE RECTUS, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN WHEN THE GREEN BUTTON WAS PRESSED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | ethicon endo-surgery | 21-01-2022 | POWERED PLUS LONG 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WAS A STAPLE MALFORMATION. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
28-12-2021 | Malfunction | covidien lp llc north haven | 21-01-2022 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE FIRST STEP IN A INGUINAL HERNIA REPAIR BY VIDEOLAPAROSCOPY, WHEN THE SURGEON TRIED TO USE THE DEVICE, NO TACKS WERE ABLE TO BE FIRED NORMALLY. THERE WAS A 30 MINUTE OR MORE SURGICAL EXTENSION TIME. TO RESOLVE THE ISSUE, THE SURGEON USED ANOTHER DEVICE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | ethicon endo-surgery | 21-01-2022 | ECHELON FLEX 60 COMPACT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
27-12-2021 | Malfunction | covidien lp llc north haven | 21-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE KNIFE DID NOT GO THROUGH AND STOPPED THE RELOAD.
27-12-2021 | Malfunction | us surgical puerto rico | 21-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE KNIFE DID NOT GO THROUGH AND STOPPED THE RELOAD.
12-12-2021 | Malfunction | covidien lp llc north haven | 21-01-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE KNIFE DID NOT GO THROUGH AND THE RELOAD STOPPED.
12-12-2021 | Malfunction | covidien lp llc north haven | 21-01-2022 | ENDO GIA | Failure to Cut | A050702 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE KNIFE DID NOT GO THROUGH AND THE RELOAD STOPPED.
12-12-2021 | Malfunction | us surgical puerto rico | 21-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE KNIFE DID NOT GO THROUGH AND THE RELOAD STOPPED.
12-12-2021 | Malfunction | us surgical puerto rico | 21-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Cut | A050702 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE KNIFE DID NOT GO THROUGH AND THE RELOAD STOPPED.
01-01-2022 | Malfunction | ethicon endo-surgery | 21-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WAS LOCKED OUT WHEN THE KNIFE WAS RETURNING TO HOME POSITION AFTER THE 2ND FIRING. THE KNIFE REVERSE SWITCH WAS USED, BUT THE KNIFE DID NOT RETURN. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 21-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WAS LOCKED OUT WHEN THE KNIFE WAS RETURNING TO HOME POSITION AFTER THE 2ND FIRING. THE KNIFE REVERSE SWITCH WAS USED, BUT THE KNIFE DID NOT RETURN. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 21-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WAS LOCKED OUT WHEN THE KNIFE WAS RETURNING TO HOME POSITION AFTER THE 2ND FIRING. THE KNIFE REVERSE SWITCH WAS USED, BUT THE KNIFE DID NOT RETURN. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
29-12-2021 | Injury | intuitive surgical | 21-01-2022 | SUREFORM | Use of Device Problem | A23 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED BARIATRIC REVISION SURGICAL PROCEDURE (FROM A BAND TO A GASTRIC BYPASS), THE SURGEON REPORTED THAT THE SUREFORM 60 (SF60) STAPLER INSTRUMENT WITH A BLUE SUREFORM 60 RELOAD WAS INSTALLED DURING THE FIRST FIRE. THE STAPLER CLAMPED DOWN ON THE STOMACH WITH NO ISSUES. HOWEVER, WHEN HE PRESSED THE YELLOW PEDAL, THE BLADE JAMMED AND DRAGGED THE TISSUE AND DID NOT CUT TISSUE. THE STAPLES DID NOT FORM CORRECTLY. THERE WAS A THREE CENTIMETER MYOTOMY (HOLE) ON BOTH SIDES OF THE STAPLE LINE (THE GASTRIC POUCH AND THE REMNANT GASTRIC TISSUE). THE SURGEON NOTICED A CLIP FROM THE PREVIOUS LAP BAND SURGERY EMBEDDED ON THE TISSUE HE WAS STAPLING. THE SURGEON THINKS THE CLIP MIGHT HAVE GOTTEN CAUGHT IN FRONT OF THE BLADE AND CAUSED THE STAPLING ISSUE. THE SURGEON REPAIRED THE DEFECT ON THE GASTRIC POUCH, STAPLING MEDIAL TO THE INCOMPLETE STAPLE LINE WITH THE SAME SUREFORM 60 STAPLER INSTRUMENT AND A NEW BLUE SUREFORM 60 RELOAD. AS A RESULT, THE SIZE OF THE POUCH WAS REDUCED. THE SURGEON OVERSEWED THE DEFECT ON THE REMNANT GASTRIC TISSUE. ALTHOUGH THIS STAPLING ISSUE RESULTED IN A SMALLER POUCH, HE DOES NOT BELIEVE THAT WILL AFFECT THE SURGICAL OUTCOME OR IF THE PATIENT WILL HAVE ANY LONG-TERM EFFECT. THE PATIENT IS REPORTED TO BE RECOVERING WELL. THE SURGEON CONFIRMED THE FOLLOWING: THE RELOAD WAS DISCARDED, HE DID NOT USE BUTTRESS MATERIAL, THE (B)(4) STAPLER INSTRUMENT WAS USED FOR THE REMAINDER OF THE PROCEDURE WITH NO ISSUES, THERE IS NO VIDEO RECORDING; HOWEVER, THERE IS AN IMAGE THAT WAS PROVIDED FOR REVIEW.
29-12-2021 | Injury | intuitive surgical | 21-01-2022 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED BARIATRIC REVISION SURGICAL PROCEDURE (FROM A BAND TO A GASTRIC BYPASS), THE SURGEON REPORTED THAT THE SUREFORM 60 (SF60) STAPLER INSTRUMENT WITH A BLUE SUREFORM 60 RELOAD WAS INSTALLED DURING THE FIRST FIRE. THE STAPLER CLAMPED DOWN ON THE STOMACH WITH NO ISSUES. HOWEVER, WHEN HE PRESSED THE YELLOW PEDAL, THE BLADE JAMMED AND DRAGGED THE TISSUE AND DID NOT CUT TISSUE. THE STAPLES DID NOT FORM CORRECTLY. THERE WAS A THREE CENTIMETER MYOTOMY (HOLE) ON BOTH SIDES OF THE STAPLE LINE (THE GASTRIC POUCH AND THE REMNANT GASTRIC TISSUE). THE SURGEON NOTICED A CLIP FROM THE PREVIOUS LAP BAND SURGERY EMBEDDED ON THE TISSUE HE WAS STAPLING. THE SURGEON THINKS THE CLIP MIGHT HAVE GOTTEN CAUGHT IN FRONT OF THE BLADE AND CAUSED THE STAPLING ISSUE. THE SURGEON REPAIRED THE DEFECT ON THE GASTRIC POUCH, STAPLING MEDIAL TO THE INCOMPLETE STAPLE LINE WITH THE SAME SUREFORM 60 STAPLER INSTRUMENT AND A NEW BLUE SUREFORM 60 RELOAD. AS A RESULT, THE SIZE OF THE POUCH WAS REDUCED. THE SURGEON OVERSEWED THE DEFECT ON THE REMNANT GASTRIC TISSUE. ALTHOUGH THIS STAPLING ISSUE RESULTED IN A SMALLER POUCH, HE DOES NOT BELIEVE THAT WILL AFFECT THE SURGICAL OUTCOME OR IF THE PATIENT WILL HAVE ANY LONG-TERM EFFECT. THE PATIENT IS REPORTED TO BE RECOVERING WELL. THE SURGEON CONFIRMED THE FOLLOWING: THE RELOAD WAS DISCARDED, HE DID NOT USE BUTTRESS MATERIAL, THE (B)(4) STAPLER INSTRUMENT WAS USED FOR THE REMAINDER OF THE PROCEDURE WITH NO ISSUES, THERE IS NO VIDEO RECORDING; HOWEVER, THERE IS AN IMAGE THAT WAS PROVIDED FOR REVIEW.
29-12-2021 | Injury | intuitive surgical | 21-01-2022 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED BARIATRIC REVISION SURGICAL PROCEDURE (FROM A BAND TO A GASTRIC BYPASS), THE SURGEON REPORTED THAT THE SUREFORM 60 (SF60) STAPLER INSTRUMENT WITH A BLUE SUREFORM 60 RELOAD WAS INSTALLED DURING THE FIRST FIRE. THE STAPLER CLAMPED DOWN ON THE STOMACH WITH NO ISSUES. HOWEVER, WHEN HE PRESSED THE YELLOW PEDAL, THE BLADE JAMMED AND DRAGGED THE TISSUE AND DID NOT CUT TISSUE. THE STAPLES DID NOT FORM CORRECTLY. THERE WAS A THREE CENTIMETER MYOTOMY (HOLE) ON BOTH SIDES OF THE STAPLE LINE (THE GASTRIC POUCH AND THE REMNANT GASTRIC TISSUE). THE SURGEON NOTICED A CLIP FROM THE PREVIOUS LAP BAND SURGERY EMBEDDED ON THE TISSUE HE WAS STAPLING. THE SURGEON THINKS THE CLIP MIGHT HAVE GOTTEN CAUGHT IN FRONT OF THE BLADE AND CAUSED THE STAPLING ISSUE. THE SURGEON REPAIRED THE DEFECT ON THE GASTRIC POUCH, STAPLING MEDIAL TO THE INCOMPLETE STAPLE LINE WITH THE SAME SUREFORM 60 STAPLER INSTRUMENT AND A NEW BLUE SUREFORM 60 RELOAD. AS A RESULT, THE SIZE OF THE POUCH WAS REDUCED. THE SURGEON OVERSEWED THE DEFECT ON THE REMNANT GASTRIC TISSUE. ALTHOUGH THIS STAPLING ISSUE RESULTED IN A SMALLER POUCH, HE DOES NOT BELIEVE THAT WILL AFFECT THE SURGICAL OUTCOME OR IF THE PATIENT WILL HAVE ANY LONG-TERM EFFECT. THE PATIENT IS REPORTED TO BE RECOVERING WELL. THE SURGEON CONFIRMED THE FOLLOWING: THE RELOAD WAS DISCARDED, HE DID NOT USE BUTTRESS MATERIAL, THE (B)(4) STAPLER INSTRUMENT WAS USED FOR THE REMAINDER OF THE PROCEDURE WITH NO ISSUES, THERE IS NO VIDEO RECORDING; HOWEVER, THERE IS AN IMAGE THAT WAS PROVIDED FOR REVIEW.
06-01-2022 | Malfunction | us surgical puerto rico | 21-01-2022 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL GASTRECTOMY, WHILE CUTTING AND CLOSING THE DISTAL STOMACH, THE DEVICE PARTIALLY FIRED. THE SURGEON USED A NEW DEVICE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2021 | Malfunction | covidien lp llc north haven | 21-01-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
20-12-2021 | Malfunction | covidien lp llc north haven | 21-01-2022 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE.
05-01-2022 | Malfunction | covidien medical products | 21-01-2022 | PREMIUM PLUS CEEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING ANASTOMOSIS OF STOMACH TO ESOPHAGUS IN AN OPEN ESOPHAGECTOMY, THE ANVIL CAME OFF AS THE DEVICE WAS BEING WITHDRAWN. THE SURGEON NEED TO MANUALLY "MILK" THE ANVIL THROUGH ANASTOMOSIS AND IS REMOVED WITH DIGITAL MANIPULATION TO THE OUTSIDE OF THE ESOPHAGUS AND GASTROTOMY. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | covidien lp llc north haven | 21-01-2022 | EEA ORVIL | Product Quality Problem | A0201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE GASTRECTOMY WHILE IN ESOPHAGOJEJUNOSTOMY, THE SURGEON COMPLAINT ABOUT THE LABELING. SURGEON USED A STANDARD SIZED DEVICE DUE TO SMALL SIZED PATIENT AND DID NOT KNOW THAT THE STAPLER WAS FOR XL SIZE. THE ANVIL WOULD NOT MATE INTO THE HANDLE AND THE ANVIL LEGS BENT. THE SURGEON OPENED A NEW ONE AND IT ALSO DID NOT WORK. TO RESOLVE THE ISSUE, THE SURGEON SWITCHED STANDARD HANDLE TO XL SIZED AND THERE WAS AN EXTENDED SURGICAL EXTENSION OF 2 HOURS OR MORE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | covidien lp llc north haven | 21-01-2022 | EEA ORVIL | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE GASTRECTOMY WHILE IN ESOPHAGOJEJUNOSTOMY, THE SURGEON COMPLAINT ABOUT THE LABELING. SURGEON USED A STANDARD SIZED DEVICE DUE TO SMALL SIZED PATIENT AND DID NOT KNOW THAT THE STAPLER WAS FOR XL SIZE. THE ANVIL WOULD NOT MATE INTO THE HANDLE AND THE ANVIL LEGS BENT. THE SURGEON OPENED A NEW ONE AND IT ALSO DID NOT WORK. TO RESOLVE THE ISSUE, THE SURGEON SWITCHED STANDARD HANDLE TO XL SIZED AND THERE WAS AN EXTENDED SURGICAL EXTENSION OF 2 HOURS OR MORE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | covidien lp llc north haven | 21-01-2022 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE TACKS WERE NOT FIXING IN CONTACT WITH TISSUE AS USUAL PRESSURE. NOTHING WAS DONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Injury | us surgical puerto rico | 21-01-2022 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC ROBOTIC ASSISTED CONVERTED TO OPEN COLOSTOMY CLOSURE, DURING THE ANASTOMOSIS OF THE PROXIMAL AND DISTAL COLON. THERE WAS STAPLE LINE LEAK. THE STAPLE LINE WAS INCOMPLETE, AND THE SURGEON HAD TO OVER SEW THE STAPLE LINE TO FIX THE ISSUE.
06-01-2022 | Injury | us surgical puerto rico | 21-01-2022 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC ROBOTIC ASSISTED CONVERTED TO OPEN COLOSTOMY CLOSURE, DURING THE ANASTOMOSIS OF THE PROXIMAL AND DISTAL COLON. THERE WAS STAPLE LINE LEAK. THE STAPLE LINE WAS INCOMPLETE, AND THE SURGEON HAD TO OVER SEW THE STAPLE LINE TO FIX THE ISSUE.
06-01-2022 | Injury | us surgical puerto rico | 21-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC ROBOTIC ASSISTED CONVERTED TO OPEN COLOSTOMY CLOSURE, DURING THE ANASTOMOSIS OF THE PROXIMAL AND DISTAL COLON. THERE WAS STAPLE LINE LEAK. THE STAPLE LINE WAS INCOMPLETE, AND THE SURGEON HAD TO OVER SEW THE STAPLE LINE TO FIX THE ISSUE.
06-01-2022 | Injury | us surgical puerto rico | 21-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC ROBOTIC ASSISTED CONVERTED TO OPEN COLOSTOMY CLOSURE, DURING THE ANASTOMOSIS OF THE PROXIMAL AND DISTAL COLON. THERE WAS STAPLE LINE LEAK. THE STAPLE LINE WAS INCOMPLETE, AND THE SURGEON HAD TO OVER SEW THE STAPLE LINE TO FIX THE ISSUE.
28-12-2021 | Malfunction | intuitive surgical | 22-01-2022 | SUREFORM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED DUODENAL SWITCH SURGICAL PROCEDURE, AT THE PROXIMAL END, THE SUREFORM STAPLER PAUSED FOR COMPRESSION, THEN CONTINUED TO ADVANCE AND PAUSED AGAIN. IT WAS NOTICED THAT THE STAPLE LINE WAS OPENED. THE CUSTOMER PULLED OUT THE STAPLER, SWITCHED FROM A BLUE LOAD TO A GREEN LOAD IN THE SAME STAPLER, AND THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) HAS PERFORMED FOLLOW-UPS TO ATTEMPT TO OBTAIN ADDITIONAL INFORMATION. HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
12-08-2021 | Malfunction | us surgical puerto rico | 24-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC SURGERY FOR LUNG CANCER, WHEN CUTTING OFF LUNG TISSUE, THE DEVICE DID NOT FIRE. BOTH THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-08-2021 | Malfunction | covidien lp llc north haven | 24-01-2022 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC SURGERY FOR LUNG CANCER, WHEN CUTTING OFF LUNG TISSUE, THE DEVICE DID NOT FIRE. BOTH THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-09-2021 | Malfunction | us surgical puerto rico | 24-01-2022 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPY PROCEDURE FOR LUNG CANCER, AFTER USING A NUMBER OF RELOADS, WHILE STAPLING THE LUNG TISSUE AGAIN, THE JAWS OF THE DEVICE COULD NOT BE CLAMPED AT ALL. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | us surgical puerto rico | 24-01-2022 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS OPENED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2021 | Malfunction | us surgical puerto rico | 24-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, THE DEVICE DID NOT FIRE EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE COULD NOT BE UNLOAD AS WELL. BOTH HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2021 | Malfunction | us surgical puerto rico | 24-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, THE DEVICE DID NOT FIRE EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE COULD NOT BE UNLOAD AS WELL. BOTH HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2021 | Malfunction | covidien lp llc north haven | 24-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, THE DEVICE DID NOT FIRE EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE COULD NOT BE UNLOAD AS WELL. BOTH HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2021 | Malfunction | covidien lp llc north haven | 24-01-2022 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY, THE DEVICE DID NOT FIRE EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE COULD NOT BE UNLOAD AS WELL. BOTH HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | celestica de monterrey s.a de c.v | 24-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR LUNG CANCER, WHILE DETACHING THE BRONCHI, THE STAPLER WAS ABLE TO FIRE, THE STAPLES DID NOT FORM PROPERLY. THE SURGEON USED A NEW CARTRIDGE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | celestica de monterrey s.a de c.v | 24-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL SURGERY FOR RECTAL CANCER, WHEN CUTTING OFF THE RECTUM, SOME STAPLES WERE NOT FORMED. TO RESOLVE THE ISSUE, THE SURGEON OPENED A NEW RELOAD. THERE WAS NO PATIENT INJURY.
07-01-2022 | Injury | celestica de monterrey s.a de c.v | 24-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RADICAL SURGERY FOR LUNG CANCER, DURING THE DETACHMENT OF LUNG PARENCHYMA, SOME STAPLES WERE NOT FORMED AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE SURGEON HAND-SEWED THE STAPLE LINE TO FIX THE ISSUE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Injury | celestica de monterrey s.a de c.v | 24-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC RADICAL SURGERY FOR LUNG CANCER, DURING THE DETACHMENT OF LUNG PARENCHYMA, SOME STAPLES WERE NOT FORMED AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE SURGEON HAND-SEWED THE STAPLE LINE TO FIX THE ISSUE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | ethicon endo-surgery | 24-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER JAWS DID NOT OPEN. MAY HAVE BEEN PARTIAL FIRE. THE CASE WAS COMPLETED WITH NO PATIENT CONSEQUENCES.
05-01-2022 | Malfunction | ethicon endo-surgery | 24-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER JAWS DID NOT OPEN. MAY HAVE BEEN PARTIAL FIRE. THE CASE WAS COMPLETED WITH NO PATIENT CONSEQUENCES.
07-01-2022 | Malfunction | us surgical puerto rico | 24-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, THE RELOAD LOCKED ON TISSUE AND WAS DIFFICULT TO UNLOAD. WHEN IT WAS ABLE TO BE REMOVED, ONLY HALF OF THE STAPLER WAS ON THE INSTRUMENT TABLE. THE OTHER HALF WAS SUSPECTED TO BE IN THE PATIENT'S CHEST CAVITY SO THE STAFF LOOKED FOR IT BUT DID NOT FIND IT INITIALLY BUT LATER FOUND IT ON THE SAME INSTRUMENT TABLE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | us surgical puerto rico | 24-01-2022 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, THE RELOAD LOCKED ON TISSUE AND WAS DIFFICULT TO UNLOAD. WHEN IT WAS ABLE TO BE REMOVED, ONLY HALF OF THE STAPLER WAS ON THE INSTRUMENT TABLE. THE OTHER HALF WAS SUSPECTED TO BE IN THE PATIENT'S CHEST CAVITY SO THE STAFF LOOKED FOR IT BUT DID NOT FIND IT INITIALLY BUT LATER FOUND IT ON THE SAME INSTRUMENT TABLE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | us surgical puerto rico | 24-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, THE RELOAD LOCKED ON TISSUE AND WAS DIFFICULT TO UNLOAD. WHEN IT WAS ABLE TO BE REMOVED, ONLY HALF OF THE STAPLER WAS ON THE INSTRUMENT TABLE. THE OTHER HALF WAS SUSPECTED TO BE IN THE PATIENT'S CHEST CAVITY SO THE STAFF LOOKED FOR IT BUT DID NOT FIND IT INITIALLY BUT LATER FOUND IT ON THE SAME INSTRUMENT TABLE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | covidien lp llc north haven | 24-01-2022 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, THE RELOAD LOCKED ON TISSUE AND WAS DIFFICULT TO UNLOAD. WHEN IT WAS ABLE TO BE REMOVED, ONLY HALF OF THE STAPLER WAS ON THE INSTRUMENT TABLE. THE OTHER HALF WAS SUSPECTED TO BE IN THE PATIENT'S CHEST CAVITY SO THE STAFF LOOKED FOR IT BUT DID NOT FIND IT INITIALLY BUT LATER FOUND IT ON THE SAME INSTRUMENT TABLE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | covidien lp llc north haven | 24-01-2022 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, THE RELOAD LOCKED ON TISSUE AND WAS DIFFICULT TO UNLOAD. WHEN IT WAS ABLE TO BE REMOVED, ONLY HALF OF THE STAPLER WAS ON THE INSTRUMENT TABLE. THE OTHER HALF WAS SUSPECTED TO BE IN THE PATIENT'S CHEST CAVITY SO THE STAFF LOOKED FOR IT BUT DID NOT FIND IT INITIALLY BUT LATER FOUND IT ON THE SAME INSTRUMENT TABLE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | covidien lp llc north haven | 24-01-2022 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, THE RELOAD LOCKED ON TISSUE AND WAS DIFFICULT TO UNLOAD. WHEN IT WAS ABLE TO BE REMOVED, ONLY HALF OF THE STAPLER WAS ON THE INSTRUMENT TABLE. THE OTHER HALF WAS SUSPECTED TO BE IN THE PATIENT'S CHEST CAVITY SO THE STAFF LOOKED FOR IT BUT DID NOT FIND IT INITIALLY BUT LATER FOUND IT ON THE SAME INSTRUMENT TABLE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | celestica de monterrey s.a de c.v | 24-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPY FOR LUNG CANCER RADICAL SURGERY, WHILE DETACHING THE PULMONARY FISSURE, SOME STAPLES DID NOT FORMED PROPERLY. THE SURGEON USED A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2021 | Injury | ethicon endo-surgery | 24-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER USE OF THE STAPLER, THERE WAS ANASTOMOTIC BLEEDING OF THE PATIENT (POST-OP BLEEDING DURING DAY 1 OR 2). NO MAJOR ISSUES. NO FURTHER INFORMATION AVAILABLE. PROCEDURE: RIGHT COLECTOMY.
06-01-2022 | Malfunction | covidien lp llc north haven | 24-01-2022 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN APPLYING THE RELOAD TO THE STOMACH TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. THE DEVICE ALSO HAD UNLOADING PROBLEM. ANOTHER STAPLER WAS USED BUT THE DEVICE WAS UNABLE TO FIRE AS WELL. THE RELOADS WERE REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | covidien lp llc north haven | 24-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN APPLYING THE RELOAD TO THE STOMACH TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. THE DEVICE ALSO HAD UNLOADING PROBLEM. ANOTHER STAPLER WAS USED BUT THE DEVICE WAS UNABLE TO FIRE AS WELL. THE RELOADS WERE REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | covidien lp llc north haven | 24-01-2022 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN APPLYING THE RELOAD TO THE STOMACH TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. THE DEVICE ALSO HAD UNLOADING PROBLEM. ANOTHER STAPLER WAS USED BUT THE DEVICE WAS UNABLE TO FIRE AS WELL. THE RELOADS WERE REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | us surgical puerto rico | 24-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN APPLYING THE RELOAD TO THE STOMACH TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. THE DEVICE ALSO HAD UNLOADING PROBLEM. ANOTHER STAPLER WAS USED BUT DEVICE WAS UNABLE TO FIRE AS WELL. THE RELOADS WERE REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | us surgical puerto rico | 24-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN APPLYING THE RELOAD TO THE STOMACH TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. THE DEVICE ALSO HAD UNLOADING PROBLEM. ANOTHER STAPLER WAS USED BUT DEVICE WAS UNABLE TO FIRE AS WELL. THE RELOADS WERE REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | us surgical puerto rico | 24-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN APPLYING THE RELOAD TO THE STOMACH TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. THE DEVICE ALSO HAD UNLOADING PROBLEM. ANOTHER STAPLER WAS USED BUT DEVICE WAS UNABLE TO FIRE AS WELL. THE RELOADS WERE REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | covidien lp llc north haven | 24-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, WHEN APPLYING THE RELOAD TO THE STOMACH TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE BUT WAS ABLE TO BE REMOVED. THE DEVICE ALSO HAD UNLOADING PROBLEM. ANOTHER STAPLER WAS USED BUT DEVICE WAS UNABLE TO FIRE AS WELL. THE RELOADS WERE REPLACED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-11-2021 | Injury | ethicon endo-surgery | 24-01-2022 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT UNDER LAPAROSCOPE, AN INCISION ABOUT 10 CM IN LENGTH WAS MADE IN THE MIDDLE ABDOMEN, THE PROXIMAL INTESTINAL CANAL WAS PULLED OUT OF THE ABDOMINAL CAVITY, THE PURSE-STRING FORCEPS WAS USED TO CUT OFF AT THE PREDETERMINED TANGENT LINE OF SIGMOID COLON, THE MUSHROOM HEAD OF THE DEVICE WAS PLACED AFTER CUTTING OFF, THE PERINEAL ANUS WAS DISINFECTED, THE DEVICE ENTERED THE BODY FROM THE ANUS, THE INTESTINAL SEGMENT WAS CLOSED BY ROTATING THE DEVICE SWITCH AFTER CONNECTING WITH THE INTRAPERITONEAL MUSHROOM HEAD, THE DEVICE WAS ROTATED TO THE GREEN AREA, AFTER FIRING, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. IT WAS CONVERTED TO LAPAROTOMY AND CHANGED ANOTHER DEVICE TO COMPLETE THE SURGERY. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-12-2021 | Malfunction | ethicon endo-surgery | 24-01-2022 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, THE CIRCULAR STAPLER DID NOT FIRE PROPERLY. THE SURGERY WAS DELAYED BY 60-MINUTES. ANOTHER CDH STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. NO FRAGMENTS WERE GENERATED. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
03-12-2021 | Malfunction | ethicon endo-surgery | 24-01-2022 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, THE CIRCULAR STAPLER DID NOT FIRE PROPERLY. THE SURGERY WAS DELAYED BY 60-MINUTES. ANOTHER CDH STAPLER WAS USED TO COMPLETE THE ANASTOMOSIS. NO FRAGMENTS WERE GENERATED. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
08-01-2022 | Injury | covidien lp llc north haven | 24-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (RYGB), WHILE CREATING A POUCH, IN THREE DEVICES, THERE WAS AN EXCESSIVE BLEEDING. THE SURGEON SUTURED THE STAPLE LINE AND EXTENDED THE SURGERY TIME FOR OVER 30 MINUTES TO RESOLVE THE ISSUE.
08-01-2022 | Injury | covidien lp llc north haven | 24-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (RYGB), WHILE CREATING A POUCH, IN THREE DEVICES, THERE WAS AN EXCESSIVE BLEEDING. THE SURGEON SUTURED THE STAPLE LINE AND EXTENDED THE SURGERY TIME FOR OVER 30 MINUTES TO RESOLVE THE ISSUE.
15-12-2021 | Malfunction | ethicon endo-surgery | 24-01-2022 | PROXIMATE* LINEAR CUTTER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTECTOMY, AT THE FIRST USE THE BLADE CUTS BUT THE POINTS IN THE PROXIMAL PART DO NOT FORM COMPLETELY. THERE WAS NO DELAY TO THE SURGERY OR FRAGMENTS GENERATED. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
03-01-2022 | Malfunction | ethicon endo-surgery | 24-01-2022 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. PROCEDURE: RIGHT HEMICOLECTOMY
03-01-2022 | Malfunction | ethicon endo-surgery | 24-01-2022 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. PROCEDURE: RIGHT HEMICOLECTOMY.
07-01-2022 | Malfunction | ethicon endo-surgery | 24-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING AN LAR VERY THICK RADIATED COLON TISSUE. CREATED TWO STAPLER MALFUNCTIONS. THE CONTOUR HANDLE BROKE AND THE GST 60 WOULDN¿T OPEN UP TO RELOAD. REPLACED WITH SAME INSTRUMENTS. SURGEON SAID THE DEVICES DID NOT NEED TO BE SENT IN FOR ANALYSIS IT WAS DUE TO TISSUE BEING SO THICK. THERE WERE NO PATIENT CONSEQUENCES.
07-01-2022 | Malfunction | ethicon endo-surgery | 24-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THAT DURING AN LAR VERY THICK RADIATED COLON TISSUE. CREATED TWO STAPLER MALFUNCTIONS. THE CONTOUR HANDLE BROKE AND THE GST 60 WOULDN¿T OPEN UP TO RELOAD. REPLACED WITH SAME INSTRUMENTS. SURGEON SAID THE DEVICES DID NOT NEED TO BE SENT IN FOR ANALYSIS IT WAS DUE TO TISSUE BEING SO THICK. THERE WERE NO PATIENT CONSEQUENCES.
29-12-2021 | Malfunction | ethicon endo-surgery | 24-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID PROCEDURE THAT THE DEVICE LIT UP AS IF IT WAS READY TO FIRE BUT THE DEVICE WOULD NOT FIRE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
13-12-2021 | Injury | ethicon endo-surgery | 24-01-2022 | 75MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE NTLC75 WAS USED IN A WHIPPLE¿S OPERATION. THE NTLC75 WAS USED TO CREATE THE GASTROJEJUNOSTOMY ANASTOMOSIS AS PART OF THE WHIPPLE¿S OPERATION, THIS WAS PERFORMED USING THE BLUE SELECTABLE STAPLE HEIGHT OF THE NTLC75. AFTER THE OPERATION WAS COMPLETED, PATIENT WAS FOUND TO HAVE POST-OP BLEEDING ORIGINATING FROM THE STAPLE LINE OF THE GASTROJEJUNOSTOMY. THE PATIENT HAD TO BE SENT BACK TO THE OPERATION THEATER AND BE RE-OPERATED ON TO ADDRESS THE BLEEDING, LESS THAN 12 HOURS POST-OP. PATIENT WAS DESCRIBED TO HAVE SEVERE BLOOD LOSS AMOUNTING TO APPROXIMATELY 2 LITERS OF BLOOD. SURGICAL INTERVENTION WAS DONE TO ACHIEVE HEMOSTASIS AND HEMOSTASIS WAS ACHIEVED. AFTER A PERIOD OF RECOVERY, PATIENT WAS SENT BACK HOME. PATIENT WAS DESCRIBED TO BE MOSTLY HEALTHY AND WITH NO COMORBIDITIES. AFTER A PERIOD OF ABOUT 1.5 WEEKS, PATIENT HAD RETURNED AND WAS PRESENTED WITH ANASTOMOTIC LEAK. THE LOCATION OF THE LEAK WAS A SUTURE SITE (CLOSURE OF GASTROTOMY AND ENTEROTOMY) THAT HAD TO BE OPENED TO ADDRESS THE BLEEDING IN THE EARLIER REOPERATION. SURGEON MENTIONED THAT THE ANASTOMOTIC LEAK OCCURRED BECAUSE THE SUTURE SITE HAD BECOME WEAKENED DUE TO THE EARLIER REOPERATION. HENCE, PATIENT WAS RE-OPERATED ON TO ADDRESS THE ANASTOMOTIC LEAK. PATIENT IS CURRENTLY RECOVERING IN WARD.
13-12-2021 | Injury | ethicon endo-surgery | 24-01-2022 | 75MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT THE NTLC75 WAS USED IN A WHIPPLE¿S OPERATION. THE NTLC75 WAS USED TO CREATE THE GASTROJEJUNOSTOMY ANASTOMOSIS AS PART OF THE WHIPPLE¿S OPERATION, THIS WAS PERFORMED USING THE BLUE SELECTABLE STAPLE HEIGHT OF THE NTLC75. AFTER THE OPERATION WAS COMPLETED, PATIENT WAS FOUND TO HAVE POST-OP BLEEDING ORIGINATING FROM THE STAPLE LINE OF THE GASTROJEJUNOSTOMY. THE PATIENT HAD TO BE SENT BACK TO THE OPERATION THEATER AND BE RE-OPERATED ON TO ADDRESS THE BLEEDING, LESS THAN 12 HOURS POST-OP. PATIENT WAS DESCRIBED TO HAVE SEVERE BLOOD LOSS AMOUNTING TO APPROXIMATELY 2 LITERS OF BLOOD. SURGICAL INTERVENTION WAS DONE TO ACHIEVE HEMOSTASIS AND HEMOSTASIS WAS ACHIEVED. AFTER A PERIOD OF RECOVERY, PATIENT WAS SENT BACK HOME. PATIENT WAS DESCRIBED TO BE MOSTLY HEALTHY AND WITH NO COMORBIDITIES. AFTER A PERIOD OF ABOUT 1.5 WEEKS, PATIENT HAD RETURNED AND WAS PRESENTED WITH ANASTOMOTIC LEAK. THE LOCATION OF THE LEAK WAS A SUTURE SITE (CLOSURE OF GASTROTOMY AND ENTEROTOMY) THAT HAD TO BE OPENED TO ADDRESS THE BLEEDING IN THE EARLIER REOPERATION. SURGEON MENTIONED THAT THE ANASTOMOTIC LEAK OCCURRED BECAUSE THE SUTURE SITE HAD BECOME WEAKENED DUE TO THE EARLIER REOPERATION. HENCE, PATIENT WAS RE-OPERATED ON TO ADDRESS THE ANASTOMOTIC LEAK. PATIENT IS CURRENTLY RECOVERING IN WARD.
24-01-2022 | Malfunction | ethicon | 24-01-2022 | ETHICON SECURESTRAP | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND THE ABSORBABLE STRAP WAS USED. DURING SURGERY, THE STRAP COULD NOT BE DEPLOYED. IT DID NOT STICK TO THE TARGET TISSUE TIGHTLY AS WELL. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION WAS REQUESTED.
07-06-2021 | Injury | covidien lp llc north haven | 24-01-2022 | UNKNOWN ABSORBATACK | Adverse Event Without Identified Device or Use Problem | A24 | Prolapse | Event Description: ACCORDING TO LITERATURE SOURCE OF STUDY PERFORMED BETWEEN MARCH 2016 AND JANUARY 2020, OUT OF THE 37 STUDY PATIENTS WHO UNDERWENT LAPAROSCOPIC LATERAL SUSPENSION PROCEDURE FOR ADVANCED STAGE PELVIC ORGAN PROLAPSE, POSTOPERATIVE RECURRENCE (ANTERIOR COMPARTMENT) WAS REPORTED ON 2 PATIENTS SO REOPERATION WAS PERFORMED.
07-01-2022 | Malfunction | covidien lp llc north haven | 24-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DISTAL PANCREATIC RESECTION, AFTER CLAMPING THE PANCREAS WITH A STAPLER AND TRIED TO FIRE AFTER A WHILE, BUT THE KNIFE COULD NOT BE ADVANCED AND THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO ATTEMPT FIRING AGAIN, THERE WAS NO PROBLEM RESECTING. AFTER THE OPERATION, THE POWERED HANDLE WAS CHECKED, BUT THERE WAS NO PROBLEM WITH THE OPERATION. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE COULD NOT BE FIRED AT THE 2ND FIRING. IT WAS FOUND THAT THE ANVIL JAW WAS DAMAGED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
06-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE COULD NOT BE FIRED AT THE 2ND FIRING. IT WAS FOUND THAT THE ANVIL JAW WAS DAMAGED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ROBOTIC LAPAROSCOPIC TOTAL GASTRECTOMY, WHILE ANASTOMOSING THE ESOPHAGUS WITH THE JEJUNUM, THE DEVICE PARTIALLY FIRED. THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE.
07-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ROBOTIC LAPAROSCOPIC TOTAL GASTRECTOMY, WHILE ANASTOMOSING THE ESOPHAGUS WITH THE JEJUNUM, THE DEVICE PARTIALLY FIRED. THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE.
07-01-2022 | Malfunction | covidien | 25-01-2022 | ENDO GIA ULTRA UNIVERSAL STAPLER 12 MM | Break | A0401 | Insufficient Information | Event Description: SURGICAL STAFF UTILIZING STAPLER INTRAOPERATIVELY AND IT BROKE DURING FIRING THE LOAD. THE HANDLE TO RETRIEVE THE BLADE AND OPEN THE STAPLER BROKE WHILE OPENING THE STAPLER NO NOTED PATIENT HARM.
04-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2022 | PWRD 23MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE GASTRIC PROCEDURE SURGERY, COULD NOT FIRE ON THE FIRST FIRING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ANTERIOR RESECTION, THE DEVICE COULD NOT BE FIRED. THE INDICATOR WAS WITHIN THE GREEN RANGE. THE BATTERY WAS REINSERTED, BUT THE ISSUE DID NOT IMPROVE. AN ILS DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS NOT CONVERTED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THE PATIENT¿S CONDITION IS STABLE. NO FURTHER INFORMATION IS AVAILABLE.
03-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. PROCEDURE: LOW ANTERIOR RESECTION.
10-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2022 | FLEX 60 ARTICULATING | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE THAT THE STAPLER GOT STUCK ON THE BOWEL AND COULDN'T BE RELEASED. DEVICE WAS STUCK IN ARTICULATING POSITION, AND NEEDED TO BE REMOVED WITH TROCAR. UNKNOWN AS TO HOW THE DEVICE WAS REMOVED OR HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
10-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2022 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOID COLECTOMY PROCEDURE THAT THE STAPLER GOT STUCK ON THE BOWEL AND COULDN'T BE RELEASED. DEVICE WAS STUCK IN ARTICULATING POSITION, AND NEEDED TO BE REMOVED WITH TROCAR. UNKNOWN AS TO HOW THE DEVICE WAS REMOVED OR HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
11-11-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | ECHELON FLEX 60 COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN COLECTOMY, THE KNIFE DID NOT MOVE FORWARD AT THE FIRING OF FEEA. THE SURGEON COMMENTED THAT BECAUSE IT WAS AN EMERGENCY SURGERY, HE WAS IN A HURRY AND THERE IS A POSSIBILITY THAT THE DEVICE WAS USED IN A WRONG WAY. THE DEVICE WAS USED ON THE COLON. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-11-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC COLORECTAL SURGERY, COULD NOT FIRE WHEN SEVERING RECTUM TISSUE ON THE 2TH FIRING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-11-2021 | Injury | us surgical puerto rico | 25-01-2022 | EEA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A COLON PROCEDURE, AFTER CLOSING THE CLAMP AND FIRING THE DEVICE, THERE WAS POOR STAPLE FORMATION, INCOMPLETE STAPLE LINE AND DONUTS. THE SURGERY WAS EXTENDED BY ONE HOUR TO SUTURE THE ANASTOMOSIS WHERE IT WAS NOTED TO BE INCOMPLETE AND HAD NO STAPLES TO RESOLVE THE ISSUE.
22-11-2021 | Injury | us surgical puerto rico | 25-01-2022 | EEA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A COLON PROCEDURE, AFTER CLOSING THE CLAMP AND FIRING THE DEVICE, THERE WAS POOR STAPLE FORMATION, INCOMPLETE STAPLE LINE AND DONUTS. THE SURGERY WAS EXTENDED BY ONE HOUR TO SUTURE THE ANASTOMOSIS WHERE IT WAS NOTED TO BE INCOMPLETE AND HAD NO STAPLES TO RESOLVE THE ISSUE.
22-11-2021 | Injury | us surgical puerto rico | 25-01-2022 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A COLON PROCEDURE, AFTER CLOSING THE CLAMP AND FIRING THE DEVICE, THERE WAS POOR STAPLE FORMATION, INCOMPLETE STAPLE LINE AND DONUTS. THE SURGERY WAS EXTENDED BY ONE HOUR TO SUTURE THE ANASTOMOSIS WHERE IT WAS NOTED TO BE INCOMPLETE AND HAD NO STAPLES TO RESOLVE THE ISSUE.
22-11-2021 | Injury | us surgical puerto rico | 25-01-2022 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A COLON PROCEDURE, AFTER CLOSING THE CLAMP AND FIRING THE DEVICE, THERE WAS POOR STAPLE FORMATION, INCOMPLETE STAPLE LINE AND DONUTS. THE SURGERY WAS EXTENDED BY ONE HOUR TO SUTURE THE ANASTOMOSIS WHERE IT WAS NOTED TO BE INCOMPLETE AND HAD NO STAPLES TO RESOLVE THE ISSUE.
27-12-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE CARTRIDGE LOADED ON A PCEE60A FELL OFF IN THE PATIENT WHEN THE DEVICE WAS INSERTED INTO THE PATIENT AT THE 1ST FIRING. THE CARTRIDGE WAS RETRIEVED BY A FORCEPS AND NOT USED FOR THE PATIENT. THE 2ND AND 3RD FIRINGS WERE COMPLETED WITH NO PROBLEM. AFTER THAT, A 2 TO 3 MM WHITE PART WAS FOUND INSIDE THE PATIENT. IT WAS RETRIEVED. THE DEVICE WAS USED ON THE LUNG. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
27-12-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | GST GOLD RELOAD, 60MM, 6 ROW | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE CARTRIDGE LOADED ON A PCEE60A FELL OFF IN THE PATIENT WHEN THE DEVICE WAS INSERTED INTO THE PATIENT AT THE 1ST FIRING. THE CARTRIDGE WAS RETRIEVED BY A FORCEPS AND NOT USED FOR THE PATIENT. THE 2ND AND 3RD FIRINGS WERE COMPLETED WITH NO PROBLEM. AFTER THAT, A 2 TO 3 MM WHITE PART WAS FOUND INSIDE THE PATIENT. IT WAS RETRIEVED. THE DEVICE WAS USED ON THE LUNG. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
29-12-2021 | Injury | ethicon endo-surgery | 25-01-2022 | ECHELON FLEX 60 COMPACT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING THE LOBECTOMY SURGERY, THE SURGEON DID NOT OBSERVE ANY ABNORMAL SITUATION IN THE OPERATION AND STAPLE LINE WAS GOOD. BUT THE PATIENT HAD ABDOMINAL PAIN AND THORACIC FLUID INCREASED IN THE EVENING, THEN THE SURGEON DID 2ND SURGERY AND NOTED THAT VESSEL OF ANASTOMOTIC SITE WAS BLEEDING. THE SURGEON USED SUTURE TO OVER SEW. THE PATIENT'S CONDITION IS CURRENTLY STABLE.
29-12-2021 | Injury | ethicon endo-surgery | 25-01-2022 | ECHELON FLEX 60 COMPACT | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: IT WAS REPORTED THAT DURING THE LOBECTOMY SURGERY, THE SURGEON DID NOT OBSERVE ANY ABNORMAL SITUATION IN THE OPERATION AND STAPLE LINE WAS GOOD. BUT THE PATIENT HAD ABDOMINAL PAIN AND THORACIC FLUID INCREASED IN THE EVENING, THEN THE SURGEON DID 2ND SURGERY AND NOTED THAT VESSEL OF ANASTOMOTIC SITE WAS BLEEDING. THE SURGEON USED SUTURE TO OVER SEW. THE PATIENT'S CONDITION IS CURRENTLY STABLE.
29-12-2021 | Injury | ethicon endo-surgery | 25-01-2022 | ECHELON FLEX 60 COMPACT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE LOBECTOMY SURGERY, THE SURGEON DID NOT OBSERVE ANY ABNORMAL SITUATION IN THE OPERATION AND STAPLE LINE WAS GOOD. BUT THE PATIENT HAD ABDOMINAL PAIN AND THORACIC FLUID INCREASED IN THE EVENING, THEN THE SURGEON DID 2ND SURGERY AND NOTED THAT VESSEL OF ANASTOMOTIC SITE WAS BLEEDING. THE SURGEON USED SUTURE TO OVER SEW. THE PATIENT'S CONDITION IS CURRENTLY STABLE.
18-11-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | ECHELON FLEX 60 COMPACT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC RADICAL RESECTION OF LUNG CANCER SURGERY. AFTER FIRED, THE SURGEON NOTED THAT THE TISSUE WAS NOT CUT AND STAPLES DEPLOYED ON THE TISSUE, CHANGED ANOTHER DEVICE TO COMPLETE THE PROCEDURE, THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-11-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DR USED 3 MANUAL GUN CONTINUOUSLY DURING SURGERY, BUT THE RELOADS GOT STUCK WHEN FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
06-11-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DR USED 3 MANUAL GUN CONTINUOUSLY DURING SURGERY, BUT THE RELOADS GOT STUCK WHEN FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
30-12-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE THAT THE DEVICE WAS INTRODUCED TO THE PATIENT AND CONNECTED TO THE ANVIL, AFTER DIALING DOWN THE ANVIL THE DEVICE WOULD NOT FIRE. THE DEVICE WAS LIT UP LIKE IT WAS READY TO FIRE BUT NOTHING HAPPENED WHEN ATTEMPTED. SURGEON REMOVED THE BATTERY AND REINSERTED WITH NO CHANGE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. A LEAK TEST WAS PREFORMED, AND NO LEAKS WERE FOUND HOWEVER THE DISTAL DONUT WAS NOT INTACT. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
30-12-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE THAT THE DEVICE WAS INTRODUCED TO THE PATIENT AND CONNECTED TO THE ANVIL, AFTER DIALING DOWN THE ANVIL THE DEVICE WOULD NOT FIRE. THE DEVICE WAS LIT UP LIKE IT WAS READY TO FIRE BUT NOTHING HAPPENED WHEN ATTEMPTED. SURGEON REMOVED THE BATTERY AND REINSERTED WITH NO CHANGE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. A LEAK TEST WAS PREFORMED, AND NO LEAKS WERE FOUND HOWEVER THE DISTAL DONUT WAS NOT INTACT. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
17-11-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SLEEVE GASTRECTOMY SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND . THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL COULD NOT FIRE AND WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
03-05-2022 | Malfunction | covidien lp llc north haven | 25-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY, THE DEVICE WAS NOT ABLE TO FIRE. DEVICE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
11-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2022 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-11-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE KNIFE DID NOT MOVE IN FRONT WHILE FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
06-11-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT REPORTED THAT THE DR USED 3 MANUAL GUN CONTINUOUSLY DURING SURGERY, BUT THE RELOADS GOT STUCK WHEN FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
20-04-2022 | Malfunction | us surgical puerto rico | 25-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY, THE DEVICE WAS NOT ABLE TO FIRE. DEVICE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
09-01-2022 | Malfunction | us surgical puerto rico | 25-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE CARTRIDGE WAS NOT FIRED. THERE WAS NO PATIENT INJURY.
24-11-2021 | Malfunction | us surgical puerto rico | 25-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE RADICAL GASTRECTOMY FOR GASTRIC CANCER WHILE CUTTING OFF THE STOMACH TISSUE, THE RELOAD WAS INSTALLED DURING THE OPERATION AND COULD NOT BE FIRED. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
14-02-2022 | Malfunction | covidien lp llc north haven | 25-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY, THE DEVICE WAS NOT ABLE TO FIRE. DEVICE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-02-2022 | Malfunction | covidien lp llc north haven | 25-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY, THE DEVICE WAS NOT ABLE TO FIRE. DEVICE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
29-11-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE VIDEO-ASSISTED THORACOSCOPIC SEGMENTECTOMY SURGERY, COULD NOT FIRE WHEN SEVERING BRONCHIAL TISSUE ON THE 3TH FIRING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | PREMIUM PLUS CEEA | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGECTOMY PROCEDURE, WHILE THE ESOPHAGUS WAS BEING ANASTOMOSED, THE STAPLES DID NOT FORM PROPERLY AND THERE WAS AN ANASTOMOTIC BLEEDING. THE ANVIL DOES NOT TILT AFTER FIRING. SUTURE REINFORCEMENT WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
07-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | PREMIUM PLUS CEEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGECTOMY PROCEDURE, WHILE THE ESOPHAGUS WAS BEING ANASTOMOSED, THE STAPLES DID NOT FORM PROPERLY AND THERE WAS AN ANASTOMOTIC BLEEDING. THE ANVIL DOES NOT TILT AFTER FIRING. SUTURE REINFORCEMENT WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
07-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | PREMIUM PLUS CEEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGECTOMY PROCEDURE, WHILE THE ESOPHAGUS WAS BEING ANASTOMOSED, THE STAPLES DID NOT FORM PROPERLY AND THERE WAS AN ANASTOMOTIC BLEEDING. THE ANVIL DOES NOT TILT AFTER FIRING. SUTURE REINFORCEMENT WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
07-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | PREMIUM PLUS CEEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGECTOMY PROCEDURE, WHILE THE ESOPHAGUS WAS BEING ANASTOMOSED, THE STAPLES DID NOT FORM PROPERLY AND THERE WAS AN ANASTOMOTIC BLEEDING. THE ANVIL DOES NOT TILT AFTER FIRING. SUTURE REINFORCEMENT WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
07-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | PREMIUM PLUS CEEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGECTOMY PROCEDURE, WHILE THE ESOPHAGUS WAS BEING ANASTOMOSED, THE STAPLES DID NOT FORM PROPERLY AND THERE WAS AN ANASTOMOTIC BLEEDING. THE ANVIL DOES NOT TILT AFTER FIRING. SUTURE REINFORCEMENT WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
07-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | PREMIUM PLUS CEEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL ESOPHAGECTOMY PROCEDURE, WHILE THE ESOPHAGUS WAS BEING ANASTOMOSED, THE STAPLES DID NOT FORM PROPERLY AND THERE WAS AN ANASTOMOTIC BLEEDING. THE ANVIL DOES NOT TILT AFTER FIRING. SUTURE REINFORCEMENT WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
09-01-2022 | Malfunction | us surgical puerto rico | 25-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE CARTRIDGE WAS NOT FIRED. THERE WAS NO PATIENT INJURY.
09-01-2022 | Malfunction | us surgical puerto rico | 25-01-2022 | GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE CARTRIDGE WAS NOT FIRED. THERE WAS NO PATIENT INJURY.
10-12-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE FARTHER AFTER FIRED 1/3 WHEN SEVERING RECTAL TISSUE ON THE 1TH FIRING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-11-2021 | Malfunction | us surgical puerto rico | 25-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE PROCEDURE, THE HANDLE AND TWO RELOADS HAD AN UNKNOWN ISSUE. TO RESOLVE THE ISSUE, A NEW RELOAD AND HANDLE WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT REVEALED SHEARED TEETH ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
01-11-2021 | Malfunction | us surgical puerto rico | 25-01-2022 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE PROCEDURE, THE HANDLE AND TWO RELOADS HAD AN UNKNOWN ISSUE. TO RESOLVE THE ISSUE, A NEW RELOAD AND HANDLE WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT REVEALED SHEARED TEETH ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
07-01-2022 | Malfunction | us surgical puerto rico | 25-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT LUNG SEGMENTECTOMY, WHILE CUTTING AND CLOSING THE LUNG SEGMENT, THE DEVICE DID NOT FIRE. THE SURGEON USED A NEW HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-11-2021 | Malfunction | covidien lp llc north haven | 25-01-2022 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE PROCEDURE, THE HANDLE AND TWO RELOADS HAD AN UNKNOWN ISSUE. TO RESOLVE THE ISSUE, A NEW RELOAD AND HANDLE WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT THE RELOAD WAS PARTIALLY FIRED TO THE 3CM CUT LINE.
01-11-2021 | Malfunction | covidien lp llc north haven | 25-01-2022 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE PROCEDURE, THE HANDLE AND TWO RELOADS HAD AN UNKNOWN ISSUE. TO RESOLVE THE ISSUE, A NEW RELOAD AND HANDLE WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT THE RELOAD WAS PARTIALLY FIRED TO THE 3CM CUT LINE.
10-01-2022 | Malfunction | covidien lp llc north haven | 25-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC REVISION BYPASS, WHEN STAPLING THE STOMACH, THE KNIFE OF THE DEVICE WAS UNABLE TO CUT THE STOMACH AND NO STAPLE LINE WAS FORMED. TO RESOLVE THE ISSUE, THE SURGEON OPENED A MANUAL STAPLER. THERE WAS NO PATIENT INJURY.
08-01-2022 | Injury | celestica de monterrey s.a de c.v | 25-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC CANCER RADICAL SURGERY, ON DETACHMENT OF STOMACH TISSUE, THE STAPLER WAS ABLE TO FIRE, SOME STAPLES WERE NOT FORMED AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE SURGEON REPAIRED WITH SUTURES TO FIX THE STAPLE LINE AND CHANGED THE RELOAD TO RESOLVE THE ISSUE.
08-01-2022 | Injury | celestica de monterrey s.a de c.v | 25-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC CANCER RADICAL SURGERY, ON DETACHMENT OF STOMACH TISSUE, THE STAPLER WAS ABLE TO FIRE, SOME STAPLES WERE NOT FORMED AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE SURGEON REPAIRED WITH SUTURES TO FIX THE STAPLE LINE AND CHANGED THE RELOAD TO RESOLVE THE ISSUE.
13-12-2021 | Malfunction | covidien lp llc north haven | 25-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC ROUX-EN-Y PROCEDURE, WHILE TRANSECTING THE STOMACH, THE RELOAD PARTIALLY FIRED. A NEW RELOAD WAS OPENED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 25-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC REVISION BYPASS, WHEN STAPLING THE STOMACH, THE KNIFE OF THE DEVICE WAS UNABLE TO CUT THE STOMACH AND NO STAPLE LINE WAS FORMED. TO RESOLVE THE ISSUE, THE SURGEON OPENED A MANUAL STAPLER. THERE WAS NO PATIENT INJURY.
26-10-2021 | Injury | us surgical puerto rico | 25-01-2022 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RADICAL RESECTION OF LUNG CANCER, WHILE DETACHING THE TRACHEA, SOME OF THE STAPLES DID NOT FORM CORRECTLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE STAPLE LINE WAS SUTURED MANUALLY AND ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
10-01-2022 | Malfunction | covidien lp llc north haven | 25-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASCENDING AORTA AND TOTAL ARCH REPLACEMENT PROCEDURE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE PRODUCT WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 25-01-2022 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC REVISION BYPASS, WHEN STAPLING THE STOMACH, THE KNIFE OF THE DEVICE WAS UNABLE TO CUT THE STOMACH AND NO STAPLE LINE WAS FORMED. TO RESOLVE THE ISSUE, THE SURGEON OPENED A MANUAL STAPLER. THERE WAS NO PATIENT INJURY.
10-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | PREMIUM PLUS CEEA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND HAD POOR STAPLE FORMATION. PART OF JEJUNUM WAS RESECTED AND NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
10-01-2022 | Injury | us surgical puerto rico | 25-01-2022 | PREMIUM PLUS CEEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, THE DEVICE PARTIALLY FIRED AND HAD POOR STAPLE FORMATION. PART OF JEJUNUM WAS RESECTED AND NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
01-01-2021 | Injury | ethicon endo-surgery | 25-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER USE OF THE STAPLER IN A RIGHT COLECTOMY PROCEDURE, ANASTOMOTIC BLEEDING OF THE PATIENT BETWEEN DAY 1 OR DAY 2 POST-OPERATIVELY. NO MAJOR ISSUES. NO FURTHER INFORMATION AVAILABLE.
19-12-2021 | Malfunction | ethicon endo-surgery | 25-01-2022 | PWRD ECH FLEX 45MM 34CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY,AFTER FIRED, COULD NOT OPEN THE JAW. AS THE TISSUE WAS CUT OFF, REMOVED THE DEVICE FROM THE PATIENT THROUGH TISSUE GAP. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2021 | Injury | ethicon endo-surgery | 25-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER USE OF THE STAPLER IN A RIGHT COLECTOMY PROCEDURE, ANASTOMOTIC BLEEDING OF THE PATIENT BETWEEN DAY 1 OR DAY 2 POST-OPERATIVELY. NO MAJOR ISSUES. NO FURTHER INFORMATION AVAILABLE.
01-01-2021 | Injury | ethicon endo-surgery | 25-01-2022 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER USE OF THE STAPLER IN A RIGHT COLECTOMY PROCEDURE, ANASTOMOTIC BLEEDING OF THE PATIENT BETWEEN DAY 1 OR DAY 2 POST-OPERATIVELY. NO MAJOR ISSUES. NO FURTHER INFORMATION AVAILABLE.
05-11-2021 | Malfunction | us surgical puerto rico | 25-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HEMICOLECTOMY, THE DEVICE WAS UNABLE TO FIRE AND THE GREEN BUTTON WAS PRESSED BUT HANDLE COULD NOT BE SQUEEZED. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. PLI20: 2 DEVICES WERE RETURNED WITHOUT ANY ACCOMPANYING INFORMATION. PLI20: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH ON THE FIRING RACK.
10-01-2022 | Malfunction | us surgical puerto rico | 25-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASCENDING AORTA AND TOTAL ARCH REPLACEMENT PROCEDURE, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE PRODUCT WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | us surgical puerto rico | 25-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASCENDING AORTA AND TOTAL ARCH REPLACEMENT PROCEDURE, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE PRODUCT WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | us surgical puerto rico | 25-01-2022 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN RADICAL GASTRECTOMY FOR GASTRIC CANCER, DURING THE TRANSECTION OF THE STOMACH, THE DEVICE DID NOT FIRE. THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-01-2022 | Malfunction | intuitive surgical | 25-01-2022 | SURE FORM 60 ROBOTIC STAPLER | Display or Visual Feedback Problem | A0902 | Insufficient Information | Event Description: STAPLER INDICATED "STAPLER INSTALLATION FAILED" MESSAGE FOLLOWING 4TH FIRING. THE STAPLER WAS ALSO NOTED TO BE CLOSING ONLY APPROXIMATELY 1/2 WAY AT THE TIME OF THE MESSAGE. FDA SAFETY REPORT ID# (B)(4).
20-01-2022 | Malfunction | intuitive surgical | 25-01-2022 | SURE FORM 60 ROBOTIC STAPLER | Failure to Form Staple | A050704 | Insufficient Information | Event Description: STAPLER INDICATED "STAPLER INSTALLATION FAILED" MESSAGE FOLLOWING 4TH FIRING. THE STAPLER WAS ALSO NOTED TO BE CLOSING ONLY APPROXIMATELY 1/2 WAY AT THE TIME OF THE MESSAGE. FDA SAFETY REPORT ID# (B)(4).
15-12-2021 | Injury | medtronic / covidien | 25-01-2022 | COVIDIEN STAPLER | Use of Device Problem | A23 | Insufficient Information | Event Description: COVIDIEN STAPLER BY MEDTRONIC, REF# (B)(4), LOT# P1A0337 MALFUNCTION OF STAPLER NOT FIRING STAPLES INTO TISSUE. FDA SAFETY REPORT ID# (B)(4).
15-12-2021 | Injury | medtronic / covidien | 25-01-2022 | COVIDIEN STAPLER | Failure to Fire | A050501 | Insufficient Information | Event Description: COVIDIEN STAPLER BY MEDTRONIC, REF# (B)(4), LOT# P1A0337 MALFUNCTION OF STAPLER NOT FIRING STAPLES INTO TISSUE. FDA SAFETY REPORT ID# (B)(4).
10-10-2021 | Malfunction | us surgical puerto rico | 26-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING LAPAROSCOPIC SEGMENTECTOMY WHILE BEING APPLIED ON THE LUNGS, THE STAPLER FAILED (BROKE) ON FIRST USE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-10-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | UNKNOWN ENDO GIA SULU | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING LAPAROSCOPIC SEGMENTECTOMY WHILE BEING APPLIED ON THE LUNGS, THE STAPLER FAILED (BROKE) ON FIRST USE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-10-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING LAPAROSCOPIC SEGMENTECTOMY WHILE BEING APPLIED ON THE LUNGS, THE STAPLER FAILED (BROKE) ON FIRST USE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-04-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, WHILE STAPLING THE STOMACH, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-01-2022 | Malfunction | us surgical puerto rico | 26-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, WHILE STAPLING THE STOMACH, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-10-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, ON A SLEEVE GASTRECTOMY, THREE HANDLES WOULD NOT CHARGE, THE CHARGER RED LIGHT INDICATOR WAS ON. THERE WAS ALSO A BLACK GREASY SUBSTANCE AND CORROSION ON THE BASE OF THE HANDLES AND IN THE FOUR CHARGING BAYS. IT WAS NOTED THE CHARGERS HAD SECRETION LIQUID AT THE BOTTOM OF THE TERMINALS WHERE THE HANDLE SITS. ALL THREE HANDLES WENT DOWN ROUGHLY THE SAME TIME. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
07-10-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, ON A SLEEVE GASTRECTOMY, THREE HANDLES WOULD NOT CHARGE, THE CHARGER RED LIGHT INDICATOR WAS ON. THERE WAS ALSO A BLACK GREASY SUBSTANCE AND CORROSION ON THE BASE OF THE HANDLES AND IN THE FOUR CHARGING BAYS. IT WAS NOTED THE CHARGERS HAD SECRETION LIQUID AT THE BOTTOM OF THE TERMINALS WHERE THE HANDLE SITS. ALL THREE HANDLES WENT DOWN ROUGHLY THE SAME TIME. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
07-10-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, ON A SLEEVE GASTRECTOMY, THREE HANDLES WOULD NOT CHARGE, THE CHARGER RED LIGHT INDICATOR WAS ON. THERE WAS ALSO A BLACK GREASY SUBSTANCE AND CORROSION ON THE BASE OF THE HANDLES AND IN THE FOUR CHARGING BAYS. IT WAS NOTED THE CHARGERS HAD SECRETION LIQUID AT THE BOTTOM OF THE TERMINALS WHERE THE HANDLE SITS. ALL THREE HANDLES WENT DOWN ROUGHLY THE SAME TIME. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
05-01-2022 | Malfunction | us surgical puerto rico | 26-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC PERCUTANEOUS FISTULA CLOSURE, WHILE FIRING ACROSS THE STOMACH, THE TISSUE WAS VERY THICK AND THE STAPLE FIRE NEVER ENGAGED FULLY TO FIRE AND THE HANDLE MADE A CRACKING SOUND. THE SURGEON THEN USED A BIGGER RELOAD AND A NEW HANDLE WHERE WHEN FIRED THE HANDLE CRACKED AGAIN BUT WERE ABLE TO CONTINUE THE FIRING THROUGH TO THE END. THERE WAS NO PATIENT INJURY.
05-01-2022 | Malfunction | us surgical puerto rico | 26-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC PERCUTANEOUS FISTULA CLOSURE, WHILE FIRING ACROSS THE STOMACH, THE TISSUE WAS VERY THICK AND THE STAPLE FIRE NEVER ENGAGED FULLY TO FIRE AND THE HANDLE MADE A CRACKING SOUND. THE SURGEON THEN USED A BIGGER RELOAD AND A NEW HANDLE WHERE WHEN FIRED THE HANDLE CRACKED AGAIN BUT WERE ABLE TO CONTINUE THE FIRING THROUGH TO THE END. THERE WAS NO PATIENT INJURY.
20-05-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, THE HANDLE WAS SHOWING A "HANDLE INSIDE THE NO SYMBOL" ERROR. ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT.
20-05-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, THE HANDLE WAS SHOWING A "HANDLE INSIDE THE NO SYMBOL" ERROR. ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MINI GASTRIC BYPASS PROCEDURE, WHILE CREATING POUCH, THE STAPLES DID NOT FORM PROPERLY. THE DEVICE WAS STUCK AT THE PROXIMAL LEVEL WHICH RESULTS TO A STAPLE LINE PROXIMALLY INCOMPLETE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MINI GASTRIC BYPASS PROCEDURE, WHILE CREATING POUCH, THE STAPLES DID NOT FORM PROPERLY. THE DEVICE WAS STUCK AT THE PROXIMAL LEVEL WHICH RESULTS TO A STAPLE LINE PROXIMALLY INCOMPLETE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-06-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE HANDLE HAD AN ERROR MESSAGE, AND THE CHARGER STARTS WITH BLINKING GREEN AND CHANGES TO BLINKING RED. THE ISSUE WAS NOT RESOLVED. ANOTHER HANDLE WAS USED INSTEAD. THERE WAS NO PATIENT INVOLVEMENT.
17-06-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE HANDLE HAD AN ERROR MESSAGE, AND THE CHARGER STARTS WITH BLINKING GREEN AND CHANGES TO BLINKING RED. THE ISSUE WAS NOT RESOLVED. ANOTHER HANDLE WAS USED INSTEAD. THERE WAS NO PATIENT INVOLVEMENT.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC SURGERY, THE RELOAD WOULD NOT TOTALLY CLOSE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | TRI-STAPLE 2.0 | Component or Accessory Incompatibility | A1701 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY , THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | TRI-STAPLE 2.0 | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY , THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY , THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY , THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE ANTEGRADE PANCREATICOSPLENECTOMY WITH HEPATICOJEJUNOSTOMY WHILE GETTING READY TO DO HEPATIC. THERE WAS PROBLEM IN LOADING THE DEVICE. DEVICE WOULDN'T CYCLE WHEN RELOAD WAS ATTACHED. TO RESOLVE THE ISSUE, THE SURGEON GOT A NEW RELOAD FOR THE STAPLER. THERE WAS NO PATIENT INJURY.
18-12-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE ALTHOUGH THE USER WAS ABLE TO PRESS THE GREEN BUTTON AND WAS ABLE TO SQUEEZED THE HANDLE. THE STAPLER CLICKED AND CRASHED. LOAD LOOSENED SOME STAPLES AND DID NOT STAPLE. TO RESOLVE THE ISSUE, THE PRODUCT WAS REPLACED. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE OPERATOR FIRED A TOTAL 3 OF RELOADS SUCCESSFULLY WITH THE POWER HANDLE SET. WHILE REPLACING THE USED RELOADS BY THE SCRUB NURSE, THE DEVICE WAS DIFFICULT TO UNLOAD; THE USER TRIED TO DISCONNECT THE FIRED RELOADS WITH THE ADAPTER, BUT IT COULD NOT WORK AT ALL. THE CONNECTING PART OF THE RELOADS WAS STUCK STRONGLY WITH THE PROXIMAL PART OF THE ADAPTER. AND THEN, AT A CERTAIN POINT, THE CONNECTING PART OF RELOADS WAS BROKEN AND DETACHED COMPLETELY WITH THE ADAPTER. AT THE TIME, SOME PARTICLE OF THE CONNECTING PART REMAINED TO BE STUCK WITH THE ADAPTER. A NEW POWER HANDLE SET WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE OPERATOR FIRED A TOTAL 3 OF RELOADS SUCCESSFULLY WITH THE POWER HANDLE SET. WHILE REPLACING THE USED RELOADS BY THE SCRUB NURSE, THE DEVICE WAS DIFFICULT TO UNLOAD; THE USER TRIED TO DISCONNECT THE FIRED RELOADS WITH THE ADAPTER, BUT IT COULD NOT WORK AT ALL. THE CONNECTING PART OF THE RELOADS WAS STUCK STRONGLY WITH THE PROXIMAL PART OF THE ADAPTER. AND THEN, AT A CERTAIN POINT, THE CONNECTING PART OF RELOADS WAS BROKEN AND DETACHED COMPLETELY WITH THE ADAPTER. AT THE TIME, SOME PARTICLE OF THE CONNECTING PART REMAINED TO BE STUCK WITH THE ADAPTER. A NEW POWER HANDLE SET WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-04-2022 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-LAPAROSCOPIC BYPASS, THE HANDLE WAS NOT ABLE TO RECOGNIZE THE RELOAD AND THERE WAS INCOMPLETE STAPLE LINE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-04-2022 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-LAPAROSCOPIC BYPASS, THE HANDLE WAS NOT ABLE TO RECOGNIZE THE RELOAD AND THERE WAS INCOMPLETE STAPLE LINE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-04-2022 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-LAPAROSCOPIC BYPASS, THE HANDLE WAS NOT ABLE TO RECOGNIZE THE RELOAD AND THERE WAS INCOMPLETE STAPLE LINE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
11-01-2022 | Injury | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
11-01-2022 | Injury | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
11-01-2022 | Injury | covidien lp llc north haven | 26-01-2022 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
11-01-2022 | Injury | covidien lp llc north haven | 26-01-2022 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
11-01-2022 | Injury | covidien lp llc north haven | 26-01-2022 | SIGNIA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
11-01-2022 | Injury | covidien lp llc north haven | 26-01-2022 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
11-01-2022 | Injury | covidien lp llc north haven | 26-01-2022 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
11-01-2022 | Injury | covidien lp llc north haven | 26-01-2022 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
11-01-2022 | Injury | covidien lp llc north haven | 26-01-2022 | SIGNIA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
11-01-2022 | Injury | covidien lp llc north haven | 26-01-2022 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, ON THE SECOND FIRING FOR CREATING THE POUCH, THE STAPLER WAS ABLE TO FIRE, WHILE THE KNIFE WAS PASSING THROUGH, IT PUSHED ALL THE STAPLES TO THE END OF THE RELOAD AND NONE OF THE STAPLE WERE FORMED IN THE TISSUE, SO THE RELOAD CUT THE TISSUE AND DID NOT STAPLE IT AT ALL. THE SURGEON HAD TO CHANGE THE TECHNIQUE OF THE SURGERY AND DO A GASTRIC SLEEVE WITH THE BYPASS AND RESECT THE REMNANT STOMACH. THERE WAS TISSUE DAMAGE. THE OPERATION WAS EXTENDED BY 1 HOUR MORE THAN USUAL. IT WAS NOTED THAT THE PATIENT WENT TO THE UPC (CRITICAL PATIENT UNIT) FOR PREVENTION.
29-07-2021 | Malfunction | ethicon endo-surgery | 26-01-2022 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE GASTRIC CANCER SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE RESECTING THE STOMACH, THE DEVICE STOPPED REPEATEDLY AND INDICATED THAT THE TISSUE IS TOO THICK ON THE LIGHT-EMITTING DIODE SCREEN OF THE HANDLE. THE SURGEON WOULD STOP THE FIRING, WAITED FOR FIFTEEN TO TWENTY SECONDS TO RESUME FIRING. THE DEVICE WOULD ADVANCE FOR A FEW MILLIMETERS BEFORE STOPPING AGAIN. THE PROCESS REPEATED FOR SIX TIMES BEFORE THE RELOAD IS COMPLETELY FIRED AND OPENED WITHOUT ANY ISSUE. THE TISSUE APPEARED TO BE NORMAL. IT WAS NOTED THAT THE HANDLE WAS USED WITH NO OTHER ISSUES IN THE REMAINDER OF THE PROCEDURE. THE STAPLE LINE HAD NO ISSUE AND THERE WAS NO PATIENT INJURY.
22-09-2021 | Malfunction | us surgical puerto rico | 26-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE IN THE STOMACH, THE SECOND RELOAD WAS INSERTED INTO THE HANDLE THEN TO THE PATIENT'S CAVITY, THE GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED BUT THE STAPLER DID NOT FIRE. CHECKED IF THE RELOAD WAS PROPERLY MOUNTED BY OPENING AND CLOSING THE STAPLER. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-09-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE IN THE STOMACH, THE SECOND RELOAD WAS INSERTED INTO THE HANDLE THEN TO THE PATIENT'S CAVITY, THE GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED BUT THE STAPLER DID NOT FIRE. CHECKED IF THE RELOAD WAS PROPERLY MOUNTED BY OPENING AND CLOSING THE STAPLER. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-09-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE IN THE STOMACH, THE SECOND RELOAD WAS INSERTED INTO THE HANDLE THEN TO THE PATIENT'S CAVITY, THE GREEN BUTTON WAS PRESSED AND HANDLE WAS SQUEEZED BUT THE STAPLER DID NOT FIRE. CHECKED IF THE RELOAD WAS PROPERLY MOUNTED BY OPENING AND CLOSING THE STAPLER. THE STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-07-2019 | Injury | smith & nephew | 26-01-2022 | TENDON ANCHORS 8 | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: IT WAS REPORTED THAT, AFTER TWO YEARS OF A ROTATOR CUFF REPAIR USING 'TENDON ANCHORS, BONE ANCHORS AND THE BIOINDUCTIVE IMPLANT', A REVISION SURGERY WAS PERFORMED TO GRAB IMPLANT BECAUSE IT DIDN'T WORK PROPERLY IN THE BONE. THE PATIENT CURRENT STATUS IS UNKNOWN.
01-07-2019 | Injury | smith & nephew | 26-01-2022 | TENDON ANCHORS 8 | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, AFTER TWO YEARS OF A ROTATOR CUFF REPAIR USING 'TENDON ANCHORS, BONE ANCHORS AND THE BIOINDUCTIVE IMPLANT', A REVISION SURGERY WAS PERFORMED TO GRAB IMPLANT BECAUSE IT DIDN'T WORK PROPERLY IN THE BONE. THE PATIENT CURRENT STATUS IS UNKNOWN.
15-12-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE WOULD NOT FIRE.
15-12-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE WOULD NOT FIRE.
15-12-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | UNKNOWN STAPLING DEVICE | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE WOULD NOT FIRE.
15-12-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | UNKNOWN STAPLING DEVICE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE WOULD NOT FIRE.
15-12-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | UNKNOWN STAPLING DEVICE | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE WOULD NOT FIRE.
15-12-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | UNKNOWN STAPLING DEVICE | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE WOULD NOT FIRE.
24-09-2021 | Malfunction | us surgical puerto rico | 26-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE, WHEN BEING APPLIED ON THE GASTRIC TISSUE, THE FIRST RELOAD WAS ABLE TO FIRE BUT A LOUD BREAKING NOISE WAS HEARD. WITH THE SECOND RELOAD, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-09-2021 | Malfunction | us surgical puerto rico | 26-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE, WHEN BEING APPLIED ON THE GASTRIC TISSUE, THE FIRST RELOAD WAS ABLE TO FIRE BUT A LOUD BREAKING NOISE WAS HEARD. WITH THE SECOND RELOAD, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-09-2021 | Malfunction | covidien lp llc north haven | 26-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTRIC SLEEVE, WHEN BEING APPLIED ON THE GASTRIC TISSUE, THE FIRST RELOAD WAS ABLE TO FIRE BUT A LOUD BREAKING NOISE WAS HEARD. WITH THE SECOND RELOAD, THE DEVICE WAS UNABLE TO FIRE. THE SURGEON USED A NEW STAPLER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-01-2022 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC BYPASS PROCEDURE, THE STAPLE LINE WAS INCOMPLETE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-04-2022 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, THE STAPLE HAD POOR FORMATION. THE RELOAD COULD NOT COMPLETE THE TRIGGER WHICH RESULTS IN A INCOMPLETE STAPLE LINE. THE DEVICE WAS DIFFICULT TO UNLOAD AS WELL. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-04-2022 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, THE STAPLE HAD POOR FORMATION. THE RELOAD COULD NOT COMPLETE THE TRIGGER WHICH RESULTS IN A INCOMPLETE STAPLE LINE. THE DEVICE WAS DIFFICULT TO UNLOAD AS WELL. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-04-2022 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, THE STAPLE HAD POOR FORMATION. THE RELOAD COULD NOT COMPLETE THE TRIGGER WHICH RESULTS IN A INCOMPLETE STAPLE LINE. THE DEVICE WAS DIFFICULT TO UNLOAD AS WELL. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-04-2022 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2022 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, THE STAPLE HAD POOR FORMATION. THE RELOAD COULD NOT COMPLETE THE TRIGGER WHICH RESULTS IN A INCOMPLETE STAPLE LINE. THE DEVICE WAS DIFFICULT TO UNLOAD AS WELL. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-03-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, THE STAPLE HAD POOR FORMATION. THE RELOAD COULD NOT COMPLETE THE TRIGGER WHICH RESULTS IN A INCOMPLETE STAPLE LINE. THE DEVICE WAS DIFFICULT TO UNLOAD AS WELL. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-03-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, THE STAPLE HAD POOR FORMATION. THE RELOAD COULD NOT COMPLETE THE TRIGGER WHICH RESULTS IN A INCOMPLETE STAPLE LINE. THE DEVICE WAS DIFFICULT TO UNLOAD AS WELL. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-03-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, THE STAPLE HAD POOR FORMATION. THE RELOAD COULD NOT COMPLETE THE TRIGGER WHICH RESULTS IN A INCOMPLETE STAPLE LINE. THE DEVICE WAS DIFFICULT TO UNLOAD AS WELL. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, WHILE ON CHARGER, THE HANDLE FAILED TO CHARGE. IT WAS NOTED THAT THE LIGHT-EMITTING DIODE DISPLAY WAS BROKEN. THE HANDLE SWITCHED OFF DURING RECHARGE. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | No Display/Image | A090206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, WHILE ON CHARGER, THE HANDLE FAILED TO CHARGE. IT WAS NOTED THAT THE LIGHT-EMITTING DIODE DISPLAY WAS BROKEN. THE HANDLE SWITCHED OFF DURING RECHARGE. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
11-01-2022 | Malfunction | covidien lp llc north haven | 26-01-2022 | SIGNIA | Product Quality Problem | A0201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, WHILE ON CHARGER, THE HANDLE FAILED TO CHARGE. IT WAS NOTED THAT THE LIGHT-EMITTING DIODE DISPLAY WAS BROKEN. THE HANDLE SWITCHED OFF DURING RECHARGE. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
10-01-2022 | Injury | covidien lp llc north haven | 27-01-2022 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOIDECTOMY, WHILE ON THE HEMORRHOID, THE STAPLER PARTIALLY FIRED AND THERE WAS NO PROPER STAPLE FORMATION. STAPLER PIN FELL INTO THE PATIENT'S CAVITY AND WAS RETRIEVED USING FORCEPS. SUTURING WAS PERFORMED TO RESOLVE THE ISSUE.
10-01-2022 | Injury | covidien lp llc north haven | 27-01-2022 | EEA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOIDECTOMY, WHILE ON THE HEMORRHOID, THE STAPLER PARTIALLY FIRED AND THERE WAS NO PROPER STAPLE FORMATION. STAPLER PIN FELL INTO THE PATIENT'S CAVITY AND WAS RETRIEVED USING FORCEPS. SUTURING WAS PERFORMED TO RESOLVE THE ISSUE.
05-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL CLOSURE AND DETACHMENT IN A TOTAL MESORECTAL EXCISION, THE DEVICE WAS UNABLE TO FIRE NORMALLY. THE SURGEON USED A NEW HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMIHEPATECTOMY PROCEDURE, THE SURGEON HAD TROUBLE REMOVING RELOAD FROM STAPLER AFTER IT WAS FIRED. A NEW RELOAD WAS ATTACHED BUT FELT CLUNKY. CLOSED JAWS TO CYCLE STAPLER BUT THEY JAMMED SHUT AND UNABLE TO OPEN THEM. SURGEON ATTEMPTED TO FIRE STAPLER OUTSIDE OF PATIENT TO TEST DEVICE BUT STAPLER JAMMED. A NEW HANDLE AND RELOAD WERE OPENED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMIHEPATECTOMY PROCEDURE, THE SURGEON HAD TROUBLE REMOVING RELOAD FROM STAPLER AFTER IT WAS FIRED. A NEW RELOAD WAS ATTACHED BUT FELT CLUNKY. CLOSED JAWS TO CYCLE STAPLER BUT THEY JAMMED SHUT AND UNABLE TO OPEN THEM. SURGEON ATTEMPTED TO FIRE STAPLER OUTSIDE OF PATIENT TO TEST DEVICE BUT STAPLER JAMMED. A NEW HANDLE AND RELOAD WERE OPENED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMIHEPATECTOMY PROCEDURE, THE SURGEON HAD TROUBLE REMOVING RELOAD FROM STAPLER AFTER IT WAS FIRED. A NEW RELOAD WAS ATTACHED BUT FELT CLUNKY. CLOSED JAWS TO CYCLE STAPLER BUT THEY JAMMED SHUT AND UNABLE TO OPEN THEM. SURGEON ATTEMPTED TO FIRE STAPLER OUTSIDE OF PATIENT TO TEST DEVICE BUT STAPLER JAMMED. A NEW HANDLE AND RELOAD WERE OPENED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMIHEPATECTOMY PROCEDURE, THE SURGEON HAD TROUBLE REMOVING RELOAD FROM STAPLER AFTER IT WAS FIRED. A NEW RELOAD WAS ATTACHED BUT FELT CLUNKY. CLOSED JAWS TO CYCLE STAPLER BUT THEY JAMMED SHUT AND UNABLE TO OPEN THEM. SURGEON ATTEMPTED TO FIRE STAPLER OUTSIDE OF PATIENT TO TEST DEVICE BUT STAPLER JAMMED. A NEW HANDLE AND RELOAD WERE OPENED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMIHEPATECTOMY PROCEDURE, THE SURGEON HAD TROUBLE REMOVING RELOAD FROM STAPLER AFTER IT WAS FIRED. A NEW RELOAD WAS ATTACHED BUT FELT CLUNKY. CLOSED JAWS TO CYCLE STAPLER BUT THEY JAMMED SHUT AND UNABLE TO OPEN THEM. SURGEON ATTEMPTED TO FIRE STAPLER OUTSIDE OF PATIENT TO TEST DEVICE BUT STAPLER JAMMED. A NEW HANDLE AND RELOAD WERE OPENED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMIHEPATECTOMY PROCEDURE, THE SURGEON HAD TROUBLE REMOVING RELOAD FROM STAPLER AFTER IT WAS FIRED. A NEW RELOAD WAS ATTACHED BUT FELT CLUNKY. CLOSED JAWS TO CYCLE STAPLER BUT THEY JAMMED SHUT AND UNABLE TO OPEN THEM. SURGEON ATTEMPTED TO FIRE STAPLER OUTSIDE OF PATIENT TO TEST DEVICE BUT STAPLER JAMMED. A NEW HANDLE AND RELOAD WERE OPENED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-02-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-LAPAROSCOPIC BYPASS, THE HANDLE WAS NOT ABLE TO RECOGNIZE THE RELOAD AND THERE WAS INCOMPLETE STAPLE LINE INCOMPLETE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
11-02-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-LAPAROSCOPIC BYPASS, THE HANDLE WAS NOT ABLE TO RECOGNIZE THE RELOAD AND THERE WAS INCOMPLETE STAPLE LINE INCOMPLETE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
11-02-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-LAPAROSCOPIC BYPASS, THE HANDLE WAS NOT ABLE TO RECOGNIZE THE RELOAD AND THERE WAS INCOMPLETE STAPLE LINE INCOMPLETE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
27-01-2022 | Malfunction | celestica de monterrey s.a de c.v | 27-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BYPASS, THE DEVICE DID NOT FIRE AND WAS DIFFICULT TO UNLOAD. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-01-2022 | Malfunction | celestica de monterrey s.a de c.v | 27-01-2022 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BYPASS, THE DEVICE DID NOT FIRE AND WAS DIFFICULT TO UNLOAD. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-02-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BYPASS, THE DEVICE DID NOT FIRE AND WAS DIFFICULT TO UNLOAD. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-02-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BYPASS, THE DEVICE DID NOT FIRE AND WAS DIFFICULT TO UNLOAD. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-02-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BYPASS, THE DEVICE DID NOT FIRE AND WAS DIFFICULT TO UNLOAD. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Injury | us surgical puerto rico | 27-01-2022 | GIA | Failure to Form Staple | A050704 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING OPEN HEMICOLECTOMY, WHEN TRANSECTING COLON, FIRST FIRING WAS FINE. WHEN RELOAD WAS REPLACED, AFTER FIRING THE SURGEON WAS ABLE TO SEE STAPLES ONLY IN THE BEGINNING AND END OF STAPLE LINE. THE SURGEON OVER SEWED THE STAPLE LINE BY HAND SUTURING.
11-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHILE ON THE DETACHMENT OF THE INTERLOBULAR FISSURE, THE HANDLE COULD NOT BE PRESSED AND THE TRIGGER COULD NOT BE FIRED AS IF IT WAS STUCK. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHILE ON THE DETACHMENT OF THE INTERLOBULAR FISSURE, THE HANDLE COULD NOT BE PRESSED AND THE TRIGGER COULD NOT BE FIRED AS IF IT WAS STUCK. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-08-2021 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, WHEN PROCEEDING WITH THE FIRST SHOT OF THE STAPLER TO THE STOMACH, WHEN ENTERING THE CAVITY, STAPLER DID NOT FIRE. THE RELOAD REMAINED COMPLETELY CLOSE AND CANNOT BE REOPENED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-08-2021 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, WHEN PROCEEDING WITH THE FIRST SHOT OF THE STAPLER TO THE STOMACH, WHEN ENTERING THE CAVITY, STAPLER DID NOT FIRE. THE RELOAD REMAINED COMPLETELY CLOSE AND CANNOT BE REOPENED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE OPERATOR FIRED A TOTAL THREE OF RELOADS SUCCESSFULLY WITH THE POWER HANDLE SET. WHILE REPLACING THE USED RELOADS BY THE SCRUB NURSE, THE DEVICE WAS DIFFICULT TO UNLOAD; THE USER TRIED TO DISCONNECT THE FIRED RELOADS WITH THE ADAPTER, BUT IT COULD NOT WORK AT ALL. THE CONNECTING PART OF THE RELOADS WAS STUCK STRONGLY WITH THE PROXIMAL PART OF THE ADAPTER. AND THEN, AT A CERTAIN POINT, THE CONNECTING PART OF RELOADS WAS BROKEN AND DETACHED COMPLETELY WITH THE ADAPTER. AT THE TIME, SOME PARTICLE OF THE CONNECTING PART REMAINED TO BE STUCK WITH THE ADAPTER. A NEW POWER HANDLE SET WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE OPERATOR FIRED A TOTAL THREE OF RELOADS SUCCESSFULLY WITH THE POWER HANDLE SET. WHILE REPLACING THE USED RELOADS BY THE SCRUB NURSE, THE DEVICE WAS DIFFICULT TO UNLOAD; THE USER TRIED TO DISCONNECT THE FIRED RELOADS WITH THE ADAPTER, BUT IT COULD NOT WORK AT ALL. THE CONNECTING PART OF THE RELOADS WAS STUCK STRONGLY WITH THE PROXIMAL PART OF THE ADAPTER. AND THEN, AT A CERTAIN POINT, THE CONNECTING PART OF RELOADS WAS BROKEN AND DETACHED COMPLETELY WITH THE ADAPTER. AT THE TIME, SOME PARTICLE OF THE CONNECTING PART REMAINED TO BE STUCK WITH THE ADAPTER. A NEW POWER HANDLE SET WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
11-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE CONNECTION TO THE HANDLE WAS HARD, AND THE PARTS TO PREVENT ACCIDENTAL FIRING BROKE OFF.
04-10-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING SERVICING, WHEN CHARGING THE HANDLE, THE GREEN LAMP ON THE CHARGER WAS FOUND TO BE CAME OFF. THE POWER HANDLE WAS REMOVED FROM THE CHARGER, BUT IT DID NOT START UP. WHEN THE POWER HANDLE WAS RECONNECTED TO THE CHARGER, THE GREEN LAMP ON THE CHARGER STARTED FLASHING. IT WAS ALSO REPORTED THAT THE BATTERY CHARGER WAS FLASHING RED. THERE WAS NO PATIENT INVOLVEMENT.
04-10-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING SERVICING, WHEN CHARGING THE HANDLE, THE GREEN LAMP ON THE CHARGER WAS FOUND TO BE CAME OFF. THE POWER HANDLE WAS REMOVED FROM THE CHARGER, BUT IT DID NOT START UP. WHEN THE POWER HANDLE WAS RECONNECTED TO THE CHARGER, THE GREEN LAMP ON THE CHARGER STARTED FLASHING. IT WAS ALSO REPORTED THAT THE BATTERY CHARGER WAS FLASHING RED. THERE WAS NO PATIENT INVOLVEMENT.
04-10-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING SERVICING, WHEN CHARGING THE HANDLE, THE GREEN LAMP ON THE CHARGER WAS FOUND TO BE CAME OFF. THE POWER HANDLE WAS REMOVED FROM THE CHARGER, BUT IT DID NOT START UP. WHEN THE POWER HANDLE WAS RECONNECTED TO THE CHARGER, THE GREEN LAMP ON THE CHARGER STARTED FLASHING. IT WAS ALSO REPORTED THAT THE BATTERY CHARGER WAS FLASHING RED. THERE WAS NO PATIENT INVOLVEMENT.
30-08-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, BOTH HANDLES WERE NOT ABLE TO CHARGE. IN ADDITION, THE CHARGER WAS FLASHING RED WHEN BOTH HANDLES WERE INSERTED TO CHARGE. THE SURGEON USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
30-08-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, BOTH HANDLES WERE NOT ABLE TO CHARGE. IN ADDITION, THE CHARGER WAS FLASHING RED WHEN BOTH HANDLES WERE INSERTED TO CHARGE. THE SURGEON USED ANOTHER DEVICE TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT.
28-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING PRIOR TO A PROCEDURE, HANDLE COULD NOT BE CHARGED. THERE WAS DEFORMATION OF THE CHARGING PART ON THE BASE OF THE HANDLE. NO PATIENT INVOLVEMENT.
28-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Product Quality Problem | A0201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING PRIOR TO A PROCEDURE, HANDLE COULD NOT BE CHARGED. THERE WAS DEFORMATION OF THE CHARGING PART ON THE BASE OF THE HANDLE. NO PATIENT INVOLVEMENT.
28-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING PRIOR TO A PROCEDURE, HANDLE COULD NOT BE CHARGED. THERE WAS DEFORMATION OF THE CHARGING PART ON THE BASE OF THE HANDLE. NO PATIENT INVOLVEMENT.
05-10-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING, THE HANDLE DID NOT START UP WHEN REMOVED FROM THE CHARGER. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INVOLVEMENT.
05-10-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING, THE HANDLE DID NOT START UP WHEN REMOVED FROM THE CHARGER. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INVOLVEMENT.
29-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING FOR APPROXIMATELY ONE HOUR, THE DEVICE DID NOT CHARGE AND DID NOT POWER ON. OTHER CHARGERS WERE TRIED AND A REBOOT WAS MADE BUT THE SAME ISSUE OCCURRED. THERE WAS NO PATIENT INVOLVEMENT.
29-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING FOR APPROXIMATELY ONE HOUR, THE DEVICE DID NOT CHARGE AND DID NOT POWER ON. OTHER CHARGERS WERE TRIED AND A REBOOT WAS MADE BUT THE SAME ISSUE OCCURRED. THERE WAS NO PATIENT INVOLVEMENT.
12-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE. TO COMPLETE THE CASE, THE SURGEON REPLACED WITH A NEW PRODUCT. THERE WAS NO PATIENT INJURY.
01-10-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE HANDLE WOULD NOT CHARGE AND THE GREEN LIGHT DID NOT ILLUMINATE. THE HANDLE WAS TESTED WITH ANOTHER CHARGER AND IT FLASHED GREEN INITIALLY BUT FLASHED RED LIGHT AFTER TWO MINUTES. BOTH CHARGERS WORKED FINE WITH ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
01-10-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE HANDLE WOULD NOT CHARGE AND THE GREEN LIGHT DID NOT ILLUMINATE. THE HANDLE WAS TESTED WITH ANOTHER CHARGER AND IT FLASHED GREEN INITIALLY BUT FLASHED RED LIGHT AFTER TWO MINUTES. BOTH CHARGERS WORKED FINE WITH ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
01-01-2021 | Malfunction | ethicon endo-surgery | 27-01-2022 | POWERED PLUS STANDARD 45 | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DESCRIPTION OF THE DEVIATION: PACKAGING IS BROKEN, SO THE PRODUCT CANNOT BE USED. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
04-10-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 27-01-2022 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A TRANSABDOMINAL PREPERITONEAL (TAPP) PROCEDURE ON (B)(6) 2021, THE BARD/DAVOL OPTIFIX FIXATION DEVICE WAS USED TO FIXATE A BARD/DAVOL 3DMAX MESH. IT WAS REPORTED THAT A BROKEN FASTENER WAS DEPLOYED AND WHILE DEPLOYING THE SECOND FASTENER NEAR THE COOPER'S LIGAMENT, THE BARREL INSERT FROM THE SHAFT TIP OF THE FIXATION DEVICE BECAME DETACHED. AS REPORTED, THE BROKEN/DETACHED COMPONENTS WERE REMOVED FROM THE PATIENT'S BODY. AS REPORTED, COUNTER-PRESSURE WAS NOT APPLIED SINCE FIXATION WAS BEING APPLIED AT COOPER'S LIGAMENT. THERE WAS A PROCEDURAL DELAY DUE TO THE REPORTED EVENT AND NO REPORTED PATIENT INJURY.
04-10-2021 | Malfunction | davol inc sub c.r bard inc -1213643 | 27-01-2022 | OPTIFIX STRAIGHT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A TRANSABDOMINAL PREPERITONEAL (TAPP) PROCEDURE ON (B)(6) 2021, THE BARD/DAVOL OPTIFIX FIXATION DEVICE WAS USED TO FIXATE A BARD/DAVOL 3DMAX MESH. IT WAS REPORTED THAT A BROKEN FASTENER WAS DEPLOYED AND WHILE DEPLOYING THE SECOND FASTENER NEAR THE COOPER'S LIGAMENT, THE BARREL INSERT FROM THE SHAFT TIP OF THE FIXATION DEVICE BECAME DETACHED. AS REPORTED, THE BROKEN/DETACHED COMPONENTS WERE REMOVED FROM THE PATIENT'S BODY. AS REPORTED, COUNTER-PRESSURE WAS NOT APPLIED SINCE FIXATION WAS BEING APPLIED AT COOPER'S LIGAMENT. THERE WAS A PROCEDURAL DELAY DUE TO THE REPORTED EVENT AND NO REPORTED PATIENT INJURY.
01-12-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC GASTRIC BYPASS, THE RELOAD WAS INSERTED TO THE HANDLE BUT THE TRIGGER COULD NOT BE PULLED. ANOTHER RELOAD WAS USED AND THE DEVICE WAS FIRED NORMALLY. THERE WAS NO PATIENT INJURY.
12-10-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE HANDLE WOULD NOT CHARGE. THE DEVICE WAS TESTED WITH ANOTHER CHARGER BUT HAD SAME RESULTS. THE CHARGER WORKED THOUGH WITH ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
12-10-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE HANDLE WOULD NOT CHARGE. THE DEVICE WAS TESTED WITH ANOTHER CHARGER BUT HAD SAME RESULTS. THE CHARGER WORKED THOUGH WITH ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
22-12-2021 | Injury | us surgical puerto rico | 27-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING RECTOSIGMOIDECTOMY, THE STAPLE LINE DID NOT HOLD THE TISSUE TIGHTLY AND THE STAPLE LINE OPENED.
22-12-2021 | Injury | us surgical puerto rico | 27-01-2022 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING RECTOSIGMOIDECTOMY, THE STAPLE LINE DID NOT HOLD THE TISSUE TIGHTLY AND THE STAPLE LINE OPENED.
20-08-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHILE THE HANDLE WAS ON THE CHARGER, THE HANDLE FAILED TO CHARGE AND DID NOT ACTIVATE. A MANUAL STAPLER WAS TAKEN OUT INSTEAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
20-08-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHILE THE HANDLE WAS ON THE CHARGER, THE HANDLE FAILED TO CHARGE AND DID NOT ACTIVATE. A MANUAL STAPLER WAS TAKEN OUT INSTEAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
21-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEVICE PREPARATION, THE POWERED HANDLE DID NOT TURN ON WHEN THE DEVICE WAS REMOVED FROM THE CHARGER. THE HANDLE WAS PLACED BACK ONTO THE CHARGER, BUT THE DEVICE DID NOT CHARGE. ANOTHER CHARGER WAS USED, BUT THE ISSUE DID NOT RESOLVE. THERE WAS NO PATIENT INVOLVED.
21-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEVICE PREPARATION, THE POWERED HANDLE DID NOT TURN ON WHEN THE DEVICE WAS REMOVED FROM THE CHARGER. THE HANDLE WAS PLACED BACK ONTO THE CHARGER, BUT THE DEVICE DID NOT CHARGE. ANOTHER CHARGER WAS USED, BUT THE ISSUE DID NOT RESOLVE. THERE WAS NO PATIENT INVOLVED.
30-06-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION/PRIOR TO PATIENT CONTACT, HANDLE COULD NOT BE FULLY CHARGED SEVERAL TIMES, WHEN CLOSING THE SHELL, THE BATTERY SUDDENLY SHOWED NO CHARGE. IT WAS NOTED THAT THE CHARGER LIGHT INDICATED THAT THE HANDLE IS FULLY CHARGED, BUT THIS WAS NOT THE CASE. THE HANDLE ALSO HAS WEAK DISPLAY. DEVICE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT INVOLVEMENT.
30-06-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION/PRIOR TO PATIENT CONTACT, HANDLE COULD NOT BE FULLY CHARGED SEVERAL TIMES, WHEN CLOSING THE SHELL, THE BATTERY SUDDENLY SHOWED NO CHARGE. IT WAS NOTED THAT THE CHARGER LIGHT INDICATED THAT THE HANDLE IS FULLY CHARGED, BUT THIS WAS NOT THE CASE. THE HANDLE ALSO HAS WEAK DISPLAY. DEVICE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT INVOLVEMENT.
30-06-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION/PRIOR TO PATIENT CONTACT, HANDLE COULD NOT BE FULLY CHARGED SEVERAL TIMES, WHEN CLOSING THE SHELL, THE BATTERY SUDDENLY SHOWED NO CHARGE. IT WAS NOTED THAT THE CHARGER LIGHT INDICATED THAT THE HANDLE IS FULLY CHARGED, BUT THIS WAS NOT THE CASE. THE HANDLE ALSO HAS WEAK DISPLAY. DEVICE WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT INVOLVEMENT.
09-06-2021 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRIC SLEEVE/ REPAIR OF DIAPHRAGMATIC HERNIA, ON SECOND FIRING WITH A 60 MILLIMETER BLACK RELOAD, A VERY LOUD NOISE WAS HEARD FROM THE STAPLER AND THE RELOAD DID NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED BUT DID NOT PERFORM STAPLING AS WELL. REPLACE WITH A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-06-2021 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRIC SLEEVE/ REPAIR OF DIAPHRAGMATIC HERNIA, ON SECOND FIRING WITH A 60 MILLIMETER BLACK RELOAD, A VERY LOUD NOISE WAS HEARD FROM THE STAPLER AND THE RELOAD DID NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED BUT DID NOT PERFORM STAPLING AS WELL. REPLACE WITH A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-12-2020 | Malfunction | ethicon endo-surgery | 27-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHEN USING THE STAPLER FOR AN INTESTINAL RESECTION, IT DOES NOT STAPLE, AND IT BLOCKS.
27-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, THE DEVICE HAD NO ACTIVATION SOUND. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
27-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, THE DEVICE HAD NO ACTIVATION SOUND. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
12-01-2022 | Injury | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS PROCEDURE, WHILE BEING ANASTOMOSED, THE DEVICE WOULD NOT FIRE ALL THE WAY, RESULTING TO A DISTALLY INCOMPLETE STAPLE LINE. ANOTHER RELOAD WAS LOADED TO THE HANDLE BUT DID NOT FIRE. A THIRD RELOAD AND ANOTHER ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS ALSO A LEAK IDENTIFIED ON A LEAK TEST. THE SURGEON HAD TO OVER SEW THE LEAK TO PREVENT FURTHER DAMAGE.
12-01-2022 | Injury | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS PROCEDURE, WHILE BEING ANASTOMOSED, THE DEVICE WOULD NOT FIRE ALL THE WAY, RESULTING TO A DISTALLY INCOMPLETE STAPLE LINE. ANOTHER RELOAD WAS LOADED TO THE HANDLE BUT DID NOT FIRE. A THIRD RELOAD AND ANOTHER ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS ALSO A LEAK IDENTIFIED ON A LEAK TEST. THE SURGEON HAD TO OVER SEW THE LEAK TO PREVENT FURTHER DAMAGE.
12-01-2022 | Injury | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS PROCEDURE, WHILE BEING ANASTOMOSED, THE DEVICE WOULD NOT FIRE ALL THE WAY, RESULTING TO A DISTALLY INCOMPLETE STAPLE LINE. ANOTHER RELOAD WAS LOADED TO THE HANDLE BUT DID NOT FIRE. A THIRD RELOAD AND ANOTHER ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS ALSO A LEAK IDENTIFIED ON A LEAK TEST. THE SURGEON HAD TO OVER SEW THE LEAK TO PREVENT FURTHER DAMAGE.
12-01-2022 | Injury | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS PROCEDURE, WHILE BEING ANASTOMOSED, THE DEVICE WOULD NOT FIRE ALL THE WAY, RESULTING TO A DISTALLY INCOMPLETE STAPLE LINE. ANOTHER RELOAD WAS LOADED TO THE HANDLE BUT DID NOT FIRE. A THIRD RELOAD AND ANOTHER ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS ALSO A LEAK IDENTIFIED ON A LEAK TEST. THE SURGEON HAD TO OVER SEW THE LEAK TO PREVENT FURTHER DAMAGE.
27-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE SALES REPRESENTATIVE WAS THERE TO CHECK AND UPGRADE TO NEW SOFTWARE, THE HANDLE FAILED TO CHARGE, THE BOTTOM OF THE HANDLE AND THE CHARGER WERE ALL RUSTY AFTER STANDING AND NOT BEING USED FOR 18 MONTHS. TROUBLESHOOTING STEPS WERE PERFORMED BUT THE ISSUE WAS NOT RESOLVED. A NEW CHARGER WAS USED, BUT THE HANDLE STILL WOULD NOT CHARGE AT ALL AND THE HANDLE WAS ALSO REPLACED RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
27-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE SALES REPRESENTATIVE WAS THERE TO CHECK AND UPGRADE TO NEW SOFTWARE, THE HANDLE FAILED TO CHARGE, THE BOTTOM OF THE HANDLE AND THE CHARGER WERE ALL RUSTY AFTER STANDING AND NOT BEING USED FOR 18 MONTHS. TROUBLESHOOTING STEPS WERE PERFORMED BUT THE ISSUE WAS NOT RESOLVED. A NEW CHARGER WAS USED, BUT THE HANDLE STILL WOULD NOT CHARGE AT ALL AND THE HANDLE WAS ALSO REPLACED RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
27-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Application Program Version or Upgrade Problem | A110206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE SALES REPRESENTATIVE WAS THERE TO CHECK AND UPGRADE TO NEW SOFTWARE, THE HANDLE FAILED TO CHARGE, THE BOTTOM OF THE HANDLE AND THE CHARGER WERE ALL RUSTY AFTER STANDING AND NOT BEING USED FOR 18 MONTHS. TROUBLESHOOTING STEPS WERE PERFORMED BUT THE ISSUE WAS NOT RESOLVED. A NEW CHARGER WAS USED, BUT THE HANDLE STILL WOULD NOT CHARGE AT ALL AND THE HANDLE WAS ALSO REPLACED RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
27-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE SALES REPRESENTATIVE WAS THERE TO CHECK AND UPGRADE TO NEW SOFTWARE, THE HANDLE FAILED TO CHARGE, THE BOTTOM OF THE HANDLE AND THE CHARGER WERE ALL RUSTY AFTER STANDING AND NOT BEING USED FOR 18 MONTHS. TROUBLESHOOTING STEPS WERE PERFORMED BUT THE ISSUE WAS NOT RESOLVED. A NEW CHARGER WAS USED, BUT THE HANDLE STILL WOULD NOT CHARGE AT ALL AND THE HANDLE WAS ALSO REPLACED RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
09-06-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRIC SLEEVE/ REPAIR OF DIAPHRAGMATIC HERNIA, ON SECOND FIRING WITH A 60 MILLIMETER BLACK RELOAD, A VERY LOUD NOISE WAS HEARD FROM THE STAPLER AND THE RELOAD DID NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED BUT DID NOT PERFORM STAPLING AS WELL. REPLACE WITH A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-06-2021 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRIC SLEEVE/ REPAIR OF DIAPHRAGMATIC HERNIA, ON SECOND FIRING WITH A 60 MILLIMETER BLACK RELOAD, A VERY LOUD NOISE WAS HEARD FROM THE STAPLER AND THE RELOAD DID NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED BUT DID NOT PERFORM STAPLING AS WELL. REPLACE WITH A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-06-2021 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRIC SLEEVE/ REPAIR OF DIAPHRAGMATIC HERNIA, ON SECOND FIRING WITH A 60 MILLIMETER BLACK RELOAD, A VERY LOUD NOISE WAS HEARD FROM THE STAPLER AND THE RELOAD DID NOT FIRE. ANOTHER HANDLE AND RELOAD WAS USED BUT DID NOT PERFORM STAPLING AS WELL. REPLACE WITH A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS PROCEDURE, WHILE BEING ANASTOMOSED, THE DEVICE WOULD NOT FIRE ALL THE WAY, RESULTING TO A DISTALLY INCOMPLETE STAPLE LINE. ANOTHER RELOAD WAS LOADED TO THE HANDLE BUT DID NOT FIRE. A THIRD RELOAD AND ANOTHER ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS ALSO A LEAK IDENTIFIED ON A LEAK TEST. THE SURGEON HAD TO OVER SEW THE LEAK TO PREVENT FURTHER DAMAGE.
27-09-2021 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROINTESTINAL RECONSTRUCTION IN AND OF ROUX, THE RELOAD WAS LOADED INTO THE HANDLE, WHEN OPENING AND CLOSING IT, IT DID NOT REOPEN THE JAWS. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-09-2021 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC GASTROINTESTINAL RECONSTRUCTION IN AND OF ROUX, THE RELOAD WAS LOADED INTO THE HANDLE, WHEN OPENING AND CLOSING IT, IT DID NOT REOPEN THE JAWS. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, WHILE IN THE FREE LOBE, THE DEVICE DID NOT FIRE. EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. A NEW HANDLE FROM THE SAME MANUFACTURER WAS USED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INJURY.
21-10-2021 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LAPAROSCOPIC PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, THE HANDLE BLOCKED, THE DEVICE WORKED WITH THE FIRST 2 RELOADS AND THEN STALLED ON THE 3RD RELOAD, THE STAPLER WAS UNABLE TO FIRE, THE GREEN BUTTON WAS NOT PRESSED AND THE SURGEON COULD NOT SQUEEZE THE HANDLE. THE USER TOOK ANOTHER HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-10-2021 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LAPAROSCOPIC PROCEDURE, THERE WAS A PROBLEM LOADING THE DEVICE, THE HANDLE BLOCKED, THE DEVICE WORKED WITH THE FIRST 2 RELOADS AND THEN STALLED ON THE 3RD RELOAD, THE STAPLER WAS UNABLE TO FIRE, THE GREEN BUTTON WAS NOT PRESSED AND THE SURGEON COULD NOT SQUEEZE THE HANDLE. THE USER TOOK ANOTHER HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Malfunction | ethicon endo-surgery | 27-01-2022 | PROXIMATE*LINEAR CUTTER | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT COLECTOMY PROCEDURE, RELOAD OF THE DEVICE BROKE. ANOTHER RELOAD WAS PUT IN AND THE BUTTON TO FIRE DEVICE WOULD NOT MOVE FORWARD. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
14-12-2021 | Malfunction | ethicon endo-surgery | 27-01-2022 | PROXIMATE*LINEAR CUTTER | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT COLECTOMY PROCEDURE, RELOAD OF THE DEVICE BROKE. ANOTHER RELOAD WAS PUT IN AND THE BUTTON TO FIRE DEVICE WOULD NOT MOVE FORWARD. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
12-01-2022 | Malfunction | ethicon endo-surgery | 27-01-2022 | ECHELON FLEX 60 COMPACT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE AND THEN USED SUTURE TO OVERSEW. CHANGED THE NEW ONE TO COMPLETE SURGERY . THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND WAS NOT ABLE TO SQUEEZED THE HANDLE. TO COMPLETE THE CASE, THE SURGEON REPLACED WITH A NEW PRODUCT. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | ethicon endo-surgery | 27-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, DEVICE "MISFIRED & LOCKED AFTER 1 STEP." ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
07-01-2022 | Malfunction | ethicon endo-surgery | 27-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, DEVICE "MISFIRED & LOCKED AFTER 1 STEP." ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
07-01-2022 | Malfunction | ethicon endo-surgery | 27-01-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, DEVICE "MISFIRED & LOCKED AFTER 1 STEP." ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
27-12-2021 | Malfunction | ethicon endo-surgery | 27-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE WOULD NOT OPEN OR UNLOCK AFTER 2 TIMES. THERE WERE NO PATIENT CONSEQUENCES.
04-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOWER ANTERIOR RESECTION, DURING RECTAL DISSECTION, AT THE SECOND FIRING, A STRANGE NOISE WAS HEARD FROM THE HANDLE, AND THE HANDLE COULD NOT BE GRASPED AFTERWARDS. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD WAS USED BUT THERE WAS ALSO A STRANGE NOISE FROM THE HANDLE AND IT ALSO COULD NOT BE FIRED. ANOTHER RELOAD WAS TAKEN OUT AND THE RECTUM WAS DISSECTED TO THE END WITH NO ISSUES. THERE WAS NO PATIENT INJURY.
28-04-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | GIA | Failure to Align | A050701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE WAS DIFFICULT OR UNABLE TO CLOSE. THERE WAS NO PATIENT INVOLVED.
15-03-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD PARTIALLY FIRED. REPLACED THE PRODUCT TO RESOLVE THE ISSUE.
27-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD PARTIALLY FIRED. REPLACED THE PRODUCT TO RESOLVE THE ISSUE.
06-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN TRANSECTING THE STOMACH, WHILE FIRING THE STAPLER, MID STAPLE LINE, THE HANDLE CRACKED, THE ARTICULATION KNOB BROKE OFF AND SPLIT INTO TWO THEN FELL TO THE FLOOR. THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY. OPENED ANOTHER HANDLE AND RELOAD TO FINISH THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN TRANSECTING THE STOMACH, WHILE FIRING THE STAPLER, MID STAPLE LINE, THE HANDLE CRACKED, THE ARTICULATION KNOB BROKE OFF AND SPLIT INTO TWO THEN FELL TO THE FLOOR. THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY. OPENED ANOTHER HANDLE AND RELOAD TO FINISH THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN TRANSECTING THE STOMACH, WHILE FIRING THE STAPLER, MID STAPLE LINE, THE HANDLE CRACKED, THE ARTICULATION KNOB BROKE OFF AND SPLIT INTO TWO THEN FELL TO THE FLOOR. THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY. OPENED ANOTHER HANDLE AND RELOAD TO FINISH THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
06-01-2022 | Malfunction | covidien lp llc north haven | 27-01-2022 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN TRANSECTING THE STOMACH, WHILE FIRING THE STAPLER, MID STAPLE LINE, THE HANDLE CRACKED, THE ARTICULATION KNOB BROKE OFF AND SPLIT INTO TWO THEN FELL TO THE FLOOR. THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY. OPENED ANOTHER HANDLE AND RELOAD TO FINISH THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN TRANSECTING THE STOMACH, WHEN THE HANDLE WAS SQUEEZED AT THE END OF THE STAPLE LINE, THE ARTICULATION KNOB BROKE OFF, SPLIT IN TWO AND FELL TO THE FLOOR. OPENED ANOTHER HANDLE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | us surgical puerto rico | 27-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN TRANSECTING THE STOMACH, WHEN THE HANDLE WAS SQUEEZED AT THE END OF THE STAPLE LINE, THE ARTICULATION KNOB BROKE OFF, SPLIT IN TWO AND FELL TO THE FLOOR. OPENED ANOTHER HANDLE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
07-01-2022 | Malfunction | us surgical puerto rico | 28-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN SUBTOTAL GASTRECTOMY, WHILE DIVIDING THE STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN AFTER PRESSING THE GREEN BUTTON. ANOTHER HANDLE WITH THE SAME LOT NUMBER OPENED HAD THE SAME ISSUE. A NEW HANDLE WITH DIFFERENT LOT NUMBER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 28-01-2022 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN LIVER RESECTION, AFTER FOUR FIRINGS DURING THE FIRING ACROSS THE VESSELS, T THE DEVICE WIGGLED BACK AND FORTH THROUGHOUT THE FIRING WHEN IT WAS FULLY ARTICULATED. IT WAS STATED THAT THE STAPLING HAD GOOD OUTCOME. THE NEXT THREE FIRINGS HAD NO ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | us surgical puerto rico | 28-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC PARTIAL PULMONARY RESECTION, ON SECOND FIRING, THE DEVICE WAS UNABLE TO FIRE ON A THICK TISSUE-HANDLE COULD NOT BE SQUEEZED AND AN ABNORMAL SOUND WAS HEARD. THE JAWS WERE RELEASED AND A COMPETITOR'S STAPLER WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | us surgical puerto rico | 28-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC PARTIAL PULMONARY RESECTION, ON SECOND FIRING, THE DEVICE WAS UNABLE TO FIRE ON A THICK TISSUE-HANDLE COULD NOT BE SQUEEZED AND AN ABNORMAL SOUND WAS HEARD. THE JAWS WERE RELEASED AND A COMPETITOR'S STAPLER WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | covidien lp llc north haven | 28-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC PARTIAL PULMONARY RESECTION, ON SECOND FIRING, THE DEVICE WAS UNABLE TO FIRE ON A THICK TISSUE - HANDLE COULD NOT BE SQUEEZED AND AN ABNORMAL SOUND WAS HEARD. THE JAWS WERE RELEASED AND A COMPETITOR'S STAPLER WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2022 | Malfunction | davol inc sub c.r bard inc -1213643 | 28-01-2022 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, A DURING LAPAROSCOPIC INGUINAL HERNIA REPAIR PROCEDURE ON (B)(6) 2022, THREE OPTIFIX STRAIGHT FIXATION DEVICES WERE USED TO FIXATE A MESH (UNSPECIFIED) INTO SOFT TISSUE. UPON USING THE 1ST OPTIFIX STRAIGHT (DEVICE #1), IT WAS NOTED THE FASTENERS WERE ¿SHREDDING¿ INTO PIECES. A SECOND OPTIFIX STRAIGHT (DEVICE #2) WAS THEN USED AND IT WAS REPORTED IT WOULD NOT FIRE PROPERLY. A THIRD OPTIFIX STRAIGHT (DEVICE #3) WAS USED AND AT THE END OF THE CASE, THE HANDLE BECAME DIFFICULT TO ACTUATE. THE CONTACT CONFIRMED THAT THE BROKEN FASTENERS WERE REMOVED FROM THE PATIENT AND THERE WAS NO PATIENT INJURY AS A RESULT OF THE BROKEN FASTENER THAT PRESENTED DURING USE OF DEVICE #1.
14-01-2022 | Malfunction | ethicon endo-surgery | 28-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT IT WAS IMPOSSIBLE TO OPEN THE DEVICE DURING THE USE. ANOTHER DEVICE FROM SAME PRODUCT CODE HAS BEEN USED. AFTER THE PROCEDURE, THEY TRY AGAIN TO USE THE DEVICE WITH SUCCESS. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
11-10-2021 | Malfunction | covidien lp llc north haven | 28-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE HANDLE IS DEAD AND WOULD NOT CHARGE. IT WAS NOTED THAT BOTH DEVICES HAVE BEEN STANDING UNUSED FOR A YEAR AND A HALF DURING THE PANDEMIC. THE CHARGER WAS RUSTY. THE ISSUE WAS NOT RESOLVED. THERE WAS NO PATIENT INVOLVEMENT.
11-10-2021 | Malfunction | covidien lp llc north haven | 28-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE HANDLE IS DEAD AND WOULD NOT CHARGE. IT WAS NOTED THAT BOTH DEVICES HAVE BEEN STANDING UNUSED FOR A YEAR AND A HALF DURING THE PANDEMIC. THE CHARGER WAS RUSTY. THE ISSUE WAS NOT RESOLVED. THERE WAS NO PATIENT INVOLVEMENT.
14-11-2021 | Malfunction | ethicon endo-surgery | 28-01-2022 | GST BLUE RELOAD, 45MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE ENDOSCOPIC WEDGE RESECTION OF LUNG LOBES SURGERY, COULD NOT FIRE WHEN SEVERING LUNG TISSUE ON THE 1TH FIRING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2022 | Malfunction | ethicon endo-surgery | 28-01-2022 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTROINTESTINAL BYPASS SURGERY, IN THE PATIENT, THE CIRCULAR STAPLER FAILED AT THE TIME OF THE ESOPHAGOJEJUNAL ANASTOMOSIS, SINCE WHEN CLOSING THE STAPLER THERE WAS A CUT, BUT NOT A CORRECT CLOSURE OF THE STAPLE LINE, THEREFORE THAT DID NOT FULFILL ITS FUNCTION OF ANASTOMOSING. WE REVIEWED THE TECHNICAL STEPS AND THE TRUTH IS WE CANNOT FIND HOW IT COULD HAVE FAILED.
01-11-2021 | Malfunction | ethicon endo-surgery | 28-01-2022 | ILS 33MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY, MISS FIRING OF A CIRCULAR STAPLER IN A CANCER COLON CASE DURING THE OR TIME. THEY STARTED THE PROCESS OF THE ANASTOMOSIS AGAIN WITH A NEW CIRCULAR STAPLER. THE SURGERY WAS DELAYED BY 20-MINUTES. FRAGMENTS WERE GENERATED BUT REMOVED WITHOUT ADDITIONAL INTERVENTION. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
01-11-2021 | Malfunction | ethicon endo-surgery | 28-01-2022 | ILS 33MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY, MISS FIRING OF A CIRCULAR STAPLER IN A CANCER COLON CASE DURING THE OR TIME. THEY STARTED THE PROCESS OF THE ANASTOMOSIS AGAIN WITH A NEW CIRCULAR STAPLER. THE SURGERY WAS DELAYED BY 20-MINUTES. FRAGMENTS WERE GENERATED BUT REMOVED WITHOUT ADDITIONAL INTERVENTION. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
18-12-2021 | Injury | us surgical puerto rico | 28-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE POWER HANDLE AND FIVE STAPLE LOADS FIRED NORMALLY. THE PATIENT HAD BLOOD LOSS OF MORE THAN 500 CC AND SYNCOPAL EPISODE. BLOOD TRANSFUSION WAS DONE TO RESOLVE THE ISSUE.
18-12-2021 | Injury | us surgical puerto rico | 28-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Syncope/Fainting | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE POWER HANDLE AND FIVE STAPLE LOADS FIRED NORMALLY. THE PATIENT HAD BLOOD LOSS OF MORE THAN 500 CC AND SYNCOPAL EPISODE. BLOOD TRANSFUSION WAS DONE TO RESOLVE THE ISSUE.
18-12-2021 | Injury | us surgical puerto rico | 28-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, POST OPERATIVELY, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE POWER HANDLE AND FIVE STAPLE LOADS FIRED NORMALLY. THE PATIENT HAD BLOOD LOSS OF MORE THAN 500 CC AND SYNCOPAL EPISODE. BLOOD TRANSFUSION WAS DONE TO RESOLVE THE ISSUE.
01-12-2021 | Malfunction | ethicon endo-surgery | 28-01-2022 | ENDO ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE ONLY PERFORMED THE POSTERIOR SUTURE. THERE WERE NO PATIENT CONSEQUENCES OR CHANGE TO THE POST-OPERATIVE CARE OF THE PATIENT.
01-01-2021 | Injury | ethicon endo-surgery | 28-01-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT NINE WEEKS POST OP TO A SLEEVE GASTRECTOMY THE PATIENT HAD A LEAK. A STINT WAS PLACED AND THE PATIENT IS DOING FINE. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME.
31-12-2021 | Injury | intuitive surgical | 28-01-2022 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED BY A SURGEON THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT WITH A GREEN SUREFORM 60 RELOAD INSTALLED HAD MISFIRED. DURING THE CREATION OF THE GASTRIC POUCH INVOLVING THE THIRD CONSECUTIVE STAPLE FIRE WITH A BLUE SUREFORM60 RELOAD, A "TISSUE TOO THICK" MESSAGE WAS RECEIVED BY THE SURGEON PRIOR TO FIRING THE STAPLER. HE DOES NOT RECALL PAUSES DURING COMPRESSION; HOWEVER, THE SUREFORM60 STAPLER INSTRUMENT CLAMPED AND FIRED WITH NO ISSUES. ON THE POUCH SIDE OF THE STOMACH, THE STAPLE LINE WAS FORMED. THE SURGEON STATED THE STAPLE LINE FELL APART OR NEVER FORMED ON THE REMNANT SIDE OF THE STOMACH. MANY STAPLES HAD NO ¿B¿ SHAPED FORMATION, AND THERE WAS A BIG HOLE ON THE STAPLE LINE. THE SURGEON RESECTED AND SUTURED THE INCOMPLETE STAPLE LINE. THEY SWAPPED TO A SECOND SUREFORM60 STAPLER INSTRUMENT AND COMPLETED THE PROCEDURE. A LEAK TEST WAS PERFORMED AND FOUND TO BE NEGATIVE (THERE WAS NO LEAK DETECTED). THE SURGEON DID STATE THAT HE HAD A NEW NURSE IN THE OPERATING ROOM, AND HE DOES NOT KNOW IF THE NURSE HAD INSTALLED THE RELOADS RIGHT AND SWISHED THE INSTRUMENT IN BETWEEN FIRES. SECONDARY TO THE STAPLE LINE ISSUE, ON POSTOPERATIVE DAY # 1, THE PATIENT HAD AN UPPER GI X-RAY DONE. THE RADIOLOGIST HAD INDICATED THERE COULD BE A LEAK BUT THE LOCATION OF THE SUSPECTED LEAK WAS NOT PROVIDED. HENCE, THE PATIENT UNDERWENT A SECOND PROCEDURE. THE SURGEON PERFORMED FOUR LEAK TESTS DURING THE LAPAROSCOPIC PROCEDURE AND COULD NOT IDENTIFY A LEAK. THE PATIENT WAS DISCHARGED THE FOLLOWING DAY AND WAS RECOVERING WELL. THE SURGEON STATED THERE WERE PHOTOGRAPHIC IMAGES IN THE OPERATING ROOM REPORT; HOWEVER, ISI CANNOT GET A COPY OF THOSE FOR REVIEW.
31-12-2021 | Injury | intuitive surgical | 28-01-2022 | SUREFORM | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED BY A SURGEON THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT WITH A GREEN SUREFORM 60 RELOAD INSTALLED HAD MISFIRED. DURING THE CREATION OF THE GASTRIC POUCH INVOLVING THE THIRD CONSECUTIVE STAPLE FIRE WITH A BLUE SUREFORM60 RELOAD, A "TISSUE TOO THICK" MESSAGE WAS RECEIVED BY THE SURGEON PRIOR TO FIRING THE STAPLER. HE DOES NOT RECALL PAUSES DURING COMPRESSION; HOWEVER, THE SUREFORM60 STAPLER INSTRUMENT CLAMPED AND FIRED WITH NO ISSUES. ON THE POUCH SIDE OF THE STOMACH, THE STAPLE LINE WAS FORMED. THE SURGEON STATED THE STAPLE LINE FELL APART OR NEVER FORMED ON THE REMNANT SIDE OF THE STOMACH. MANY STAPLES HAD NO ¿B¿ SHAPED FORMATION, AND THERE WAS A BIG HOLE ON THE STAPLE LINE. THE SURGEON RESECTED AND SUTURED THE INCOMPLETE STAPLE LINE. THEY SWAPPED TO A SECOND SUREFORM60 STAPLER INSTRUMENT AND COMPLETED THE PROCEDURE. A LEAK TEST WAS PERFORMED AND FOUND TO BE NEGATIVE (THERE WAS NO LEAK DETECTED). THE SURGEON DID STATE THAT HE HAD A NEW NURSE IN THE OPERATING ROOM, AND HE DOES NOT KNOW IF THE NURSE HAD INSTALLED THE RELOADS RIGHT AND SWISHED THE INSTRUMENT IN BETWEEN FIRES. SECONDARY TO THE STAPLE LINE ISSUE, ON POSTOPERATIVE DAY # 1, THE PATIENT HAD AN UPPER GI X-RAY DONE. THE RADIOLOGIST HAD INDICATED THERE COULD BE A LEAK BUT THE LOCATION OF THE SUSPECTED LEAK WAS NOT PROVIDED. HENCE, THE PATIENT UNDERWENT A SECOND PROCEDURE. THE SURGEON PERFORMED FOUR LEAK TESTS DURING THE LAPAROSCOPIC PROCEDURE AND COULD NOT IDENTIFY A LEAK. THE PATIENT WAS DISCHARGED THE FOLLOWING DAY AND WAS RECOVERING WELL. THE SURGEON STATED THERE WERE PHOTOGRAPHIC IMAGES IN THE OPERATING ROOM REPORT; HOWEVER, ISI CANNOT GET A COPY OF THOSE FOR REVIEW.
21-10-2021 | Injury | covidien lp llc north haven | 28-01-2022 | UNKNOWN EGIA TRI STAPLE | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER MISFIRED. IT DID NOT STAPLE BUT DID CUT. THIS RESULTED IN AN OPEN BOWEL AND THE SURGERY HAD TO BE CONVERTED TO AN OPEN PROCEDURE.
12-01-2022 | Malfunction | covidien lp llc north haven | 28-01-2022 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC DUODENAL SWITCH, THE RELOAD STOPPED MIDFIRE WHILE BEING APPLIED AT THE STOMACH TISSUE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-11-2021 | Malfunction | ethicon endo-surgery | 28-01-2022 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RADICAL DISTAL GASTRECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-01-2022 | Malfunction | covidien lp llc north haven | 28-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE INITIAL STOMACH STAPLING ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS STAPLE LINE INCOMPLETE DISTALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | us surgical puerto rico | 28-01-2022 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THORACOSCOPY PROCEDURE FOR PNEUMOTHORAX, WHILE ON LOBECTOMY, PULLING OF THE BLACK RETURN KNOB WAS QUITE STIFF AFTER THE SECOND FIRING USING THE DEVICE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2021 | Malfunction | covidien lp llc north haven | 28-01-2022 | UNKNOWN ENDO GIA SULU | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THORACOSCOPY PROCEDURE FOR PNEUMOTHORAX, WHILE ON LOBECTOMY, PULLING OF THE BLACK RETURN KNOB WAS Q UITE STIFF AFTER THE SECOND FIRING USING THE DEVICE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | davis & geck caribe | 28-01-2022 | UNKNOWN ENDO STITCH SULU | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE SUTURE NEEDLE BROKE OFF. IT WAS REPORTED ALSO THAT THE DEVICE WAS DIFFICULT TO TOGGLE AND LOAD. TO RESOLVE THE ISSUE, THE SURGEON OPENED A NEW DEVICE TO COMPLETE THE CASE AND THE REMAINING NEEDLE PIECE WAS RETRIEVED. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | davis & geck caribe | 28-01-2022 | UNKNOWN ENDO STITCH SULU | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE SUTURE NEEDLE BROKE OFF. IT WAS REPORTED ALSO THAT THE DEVICE WAS DIFFICULT TO TOGGLE AND LOAD. TO RESOLVE THE ISSUE, THE SURGEON OPENED A NEW DEVICE TO COMPLETE THE CASE AND THE REMAINING NEEDLE PIECE WAS RETRIEVED. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | davis & geck caribe | 28-01-2022 | UNKNOWN ENDO STITCH SULU | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE SUTURE NEEDLE BROKE OFF. IT WAS REPORTED ALSO THAT THE DEVICE WAS DIFFICULT TO TOGGLE AND LOAD. TO RESOLVE THE ISSUE, THE SURGEON OPENED A NEW DEVICE TO COMPLETE THE CASE AND THE REMAINING NEEDLE PIECE WAS RETRIEVED. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | davis & geck caribe | 28-01-2022 | UNKNOWN ENDO STITCH SULU | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE SUTURE NEEDLE BROKE OFF. IT WAS REPORTED ALSO THAT THE DEVICE WAS DIFFICULT TO TOGGLE AND LOAD. TO RESOLVE THE ISSUE, THE SURGEON OPENED A NEW DEVICE TO COMPLETE THE CASE AND THE REMAINING NEEDLE PIECE WAS RETRIEVED. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | covidien lp llc north haven | 28-01-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION, WHILE BEING APPLIED AT THE RECTUM, THE DEVICE DID NOT FIRE. AFTER THE JAWS WERE CLOSED, THE SAFETY WAS RELEASED AND THE FIRING WAS ATTEMPTED, BUT IT COULD NOT BE SQUEEZED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
12-01-2022 | Malfunction | covidien lp llc north haven | 28-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOW ANTERIOR RESECTION, WHILE BEING APPLIED AT THE RECTUM, THE DEVICE DID NOT FIRE. AFTER THE JAWS WERE CLOSED, THE SAFETY WAS RELEASED AND THE FIRING WAS ATTEMPTED, BUT IT COULD NOT BE SQUEEZED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
13-01-2022 | Malfunction | us surgical puerto rico | 28-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE TRANSECTING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AFTER FIRING, THE HANDLE CRACKED AND THE ARTICULATION LEVER BROKE INTO TWO AND FELL TO THE FLOOR. THE STAPLE LINE WAS ALSO CENTRALLY INCOMPLETE. TO RESOLVE THE ISSUE, THE TISSUE REINFORCEMENT SYSTEM (TRS) WAS CUT WITH SCISSORS, A NEW RELOAD AND NEW HANDLE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
13-01-2022 | Malfunction | covidien lp llc north haven | 28-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE TRANSECTING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AFTER FIRING, THE HANDLE CRACKED AND THE ARTICULATION LEVER BROKE INTO TWO AND FELL TO THE FLOOR. THE STAPLE LINE WAS ALSO CENTRALLY INCOMPLETE. TO RESOLVE THE ISSUE, THE TISSUE REINFORCEMENT SYSTEM (TRS) WAS CUT WITH SCISSORS, A NEW RELOAD AND NEW HANDLE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Injury | covidien lp llc north haven | 28-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY WHILE REMOVING THE LUNG, THE STAPLE BURST OUT AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. MANUAL SUTURING WAS DONE BY THE SURGEON TO RESOLVE THE ISSUE.
16-12-2021 | Injury | us surgical puerto rico | 28-01-2022 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMICOLECTOMY AND END-TO-END COLORECTAL ANASTOMOSIS PROCEDURE, WHILE STAPLING THE COLORECTAL ANASTOMOSIS, THE DEVICE HAD A POOR STAPLE FORMATION AND THE ANVIL WAS BENT. THERE WAS AN UNANTICIPATED TISSUE LOSS. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED FOR MORE THAN 1 INCH. THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE DUE TO DEVICE PROBLEM. THE STAPLE LINE HAD TO BE RESECTED AND RE-STAPLED. THE MOST DISTAL ANASTOMOSIS WAS PERFORMED AND ANOTHER CIRCULAR STAPLER WAS USED TO COMPLETE THE CASE.
16-12-2021 | Injury | us surgical puerto rico | 28-01-2022 | EEA | Material Twisted/Bent | A040609 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMICOLECTOMY AND END-TO-END COLORECTAL ANASTOMOSIS PROCEDURE, WHILE STAPLING THE COLORECTAL ANASTOMOSIS, THE DEVICE HAD A POOR STAPLE FORMATION AND THE ANVIL WAS BENT. THERE WAS AN UNANTICIPATED TISSUE LOSS. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AND THE INCISION WAS EXTENDED FOR MORE THAN 1 INCH. THE LAPAROSCOPIC PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE DUE TO DEVICE PROBLEM. THE STAPLE LINE HAD TO BE RESECTED AND RE-STAPLED. THE MOST DISTAL ANASTOMOSIS WAS PERFORMED AND ANOTHER CIRCULAR STAPLER WAS USED TO COMPLETE THE CASE.
12-01-2022 | Malfunction | us surgical puerto rico | 28-01-2022 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE TRANSECTION OF THE STOMACH IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN THE DEVICE WAS FIRED, THE HANDLE MADE A LOUD CRACK AND THE ARTICULATION KNOB BROKE IN TWO PIECES AND FELL ON THE FLOOR. THE STAPLE LINE ON THE CENTRAL AREA WAS INCOMPLETE. THE SURGEON USED ANOTHER RELOAD AND HANDLE TO COMPLETE THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | covidien lp llc north haven | 28-01-2022 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE TRANSECTION OF THE STOMACH IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN THE DEVICE WAS FIRED, THE HANDLE MADE A LOUD CRACK AND THE ARTICULATION KNOB BROKE IN TWO PIECES AND FELL ON THE FLOOR. THE STAPLE LINE ON THE CENTRAL AREA WAS INCOMPLETE. THE SURGEON USED ANOTHER RELOAD AND HANDLE TO COMPLETE THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
23-11-2021 | Malfunction | us surgical puerto rico | 31-01-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN LOWER ANTERIOR RESECTION, THE DEVICE FIRED AND THE SURGEON WAS ABLE TO PERFORM FULL AND COMPLETE SQUEEZES OF THE HANDLE, BUT THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER DEVICE WAS USED TO CONTINUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED FIRING RACK TEETH DAMAGE.
23-11-2021 | Malfunction | us surgical puerto rico | 31-01-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN LOWER ANTERIOR RESECTION, THE DEVICE FIRED AND THE SURGEON WAS ABLE TO PERFORM FULL AND COMPLETE SQUEEZES OF THE HANDLE, BUT THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER DEVICE WAS USED TO CONTINUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED FIRING RACK TEETH DAMAGE.
14-01-2022 | Injury | celestica de monterrey s.a de c.v | 31-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING CUTTING OF THE STOMACH IN A LAPAROSCOPIC BARIATRIC SURGERY, THE STITCHES AT THE DISTAL END OF STAPLE LINE IS OOZING OR BLEEDING FOR MORE THAN 500 ML. THE RESOLVE THE ISSUE, A BLOOD TRANSFUSION AND REOPERATION WAS DONE. THE SURGICAL PROCEDURE WAS ALSO EXTENDED FOR MORE THAN THIRTY MINUTES. THE PATIENT HAD TO EXTEND HOSPITALIZATION FOR TWO MORE DAYS.
14-01-2022 | Injury | celestica de monterrey s.a de c.v | 31-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING CUTTING OF THE STOMACH IN A LAPAROSCOPIC BARIATRIC SURGERY, THE STITCHES AT THE DISTAL END OF STAPLE LINE IS OOZING OR BLEEDING FOR MORE THAN 500 ML. THE RESOLVE THE ISSUE, A BLOOD TRANSFUSION AND REOPERATION WAS DONE. THE SURGICAL PROCEDURE WAS ALSO EXTENDED FOR MORE THAN THIRTY MINUTES. THE PATIENT HAD TO EXTEND HOSPITALIZATION FOR TWO MORE DAYS.
14-01-2022 | Injury | celestica de monterrey s.a de c.v | 31-01-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING CUTTING OF THE STOMACH IN A LAPAROSCOPIC BARIATRIC SURGERY, THE STITCHES AT THE DISTAL END OF STAPLE LINE IS OOZING OR BLEEDING FOR MORE THAN 500 ML. THE RESOLVE THE ISSUE, A BLOOD TRANSFUSION AND REOPERATION WAS DONE. THE SURGICAL PROCEDURE WAS ALSO EXTENDED FOR MORE THAN THIRTY MINUTES. THE PATIENT HAD TO EXTEND HOSPITALIZATION FOR TWO MORE DAYS.
12-01-2022 | Malfunction | celestica de monterrey s.a de c.v | 31-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CUTTING OF THE TISSUE IN A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON. BUT THE HANDLE CANNOT BE SQUEEZED. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
01-01-2022 | Injury | ethicon endo-surgery | 31-01-2022 | PROXIMATE*RL LINEAR STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THERE WAS POST-OP BLEEDING ON THE STAPLE LINE. A SECOND OPERATION WAS NEEDED TO OVER-SEW THE STAPLE LINE.
12-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ESOPHAGECTOMY, THE SURGEON WOULD LIKE TO PUT THE DEVICE INTO THE PATIENT'S ESOPHAGEAL THEN TUBE WAS UNLOCK AND FALL OFF. THE ANVIL COULD NOT BE PUT INTO PATIENT'S BODY SUCCESSFULLY. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
29-11-2021 | Malfunction | ethicon endo-surgery | 31-01-2022 | 55MM SELECTABLE NEW TLC | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE BROKE.
10-01-2022 | Malfunction | us surgical puerto rico | 31-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN TRYING TO DIVIDE THE LOBE IN A LAPAROSCOPIC SEGMENTECTOMY, AFTER GRASPING THE TISSUE AND WAS READY TO FIRE, THE DEVICE HANDLE COULD NOT BE SQUEEZED. ANOTHER HANDLE WAS USED BUT THE SAME ISSUE OCCURRED. TO RESOLVE THE ISSUE, A THIRD DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | us surgical puerto rico | 31-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN TRYING TO DIVIDE THE LOBE IN A LAPAROSCOPIC SEGMENTECTOMY, AFTER GRASPING THE TISSUE AND WAS READY TO FIRE, THE DEVICE HANDLE COULD NOT BE SQUEEZED. ANOTHER HANDLE WAS USED BUT THE SAME ISSUE OCCURRED. TO RESOLVE THE ISSUE, A THIRD DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
13-01-2022 | Injury | covidien lp llc north haven | 31-01-2022 | EEA ORVIL | Mechanical Jam | A0506 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, WHEN TRYING TO POSITION THE ANVIL IN AN UMBRELLA POSITION TO PERFORM THE MECHANICAL OESO-JEJUNAL SUTURE, THE ORVIL DID NOT DEPLOY CORRECTLY ONCE INTRODUCED INTO THE ESOPHAGUS. THE SURGEON PERFORMED LAPAROTOMY TO CONTINUE THE PROCEDURE.
10-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHEN TRYING TO DIVIDE THE LOBE IN A LAPAROSCOPIC SEGMENTECTOMY, AFTER GRASPING THE TISSUE AND WAS READY TO FIRE, THE DEVICE HANDLE COULD NOT BE SQUEEZED. ANOTHER HANDLE WAS USED BUT THE SAME ISSUE OCCURRED. TO RESOLVE THE ISSUE, A THIRD DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
03-01-2022 | Injury | ethicon endo-surgery | 31-01-2022 | PROXIMATE*RL LINEAR STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED POST-OP TO AN UNKNOWN PROCEDURE, THE PATIENT HAD BLEEDING ON THE STAPLE LINE. A SECOND OPERATION WAS NEEDED TO OVER-SEW THE STAPLE LINE.
13-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED LAPAROSCOPIC RECTAL RESECTION, WHEN SEPARATING THE RECTUM, THE KNIFE STOPPED HALFWAY OF THE FIRING. THE USER TOOK OUT A NEW DEVICE AND RESECT IT AGAIN FROM WHERE IT STOPPED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED LAPAROSCOPIC RECTAL RESECTION, WHEN SEPARATING THE RECTUM, THE KNIFE STOPPED HALFWAY OF THE FIRING. THE USER TOOK OUT A NEW DEVICE AND RESECT IT AGAIN FROM WHERE IT STOPPED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2021 | Injury | ethicon endo-surgery | 31-01-2022 | 55MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE DEVICE WAS USED IN THE OPERATION. THE PATIENT HAD ABDOMINAL PAIN AND ABDOMINAL DISTENSION ON POSTOPERATIVE DAY ONE, AND DARK RED FLUID WAS OBSERVED IN ABDOMINAL DRAINAGE TUBE. ABDOMINAL CT SHOWED ABDOMINAL EFFUSION, CONSIDERING ANASTOMOTIC LEAKAGE. ONE DAY AFTER THE OPERATION, THE PATIENT'S ABDOMEN WAS EXPLORED. IT WAS FOUND THAT THE STAPLE AT THE DISTAL 1/2 OF THE INTESTINAL ANASTOMOSIS FELL OFF, BLOOD CLOTS WERE OBSERVED NEARBY, AND A TOTAL OF ABOUT 600 ML OF BILE AND BLOOD MIXTURE WAS OBSERVED IN THE ABDOMINAL CAVITY. THE SURGERY WAS PERFORMED TO FIX IT. PROCEDURE: GASTRIC RESECTION. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2021 | Injury | ethicon endo-surgery | 31-01-2022 | 55MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: IT WAS REPORTED THAT THE DEVICE WAS USED IN THE OPERATION. THE PATIENT HAD ABDOMINAL PAIN AND ABDOMINAL DISTENSION ON POSTOPERATIVE DAY ONE, AND DARK RED FLUID WAS OBSERVED IN ABDOMINAL DRAINAGE TUBE. ABDOMINAL CT SHOWED ABDOMINAL EFFUSION, CONSIDERING ANASTOMOTIC LEAKAGE. ONE DAY AFTER THE OPERATION, THE PATIENT'S ABDOMEN WAS EXPLORED. IT WAS FOUND THAT THE STAPLE AT THE DISTAL 1/2 OF THE INTESTINAL ANASTOMOSIS FELL OFF, BLOOD CLOTS WERE OBSERVED NEARBY, AND A TOTAL OF ABOUT 600 ML OF BILE AND BLOOD MIXTURE WAS OBSERVED IN THE ABDOMINAL CAVITY. THE SURGERY WAS PERFORMED TO FIX IT. PROCEDURE: GASTRIC RESECTION. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2021 | Injury | ethicon endo-surgery | 31-01-2022 | 55MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: IT WAS REPORTED THAT THE DEVICE WAS USED IN THE OPERATION. THE PATIENT HAD ABDOMINAL PAIN AND ABDOMINAL DISTENSION ON POSTOPERATIVE DAY ONE, AND DARK RED FLUID WAS OBSERVED IN ABDOMINAL DRAINAGE TUBE. ABDOMINAL CT SHOWED ABDOMINAL EFFUSION, CONSIDERING ANASTOMOTIC LEAKAGE. ONE DAY AFTER THE OPERATION, THE PATIENT'S ABDOMEN WAS EXPLORED. IT WAS FOUND THAT THE STAPLE AT THE DISTAL 1/2 OF THE INTESTINAL ANASTOMOSIS FELL OFF, BLOOD CLOTS WERE OBSERVED NEARBY, AND A TOTAL OF ABOUT 600 ML OF BILE AND BLOOD MIXTURE WAS OBSERVED IN THE ABDOMINAL CAVITY. THE SURGERY WAS PERFORMED TO FIX IT. PROCEDURE: GASTRIC RESECTION. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2021 | Injury | ethicon endo-surgery | 31-01-2022 | 55MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT THE DEVICE WAS USED IN THE OPERATION. THE PATIENT HAD ABDOMINAL PAIN AND ABDOMINAL DISTENSION ON POSTOPERATIVE DAY ONE, AND DARK RED FLUID WAS OBSERVED IN ABDOMINAL DRAINAGE TUBE. ABDOMINAL CT SHOWED ABDOMINAL EFFUSION, CONSIDERING ANASTOMOTIC LEAKAGE. ONE DAY AFTER THE OPERATION, THE PATIENT'S ABDOMEN WAS EXPLORED. IT WAS FOUND THAT THE STAPLE AT THE DISTAL 1/2 OF THE INTESTINAL ANASTOMOSIS FELL OFF, BLOOD CLOTS WERE OBSERVED NEARBY, AND A TOTAL OF ABOUT 600 ML OF BILE AND BLOOD MIXTURE WAS OBSERVED IN THE ABDOMINAL CAVITY. THE SURGERY WAS PERFORMED TO FIX IT. PROCEDURE: GASTRIC RESECTION. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2021 | Injury | ethicon endo-surgery | 31-01-2022 | 55MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Distention | Event Description: IT WAS REPORTED THAT THE DEVICE WAS USED IN THE OPERATION. THE PATIENT HAD ABDOMINAL PAIN AND ABDOMINAL DISTENSION ON POSTOPERATIVE DAY ONE, AND DARK RED FLUID WAS OBSERVED IN ABDOMINAL DRAINAGE TUBE. ABDOMINAL CT SHOWED ABDOMINAL EFFUSION, CONSIDERING ANASTOMOTIC LEAKAGE. ONE DAY AFTER THE OPERATION, THE PATIENT'S ABDOMEN WAS EXPLORED. IT WAS FOUND THAT THE STAPLE AT THE DISTAL 1/2 OF THE INTESTINAL ANASTOMOSIS FELL OFF, BLOOD CLOTS WERE OBSERVED NEARBY, AND A TOTAL OF ABOUT 600 ML OF BILE AND BLOOD MIXTURE WAS OBSERVED IN THE ABDOMINAL CAVITY. THE SURGERY WAS PERFORMED TO FIX IT. PROCEDURE: GASTRIC RESECTION. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED LAPAROSCOPIC RECTAL RESECTION, WHEN SEPARATING THE RECTUM, THE KNIFE STOPPED HALFWAY OF THE FIRING. THE USER TOOK OUT A NEW DEVICE AND RESECT IT AGAIN FROM WHERE IT STOPPED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED LAPAROSCOPIC RECTAL RESECTION, WHEN SEPARATING THE RECTUM, THE KNIFE STOPPED HALFWAY OF THE FIRING. THE USER TOOK OUT A NEW DEVICE AND RESECT IT AGAIN FROM WHERE IT STOPPED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-10-2021 | Malfunction | ethicon endo-surgery | 31-01-2022 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE FIRING KNOB WAS SO STIFF THAT THE KNIFE DID NOT MOVE FORWARD AT THE 1ST FIRING. ANOTHER ECHELON WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
27-10-2021 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO SURGERY, THE HANDLE DID NOT WORK. THE DEVICE WAS BLACK OUT AND COULD NOT BE USED FROM THE BEGINNING. ANOTHER HANDLE WAS USED TO PROCEED WITH THE CASE. THERE WAS NO PATIENT INVOLVED.
27-10-2021 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO SURGERY, THE HANDLE DID NOT WORK. THE DEVICE WAS BLACK OUT AND COULD NOT BE USED FROM THE BEGINNING. ANOTHER HANDLE WAS USED TO PROCEED WITH THE CASE. THERE WAS NO PATIENT INVOLVED.
30-11-2021 | Malfunction | ethicon endo-surgery | 31-01-2022 | 75MM SELECTABLE NEW TLC | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PANCREATICODUODENECTOMY, A PART OF THE DEVICE CAME OFF. NO PIECES WERE LEFT INSIDE THE PATIENT. THE FIRING WAS COMPLETED PROPERLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2021 | Malfunction | ethicon endo-surgery | 31-01-2022 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LEFT COLON PROCEDURE, SURGEON SAYS CONTOUR WITH BLUE LOAD WOULD NOT LOCK CLOSED OR ALLOW TO BE FIRED. THERE WERE NO PATIENT CONSEQUENCES AND PROCEDURE WAS COMPLETE BY USING ANOTHER STAPLER.
01-01-2021 | Malfunction | ethicon endo-surgery | 31-01-2022 | CONTOUR CURVED CUTTER STAPLER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LEFT COLON PROCEDURE, SURGEON SAYS CONTOUR WITH BLUE LOAD WOULD NOT LOCK CLOSED OR ALLOW TO BE FIRED. THERE WERE NO PATIENT CONSEQUENCES AND PROCEDURE WAS COMPLETE BY USING ANOTHER STAPLER.
06-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED. PROCEDURE: RECTUM PROCEDURE.
08-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLE MALFORMATION. THERE WAS NO DELAY TO THE SURGERY. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
06-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY, STAPLE MALFORMATION. THERE WAS NO DELAY TO THE SURGERY. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
10-01-2022 | Malfunction | arthrex | 31-01-2022 | TISSUETAK TENDON ANCHOR | Mechanical Problem | A05 | Insufficient Information | Event Description: ON (B)(6) 2022 IT WAS REPORTED BY A SALES REPRESENTATIVE VIA THE PHONE THAT AN AR-19021TT TENDON ANCHOR, FIRED AND LOCKED WITH THE PRONGS CROSS OVER THE TOP. THIS WAS DISCOVERED DURING USE IN A CUFFMEND PROCEDURE ON (B)(6) 2022. THE SURGEON OPENED A 2ND AR-19021TT AND 4 ANCHORS WERE INSERTED BEFORE THE DEVICE LOCKED AND WOULD NOT FIRE. THE CASE WAS COMPLETED BY OPENING A 3RD AR-19021TT SAME LOT NUMBER AND INSERTING 3 ADDITIONAL ANCHORS.
10-01-2022 | Malfunction | arthrex | 31-01-2022 | TISSUETAK TENDON ANCHOR | Failure to Fire | A050501 | Insufficient Information | Event Description: ON (B)(6) 2022 IT WAS REPORTED BY A SALES REPRESENTATIVE VIA THE PHONE THAT AN AR-19021TT TENDON ANCHOR, FIRED AND LOCKED WITH THE PRONGS CROSS OVER THE TOP. THIS WAS DISCOVERED DURING USE IN A CUFFMEND PROCEDURE ON (B)(6) 2022. THE SURGEON OPENED A 2ND AR-19021TT AND 4 ANCHORS WERE INSERTED BEFORE THE DEVICE LOCKED AND WOULD NOT FIRE. THE CASE WAS COMPLETED BY OPENING A 3RD AR-19021TT SAME LOT NUMBER AND INSERTING 3 ADDITIONAL ANCHORS.
10-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS POSSIBLE TO CLOSE THE INSTRUMENT BUT NO LONGER OPEN FOR READJUSTMENT. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED.
14-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE GOT BLOCKED WITH JAWS CLOSED ON INSERTION INTO THE TROCAR. THE SURGEON WAS UNABLE TO OPEN THE JAWS TO PERFORM THE STAPLING. NO CONSEQUENCES. USE OF ANOTHER DEVICE TO REDUCE THE RISK OF THE JAWS LOCKING IN THE CLOSED POSITION ON THE TISSUE. UNLOCKED THE JAWS TO PREVENT TEARING.
07-10-2021 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER SEGMENTAL RESECTION, WHEN THE USER APPROACHED THE GLISSON, THE GREEN BUTTON WAS NOT PRESSED, BUT IT WAS FOUND TO BE INDICATED THAT THE CARTRIDGE HAD BEEN PRE FIRED AND WAS UNABLE TO FIRE. THE CARTRIDGE PART OF THE DISPLAY WAS 0. THE OPERATOR REPLACED WITH A NEW CARTRIDGE, SHELL AND ADAPTER AND THE SAME HANDLE WERE USED AND FIRING WAS PERFORMED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-10-2021 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER SEGMENTAL RESECTION, WHEN THE USER APPROACHED THE GLISSON, THE GREEN BUTTON WAS NOT PRESSED, BUT IT WAS FOUND TO BE INDICATED THAT THE CARTRIDGE HAD BEEN PRE FIRED AND WAS UNABLE TO FIRE. THE CARTRIDGE PART OF THE DISPLAY WAS 0. THE OPERATOR REPLACED WITH A NEW CARTRIDGE, SHELL AND ADAPTER AND THE SAME HANDLE WERE USED AND FIRING WAS PERFORMED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIGESTIVE PROCEDURE, ON ANASTOMOSIS, THERE POOR STAPLE FORMATION AND AN INCOMPLETE STAPLE LINE. ONLY ONE LINE WAS COMPLETED. THE SURGEON USED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIGESTIVE PROCEDURE, ON ANASTOMOSIS, THERE POOR STAPLE FORMATION AND AN INCOMPLETE STAPLE LINE. ONLY ONE LINE WAS COMPLETED. THE SURGEON USED A NEW DEVICE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, THE SURGEON HEARD A STRANGE NOISE AND THEN THE WHOLE LOADING UNIT DID NOT FIRE, THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. ONLY ONE ROW WAS DEPLOYED ON THE VESSEL SIDE AND THREE STAPLE LINES ON THE OTHER SIDE. DUE TO NO BLEEDING, THE SURGEON DECIDED NOT TO INTERVENE AND JUST KEPT THE ONE LINE OF VESSEL CLAMPING. THE RELOAD WERE REPLACED AND THE HANDLE WORKED WELL. THERE WAS NO PATIENT INJURY.
11-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | ENDO GIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, THE SURGEON HEARD A STRANGE NOISE AND THEN THE WHOLE LOADING UNIT DID NOT FIRE, THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. ONLY ONE ROW WAS DEPLOYED ON THE VESSEL SIDE AND THREE STAPLE LINES ON THE OTHER SIDE. DUE TO NO BLEEDING, THE SURGEON DECIDED NOT TO INTERVENE AND JUST KEPT THE ONE LINE OF VESSEL CLAMPING. THE RELOAD WERE REPLACED AND THE HANDLE WORKED WELL. THERE WAS NO PATIENT INJURY.
19-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WAS A MALFUNCTION IN FIRING STAPLES. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
14-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PANCREATIC PARENCHYMAL RESECTION IN AN OPEN ABDOMINAL DISTAL PANCREATECTOMY, THE DEVICE WAS TRIED TO FIRE USING A POWERED HANDLE BUT IT DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
15-07-2021 | Malfunction | covidien lp llc north haven | 31-01-2022 | EEA ORVIL | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, TWO ANVILS WAS NOT FEASIBLE AT THE TIME OF ATTACHING TO THE STAPLER SINCE IT WAS RELEASE WITH AND COULD NOT ATTACH IN THE END. THE ANVILS WERE UNABLE TO CLOSE. THE DEVICES WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PANCREATIC PARENCHYMAL RESECTION IN AN OPEN ABDOMINAL DISTAL PANCREATECTOMY, THE DEVICE WAS TRIED TO FIRE USING A POWERED HANDLE BUT IT DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
14-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PANCREATIC PARENCHYMAL RESECTION IN AN OPEN ABDOMINAL DISTAL PANCREATECTOMY, THE DEVICE WAS TRIED TO FIRE USING A POWERED HANDLE BUT IT DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
22-07-2021 | Malfunction | covidien lp llc north haven | 31-01-2022 | UNKNOWN ENDO GIA SULU | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING ANASTOMOSIS DURING THE SUBTOTAL LAPAROSCOPY GASTRECTOMY, AT THE MOMENT OF FIRING, IT DOES NOT ALLOW THE PASSAGE OF THE RELOADING BLADE. THERE WAS ALSO PROBLEM LOADING AND UNLOADING THE DEVICE. THE PASSAGE OF THE STAPLES IS VISIBLE BUT THE BLADE DOES NOT ACTIVATE. TO RESOLVE THE ISSUE, THE USER PERFORMED THE STEP OF ANOTHER HANDLE. THERE WAS NO PATIENT INJURY.
22-07-2021 | Malfunction | covidien lp llc north haven | 31-01-2022 | UNKNOWN ENDO GIA SULU | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING ANASTOMOSIS DURING THE SUBTOTAL LAPAROSCOPY GASTRECTOMY, AT THE MOMENT OF FIRING, IT DOES NOT ALLOW THE PASSAGE OF THE RELOADING BLADE. THERE WAS ALSO PROBLEM LOADING AND UNLOADING THE DEVICE. THE PASSAGE OF THE STAPLES IS VISIBLE BUT THE BLADE DOES NOT ACTIVATE. TO RESOLVE THE ISSUE, THE USER PERFORMED THE STEP OF ANOTHER HANDLE. THERE WAS NO PATIENT INJURY.
22-07-2021 | Malfunction | covidien lp llc north haven | 31-01-2022 | UNKNOWN ENDO GIA SULU | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE PERFORMING ANASTOMOSIS DURING THE SUBTOTAL LAPAROSCOPY GASTRECTOMY, AT THE MOMENT OF FIRING, IT DOES NOT ALLOW THE PASSAGE OF THE RELOADING BLADE. THERE WAS ALSO PROBLEM LOADING AND UNLOADING THE DEVICE. THE PASSAGE OF THE STAPLES IS VISIBLE BUT THE BLADE DOES NOT ACTIVATE. TO RESOLVE THE ISSUE, THE USER PERFORMED THE STEP OF ANOTHER HANDLE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE JAW COULDN'T BE OPENED WHEN USING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
13-01-2022 | Malfunction | us surgical puerto rico | 31-01-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, WHILE ON RECTUM RESECTION, THE FIRST HANDLE AND RELOAD, DID NOT FIRE. A CRACK SOUND WAS HEARD AND THE HANDLE WAS STIFF. A NEW PAIR OF HANDLE AND RELOAD WERE USED BUT THIS TIME THE JAWS COULD NOT BE CLOSED AND THE CRACK SOUND WAS HEARD AGAIN. A NEW INSTRUMENT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2022 | Malfunction | us surgical puerto rico | 31-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, WHILE ON RECTUM RESECTION, THE FIRST HANDLE AND RELOAD, DID NOT FIRE. A CRACK SOUND WAS HEARD AND THE HANDLE WAS STIFF. A NEW PAIR OF HANDLE AND RELOAD WERE USED BUT THIS TIME THE JAWS COULD NOT BE CLOSED AND THE CRACK SOUND WAS HEARD AGAIN. A NEW INSTRUMENT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | PROXIMATE*CUTTER RELOAD UNIT | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN COLECTOMY, IT WAS FOUND THAT THE STAPLES WERE PROTRUDED BEFORE USE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
13-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, WHILE ON RECTUM RESECTION, THE FIRST HANDLE AND RELOAD, DID NOT FIRE. A CRACK SOUND WAS HEARD AND THE HANDLE WAS STIFF. A NEW PAIR OF HANDLE AND RELOAD WERE USED BUT THIS TIME THE JAWS COULD NOT BE CLOSED AND THE CRACK SOUND WAS HEARD AGAIN. A NEW INSTRUMENT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE THE DOCTOR WENT TO FIRE THE POWERED CIRCULAR STAPLER AND IT WOULDN'T FIRE. THE BATTERY WAS FULLY INSERTED, THE ANVIL WAS FULLY CLICKED/SECURE, THE COMPRESSION SCALE LIT UP AND THE COMPRESSION WAS WITHIN THE PROPER FIRING AREA. WHEN THE DOCTOR RELEASED THE SAFETY AND TRIED TO FIRE, THE STAPLER DID NOT FIRE NOR DID IT MAKE A NOISE THAT IT WAS TRYING TO FIRE. DOCTOR OPENED A NEW STAPLER AND TRIED THAT BATTERY AND AGAIN, THE SCALE LIT UP BUT DID NOT FIRE. THE STAFF DID NOT ACCIDENTALLY FIRE WHEN SETTING UP THE STAPLER, NOR WAS THE CHECKMARK LIT UP PRIOR TO ATTEMPT TO FIRE. DOCTOR REMOVED THE STAPLER FROM THE PATIENT AND USED A NEW ONE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT AND THE SURGERY WAS COMPLETED SUCCESSFULLY.
02-11-2021 | Malfunction | us surgical puerto rico | 31-01-2022 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THE STAPLES DID NOT COME OUT AT THE TIME OF COUPLING WITH THE ANVIL. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-11-2021 | Malfunction | us surgical puerto rico | 31-01-2022 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THE STAPLES DID NOT COME OUT AT THE TIME OF COUPLING WITH THE ANVIL. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2022 | Malfunction | us surgical puerto rico | 31-01-2022 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, WHILE ON RECTUM RESECTION, THE FIRST HANDLE AND RELOAD, DID NOT FIRE. A CRACK SOUND WAS HEARD AND THE HANDLE WAS STIFF. A NEW PAIR OF HANDLE AND RELOAD WERE USED BUT THIS TIME THE JAWS COULD NOT BE CLOSED AND THE CRACK SOUND WAS HEARD AGAIN. A NEW INSTRUMENT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2022 | Malfunction | us surgical puerto rico | 31-01-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, WHILE ON RECTUM RESECTION, THE FIRST HANDLE AND RELOAD, DID NOT FIRE. A CRACK SOUND WAS HEARD AND THE HANDLE WAS STIFF. A NEW PAIR OF HANDLE AND RELOAD WERE USED BUT THIS TIME THE JAWS COULD NOT BE CLOSED AND THE CRACK SOUND WAS HEARD AGAIN. A NEW INSTRUMENT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, WHILE ON RECTUM RESECTION, THE FIRST HANDLE AND RELOAD, DID NOT FIRE. A CRACK SOUND WAS HEARD AND THE HANDLE WAS STIFF. A NEW PAIR OF HANDLE AND RELOAD WERE USED BUT THIS TIME THE JAWS COULD NOT BE CLOSED AND THE CRACK SOUND WAS HEARD AGAIN. A NEW INSTRUMENT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PULMONARY RESECTION, WHEN THE CARTRIDGE WAS ATTEMPTED TO BE USED ON THE VESSEL, WHILE THE USER LOADED THE CARTRIDGE AND CHECKED IT, A YELLOW ONE WAS FOUND ATTACHED, THERE WAS DAMAGE TO THE YELLOW TAB. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PANCREATIC PARENCHYMAL RESECTION IN AN OPEN ABDOMINAL DISTAL PANCREATECTOMY, THE DEVICE WAS TRIED TO FIRE USING A POWERED HANDLE BUT IT DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
14-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PANCREATIC PARENCHYMAL RESECTION IN AN OPEN ABDOMINAL DISTAL PANCREATECTOMY, THE DEVICE WAS TRIED TO FIRE USING A POWERED HANDLE BUT IT DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
13-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED LAPAROSCOPIC RECTAL RESECTION, WHEN SEPARATING THE RECTUM, THE KNIFE STOPPED HALFWAY OF THE FIRING. THE USER TOOK OUT A NEW DEVICE AND RESECT IT AGAIN FROM WHERE IT STOPPED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | CIRCULAR MECH XL SEAL 29MM | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING UNKNOWN PROCEDURE STAPLES WERE ¿PRE-DEPLOYED¿ WHEN OPENING FROM STERILE PACKAGING. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
19-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. IT STILL WOULD NOT RETURN AFTER USING MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. IT STILL WOULD NOT RETURN AFTER USING MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. IT STILL WOULD NOT RETURN AFTER USING MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON WOULD NOT WORK. IT STILL WOULD NOT RETURN AFTER USING MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A COLORECTAL PROCEDURE, AFTER SHOOTING ONE STAPLE LINE ON THE TISSUE WITHOUT PROBLEMS, BEFORE SHOOTING ANOTHER ONE, THE POWER HANDLE SUDDENLY TURNED OFF AND WOULD NOT TURN ON AGAIN. IT WAS NOTED THAT THE POWER HANDLE WOULD NOT CHARGE, THE CHARGER FAILED AND CAUSED BATTERY ISSUE ON THE HANDLE. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A COLORECTAL PROCEDURE, AFTER SHOOTING ONE STAPLE LINE ON THE TISSUE WITHOUT PROBLEMS, BEFORE SHOOTING ANOTHER ONE, THE POWER HANDLE SUDDENLY TURNED OFF AND WOULD NOT TURN ON AGAIN. IT WAS NOTED THAT THE POWER HANDLE WOULD NOT CHARGE, THE CHARGER FAILED AND CAUSED BATTERY ISSUE ON THE HANDLE. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY WHILE FIRING ACROSS THE CECUM MIDWAY THROUGH THE TISSUE BEGAN SLIDING OUT OF THE JAWS OF THE DEVICE CAUSING THE STAPLES TO MALFORM. THE SURGEON OVERSEWED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
19-01-2022 | Malfunction | ethicon endo-surgery | 31-01-2022 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY WHILE FIRING ACROSS THE CECUM MIDWAY THROUGH THE TISSUE BEGAN SLIDING OUT OF THE JAWS OF THE DEVICE CAUSING THE STAPLES TO MALFORM. THE SURGEON OVERSEWED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
05-01-2022 | Malfunction | covidien lp llc north haven | 31-01-2022 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE YELLOW TAB WAS ALREADY DAMAGED. THERE WAS NO PATIENT INVOLVEMENT.
14-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PARTIAL THORACOSCOPIC PNEUMONECTOMY, AFTER UNPACKING AND LOADING THE HANDLE WITH THE RELOAD, IT WAS FOUND OUT THAT THE RELOAD COULD NOT BE CLOSED TIGHTLY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2022 | Malfunction | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BAND, THE RELOAD'S JAW GOT DAMAGED AND ALMOST BROKE OFF. THE SURGEON PULLED OUT THE TROCAR TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE JAWS OF THE RELOAD WILL NOT CLOSE WHILE LOADING. THERE WAS NO PATIENT INVOLVED.
15-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE JAWS OF THE RELOAD WILL NOT CLOSE WHILE LOADING. THERE WAS NO PATIENT INVOLVED.
18-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE PRODUCED HALF OF THE STAPLE LINE AND THEN JAMMED.
18-01-2022 | Injury | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE GOLD RELOAD ECR60D DID NOT GRAB THE TISSUE. KNIFE PUSHED TISSUE FORWARD AND BACK AGAIN. STAPLES DID NOT GO INTO TISSUE. TISSUE ON THE STOMACH WAS TORN AND HAD TO BE SUTURED OVER. SURGEON CONTINUED WITH NEW GOLD RELOAD GST60D WITH SAME RESULT. NEW STAPLER WAS OPENED. WHITE AND GREEN RELOADS WORKED WITHOUT ISSUES. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
18-01-2022 | Injury | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Cut | A050702 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE GOLD RELOAD ECR60D DID NOT GRAB THE TISSUE. KNIFE PUSHED TISSUE FORWARD AND BACK AGAIN. STAPLES DID NOT GO INTO TISSUE. TISSUE ON THE STOMACH WAS TORN AND HAD TO BE SUTURED OVER. SURGEON CONTINUED WITH NEW GOLD RELOAD GST60D WITH SAME RESULT. NEW STAPLER WAS OPENED. WHITE AND GREEN RELOADS WORKED WITHOUT ISSUES. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
18-01-2022 | Injury | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE GOLD RELOAD ECR60D DID NOT GRAB THE TISSUE. KNIFE PUSHED TISSUE FORWARD AND BACK AGAIN. STAPLES DID NOT GO INTO TISSUE. TISSUE ON THE STOMACH WAS TORN AND HAD TO BE SUTURED OVER. SURGEON CONTINUED WITH NEW GOLD RELOAD GST60D WITH SAME RESULT. NEW STAPLER WAS OPENED. WHITE AND GREEN RELOADS WORKED WITHOUT ISSUES. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
19-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WERE NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT FIT INTO USUALLY USER TROCAR. IT FIT INTO TROCAR, BUT COULD NOT BE BROUGHT TO HOME POSITION. NO HARM TO PATIENT. PROCEDURE: SLEEVE RESECTION.
01-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT FIT INTO USUALLY USER TROCAR. IT FIT INTO TROCAR, BUT COULD NOT BE BROUGHT TO HOME POSITION. NO HARM TO PATIENT. PROCEDURE: SLEEVE RESECTION.
30-09-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER PRIOR TO USE, THE HANDLE DOES NOT CHARGED AND DID NOT WORK. THERE WAS NO PATIENT INVOLVEMENT.
30-09-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | SIGNIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER PRIOR TO USE, THE HANDLE DOES NOT CHARGED AND DID NOT WORK. THERE WAS NO PATIENT INVOLVEMENT.
30-09-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER PRIOR TO USE, THE HANDLE DOES NOT CHARGED AND DID NOT WORK. THERE WAS NO PATIENT INVOLVEMENT.
18-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC DONOR NEPHRECTOMY PROCEDURE THAT WHILE FIRING ON THE RENAL ARTERY THE STAPLES DID NOT FORM. STAPLES WERE GOAL POSTED. SURGEON THEN CLIPPED THE ARTERY. SAME STAPLER WITH A NEW RELOAD WAS THEN FIRED ON THE RENAL VEIN WITH NO ADDITIONAL ISSUES. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGEON PERFORMED MANUAL ANASTOMOSIS TO RESOLVE THE ISSUE.
14-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGEON PERFORMED MANUAL ANASTOMOSIS TO RESOLVE THE ISSUE.
14-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGEON PERFORMED MANUAL ANASTOMOSIS TO RESOLVE THE ISSUE.
14-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | UNKNOWN ENDO GIA INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGEON PERFORMED MANUAL ANASTOMOSIS TO RESOLVE THE ISSUE.
14-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGEON PERFORMED MANUAL ANASTOMOSIS TO RESOLVE THE ISSUE.
14-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE HAD AN UNKNOWN ISSUE. THE SURGEON PERFORMED MANUAL ANASTOMOSIS TO RESOLVE THE ISSUE.
21-09-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | TRI-STAPLE 2.0 | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE RELOADS WOULD NOT FIRE COMPLETELY. MULTIPLE RELOADS WERE OPEN WITH NO SUCCESS. IT WAS NOTED THAT THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. MANUAL SUTURING WAS PERFORMED TO FIX THE STAPLE LINE. THE SURGEON OPENED NEW STAPLES AND ADJUSTED THE DIRECTION OF THE STAPLE LINE AND THE PROCEDURE WAS COMPLETED WITH POSITIVE RESULTS. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED. STAPLE PUSHERS WERE VISIBLE A QUARTER OF THE WAY UP THE CUT LINE.
21-09-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, WHILE STAPLING THE STOMACH DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE RELOADS WOULD NOT FIRE COMPLETELY. MULTIPLE RELOADS WERE OPEN WITH NO SUCCESS. IT WAS NOTED THAT THERE WAS AN INCOMPLETE STAPLE LINE PROXIMALLY. MANUAL SUTURING WAS PERFORMED TO FIX THE STAPLE LINE. THE SURGEON OPENED NEW STAPLES AND ADJUSTED THE DIRECTION OF THE STAPLE LINE AND THE PROCEDURE WAS COMPLETED WITH POSITIVE RESULTS. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED. STAPLE PUSHERS WERE VISIBLE A QUARTER OF THE WAY UP THE CUT LINE.
19-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE PLEE60A JAW WAS NOT OPENING AND THE ARTICULATION IS WAS NOT WORKING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
19-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE PLEE60A JAW WAS NOT OPENING AND THE ARTICULATION IS WAS NOT WORKING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
27-09-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BYPASS GASTROPLASTY, WHEN STAPLING, THE REINFORCEMENT FORMED THE STAPLE LINE ON ITSELF WITHOUT CUTTING OR STAPLING THE STOMACH. A NEW LOAD WITH SAME CODE WAS OPENED TO CONTINUE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT ALL THE RELOAD PUSHERS WERE FIRED, HOWEVER, THE SLED WAS NOT FULLY ADVANCED
27-09-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Separation Failure | A150301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BYPASS GASTROPLASTY, WHEN STAPLING, THE REINFORCEMENT FORMED THE STAPLE LINE ON ITSELF WITHOUT CUTTING OR STAPLING THE STOMACH. A NEW LOAD WITH SAME CODE WAS OPENED TO CONTINUE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT ALL THE RELOAD PUSHERS WERE FIRED, HOWEVER, THE SLED WAS NOT FULLY ADVANCED
12-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CUTTING OF THE TISSUE IN A LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON. BUT THE HANDLE CANNOT BE SQUEEZED. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
06-01-2022 | Injury | ethicon endo-surgery | 01-02-2022 | ILS 33MM, CURVED | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY PROCEDURE THAT AFTER COMPLETING THE FIRING SEQUENCE THE STAPLER SEEMED TO FIRE AS NORMAL AND THERE WERE TWO INTACT DONUTS. WHEN THEY WENT TO LOOK IN THE ABDOMEN THERE WAS ¿STOOL EVERYWHERE¿, THE STAPLES NEVER GRABBED THE TISSUE TO DO THE ANASTOMOSIS. SURGEON ENDED UP PREFORMED A COLOSTOMY ON THE PATIENT.
06-01-2022 | Injury | ethicon endo-surgery | 01-02-2022 | ILS 33MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY PROCEDURE THAT AFTER COMPLETING THE FIRING SEQUENCE THE STAPLER SEEMED TO FIRE AS NORMAL AND THERE WERE TWO INTACT DONUTS. WHEN THEY WENT TO LOOK IN THE ABDOMEN THERE WAS ¿STOOL EVERYWHERE¿, THE STAPLES NEVER GRABBED THE TISSUE TO DO THE ANASTOMOSIS. SURGEON ENDED UP PREFORMED A COLOSTOMY ON THE PATIENT.
16-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING RECTAL TUMOR RESECTION IN A LAPAROSCOPIC LOW PROCTECTOMY, THE DEVICE WAS UNABLE TO FIRE AFTER PRE SSING THE GREEN BUTTON. THE HANDLE CANNOT BE SQUEEZED. A NEW HANDLE IS USED TO CONTINUE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
17-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED AND THE STAPLES DID NOT FORM PROPERLY. THE SURGEON SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE.
17-01-2022 | Injury | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED AND THE STAPLES DID NOT FORM PROPERLY. THE SURGEON SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE.
20-01-2022 | Malfunction | ethicon endo-surgery | 01-02-2022 | ECHELON FLEX 60 COMPACT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE JAW WOULD NOT OPEN AFTER FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
17-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE.
17-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
17-01-2022 | Injury | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
17-01-2022 | Injury | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
17-01-2022 | Injury | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
17-01-2022 | Injury | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE STAPLES DID NOT FORM PROPERLY AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. SUTURE REINFORCEMENT WAS DONE TO RESOLVE THE ISSUE.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC RESECTION, ON GASTRIC PARENCHYMA RESECTION, THE FIRING COULD NOT BE PERFORMED AND THE RELOAD CANNOT BE REMOVED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BT UNABLE TO SQUEEZE THE HANDLE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC RESECTION, ON GASTRIC PARENCHYMA RESECTION, THE FIRING COULD NOT BE PERFORMED AND THE RELOAD CANNOT BE REMOVED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BT UNABLE TO SQUEEZE THE HANDLE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC RESECTION, ON GASTRIC PARENCHYMA RESECTION, THE FIRING COULD NOT BE PERFORMED AND THE RELOAD CANNOT BE REMOVED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BT UNABLE TO SQUEEZE THE HANDLE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC RESECTION, ON GASTRIC PARENCHYMA RESECTION, THE FIRING COULD NOT BE PERFORMED AND THE RELOAD CANNOT BE REMOVED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BT UNABLE TO SQUEEZE THE HANDLE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, THERE WAS PROBLEM UNLOADING. THE RELOAD COULD NOT BE REMOVED. THE CONNECTION BETWEEN THE ADAPTER AND THE RELOAD WAS FOUND TO BE DAMAGED AND A GAP WAS FOUND IN THE CONNECTION PART. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, THERE WAS PROBLEM UNLOADING. THE RELOAD COULD NOT BE REMOVED. THE CONNECTION BETWEEN THE ADAPTER AND THE RELOAD WAS FOUND TO BE DAMAGED AND A GAP WAS FOUND IN THE CONNECTION PART. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Dehydration | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Diarrhea | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Purulent Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Dysphagia/ Odynophagia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Dyspnea | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Emotional Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Erythema | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Heartburn/Indigestion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Memory Loss/Impairment | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Perforation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Vomiting | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Chills | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Discomfort | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fibrosis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Swollen Lymph Nodes/Glands | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Nodule | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
23-07-2019 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, THERE WAS PROBLEM UNLOADING. THE RELOAD COULD NOT BE REMOVED. THE CONNECTION BETWEEN THE ADAPTER AND THE RELOAD WAS FOUND TO BE DAMAGED AND A GAP WAS FOUND IN THE CONNECTION PART. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, THERE WAS PROBLEM UNLOADING. THE RELOAD COULD NOT BE REMOVED. THE CONNECTION BETWEEN THE ADAPTER AND THE RELOAD WAS FOUND TO BE DAMAGED AND A GAP WAS FOUND IN THE CONNECTION PART. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING THE FIFTH BLACK RELOAD IN THE BRONCHUS IN A LAPAROSCOPIC PNEUMONECTOMY, IT WAS POSSIBLE TO STAPLE TO THE END AND RETRACT THE BLADE BUT HAS A POOR STAPLE FORMATION. THE DEVICE BROKE, MADE A LOUD NOISE WHEN FIRED AND NO LONGER WORKED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING THE FIFTH BLACK RELOAD IN THE BRONCHUS IN A LAPAROSCOPIC PNEUMONECTOMY, IT WAS POSSIBLE TO STAPLE TO THE END AND RETRACT THE BLADE BUT HAS A POOR STAPLE FORMATION. THE DEVICE BROKE, MADE A LOUD NOISE WHEN FIRED AND NO LONGER WORKED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-02-2022 | Injury | covidien lp llc north haven | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
26-04-2019 | Malfunction | ethicon endo-surgery | 01-02-2022 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER ONLY FIRED HALF OF THE STAPLE CARTRIDGE THEN WOULDN¿T STAPLE FURTHER. NEW STAPLER OPENED. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
17-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, WHILE FIRING THE FIFTH BLACK RELOAD IN THE BRONCHUS IN A LAPAROSCOPIC PNEUMONECTOMY, IT WAS POSSIBLE TO STAPLE TO THE END AND RETRACT THE BLADE BUT HAS A POOR STAPLE FORMATION. THE DEVICE BROKE, MADE A LOUD NOISE WHEN FIRED AND NO LONGER WORKED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, WHILE AT THE PULMONARY ARTERY, THE DEVICE DID NOT FIRE. EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE AND THE INSTRUMENT WAS FORCE REMOVED. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE. THE DEVICE WAS EXCHANGED TO ANOTHER ONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
17-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | ENDO GIA ULTRA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, WHILE AT THE PULMONARY ARTERY, THE DEVICE DID NOT FIRE. EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE AND THE INSTRUMENT WAS FORCE REMOVED. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE. THE DEVICE WAS EXCHANGED TO ANOTHER ONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
17-01-2022 | Injury | covidien lp llc north haven | 01-02-2022 | UNKNOWN EGIA TRI STAPLE | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, WHILE AT THE PULMONARY ARTERY, THE DEVICE DID NOT FIRE. EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE AND THE INSTRUMENT WAS FORCE REMOVED. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE. THE DEVICE WAS EXCHANGED TO ANOTHER ONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
17-01-2022 | Injury | covidien lp llc north haven | 01-02-2022 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, WHILE AT THE PULMONARY ARTERY, THE DEVICE DID NOT FIRE. EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE AND THE INSTRUMENT WAS FORCE REMOVED. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE. THE DEVICE WAS EXCHANGED TO ANOTHER ONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
17-01-2022 | Injury | covidien lp llc north haven | 01-02-2022 | UNKNOWN EGIA TRI STAPLE | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, WHILE AT THE PULMONARY ARTERY, THE DEVICE DID NOT FIRE. EVEN THOUGH THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE AND THE INSTRUMENT WAS FORCE REMOVED. THERE WILL BE AN ADDITIONAL OPERATION TO BE PERFORMED TO CORRECT THE ISSUE. THE DEVICE WAS EXCHANGED TO ANOTHER ONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
17-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE SURGERY, THE FIRINGS WITH THE PURPLE RELOADS WERE OOZY, AND NEEDED TO BE CLIPPED. THE SURGEON USED A CLIP APPLIER TO REINFORCE THE STAPLE LINE TO RESOLVE THE ISSUE.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Failure to Anastomose | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Abdominal Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Adhesion(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Sepsis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Tachycardia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Obstruction/Occlusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Rectal Anastomotic Leakage | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Insufficient Information | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Leak/Splash | A0504 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Adhesion(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Tachycardia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Obstruction/Occlusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Rectal Anastomotic Leakage | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
14-02-2019 | Injury | us surgical puerto rico | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
01-02-2022 | Injury | covidien lp llc north haven | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
01-02-2022 | Injury | covidien lp llc north haven | 01-02-2022 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED A LIFE THREATENING SERIOUS INJURY THAT RESULTED FROM USE OF THE DEVICE. TREATMENT INCLUDED INTENSIVE CARE TREATMENT, EXTENDED HOSPITALIZATION, AND SUBSEQUENT SURGERIES TO CORRECT THE INJURIES.
15-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE JAWS OF THE RELOAD WILL NOT CLOSE WHILE LOADING. THERE WAS NO PATIENT INVOLVED.
15-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE JAWS OF THE RELOAD WILL NOT CLOSE WHILE LOADING. THERE WAS NO PATIENT INVOLVED.
15-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE JAWS OF THE RELOAD WILL NOT CLOSE WHILE LOADING. THERE WAS NO PATIENT INVOLVED.
15-01-2022 | Malfunction | us surgical puerto rico | 01-02-2022 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE JAWS OF THE RELOAD WILL NOT CLOSE WHILE LOADING. THERE WAS NO PATIENT INVOLVED.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Purulent Discharge | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hypoxia | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Inflammation | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Septic Shock | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Seroma | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tachycardia | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Vomiting | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Ulcer | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Anxiety | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Distress | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Discomfort | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Impaired Healing | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fungal Infection | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Post Operative Wound Infection | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Respiratory Failure | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Ascites | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Distention | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Rectal Anastomotic Leakage | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Swelling/ Edema | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Purulent Discharge | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hypoxia | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Inflammation | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Septic Shock | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Seroma | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tachycardia | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Vomiting | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Ulcer | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Anxiety | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Distress | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Discomfort | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Impaired Healing | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fungal Infection | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Post Operative Wound Infection | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Respiratory Failure | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Ascites | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Distention | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Rectal Anastomotic Leakage | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Swelling/ Edema | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
27-12-2018 | Injury | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, THE PATIENT UNDERWENT A SLEEVE GASTRECTOMY PROCEDURE WITH NO INITIAL FAULTS. POST-OPERATIVELY, THE PATIENT WAS READMITTED TO THE EMERGENCY ROOM DUE TO AN ANASTOMOTIC LEAK. THE PATIENT EXPERIENCED CONTAMINATION OF THE ABDOMINAL CAVITY, BILE LEAKAGE, ABDOMINAL PAIN, NAUSEA, VOMITING, POST-REVISION SURGERY WOUND INFECTION, ABSCESS, SEPTIC SHOCK, ACUTE RESPIRATORY FAILURE WITH HYPOXIA, EXTRALUMINAL AIR IN THE GALLBLADDER, SURGICAL DEBRIDEMENT, TISSUE INFECTION, AND FOCAL DEFECTS. TREATMENT INCLUDED ADDITIONAL HOSPITALIZATION, SURGICAL INTERVENTION, SUBSEQUENT SURGERIES, IMAGING, IRRIGATION OF WOUNDS VIA PULSE LAVAGE, INSTALLATION OF A PENROSE DRAIN, WOUND REPACKING, WOUND VAC PLACEMENT, DEBRIDEMENT AND EXPLORATION OF ABDOMINAL WOUNDS, AND FISTULA.
10-04-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD WAS FIRED UP TO 30MM ONLY THEN AFTER IT GOT STUCK AND THE STAPLER COULD NOT BE FIRED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
10-04-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD WAS FIRED UP TO 30MM ONLY THEN AFTER IT GOT STUCK AND THE STAPLER COULD NOT BE FIRED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
10-04-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD WAS FIRED UP TO 30MM ONLY THEN AFTER IT GOT STUCK AND THE STAPLER COULD NOT BE FIRED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
10-04-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE RELOAD WAS FIRED UP TO 30MM ONLY THEN AFTER IT GOT STUCK AND THE STAPLER COULD NOT BE FIRED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
16-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | EEA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOID SURGERY WHEN REMOVING THE HEMORRHOIDS THE DEVICE BROKE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
16-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOID SURGERY WHEN REMOVING THE HEMORRHOIDS THE DEVICE BROKE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
16-01-2022 | Malfunction | covidien lp llc north haven | 01-02-2022 | EEA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOID SURGERY WHEN REMOVING THE HEMORRHOIDS THE DEVICE BROKE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
17-01-2022 | Injury | us surgical puerto rico | 01-02-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER , INTRAOPERATIVELY, THERE WAS POOR STAPLE FORMATION AND STAPLE LINE WAS INCOMPLETELY DISTALLY. TO RESOLVE THE ISSUE, SUTURE REINFORCEMENT WAS DONE BY THE SURGEON.
17-01-2022 | Injury | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THERE WAS POOR STAPLE FORMATION AND STAPLE LINE WAS INCOMPLETELY DISTALLY. TO RESOLVE THE ISSUE, SUTURE REINFORCEMENT WAS DONE BY THE SURGEON.
17-01-2022 | Injury | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THERE WAS POOR STAPLE FORMATION AND STAPLE LINE WAS INCOMPLETELY DISTALLY. TO RESOLVE THE ISSUE, SUTURE REINFORCEMENT WAS DONE BY THE SURGEON.
17-01-2022 | Injury | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THERE WAS POOR STAPLE FORMATION AND STAPLE LINE WAS INCOMPLETELY DISTALLY. TO RESOLVE THE ISSUE, SUTURE REINFORCEMENT WAS DONE BY THE SURGEON.
17-01-2022 | Injury | celestica de monterrey s.a de c.v | 01-02-2022 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THERE WAS POOR STAPLE FORMATION AND STAPLE LINE WAS INCOMPLETELY DISTALLY. TO RESOLVE THE ISSUE, SUTURE REINFORCEMENT WAS DONE BY THE SURGEON.
15-01-2021 | Malfunction | covidien lp llc north haven | 01-02-2022 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTOSIGMOIDECTOMY PROCEDURE, WHILE STAPLING, THE LOAD STOPPED AT THE BEGINNING OF THE CRIMPING AND ONLY STAPLED THE TISSUE PROXIMALLY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-01-2022 | Malfunction | 3m health care / 3m | 01-02-2022 | PRECISE SKIN STAPLER | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: SKIN STAPLER JAMMED AS ATTEMPTING FIRST FIRING. FDA SAFETY REPORT ID# (B)(4).
02-08-2021 | Malfunction | covidien lp llc north haven | 02-02-2022 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN DEVICE WAS CHECKED DURING PREPARATION PRIOR TO THE START OF THE OPERATION, THE BATTERY RECHARGER FLASHED RED, AND FLUID LEAKAGE WAS OBSERVED IN THE HANDLE AND CHARGER. THE HANDLE WAS TRIED TO BE INSERTED INTO THE SHELL AND FOUND OUT THAT IT WAS NOT RECOGNIZED. THE DEVICE WAS NOT USED DURING SURGERY AND THE SURGERY WAS PERFORMED WITH A MANUAL HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
02-08-2021 | Malfunction | covidien lp llc north haven | 02-02-2022 | SIGNIA | Imprecision | A0803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN DEVICE WAS CHECKED DURING PREPARATION PRIOR TO THE START OF THE OPERATION, THE BATTERY RECHARGER FLASHED RED, AND FLUID LEAKAGE WAS OBSERVED IN THE HANDLE AND CHARGER. THE HANDLE WAS TRIED TO BE INSERTED INTO THE SHELL AND FOUND OUT THAT IT WAS NOT RECOGNIZED. THE DEVICE WAS NOT USED DURING SURGERY AND THE SURGERY WAS PERFORMED WITH A MANUAL HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
02-08-2021 | Malfunction | covidien lp llc north haven | 02-02-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN DEVICE WAS CHECKED DURING PREPARATION PRIOR TO THE START OF THE OPERATION, THE BATTERY RECHARGER FLASHED RED, AND FLUID LEAKAGE WAS OBSERVED IN THE HANDLE AND CHARGER. THE HANDLE WAS TRIED TO BE INSERTED INTO THE SHELL AND FOUND OUT THAT IT WAS NOT RECOGNIZED. THE DEVICE WAS NOT USED DURING SURGERY AND THE SURGERY WAS PERFORMED WITH A MANUAL HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
23-11-2021 | Malfunction | us surgical puerto rico | 02-02-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN LOBECTOMY, WHILE IN THE LOBAR FISSURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
23-11-2021 | Malfunction | us surgical puerto rico | 02-02-2022 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN LOBECTOMY, WHILE IN THE LOBAR FISSURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
26-04-2021 | Malfunction | covidien lp llc north haven | 02-02-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WAS SET IN THE CHARGER OVERNIGHT AND HAD SWITCHED TO GREEN ILLUMINATION, SO THE HANDLE WAS REMOVED FROM THE CHARGER AND ACTIVATED, BUT THE REMAINING BATTERY LEVEL DISPLAYED WAS NOT FULL. THE DEVICE COULD NOT BE RECHARGED AFTER BEING DISCHARGED. THE LAMP DISPLAYED ON THE CHARGER FLASHED GREEN ,THEN, ILLUMINATED GREEN AFTER 2 TO 3 MINUTES, AND RED LAMP ILLUMINATED AFTER ABOUT 30 MINUTES. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
26-04-2021 | Malfunction | covidien lp llc north haven | 02-02-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WAS SET IN THE CHARGER OVERNIGHT AND HAD SWITCHED TO GREEN ILLUMINATION, SO THE HANDLE WAS REMOVED FROM THE CHARGER AND ACTIVATED, BUT THE REMAINING BATTERY LEVEL DISPLAYED WAS NOT FULL. THE DEVICE COULD NOT BE RECHARGED AFTER BEING DISCHARGED. THE LAMP DISPLAYED ON THE CHARGER FLASHED GREEN ,THEN, ILLUMINATED GREEN AFTER 2 TO 3 MINUTES, AND RED LAMP ILLUMINATED AFTER ABOUT 30 MINUTES. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
12-12-2021 | Malfunction | us surgical puerto rico | 02-02-2022 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DUODENOPANCREATECTOMY, THE DEVICE HAD AN INTERNAL CRACK AND THE STAPLER WOULD NOT STAPLE ANYMORE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2022 | Malfunction | ethicon endo-surgery | 02-02-2022 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RECTUM SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-09-2021 | Malfunction | covidien lp llc north haven | 02-02-2022 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CORRECTIVE MAINTENANCE, THE HANDLE DID NOT KEEP CHARGE EVEN ON ANOTHER CHARGER. THERE WAS NO PATIENT INVOLVEMENT.
28-09-2021 | Malfunction | covidien lp llc north haven | 02-02-2022 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CORRECTIVE MAINTENANCE, THE HANDLE DID NOT KEEP CHARGE EVEN ON ANOTHER CHARGER. THERE WAS NO PATIENT INVOLVEMENT.
18-10-2022 | Malfunction | ethicon endo-surgery | 03-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHEN THE GST60B RELOAD WAS INSERTED INTO THE GUN EC60A, THE GUN COULD NOT FIRE, STILL LOCKED. THIS SITUATION HAPPENED WITH BOTH 2 RELOADS. WHEN SURGEON REPLACED ANOTHER NEW GUN, IT LOCKED AGAIN. THE GUN JUST COULD BE FIRED UNTIL THE SURGEON REPLACED ANOTHER NEW GST60B RELOAD. NO IMPACT TO PATIENT.
04-12-2022 | Malfunction | covidien lp | 03-01-2023 | TA AUTO SUTURE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: JAMMED. NOT DISPENSING. NEW STAPLER ACQUIRED TO FINISH PROCEDURE.
04-12-2022 | Malfunction | covidien lp | 03-01-2023 | TA AUTO SUTURE | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: JAMMED. NOT DISPENSING. NEW STAPLER ACQUIRED TO FINISH PROCEDURE.
18-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY, WHILE AT THE EMPHYSEMA LUNG PARENCHYMA, THE FIRST STAPLE SEAM COULD BE PLACED, BUT THE OVERLOAD PROTECTION HAS STRUCK. ANOTHER STAPLE SEAM COULD NO LONGER BE PLACED WITH THE DEVICE. THE SURGEON THEN INSERTED A NEW DEVICE, BUT IT HAS THE SAME ISSUE. BOTH DEVICE DID NOT FIRE. THE SURGEON PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZED THE HANDLE. THEN A THIRD HANDLE WAS USED WHERE THE OVERLOAD PROTECTION RESPONDED. THE TWO RELOADS WERE ABLE TO BE TRIGGERED AFTER EXCHANGING THE HANDLE. NO PATIENT INJURY. PLI 20 AND PLI 40: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT AN UNREPORTED CONDITION OF DEFORMATION OF CLAMPING MECHANISM THAT HAS NO RELATIONSHIP TO THE REPORTED CONDITION. STAPLE PUSHERS WERE PARTIALLY FLUSH WITH THE CARTRIDGE. PLI 30: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE.
18-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY, WHILE AT THE EMPHYSEMA LUNG PARENCHYMA, THE FIRST STAPLE SEAM COULD BE PLACED, BUT THE OVERLOAD PROTECTION HAS STRUCK. ANOTHER STAPLE SEAM COULD NO LONGER BE PLACED WITH THE DEVICE. THE SURGEON THEN INSERTED A NEW DEVICE, BUT IT HAS THE SAME ISSUE. BOTH DEVICE DID NOT FIRE. THE SURGEON PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZED THE HANDLE. THEN A THIRD HANDLE WAS USED WHERE THE OVERLOAD PROTECTION RESPONDED. THE TWO RELOADS WERE ABLE TO BE TRIGGERED AFTER EXCHANGING THE HANDLE. NO PATIENT INJURY. PLI 20 AND PLI 40: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT AN UNREPORTED CONDITION OF DEFORMATION OF CLAMPING MECHANISM THAT HAS NO RELATIONSHIP TO THE REPORTED CONDITION. STAPLE PUSHERS WERE PARTIALLY FLUSH WITH THE CARTRIDGE. PLI 30: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE.
18-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY, WHILE AT THE EMPHYSEMA LUNG PARENCHYMA, THE FIRST STAPLE SEAM COULD BE PLACED, BUT THE OVERLOAD PROTECTION HAS STRUCK. ANOTHER STAPLE SEAM COULD NO LONGER BE PLACED WITH THE DEVICE. THE SURGEON THEN INSERTED A NEW DEVICE, BUT IT HAS THE SAME ISSUE. BOTH DEVICE DID NOT FIRE. THE SURGEON PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZED THE HANDLE. THEN A THIRD HANDLE WAS USED WHERE THE OVERLOAD PROTECTION RESPONDED. THE TWO RELOADS WERE ABLE TO BE TRIGGERED AFTER EXCHANGING THE HANDLE. NO PATIENT INJURY.
18-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY, WHILE AT THE EMPHYSEMA LUNG PARENCHYMA, THE FIRST STAPLE SEAM COULD BE PLACED, BUT THE OVERLOAD PROTECTION HAS STRUCK. ANOTHER STAPLE SEAM COULD NO LONGER BE PLACED WITH THE DEVICE. THE SURGEON THEN INSERTED A NEW DEVICE, BUT IT HAS THE SAME ISSUE. BOTH DEVICE DID NOT FIRE. THE SURGEON PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZED THE HANDLE. THEN A THIRD HANDLE WAS USED WHERE THE OVERLOAD PROTECTION RESPONDED. THE TWO RELOADS WERE ABLE TO BE TRIGGERED AFTER EXCHANGING THE HANDLE. NO PATIENT INJURY.
16-12-2022 | Injury | ethicon endo-surgery | 03-01-2023 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A DST, THE KNIFE STOPPED ON THE WAY OF 1ST FIRING. KNIFE REVERSE SWITCH AND MANUAL OVERRIDE COULD NOT BE USED, AND THE JAWS DID NOT OPEN. THE DEVICE REMOVED WITH THE TROCAR FORCIBLY. ADDITIONAL SUTURE WAS PERFORMED, AND SUREFORM 45G WAS USED FOR 3 TO 4 FIRINGS TO CUT. THE SURGEON DECIDED THAT IT WILL BE DIFFICULT TO PERFORM DST, SO COLOSTOMY WAS PERFORMED. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS DONE AFTER THE PATIENT HAD RECEIVED CHEMOTHERAPY AND THE SITE OF THE CANCER WAS VERY LOW, SO THE SURGEONS WERE PLANNING THE APR AT FIRST, BUT DURING THE OPERATION THEY THOUGHT THEY COULD LEAVE THE ANUS, SO THE OPERATION CHANGED TO THE DST. SO, THE RECTUM RESECTION WAS FIRST PERFORMED BY SUREFORM45G FOR 3 FIRINGS, BUT DUE TO THE CHEMOTHERAPY, THE RECTUM WAS SO STIFF, SO THE DEVICE WAS CHANGED AT 4TH FIRING, AND THEN THE ISSUE OCCURRED. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2022 | Injury | ethicon endo-surgery | 03-01-2023 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A DST, THE KNIFE STOPPED ON THE WAY OF 1ST FIRING. KNIFE REVERSE SWITCH AND MANUAL OVERRIDE COULD NOT BE USED, AND THE JAWS DID NOT OPEN. THE DEVICE REMOVED WITH THE TROCAR FORCIBLY. ADDITIONAL SUTURE WAS PERFORMED, AND SUREFORM 45G WAS USED FOR 3 TO 4 FIRINGS TO CUT. THE SURGEON DECIDED THAT IT WILL BE DIFFICULT TO PERFORM DST, SO COLOSTOMY WAS PERFORMED. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS DONE AFTER THE PATIENT HAD RECEIVED CHEMOTHERAPY AND THE SITE OF THE CANCER WAS VERY LOW, SO THE SURGEONS WERE PLANNING THE APR AT FIRST, BUT DURING THE OPERATION THEY THOUGHT THEY COULD LEAVE THE ANUS, SO THE OPERATION CHANGED TO THE DST. SO, THE RECTUM RESECTION WAS FIRST PERFORMED BY SUREFORM45G FOR 3 FIRINGS, BUT DUE TO THE CHEMOTHERAPY, THE RECTUM WAS SO STIFF, SO THE DEVICE WAS CHANGED AT 4TH FIRING, AND THEN THE ISSUE OCCURRED. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2022 | Injury | ethicon endo-surgery | 03-01-2023 | POWERED PLUS STANDARD 45 | Activation Failure | A150101 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A DST, THE KNIFE STOPPED ON THE WAY OF 1ST FIRING. KNIFE REVERSE SWITCH AND MANUAL OVERRIDE COULD NOT BE USED, AND THE JAWS DID NOT OPEN. THE DEVICE REMOVED WITH THE TROCAR FORCIBLY. ADDITIONAL SUTURE WAS PERFORMED, AND SUREFORM 45G WAS USED FOR 3 TO 4 FIRINGS TO CUT. THE SURGEON DECIDED THAT IT WILL BE DIFFICULT TO PERFORM DST, SO COLOSTOMY WAS PERFORMED. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS DONE AFTER THE PATIENT HAD RECEIVED CHEMOTHERAPY AND THE SITE OF THE CANCER WAS VERY LOW, SO THE SURGEONS WERE PLANNING THE APR AT FIRST, BUT DURING THE OPERATION THEY THOUGHT THEY COULD LEAVE THE ANUS, SO THE OPERATION CHANGED TO THE DST. SO, THE RECTUM RESECTION WAS FIRST PERFORMED BY SUREFORM45G FOR 3 FIRINGS, BUT DUE TO THE CHEMOTHERAPY, THE RECTUM WAS SO STIFF, SO THE DEVICE WAS CHANGED AT 4TH FIRING, AND THEN THE ISSUE OCCURRED. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2022 | Injury | ethicon endo-surgery | 03-01-2023 | POWERED PLUS STANDARD 45 | Activation Problem | A1501 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A DST, THE KNIFE STOPPED ON THE WAY OF 1ST FIRING. KNIFE REVERSE SWITCH AND MANUAL OVERRIDE COULD NOT BE USED, AND THE JAWS DID NOT OPEN. THE DEVICE REMOVED WITH THE TROCAR FORCIBLY. ADDITIONAL SUTURE WAS PERFORMED, AND SUREFORM 45G WAS USED FOR 3 TO 4 FIRINGS TO CUT. THE SURGEON DECIDED THAT IT WILL BE DIFFICULT TO PERFORM DST, SO COLOSTOMY WAS PERFORMED. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THE PROCEDURE WAS DONE AFTER THE PATIENT HAD RECEIVED CHEMOTHERAPY AND THE SITE OF THE CANCER WAS VERY LOW, SO THE SURGEONS WERE PLANNING THE APR AT FIRST, BUT DURING THE OPERATION THEY THOUGHT THEY COULD LEAVE THE ANUS, SO THE OPERATION CHANGED TO THE DST. SO, THE RECTUM RESECTION WAS FIRST PERFORMED BY SUREFORM45G FOR 3 FIRINGS, BUT DUE TO THE CHEMOTHERAPY, THE RECTUM WAS SO STIFF, SO THE DEVICE WAS CHANGED AT 4TH FIRING, AND THEN THE ISSUE OCCURRED. NO FURTHER INFORMATION IS AVAILABLE.
12-12-2022 | Malfunction | davol inc sub c.r bard inc -1213643 | 03-01-2023 | SORBAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A VIDEO LAPAROSCOPY INGUINAL HERNIA REPAIR PROCEDURE, THE SURGEON TRIED TO FIXATE A BARD/DAVOL POLYPROPYLENE MESH (UNSPECIFIED) USING A SORBAFIX ENHANCED FIXATION DEVICE, THE FASTENERS DID NOT PENETRATE THE TISSUE. AS REPORTED, THERE WERE NO REMNANTS OF THE DEVICE LEFT INSIDE THE PATIENT AND THERE WAS NO REPORTED PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE JAWS OF THE DEVICE DID NOT CLOSE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-07-2021 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, DURING PREPARATION THE DEVICE ROTATION WAS NOT POSSIBLE EVEN WHEN THE LEFT UPPER/LOWER ROTATION BUTTON WAS PRESSED IN THE DIRECTION OF THE CARTRIDGE. ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS LOOSE MOTOR SCREWS.
28-07-2021 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, DURING PREPARATION THE DEVICE ROTATION WAS NOT POSSIBLE EVEN WHEN THE LEFT UPPER/LOWER ROTATION BUTTON WAS PRESSED IN THE DIRECTION OF THE CARTRIDGE. ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS LOOSE MOTOR SCREWS.
17-01-2023 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-01-2023 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE HAD AN UNKNOWN ISSUE. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOT-ASSISTED LAPAROSCOPIC DA VINCI TOTAL CYSTECTOMY, ON INTESTINAL ANASTOMOSIS AFTER ILEAL CONDUIT CREATION, AFTER THE THIRD RELOAD REPLACEMENT, THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE DONE. A NEW HANDLE WAS USED TO COMPLETE THE CASE. THE RELOAD COULD BE USED AND WAS NOT REPLACED. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROTOMY OF THE ESOPHAGUS, WHILE ON TISSUE RESECTION, THE GEARING KNOB DID NOT ADVANCE WHEN PUSHED FORWARD, THUS FIRING WAS NOT PERFORMED. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | EEA | Device Slipped | A051204 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTROINTESTINAL PROCEDURE, WHEN THE SURGEON WANTED TO CLOSE THE DEVICE, THE JAWS DID NOT CLOSE ON THE TISSUE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED, THE JAWS WERE STILL OPEN AND WERE NOT ABLE TO FIRE; THEREFORE, THE TISSUE COULD NOT BE STAPLED. THE GREEN BUTTON DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
02-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | EEA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTROINTESTINAL PROCEDURE, WHEN THE SURGEON WANTED TO CLOSE THE DEVICE, THE JAWS DID NOT CLOSE ON THE TISSUE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED, THE JAWS WERE STILL OPEN AND WERE NOT ABLE TO FIRE; THEREFORE, THE TISSUE COULD NOT BE STAPLED. THE GREEN BUTTON DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
02-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE JAWS WERE IMPOSSIBLE TO CLOSE ON THE TISSUE, THUS THE GREEN BUTTON COULD NOT BE PRESSED TO FIRE THE DEVICE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED, THE JAWS WERE STILL OPEN AND WERE NOT ABLE TO FIRE; THEREFORE, THE TISSUE COULD NOT BE STAPLED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE JAWS WERE IMPOSSIBLE TO CLOSE ON THE TISSUE, THUS THE GREEN BUTTON COULD NOT BE PRESSED TO FIRE THE DEVICE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED, THE JAWS WERE STILL OPEN AND WERE NOT ABLE TO FIRE; THEREFORE, THE TISSUE COULD NOT BE STAPLED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A BYPASS OPERATION, WHILE STAPLING THE LITTLE STOMACH, THE RELOAD WITH THE POWERED STAPLER HAD A STAPLE PROBLEM, SEVERAL STAPLES REMAINED ON THE RELOAD AND CAUSED A SMALL WOUND IN THE STOMACH. THE STAPLE WAS POORLY FORMED. COMPETITOR'S SUTURE WAS NEEDED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
01-01-2022 | Malfunction | ethicon | 03-01-2023 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2022 AND ABSORBABLE STRAPS WERE USED. THE STAPLE DID NOT CLOSE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION WAS PROVIDED.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Device Displays Incorrect Message | A090201 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Device Displays Incorrect Message | A090201 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
03-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DEVICE FOLLOW UP, WHILE ON CHARGER, THE HANDLE FAILED TO CHARGE. THE HANDLE WAS BLACK WITHOUT POSSIBILITY TO CHARGE. ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE CELL VENTING IS AN INTEGRATED DESIGN FEATURE OF THE CELL THAT PRODUCES A DISCHARGE OF THE ELECTROLYTE FLUID AND RENDERS THE BATTERY CELL NONFUNCTIONAL AND NON-CHARGEABLE WHEN BATTERY CELL LIFE HAS BEEN EXCEEDED. ROOT CAUSE OF THE OBSERVED VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE). DEVICE BATTERY MANAGEMENT ENHANCEMENTS HAVE BEEN IMPLEMENTED TO EXTEND BATTERY LIFE AND ENHANCE BATTERY HEALTH MONITORING.
03-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DEVICE FOLLOW UP, WHILE ON CHARGER, THE HANDLE FAILED TO CHARGE. THE HANDLE WAS BLACK WITHOUT POSSIBILITY TO CHARGE. ANOTHER HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE CELL VENTING IS AN INTEGRATED DESIGN FEATURE OF THE CELL THAT PRODUCES A DISCHARGE OF THE ELECTROLYTE FLUID AND RENDERS THE BATTERY CELL NONFUNCTIONAL AND NON-CHARGEABLE WHEN BATTERY CELL LIFE HAS BEEN EXCEEDED. ROOT CAUSE OF THE OBSERVED VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE). DEVICE BATTERY MANAGEMENT ENHANCEMENTS HAVE BEEN IMPLEMENTED TO EXTEND BATTERY LIFE AND ENHANCE BATTERY HEALTH MONITORING.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
29-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE RELOAD FIRED AND WOULD NOT CONTINUE STAPLING. THE WAS AN INCOMPLETE STAPLE LINE DISTALLY. THE RELOAD HAD TO BE REMOVED AND REMAINING TISSUE WAS RESECTED. THEN, THE SHELL SAID THAT IT HAD "0" USES LEFT SO IT HAD TO BE OPENED. A NEW SHELL WHICH WAS RECOGNIZED PERFECTLY AND ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
03-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING, THE DEVICE FAILED TO CHARGE. THE HANDLE WAS BLACK WITH NO POSSIBILITY TO START AGAIN. NOTHING HAPPENED WHEN IT WAS TAKEN OUT OF THE CHARGER, SO ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THE CHARGER WORKED WITH THE OTHER HANDLES. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT ONE BATTERY CELL WAS FOUND TO BE VENTED.
03-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CHARGING, THE DEVICE FAILED TO CHARGE. THE HANDLE WAS BLACK WITH NO POSSIBILITY TO START AGAIN. NOTHING HAPPENED WHEN IT WAS TAKEN OUT OF THE CHARGER, SO ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THE CHARGER WORKED WITH THE OTHER HANDLES. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT ONE BATTERY CELL WAS FOUND TO BE VENTED.
14-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | GIA | Protective Measures Problem | A16 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RIGHT HEMICOLECTOMY PROCEDURE, BLADE EJECTION CANNOT BE LOADED AFTER PRODUCT PACKAGE WAS DISASSEMBLED. THE DEVICE WAS UNABLE TO USE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE CHARGER WAS BLINKING RED AND THE HANDLE WOULD NOT TURN ON. EVEN WHEN CHARGER WAS REPLACED, HANDLE STILL WOULD NOT CHARGE. THE SECOND CHARGER WORKED FINE WITH ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
03-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE CHARGER WAS BLINKING RED AND THE HANDLE WOULD NOT TURN ON. EVEN WHEN CHARGER WAS REPLACED, HANDLE STILL WOULD NOT CHARGE. THE SECOND CHARGER WORKED FINE WITH ANOTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT.
28-09-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE WAS REMOVED FROM THE CHARGER TO PREPARE FOR OPERATION, BUT THE DEVICE DID NOT START UP. THE DISPLAY WAS ALSO BLACKED OUT. THE HANDLE WAS ATTEMPTED TO CHARGE AGAIN, BUT IT WAS NOT CHARGED. THE USE OF THE DEVICE WAS STOPPED, AND THE SURGERY WAS COMPLETED WITHOUT ANY DELAY USING ANOTHER COMPANY'S DEVICE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
28-09-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE WAS REMOVED FROM THE CHARGER TO PREPARE FOR OPERATION, BUT THE DEVICE DID NOT START UP. THE DISPLAY WAS ALSO BLACKED OUT. THE HANDLE WAS ATTEMPTED TO CHARGE AGAIN, BUT IT WAS NOT CHARGED. THE USE OF THE DEVICE WAS STOPPED, AND THE SURGERY WAS COMPLETED WITHOUT ANY DELAY USING ANOTHER COMPANY'S DEVICE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
05-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE FIRING THE FOURTH RELOAD AT THE TIME OF COLON RESECTION, A BIG CLANK SOUND FROM THE HANDLE WAS HEARD AND THEN NO STAPLING WAS POSSIBLE; THE SURGEON COULD NOT SQUEEZE THE HANDLE EVEN WHEN THE GREEN BUTTON WAS PRESSED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. ANOTHER HANDLE THAT WAS USED DID NOT WORK. THERE WAS NO PATIENT INJURY.
05-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE FIRING THE FOURTH RELOAD AT THE TIME OF COLON RESECTION, A BIG CLANK SOUND FROM THE HANDLE WAS HEARD AND THEN NO STAPLING WAS POSSIBLE; THE SURGEON COULD NOT SQUEEZE THE HANDLE EVEN WHEN THE GREEN BUTTON WAS PRESSED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. ANOTHER HANDLE THAT WAS USED DID NOT WORK. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, THE GASTRIC BODY PART WAS INSERTED INTO JAWS OF THE RELEVANT RELOAD, THE JAWS WERE CLOSED, THE GREEN BUTTON WAS PRESSED AND AS SOON AS THE HANDLE WAS SQUEEZED, IT BECAME UNABLE TO SQUEEZE AND THE DEVICE DID NOT FIRE. FOR THAT REASON, THE HANDLE WAS SQUEEZED STRONGLY BUT THE SITUATION DID NOT IMPROVE, SO THE RELOAD AND THE STAPLER WERE REPLACED, AND THE DEVICE WAS USED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
05-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHILE FIRING THE FOURTH RELOAD AT THE TIME OF COLON RESECTION, A BIG CLANK SOUND FROM THE HANDLE WAS HEARD AND THEN NO STAPLING WAS POSSIBLE; THE SURGEON COULD NOT SQUEEZE THE HANDLE EVEN WHEN THE GREEN BUTTON WAS PRESSED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. ANOTHER HANDLE THAT WAS USED DID NOT WORK. THERE WAS NO PATIENT INJURY.
05-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL ESOPHAGO-GASTRIC PROCEDURE, THE JAWS DID NOT CLOSE ON THE TISSUE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED, THE JAWS WERE STILL OPEN AND WAS NOT ABLE TO FIRE THEREFORE THE TISSUE COULD NOT BE STAPLED. THE GREEN BUTTON DID NOT FIRE. IT WAS ALSO NOTED THAT THERE WAS DIFFICULTY UNLOADING THE RELOAD. BOTH THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL ESOPHAGO-GASTRIC PROCEDURE, THE JAWS DID NOT CLOSE ON THE TISSUE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED, THE JAWS WERE STILL OPEN AND WAS NOT ABLE TO FIRE THEREFORE THE TISSUE COULD NOT BE STAPLED. THE GREEN BUTTON DID NOT FIRE. IT WAS ALSO NOTED THAT THERE WAS DIFFICULTY UNLOADING THE RELOAD. BOTH THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE POWER HANDLE DID NOT START UP AND THE POWER SHELL COULD NOT BE USED. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
24-10-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE POWER HANDLE DID NOT START UP AND THE POWER SHELL COULD NOT BE USED. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
05-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL ESOPHAGO-GASTRIC PROCEDURE, THE JAWS DID NOT CLOSE ON THE TISSUE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED, THE JAWS WERE STILL OPEN AND WAS NOT ABLE TO FIRE THEREFORE THE TISSUE COULD NOT BE STAPLED. THE GREEN BUTTON DID NOT FIRE. IT WAS ALSO NOTED THAT THERE WAS DIFFICULTY UNLOADING THE RELOAD. BOTH THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | UNKNOWN ENDO GIA SULU | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL ESOPHAGO-GASTRIC PROCEDURE, THE JAWS DID NOT CLOSE ON THE TISSUE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED, THE JAWS WERE STILL OPEN AND WAS NOT ABLE TO FIRE THEREFORE THE TISSUE COULD NOT BE STAPLED. THE GREEN BUTTON DID NOT FIRE. IT WAS ALSO NOTED THAT THERE WAS DIFFICULTY UNLOADING THE RELOAD. BOTH THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ABSORBATACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC TRANSABDOMINAL PREPERITONEAL (TAPP), THE PART OF THE SHAFT CONNECTED TO THE BODY OF THE HANDLE BROKE WHILE THE SURGEON WAS TRYING TO CLOSE THE PERITONEAL LAYER. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
02-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE JAWS WERE IMPOSSIBLE TO CLOSE ON THE TISSUE, THUS THE GREEN BUTTON COULD NOT BE PRESSED TO FIRE THE DEVICE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED ,THE JAWS WAS STILL OPEN AND WAS NOT ABLE TO FIRE THEREFORE THE TISSUE COULD NOT BE STAPLED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, THE JAWS WERE IMPOSSIBLE TO CLOSE ON THE TISSUE, THUS THE GREEN BUTTON COULD NOT BE PRESSED TO FIRE THE DEVICE. WHEN THE HAND OF THE SURGEON RELAXED AND OPENED ,THE JAWS WAS STILL OPEN AND WAS NOT ABLE TO FIRE THEREFORE THE TISSUE COULD NOT BE STAPLED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-11-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE GASTROPLASTY SURGERY, THE RELOAD DID NOT WORK, AS IT DID NOT WANT TO CLOSE AND FIRE. ANOTHER RELOAD WAS USED TO COMPLETE CASE. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, THE GASTRIC BODY PART WAS INSERTED INTO JAWS OF THE RELEVANT RELOAD, THE JAWS WERE CLOSED, THE GREEN BUTTON WAS PRESSED AND AS SOON AS THE HANDLE WAS SQUEEZED, IT BECAME UNABLE TO SQUEEZE AND THE DEVICE DID NOT FIRE. FOR THAT REASON, THE HANDLE WAS SQUEEZED STRONGLY BUT THE SITUATION DID NOT IMPROVE, SO THE RELOAD AND THE STAPLER WERE REPLACED, AND THE DEVICE WAS USED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION SURGERY, WHILE AT RECTAL RESECTION, CLAMP COULD NOT BE DONE FOR THE 2 RELOADS AND HANDLE, ALTHOUGH THERE WAS NO PROBLEM IN CONFIRMING THE OPENING AND CLOSING IN PREPARATION. EVEN AFTER SOMEHOW TRYING THE FIRING, IT COULD NOT BE DONE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LEFT UPPER QUADRANT RESECTION, THERE WAS A STAPLE LINE INCOMPLETE AT THE PROXIMAL SIDE AND POOR STAPLE FORMATION ON BOTH RELOAD. ON THE FIRST FIRING, THE BLOOD WAS OOZING AND HEMOSTASIS BY COMPRESSION WAS DONE. ON THE SECOND FIRING, USING THE SECOND RELOAD FROM A DIFFERENT BATCH, WHILE AT THE SUPERIOR PULMONARY VEIN, A FIRE WAS PERFORMED AT A LOCATION WHERE THE CROTCH OF THE BIFURCATION WAS JUST OVER THE KNIFE LINE. BLEEDING FROM THE TWO BRANCH VESSELS DID NOT STOP. SUTURE LIGATION WAS PERFORMED FOR HEMOSTASIS ON THE SUPERIOR PULMONARY VEIN WITH EXAMINATION THREAD TO STOP TWO BRANCH FROM BLEEDING. THE STAPLE LINE ON THE REMAINING TRUNK SIDE OF THE PATIENT WAS NOT CLEAN, AND BLEEDING FROM THE DISCONNECTED AREA DID NOT STOP. THERE WAS 30CC OF BLEEDING AND 30 MINUTES OF COMPRESSION. HEMOSTASIS WITH SEAL AFTER COMPRESSION WAS DONE. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. A NEW RELOAD WAS USED TO COMPLETE THE CASE.
14-12-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LEFT UPPER QUADRANT RESECTION, THERE WAS A STAPLE LINE INCOMPLETE AT THE PROXIMAL SIDE AND POOR STAPLE FORMATION ON BOTH RELOAD. ON THE FIRST FIRING, THE BLOOD WAS OOZING AND HEMOSTASIS BY COMPRESSION WAS DONE. ON THE SECOND FIRING, USING THE SECOND RELOAD FROM A DIFFERENT BATCH, WHILE AT THE SUPERIOR PULMONARY VEIN, A FIRE WAS PERFORMED AT A LOCATION WHERE THE CROTCH OF THE BIFURCATION WAS JUST OVER THE KNIFE LINE. BLEEDING FROM THE TWO BRANCH VESSELS DID NOT STOP. SUTURE LIGATION WAS PERFORMED FOR HEMOSTASIS ON THE SUPERIOR PULMONARY VEIN WITH EXAMINATION THREAD TO STOP TWO BRANCH FROM BLEEDING. THE STAPLE LINE ON THE REMAINING TRUNK SIDE OF THE PATIENT WAS NOT CLEAN, AND BLEEDING FROM THE DISCONNECTED AREA DID NOT STOP. THERE WAS 30CC OF BLEEDING AND 30 MINUTES OF COMPRESSION. HEMOSTASIS WITH SEAL AFTER COMPRESSION WAS DONE. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. A NEW RELOAD WAS USED TO COMPLETE THE CASE.
14-12-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LEFT UPPER QUADRANT RESECTION, THERE WAS A STAPLE LINE INCOMPLETE AT THE PROXIMAL SIDE AND POOR STAPLE FORMATION ON BOTH RELOAD. ON THE FIRST FIRING, THE BLOOD WAS OOZING AND HEMOSTASIS BY COMPRESSION WAS DONE. ON THE SECOND FIRING, USING THE SECOND RELOAD FROM A DIFFERENT BATCH, WHILE AT THE SUPERIOR PULMONARY VEIN, A FIRE WAS PERFORMED AT A LOCATION WHERE THE CROTCH OF THE BIFURCATION WAS JUST OVER THE KNIFE LINE. BLEEDING FROM THE TWO BRANCH VESSELS DID NOT STOP. SUTURE LIGATION WAS PERFORMED FOR HEMOSTASIS ON THE SUPERIOR PULMONARY VEIN WITH EXAMINATION THREAD TO STOP TWO BRANCH FROM BLEEDING. THE STAPLE LINE ON THE REMAINING TRUNK SIDE OF THE PATIENT WAS NOT CLEAN, AND BLEEDING FROM THE DISCONNECTED AREA DID NOT STOP. THERE WAS 30CC OF BLEEDING AND 30 MINUTES OF COMPRESSION. HEMOSTASIS WITH SEAL AFTER COMPRESSION WAS DONE. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. A NEW RELOAD WAS USED TO COMPLETE THE CASE.
14-12-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LEFT UPPER QUADRANT RESECTION, THERE WAS A STAPLE LINE INCOMPLETE AT THE PROXIMAL SIDE AND POOR STAPLE FORMATION ON BOTH RELOAD. ON THE FIRST FIRING, THE BLOOD WAS OOZING AND HEMOSTASIS BY COMPRESSION WAS DONE. ON THE SECOND FIRING, USING THE SECOND RELOAD FROM A DIFFERENT BATCH, WHILE AT THE SUPERIOR PULMONARY VEIN, A FIRE WAS PERFORMED AT A LOCATION WHERE THE CROTCH OF THE BIFURCATION WAS JUST OVER THE KNIFE LINE. BLEEDING FROM THE TWO BRANCH VESSELS DID NOT STOP. SUTURE LIGATION WAS PERFORMED FOR HEMOSTASIS ON THE SUPERIOR PULMONARY VEIN WITH EXAMINATION THREAD TO STOP TWO BRANCH FROM BLEEDING. THE STAPLE LINE ON THE REMAINING TRUNK SIDE OF THE PATIENT WAS NOT CLEAN, AND BLEEDING FROM THE DISCONNECTED AREA DID NOT STOP. THERE WAS 30CC OF BLEEDING AND 30 MINUTES OF COMPRESSION. HEMOSTASIS WITH SEAL AFTER COMPRESSION WAS DONE. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. A NEW RELOAD WAS USED TO COMPLETE THE CASE.
14-12-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LEFT UPPER QUADRANT RESECTION, THERE WAS A STAPLE LINE INCOMPLETE AT THE PROXIMAL SIDE AND POOR STAPLE FORMATION ON BOTH RELOAD. ON THE FIRST FIRING, THE BLOOD WAS OOZING AND HEMOSTASIS BY COMPRESSION WAS DONE. ON THE SECOND FIRING, USING THE SECOND RELOAD FROM A DIFFERENT BATCH, WHILE AT THE SUPERIOR PULMONARY VEIN, A FIRE WAS PERFORMED AT A LOCATION WHERE THE CROTCH OF THE BIFURCATION WAS JUST OVER THE KNIFE LINE. BLEEDING FROM THE TWO BRANCH VESSELS DID NOT STOP. SUTURE LIGATION WAS PERFORMED FOR HEMOSTASIS ON THE SUPERIOR PULMONARY VEIN WITH EXAMINATION THREAD TO STOP TWO BRANCH FROM BLEEDING. THE STAPLE LINE ON THE REMAINING TRUNK SIDE OF THE PATIENT WAS NOT CLEAN, AND BLEEDING FROM THE DISCONNECTED AREA DID NOT STOP. THERE WAS 30CC OF BLEEDING AND 30 MINUTES OF COMPRESSION. HEMOSTASIS WITH SEAL AFTER COMPRESSION WAS DONE. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. A NEW RELOAD WAS USED TO COMPLETE THE CASE.
14-12-2022 | Injury | covidien lp llc north haven | 03-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LEFT UPPER QUADRANT RESECTION, THERE WAS A STAPLE LINE INCOMPLETE AT THE PROXIMAL SIDE AND POOR STAPLE FORMATION ON BOTH RELOAD. ON THE FIRST FIRING, THE BLOOD WAS OOZING AND HEMOSTASIS BY COMPRESSION WAS DONE. ON THE SECOND FIRING, USING THE SECOND RELOAD FROM A DIFFERENT BATCH, WHILE AT THE SUPERIOR PULMONARY VEIN, A FIRE WAS PERFORMED AT A LOCATION WHERE THE CROTCH OF THE BIFURCATION WAS JUST OVER THE KNIFE LINE. BLEEDING FROM THE TWO BRANCH VESSELS DID NOT STOP. SUTURE LIGATION WAS PERFORMED FOR HEMOSTASIS ON THE SUPERIOR PULMONARY VEIN WITH EXAMINATION THREAD TO STOP TWO BRANCH FROM BLEEDING. THE STAPLE LINE ON THE REMAINING TRUNK SIDE OF THE PATIENT WAS NOT CLEAN, AND BLEEDING FROM THE DISCONNECTED AREA DID NOT STOP. THERE WAS 30CC OF BLEEDING AND 30 MINUTES OF COMPRESSION. HEMOSTASIS WITH SEAL AFTER COMPRESSION WAS DONE. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. A NEW RELOAD WAS USED TO COMPLETE THE CASE.
06-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE DID NOT FIRE.
06-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE DID NOT FIRE.
14-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, ARTICULATION WAS DIFFICULT. WHEN THE RELOAD WAS ARTICULATED OUTSIDE OF THE BODY, THE ARTICULATION PART WAS DAMAGED, AND THE PARTS FLEW INTO THE OPERATING ROOM. HANDLE, RELOAD AND ADAPTER WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, ARTICULATION WAS DIFFICULT. WHEN THE RELOAD WAS ARTICULATED OUTSIDE OF THE BODY, THE ARTICULATION PART WAS DAMAGED, AND THE PARTS FLEW INTO THE OPERATING ROOM. HANDLE, RELOAD AND ADAPTER WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Positioning Problem | A1502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, ARTICULATION WAS DIFFICULT. WHEN THE RELOAD WAS ARTICULATED OUTSIDE OF THE BODY, THE ARTICULATION PART WAS DAMAGED, AND THE PARTS FLEW INTO THE OPERATING ROOM. HANDLE, RELOAD AND ADAPTER WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-11-2022 | Injury | celestica de monterrey s.a de c.v | 03-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY FOR MORBID OBESITY FOR VIDEOLAPAROSCOPY, SOME STAPLES DID NOT CLOSED PROPERLY; THE TISSUE WAS OPEN AT SOME POINTS AND SLIGHT BLEEDING WAS OBSERVED. THE SURGEON SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE.
29-11-2022 | Injury | celestica de monterrey s.a de c.v | 03-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY FOR MORBID OBESITY FOR VIDEOLAPAROSCOPY, SOME STAPLES DID NOT CLOSED PROPERLY; THE TISSUE WAS OPEN AT SOME POINTS AND SLIGHT BLEEDING WAS OBSERVED. THE SURGEON SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE.
29-11-2022 | Injury | celestica de monterrey s.a de c.v | 03-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY FOR MORBID OBESITY FOR VIDEOLAPAROSCOPY, SOME STAPLES DID NOT CLOSED PROPERLY; THE TISSUE WAS OPEN AT SOME POINTS AND SLIGHT BLEEDING WAS OBSERVED. THE SURGEON SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE.
29-11-2022 | Injury | celestica de monterrey s.a de c.v | 03-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY FOR MORBID OBESITY FOR VIDEOLAPAROSCOPY, SOME STAPLES DID NOT CLOSED PROPERLY; THE TISSUE WAS OPEN AT SOME POINTS AND SLIGHT BLEEDING WAS OBSERVED. THE SURGEON SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE.
14-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SEPTATION, WHILE STAPLING THE STOMACH, THE HANDLE LOCKED AFTER THE SECOND SHOT, HENCE COULD NOT BE USED. THE RELOAD WAS CHANGED AND WAS CONNECTED TO THE HANDLE BUT THE HANDLE DID NOT GIVE A GREEN LIGHT SIGNAL. THE ADAPTER WAS THEN CHANGED BUT STILL THE DEFECT PERSISTED. IT WAS ALSO NOTED THAT THE HANDLE, EVEN WITHOUT THE SHELL CONTINUED IN SHOOTING MODE. THE MANEUVER OF RESTARTING THE HANDLE WAS CARRIED OUT TO RETURN TO NORMAL OPERATION. AFTER LOCKING THE HANDLE AND CHANGING THE STEM, A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LUNG SEGMENTECTOMY, WHILE AT PARENCHYMAL RESECTION, THE RELOAD WAS LOADED INT O THE HANDLE AND USED FOR STAPLING, THE HANDLE BECAME HEAVY AND MADE A "CRACKING" SOUND. THE SOUND CONTINUED TWICE, DEVICE WAS SLIGHTLY ABLE TO BE STAPLED, BUT IT WAS UNKNOWN HOW LONG (CM) IT WAS ABLE TO BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND IT WAS NOTED THAT DUE TO THE THICKNESS OF THE TISSUE, THE GEARS INSIDE THE HANDLE WERE BROKEN IN SEVERAL PLACES, AND IT BECAME IMPOSSIBLE TO SQUEEZE THE HANDLE (NOT CAUGHT). TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THE PHYSICIAN DECIDED TO CHANGE THE HANDLE TO A NEW ONE, AND A DIFFERENT RELOAD WAS TAKEN OUT AND FIRED. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LUNG SEGMENTECTOMY, WHILE AT PARENCHYMAL RESECTION, THE RELOAD WAS LOADED INT O THE HANDLE AND USED FOR STAPLING, THE HANDLE BECAME HEAVY AND MADE A "CRACKING" SOUND. THE SOUND CONTINUED TWICE, DEVICE WAS SLIGHTLY ABLE TO BE STAPLED, BUT IT WAS UNKNOWN HOW LONG (CM) IT WAS ABLE TO BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND IT WAS NOTED THAT DUE TO THE THICKNESS OF THE TISSUE, THE GEARS INSIDE THE HANDLE WERE BROKEN IN SEVERAL PLACES, AND IT BECAME IMPOSSIBLE TO SQUEEZE THE HANDLE (NOT CAUGHT). TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THE PHYSICIAN DECIDED TO CHANGE THE HANDLE TO A NEW ONE, AND A DIFFERENT RELOAD WAS TAKEN OUT AND FIRED. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LUNG SEGMENTECTOMY, WHILE AT PARENCHYMAL RESECTION, THE RELOAD WAS LOADED INT O THE HANDLE AND USED FOR STAPLING, THE HANDLE BECAME HEAVY AND MADE A "CRACKING" SOUND. THE SOUND CONTINUED TWICE, DEVICE WAS SLIGHTLY ABLE TO BE STAPLED, BUT IT WAS UNKNOWN HOW LONG (CM) IT WAS ABLE TO BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND IT WAS NOTED THAT DUE TO THE THICKNESS OF THE TISSUE, THE GEARS INSIDE THE HANDLE WERE BROKEN IN SEVERAL PLACES, AND IT BECAME IMPOSSIBLE TO SQUEEZE THE HANDLE (NOT CAUGHT). TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THE PHYSICIAN DECIDED TO CHANGE THE HANDLE TO A NEW ONE, AND A DIFFERENT RELOAD WAS TAKEN OUT AND FIRED. THERE WAS NO PATIENT INJURY.
09-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDE ASSISTED THORACOSCOPIC PARTIAL PULMONARY RESECTION, WHEN THE DEVICE WAS FIRED, A METAL STRING-SHAPED OBJECT FROM THE RELOAD FELL INTO THE SURGICAL FIELD. THE DEVICE HAD NO PROBLEM WITH THE FIRING SO IT WAS USED AS IT IS. NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 03-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LUNG SEGMENTECTOMY, WHILE AT PARENCHYMAL RESECTION, THE RELOAD WAS LOADED INT O THE HANDLE AND USED FOR STAPLING, THE HANDLE BECAME HEAVY AND MADE A "CRACKING" SOUND. THE SOUND CONTINUED TWICE, DEVICE WAS SLIGHTLY ABLE TO BE STAPLED, BUT IT WAS UNKNOWN HOW LONG (CM) IT WAS ABLE TO BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND IT WAS NOTED THAT DUE TO THE THICKNESS OF THE TISSUE, THE GEARS INSIDE THE HANDLE WERE BROKEN IN SEVERAL PLACES, AND IT BECAME IMPOSSIBLE TO SQUEEZE THE HANDLE (NOT CAUGHT). TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THE PHYSICIAN DECIDED TO CHANGE THE HANDLE TO A NEW ONE, AND A DIFFERENT RELOAD WAS TAKEN OUT AND FIRED. THERE WAS NO PATIENT INJURY.
01-12-2022 | Injury | covidien lp llc north haven | 03-01-2023 | UNKNOWN EGIA TRI STAPLE | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, IN STAPLING ACROSS STOMACH, THE HANDLE MADE A CRACKING SOUND EACH TIME THE SURGEON TRIED TO FIRE. THE STAPLE LINE WAS ALSO INCOMPLETE. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED.
12-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCOAST TUMOR RESECTION, IN STAPLING THE BRONCHUS, THE DEVICE STOPPED FIRING. THE SURGEON WAS NOT ABLE TO SEE IF THE I BEAM WAS ADVANCING AS IT WAS AN OPEN PROCEDURE. THE DEVICE STAPLED WITHOUT CUTTING THE TISSUE. AFTER THE CASE, VISUALLY, THE RELOAD DID NOT LOOK LIKE IT WAS FULLY FIRED. IN THE MIDDLE OF THE RELOAD STAPLES WERE TRYING TO COME OUT AND THEN THE REST OF THE PUSH PINS THAT SHOW WHEN STAPLES HAVE BEEN FIRED WERE NOT VISIBLE AT THE DISTAL TIP OF THE RELOAD. VISUALLY, THE RELOAD LOOKS LIKE IT STARTED TO FIRE AND THEN WAS OPENED BEFORE IT COMPLETED THE FIRING. TO RESOLVE THE ISSUE, AN OPEN STAPLER WAS USED TO COMPLETE THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCOAST TUMOR RESECTION, IN STAPLING THE BRONCHUS, THE DEVICE STOPPED FIRING. THE SURGEON WAS NOT ABLE TO SEE IF THE I BEAM WAS ADVANCING AS IT WAS AN OPEN PROCEDURE. THE DEVICE STAPLED WITHOUT CUTTING THE TISSUE. AFTER THE CASE, VISUALLY, THE RELOAD DID NOT LOOK LIKE IT WAS FULLY FIRED. IN THE MIDDLE OF THE RELOAD STAPLES WERE TRYING TO COME OUT AND THEN THE REST OF THE PUSH PINS THAT SHOW WHEN STAPLES HAVE BEEN FIRED WERE NOT VISIBLE AT THE DISTAL TIP OF THE RELOAD. VISUALLY, THE RELOAD LOOKS LIKE IT STARTED TO FIRE AND THEN WAS OPENED BEFORE IT COMPLETED THE FIRING. TO RESOLVE THE ISSUE, AN OPEN STAPLER WAS USED TO COMPLETE THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
02-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | ABSORBATACK | Loss of or Failure to Bond | A040102 | Bowel Perforation | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH (IPOM) TECHNIQUE FOR HERNIA PROCEDURE, WHILE ON MESH FIXATION, THE TACKS DID NOT HOLD AND HAD REMAINED LOOSE IN THE SITUS RESULTED TO AN INTESTINAL PERFORATION. THE ISSUE RESULTED TO AN EXTENDED SURGICAL TIME AND HOSPITALIZATION, A SECONDARY PROCEDURE-PARTIAL BOWEL RESECTION, TISSUE LOSS AND DAMAGE.
02-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | ABSORBATACK | Loss of or Failure to Bond | A040102 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH (IPOM) TECHNIQUE FOR HERNIA PROCEDURE, WHILE ON MESH FIXATION, THE TACKS DID NOT HOLD AND HAD REMAINED LOOSE IN THE SITUS RESULTED TO AN INTESTINAL PERFORATION. THE ISSUE RESULTED TO AN EXTENDED SURGICAL TIME AND HOSPITALIZATION, A SECONDARY PROCEDURE-PARTIAL BOWEL RESECTION, TISSUE LOSS AND DAMAGE.
02-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | ABSORBATACK | Detachment of Device or Device Component | A0501 | Bowel Perforation | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH (IPOM) TECHNIQUE FOR HERNIA PROCEDURE, WHILE ON MESH FIXATION, THE TACKS DID NOT HOLD AND HAD REMAINED LOOSE IN THE SITUS RESULTED TO AN INTESTINAL PERFORATION. THE ISSUE RESULTED TO AN EXTENDED SURGICAL TIME AND HOSPITALIZATION, A SECONDARY PROCEDURE-PARTIAL BOWEL RESECTION, TISSUE LOSS AND DAMAGE.
02-11-2022 | Injury | covidien lp llc north haven | 03-01-2023 | ABSORBATACK | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH (IPOM) TECHNIQUE FOR HERNIA PROCEDURE, WHILE ON MESH FIXATION, THE TACKS DID NOT HOLD AND HAD REMAINED LOOSE IN THE SITUS RESULTED TO AN INTESTINAL PERFORATION. THE ISSUE RESULTED TO AN EXTENDED SURGICAL TIME AND HOSPITALIZATION, A SECONDARY PROCEDURE-PARTIAL BOWEL RESECTION, TISSUE LOSS AND DAMAGE.
07-12-2022 | Malfunction | ethicon endo-surgery | 03-01-2023 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AND UNKNOWN PROCEDURE, SOME PARTS OF STAPLES WERE NOT FIRED AFTER FIRING THE DEVICE. NO HARM ON THE PATIENT. THERE WERE NO PATIENT CONSEQUENCES.
08-12-2022 | Malfunction | ethicon endo-surgery | 03-01-2023 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE DID NOT FIRE.
26-10-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN ATTEMPTED TO FIRE THE DEVICE, THE HANDLE BROKE. DURING THE ARTICULATION, THE PART OF JOINT CONNECTED TO THE RELOAD WAS BROKEN. THE DEVICE COULD NOT BE FIRED. A NEW HANDLE FROM A DIFFERENT MANUFACTURER WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
01-01-2021 | Injury | ethicon | 03-01-2023 | SECURESTRAP UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Hernia | Event Description: TITLE: VARIABLES IN THE DANISH HERNIA DATABASES: INGUINAL AND VENTRAL THE AIM OF THE CURRENT PAPER WAS TO GIVE AN UPDATE ON THE REGISTERED PARAMETERS IN THE TWO DANISH HERNIA DATABASES, THE DANISH INGUINAL HERNIA DATABASE AND THE DANISH VENTRAL HERNIA DATABASE, AS WELL AS GIVE AN OVERVIEW OF THE AVAILABLE DATA. DURING THE PERIOD JANUARY 1, 1998 TO DECEMBER 31, 2019, A TOTAL OF 181,715 PATIENTS WERE REGISTERED IN THE TWO DATABASES. THE DANISH INGUINAL DATABASE WAS LAUNCHED IN 1997, AND THE DANISH VENTRAL HERNIA DATABASE IN 2007. FOR INGUINAL HERNIA REPAIR THROUGH THE LIFESPAN OF THE DATABASE, 22% WERE DONE BY LAPAROSCOPY AND 62% BY LICHTENSTEIN REPAIR. VARIABLES IN THE DANISH INGUINAL HERNIA DATABASE INCLUDES: IF TACKS, SPECIFY: ABSORBATACK, OPTIFIX, PROTACK, CAPSURE, RELIATACK, SECURESTRAP, OTHER (TEXT). IF SUTURE, SPECIFY: NOT SELF-LOCKING, V-LOC, STRATAFIX, OTHER SELF-LOCKING. VARIABLES IN THE DANISH VENTRAL HERNIA DATABASE INCLUDES: IF SUTURE CLOSURE, SPECIFY: (I) MONO- OR MULTIFILAMENTOUS SUTURE. (II) QUICKLY RESORBABLE (SPECIFY VICRYL, DEXON, MONOCRYL), SLOWLY RESORBABLE (SPECIFY PDS, MAXON, MONOMAX), NON-RESORBABLE [SPECIFY PROLENE, NYLON, ETHIBOND, MERSILENE, OTHER (TEXT)]. (III) NON-LOCKING OR LOCKING SUTURE (SPECIFY V-LOC, STRATAFIX, OTHER). IF OTHER TYPE OF MESH, SPECIFY MESH (TEXT). IF SUTURE, SPECIFY A. MONOFILAMENTOUS OR MULTIFILAMENTOUS, B. QUICKY RESTORABLE (SPECIFY VICRYL, DEXON, MONOCRYL), SLOWLY RESORBABLE (SPECIFY PDS, MAXON, MONOMAX), NONRESORBABLE (SPECIFY PROLENE, NYLON, ETHIBOND, MERSILENE, OTHER (TEXT)), C.NON-LOCKING OR LOCKING SUTURE (SPECIFY V-LOC, STRATAFIX, OTHER). IF TACKS, SPECIFY: ABSORBABLE, NONABSORBABLE. IF TACKS, SPECIFY: ABSORBATACK, CAPSURE, OPTIFIX, PROTACK, RELIATACK, SECURESTRAP, OTHER (TEXT). REPORTED COMPLICATIONS INCLUDE RECURRENCE REQUIRING REOPERATION (N=?). IN CONCLUSION, THE TWO DANISH HERNIA DATABASES REGISTER ALL HERNIA REPAIRS IN DENMARK WHERE IT IS MANDATORY FOR SURGEONS TO ENTER OPERATIVE DATA AFTER EACH OPERATION. VARIABLES HAVE BEEN UPDATED SEVERAL TIMES OVER THE YEARS WHEN NEW SCIENTIFIC QUESTIONS NEEDED ANSWERS. NATIONAL DATABASES GIVE THE OPPORTUNITY TO FOLLOW DEVELOPMENT IN DAILY CLINICAL PRACTICE AND TO ANSWER IMPORTANT SCIENTIFIC QUESTIONS.
14-12-2022 | Malfunction | us surgical puerto rico | 03-01-2023 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON PERFORM FULL, COMPLETE SQUEEZES OF THE HANDLE. THE HANDLE DID NOT REMAIN IN THE CLOSED POSITION FOLLOWING THE SECOND COMPLETE SQUEEZE OF THE HANDLE. REPLACED A NEW PRODUCT TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN LAP SIGMOID RESECTION PROCEDURE THAT THEY INSERTED THE STAPLER ACCORDING TO THE IFU. THE MEASUREMENT INDICATOR LIGHT WAS ON. THEY REMOVED THE SAFETY AND WHEN THEY ENGAGED THE ENERGY TRIGGER, NOTHING HAPPENED. THERE WAS NO MOTOR SOUND OR STAPLER ADVANCE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
11-10-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED HALF. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ESOPHAGEAL SURGERY, THE DEVICE COULD NOT FIRE. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2022 | Injury | ethicon endo-surgery | 04-01-2023 | UNK_ECHELON CIRCULAR POWERED STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A SIGMOID COLECTOMY PROCEDURE, THE PATIENT EXPERIENCED A POST-OP ANASTOMOTIC LEAK AND REQUIRED REOPERATION ON (B)(6) 2022. THE PATIENT WAS GIVEN A PERMANENT COLOSTOMY AND IS STABLE. THE ECHELON POWERED CIRCULAR STAPLER WAS USED DURING THE COLECTOMY PROCEDURE TO CREATE THE ANASTOMOSIS; THE EXACT DIAMETER OF THE DEVICE IS UNKNOWN. THE SURGEON CLAIMS THE COLECTOMY PROCEDURE WAS SUCCESSFUL AND THERE WERE NO ISSUES WITH THE DEVICE. THE SOURCE AND REASON FOR THE ANASTOMOTIC LEAK ARE UNKNOWN.
08-12-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | CIRCULAR MECH STD 29MM | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ANVIL OF THE CIRCULAR STAPLER WOULD NOT CONNECT TO THE TROCAR SPIKE DURING THE PROCEDURE. THE PRODUCT WAS NOT USED IN A CLINICAL TRIAL. A NEW CIRCULAR STAPLER WAS OPENED AND IT WORKED FINE. THERE WERE NO CONSEQUENCE TO THE PATIENT DUE TO THE EVENT. THE SURGERY WAS NOT PROLONGED DUE TO THE EVENT.
14-12-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | ENDO ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE COULD NOT FINISH THE FIRING SEQUENCE.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATECTOMY, IN CLOSING THE BLOOD VESSEL, ON TWO HANDLES, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE DEVICES DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATECTOMY, IN CLOSING THE BLOOD VESSEL, ON TWO HANDLES, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE DEVICES DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATECTOMY, IN CLOSING THE BLOOD VESSEL, ON TWO HANDLES, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE DEVICES DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOBECTOMY, WHEN FIRING THE STAPLER WITH THE BLACK RELOAD, IT MADE A CRACKING SOUND TWICE AND FAILED TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOBECTOMY, WHEN FIRING THE STAPLER WITH THE BLACK RELOAD, IT MADE A CRACKING SOUND TWICE AND FAILED TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-11-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A RESECTION OF LUNG PARENCHYMA LOBECTOMY, ON THE FIRST BLACK RELOAD, THE STAPLER DID NOT FIRE. A PURPLE RELOAD WAS USED BUT ALSO FAILED. FAILURE DESPITE SEVERAL ATTEMPTS TO UNDO AND REPLACE THE RELOAD. ONLY THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE.
15-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Component Missing | A020602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CARTRIDGE JAMMED ON TWO RELOADS. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CARTRIDGE JAMMED ON TWO RELOADS. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CARTRIDGE JAMMED ON TWO RELOADS. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
26-08-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE HANDLE WAS BROKEN DURING USE. LOUD NOISE WAS HEARD WHEN BREAKING. A NEW DEVICE WAS USED IN ORDER TO RESOLVE THE ISSUE. PLI 10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE.
26-08-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE HANDLE WAS BROKEN DURING USE. LOUD NOISE WAS HEARD WHEN BREAKING. A NEW DEVICE WAS USED IN ORDER TO RESOLVE THE ISSUE. PLI 10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE.
15-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Component Missing | A020602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CARTRIDGE JAMMED ON TWO RELOADS. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CARTRIDGE JAMMED ON TWO RELOADS. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE CARTRIDGE JAMMED ON TWO RELOADS. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
12-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHILE AT GASTRIC PARENCHYMA RESECTION, A DEFECT OCCURRED WITH THE SECOND FIRING OF STAPLER, IT WAS FIRED AS USUAL, BUT COULD NOT PERFORM THE STAPLING FROM MIDWAY, THE KNIFE WAS ALSO NOT RUNNING AND STAPLING NEAR THE CUT LINE WAS NOT DONE IN ALL THREE ROWS. IT WAS NOTED THAT THERE WAS NO ERROR DISPLAYED ON DISPLAY AND THERE WERE NO PROBLEMS WITH THE FIRST FIRING OF STAPLER. AFTER CONNECTING THE 3RD RELOAD AND RE-STAPLING, THE STAPLING WAS ABLE TO BE PERFORMED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SEGMENT SURGERY, WHILE STAPLING THE LARGE VESSELS, THE STAPLE LINE FOR BOTH RELOAD HAD POOR FORMATION AND WAS OOZING BLOOD. THE PULMONARY ARTERY WAS DAMAGED AND HAD TO OPEN THE CHEST AND BE REPAIRED BY A CARDIAC SURGEON. THERE WAS A BLOOD LOSS OF 2500 ML. THE INCISION WAS EXTENDED FOR MORE THAN ONE INCH AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. AN ADDITIONAL OPERATION WILL BE PERFORMED TO CORRECT THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SEGMENT SURGERY, WHILE STAPLING THE LARGE VESSELS, THE STAPLE LINE FOR BOTH RELOAD HAD POOR FORMATION AND WAS OOZING BLOOD. THE PULMONARY ARTERY WAS DAMAGED AND HAD TO OPEN THE CHEST AND BE REPAIRED BY A CARDIAC SURGEON. THERE WAS A BLOOD LOSS OF 2500 ML. THE INCISION WAS EXTENDED FOR MORE THAN ONE INCH AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. AN ADDITIONAL OPERATION WILL BE PERFORMED TO CORRECT THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SEGMENT SURGERY, WHILE STAPLING THE LARGE VESSELS, THE STAPLE LINE FOR BOTH RELOAD HAD POOR FORMATION AND WAS OOZING BLOOD. THE PULMONARY ARTERY WAS DAMAGED AND HAD TO OPEN THE CHEST AND BE REPAIRED BY A CARDIAC SURGEON. THERE WAS A BLOOD LOSS OF 2500 ML. THE INCISION WAS EXTENDED FOR MORE THAN ONE INCH AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. AN ADDITIONAL OPERATION WILL BE PERFORMED TO CORRECT THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SEGMENT SURGERY, WHILE STAPLING THE LARGE VESSELS, THE STAPLE LINE FOR BOTH RELOAD HAD POOR FORMATION AND WAS OOZING BLOOD. THE PULMONARY ARTERY WAS DAMAGED AND HAD TO OPEN THE CHEST AND BE REPAIRED BY A CARDIAC SURGEON. THERE WAS A BLOOD LOSS OF 2500 ML. THE INCISION WAS EXTENDED FOR MORE THAN ONE INCH AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. AN ADDITIONAL OPERATION WILL BE PERFORMED TO CORRECT THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SEGMENT SURGERY, WHILE STAPLING THE LARGE VESSELS, THE STAPLE LINE FOR BOTH RELOAD HAD POOR FORMATION AND WAS OOZING BLOOD. THE PULMONARY ARTERY WAS DAMAGED AND HAD TO OPEN THE CHEST AND BE REPAIRED BY A CARDIAC SURGEON. THERE WAS A BLOOD LOSS OF 2500 ML. THE INCISION WAS EXTENDED FOR MORE THAN ONE INCH AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. AN ADDITIONAL OPERATION WILL BE PERFORMED TO CORRECT THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SEGMENT SURGERY, WHILE STAPLING THE LARGE VESSELS, THE STAPLE LINE FOR BOTH RELOAD HAD POOR FORMATION AND WAS OOZING BLOOD. THE PULMONARY ARTERY WAS DAMAGED AND HAD TO OPEN THE CHEST AND BE REPAIRED BY A CARDIAC SURGEON. THERE WAS A BLOOD LOSS OF 2500 ML. THE INCISION WAS EXTENDED FOR MORE THAN ONE INCH AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. AN ADDITIONAL OPERATION WILL BE PERFORMED TO CORRECT THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SEGMENT SURGERY, WHILE STAPLING THE LARGE VESSELS, THE STAPLE LINE FOR BOTH RELOAD HAD POOR FORMATION AND WAS OOZING BLOOD. THE PULMONARY ARTERY WAS DAMAGED AND HAD TO OPEN THE CHEST AND BE REPAIRED BY A CARDIAC SURGEON. THERE WAS A BLOOD LOSS OF 2500 ML. THE INCISION WAS EXTENDED FOR MORE THAN ONE INCH AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. AN ADDITIONAL OPERATION WILL BE PERFORMED TO CORRECT THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SEGMENT SURGERY, WHILE STAPLING THE LARGE VESSELS, THE STAPLE LINE FOR BOTH RELOAD HAD POOR FORMATION AND WAS OOZING BLOOD. THE PULMONARY ARTERY WAS DAMAGED AND HAD TO OPEN THE CHEST AND BE REPAIRED BY A CARDIAC SURGEON. THERE WAS A BLOOD LOSS OF 2500 ML. THE INCISION WAS EXTENDED FOR MORE THAN ONE INCH AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. AN ADDITIONAL OPERATION WILL BE PERFORMED TO CORRECT THE ISSUE.
08-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | EEA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION (LAR), THE STAPLE WAS MALFORMED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SEPTATION, WHILE STAPLING THE STOMACH, THE HANDLE LOCKED AFTER THE SECOND SHOT, HENCE COULD NOT BE USED. THE RELOAD WAS CHANGED AND WAS CONNECTED TO THE HANDLE BUT THE HANDLE DID NOT GIVE A GREEN LIGHT SIGNAL. THE ADAPTER WAS THEN CHANGED BUT STILL THE DEFECT PERSISTED. IT WAS ALSO NOTED THAT THE HANDLE, EVEN WITHOUT THE SHELL CONTINUED IN SHOOTING MODE. THE MANEUVER OF RESTARTING THE HANDLE WAS CARRIED OUT TO RETURN TO NORMAL OPERATION. AFTER LOCKING THE HANDLE ANDCHANGING THE STEM, A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SEPTATION, WHILE STAPLING THE STOMACH, THE HANDLE LOCKED AFTER THE SECOND SHOT, HENCE COULD NOT BE USED. THE RELOAD WAS CHANGED AND WAS CONNECTED TO THE HANDLE BUT THE HANDLE DID NOT GIVE A GREEN LIGHT SIGNAL. THE ADAPTER WAS THEN CHANGED BUT STILL THE DEFECT PERSISTED. IT WAS ALSO NOTED THAT THE HANDLE, EVEN WITHOUT THE SHELL CONTINUED IN SHOOTING MODE. THE MANEUVER OF RESTARTING THE HANDLE WAS CARRIED OUT TO RETURN TO NORMAL OPERATION. AFTER LOCKING THE HANDLE ANDCHANGING THE STEM, A MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN COLECTOMY, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A UROLOGY PROCEDURE, WHEN THE HANDLE WAS SQUEEZED, THE TIP WAS UNRESPONSIVE AND COULD NOT BE FIRED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC JEJUNAL BYPASS PROCEDURE, WHILE RESECTIONING THE STOMACH AND PERFORMING GASTRIC JEJUNAL SIDE-TO-SIDE ANASTOMOSIS, AT THE TIME OF PERFORMING LATERAL ANASTOMOSIS WHEN CREATING THE Y-LEG, WHEN THE FIRING HAD STOPPED ON THE WAY OF FIRING OPERATION, THE SURGEON RELEASED THE DEVICE AND IT WAS OBSERVED THAT THE TIP OF THE RELOAD WAS INCOMPLETE. THE ANASTOMOSIS BECAME SMALLER, BUT THE LUMEN WAS SECURED, SO THE INSERTION OPENING WAS CLOSED WITHOUT ANY ADDITIONAL RESECTION AND NO ADDITIONAL PROCEDURE WAS PERFORMED. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC JEJUNAL BYPASS PROCEDURE, WHILE RESECTIONING THE STOMACH AND PERFORMING GASTRIC JEJUNAL SIDE-TO-SIDE ANASTOMOSIS, AT THE TIME OF PERFORMING LATERAL ANASTOMOSIS WHEN CREATING THE Y-LEG, WHEN THE FIRING HAD STOPPED ON THE WAY OF FIRING OPERATION, THE SURGEON RELEASED THE DEVICE AND IT WAS OBSERVED THAT THE TIP OF THE RELOAD WAS INCOMPLETE. THE ANASTOMOSIS BECAME SMALLER, BUT THE LUMEN WAS SECURED, SO THE INSERTION OPENING WAS CLOSED WITHOUT ANY ADDITIONAL RESECTION AND NO ADDITIONAL PROCEDURE WAS PERFORMED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | TA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN EN BLOC RIB RESECTION, THE DEVICE DID NOT FIRE. THERE WAS AN ACTIVE HEMORRHAGE, BLOOD LOSS OF 500CC OR MORE. THE VESSELS AND BRONCHUS WERE HAND LIGATED AND THE LUNG WAS OVERSEWED TO RESOLVE THE ISSUE. NO BLOOD PRODUCTS WERE NEEDED. IT WAS STATED THAT THE ISSUE OCCURRED ON THE SUCCEEDING RELOAD. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AS A RESULT OF THE EVENT.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | TA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN EN BLOC RIB RESECTION, THE DEVICE DID NOT FIRE. THERE WAS AN ACTIVE HEMORRHAGE, BLOOD LOSS OF 500CC OR MORE. THE VESSELS AND BRONCHUS WERE HAND LIGATED AND THE LUNG WAS OVERSEWED TO RESOLVE THE ISSUE. NO BLOOD PRODUCTS WERE NEEDED. IT WAS STATED THAT THE ISSUE OCCURRED ON THE SUCCEEDING RELOAD. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AS A RESULT OF THE EVENT.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN EN BLOC RIB RESECTION, THE DEVICE DID NOT FIRE. THERE WAS AN ACTIVE HEMORRHAGE, BLOOD LOSS OF 500CC OR MORE. THE VESSELS AND BRONCHUS WERE HAND LIGATED AND THE LUNG WAS OVERSEWED TO RESOLVE THE ISSUE. NO BLOOD PRODUCTS WERE NEEDED. IT WAS STATED THAT THE ISSUE OCCURRED ON THE SUCCEEDING RELOAD. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AS A RESULT OF THE EVENT.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN EN BLOC RIB RESECTION, THE DEVICE DID NOT FIRE. THERE WAS AN ACTIVE HEMORRHAGE, BLOOD LOSS OF 500CC OR MORE. THE VESSELS AND BRONCHUS WERE HAND LIGATED AND THE LUNG WAS OVERSEWED TO RESOLVE THE ISSUE. NO BLOOD PRODUCTS WERE NEEDED. IT WAS STATED THAT THE ISSUE OCCURRED ON THE SUCCEEDING RELOAD. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AS A RESULT OF THE EVENT.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Entrapment of Device | A150208 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT UPPER LOBECTOMY, DURING SECOND FIRING IN A RIGHT UPPER LOBE PARTIAL RESECTION, THE TISSUE WAS CLAMPED NORMALLY, BUT THE STAPLING STOPPED DURING FIRING. THE TISSUE COULD NOT BE RESECTED OR STAPLED. STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL SIDE. THE STAPLES FELL INTO PATIENT CAVITY DUE TO POOR FORMATION. THE SURGEON COULD NOT DO A RESECTION SO A NEW RELOAD WAS USED TO RE-STAPLE THE RESECTION.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT UPPER LOBECTOMY, DURING SECOND FIRING IN A RIGHT UPPER LOBE PARTIAL RESECTION, THE TISSUE WAS CLAMPED NORMALLY, BUT THE STAPLING STOPPED DURING FIRING. THE TISSUE COULD NOT BE RESECTED OR STAPLED. STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL SIDE. THE STAPLES FELL INTO PATIENT CAVITY DUE TO POOR FORMATION. THE SURGEON COULD NOT DO A RESECTION SO A NEW RELOAD WAS USED TO RE-STAPLE THE RESECTION.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Misfire | A050502 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT UPPER LOBECTOMY, DURING SECOND FIRING IN A RIGHT UPPER LOBE PARTIAL RESECTION, THE TISSUE WAS CLAMPED NORMALLY, BUT THE STAPLING STOPPED DURING FIRING. THE TISSUE COULD NOT BE RESECTED OR STAPLED. STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL SIDE. THE STAPLES FELL INTO PATIENT CAVITY DUE TO POOR FORMATION. THE SURGEON COULD NOT DO A RESECTION SO A NEW RELOAD WAS USED TO RE-STAPLE THE RESECTION.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT UPPER LOBECTOMY, DURING SECOND FIRING IN A RIGHT UPPER LOBE PARTIAL RESECTION, THE TISSUE WAS CLAMPED NORMALLY, BUT THE STAPLING STOPPED DURING FIRING. THE TISSUE COULD NOT BE RESECTED OR STAPLED. STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL SIDE. THE STAPLES FELL INTO PATIENT CAVITY DUE TO POOR FORMATION. THE SURGEON COULD NOT DO A RESECTION SO A NEW RELOAD WAS USED TO RE-STAPLE THE RESECTION.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT UPPER LOBECTOMY, DURING SECOND FIRING IN A RIGHT UPPER LOBE PARTIAL RESECTION, THE TISSUE WAS CLAMPED NORMALLY, BUT THE STAPLING STOPPED DURING FIRING. THE TISSUE COULD NOT BE RESECTED OR STAPLED. STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL SIDE. THE STAPLES FELL INTO PATIENT CAVITY DUE TO POOR FORMATION. THE SURGEON COULD NOT DO A RESECTION SO A NEW RELOAD WAS USED TO RE-STAPLE THE RESECTION.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT UPPER LOBECTOMY, DURING SECOND FIRING IN A RIGHT UPPER LOBE PARTIAL RESECTION, THE TISSUE WAS CLAMPED NORMALLY, BUT THE STAPLING STOPPED DURING FIRING. THE TISSUE COULD NOT BE RESECTED OR STAPLED. STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL SIDE. THE STAPLES FELL INTO PATIENT CAVITY DUE TO POOR FORMATION. THE SURGEON COULD NOT DO A RESECTION SO A NEW RELOAD WAS USED TO RE-STAPLE THE RESECTION.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT UPPER LOBECTOMY, DURING SECOND FIRING IN A RIGHT UPPER LOBE PARTIAL RESECTION, THE TISSUE WAS CLAMPED NORMALLY, BUT THE STAPLING STOPPED DURING FIRING. THE TISSUE COULD NOT BE RESECTED OR STAPLED. STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL SIDE. THE STAPLES FELL INTO PATIENT CAVITY DUE TO POOR FORMATION. THE SURGEON COULD NOT DO A RESECTION SO A NEW RELOAD WAS USED TO RE-STAPLE THE RESECTION.
12-12-2022 | Injury | covidien lp llc north haven | 04-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT UPPER LOBECTOMY, DURING SECOND FIRING IN A RIGHT UPPER LOBE PARTIAL RESECTION, THE TISSUE WAS CLAMPED NORMALLY, BUT THE STAPLING STOPPED DURING FIRING. THE TISSUE COULD NOT BE RESECTED OR STAPLED. STAPLE LINE WAS INCOMPLETE AT THE PROXIMAL SIDE. THE STAPLES FELL INTO PATIENT CAVITY DUE TO POOR FORMATION. THE SURGEON COULD NOT DO A RESECTION SO A NEW RELOAD WAS USED TO RE-STAPLE THE RESECTION.
12-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PARTIAL PULMONARY RESECTION, AT PULMONARY PARENCHYMA APPROACH, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE PRESSED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
08-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON TRIED TO FIRE BUT COULD NOT, NO STAPLES WERE DEPLOYED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHEN THE SURGEON WANTED TO CUT THE GASTRIC TISSUE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS PRESSED AND THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN TOTAL GASTRECTOMY, THE DEVICE DID NOT FIRE - THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. BOTH THE PRELOADED CARTRIDGE AND THE SEPARATE RELOAD WERE UNABLE TO FIRE. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | UNKNOWN GIA PRODUCT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN TOTAL GASTRECTOMY, THE DEVICE DID NOT FIRE - THE BLACK PUSHER THUMB PIECE COULD NOT BE MOVED AT ALL. BOTH THE PRELOADED CARTRIDGE AND THE SEPARATE RELOAD WERE UNABLE TO FIRE. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
14-01-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | ECHELON 3000 60MM COMPACT | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE THAT THEY FIRED TWO FIRINGS ON THE STOMACH WITH A BLUE RELOAD. THE SURGEON STATED THAT THE STAPLE LINES LOOKED A LITTLE OOZY. SUBSEQUENTLY AT THE END OF THE CASE, THEY LOOKED AT THE STAPLE LINES AGAIN AND IT LOOKED LIKE THERE WAS A LEAK. THEY FIRED A COVIDIEN STAPLER TO REMOVE OUR PORTION OF THE STAPLE LINE THAT WAS LEAKING. THIS ADDED 30 MINUTES TO THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-01-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | ECHELON 3000 60MM COMPACT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE THAT THEY FIRED TWO FIRINGS ON THE STOMACH WITH A BLUE RELOAD. THE SURGEON STATED THAT THE STAPLE LINES LOOKED A LITTLE OOZY. SUBSEQUENTLY AT THE END OF THE CASE, THEY LOOKED AT THE STAPLE LINES AGAIN AND IT LOOKED LIKE THERE WAS A LEAK. THEY FIRED A COVIDIEN STAPLER TO REMOVE OUR PORTION OF THE STAPLE LINE THAT WAS LEAKING. THIS ADDED 30 MINUTES TO THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-01-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | ECHELON 3000 60MM COMPACT | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE THAT THEY FIRED TWO FIRINGS ON THE STOMACH WITH A BLUE RELOAD. THE SURGEON STATED THAT THE STAPLE LINES LOOKED A LITTLE OOZY. SUBSEQUENTLY AT THE END OF THE CASE, THEY LOOKED AT THE STAPLE LINES AGAIN AND IT LOOKED LIKE THERE WAS A LEAK. THEY FIRED A COVIDIEN STAPLER TO REMOVE OUR PORTION OF THE STAPLE LINE THAT WAS LEAKING. THIS ADDED 30 MINUTES TO THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-01-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | ECHELON 3000 60MM COMPACT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE THAT THEY FIRED TWO FIRINGS ON THE STOMACH WITH A BLUE RELOAD. THE SURGEON STATED THAT THE STAPLE LINES LOOKED A LITTLE OOZY. SUBSEQUENTLY AT THE END OF THE CASE, THEY LOOKED AT THE STAPLE LINES AGAIN AND IT LOOKED LIKE THERE WAS A LEAK. THEY FIRED A COVIDIEN STAPLER TO REMOVE OUR PORTION OF THE STAPLE LINE THAT WAS LEAKING. THIS ADDED 30 MINUTES TO THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
15-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HERNIA REPAIR, WHEN ANCHORING MESH, THE TACKER FIRED FOUR TACKS AND THEN JAMMED. THERE WAS SMALL SURGICAL TIME EXTENSION TO REPLACE THE TACKER. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | covidien lp llc north haven | 04-01-2023 | ABSORBATACK | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC HERNIA REPAIR, WHEN ANCHORING MESH, THE TACKER FIRED FOUR TACKS AND THEN JAMMED. THERE WAS SMALL SURGICAL TIME EXTENSION TO REPLACE THE TACKER. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND THERE WAS SLIGHT OOZING. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND THERE WAS SLIGHT OOZING. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, IN STAPLING, THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. HOWEVER, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, THE JAWS DID NOT OPEN; THE DEVICE LOCKED ON TISSUE. THE SURGEON MANIPULATED THE TISSUE TO REMOVE IT FROM THE JAWS AND THE STAPLES WERE NOT FIRED. THE DEVICE STILL IN STUCK POSITION AND A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND THERE WAS SLIGHT OOZING . SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2022 | Malfunction | ethicon endo-surgery | 04-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND THERE WAS SLIGHT OOZING . SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, THE JAWS DID NOT OPEN; THE DEVICE LOCKED ON TISSUE. THE DEVICE WAS STILL IN STUCK POSITION AND A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RIGHT HEMIHEPATECTOMY, WHEN THE SURGEON UNDERWENT LOBAR RESECTION, IT WAS FOUND THAT THE ADJUSTABLE EXCITATION KNOB COULD NOT BE PUSHED FORWARD AFTER HALF OF IT WAS PUSHED. THE RELOAD WAS REPLACED TO TRY FIRING AGAIN, AND STILL THE ENCAPSULANT COULD NOT BE PUSHED. A NEW MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 04-01-2023 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN RIGHT HEMIHEPATECTOMY, WHEN THE SURGEON UNDERWENT LOBAR RESECTION, IT WAS FOUND THAT THE ADJUSTABLE EXCITATION KNOB COULD NOT BE PUSHED FORWARD AFTER HALF OF IT WAS PUSHED. THE RELOAD WAS REPLACED TO TRY FIRING AGAIN, AND STILL THE ENCAPSULANT COULD NOT BE PUSHED. A NEW MANUAL STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2022 | Injury | us surgical puerto rico | 04-01-2023 | EEA | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ESOPHAGUS RESECTION, IT WAS DIFFICULT TO REMOVE THE STAPLER FROM THE CAVITY. THE ANVIL DID NOT RETRACT AFTER THE ANASTOMOSIS. IT WAS IMPOSSIBLE TO MOVE THE STAPLER TO THE ESOPHAGUS, SO THE ANASTOMOSIS WAS DESTROYED. IT WAS ALSO MENTIONED THAT THE ANVIL DOES NOT TILT AFTER FIRING. THE SURGEON SUTURED THE ANASTOMOSIS TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AS A RESULT.
08-12-2022 | Injury | us surgical puerto rico | 04-01-2023 | EEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ESOPHAGUS RESECTION, IT WAS DIFFICULT TO REMOVE THE STAPLER FROM THE CAVITY. THE ANVIL DID NOT RETRACT AFTER THE ANASTOMOSIS. IT WAS IMPOSSIBLE TO MOVE THE STAPLER TO THE ESOPHAGUS, SO THE ANASTOMOSIS WAS DESTROYED. IT WAS ALSO MENTIONED THAT THE ANVIL DOES NOT TILT AFTER FIRING. THE SURGEON SUTURED THE ANASTOMOSIS TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AS A RESULT.
08-12-2022 | Injury | us surgical puerto rico | 04-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ESOPHAGUS RESECTION, IT WAS DIFFICULT TO REMOVE THE STAPLER FROM THE CAVITY. THE ANVIL DID NOT RETRACT AFTER THE ANASTOMOSIS. IT WAS IMPOSSIBLE TO MOVE THE STAPLER TO THE ESOPHAGUS, SO THE ANASTOMOSIS WAS DESTROYED. IT WAS ALSO MENTIONED THAT THE ANVIL DOES NOT TILT AFTER FIRING. THE SURGEON SUTURED THE ANASTOMOSIS TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AS A RESULT.
08-12-2022 | Injury | us surgical puerto rico | 04-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ESOPHAGUS RESECTION, IT WAS DIFFICULT TO REMOVE THE STAPLER FROM THE CAVITY. THE ANVIL DID NOT RETRACT AFTER THE ANASTOMOSIS. IT WAS IMPOSSIBLE TO MOVE THE STAPLER TO THE ESOPHAGUS, SO THE ANASTOMOSIS WAS DESTROYED. IT WAS ALSO MENTIONED THAT THE ANVIL DOES NOT TILT AFTER FIRING. THE SURGEON SUTURED THE ANASTOMOSIS TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE AS A RESULT.
05-12-2022 | Injury | intuitive surgical | 04-01-2023 | SUREFORM | Misfire | A050502 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS (ROUX-EN-Y) PROCEDURE, THERE WAS A TISSUE PUSHOUT EVENT INVOLVING A BLUE SUREFORM 60 RELOAD AND A SUREFORM 60 STAPLER INSTRUMENT. PER THE SURGEON, ABOUT HALFWAY THROUGH THE SECOND FIRING SEQUENCE, THE TARGET TISSUE WAS PUSHED FORWARD/DRAGGED, RESULTING IN MISSING STAPLES IN THE STAPLE LINE AND "NOT A LOT" OF BLEEDING IN THE AREA OF THE MISSHAPED STAPLES. THE SURGEON CREATED A NEW STAPLE LINE MEDIAL TO THE INCOMPLETE LINE. AT THE TIME OF THE EVENT, THE SURGEON WAS CREATING THE GASTRIC POUCH; THE TARGET TISSUE WAS NORMAL GASTRIC TISSUE, WITH NO CALCIFICATIONS, NO EXPOSURE TO RADIATION OR CHEMOTHERAPY, AND NO TISSUE TENSION OR BUNCHING. THE BLUE SUREFORM 60 STAPLER RELOAD IS THE SURGEON'S TYPICAL CHOICE FOR THIS PART OF THE PROCEDURE. THE INSTRUMENT WAS INSPECTED PRIOR TO USE, WITH NO DAMAGE OBSERVED. NO ERROR MESSAGES WERE GENERATED AT THE TIME OF THE EVENT, AND THERE WERE NO CLAMPING ISSUES PRIOR TO THE FIRING SEQUENCE IN QUESTION. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE, THERE WERE NO OBSTRUCTIONS BETWEEN THE JAWS OF THE INSTRUMENT WHEN THE ISSUE OCCURRED, AND BUTTRESS MATERIAL WAS NOT USED. THE PROCEDURE WAS COMPLETED ROBOTICALLY.
05-12-2022 | Injury | intuitive surgical | 04-01-2023 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS (ROUX-EN-Y) PROCEDURE, THERE WAS A TISSUE PUSHOUT EVENT INVOLVING A BLUE SUREFORM 60 RELOAD AND A SUREFORM 60 STAPLER INSTRUMENT. PER THE SURGEON, ABOUT HALFWAY THROUGH THE SECOND FIRING SEQUENCE, THE TARGET TISSUE WAS PUSHED FORWARD/DRAGGED, RESULTING IN MISSING STAPLES IN THE STAPLE LINE AND "NOT A LOT" OF BLEEDING IN THE AREA OF THE MISSHAPED STAPLES. THE SURGEON CREATED A NEW STAPLE LINE MEDIAL TO THE INCOMPLETE LINE. AT THE TIME OF THE EVENT, THE SURGEON WAS CREATING THE GASTRIC POUCH; THE TARGET TISSUE WAS NORMAL GASTRIC TISSUE, WITH NO CALCIFICATIONS, NO EXPOSURE TO RADIATION OR CHEMOTHERAPY, AND NO TISSUE TENSION OR BUNCHING. THE BLUE SUREFORM 60 STAPLER RELOAD IS THE SURGEON'S TYPICAL CHOICE FOR THIS PART OF THE PROCEDURE. THE INSTRUMENT WAS INSPECTED PRIOR TO USE, WITH NO DAMAGE OBSERVED. NO ERROR MESSAGES WERE GENERATED AT THE TIME OF THE EVENT, AND THERE WERE NO CLAMPING ISSUES PRIOR TO THE FIRING SEQUENCE IN QUESTION. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE, THERE WERE NO OBSTRUCTIONS BETWEEN THE JAWS OF THE INSTRUMENT WHEN THE ISSUE OCCURRED, AND BUTTRESS MATERIAL WAS NOT USED. THE PROCEDURE WAS COMPLETED ROBOTICALLY.
03-11-2022 | Injury | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT 1 WEEK AFTER LAPAROSCOPIC ANTERIOR RESECTION, ANASTOMOTIC LEAKAGE OCCURRED SOON AFTER THE LIQUID FOOD WAS STARTED. IT WAS MINOR LEAK GRADE 2 AND IT DID NOT NEED COLOSTOMY. THE HYPERALIMENTATION WAS USED, AND THE PATIENT HAS BEEN FOLLOWED UP. IN THE INITIAL OPERATION, THE RECTUM RA RB WAS CUT WITH PSEE60A AND GOLD CARTRIDGE FOR 2 TIMES, AND CDH25A WAS USED TO ANASTOMOSE. DST RECONSTRUCTION WAS PERFORMED. THERE WAS NO PROBLEM TO THE DEVICE OR THE STAPLING DURING THE INITIAL OPERATION. THE DEVICE WAS USED BETWEEN COLON AND RECTUM. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Malfunction | ethicon endo-surgery | 05-01-2023 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: THE BLADE JAMMED HALFWAY THROUGH FIRING, THE FRONT MOUTH OF THE GUN BODY CANNOT BE OPENED, IT CAN ONLY BE DESTROYED BY FORCE DID THE EVENT OCCUR DURING STERILE PROCESSING? UNKNOWN, DID THE EVENT OCCUR DURING FIELD INSPECTION? UNKNOWN, DID THE EVENT OCCUR DURING INTERNAL SERVICE ACTIVITIES SUCH AS CALIBRATION? UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES: NO, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN, IS THE PATIENT PART OF A CLINICAL STUDY: UNKNOWN, (B)(4). DEVICE PROPERTY OF: NONE, DEVICE IN POSSESSION OF: NONE. BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST. TRUE.
01-01-2023 | Malfunction | ethicon endo-surgery | 05-01-2023 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: THE BLADE JAMMED HALFWAY THROUGH FIRING, THE FRONT MOUTH OF THE GUN BODY CANNOT BE OPENED, IT CAN ONLY BE DESTROYED BY FORCE DID THE EVENT OCCUR DURING STERILE PROCESSING? UNKNOWN, DID THE EVENT OCCUR DURING FIELD INSPECTION? UNKNOWN, DID THE EVENT OCCUR DURING INTERNAL SERVICE ACTIVITIES SUCH AS CALIBRATION? UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES: NO, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN, IS THE PATIENT PART OF A CLINICAL STUDY: UNKNOWN, (B)(4). DEVICE PROPERTY OF: NONE, DEVICE IN POSSESSION OF: NONE. BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST. TRUE.
01-01-2023 | Malfunction | ethicon endo-surgery | 05-01-2023 | FLEX 60 ARTICULATING | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: THE BLADE JAMMED HALFWAY THROUGH FIRING, THE FRONT MOUTH OF THE GUN BODY CANNOT BE OPENED, IT CAN ONLY BE DESTROYED BY FORCE DID THE EVENT OCCUR DURING STERILE PROCESSING? UNKNOWN, DID THE EVENT OCCUR DURING FIELD INSPECTION? UNKNOWN, DID THE EVENT OCCUR DURING INTERNAL SERVICE ACTIVITIES SUCH AS CALIBRATION? UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES: NO, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN, IS THE PATIENT PART OF A CLINICAL STUDY: UNKNOWN, (B)(4). DEVICE PROPERTY OF: NONE, DEVICE IN POSSESSION OF: NONE. BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST. TRUE.
28-10-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN SUBTOTAL COLECTOMY, END ILEOSTOMY PROCEDURE; ECHELON CONTOUR USED TO CUT COLON, COLON CUT, STAPLES NOT DEPLOYED. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE BLADE JAMMED HALFWAY THROUGH FIRING, THE FRONT MOUTH OF THE GUN BODY CANNOT BE OPENED, IT CAN ONLY BE DESTROYED BY FORCE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE BLADE JAMMED HALFWAY THROUGH FIRING, THE FRONT MOUTH OF THE GUN BODY CANNOT BE OPENED, IT CAN ONLY BE DESTROYED BY FORCE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | FLEX 60 ARTICULATING | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE BLADE JAMMED HALFWAY THROUGH FIRING, THE FRONT MOUTH OF THE GUN BODY CANNOT BE OPENED, IT CAN ONLY BE DESTROYED BY FORCE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN BARIATRIC PROCEDURE, AT THE TIME OF STAPLING, THE LOAD WAS LOCKED. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. THERE WAS ALSO A PROBLEM UNLOADING THE DEVICE. THE RELOAD WAS CHANGED AND WAS USED WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN BARIATRIC PROCEDURE, AT THE TIME OF STAPLING, THE LOAD WAS LOCKED. THE SURGEON WAS ABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE, HOWEVER, THE DEVICE DID NOT FIRE. THERE WAS ALSO A PROBLEM UNLOADING THE DEVICE. THE RELOAD WAS CHANGED AND WAS USED WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
24-11-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Break | A0401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, THE STAPLER FIRED BUT THERE WAS NO STAPLING DONE. IT WAS ALSO NOTED THAT THE RELOAD CRASHED. A NEW RELOAD WAS USED TO COMPLETE THE CASE.
24-11-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, THE STAPLER FIRED BUT THERE WAS NO STAPLING DONE. IT WAS ALSO NOTED THAT THE RELOAD CRASHED. A NEW RELOAD WAS USED TO COMPLETE THE CASE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. A NEW HANDLE WAS USED WITH THE SAME RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON. A NEW HANDLE WAS USED WITH THE SAME RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A COLORECTAL TRANSECTING/LIGATING SMALL COLON, THE STAPLER CUT UNTIL THE LAST PART OF SMALL BOWEL. THERE WAS NO PATIENT INJURY.
19-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY RESECTION PROCEDURE, WHILE RESECTING THE BLOOD VESSEL, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, EVEN WHEN THE GREEN BUTTON WAS PRESSED. THE HANDLE WAS TOO HARD TO SQUEEZE; THUS, THE KNIFE DID NOT ADVANCE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2022 | Malfunction | intuitive surgical | 05-01-2023 | SUREFORM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LUNG PROCEDURE, THE SUREFORM STAPLER RELOAD DID NOT FORM A COMPLETE STAPLE LINE. A BACKUP INSTRUMENT OF SAME KIND AND A DIFFERENT RELOAD WAS USED TO COMPLETE THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION: THE INSTRUMENT AND ACCESSORY WERE INSPECTED PRIOR TO USE AND NO DAMAGE WAS NOTED. WHEN THE STAPLING ISSUE OCCURRED THE CUSTOMER RECEIVED AN ERROR MESSAGE AS WOULD NOT REGISTER THE STAPLER. THERE WAS NO OBSTRUCTION BETWEEN INSTRUMENT JAWS. NO CLAMPING ISSUE WAS NOTED. THE PROCEDURE WAS COMPLETED AS PLANNED WITH A BACKUP STAPLER FROM A DIFFERENT LOT. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE DID NOT RETURN WHEN FIRING. KNIFE REVERSE SWITCH WAS USED, BUT THE ISSUE DID NOT IMPROVE. MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE DID NOT RETURN WHEN FIRING. KNIFE REVERSE SWITCH WAS USED, BUT THE ISSUE DID NOT IMPROVE. MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE DID NOT RETURN WHEN FIRING. KNIFE REVERSE SWITCH WAS USED, BUT THE ISSUE DID NOT IMPROVE. MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
17-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND THERE WAS SLIGHT OOZING. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
17-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND THERE WAS SLIGHT OOZING. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
16-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. NO PATIENT INVOLVED.
16-12-2022 | Injury | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE RELOADS HAVE BEEN MISFIRED INTRA-OPERATIVE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2022 | Injury | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Firing Problem | A0505 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND SECOND HANDLE WAS USED WITH SAME RELOAD. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RE-CLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON RE-ENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A SECOND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. NO PATIENT INVOLVED.
03-11-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AFTER CLOSING THE STAPLER IT COULD NOT BE TRIGGERED. THE HANDLE COULD NOT BE RELEASED. THERE WAS NO PATIENT HARM.
03-11-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AFTER CLOSING THE STAPLER IT COULD NOT BE TRIGGERED. THE HANDLE COULD NOT BE RELEASED. THERE WAS NO PATIENT HARM.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. NO PATIENT INVOLVED.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER JAMMED AND WOULD NOT FIRE. THE STAPLER JAMMED WHILE IT WAS ON TISSUE. THE SURGEON WAS ABLE TO RELEASE IT AND REQUESTED A NEW STAPLER AND CARTRIDGE. THERE WERE NO PATIENT CONSEQUENCES OR CHANGE IN THE POST-OP CARE.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. NO PATIENT INVOLVED.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | FLEX 60 ARTICULATING | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE. WHEN RELOADING THE STAPLER AND DESPITE THE AUDIBLE FEEDBACK HAVING BEEN HEARD, THE NEW RELOADS KEPT COMING OFF WHILE HANDLING THE DEVICE. A NEW ENDOCUTTER HAD TO BE USED, THERE IS NO REPORTED HARM TO THE PATIENT. THE DEVICE IS AVAILABLE TO BE SENT FOR INSPECTION.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. NO PATIENT INVOLVED.
10-12-2022 | Injury | ethicon endo-surgery | 05-01-2023 | ILS 33MM, CURVED | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT THE CIRCULAR STAPLER DID NOT WORK CORRECTLY DURING THE OPERATION. THE SEWING HANDLE DID NOT CLOSE COMPLETELY, AND THERE WAS NO CLEAR SEWING SOUND. DURING THE REVISION, IT TURNED OUT THAT SOME OF THE STAPLES WERE ON THE FABRIC, SOME REMAINED IN THE DEVICE - THE FABRIC WAS NOT STITCHED BUT CUT OFF. THE OPERATION WENT WITHOUT COMPLICATIONS, BUT INSTEAD OF THE PLANNED ANASTOMOSIS, I HAD TO IMPOSE A STOMA. THE PATIENT WAS NOT INJURED, THERE WERE NO DELAYS. THE PROBLEM WAS IDENTIFIED IN TIME AND APPROPRIATE MEASURES WERE TAKEN (IMPOSITION OF A COLOSTOMY). PROCEDURE: UNKNOWN.
10-12-2022 | Injury | ethicon endo-surgery | 05-01-2023 | ILS 33MM, CURVED | Difficult to Open or Close | A051104 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT THE CIRCULAR STAPLER DID NOT WORK CORRECTLY DURING THE OPERATION. THE SEWING HANDLE DID NOT CLOSE COMPLETELY, AND THERE WAS NO CLEAR SEWING SOUND. DURING THE REVISION, IT TURNED OUT THAT SOME OF THE STAPLES WERE ON THE FABRIC, SOME REMAINED IN THE DEVICE - THE FABRIC WAS NOT STITCHED BUT CUT OFF. THE OPERATION WENT WITHOUT COMPLICATIONS, BUT INSTEAD OF THE PLANNED ANASTOMOSIS, I HAD TO IMPOSE A STOMA. THE PATIENT WAS NOT INJURED, THERE WERE NO DELAYS. THE PROBLEM WAS IDENTIFIED IN TIME AND APPROPRIATE MEASURES WERE TAKEN (IMPOSITION OF A COLOSTOMY). PROCEDURE: UNKNOWN.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
15-11-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY THE GREEN CONTOUR CURVED STAPLER HAD A MISFIRE- IT DID NOT CUT, STAPLE, OR RELEASE- MUST HAVE EVENTUALLY OPENED TO REMOVE FROM PATIENT BEFORE OPENING A SECOND DEVICE. THE CASE WAS COMPLETED WITH NO PATIENT CONSEQUENCES.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD JAM.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD JAM.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Mechanical Jam | A0506 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD JAM.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD JAM.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
28-06-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, ECHELON CIRCULAR POWERED STAPLER SPONTANEOUSLY FIRED ON ITS OWN WHILE INSERTED IN THE RECTUM OF THE PATIENT. NO PATIENT HARM WAS REPORTED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE TWO RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE TWO RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE TWO RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE TWO RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
12-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY PROCEDURE, AFTER ATTACHING THE ANVIL, THEY REMOVED THE SAFETY AND ATTEMPTED TO FIRE THE DEVICE AND IT WOULD NOT FIRE. THEY REMOVED THE BATTERY AND REATTACHED AND IT STILL WOULD NOT FIRE. THEY OPENED ANOTHER STAPLER, TOOK THAT BATTERY OUT OF THE NEW STAPLER. DEVICE STILL WOULD NOT FIRE. THEY REMOVED THE STAPLER FROM THE PATIENT, USED THE NEW STAPLER AND IT WORKED JUST FINE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. THE PATIENT IS DOING WELL.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE LOAD DID NOT COME OUT. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE LOAD DID NOT COME OUT. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
12-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RADICAL NEPHRECTOMY THE ANVIL ON THE DEVICE WAS BOWED AND WOULD NOT CLOSE FLUSH WITH THE STAPLE LOAD. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
12-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | ECHELON FLEX 45 STANDARD | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RADICAL NEPHRECTOMY THE ANVIL ON THE DEVICE WAS BOWED AND WOULD NOT CLOSE FLUSH WITH THE STAPLE LOAD. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
02-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBE RESECTION, WHEN THE PULMONARY PARENCHYMA WAS ATTEMPTED TO BE RESECTED, FIRING COULD NOT BE DONE. IT WAS POSSIBLE THAT THE GREEN WAS PRESSED AND THEN THE HANDLE WAS SQUEEZED, RESULTING TO PRE-FIRING. IT WAS ALSO NOTED THAT THE JAWS WERE LOCKED ON TISSUE. TO REMOVE THE DEVICE, THE TISSUE WAS RESECTED USING A NEW STAPLE.
02-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBE RESECTION, WHEN THE PULMONARY PARENCHYMA WAS ATTEMPTED TO BE RESECTED, FIRING COULD NOT BE DONE. IT WAS POSSIBLE THAT THE GREEN WAS PRESSED AND THEN THE HANDLE WAS SQUEEZED, RESULTING TO PRE-FIRING. IT WAS ALSO NOTED THAT THE JAWS WERE LOCKED ON TISSUE. TO REMOVE THE DEVICE, THE TISSUE WAS RESECTED USING A NEW STAPLE.
02-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Expulsion | A010401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBE RESECTION, WHEN THE PULMONARY PARENCHYMA WAS ATTEMPTED TO BE RESECTED, FIRING COULD NOT BE DONE. IT WAS POSSIBLE THAT THE GREEN WAS PRESSED AND THEN THE HANDLE WAS SQUEEZED, RESULTING TO PRE-FIRING. IT WAS ALSO NOTED THAT THE JAWS WERE LOCKED ON TISSUE. TO REMOVE THE DEVICE, THE TISSUE WAS RESECTED USING A NEW STAPLE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE SURGEON WAS UNABLE TO PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE. THE DEVICE DID NOT FIRE.
01-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2022 | Injury | ethicon endo-surgery | 05-01-2023 | CIRCULAR MECH XL SEAL 29MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC ASSISTED SIGMOID RESECTION, AFTER FIRING OF THE DEVICE AND PERFORMING AN AIR LEAK TEST, A LEAK WAS DETECTED FROM AIR BUBBLES ESCAPING. THE LEAK WAS SEWN CLOSED WITH SUTURE AND ANOTHER LEAK TEST WAS PERFORMED WITH SATISFACTORY RESULTS. THE PROCEDURE WAS DELAYED BY 60 MINUTES. A DIVERTING ILEOSTOMY WAS CREATED DUE TO THE LEAK.
01-01-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS FAULTY, SHE DOES NOT KNOW WHOSE IT IS AND WHAT THE SPECIFIC CASE IS. IT GOT STUCK WHILE FIRING AND SEND IT TO US FOR INSPECTION. THERE IS NO HARM TO THE PATIENT OR EXTENSION OF SURGERY TIME.
16-12-2022 | Injury | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND THE ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND 2ND HANDLE WAS USED WITH SAME RELOAD WAS USED. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RECLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON REENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A 2ND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND THE ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND 2ND HANDLE WAS USED WITH SAME RELOAD WAS USED. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RECLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON REENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A 2ND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
17-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC COLONOSCOPY -ENDOSCOPIC AND LYMPH NODE DISSECTION PROCEDURE, WHEN STAPLER WAS INS ERTED TO CUT THE TISSUE, THE STAPLER FIRED BUT IT COULD NOT CUT. THE HANDLE WAS REPLACED BUT IT WAS STILL NOT CUT. THEN, A NEW HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, IN THE COLON, THE DEVICE PARTIALLY FIRED. ADDITIONALLY, THE STAPLES DID NOT MAKE FULL B SHAPE. THIS RESULTED IN TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS THE SURGEON NEED TO REDO THE ANASTOMOSIS BY RE-STAPLING THE DISTAL END OF THE ANASTOMOSIS WITH ANOTHER CIRCULAR STAPLER. THE PROCEDURE WAS EXTENDED FOR MORE THAN THIRTY MINUTES.
15-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, IN THE COLON, THE DEVICE PARTIALLY FIRED. ADDITIONALLY, THE STAPLES DID NOT MAKE FULL B SHAPE. THIS RESULTED IN TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS THE SURGEON NEED TO REDO THE ANASTOMOSIS BY RE-STAPLING THE DISTAL END OF THE ANASTOMOSIS WITH ANOTHER CIRCULAR STAPLER. THE PROCEDURE WAS EXTENDED FOR MORE THAN THIRTY MINUTES.
15-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, IN THE COLON, THE DEVICE PARTIALLY FIRED. ADDITIONALLY, THE STAPLES DID NOT MAKE FULL B SHAPE. THIS RESULTED IN TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS THE SURGEON NEED TO REDO THE ANASTOMOSIS BY RE-STAPLING THE DISTAL END OF THE ANASTOMOSIS WITH ANOTHER CIRCULAR STAPLER. THE PROCEDURE WAS EXTENDED FOR MORE THAN THIRTY MINUTES.
15-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, IN THE COLON, THE DEVICE PARTIALLY FIRED. ADDITIONALLY, THE STAPLES DID NOT MAKE FULL B SHAPE. THIS RESULTED IN TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS THE SURGEON NEED TO REDO THE ANASTOMOSIS BY RE-STAPLING THE DISTAL END OF THE ANASTOMOSIS WITH ANOTHER CIRCULAR STAPLER. THE PROCEDURE WAS EXTENDED FOR MORE THAN THIRTY MINUTES.
17-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC COLONOSCOPY -ENDOSCOPIC AND LYMPH NODE DISSECTION PROCEDURE, WHEN STAPLER WAS INS ERTED TO CUT THE TISSUE, THE STAPLER FIRED BUT IT COULD NOT CUT. THE HANDLE WAS REPLACED BUT IT WAS STILL NOT CUT. THEN, A NEW HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC COLONOSCOPY ENDOSCOPIC AND LYMPH NODE DISSECTION PROCEDURE, WHEN STAPLER WAS INS ERTED TO CUT THE TISSUE, THE STAPLER FIRED BUT IT COULD NOT CUT. THE HANDLE WAS REPLACED BUT IT WAS STILL NOT CUT. THEN, A NEW HANDLE AND RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PLEURAL BIOPSY AND RIGHT WEDGE RESECTION, WHEN THE SURGEON STARTED SQUEEZING THE HANDLE TO FIRE THE FIRST RELOAD ON THE RIGHT LOBE, A CRACK WAS HEARD ON THE FIRST SQUEEZE AND THEN THE STAPLER WOULD NOT ALLOW THE FIRE TO CONTINUE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PLEURAL BIOPSY AND RIGHT WEDGE RESECTION, WHEN THE SURGEON STARTED SQUEEZING THE HANDLE TO FIRE THE FIRST RELOAD ON THE RIGHT LOBE, A CRACK WAS HEARD ON THE FIRST SQUEEZE AND THEN THE STAPLER WOULD NOT ALLOW THE FIRE TO CONTINUE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PLEURAL BIOPSY AND RIGHT WEDGE RESECTION, WHEN THE SURGEON STARTED SQUEEZING THE HANDLE TO FIRE THE FIRST RELOAD ON THE RIGHT LOBE, A CRACK WAS HEARD ON THE FIRST SQUEEZE AND THEN THE STAPLER WOULD NOT ALLOW THE FIRE TO CONTINUE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-11-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, WHILE ANASTOMOSING, APPARENTLY THE DEVICE LOADED PROPERLY, AND IT ALL PERFORMED AS NORMAL UNTIL THE SURGEON WENT TO FIRE THE DEVICE AND IT DID NOT FIRE. THE DEVICE WAS UNLOADED AND LOADED WITH ANOTHER RELOAD TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC SEGMENTECTOMY, IN DETACHING THE LUNG PARENCHYMA USING THE FIRST RELOAD, THE STAPLES WERE POORLY FORMED. THE STAPLES DID NOT HOLD THE TISSUES AND THE STAPLE LINE OPENS UP. A BLACK SIXTY MILLIMETER (60MM) RELOAD WAS THEN USED, WHICH SUCCESSFULLY DETACHED THE LUNG. WHEN A NEW PURPLE RELOAD WAS USED TO DETACH THE LUNG LOBE BRONCHI, THE STAPLES WERE POORLY FORMED. SUTURES WERE ADDED FOR STITCHING SUBSEQUENTLY.
19-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC SEGMENTECTOMY, IN DETACHING THE LUNG PARENCHYMA USING THE FIRST RELOAD, THE STAPLES WERE POORLY FORMED. THE STAPLES DID NOT HOLD THE TISSUES AND THE STAPLE LINE OPENS UP. A BLACK SIXTY MILLIMETER (60MM) RELOAD WAS THEN USED, WHICH SUCCESSFULLY DETACHED THE LUNG. WHEN A NEW PURPLE RELOAD WAS USED TO DETACH THE LUNG LOBE BRONCHI, THE STAPLES WERE POORLY FORMED. SUTURES WERE ADDED FOR STITCHING SUBSEQUENTLY.
19-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC SEGMENTECTOMY, IN DETACHING THE LUNG PARENCHYMA USING THE FIRST RELOAD, THE STAPLES WERE POORLY FORMED. THE STAPLES DID NOT HOLD THE TISSUES AND THE STAPLE LINE OPENS UP. A BLACK SIXTY MILLIMETER (60MM) RELOAD WAS THEN USED, WHICH SUCCESSFULLY DETACHED THE LUNG. WHEN A NEW PURPLE RELOAD WAS USED TO DETACH THE LUNG LOBE BRONCHI, THE STAPLES WERE POORLY FORMED. SUTURES WERE ADDED FOR STITCHING SUBSEQUENTLY.
19-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC SEGMENTECTOMY, IN DETACHING THE LUNG PARENCHYMA USING THE FIRST RELOAD, THE STAPLES WERE POORLY FORMED. THE STAPLES DID NOT HOLD THE TISSUES AND THE STAPLE LINE OPENS UP. A BLACK SIXTY MILLIMETER (60MM) RELOAD WAS THEN USED, WHICH SUCCESSFULLY DETACHED THE LUNG. WHEN A NEW PURPLE RELOAD WAS USED TO DETACH THE LUNG LOBE BRONCHI, THE STAPLES WERE POORLY FORMED. SUTURES WERE ADDED FOR STITCHING SUBSEQUENTLY.
19-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC SEGMENTECTOMY, IN DETACHING THE LUNG PARENCHYMA USING THE FIRST RELOAD, THE STAPLES WERE POORLY FORMED. THE STAPLES DID NOT HOLD THE TISSUES AND THE STAPLE LINE OPENS UP. A BLACK SIXTY MILLIMETER (60MM) RELOAD WAS THEN USED, WHICH SUCCESSFULLY DETACHED THE LUNG. WHEN A NEW PURPLE RELOAD WAS USED TO DETACH THE LUNG LOBE BRONCHI, THE STAPLES WERE POORLY FORMED. SUTURES WERE ADDED FOR STITCHING SUBSEQUENTLY
03-08-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED ANTERIOR RESECTION PLUS STOMA FORMATION, THE DEVICE WAS LOADED. THE HANDLE BROKE OFF AFTER FIRING THE DEVICE. TO RESOLVE THE ISSUE, A STAPLER FROM A DIFFERENT MANUFACTURER WAS USED. THERE WAS NO PATIENT INJURY.
03-08-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED ANTERIOR RESECTION PLUS STOMA FORMATION, THE DEVICE WAS LOADED. THE HANDLE BROKE OFF AFTER FIRING THE DEVICE. TO RESOLVE THE ISSUE, A STAPLER FROM A DIFFERENT MANUFACTURER WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE TWO RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE TWO RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS DID NOT OPEN WHEN THE ECH60R WAS ATTEMPTED TO LOAD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
17-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC RESECTION, WHILE DEALING WITH INTERLOBAR FISSURES IN THE LOWER LOBE OF THE RIGHT LUNG, AFTER THE FIRST SHOT WAS CHALLENGED WITH THE FIRST RELOAD, THE STAPLES DID NOT HAVE GOOD B-SHAPED SHAPE AND BLEEDING OCCURRED. THE OOZING TISSUE WAS SEWN AND REPLACED WITH A NEW GUN AND 2 NEW RELOADS. WHEN 2 NEW RELOADS WITH THE NEW GUN WERE TRIED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HANDLE WAS SQUEEZED AND THE SECOND AND THIRD RELOAD DID NOT FIRE. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. WHEN THE ARTICULATION WAS ATTEMPTED TO BE ARTICULATED TO THE RIGHT, THE JAWS WERE CLOSING AND WAS MAKING UNNATURAL MOVEMENTS. THE RELOAD COULD NOT BE REMOVED SMOOTHLY. THIS HAPPENED ON 2 RELOADS.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Unintended Movement | A0512 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. WHEN THE ARTICULATION WAS ATTEMPTED TO BE ARTICULATED TO THE RIGHT, THE JAWS WERE CLOSING AND WAS MAKING UNNATURAL MOVEMENTS. THE RELOAD COULD NOT BE REMOVED SMOOTHLY. THIS HAPPENED ON 2 RELOADS.
13-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION SURGERY, WHILE AT RECTAL RESECTION, CLAMP COULD NOT BE DONE FOR THE 2 RELOADS AND HANDLE, ALTHOUGH THERE WAS NO PROBLEM IN CONFIRMING THE OPENING AND CLOSING IN PREPARATION. EVEN AFTER SOMEHOW TRYING THE FIRING, IT COULD NOT BE DONE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC RESECTION, WHILE DEALING WITH INTERLOBAR FISSURES IN THE LOWER LOBE OF THE RIGHT LUNG, AFTER THE FIRST SHOT WAS CHALLENGED WITH THE FIRST RELOAD, THE STAPLES DID NOT HAVE GOOD B-SHAPED SHAPE AND BLEEDING OCCURRED. THE OOZING TISSUE WAS SEWN AND REPLACED WITH A NEW GUN AND 2 NEW RELOADS. WHEN 2 NEW RELOADS WITH THE NEW GUN WERE TRIED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HANDLE WAS SQUEEZED AND THE SECOND AND THIRD RELOAD DID NOT FIRE. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
16-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC RESECTION, WHILE DEALING WITH INTERLOBAR FISSURES IN THE LOWER LOBE OF THE RIGHT LUNG, AFTER THE FIRST SHOT WAS CHALLENGED WITH THE FIRST RELOAD, THE STAPLES DID NOT HAVE GOOD B-SHAPED SHAPE AND BLEEDING OCCURRED. THE OOZING TISSUE WAS SEWN AND REPLACED WITH A NEW GUN AND 2 NEW RELOADS. WHEN 2 NEW RELOADS WITH THE NEW GUN WERE TRIED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HANDLE WAS SQUEEZED AND THE SECOND AND THIRD RELOAD DID NOT FIRE. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
16-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC RESECTION, WHILE DEALING WITH INTERLOBAR FISSURES IN THE LOWER LOBE OF THE RIGHT LUNG, AFTER THE FIRST SHOT WAS CHALLENGED WITH THE FIRST RELOAD, THE STAPLES DID NOT HAVE GOOD B-SHAPED SHAPE AND BLEEDING OCCURRED. THE OOZING TISSUE WAS SEWN AND REPLACED WITH A NEW GUN AND 2 NEW RELOADS. WHEN 2 NEW RELOADS WITH THE NEW GUN WERE TRIED,THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HANDLE WAS SQUEEZED AND THE SECOND AND THIRD RELOAD DID NOT FIRE. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC RESECTION, WHILE DEALING WITH INTERLOBAR FISSURES IN THE LOWER LOBE OF THE RIGHT LUNG, AFTER THE FIRST SHOT WAS CHALLENGED WITH THE FIRST RELOAD, THE STAPLES DID NOT HAVE GOOD B-SHAPED SHAPE AND BLEEDING OCCURRED. THE OOZING TISSUE WAS SEWN AND REPLACED WITH A NEW GUN AND 2 NEW RELOADS. WHEN 2 NEW RELOADS WITH THE NEW GUN WERE TRIED,THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HANDLE WAS SQUEEZED AND THE SECOND AND THIRD RELOAD DID NOT FIRE. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
15-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID COLON RESECTION PROCEDURE, WHILE RESECTIONING THE SIGMOID COLON, THE BLADE STOPPED DURING STAPLING. WHEN THE JAWS WERE OPENED, IT WAS OBSERVED THAT THE BLADE DID NOT RUN UNTIL THE TIP AND THE STAPLING STOPPED HALFWAY THROUGH. A NEW RELOAD WAS USED TO CONTINUE AND COMPLETE THE STAPLE LINE. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION SURGERY, WHILE AT RECTAL RESECTION, CLAMP COULD NOT BE DONE FOR THE 2 RELOADS AND HANDLE, ALTHOUGH THERE WAS NO PROBLEM IN CONFIRMING THE OPENING AND CLOSING IN PREPARATION. EVEN AFTER SOMEHOW TRYING THE FIRING, IT COULD NOT BE DONE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE DONE WHEN IT WAS ATTEMPTED TO USE THIS STAPLER. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE HANDLE WAS NOT SQUEEZED, GREEN BUTTON WAS NOT PRESSED, AND THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND THERE WAS SLIGHT OOZING . SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND THERE WAS SLIGHT OOZING . SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. WHEN THE ARTICULATION WAS ATTEMPTED TO BE ARTICULATED TO THE RIGHT, THE JAWS WERE CLOSING AND WAS MAKING UNNATURAL MOVEMENTS. THE RELOAD COULD NOT BE REMOVED SMOOTHLY. THIS HAPPENED ON 2 RELOADS.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. WHEN THE ARTICULATION WAS ATTEMPTED TO BE ARTICULATED TO THE RIGHT, THE JAWS WERE CLOSING AND WAS MAKING UNNATURAL MOVEMENTS. THE RELOAD COULD NOT BE REMOVED SMOOTHLY. THIS HAPPENED ON 2 RELOADS.
16-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Unintended Movement | A0512 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. WHEN THE ARTICULATION WAS ATTEMPTED TO BE ARTICULATED TO THE RIGHT, THE JAWS WERE CLOSING AND WAS MAKING UNNATURAL MOVEMENTS. THE RELOAD COULD NOT BE REMOVED SMOOTHLY. THIS HAPPENED ON 2 RELOADS.
16-12-2022 | Malfunction | us surgical puerto rico | 05-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RADICAL RECTAL CANCER PROCEDURE, WEN THE RECTUM WAS DETACHED, THE SOUND OF RATTLING WAS HEARD, BUT THE STAPLE WAS FIRED SMOOTHLY. WHEN THE NEXT STAPLE WAS USED, IT COULD NOT BE FIRED, AND ANOTHER HANDLE WAS USED, BUT THE RATTLING SOUND STILL APPEARED. THERE WAS NO RESPONSE TO STAPLE FIRING. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. REPLACED WITH THE THIRD HANDLE AND THE STAPLE WAS FIRED SMOOTHLY. THERE WAS NO PATIENT INJURY.
17-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Misfire | A050502 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC GASTRECTOMY PROCEDURE, IN STOMACH SECTION, THE STAPLE THAT DEPLOY ON THE PROXIMAL SIDE DID NOT CLOSE. ADDITIONALLY, THE DEVICE DID NOT FIRE COMPLETELY AND THE KNIFE RETRACTED BACK. AN OPEN STAPLE FELL INTO THE PATIENTS CAVITY. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED WITH THE SAME STAPLER. THE SURGEON ALSO PERFORMED A RAFFIA WITH SUTURE. IT WAS ALSO NOTED THAT AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
17-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC GASTRECTOMY PROCEDURE, IN STOMACH SECTION, THE STAPLE THAT DEPLOY ON THE PROXIMAL SIDE DID NOT CLOSE. ADDITIONALLY, THE DEVICE DID NOT FIRE COMPLETELY AND THE KNIFE RETRACTED BACK. AN OPEN STAPLE FELL INTO THE PATIENTS CAVITY. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED WITH THE SAME STAPLER. THE SURGEON ALSO PERFORMED A RAFFIA WITH SUTURE. IT WAS ALSO NOTED THAT AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
17-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC GASTRECTOMY PROCEDURE, IN STOMACH SECTION, THE STAPLE THAT DEPLOY ON THE PROXIMAL SIDE DID NOT CLOSE. ADDITIONALLY, THE DEVICE DID NOT FIRE COMPLETELY AND THE KNIFE RETRACTED BACK. AN OPEN STAPLE FELL INTO THE PATIENTS CAVITY. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED WITH THE SAME STAPLER. THE SURGEON ALSO PERFORMED A RAFFIA WITH SUTURE. IT WAS ALSO NOTED THAT AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
17-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC GASTRECTOMY PROCEDURE, IN STOMACH SECTION, THE STAPLE THAT DEPLOY ON THE PROXIMAL SIDE DID NOT CLOSE. ADDITIONALLY, THE DEVICE DID NOT FIRE COMPLETELY AND THE KNIFE RETRACTED BACK. AN OPEN STAPLE FELL INTO THE PATIENTS CAVITY. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED WITH THE SAME STAPLER. THE SURGEON ALSO PERFORMED A RAFFIA WITH SUTURE. IT WAS ALSO NOTED THAT AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
17-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC GASTRECTOMY PROCEDURE, IN STOMACH SECTION, THE STAPLE THAT DEPLOY ON THE PROXIMAL SIDE DID NOT CLOSE. ADDITIONALLY, THE DEVICE DID NOT FIRE COMPLETELY AND THE KNIFE RETRACTED BACK. AN OPEN STAPLE FELL INTO THE PATIENTS CAVITY. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED WITH THE SAME STAPLER. THE SURGEON ALSO PERFORMED A RAFFIA WITH SUTURE. IT WAS ALSO NOTED THAT AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
17-12-2022 | Injury | us surgical puerto rico | 05-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC GASTRECTOMY PROCEDURE, IN STOMACH SECTION, THE STAPLE THAT DEPLOY ON THE PROXIMAL SIDE DID NOT CLOSE. ADDITIONALLY, THE DEVICE DID NOT FIRE COMPLETELY AND THE KNIFE RETRACTED BACK. AN OPEN STAPLE FELL INTO THE PATIENTS CAVITY. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED WITH THE SAME STAPLER. THE SURGEON ALSO PERFORMED A RAFFIA WITH SUTURE. IT WAS ALSO NOTED THAT AN ADDITIONAL OPERATION WAS PERFORMED TO CORRECT THE ISSUE.
02-11-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IN THE STOMACH, THE STAPLE RELOAD WAS OPENED AND LOADED ONTO STAPLER. THE SURGEON WENT TO GRASP TISSUE AND NOTICED THAT THE SUTURE ON THE TOP OF THE CARTRIDGE WHICH HOLDS THE REINFORCEMENT MATERIAL IN PLACE WAS MISSING. THIS CAUSED THE MATERIAL TO FLOP OFF THE RELOAD. THE RELOAD WAS REMOVED AND REPLACED WITH NEW RELOAD WITH REINFORCEMENT. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT THE DISTAL CARTRIDGE SUTURE WAS MISSING. WITNESS MARKS ON THE DISTAL CARTRIDGE REINFORCEMENT MATERIAL INDICATE THAT THE SUTURE WAS ANCHORED AT SOME POINT.
02-11-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Component Missing | A020602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IN THE STOMACH, THE STAPLE RELOAD WAS OPENED AND LOADED ONTO STAPLER. THE SURGEON WENT TO GRASP TISSUE AND NOTICED THAT THE SUTURE ON THE TOP OF THE CARTRIDGE WHICH HOLDS THE REINFORCEMENT MATERIAL IN PLACE WAS MISSING. THIS CAUSED THE MATERIAL TO FLOP OFF THE RELOAD. THE RELOAD WAS REMOVED AND REPLACED WITH NEW RELOAD WITH REINFORCEMENT. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT THE DISTAL CARTRIDGE SUTURE WAS MISSING. WITNESS MARKS ON THE DISTAL CARTRIDGE REINFORCEMENT MATERIAL INDICATE THAT THE SUTURE WAS ANCHORED AT SOME POINT.
05-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, WHEN THE CARTRIDGE WAS CONNECTED AND THE OPENING AND CLOSING WAS CHECKED, THE JAW DID NOT OPEN. AFTERWARDS, WHEN VARIOUS THINGS WERE TRIED ON THE MACHINE TABLE, THE GREEN BUTTON LIT UP SO FIRING WAS TESTED, AND THE DEVICE WAS ABLE TO FIRE UNTIL THE END. THE CLAMP COVER RETURNED TO NEAR THE PROXIMAL SIDE, BUT THE JAWS DID NOT OPEN, AND THE CARTRIDGE COULD NOT BE REMOVED FROM THE ADAPTER. THE RELEASE BUTTON MOVED UP BUT NOT IN A STATE IN WHICH THE RED BAR WAS VISIBLE; BUT IN A MOVABLE STATE. AFTER THAT, THE CARTRIDGE CAME OFF, BUT THE RESIN PART AT THE PROXIMAL SIDE BROKE, SO SOME PARTS REMAINED INSIDE THE ADAPTER. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INVOLVED.
05-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, WHEN THE CARTRIDGE WAS CONNECTED AND THE OPENING AND CLOSING WAS CHECKED, THE JAW DID NOT OPEN. AFTERWARDS, WHEN VARIOUS THINGS WERE TRIED ON THE MACHINE TABLE, THE GREEN BUTTON LIT UP SO FIRING WAS TESTED, AND THE DEVICE WAS ABLE TO FIRE UNTIL THE END. THE CLAMP COVER RETURNED TO NEAR THE PROXIMAL SIDE, BUT THE JAWS DID NOT OPEN, AND THE CARTRIDGE COULD NOT BE REMOVED FROM THE ADAPTER. THE RELEASE BUTTON MOVED UP BUT NOT IN A STATE IN WHICH THE RED BAR WAS VISIBLE; BUT IN A MOVABLE STATE. AFTER THAT, THE CARTRIDGE CAME OFF, BUT THE RESIN PART AT THE PROXIMAL SIDE BROKE, SO SOME PARTS REMAINED INSIDE THE ADAPTER. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INVOLVED.
05-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, WHEN THE CARTRIDGE WAS CONNECTED AND THE OPENING AND CLOSING WAS CHECKED, THE JAW DID NOT OPEN. AFTERWARDS, WHEN VARIOUS THINGS WERE TRIED ON THE MACHINE TABLE, THE GREEN BUTTON LIT UP SO FIRING WAS TESTED, AND THE DEVICE WAS ABLE TO FIRE UNTIL THE END. THE CLAMP COVER RETURNED TO NEAR THE PROXIMAL SIDE, BUT THE JAWS DID NOT OPEN, AND THE CARTRIDGE COULD NOT BE REMOVED FROM THE ADAPTER. THE RELEASE BUTTON MOVED UP BUT NOT IN A STATE IN WHICH THE RED BAR WAS VISIBLE; BUT IN A MOVABLE STATE. AFTER THAT, THE CARTRIDGE CAME OFF, BUT THE RESIN PART AT THE PROXIMAL SIDE BROKE, SO SOME PARTS REMAINED INSIDE THE ADAPTER. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INVOLVED.
05-12-2022 | Malfunction | covidien lp llc north haven | 05-01-2023 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, WHEN THE CARTRIDGE WAS CONNECTED AND THE OPENING AND CLOSING WAS CHECKED, THE JAW DID NOT OPEN. AFTERWARDS, WHEN VARIOUS THINGS WERE TRIED ON THE MACHINE TABLE, THE GREEN BUTTON LIT UP SO FIRING WAS TESTED, AND THE DEVICE WAS ABLE TO FIRE UNTIL THE END. THE CLAMP COVER RETURNED TO NEAR THE PROXIMAL SIDE, BUT THE JAWS DID NOT OPEN, AND THE CARTRIDGE COULD NOT BE REMOVED FROM THE ADAPTER. THE RELEASE BUTTON MOVED UP BUT NOT IN A STATE IN WHICH THE RED BAR WAS VISIBLE; BUT IN A MOVABLE STATE. AFTER THAT, THE CARTRIDGE CAME OFF, BUT THE RESIN PART AT THE PROXIMAL SIDE BROKE, SO SOME PARTS REMAINED INSIDE THE ADAPTER. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INVOLVED.
13-12-2022 | Malfunction | davol inc sub c.r bard inc -1213643 | 05-01-2023 | SORBAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING IPOM PROCEDURE THE SORBAFIX ENHANCED WAS USED TO FIXATE THE VENTRALIGHT ST W/ECHO. IT WAS REPORTED THAT WHEN DEPLOYING THE FIRST FASTENER THE DEVICE REQUIRED EXTRA FORCE TO RELEASE THE FASTENER. IT WAS ALSO REPORTED THAT THE FASTENERS WERE STUCK WHILE FIRING AND SOME OF THE FASTENERS WERE LOOSE. THE LOOSE FASTENERS WERE REMOVED FROM PATIENT'S BODY. THERE WERE NO BROKEN FASTENERS AND THE PROCEDURE WAS COMPLETED USING ANOTHER NON BARD/DAVOL FIXATION DEVICE. THERE WAS NO REPORTED PATIENT INJURY.
19-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT SURGEON FIRED ECHELON+ ON SECOND FIRING IN SLEEVE CASE USING A GREEN RELOAD. NO STAPLES DEPLOYED ON STOMACH SIDE AND SIGNIFICANT HOLE AND BLEEDING CREATED IN PATIENTS STOMACH. SURGEON HAD TO HAND SEW HOLE CLOSED. SURGERY DELAYED DUE TO THE REPORTED EVENT BY 60 MINUTES. PROCEDURE WAS SUCCESSFULLY COMPLETED. THE PATIENTS RECOVERY FROM THE SURGERY WILL TAKE A LOT LONGER THEN IT SHOULD.
19-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT SURGEON FIRED ECHELON+ ON SECOND FIRING IN SLEEVE CASE USING A GREEN RELOAD. NO STAPLES DEPLOYED ON STOMACH SIDE AND SIGNIFICANT HOLE AND BLEEDING CREATED IN PATIENTS STOMACH. SURGEON HAD TO HAND SEW HOLE CLOSED. SURGERY DELAYED DUE TO THE REPORTED EVENT BY 60 MINUTES. PROCEDURE WAS SUCCESSFULLY COMPLETED. THE PATIENTS RECOVERY FROM THE SURGERY WILL TAKE A LOT LONGER THEN IT SHOULD.
19-12-2022 | Malfunction | ethicon endo-surgery | 05-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT SURGEON FIRED ECHELON+ ON SECOND FIRING IN SLEEVE CASE USING A GREEN RELOAD. NO STAPLES DEPLOYED ON STOMACH SIDE AND SIGNIFICANT HOLE AND BLEEDING CREATED IN PATIENTS STOMACH. SURGEON HAD TO HAND SEW HOLE CLOSED. SURGERY DELAYED DUE TO THE REPORTED EVENT BY 60 MINUTES. PROCEDURE WAS SUCCESSFULLY COMPLETED. THE PATIENTS RECOVERY FROM THE SURGERY WILL TAKE A LOT LONGER THEN IT SHOULD.
14-12-2022 | Injury | covidien lp llc north haven | 05-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PYLORUS-PRESERVING PANCREATICODUODENECTOMY AND GASTRECTOMY, THE DEVICE STAPLED ONLY HALF OF THE SUTURE LENGTH. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE, BUT ADDITIONAL RESECTION WAS PERFORMED FOR THE PART THAT WAS FAILED TO BE STAPLED. INCISION WAS EXTENDED BUT DID NOT EXCEED ONE INCH.
06-12-2022 | Injury | intuitive surgical | 05-01-2023 | SUREFORM | Failure to Cut | A050702 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, A SUREFORM 60 STAPLER INSTRUMENT FIRED A RELOAD AND THE TARGET TISSUE WAS PUSHED AND BUNCHED. THERE WAS NO REPORTED PATIENT INJURY. THE PROCEDURE WAS COMPLETED. THE INTUITIVE SURGICAL INC. (ISI) TECHNICAL SUPPORT ENGINEER (TSE) REVIEWED THE SYSTEM LOGS FOR THIS PROCEDURE AND FOUND NO RELEVANT ERRORS. ISI CONTACTED THE INITIAL REPORTER FOR THIS COMPLAINT AND ADDITIONAL INFORMATION WAS OBTAINED: THE REPORTED STAPLING EVENT OCCURRED WHILE FIRING ON THE STOMACH TO CREATE THE SLEEVE. THE INITIAL REPORTER SAID SHE WAS GIVEN THE SUREFORM 60 STAPLER INSTRUMENT FIRED DURING THIS EVENT AND IT HAD A BLUE SUREFORM 60 RELOAD INSTALLED IN THE JAWS WHEN SHE PACKAGED IT TO BE RETURNED. THE CUSTOMER REPORTED THAT THEY WERE ADVISED TO RETURN ALL THE INSTRUMENTS USED DURING THIS EVENT BY THE SITE'S ISI CLINICAL SALES REPRESENTATIVE (CSR). THE SITE CSR WAS CONTACTED AND REPORTED THAT THIS EVENT OCCURRED WITH A GREEN SUREFORM 60 RELOAD. AN IMAGE OF THIS EVENT WAS PROVIDED FROM THE SURGEON TO THE CSR. ISI HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION FROM THE SURGEON OF THIS PROCEDURE. HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
01-12-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP ANTERIOR RESECTION THE POWERED CIRCULAR DID NOT WORK. THEY PREPARED THE DEVICE, INSERTED BATTERY AND THEN WHEN THE TIME CAME FOR CIRCULAR STAPLER ANASTASIS CREATION STAGE IT DID NOT FIRE. INDICATOR WAS IN GREEN STAGE ACCORDING TO THEM, BUT AFTER TRIGGER WAS PUSHED DEVICE DID NOT WORKED. THE SURGERY WAS DELAYED BY 10 MINUTES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. NO INFORMATION ABOUT THE CONSEQUENCES TO THE PATIENT OR ABOUT POST-OPERATIVE CARE AS A RESULT OF THE EVENT.
14-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE HANDLE WOULD NOT FIRE RELOADS WHEN LOADED. HAND SUTURING WAS PERFORMED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, ESOPHAGEAL JEJUNUM FUNCTIONAL END-TO-END RECONSTRUCTION, WHILE INSERTING INTO THE ABDOMINAL CAVITY, AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT THE JAWS COULD NOT BE OPENED. IT WAS NOTED THAT IT WAS WHILE USING THE HANDLE FOR DUODENAL TRANSECTION, ESOPHAGEAL TRANSECTION, JEJUNAL TRANSECTION, Y SIDE-TO-SIDE ANASTOMOSIS , JEJUNUM RESECTION AND SIDE -TO-SIDE ANASTOMOSIS OF ESOPHAGEAL JEJUNUM WERE PERFORMED, THEN USED THE REPORTED RELOAD FOR THE SEVENTH FIRING TO PERFORM ENTRY HOLE CLOSURE. THE NURSE CONNECTED THE RELOAD SUCCESSFULLY, WITH CONFIRMATION OF OPENING AND CLOSING THAT WAS ALSO NORMAL (ALSO THE DISPLAY). WHEN THE PHYSICIAN WAS APPLYING THE SUPPORT SUTURE TO THE TISSUE, THE CARTRIDGE WAS CLOSED AND WAS ON STAND-BY ON THE INSTRUMENT TABLE. AFTER THAT, EVEN IF THE RELEVANT POWER HANDLE WAS FORCIBLY RESET, THE POWER HANDLE ERROR WAS SHOWING. RECONSTRUCTION OUTSIDE THE ABDOMINAL CAVITY WAS COMPLETED WITHOUT PROBLEMS BY HAND-SEWING THE COMMON OPENING. THE ANASTOMOSIS O PERATION WAS EXTENDED FOR ABOUT 15 MINUTES BY MANUAL SUTURING. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, ESOPHAGEAL JEJUNUM FUNCTIONAL END-TO-END RECONSTRUCTION, WHILE INSERTING INTO THE ABDOMINAL CAVITY, AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT THE JAWS COULD NOT BE OPENED. IT WAS NOTED THAT IT WAS WHILE USING THE HANDLE FOR DUODENAL TRANSECTION, ESOPHAGEAL TRANSECTION, JEJUNAL TRANSECTION, Y SIDE-TO-SIDE ANASTOMOSIS , JEJUNUM RESECTION AND SIDE -TO-SIDE ANASTOMOSIS OF ESOPHAGEAL JEJUNUM WERE PERFORMED, THEN USED THE REPORTED RELOAD FOR THE SEVENTH FIRING TO PERFORM ENTRY HOLE CLOSURE. THE NURSE CONNECTED THE RELOAD SUCCESSFULLY, WITH CONFIRMATION OF OPENING AND CLOSING THAT WAS ALSO NORMAL (ALSO THE DISPLAY). WHEN THE PHYSICIAN WAS APPLYING THE SUPPORT SUTURE TO THE TISSUE, THE CARTRIDGE WAS CLOSED AND WAS ON STAND-BY ON THE INSTRUMENT TABLE. AFTER THAT, EVEN IF THE RELEVANT POWER HANDLE WAS FORCIBLY RESET, THE POWER HANDLE ERROR WAS SHOWING. RECONSTRUCTION OUTSIDE THE ABDOMINAL CAVITY WAS COMPLETED WITHOUT PROBLEMS BY HAND-SEWING THE COMMON OPENING. THE ANASTOMOSIS O PERATION WAS EXTENDED FOR ABOUT 15 MINUTES BY MANUAL SUTURING. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, ESOPHAGEAL JEJUNUM FUNCTIONAL END-TO-END RECONSTRUCTION, WHILE INSERTING INTO THE ABDOMINAL CAVITY, AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT THE JAWS COULD NOT BE OPENED. IT WAS NOTED THAT IT WAS WHILE USING THE HANDLE FOR DUODENAL TRANSECTION, ESOPHAGEAL TRANSECTION, JEJUNAL TRANSECTION, Y SIDE-TO-SIDE ANASTOMOSIS , JEJUNUM RESECTION AND SIDE -TO-SIDE ANASTOMOSIS OF ESOPHAGEAL JEJUNUM WERE PERFORMED, THEN USED THE REPORTED RELOAD FOR THE SEVENTH FIRING TO PERFORM ENTRY HOLE CLOSURE. THE NURSE CONNECTED THE RELOAD SUCCESSFULLY, WITH CONFIRMATION OF OPENING AND CLOSING THAT WAS ALSO NORMAL (ALSO THE DISPLAY). WHEN THE PHYSICIAN WAS APPLYING THE SUPPORT SUTURE TO THE TISSUE, THE CARTRIDGE WAS CLOSED AND WAS ON STAND-BY ON THE INSTRUMENT TABLE. AFTER THAT, EVEN IF THE RELEVANT POWER HANDLE WAS FORCIBLY RESET, THE POWER HANDLE ERROR WAS SHOWING. RECONSTRUCTION OUTSIDE THE ABDOMINAL CAVITY WAS COMPLETED WITHOUT PROBLEMS BY HAND-SEWING THE COMMON OPENING. THE ANASTOMOSIS OPERATION WAS EXTENDED FOR ABOUT 15 MINUTES BY MANUAL SUTURING. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, ESOPHAGEAL JEJUNUM FUNCTIONAL END-TO-END RECONSTRUCTION, WHILE INSERTING INTO THE ABDOMINAL CAVITY, AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT THE JAWS COULD NOT BE OPENED. IT WAS NOTED THAT IT WAS WHILE USING THE HANDLE FOR DUODENAL TRANSECTION, ESOPHAGEAL TRANSECTION, JEJUNAL TRANSECTION, Y SIDE-TO-SIDE ANASTOMOSIS , JEJUNUM RESECTION AND SIDE -TO-SIDE ANASTOMOSIS OF ESOPHAGEAL JEJUNUM WERE PERFORMED, THEN USED THE REPORTED RELOAD FOR THE SEVENTH FIRING TO PERFORM ENTRY HOLE CLOSURE. THE NURSE CONNECTED THE RELOAD SUCCESSFULLY, WITH CONFIRMATION OF OPENING AND CLOSING THAT WAS ALSO NORMAL (ALSO THE DISPLAY). WHEN THE PHYSICIAN WAS APPLYING THE SUPPORT SUTURE TO THE TISSUE, THE CARTRIDGE WAS CLOSED AND WAS ON STAND-BY ON THE INSTRUMENT TABLE. AFTER THAT, EVEN IF THE RELEVANT POWER HANDLE WAS FORCIBLY RESET, THE POWER HANDLE ERROR WAS SHOWING. RECONSTRUCTION OUTSIDE THE ABDOMINAL CAVITY WAS COMPLETED WITHOUT PROBLEMS BY HAND-SEWING THE COMMON OPENING. THE ANASTOMOSIS OPERATION WAS EXTENDED FOR ABOUT 15 MINUTES BY MANUAL SUTURING. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, ESOPHAGEAL JEJUNUM FUNCTIONAL END-TO-END RECONSTRUCTION, WHILE INSERTING INTO THE ABDOMINAL CAVITY, AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT THE JAWS COULD NOT BE OPENED. IT WAS NOTED THAT IT WAS WHILE USING THE HANDLE FOR DUODENAL TRANSECTION, ESOPHAGEAL TRANSECTION, JEJUNAL TRANSECTION, Y SIDE-TO-SIDE ANASTOMOSIS , JEJUNUM RESECTION AND SIDE -TO-SIDE ANASTOMOSIS OF ESOPHAGEAL JEJUNUM WERE PERFORMED, THEN USED THE REPORTED RELOAD FOR THE SEVENTH FIRING TO PERFORM ENTRY HOLE CLOSURE. THE NURSE CONNECTED THE RELOAD SUCCESSFULLY, WITH CONFIRMATION OF OPENING AND CLOSING THAT WAS ALSO NORMAL (ALSO THE DISPLAY). WHEN THE PHYSICIAN WAS APPLYING THE SUPPORT SUTURE TO THE TISSUE, THE CARTRIDGE WAS CLOSED AND WAS ON STAND-BY ON THE INSTRUMENT TABLE. AFTER THAT, EVEN IF THE RELEVANT POWER HANDLE WAS FORCIBLY RESET, THE POWER HANDLE ERROR WAS SHOWING. RECONSTRUCTION OUTSIDE THE ABDOMINAL CAVITY WAS COMPLETED WITHOUT PROBLEMS BY HAND-SEWING THE COMMON OPENING. THE ANASTOMOSIS O PERATION WAS EXTENDED FOR ABOUT 15 MINUTES BY MANUAL SUTURING. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, IN STAPLING, THERE WAS DIFFICULTY ARTICULATING THE DEVICE. THERE WAS ALSO A UNEXPECTED ARTICULATION OR ROTATION. ADDITIONALLY, THERE WAS DIFFICULTY CLOSING THE JAWS OR POSITIONING THE DEVICE. THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE AS IT WOULD NOT ENTER FIRING MODE. THE RELOAD GOT STUCK IN THE ADAPTER AND COULD NOT BE UNLOAD. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, IN STAPLING, THERE WAS DIFFICULTY ARTICULATING THE DEVICE. THERE WAS ALSO A UNEXPECTED ARTICULATION OR ROTATION. ADDITIONALLY, THERE WAS DIFFICULTY CLOSING THE JAWS OR POSITIONING THE DEVICE. THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE AS IT WOULD NOT ENTER FIRING MODE. THE RELOAD GOT STUCK IN THE ADAPTER AND COULD NOT BE UNLOAD. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, IN STAPLING, THERE WAS DIFFICULTY ARTICULATING THE DEVICE. THERE WAS ALSO A UNEXPECTED ARTICULATION OR ROTATION. ADDITIONALLY, THERE WAS DIFFICULTY CLOSING THE JAWS OR POSITIONING THE DEVICE. THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE AS IT WOULD NOT ENTER FIRING MODE. THE RELOAD GOT STUCK IN THE ADAPTER AND COULD NOT BE UNLOAD. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, IN STAPLING, THERE WAS DIFFICULTY ARTICULATING THE DEVICE. THERE WAS ALSO A UNEXPECTED ARTICULATION OR ROTATION. ADDITIONALLY, THERE WAS DIFFICULTY CLOSING THE JAWS OR POSITIONING THE DEVICE. THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE AS IT WOULD NOT ENTER FIRING MODE. THE RELOAD GOT STUCK IN THE ADAPTER AND COULD NOT BE UNLOAD. IT WAS ALSO MENTIONED THAT THE ANOTHER ADAPTER DID NOT CLOSE. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, IN STAPLING, THERE WAS DIFFICULTY ARTICULATING THE DEVICE. THERE WAS ALSO A UNEXPECTED ARTICULATION OR ROTATION. ADDITIONALLY, THERE WAS DIFFICULTY CLOSING THE JAWS OR POSITIONING THE DEVICE. THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE AS IT WOULD NOT ENTER FIRING MODE. THE RELOAD GOT STUCK IN THE ADAPTER AND COULD NOT BE UNLOAD. IT WAS ALSO MENTIONED THAT THE ANOTHER ADAPTER DID NOT CLOSE. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, IN STAPLING, THERE WAS DIFFICULTY ARTICULATING THE DEVICE. THERE WAS ALSO A UNEXPECTED ARTICULATION OR ROTATION. ADDITIONALLY, THERE WAS DIFFICULTY CLOSING THE JAWS OR POSITIONING THE DEVICE. THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE AS IT WOULD NOT ENTER FIRING MODE. THE RELOAD GOT STUCK IN THE ADAPTER AND COULD NOT BE UNLOAD. IT WAS ALSO MENTIONED THAT THE ANOTHER ADAPTER DID NOT CLOSE. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Positioning Problem | A1502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, IN STAPLING, THERE WAS DIFFICULTY ARTICULATING THE DEVICE. THERE WAS ALSO A UNEXPECTED ARTICULATION OR ROTATION. ADDITIONALLY, THERE WAS DIFFICULTY CLOSING THE JAWS OR POSITIONING THE DEVICE. THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE AS IT WOULD NOT ENTER FIRING MODE. THE RELOAD GOT STUCK IN THE ADAPTER AND COULD NOT BE UNLOAD. IT WAS ALSO MENTIONED THAT THE ANOTHER ADAPTER DID NOT CLOSE. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, IN STAPLING, THERE WAS DIFFICULTY ARTICULATING THE DEVICE. THERE WAS ALSO A UNEXPECTED ARTICULATION OR ROTATION. ADDITIONALLY, THERE WAS DIFFICULTY CLOSING THE JAWS OR POSITIONING THE DEVICE. THE SURGEON WAS NOT ABLE TO FIRE THE DEVICE AS IT WOULD NOT ENTER FIRING MODE. THE RELOAD GOT STUCK IN THE ADAPTER AND COULD NOT BE UNLOAD. IT WAS ALSO MENTIONED THAT THE ANOTHER ADAPTER DID NOT CLOSE. TO RESOLVE THE ISSUE, A NEW ADAPTER AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
10-11-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS GASTROPLASTY, ON STOMACH, THE STAPLER DID NOT SHOOT THE LOAD. A NEW HANDLE HAS BEEN OPENED TO COMPLETE THE CASE. THE RELOAD WORKED ON THE NEW HANDLE. THERE WAS NO PATIENT INJURY.
10-11-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS GASTROPLASTY, ON STOMACH, THE STAPLER DID NOT SHOOT THE LOAD. A NEW HANDLE HAS BEEN OPENED TO COMPLETE THE CASE. THE RELOAD WORKED ON THE NEW HANDLE. THERE WAS NO PATIENT INJURY.
10-11-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS GASTROPLASTY, ON STOMACH, THE STAPLER DID NOT SHOOT THE LOAD. A NEW HANDLE HAS BEEN OPENED TO COMPLETE THE CASE. THE RELOAD WORKED ON THE NEW HANDLE. THERE WAS NO PATIENT INJURY.
14-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE HANDLE WOULD NOT FIRE RELOADS WHEN LOADED. HAND SUTURING WAS PERFORMED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SIGMOID COLON RESECTION, THE HANDLE WAS SQUEEZED AND THE DEVICE DID NOT FIRE. THE SHELL WAS NOT RECOGNIZED. THE RELOAD AND SHELL WERE REPLACED TO RESOLVE THE ISSUE. THE DEVICE WAS NOT USED IN PATIENT. THERE WAS NO PATIENT INJURY.
16-12-2022 | Injury | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Patient Device Interaction Problem | A01 | Hematoma | Event Description: IT WAS REPORTED THAT THE PATIENT WAS READMITTED AND RE-OPERATED AFTER INSUFFICIENT RECOVERY FROM A GASTRIC SLEEVE RESECTION. CT SCAN SHOWED 16CM HEMATOMA. PATIENT WAS OPERATED AGAIN ON FRIDAY ON (B)(6) DUE TO INFECTION PARAMETERS AND SUSPECTED INFECTION IN HEMATOMA. ADDITIONAL SURGERY TO DRAIN HEMATOMA, UNSUCCESSFUL. PATIENT WAS SENT HOME ON MONDAY ON (B)(6).
16-12-2022 | Injury | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Patient Device Interaction Problem | A01 | Unspecified Infection | Event Description: IT WAS REPORTED THAT THE PATIENT WAS READMITTED AND RE-OPERATED AFTER INSUFFICIENT RECOVERY FROM A GASTRIC SLEEVE RESECTION. CT SCAN SHOWED 16CM HEMATOMA. PATIENT WAS OPERATED AGAIN ON FRIDAY ON (B)(6) DUE TO INFECTION PARAMETERS AND SUSPECTED INFECTION IN HEMATOMA. ADDITIONAL SURGERY TO DRAIN HEMATOMA, UNSUCCESSFUL. PATIENT WAS SENT HOME ON MONDAY ON (B)(6).
07-11-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER WOULDN'T RELEASE FROM THE PATIENT AFTER FIRING THE 3RD RELOAD. NO OTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCE REPORTED.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ILEOCECAL RESECTION PROCEDURE THAT WAS NOTED FOR PATIENT WITH CECAL CANCER, THERE WAS NO PROBLEM WITH THE FIRST RESECTION NEAR THE CECUM, BUT WHEN THE ASCENDING COLON WAS RESECTED, THE TISSUE GOT STUCK IN THE STAPLE POCKET, AND THE INTESTINAL TRACT COULD NOT BE REMOVED FROM THE CARTRIDGE AFTER RESECTION. THEREFORE, THE TISSUE WAS GENTLY PULLED OUT FROM THE STAPLE POCKET, AND THE CARTRIDGE COULD BE REMOVED. A CONFIRMATION WAS REQUESTED ON WHETHER THERE WERE ANY PROBLEMS WITH THE STAPLE POCKET. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ILEOCECAL RESECTION PROCEDURE THAT WAS NOTED FOR PATIENT WITH CECAL CANCER, THERE WAS NO PROBLEM WITH THE FIRST RESECTION NEAR THE CECUM, BUT WHEN THE ASCENDING COLON WAS RESECTED, THE TISSUE GOT STUCK IN THE STAPLE POCKET, AND THE INTESTINAL TRACT COULD NOT BE REMOVED FROM THE CARTRIDGE AFTER RESECTION. THEREFORE, THE TISSUE WAS GENTLY PULLED OUT FROM THE STAPLE POCKET, AND THE CARTRIDGE COULD BE REMOVED. A CONFIRMATION WAS REQUESTED ON WHETHER THERE WERE ANY PROBLEMS WITH THE STAPLE POCKET. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ILEOCECAL RESECTION PROCEDURE THAT WAS NOTED FOR PATIENT WITH CECAL CANCER, THERE WAS NO PROBLEM WITH THE FIRST RESECTION NEAR THE CECUM, BUT WHEN THE ASCENDING COLON WAS RESECTED, THE TISSUE GOT STUCK IN THE STAPLE POCKET, AND THE INTESTINAL TRACT COULD NOT BE REMOVED FROM THE CARTRIDGE AFTER RESECTION. THEREFORE, THE TISSUE WAS GENTLY PULLED OUT FROM THE STAPLE POCKET, AND THE CARTRIDGE COULD BE REMOVED. A CONFIRMATION WAS REQUESTED ON WHETHER THERE WERE ANY PROBLEMS WITH THE STAPLE POCKET. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ILEOCECAL RESECTION PROCEDURE THAT WAS NOTED FOR PATIENT WITH CECAL CANCER, THERE WAS NO PROBLEM WITH THE FIRST RESECTION NEAR THE CECUM, BUT WHEN THE ASCENDING COLON WAS RESECTED, THE TISSUE GOT STUCK IN THE STAPLE POCKET, AND THE INTESTINAL TRACT COULD NOT BE REMOVED FROM THE CARTRIDGE AFTER RESECTION. THEREFORE, THE TISSUE WAS GENTLY PULLED OUT FROM THE STAPLE POCKET, AND THE CARTRIDGE COULD BE REMOVED. A CONFIRMATION WAS REQUESTED ON WHETHER THERE WERE ANY PROBLEMS WITH THE STAPLE POCKET. THERE WAS NO PATIENT INJURY.
23-12-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY AFTER THE FIRST FIRING THE ANVIL WAS BENT/DEFLECTED UPWARD SO THE DEVICE WOULD NOT FIT THROUGH THE TROCAR. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
23-12-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY AFTER THE FIRST FIRING THE ANVIL WAS BENT/DEFLECTED UPWARD SO THE DEVICE WOULD NOT FIT THROUGH THE TROCAR. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY: THIRD AND LAST RELOAD - GOLD WITH SINGARD - TOWARDS THE ESOPHAGUS, THE ECHELON POWERED STAPLER HAS A LOT OF DIFFICULTY ADVANCING THEN PROCEEDS EVEN IF THERE WAS SOMETHING STRANGE AND WHEN IT REOPENS THE SUTURE HAS NOT BEEN MADE BUT HAS ONLY CUT THE TISSUE AND THE RELOAD. THERE ARE SOME CROOKED OPEN CLIPS AND THE CARTRIDGE IS BROKEN. THE PROBE IS INTACT AND THEREFORE WAS NOT CAUGHT IN THE STAPLER. BEFORE THE EVENT, THE STAPLER HAD WORKED WELL: 1 BLACK AND 1 GREEN RELOADS WERE USED BOTH WITH SINGARD. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY: THIRD AND LAST RELOAD - GOLD WITH SINGARD - TOWARDS THE ESOPHAGUS, THE ECHELON POWERED STAPLER HAS A LOT OF DIFFICULTY ADVANCING THEN PROCEEDS EVEN IF THERE WAS SOMETHING STRANGE AND WHEN IT REOPENS THE SUTURE HAS NOT BEEN MADE BUT HAS ONLY CUT THE TISSUE AND THE RELOAD. THERE ARE SOME CROOKED OPEN CLIPS AND THE CARTRIDGE IS BROKEN. THE PROBE IS INTACT AND THEREFORE WAS NOT CAUGHT IN THE STAPLER. BEFORE THE EVENT, THE STAPLER HAD WORKED WELL: 1 BLACK AND 1 GREEN RELOADS WERE USED BOTH WITH SINGARD. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY: THIRD AND LAST RELOAD - GOLD WITH SINGARD - TOWARDS THE ESOPHAGUS, THE ECHELON POWERED STAPLER HAS A LOT OF DIFFICULTY ADVANCING THEN PROCEEDS EVEN IF THERE WAS SOMETHING STRANGE AND WHEN IT REOPENS THE SUTURE HAS NOT BEEN MADE BUT HAS ONLY CUT THE TISSUE AND THE RELOAD. THERE ARE SOME CROOKED OPEN CLIPS AND THE CARTRIDGE IS BROKEN. THE PROBE IS INTACT AND THEREFORE WAS NOT CAUGHT IN THE STAPLER. BEFORE THE EVENT, THE STAPLER HAD WORKED WELL: 1 BLACK AND 1 GREEN RELOADS WERE USED BOTH WITH SINGARD. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY: THIRD AND LAST RELOAD - GOLD WITH SINGARD - TOWARDS THE ESOPHAGUS, THE ECHELON POWERED STAPLER HAS A LOT OF DIFFICULTY ADVANCING THEN PROCEEDS EVEN IF THERE WAS SOMETHING STRANGE AND WHEN IT REOPENS THE SUTURE HAS NOT BEEN MADE BUT HAS ONLY CUT THE TISSUE AND THE RELOAD. THERE ARE SOME CROOKED OPEN CLIPS AND THE CARTRIDGE IS BROKEN. THE PROBE IS INTACT AND THEREFORE WAS NOT CAUGHT IN THE STAPLER. BEFORE THE EVENT, THE STAPLER HAD WORKED WELL: 1 BLACK AND 1 GREEN RELOADS WERE USED BOTH WITH SINGARD. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
08-12-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
25-12-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. USED SUTURE TO OVER-SEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
01-01-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN UROLOGY CASE, THE HANDLE WAS NOT RELEASING AFTER THE SECOND FIRING. IT BECAME JAMMED AND THEY NEEDED TO REPLACE THE DEVICE WITH A NEW ONE AND FINISH THE PROCEDURE. DEVICE IS AVAILABLE FOR RETURN. NO ADVERSE EVENTS WERE REPORTED.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
20-09-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A OPEN LOBECTOMY, AT THE TIME OF THORACIC VASCULAR APPLICATION, THE DEVICE WAS NOT ARTICULATING CORRECTLY. THE ARTICULATION LEVER MOVED THE JAWS LEFT AND RIGHT, HOWEVER, IT DID NOT HOLD ANY PREDETERMINED POSITIONS. SURGEON POSITIONED AND FIRED THE DEVICE WITHOUT ARTICULATING. ONCE STAPLER WAS FIRED AND REMOVED FROM THE TISSUE, THE STAPLE LINE WAS NOT COMPLETELY HEMOSTATIC. THE SURGEON USED A SUTURE AND VESSEL SEALING DEVICE TO CONTROL THE BLEEDING. SURGICAL TIME WAS EXTENDED FOR THREE TO FIVE MINUTES. PLI 10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE ARTICULATION FEATURE WAS NOT WORKING PROPERLY. THE INSTRUMENT WOULD ARTICULATE FULLY TO THE RIGHT AND LEFT BUT WOULD NOT LOCK INTO ANY POSITIONS.
20-09-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A OPEN LOBECTOMY, AT THE TIME OF THORACIC VASCULAR APPLICATION, THE DEVICE WAS NOT ARTICULATING CORRECTLY. THE ARTICULATION LEVER MOVED THE JAWS LEFT AND RIGHT, HOWEVER, IT DID NOT HOLD ANY PREDETERMINED POSITIONS. SURGEON POSITIONED AND FIRED THE DEVICE WITHOUT ARTICULATING. ONCE STAPLER WAS FIRED AND REMOVED FROM THE TISSUE, THE STAPLE LINE WAS NOT COMPLETELY HEMOSTATIC. THE SURGEON USED A SUTURE AND VESSEL SEALING DEVICE TO CONTROL THE BLEEDING. SURGICAL TIME WAS EXTENDED FOR THREE TO FIVE MINUTES. PLI 10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE ARTICULATION FEATURE WAS NOT WORKING PROPERLY. THE INSTRUMENT WOULD ARTICULATE FULLY TO THE RIGHT AND LEFT BUT WOULD NOT LOCK INTO ANY POSITIONS.
01-01-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE IT WASN´T POSSIBLE TO TRIGGER THE DEVICE. A NEW DEVICE WAS OPENED. THERE WERE NO CONSEQUENCES FOR THE PATIENT.
22-12-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE BLADE WAS STUCK ON FIRING.PROCEDURE WAS SUCCESSFULLY COMPLETED. NO PATIENT CONSEQUENCE.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
13-12-2022 | Malfunction | johnson & johnson medical (suzhou | 06-01-2023 | CARTRIDGE FINAL ASSY 60MM BLUE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND THE ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND 2ND HANDLE WAS USED WITH SAME RELOAD WAS USED. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RECLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON REENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A 2ND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND THE ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND 2ND HANDLE WAS USED WITH SAME RELOAD WAS USED. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RECLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON REENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A 2ND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
01-01-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE DID NOT RETURN AT 3RD FIRING. MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ONE PSEE60A AND ONE GST60D WILL BE RETURNING. NO FURTHER INFORMATION IS AVAILABLE. AFFILIATE REFERENCE NUMBER: (B)(4). PLEASE TAKE PHOTOS FOR JAPANESE CUSTOMER.
01-01-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE DID NOT RETURN AT 3RD FIRING. MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ONE PSEE60A AND ONE GST60D WILL BE RETURNING. NO FURTHER INFORMATION IS AVAILABLE. AFFILIATE REFERENCE NUMBER: (B)(4). PLEASE TAKE PHOTOS FOR JAPANESE CUSTOMER.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
24-02-2023 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION.
16-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE ADAPTER WAS BROKEN. THERE WAS NO PATIENT INVOLVEMENT.
16-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Break | A0401 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND THE ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND 2ND HANDLE WAS USED WITH SAME RELOAD WAS USED. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RECLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON REENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A 2ND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
16-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE INITIAL FIRING ON THE STOMACH NEAR THE PYLORUS, THE FIRST HANDLE WAS LOADED AND THE ABLE TO ARTICULATE THE RELOAD TO THE RIGHT, BUT NOT THE LEFT. HANDLE WAS UNLOADED AND 2ND HANDLE WAS USED WITH SAME RELOAD WAS USED. THE RELOAD WOULD NOT ARTICULATE AGAIN, SO A NEW RELOAD WAS USED AND WAS ABLE TO ARTICULATE BOTH WAYS. THE SURGEON SECURELY CLOSED THE BLACK HANDLE OVER THE STOMACH TISSUE AND PRESSED THE GREEN BUTTON TO FIRE. HOWEVER, THE HANDLE DID NOT FIRE STAPLES AND DID NOT ADVANCE THE KNIFE BLADE. THE RELOAD WAS UNCLAMPED AND REMOVED FROM THE ABDOMEN. DURING THAT TIME, A DRY FIRE WAS DONE AND THE RELOAD FIRED OVER THE FIELD. LOADED ANOTHER RELOAD AND THE SURGEON RECLAMPED THE STOMACH TISSUE WITH THE RELOAD AND PRESSED THE GREEN BUTTON TO FIRE. THE RELOAD DID NOT FIRE. THE SURGEON PULLED BACK ON THE TOP BLACK PIECE TO RESET THE HANDLE, SECURELY SQUEEZED AGAIN AND PRESSED THE GREEN BUTTON TO FIRE AND THE RELOAD FIRED SUCCESSFULLY. ANOTHER RELOAD WAS LOADED TO THE HANDLE, THE HANDLE FIRED, BUT THE HANDLE AND RELOAD WERE BROKEN AT THE JOINT OF THE RELOAD DUE TO THE SURGEON DID NOT HAVE ENOUGH ABDOMINAL SPACE AND WAS PUTTING A LOT OF PRESSURE ON THE TROCAR AND HANDLE IN ORDER TO GET THE APPROPRIATE FIRING ANGLE. ANOTHER HANDLE WAS RELOADED WITH A DIFFERENT TYPE OF RELOAD. THE SURGEON ENTERED THE ABDOMEN AND FOLLOWED ALONG THE PREVIOUS STAPLE LINE. HE SQUEEZED THE HANDLE AND SECURED THE TISSUE. HE PUSHED THE GREEN BUTTON TO FIRE AND THE RELOAD DID NOT FIRE. HE THEN PULLED IT OUT, DID A DRY FIRING OVER THE STERILE FIELD AND THE RELOAD FIRED SUCCESSFULLY. A NEW SECOND TYPE OF RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON REENTERED THE ABDOMEN, ARTICULATED THE HANDLE, AND SQUEEZED THE HANDLE TO SECURE THE RELOAD ON THE STOMACH. HE PUSHED THE GREEN BUTTON AND ATTEMPTED TO FIRE. THE HANDLE DID NOT FIRE. HANDLE WAS REMOVED TO DRY FIRE THE RELOAD AGAIN AND THE HANDLE DID SUCCESSFULLY FIRE THE RELOAD. THE SCRUB TECHNICIAN RESET THE HANDLE, SQUEEZED THE HANDLE FOR A 2ND ATTEMPT AT DRY FIRING AND IT WAS AN UNSUCCESSFUL ATTEMPT. THE HANDLE WAS RESET, SQUEEZED, GREEN BUTTON WAS PRESSED AND WAS FIRED SUCCESSFULLY. THE SURGICAL TIME WAS EXTENDED TO AN HOUR DUE TO TWO MISFIRED AND BROKEN HANDLES. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE.
13-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PULMONARY RESECTION, IN VESSEL RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PULMONARY RESECTION, IN VESSEL RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
25-12-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | BLUE RELOAD FOR ECHELON 60 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AFTER OPENING THE PACKING, NOTED THE CARTRIDGE DAMAGED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2022 | Malfunction | ethicon endo-surgery | 06-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN SLEEVE GASTRECTOMY PROCEDURE, ON THE FIST FIRING, THE TOP ANVIL BENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
20-10-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Premature End-of-Life Indicator | A160602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, A WRENCH SYMBOL APPEARED ON THE SCREEN AND THE HANDLE WAS UNABLE TO USE. REPLACEMENT WAS MADE TO FIX THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
20-10-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, A WRENCH SYMBOL APPEARED ON THE SCREEN AND THE HANDLE WAS UNABLE TO USE. REPLACEMENT WAS MADE TO FIX THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
23-11-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, THE INSTRUMENT NEEDED AN EXTRAORDINARY FORCE BUT HANDLE COULD STILL NOT BE SQUEEZED TO STAPLE THE TISSUE. ANOTHER RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
23-11-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, FIRING WAS IMPOSSIBLE, THE INSTRUMENT NEEDED AN EXTRAORDINARY FORCE FOR FIRING THE RELOAD AND STAPLING THE TISSUE. ANOTHER RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
30-09-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY BYPASS PROCEDURE, WHILE STAPLING GASTRIC, 6 FIRINGS WERE PERFORMED WITH THE HANDLE, AND BEFORE THE SEVENTH FIRING, THERE WAS LOUD STRANGE NOISE, THE HANDLE THEN GOT SOFT, THE RELOAD PRE-FIRED, BUT THE TRIGGER OF THE HANDLE BROKE AND FIRING WAS NOT POSSIBLE. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO SHEARED FIRING RACK TEETH DAMAGE.
30-09-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY BYPASS PROCEDURE, WHILE STAPLING GASTRIC, 6 FIRINGS WERE PERFORMED WITH THE HANDLE, AND BEFORE THE SEVENTH FIRING, THERE WAS LOUD STRANGE NOISE, THE HANDLE THEN GOT SOFT, THE RELOAD PRE-FIRED, BUT THE TRIGGER OF THE HANDLE BROKE AND FIRING WAS NOT POSSIBLE. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO SHEARED FIRING RACK TEETH DAMAGE.
17-10-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE HANDLE WOULD NOT CHARGE. RED LIGHT WAS NOTED ON THE CHARGER. WHEN REMOVED, GREEN LIGHT WAS FLASHING FROM THE CHARGER. THE HANDLE WAS REDOCKED ON THE CHARGER BUT THE PROBLEM WAS NOT RESOLVED. FLUID FROM THE POWER HANDLE WAS LEAKING INTO THE TERMINAL OF THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. NOTE: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
17-10-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Corroded | A040502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE HANDLE WOULD NOT CHARGE. RED LIGHT WAS NOTED ON THE CHARGER. WHEN REMOVED, GREEN LIGHT WAS FLASHING FROM THE CHARGER. THE HANDLE WAS REDOCKED ON THE CHARGER BUT THE PROBLEM WAS NOT RESOLVED. FLUID FROM THE POWER HANDLE WAS LEAKING INTO THE TERMINAL OF THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. NOTE: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
17-10-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE HANDLE WOULD NOT CHARGE. RED LIGHT WAS NOTED ON THE CHARGER. WHEN REMOVED, GREEN LIGHT WAS FLASHING FROM THE CHARGER. THE HANDLE WAS REDOCKED ON THE CHARGER BUT THE PROBLEM WAS NOT RESOLVED. FLUID FROM THE POWER HANDLE WAS LEAKING INTO THE TERMINAL OF THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. NOTE: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
19-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING AN OPEN HARTMANNS TAKEDOWN, WHILE DOING A SMALL BOWEL RESECTION, THE WHITE PLASTIC TIP ON THE ANVIL WAS NOT SMOOTHLY FIXATED TO THE METAL PORTION OF THE ANVIL AND TISSUE WAS GETTING STUCK IN THE EDGE BETWEEN THEM. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-10-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO REPORTER, DURING CHARGING, THE DEVICE WAS LOADED INTO THE BATTERY CHARGER, BUT THE GREEN LIGHT DID NOT BLINK, AND THE BATTERY COULD NOT BE CHARGED. THE GREEN LIGHT TURNED ON WHEN THE SALES DEMO UNIT HANDLE WAS LOADED. ANOTHER HANDLE FROM ANOTHER COMPANY WAS USED TO RESOLVE THE ISSUE. THE CHARGER WAS USED WITH OTHER HANDLE WITH NO PROBLEM. THERE WAS NO PATIENT INVOLVED. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
17-10-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO REPORTER, DURING CHARGING, THE DEVICE WAS LOADED INTO THE BATTERY CHARGER, BUT THE GREEN LIGHT DID NOT BLINK, AND THE BATTERY COULD NOT BE CHARGED. THE GREEN LIGHT TURNED ON WHEN THE SALES DEMO UNIT HANDLE WAS LOADED. ANOTHER HANDLE FROM ANOTHER COMPANY WAS USED TO RESOLVE THE ISSUE. THE CHARGER WAS USED WITH OTHER HANDLE WITH NO PROBLEM. THERE WAS NO PATIENT INVOLVED. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
28-09-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, ON CHARGER, THE HANDLE FAILED TO CHARGED. THE HANDLE DID NOT WORKED WITH ANY CHARGER. THE CHARGERS WORK FINE. TO RESOLVE THE ISSUE, A POWERED HANDLE FROM A DIFFERENT MANUFACTURER WAS USED AS THE USER ONLY HAVE ONE HANDLE. THERE WAS NO PATIENT INVOLVED. PLI10 - MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
28-09-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, ON CHARGER, THE HANDLE FAILED TO CHARGED. THE HANDLE DID NOT WORKED WITH ANY CHARGER. THE CHARGERS WORK FINE. TO RESOLVE THE ISSUE, A POWERED HANDLE FROM A DIFFERENT MANUFACTURER WAS USED AS THE USER ONLY HAVE ONE HANDLE. THERE WAS NO PATIENT INVOLVED. PLI10 - MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
09-12-2022 | Injury | us surgical puerto rico | 06-01-2023 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, A SUTURE TEST WAS PERFORMED AND THERE WAS CIRCUMFERENTIAL AIR LEAKAGE THROUGHOUT THE ANASTOMOSIS. IT WAS OBSERVED THAT THERE WAS A PARTIAL CLOSURE OF THE STAPLES. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
09-12-2022 | Injury | us surgical puerto rico | 06-01-2023 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, A SUTURE TEST WAS PERFORMED AND THERE WAS CIRCUMFERENTIAL AIR LEAKAGE THROUGHOUT THE ANASTOMOSIS. IT WAS OBSERVED THAT THERE WAS A PARTIAL CLOSURE OF THE STAPLES. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
09-12-2022 | Injury | us surgical puerto rico | 06-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, A SUTURE TEST WAS PERFORMED AND THERE WAS CIRCUMFERENTIAL AIR LEAKAGE THROUGHOUT THE ANASTOMOSIS. IT WAS OBSERVED THAT THERE WAS A PARTIAL CLOSURE OF THE STAPLES. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
09-12-2022 | Injury | us surgical puerto rico | 06-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, AFTER STAPLING, A SUTURE TEST WAS PERFORMED AND THERE WAS CIRCUMFERENTIAL AIR LEAKAGE THROUGHOUT THE ANASTOMOSIS. IT WAS OBSERVED THAT THERE WAS A PARTIAL CLOSURE OF THE STAPLES. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
09-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT ASSISTED LAPAROSCOPIC GASTRECTOMY, DURING PREPARATION STAGE, OPENING AND CLOSING CHECK WAS PERFORMED WITH THE YELLOW TAB ATTACHED, AND FRAGMENTS OF THE YELLOW TAB WAS ATTACHED ON THE CARTRIDGE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, WHILE ASSEMBLING THE STAPLER PRIOR TO USE ON PATIENT, THE STEM DID NOT WORK AND FIT THE RELOAD. ANOTHER MANUFACTURER PRODUCT WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, WHILE ASSEMBLING THE STAPLER PRIOR TO USE ON PATIENT, THE STEM DID NOT WORK AND FIT THE RELOAD. ANOTHER MANUFACTURER PRODUCT WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, WHILE ASSEMBLING THE STAPLER PRIOR TO USE ON PATIENT, THE STEM DID NOT WORK AND FIT THE RELOAD. ANOTHER MANUFACTURER PRODUCT WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, WHILE ASSEMBLING THE STAPLER PRIOR TO USE ON PATIENT, THE STEM DID NOT WORK AND FIT THE RELOAD. ANOTHER MANUFACTURER PRODUCT WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, WHILE ASSEMBLING THE STAPLER PRIOR TO USE ON PATIENT, THE STEM DID NOT WORK AND FIT THE RELOAD. ANOTHER MANUFACTURER PRODUCT WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, WHILE ASSEMBLING THE STAPLER PRIOR TO USE ON PATIENT, THE STEM DID NOT WORK AND FIT THE RELOAD. ANOTHER MANUFACTURER PRODUCT WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, WHILE ASSEMBLING THE STAPLER PRIOR TO USE ON PATIENT, THE STEM DID NOT WORK AND FIT THE RELOAD. ANOTHER MANUFACTURER PRODUCT WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL GASTRECTOMY, WHILE ASSEMBLING THE STAPLER PRIOR TO USE ON PATIENT, THE STEM DID NOT WORK AND FIT THE RELOAD. ANOTHER MANUFACTURER PRODUCT WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. THERE WAS NO PATIENT INJURY.
18-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY (LADG) PROCEDURE, WHILE PERFORMING GASTRIC BODY RESECTION, THE DEVICE DID NOT FIRE. THE RELOAD WAS UNLOADED AND LOADED AGAIN ON THE ADAPTER, AND WHEN IT WAS CHECKED ON THE TABLE, THE RELOAD PART AT THE RIGHT END DID NOT ILLUMINATE GREEN. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. WHEN THE INITIAL RELOAD WAS CHECKED, A TRACE OF FIRING WAS OBSERVED ON THE PROXIMAL SIDE. THERE WAS NO PATIENT INJURY.
18-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY (LADG) PROCEDURE, WHILE PERFORMING GASTRIC BODY RESECTION, THE DEVICE DID NOT FIRE. THE RELOAD WAS UNLOADED AND LOADED AGAIN ON THE ADAPTER, AND WHEN IT WAS CHECKED ON THE TABLE, THE RELOAD PART AT THE RIGHT END DID NOT ILLUMINATE GREEN. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. WHEN THE INITIAL RELOAD WAS CHECKED, A TRACE OF FIRING WAS OBSERVED ON THE PROXIMAL SIDE. THERE WAS NO PATIENT INJURY.
18-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY (LADG) PROCEDURE, WHILE PERFORMING GASTRIC BODY RESECTION, THE DEVICE DID NOT FIRE. THE RELOAD WAS UNLOADED AND LOADED AGAIN ON THE ADAPTER, AND WHEN IT WAS CHECKED ON THE TABLE, THE RELOAD PART AT THE RIGHT END DID NOT ILLUMINATE GREEN. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. WHEN THE INITIAL RELOAD WAS CHECKED, A TRACE OF FIRING WAS OBSERVED ON THE PROXIMAL SIDE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE REPORTER, THE SURGERY WAS A SUCCESS DURING A LAPAROSCOPIC GASTROPLASTY, AND ALL THE STAPLING WAS PERFECT. AFTER TWO DAYS, THE STAPLE LINE OPENED, AND THE PATIENT HAD AN ABSCESS AND INTERNAL BLEEDING. THE PATIENT WAS RETURNED TO THE OPERATING ROOM ON THE FIFTH DAY FOR RE-APPROACH. THE SURGEON RE-SUTURED MANUALLY THE STOMACH SITE THAT OPENED AND CLEANED THE ABSCESS. PATIENT HOSPITALIZATION WAS EXTENDED AND A THERE WAS A TISSUE DAMAGED. THE PATIENT HAS A MEDICAL HISTORY OF OBESITY. THE PATIENT BECAME WELL AFTER THE SECOND SURGERY.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, THE SURGERY WAS A SUCCESS DURING A LAPAROSCOPIC GASTROPLASTY, AND ALL THE STAPLING WAS PERFECT. AFTER TWO DAYS, THE STAPLE LINE OPENED, AND THE PATIENT HAD AN ABSCESS AND INTERNAL BLEEDING. THE PATIENT WAS RETURNED TO THE OPERATING ROOM ON THE FIFTH DAY FOR RE-APPROACH. THE SURGEON RE-SUTURED MANUALLY THE STOMACH SITE THAT OPENED AND CLEANED THE ABSCESS. PATIENT HOSPITALIZATION WAS EXTENDED AND A THERE WAS A TISSUE DAMAGED. THE PATIENT HAS A MEDICAL HISTORY OF OBESITY. THE PATIENT BECAME WELL AFTER THE SECOND SURGERY.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, THE SURGERY WAS A SUCCESS DURING A LAPAROSCOPIC GASTROPLASTY, AND ALL THE STAPLING WAS PERFECT. AFTER TWO DAYS, THE STAPLE LINE OPENED, AND THE PATIENT HAD AN ABSCESS AND INTERNAL BLEEDING. THE PATIENT WAS RETURNED TO THE OPERATING ROOM ON THE FIFTH DAY FOR RE-APPROACH. THE SURGEON RE-SUTURED MANUALLY THE STOMACH SITE THAT OPENED AND CLEANED THE ABSCESS. PATIENT HOSPITALIZATION WAS EXTENDED AND A THERE WAS A TISSUE DAMAGED. THE PATIENT HAS A MEDICAL HISTORY OF OBESITY. THE PATIENT BECAME WELL AFTER THE SECOND SURGERY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE FIRING, THE BUTTON WOULD NOT CLICK AND THE STAPLER WOULD NOT TURN. IT WAS ALSO NOTED THAT THE CLAMSHELL WAS NOT RECOGNIZED. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE FIRING, THE BUTTON WOULD NOT CLICK AND THE STAPLER WOULD NOT TURN. IT WAS ALSO NOTED THAT THE CLAMSHELL WAS NOT RECOGNIZED. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | SIGNIA | Material Split, Cut or Torn | A0414 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE FIRING, THE BUTTON WOULD NOT CLICK AND THE STAPLER WOULD NOT TURN. IT WAS ALSO NOTED THAT THE CLAMSHELL WAS NOT RECOGNIZED. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE FIRING, THE BUTTON WOULD NOT CLICK AND THE STAPLER WOULD NOT TURN. IT WAS ALSO NOTED THAT THE CLAMSHELL WAS NOT RECOGNIZED. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE FIRING, THE BUTTON WOULD NOT CLICK AND THE STAPLER WOULD NOT TURN. IT WAS ALSO NOTED THAT THE CLAMSHELL WAS NOT RECOGNIZED. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE FIRING, THE BUTTON WOULD NOT CLICK AND THE STAPLER WOULD NOT TURN. IT WAS ALSO NOTED THAT THE CLAMSHELL WAS NOT RECOGNIZED. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE FIRING, THE BUTTON WOULD NOT CLICK AND THE STAPLER WOULD NOT TURN. IT WAS ALSO NOTED THAT THE CLAMSHELL WAS NOT RECOGNIZED. A NEW SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | SIGNIA | Device Displays Incorrect Message | A090201 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | SIGNIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
20-12-2022 | Injury | covidien lp llc north haven | 06-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE SURGERY, ON THE SECOND FIRING ALONG THE ANTRUM OF THE STOMACH, THE PURPLE RELOAD WAS CLOSED ON TISSUE BUT WAS REMOVED NORMALLY DUE TO THE HANDLE INDICATING ZONE 3 EVEN THOUGH THE TISSUE WAS NOT EXTRA THICK. THE SURGEON THEN REMOVED THE PURPLE RELOAD AND PUT ON AN EXTRA THICK RELOAD AND SAME ADAPTER. THE STAPLER FIRED THE EXTRA THICK RELOAD BUT IT STOPPED AND ONLY FIRED HALFWAY AND THERE INCOMPLETE STAPLE LINE ON THE PROXIMAL AREA. THE JAWS ALSO LOCKED ON TISSUE AND MANUAL EXCISION WAS DONE TO REMOVE THE RELOAD. IT WAS ALSO NOTED THAT THERE WERE STAPLES THAT DID NOT FORM INTO B SHAPE. TO COMPLETE THE PROCEDURE, A MANUAL HANDLE AND RELOAD WAS USED AND ADDED A LINE OF SUTURES FOR ADDITIONAL SECURITY AND TO PREVENT ADDITIONAL POTENTIAL BLEED SURGEON ADDED A ROW OF SUTURES TO THE AFFECTED AREA. TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS WERE NOTED. THE SURGICAL TIME WAS ALSO EXTENDED FOR 30 MINUTES OR MORE.
15-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATIC BODY AND TAIL RESECTION, IN PANCREAS RESECTION, THE PANCREAS WAS GRASPED AND WAITED FOR ABOUT A MINUTE FOR 1 CM RESECTION. RESECTION BY 1 CM AFTER ABOUT 1 MINUTE WAS REPEATEDLY DONE, BUT WHILE WAITING, THE KNIFE AUTOMATICALLY RETURNED AND RELEASED EVEN WHEN IT WAS NOT OPERATED. THERE WAS NO PROBLEM WITH THE STAPLE LINE. TO RESOLVED THE ISSUE, THE RELOAD WAS REPLACED AS WELL AS THE SHELL JUST IN CASE, AND PROCEDURE WAS CONTINUED. AFTER THAT, THERE WERE NO PROBLEM. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PULMONARY RESECTION PROCEDURE, WHILE ATTEMPTING TO RESECTION THE INFERIOR PULMONARY VEIN, THE STAPLE DID NOT GO ALL THE WAY AND STIFFENED. THEN JAWS COULD NOT BE OPENED ANYMORE; THE BLACK KNOB COULD NOT BE PULLED UNTIL THE END. THE SURGEON USED FORCEPS TO FORCIBLY OPEN THE JAWS; AND A NEW HANDLE AND RELOAD WERE USED TO PERFORM THE RESECTION. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PULMONARY RESECTION PROCEDURE, WHILE ATTEMPTING TO RESECTION THE INFERIOR PULMONARY VEIN, THE STAPLE DID NOT GO ALL THE WAY AND STIFFENED. THEN JAWS COULD NOT BE OPENED ANYMORE; THE BLACK KNOB COULD NOT BE PULLED UNTIL THE END. THE SURGEON USED FORCEPS TO FORCIBLY OPEN THE JAWS; AND A NEW HANDLE AND RELOAD WERE USED TO PERFORM THE RESECTION. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LAPAROTOMY ON SMALL BOWEL RESECTION PROCEDURE, SOME FUNNY PIECE OF METAL STICKING OUT OF THE STAPLER AND COULD NOT BE REMOVED TO RELOAD ANOTHER CARTRIDGE. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | us surgical puerto rico | 06-01-2023 | GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LAPAROTOMY ON SMALL BOWEL RESECTION PROCEDURE, SOME FUNNY PIECE OF METAL STICKING OUT OF THE STAPLER AND COULD NOT BE REMOVED TO RELOAD ANOTHER CARTRIDGE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC SWITCH REVISION TO ROUX-EN-Y DUODENOENTEROSTOMY, DURING STAPLING ON SMALL BOWEL, THE DEVICE LOCKED ON TISSUE. THE JAWS DID NOT OPEN AND THE KNIFE DID NOT RETRACT AFTER FIRING. THE POWERED HANDLE WAS TURNED OFF AND THEN ON AGAIN, AND DURING THE BOOT-UP CALIBRATION, THE KNIFE BLADE WAS RETRACTED AND THE JAWS OPENED. ANOTHER RELOAD FROM THE SAME BATCH WAS FIRED IMMEDIATELY AFTERWARD WITH NO ISSUES. THERE WAS NO PATIENT INJURY.
08-12-2022 | Injury | intuitive surgical | 06-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
08-12-2022 | Injury | intuitive surgical | 06-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
08-12-2022 | Injury | intuitive surgical | 06-01-2023 | SUREFORM | Failure to Cut | A050702 | Hemorrhage/Blood Loss/Bleeding | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
08-12-2022 | Injury | intuitive surgical | 06-01-2023 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
06-12-2022 | Malfunction | intuitive surgical | 06-01-2023 | SUREFORM | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SURGICAL PROCEDURE, THE CURVED-TIP SUREFORM 45 STAPLER MOVED WHEN THE SURGEON HAD NOT MOVED IT AT THE CONSOLE. THE PROCEDURE WAS COMPLETED AS PLANNED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) MADE MULTIPLE FOLLOW-UP ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION, HOWEVER, NO FURTHER DETAILS HAVE BEEN RECEIVED AS OF THE DATE OF THIS REPORT.
26-04-2023 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND THE CUFF WAS SQUEEZED. THE EQUIPMENT THEN BROKE DOWN, AND THERE WAS NO INFORMATION ON WHICH PART BROKE DOWN OR ANY DETAILS ABOUT WHAT HAPPENED. IT WAS NOTED THAT NO PART OF THE EQUIPMENT FELL ON THE PATIENT. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE AND THE SURGERY WAS SUCCESSFUL. THE HANDLE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-04-2023 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND THE CUFF WAS SQUEEZED. THE EQUIPMENT THEN BROKE DOWN, AND THERE WAS NO INFORMATION ON WHICH PART BROKE DOWN OR ANY DETAILS ABOUT WHAT HAPPENED. IT WAS NOTED THAT NO PART OF THE EQUIPMENT FELL ON THE PATIENT. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE AND THE SURGERY WAS SUCCESSFUL. THE HANDLE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-01-2023 | Malfunction | us surgical puerto rico | 06-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE WAS BROKEN. THE ARTICULATION PART OF JOINT TO THE RELOAD WAS BROKEN.
06-12-2022 | Malfunction | covidien lp llc north haven | 06-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SIGMOID COLECTOMY, WHEN THE COLON WAS RESECTED IN THE BODY CAVITY, THE HANDLE WAS USED UP TO THE SECOND FIRING, BUT THE THIRD FIRING COULD NOT BE PERFORMED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-11-2022 | Malfunction | covidien lp llc north haven | 08-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE RELOAD WAS CONNECTED TO THE HANDLE, BUT WOULD NOT FIRE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
14-12-2022 | Malfunction | us surgical puerto rico | 08-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN THE RELOAD WAS APPLIED TO THE TISSUE, THE HANDLE DID NOT FIRE THE RELOAD. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS OPENED AND USED THE SAME RELOAD TO RESOLVE THE ISSUE. NO PATIENT INJURY.
21-12-2022 | Injury | covidien lp llc north haven | 08-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE AT BLOOD VESSEL, THE DEVICE FIRED, STAPLE LINE WAS INCOMPLETE PROXIMALLY AND THERE WAS POOR STAPLE FORMATION. THE RELOAD WAS REPLACED TO FIX THE STAPLE LINE, THAT RESOLVE THE ISSUE AND COMPLETE THE CASE.
21-12-2022 | Injury | covidien lp llc north haven | 08-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY TRANSVERSE COLONIC STOMA CLOSURE PROCEDURE, AT THE TIME OF FIRING THE SECOND RELOAD DURING THE COLONIC FIRING NEAR THE STOMA, A CRACKING SOUND WAS HEARD FROM THE HANDLE AND THEN THE FIRING STOPPED MIDWAY. WHEN THE SURGEON TRIED RELEASING THE DEVICE, IT DID NOT MOVE. AFTER SEVERAL ATTEMPTS, THE DEVICE WAS ABLE TO BE RELEASED. THE STAPLE LINE WAS THEN CHECKED AND IT WAS OBSERVED THAT THE FIRED STAPLE NAILS WERE NOT FORMED AND THE PROXIMAL END WAS INCOMPLETE. WHEN THE SPECIMEN WAS CHECKED FOR ANY FOREIGN OBJECT, IT WAS FOUND OUT THAT THE TISSUE SUDDENLY BECAME THICK FROM THIN PART IN THE DIRECTION OF THE STAPLE FIRING. THE SURGEON PERFORMED A RESECTION USING A POWERED HANDLE AND A NEW RELOAD TO RESOLVE THE ISSUE.
21-12-2022 | Injury | covidien lp llc north haven | 08-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY TRANSVERSE COLONIC STOMA CLOSURE PROCEDURE, AT THE TIME OF FIRING THE SECOND RELOAD DURING THE COLONIC FIRING NEAR THE STOMA, A CRACKING SOUND WAS HEARD FROM THE HANDLE AND THEN THE FIRING STOPPED MIDWAY. WHEN THE SURGEON TRIED RELEASING THE DEVICE, IT DID NOT MOVE. AFTER SEVERAL ATTEMPTS, THE DEVICE WAS ABLE TO BE RELEASED. THE STAPLE LINE WAS THEN CHECKED AND IT WAS OBSERVED THAT THE FIRED STAPLE NAILS WERE NOT FORMED AND THE PROXIMAL END WAS INCOMPLETE. WHEN THE SPECIMEN WAS CHECKED FOR ANY FOREIGN OBJECT, IT WAS FOUND OUT THAT THE TISSUE SUDDENLY BECAME THICK FROM THIN PART IN THE DIRECTION OF THE STAPLE FIRING. THE SURGEON PERFORMED A RESECTION USING A POWERED HANDLE AND A NEW RELOAD TO RESOLVE THE ISSUE.
21-12-2022 | Injury | covidien lp llc north haven | 08-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY TRANSVERSE COLONIC STOMA CLOSURE PROCEDURE, AT THE TIME OF FIRING THE SECOND RELOAD DURING THE COLONIC FIRING NEAR THE STOMA, A CRACKING SOUND WAS HEARD FROM THE HANDLE AND THEN THE FIRING STOPPED MIDWAY. WHEN THE SURGEON TRIED RELEASING THE DEVICE, IT DID NOT MOVE. AFTER SEVERAL ATTEMPTS, THE DEVICE WAS ABLE TO BE RELEASED. THE STAPLE LINE WAS THEN CHECKED AND IT WAS OBSERVED THAT THE FIRED STAPLE NAILS WERE NOT FORMED AND THE PROXIMAL END WAS INCOMPLETE. WHEN THE SPECIMEN WAS CHECKED FOR ANY FOREIGN OBJECT, IT WAS FOUND OUT THAT THE TISSUE SUDDENLY BECAME THICK FROM THIN PART IN THE DIRECTION OF THE STAPLE FIRING. THE SURGEON PERFORMED A RESECTION USING A POWERED HANDLE AND A NEW RELOAD TO RESOLVE THE ISSUE.
20-12-2022 | Malfunction | us surgical puerto rico | 08-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL SLEEVE GASTRECTOMY, IN STAPLING THE STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | covidien lp llc north haven | 08-01-2023 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY AND DUODENAL RESECTION PROCEDURE, WHILE RESECTIONING THE DUODENAL, THE DEVICE STOPPED FIRING MIDWAY. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2022 | Injury | us surgical puerto rico | 08-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY TRANSVERSE COLONIC STOMA CLOSURE PROCEDURE, AT THE TIME OF FIRING THE SECOND RELOAD DURING THE COLONIC FIRING NEAR THE STOMA, A CRACKING SOUND WAS HEARD FROM THE HANDLE AND THEN THE FIRING STOPPED MIDWAY. WHEN THE SURGEON TRIED RELEASING THE DEVICE, IT DID NOT MOVE. AFTER SEVERAL ATTEMPTS, THE DEVICE WAS ABLE TO BE RELEASED. THE STAPLE LINE WAS THEN CHECKED AND IT WAS OBSERVED THAT THE FIRED STAPLE NAILS WERE NOT FORMED AND THE PROXIMAL END WAS INCOMPLETE. WHEN THE SPECIMEN WAS CHECKED FOR ANY FOREIGN OBJECT, IT WAS FOUND OUT THAT THE TISSUE SUDDENLY BECAME THICK FROM THIN PART IN THE DIRECTION OF THE STAPLE FIRING. THE SURGEON PERFORMED A RESECTION USING A POWERED HANDLE AND A NEW RELOAD TO RESOLVE THE ISSUE.
21-12-2022 | Injury | us surgical puerto rico | 08-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY TRANSVERSE COLONIC STOMA CLOSURE PROCEDURE, AT THE TIME OF FIRING THE SECOND RELOAD DURING THE COLONIC FIRING NEAR THE STOMA, A CRACKING SOUND WAS HEARD FROM THE HANDLE AND THEN THE FIRING STOPPED MIDWAY. WHEN THE SURGEON TRIED RELEASING THE DEVICE, IT DID NOT MOVE. AFTER SEVERAL ATTEMPTS, THE DEVICE WAS ABLE TO BE RELEASED. THE STAPLE LINE WAS THEN CHECKED AND IT WAS OBSERVED THAT THE FIRED STAPLE NAILS WERE NOT FORMED AND THE PROXIMAL END WAS INCOMPLETE. WHEN THE SPECIMEN WAS CHECKED FOR ANY FOREIGN OBJECT, IT WAS FOUND OUT THAT THE TISSUE SUDDENLY BECAME THICK FROM THIN PART IN THE DIRECTION OF THE STAPLE FIRING. THE SURGEON PERFORMED A RESECTION USING A POWERED HANDLE AND A NEW RELOAD TO RESOLVE THE ISSUE.
21-12-2022 | Injury | us surgical puerto rico | 08-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY TRANSVERSE COLONIC STOMA CLOSURE PROCEDURE, AT THE TIME OF FIRING THE SECOND RELOAD DURING THE COLONIC FIRING NEAR THE STOMA, A CRACKING SOUND WAS HEARD FROM THE HANDLE AND THEN THE FIRING STOPPED MIDWAY. WHEN THE SURGEON TRIED RELEASING THE DEVICE, IT DID NOT MOVE. AFTER SEVERAL ATTEMPTS, THE DEVICE WAS ABLE TO BE RELEASED. THE STAPLE LINE WAS THEN CHECKED AND IT WAS OBSERVED THAT THE FIRED STAPLE NAILS WERE NOT FORMED AND THE PROXIMAL END WAS INCOMPLETE. WHEN THE SPECIMEN WAS CHECKED FOR ANY FOREIGN OBJECT, IT WAS FOUND OUT THAT THE TISSUE SUDDENLY BECAME THICK FROM THIN PART IN THE DIRECTION OF THE STAPLE FIRING. THE SURGEON PERFORMED A RESECTION USING A POWERED HANDLE AND A NEW RELOAD TO RESOLVE THE ISSUE.
20-12-2022 | Malfunction | covidien lp llc north haven | 08-01-2023 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC UMBILICAL HERNIA PROCEDURE, THE DEVICE ONLY HAD 15 TACKERS IN IT, NORMALLY IT SHOULD HAVE 30 PIECES OF TACKER IN IT. THERE WERE NO TACKS ABLE TO PENETRATE AND HOLD CORRECTLY ONTO THE TISSUE. A NEW DEVICE WITH SAME LOT WITH SAME TACKERS (15) WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 08-01-2023 | ABSORBATACK | Component Missing | A020602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC UMBILICAL HERNIA PROCEDURE, THE DEVICE ONLY HAD 15 TACKERS IN IT, NORMALLY IT SHOULD HAVE 30 PIECES OF TACKER IN IT. THERE WERE NO TACKS ABLE TO PENETRATE AND HOLD CORRECTLY ONTO THE TISSUE. A NEW DEVICE WITH SAME LOT WITH SAME TACKERS (15) WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | covidien lp llc north haven | 08-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY AND DUODENAL RESECTION PROCEDURE, WHILE RESECTIONING THE DUODENAL, THE DEVICE STOPPED FIRING MIDWAY. THE SURGEON USED A NEW HANDLE AND RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-06-2022 | Death | covidien lp llc north haven | 08-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY COMPARED OUTCOMES IN PATIENTS WHO UNDERWENT HYBRID ROBOT-ASSISTED IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2015 AND 2019. TWO GROUPS WERE COMPARED: 21 PATIENTS IN LINEAR STAPLING AND 32 PATIENTS IN CONTINUOUS SUTURE GASTROESOPHAGEAL ANASTOMOSIS. ALL PATIENTS UNDERWENT A STANDARDIZED NEOADJUVANT CHEMORADIOTHERAPY (NCRT) TREATMENT. THE GASTRIC CONDUIT WAS CREATED BY USING ENDO GIA IN ALL PATIENTS. COMPETITOR DEVICES WERE USED FOR THE GASTROESOPHAGEAL ANASTOMOSIS. THERE WERE 53 PATIENTS IN THE STUDY AND ALL OF THEM WERE ADMITTED TO THE INTENSIVE CARE UNIT POST-OPERATIVE. ONE PATIENT DEVELOPED NECROSIS OF THE CONDUIT. THE PATIENT HAD AN ANASTOMOTIC LEAK, COMPLICATED BY PLEURAL EMPYEMA AND UNDERWENT STENT APPLICATION AS WELL AS VIDEO ASSISTED THORACOSCOPIC SURGERY. THE ANASTOMOSIS WAS RESECTION DUE TO THE CONDUIT NECROSIS. THE PATIENT LATER DEVELOPED A SEVERE SEPSIS AND DECEASED DUE TO MULTIPLE ORGAN DYSFUNCTION. COMPLICATIONS RELATED TO THE GASTROESOPHAGEAL ANASTOMOSIS, WHICH INCLUDED LEAK AND STENOSIS, DYSPHAGIA, DYSPNEA AND PLEURAL EMPYEMA WERE NOT RELATED TO THE DEVICE.
24-06-2022 | Death | covidien lp llc north haven | 08-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY COMPARED OUTCOMES IN PATIENTS WHO UNDERWENT HYBRID ROBOT-ASSISTED IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2015 AND 2019. TWO GROUPS WERE COMPARED: 21 PATIENTS IN LINEAR STAPLING AND 32 PATIENTS IN CONTINUOUS SUTURE GASTROESOPHAGEAL ANASTOMOSIS. ALL PATIENTS UNDERWENT A STANDARDIZED NEOADJUVANT CHEMORADIOTHERAPY (NCRT) TREATMENT. THE GASTRIC CONDUIT WAS CREATED BY USING ENDO GIA IN ALL PATIENTS. COMPETITOR DEVICES WERE USED FOR THE GASTROESOPHAGEAL ANASTOMOSIS. THERE WERE 53 PATIENTS IN THE STUDY AND ALL OF THEM WERE ADMITTED TO THE INTENSIVE CARE UNIT POST-OPERATIVE. ONE PATIENT DEVELOPED NECROSIS OF THE CONDUIT. THE PATIENT HAD AN ANASTOMOTIC LEAK, COMPLICATED BY PLEURAL EMPYEMA AND UNDERWENT STENT APPLICATION AS WELL AS VIDEO ASSISTED THORACOSCOPIC SURGERY. THE ANASTOMOSIS WAS RESECTION DUE TO THE CONDUIT NECROSIS. THE PATIENT LATER DEVELOPED A SEVERE SEPSIS AND DECEASED DUE TO MULTIPLE ORGAN DYSFUNCTION. COMPLICATIONS RELATED TO THE GASTROESOPHAGEAL ANASTOMOSIS, WHICH INCLUDED LEAK AND STENOSIS, DYSPHAGIA, DYSPNEA AND PLEURAL EMPYEMA WERE NOT RELATED TO THE DEVICE.
24-06-2022 | Death | covidien lp llc north haven | 08-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY COMPARED OUTCOMES IN PATIENTS WHO UNDERWENT HYBRID ROBOT-ASSISTED IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2015 AND 2019. TWO GROUPS WERE COMPARED: 21 PATIENTS IN LINEAR STAPLING AND 32 PATIENTS IN CONTINUOUS SUTURE GASTROESOPHAGEAL ANASTOMOSIS. ALL PATIENTS UNDERWENT A STANDARDIZED NEOADJUVANT CHEMORADIOTHERAPY (NCRT) TREATMENT. THE GASTRIC CONDUIT WAS CREATED BY USING ENDO GIA IN ALL PATIENTS. COMPETITOR DEVICES WERE USED FOR THE GASTROESOPHAGEAL ANASTOMOSIS. THERE WERE 53 PATIENTS IN THE STUDY AND ALL OF THEM WERE ADMITTED TO THE INTENSIVE CARE UNIT POST-OPERATIVE. ONE PATIENT DEVELOPED NECROSIS OF THE CONDUIT. THE PATIENT HAD AN ANASTOMOTIC LEAK, COMPLICATED BY PLEURAL EMPYEMA AND UNDERWENT STENT APPLICATION AS WELL AS VIDEO ASSISTED THORACOSCOPIC SURGERY. THE ANASTOMOSIS WAS RESECTION DUE TO THE CONDUIT NECROSIS. THE PATIENT LATER DEVELOPED A SEVERE SEPSIS AND DECEASED DUE TO MULTIPLE ORGAN DYSFUNCTION. COMPLICATIONS RELATED TO THE GASTROESOPHAGEAL ANASTOMOSIS, WHICH INCLUDED LEAK AND STENOSIS, DYSPHAGIA, DYSPNEA AND PLEURAL EMPYEMA WERE NOT RELATED TO THE DEVICE.
14-12-2022 | Malfunction | covidien lp | 09-01-2023 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: THE SMALL GRAY TAB, WHICH SHOULD BE FIXED TO THE PURPLE SECTION OF RELOAD BY A BLUE CLIP, WAS NOT ATTACHED, OR SECURED TO THE STAPLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | PROXIMATE*RL TITAN HEAVY WIRE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOTOMY, THE DEVICE DIDN¿T MAKE THE STAPLE LINE AT ALL. THE HEART AND THORACIC SURGEON HAS USED THIS DEVICE MORE THAN 15 YEARS. THIS WAS A FIRST TIME, WHEN IT MADE THE "TRICK"! THE SURGERY WAS DELAYED BY 15-MINUTES. THEY USED AN ENDOCUTTER WHEN THIS EVENT OCCURRED. FRAGMENTS WERE GENERATED BUT REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
20-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMA RESECTION PROCEDURE, WHILE STAPLING THE SIGMA, THE SURGEON PRESSED THE GREEN BUTTON BUT THE DEVICE DID NOT FIRE. THE SAME ISSUE OCCURRED ON THE SECOND RELOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMA RESECTION PROCEDURE, WHILE STAPLING THE SIGMA, THE SURGEON PRESSED THE GREEN BUTTON BUT THE DEVICE DID NOT FIRE. THE SAME ISSUE OCCURRED ON THE SECOND RELOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
26-10-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | 55MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY, THE GUN GOT DAMAGED IN FIRST CASE, FIRING KNOB NOT WORKING - NOT COMING ABOVE WHILE THE SURGEON TRYING TO FIRE. NO FRAGMENTS WERE GENERATED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
20-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMA RESECTION PROCEDURE, WHILE STAPLING THE SIGMA, THE SURGEON PRESSED THE GREEN BUTTON BUT THE DEVICE DID NOT FIRE. THE SAME ISSUE OCCURRED ON THE SECOND RELOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
21-12-2022 | Injury | us surgical puerto rico | 09-01-2023 | GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE RADICAL CARDIACECTOMY, WHILE DETACHING THE ESOPHAGUS, THERE WAS POOR STAPLE FORMATION. THE STAPLES DID NOT FORM INTO B SHAPE. OVER SEWING USING ADDITIONAL SUTURE WAS PERFORMED TO RESOLVE THE ISSUE.
08-05-2023 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE DEVICE DID NOT FIRE. ADDITIONALLY, THE RELOAD BROKE OFF. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
08-05-2023 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE DEVICE DID NOT FIRE. ADDITIONALLY, THE RELOAD BROKE OFF. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
14-12-2022 | Injury | covidien lp llc north haven | 09-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, INTRALUMINAL GASTROINTESTINAL BLEEDING WAS NOTED THREE DAYS POSTOPERATIVELY TO AN OPEN EXTENDED RIGHT COLECTOMY, ON SIDE TO SIDE ANASTOMOSIS. THE EVENT LEAD TO AN EXTENDED HOSPITALIZATION OF TWO DAYS AND REQUIRED TRANSFUSION OF 1 UNIT OF BLOOD. THE ISSUE RESOLVED ITSELF AND NO FURTHER ACTIONS WERE NECESSARY. PATIENT HAS SINCE GONE HOME, BEEN FOLLOWED UP ON, AND NO OTHER ISSUES EXIST. THE PATIENT IS NOTED TO BE OLDER AND HAS A LOW BODY MASS INDEX WITH A COLON BEING RESECTED FOR CANCER.
14-12-2022 | Injury | covidien lp llc north haven | 09-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, INTRALUMINAL GASTROINTESTINAL BLEEDING WAS NOTED THREE DAYS POSTOPERATIVELY TO AN OPEN EXTENDED RIGHT COLECTOMY, ON SIDE TO SIDE ANASTOMOSIS. THE EVENT LEAD TO AN EXTENDED HOSPITALIZATION OF TWO DAYS AND REQUIRED TRANSFUSION OF 1 UNIT OF BLOOD. THE ISSUE RESOLVED ITSELF AND NO FURTHER ACTIONS WERE NECESSARY. PATIENT HAS SINCE GONE HOME, BEEN FOLLOWED UP ON, AND NO OTHER ISSUES EXIST. THE PATIENT IS NOTED TO BE OLDER AND HAS A LOW BODY MASS INDEX WITH A COLON BEING RESECTED FOR CANCER.
14-12-2022 | Injury | covidien lp llc north haven | 09-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Gastrointestinal Hemorrhage | Event Description: ACCORDING TO THE REPORTER, INTRALUMINAL GASTROINTESTINAL BLEEDING WAS NOTED THREE DAYS POSTOPERATIVELY TO AN OPEN EXTENDED RIGHT COLECTOMY, ON SIDE TO SIDE ANASTOMOSIS. THE EVENT LEAD TO AN EXTENDED HOSPITALIZATION OF TWO DAYS AND REQUIRED TRANSFUSION OF 1 UNIT OF BLOOD. THE ISSUE RESOLVED ITSELF AND NO FURTHER ACTIONS WERE NECESSARY. PATIENT HAS SINCE GONE HOME, BEEN FOLLOWED UP ON, AND NO OTHER ISSUES EXIST. THE PATIENT IS NOTED TO BE OLDER AND HAS A LOW BODY MASS INDEX WITH A COLON BEING RESECTED FOR CANCER.
14-12-2022 | Injury | covidien lp llc north haven | 09-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, INTRALUMINAL GASTROINTESTINAL BLEEDING WAS NOTED THREE DAYS POSTOPERATIVELY TO AN OPEN EXTENDED RIGHT COLECTOMY, ON SIDE TO SIDE ANASTOMOSIS. THE EVENT LEAD TO AN EXTENDED HOSPITALIZATION OF TWO DAYS AND REQUIRED TRANSFUSION OF 1 UNIT OF BLOOD. THE ISSUE RESOLVED ITSELF AND NO FURTHER ACTIONS WERE NECESSARY. PATIENT HAS SINCE GONE HOME, BEEN FOLLOWED UP ON, AND NO OTHER ISSUES EXIST. THE PATIENT IS NOTED TO BE OLDER AND HAS A LOW BODY MASS INDEX WITH A COLON BEING RESECTED FOR CANCER.
07-12-2022 | Injury | us surgical puerto rico | 09-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POSTOPERATIVE TO ROBOTIC ASCENDING CANCER COLECTOMY, AT RECOVERY, THE PATIENT HAD INTRALUMINAL GASTROINTESTINAL BLEED. ONE UNIT BLOOD TRANSFUSION WAS REQUIRED AND HOSPITALIZATION WAS EXTENDED BY TWO DAYS AS A RESULT OF THE EVENT. PATIENT HAD SINCE GONE HOME, BEEN FOLLOWED UP ON AND NO OTHER ISSUES EXIST.
02-12-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP SIGMA PROCEDURE, THE DEVICE DOESN'T WORK. WHEN THE SURGEON TRIED TO FIRE THE INSTRUMENT TO MAKE ANASTOMOTIC STAPLE LINE THE LIGHT WAS ON IN THE INDICATOR: THEN SURGEON PRESSED THE TRIGGER NOTHING HAPPENED. THE ANVIL WAS PROPERLY CLOCKED. FRAGMENTS WERE GENERATED BUT REMOVED EASILY WITHOUT INTERVENTION. THE PROCEDURE WAS DELAYED BY 20 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH NO PATIENT CONSEQUENCES.
10-12-2022 | Injury | covidien lp llc north haven | 09-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, THREE DAYS POST-OPERATIVELY OF ROBOTIC ASCENDING COLECTOMY WHERE THE DEVICE WAS USED FOR SIDE-TO-SIDE ANASTOMOSIS FOR COLON CANCER RESECTION, THE PATIENT HAD AN INTRALUMINAL GASTRO-INTESTINAL BLEEDING, REQUIRING ONE UNIT OF BLOOD TO BE TRANSFUSED. THE ISSUE RESOLVED ITSELF WITH NO FURTHER ACTIONS NECESSARY. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR 2 DAYS.
14-12-2022 | Injury | ethicon endo-surgery | 09-01-2023 | ILS 29MM, CURVED | Failure to Cut | A050702 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT A BRAND NEW CDH29A STAPLER WAS OPENED AND USED FOR THE CIRCULAR ANASTOMOSIS DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE. SALES REP WAS PRESENT DURING THE PROCEDURE TO GUIDE THE SURGEONS THROUGH THE PROCESS OF USING THE CIRCULAR STAPLER. DURING THE FIRING SEQUENCE, THE SURGEON FIRING THE DEVICE WAS INSTRUCTED TO SQUEEZE THE STAPLER FIRING HANDLE UNTIL IT TOUCHES THE BODY OF THE DEVICE (PLASTIC TO PLASTIC), AND TO ONLY RELEASE THE FIRING HANDLE AFTER A CRUNCH NOISE IS HEARD INDICATING THE BREAKAWAY WASHER HAS BEEN CUT THROUGH. WHEN THE SURGEON FIRED THE DEVICE, THE SALES REP DID NOT HEAR A CRUNCH NOISE AND ALERTED THE SURGEON TO SQUEEZE HARDER UNTIL A CRUNCH IS HEARD BEFORE RELEASING THE FIRING HANDLE. HOWEVER, THE SURGEON SAID SHE HEARD THE CRUNCH NOISE AND FELT IT AND WAS UNDER THE IMPRESSION THAT THE FIRING SEQUENCE WAS COMPLETE. THE STAPLER WAS THEN OPENED AND REMOVED FROM THE PATIENT. WHEN THE DONUTS WERE INSPECTED, IT WAS NOTED THAT ONLY THE PROXIMAL DONUT WAS PRESENT, AND THE BREAKAWAY WASHER WAS STILL NOT CUT THROUGH. AN AIR LEAK TEST WAS PERFORMED, AND BUBBLES WERE PRESENT INDICATING A LEAK. SURGEON THEN PROCEEDED TO HAND SEW THE ANASTOMOSIS AS THERE WAS NOT ENOUGH MARGIN FOR THE SURGEON TO FIRE ANOTHER CIRCULAR STAPLER AS THE RESECTED TUMOR WAS VERY LOW. DRY FIRING OF THE STAPLER WAS PERFORMED OUTSIDE THE PATIENT AND THE STAPLER COULD BE FIRED OPTIMALLY, CUTTING THROUGH THE BREAKAWAY WASHER WITHOUT EXCESSIVE FORCE NEEDED. AFTER THE ANASTOMOSIS WAS SUCCESSFULLY HAND SEWN AND TESTED FOR AIR LEAKS, THE PATIENT WAS GIVEN A COLOSTOMY BAG. THE SURGERY WAS DELAYED BY 120 MINUTES.
14-12-2022 | Injury | ethicon endo-surgery | 09-01-2023 | ILS 29MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT A BRAND NEW CDH29A STAPLER WAS OPENED AND USED FOR THE CIRCULAR ANASTOMOSIS DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE. SALES REP WAS PRESENT DURING THE PROCEDURE TO GUIDE THE SURGEONS THROUGH THE PROCESS OF USING THE CIRCULAR STAPLER. DURING THE FIRING SEQUENCE, THE SURGEON FIRING THE DEVICE WAS INSTRUCTED TO SQUEEZE THE STAPLER FIRING HANDLE UNTIL IT TOUCHES THE BODY OF THE DEVICE (PLASTIC TO PLASTIC), AND TO ONLY RELEASE THE FIRING HANDLE AFTER A CRUNCH NOISE IS HEARD INDICATING THE BREAKAWAY WASHER HAS BEEN CUT THROUGH. WHEN THE SURGEON FIRED THE DEVICE, THE SALES REP DID NOT HEAR A CRUNCH NOISE AND ALERTED THE SURGEON TO SQUEEZE HARDER UNTIL A CRUNCH IS HEARD BEFORE RELEASING THE FIRING HANDLE. HOWEVER, THE SURGEON SAID SHE HEARD THE CRUNCH NOISE AND FELT IT AND WAS UNDER THE IMPRESSION THAT THE FIRING SEQUENCE WAS COMPLETE. THE STAPLER WAS THEN OPENED AND REMOVED FROM THE PATIENT. WHEN THE DONUTS WERE INSPECTED, IT WAS NOTED THAT ONLY THE PROXIMAL DONUT WAS PRESENT, AND THE BREAKAWAY WASHER WAS STILL NOT CUT THROUGH. AN AIR LEAK TEST WAS PERFORMED, AND BUBBLES WERE PRESENT INDICATING A LEAK. SURGEON THEN PROCEEDED TO HAND SEW THE ANASTOMOSIS AS THERE WAS NOT ENOUGH MARGIN FOR THE SURGEON TO FIRE ANOTHER CIRCULAR STAPLER AS THE RESECTED TUMOR WAS VERY LOW. DRY FIRING OF THE STAPLER WAS PERFORMED OUTSIDE THE PATIENT AND THE STAPLER COULD BE FIRED OPTIMALLY, CUTTING THROUGH THE BREAKAWAY WASHER WITHOUT EXCESSIVE FORCE NEEDED. AFTER THE ANASTOMOSIS WAS SUCCESSFULLY HAND SEWN AND TESTED FOR AIR LEAKS, THE PATIENT WAS GIVEN A COLOSTOMY BAG. THE SURGERY WAS DELAYED BY 120 MINUTES.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC-ASSISTED LOBE RESECTION, WHEN LOBE RESECTION WAS ATTEMPTED USING RELOAD, THE KNIFE DID NOT REACH THE ANCHORING SUTURE AT THE TIP, AND THE FIRING STOPPED JUST BEFORE THAT POINT, SO ALL OF THE TISSUE COULD NOT BE RESECTED. ANOTHER RELOAD WAS TAKEN OUT AND ADDITIONAL RESECTION WAS PERFORMED. THERE WAS NO PATIENT INJURY.
21-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE.
21-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE.
05-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION, DESPITE BEING CONNECTED TO THE HANDLE WITH THE YELLOW TAB STILL ATTACHED, THE RELOAD WAS PRE-FIRING. IT WAS NOTICED WHEN THE PATIENT WAS BEING APPROACHED, SO THE USE OF THE RELOAD WAS DISCONTINUED. A NEW CARTRIDGE WAS USED AND PERFORMED WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE RELOAD WAS PARTIALLY FIRED AND THE INTERLOCK WAS ENGAGED. THE SLED AND STAPLES WERE VISIBLE AT THE 3.5CM CUT LINE. STAPLE PUSHERS WERE VISIBLE AT THE 4CM CUT LINE.
05-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION, DESPITE BEING CONNECTED TO THE HANDLE WITH THE YELLOW TAB STILL ATTACHED, THE RELOAD WAS PRE-FIRING. IT WAS NOTICED WHEN THE PATIENT WAS BEING APPROACHED, SO THE USE OF THE RELOAD WAS DISCONTINUED. A NEW CARTRIDGE WAS USED AND PERFORMED WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE RELOAD WAS PARTIALLY FIRED AND THE INTERLOCK WAS ENGAGED. THE SLED AND STAPLES WERE VISIBLE AT THE 3.5CM CUT LINE. STAPLE PUSHERS WERE VISIBLE AT THE 4CM CUT LINE.
01-01-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE.
22-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | TRI-STAPLE 2.0 | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC SURGERY, WHILE RESECTIONING THE BLOOD VESSEL, THE RELOAD WAS SUCCESSFULLY FIRED, BUT THE KNIFE DID NOT RETURN AFTER FIRING AND AN ERROR WAS DISPLAYED ON THE DISPLAY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | TRI-STAPLE 2.0 | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC SURGERY, WHILE RESECTIONING THE BLOOD VESSEL, THE RELOAD WAS SUCCESSFULLY FIRED, BUT THE KNIFE DID NOT RETURN AFTER FIRING AND AN ERROR WAS DISPLAYED ON THE DISPLAY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN PROCEDURE, THE PRODUCT COULD NOT BE USED AND FIRING IN THE SURGERY.
17-11-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE PREPARING FOR THE OPERATION, WHEN THE NURSE LOADED THE STAPLER WITH THE LOADING UNIT, HOWEVER, THE JAWS OF THE RELOAD WAS UNABLE TO CLOSE. THIS HAPPENED ON 4 RELOADS. ANOTHER BRAND PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-11-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | ILS 33MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE AFTER THE DEVICE WAS FIRED IT WAS NOTED THE STAPLES MALFORMED AND THE DONUTS WERE INCOMPLETE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-11-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | ILS 33MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE AFTER THE DEVICE WAS FIRED IT WAS NOTED THE STAPLES MALFORMED AND THE DONUTS WERE INCOMPLETE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-11-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE PREPARING FOR THE OPERATION, WHEN THE NURSE LOADED THE STAPLER WITH THE LOADING UNIT, HOWEVER, THE JAWS OF THE RELOAD WAS UNABLE TO CLOSE. THIS HAPPENED ON 4 RELOADS. ANOTHER BRAND PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-11-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GENERAL SURGERY THE HANDPIECE DID NOT WORK AT ALL. THE NURSE THAT WAS PRESENT DURING THE INCIDENT SAID THAT THE DEVICE COULD NOT BE FIRED AT ALL. THE SURGERY WAS DELAYED FOR AN UNKNOWN NUMBER OF MINUTES. ANOTHER DEVICE WAS OPENED TO FINISH THE PROCEDURE SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
17-11-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE PREPARING FOR THE OPERATION, WHEN THE NURSE LOADED THE THE STAPLER WITH THE LOADING UNIT, HOWEVER, THE JAWS OF THE RELOAD WAS UNABLE TO CLOSE. THIS HAPPENED ON 4 RELOADS. ANOTHER BRAND PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-11-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GENERAL SURGERY THE HANDPIECE DID NOT WORK AT ALL. THE NURSE THAT WAS PRESENT DURING THE INCIDENT SAID THAT THE DEVICE COULD NOT BE FIRED AT ALL. THE SURGERY WAS DELAYED FOR AN UNKNOWN NUMBER OF MINUTES. ANOTHER DEVICE WAS OPENED TO FINISH THE PROCEDURE SUCCESSFULLY WITH NO PATIENT CONSEQUENCES.
17-11-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE PREPARING FOR THE OPERATION, WHEN THE NURSE LOADED THE THE STAPLER WITH THE LOADING UNIT, HOWEVER, THE JAWS OF THE RELOAD WAS UNABLE TO CLOSE. THIS HAPPENED ON 4 RELOADS. ANOTHER BRAND PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-11-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE PREPARING FOR THE OPERATION, WHEN THE NURSE LOADED THE STAPLER WITH THE LOADING UNIT, HOWEVER, THE JAWS OF THE RELOAD WAS UNABLE TO CLOSE. THIS HAPPENED ON 4 RELOADS. ANOTHER BRAND PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | CIRCULAR MECH STD 29MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY THE SURGEON CLAIMS THE DEVICE FIRED PROPERLY, BUT ONLY CLIPS APPLIED WITHOUT CUTTING THE TISSUE. HE HAD TO CREATE NEW ANASTOMOSIS, USING NEW DEVICE (CDH29B). THE SURGERY WAS PROLONGED BUT SUCCESSFULLY COMPLETED WITH NO PATIENT CONSEQUENCES.
16-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY, IN GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. THE DEVICE LOCKED ON THE TISSUE AS WELL BUT WAS REMOVED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAP ANTERIOR RESECTION PROCEDURE, IT WAS REPORTED THAT DURING THE FIRST FIRING, THE DEVICE WOULD NOT FIRE. A NEW DEVICE WAS OPENED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
15-12-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | PROXIMATE*LINEAR CUTTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT ONLY 4 STAPLES DEPLOYED, AND THE REST DID NOT. TO COMPLETE THE TRANS-SECTION, THE STAFF OPENED ANOTHER PROXIMATE LINEAR STAPLER WHICH WAS FUNCTIONING FINE. THERE WAS NO PATIENT INJURY, NO DELAY TO OR TIME, AND NO IMPACT TO THE TEAM/THEATRE ENVIRONMENT. THE SURGEON IS VERY FAMILIAR WITH OUR PRODUCT AND A REGULAR USER OF THIS PRODUCT (PER CLINICAL STAFF ADVISE). PROCEDURE: UNKNOWN.
16-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY, IN GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. THE DEVICE LOCKED ON THE TISSUE AS WELL BUT WAS REMOVED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, WHILE RESECTING PART OF THE LUNG, THE HANDLE DID NOT ALLOW TO FIRE THE RELOAD ON THE LUNG, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS NOT SQUEEZED. THE NEXT HANDLE WAS TAKEN, SECOND ONE FIRED ONCE PROPERLY AND DID NOT ALLOW TO FIRE THE NEXT TIME. A THIRD HANDLE WAS TAKEN AND FIRED PROPERLY ABOUT 3 TIMES THAT RESOLVE THE ISSUE AND COMPLETE THE CASE. IT WAS NOTED THAT WHENEVER THE RELOAD WAS TAKEN AND DID NOT FIRE WITH ONE HANDLE, IT FIRED PROPERLY WITH THE NEXT ONE, SO THERE WAS NO RELOAD ISSUE, THEY WORKED PROPERLY. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THE DEVICE WAS PROPERLY TRIGGERED, VERY EXPERIENCED USER. UPON REVIEWING THE ANASTOMOSIS, IT WAS FOUND THAT THE DEVICE FAILED TO STAPLE APPROXIMATELY 1-1.5 CM ON ONE SIDE. THE ANASTOMOSIS HAD TO BE REVISED AND PERFORMED WITH A SECOND CDH29P. THE PATIENT WAS NOT HARMED. UNFORTUNATELY, THE SURGERY TIME WAS DELAYED BY ABOUT 1 HOUR. RECTUM PROCEDURE.
20-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY PROCEDURE, WHILE TRYING TO TAKE THE VESSEL, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. WHILE FIRING THE SECOND RELOAD, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. BOTH ALSO LOCKED DOWN ON THE TISSUE AND WOULD NOT RELEASE. A GRASPER WAS USED TO PUSH THE JAWS OPEN ON BOTH INSTANCES. A NEW HANDLE AND NEW RELOADS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY PROCEDURE, WHILE TRYING TO TAKE THE VESSEL, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. WHILE FIRING THE SECOND RELOAD, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. BOTH ALSO LOCKED DOWN ON THE TISSUE AND WOULD NOT RELEASE. A GRASPER WAS USED TO PUSH THE JAWS OPEN ON BOTH INSTANCES. A NEW HANDLE AND NEW RELOADS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY PROCEDURE, WHILE TRYING TO TAKE THE VESSEL, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. WHILE FIRING THE SECOND RELOAD, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. BOTH ALSO LOCKED DOWN ON THE TISSUE AND WOULD NOT RELEASE. A GRASPER WAS USED TO PUSH THE JAWS OPEN ON BOTH INSTANCES. A NEW HANDLE AND NEW RELOADS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY PROCEDURE, WHILE TRYING TO TAKE THE VESSEL, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. WHILE FIRING THE SECOND RELOAD, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. BOTH ALSO LOCKED DOWN ON THE TISSUE AND WOULD NOT RELEASE. A GRASPER WAS USED TO PUSH THE JAWS OPEN ON BOTH INSTANCES. A NEW HANDLE AND NEW RELOADS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | ethicon endo-surgery | 09-01-2023 | UNKNOWN CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE WAS FIRED. THERE WAS NO ISSUE WITH THE DEVICE. THE ANASTOMOSIS WAS COMPLETED, HOWEVER, ON INSPECTION OF THE DONUTS, ONE SEEMED SLIGHTLY THIN IN ONE PART. THERE WAS ALSO TWO MALFORMED STAPLES. THE SURGEON STATED THESE WERE FROM THE TRANSVERSE FIRING WITH THE CONTOUR. THE SURGEON SUTURED THE ANASTOMOSIS TOO SO AS TO STABILIZE THE ANASTOMOSIS. THERE WERE NO PATIENT CONSEQUENCES.
20-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY PROCEDURE, WHILE TRYING TO TAKE THE VESSEL, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. WHILE FIRING THE SECOND RELOAD, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. BOTH ALSO LOCKED DOWN ON THE TISSUE AND WOULD NOT RELEASE. A GRASPER WAS USED TO PUSH THE JAWS OPEN ON BOTH INSTANCES. A NEW HANDLE AND NEW RELOADS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY PROCEDURE, WHILE TRYING TO TAKE THE VESSEL, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. WHILE FIRING THE SECOND RELOAD, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. BOTH ALSO LOCKED DOWN ON THE TISSUE AND WOULD NOT RELEASE. A GRASPER WAS USED TO PUSH THE JAWS OPEN ON BOTH INSTANCES. A NEW HANDLE AND NEW RELOADS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC NEPHRECTOMY PROCEDURE, WHILE TRYING TO TAKE THE VESSEL, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. WHILE FIRING THE SECOND RELOAD, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. BOTH ALSO LOCKED DOWN ON THE TISSUE AND WOULD NOT RELEASE. A GRASPER WAS USED TO PUSH THE JAWS OPEN ON BOTH INSTANCES. A NEW HANDLE AND NEW RELOADS WERE USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE ECHELON CIRCULAR DEVICE WAS FIRED. THERE WAS NO ISSUE WITH THE DEVICE. THE ANASTOMOSIS WAS COMPLETED, HOWEVER, ON INSPECTION OF THE DONUTS, ONE SEEMED SLIGHTLY THIN IN ONE PART. THERE WAS ALSO TWO MALFORMED STAPLES. THE SURGEON STATED THESE WERE FROM THE TRANSVERSE FIRING WITH THE CONTOUR. THE SURGEON SUTURED THE ANASTOMOSIS TOO SO AS TO STABILIZE THE ANASTOMOSIS. THERE WERE NO PATIENT CONSEQUENCES.
16-12-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE ECHELON CIRCULAR DEVICE WAS FIRED. THERE WAS NO ISSUE WITH THE DEVICE. THE ANASTOMOSIS WAS COMPLETED, HOWEVER, ON INSPECTION OF THE DONUTS, ONE SEEMED SLIGHTLY THIN IN ONE PART. THERE WAS ALSO TWO MALFORMED STAPLES. THE SURGEON STATED THESE WERE FROM THE TRANSVERSE FIRING WITH THE CONTOUR. THE SURGEON SUTURED THE ANASTOMOSIS TOO SO AS TO STABILIZE THE ANASTOMOSIS. THERE WERE NO PATIENT CONSEQUENCES.
12-10-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA PROCEDURE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-04-2023 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, AT THE END OF THE PROCEDURE, THE SURGEON FOUND A PIECE OF YELLOW PLASTIC FROM THE RELOAD INTO THE PATIENT CAVITY. THE SURGEON RETRIEVED THE YELLOW FRAGMENT BY JUST LOOKING THROUGH THE CAMERA. THERE WAS NO PATIENT INJURY.
05-04-2023 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, AT THE END OF THE PROCEDURE, THE SURGEON FOUND A PIECE OF YELLOW PLASTIC FROM THE RELOAD INTO THE PATIENT CAVITY. THE SURGEON RETRIEVED THE YELLOW FRAGMENT BY JUST LOOKING THROUGH THE CAMERA. THERE WAS NO PATIENT INJURY.
09-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLORECTAL SURGERY, THE DEVICE DID NOT FIRE. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
13-10-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | 75MM SELECTABLE NEW TLC | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CUTTER BROKE WHEN LOADED. PRODUCT FAILURE. ANOTHER DEVICE WAS USED INSTEAD.
20-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, WHILE RESECTING PART OF THE LUNG, THE HANDLE DID NOT ALLOW TO FIRE THE RELOAD ON THE LUNG, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS NOT SQUEEZED. THE NEXT HANDLE WAS TAKEN, SECOND ONE FIRED ONCE PROPERLY AND DID NOT ALLOW TO FIRE THE NEXT TIME. A THIRD HANDLE WAS TAKEN AND FIRED PROPERLY ABOUT 3 TIMES THAT RESOLVE THE ISSUE AND COMPLETE THE CASE. IT WAS NOTED THAT WHENEVER THE RELOAD WAS TAKEN AND DID NOT FIRE WITH ONE HANDLE, IT FIRED PROPERLY WITH THE NEXT ONE, SO THERE WAS NO RELOAD ISSUE, THEY WORKED PROPERLY. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | ENDO ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | GST BLUE RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2022 | Injury | ethicon endo-surgery | 09-01-2023 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING AN UNKNOW PROCEDURE THE STAPLE FAILED TO FIRE FOR SECOND STAPLER RELOAD. THE CASE WAS LAPAROSCOPIC AND DUE TO FAULTY STAPLE IT WAS CONVERTED TO OPEN.
22-12-2022 | Injury | ethicon endo-surgery | 09-01-2023 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING AN UNKNOW PROCEDURE THE STAPLE FAILED TO FIRE FOR SECOND STAPLER RELOAD. THE CASE WAS LAPAROSCOPIC AND DUE TO FAULTY STAPLE IT WAS CONVERTED TO OPEN.
11-11-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | 55MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE GASTRECTOMY PROCEDURE, THE GUN¿S FIRING KNOB TURNING UP WHILE THE SURGEON TRIED TO FIRE THE DEVICE. THERE WAS NO SURGICAL DELAY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN TRANSECTING JEJUNUM, THE I-BEAM DID NOT RETREAT, THEN A POWER HANDLE ERROR APPEARED. AT THE SAME TIME, MESSAGE APPEARED 'TO PRESS THE READY-TO-FIRE BUTTONS FOR 10 SECONDS'. THE USER OPENED THE JAWS AFTER PRESSING THE READY-TO-FIRE BUTTONS FOR 10 SECONDS AND ANOTHER HANDLE AND ADAPTER WERE USED TO CONTINUE THE CASE.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Retraction Problem | A0510 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN TRANSECTING JEJUNUM, THE I-BEAM DID NOT RETREAT, THEN A POWER HANDLE ERROR APPEARED. AT THE SAME TIME, MESSAGE APPEARED 'TO PRESS THE READY-TO-FIRE BUTTONS FOR 10 SECONDS'. THE USER OPENED THE JAWS AFTER PRESSING THE READY-TO-FIRE BUTTONS FOR 10 SECONDS AND ANOTHER HANDLE AND ADAPTER WERE USED TO CONTINUE THE CASE.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN TRANSECTING JEJUNUM, THE I-BEAM DID NOT RETREAT, THEN A POWER HANDLE ERROR APPEARED. AT THE SAME TIME, MESSAGE APPEARED 'TO PRESS THE READY-TO-FIRE BUTTONS FOR 10 SECONDS'. THE USER OPENED THE JAWS AFTER PRESSING THE READY-TO-FIRE BUTTONS FOR 10 SECONDS AND ANOTHER HANDLE AND ADAPTER WERE USED TO CONTINUE THE CASE.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | UNKNOWN ENDO GIA SULU | Retraction Problem | A0510 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN TRANSECTING JEJUNUM, THE I-BEAM DID NOT RETREAT, THEN A POWER HANDLE ERROR APPEARED. AT THE SAME TIME, MESSAGE APPEARED 'TO PRESS THE READY-TO-FIRE BUTTONS FOR 10 SECONDS'. THE USER OPENED THE JAWS AFTER PRESSING THE READY-TO-FIRE BUTTONS FOR 10 SECONDS AND ANOTHER HANDLE AND ADAPTER WERE USED TO CONTINUE THE CASE.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN TRANSECTING JEJUNUM, THE I-BEAM DID NOT RETREAT, THEN A POWER HANDLE ERROR APPEARED. AT THE SAME TIME, MESSAGE APPEARED 'TO PRESS THE READY-TO-FIRE BUTTONS FOR 10 SECONDS'. THE USER OPENED THE JAWS AFTER PRESSING THE READY-TO-FIRE BUTTONS FOR 10 SECONDS AND ANOTHER HANDLE AND ADAPTER WERE USED TO CONTINUE THE CASE.
21-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, IN INTESTINAL RESECTION, THE LOADING COULD BE DONE WITHOUT ANY PROBLEMS. THE OPENING AND CLOSING OF THE JAWS WERE CHECKED. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT WHEN THE HANDLE WAS TRIED TO BE SQUEEZED, THE HANDLE WAS LOCKED. THE DEVICE DID NOT FIRE. A DIFFERENT RELOAD WAS ATTACHED TO THE SAME HANDLE, AND IT COULD BE USED. THERE WAS NO PATIENT INJURY.
01-03-2022 | Injury | us surgical puerto rico | 09-01-2023 | UNKNOWN TA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY EVALUATED OUTCOMES IN PATIENTS WHO UNDERWENT RESTORATIVE PROCTECTOMY FOR ULCERATIVE COLITIS BETWEEN OCTOBER 2017 AND NOVEMBER 2020. A DOUBLE STAPLED ILEAL-ANAL ANASTOMOSIS WAS PERFORMED WITH A 30MM TA STAPLER USED FOR RECTAL DIVISION VIA LAPAROSCOPIC APPROACH. THERE WERE 207 PATIENTS (161 LAPAROSCOPIC, 46 OPEN). POSTOPERATIVE COMPLICATIONS INCLUDED: ANASTOMOTIC LEAK, PELVIC COLLECTION, ILEUS, AND BLEEDING. REOPERATIONS AND PROLONGED HOSPITALIZATION WERE REPORTED BUT REASONS WERE NOT SPECIFIED. NO ADDITIONAL INFORMATION WAS PROVIDED.
01-03-2022 | Injury | us surgical puerto rico | 09-01-2023 | UNKNOWN TA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY EVALUATED OUTCOMES IN PATIENTS WHO UNDERWENT RESTORATIVE PROCTECTOMY FOR ULCERATIVE COLITIS BETWEEN OCTOBER 2017 AND NOVEMBER 2020. A DOUBLE STAPLED ILEAL-ANAL ANASTOMOSIS WAS PERFORMED WITH A 30MM TA STAPLER USED FOR RECTAL DIVISION VIA LAPAROSCOPIC APPROACH. THERE WERE 207 PATIENTS (161 LAPAROSCOPIC, 46 OPEN). POSTOPERATIVE COMPLICATIONS INCLUDED: ANASTOMOTIC LEAK, PELVIC COLLECTION, ILEUS, AND BLEEDING. REOPERATIONS AND PROLONGED HOSPITALIZATION WERE REPORTED BUT REASONS WERE NOT SPECIFIED. NO ADDITIONAL INFORMATION WAS PROVIDED.
01-03-2022 | Injury | us surgical puerto rico | 09-01-2023 | UNKNOWN TA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY EVALUATED OUTCOMES IN PATIENTS WHO UNDERWENT RESTORATIVE PROCTECTOMY FOR ULCERATIVE COLITIS BETWEEN OCTOBER 2017 AND NOVEMBER 2020. A DOUBLE STAPLED ILEAL-ANAL ANASTOMOSIS WAS PERFORMED WITH A 30MM TA STAPLER USED FOR RECTAL DIVISION VIA LAPAROSCOPIC APPROACH. THERE WERE 207 PATIENTS (161 LAPAROSCOPIC, 46 OPEN). POSTOPERATIVE COMPLICATIONS INCLUDED: ANASTOMOTIC LEAK, PELVIC COLLECTION, ILEUS, AND BLEEDING. REOPERATIONS AND PROLONGED HOSPITALIZATION WERE REPORTED BUT REASONS WERE NOT SPECIFIED. NO ADDITIONAL INFORMATION WAS PROVIDED.
01-12-2022 | Malfunction | ethicon endo-surgery | 09-01-2023 | CIRCULAR MECH STD 29MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE WHEN THE SHOT WAS FIRED PROPERLY, THE TISSUE WAS CUT, BUT ONLY HALF THE CIRCUMFERENCE OF THE LARGE INTESTINE WAS FUSED.
19-12-2022 | Malfunction | us surgical puerto rico | 09-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE STARTING TO CUT TISSUE, THE DEVICE DID NOT FIRE. ANOTHER HANDLE WA S USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | SIGNIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN TRANSECTING JEJUNUM, THE I-BEAM DID NOT RETREAT, THEN A POWER HANDLE ERROR APPEARED. AT THE SAME TIME, MESSAGE APPEARED TO PRESS THE READY-TO-FIRE BUTTONS FOR 10 SECONDS. THE USER OPENED THE JAWS AFTER PRESSING THE READY-TO-FIRE BUTTONS FOR 10 SECONDS AND ANOTHER SYSTEM WAS USED TO CONTINUE THE CASE.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | SIGNIA | Retraction Problem | A0510 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN TRANSECTING JEJUNUM, THE I-BEAM DID NOT RETREAT, THEN A POWER HANDLE ERROR APPEARED. AT THE SAME TIME, MESSAGE APPEARED TO PRESS THE READY-TO-FIRE BUTTONS FOR 10 SECONDS. THE USER OPENED THE JAWS AFTER PRESSING THE READY-TO-FIRE BUTTONS FOR 10 SECONDS AND ANOTHER SYSTEM WAS USED TO CONTINUE THE CASE.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | SIGNIA | Device Displays Incorrect Message | A090201 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN TRANSECTING JEJUNUM, THE I-BEAM DID NOT RETREAT, THEN A POWER HANDLE ERROR APPEARED. AT THE SAME TIME, MESSAGE APPEARED TO PRESS THE READY-TO-FIRE BUTTONS FOR 10 SECONDS. THE USER OPENED THE JAWS AFTER PRESSING THE READY-TO-FIRE BUTTONS FOR 10 SECONDS AND ANOTHER SYSTEM WAS USED TO CONTINUE THE CASE.
14-12-2022 | Malfunction | covidien lp llc north haven | 09-01-2023 | SIGNIA | Mechanics Altered | A0507 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TOTAL LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN TRANSECTING JEJUNUM, THE I-BEAM DID NOT RETREAT, THEN A POWER HANDLE ERROR APPEARED. AT THE SAME TIME, MESSAGE APPEARED TO PRESS THE READY-TO-FIRE BUTTONS FOR 10 SECONDS. THE USER OPENED THE JAWS AFTER PRESSING THE READY-TO-FIRE BUTTONS FOR 10 SECONDS AND ANOTHER SYSTEM WAS USED TO CONTINUE THE CASE.
15-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY SIGMOID WITH END TO END ANASTOMOSIS PROCEDURE, THE TWO STAPLERS DID NOT FIRE. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC OPEN COLON RESECTION THEY FIRED THE DEVICE, AND NO STAPLES CAME OUT. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
29-12-2022 | Malfunction | covidien | 10-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ATTEMPTED USE OF COVIDIEN TRI-STAPLE GRAY CURVED TIP 30MM EXTRA THIN/VASCULAR LOAD ON APPENDIX. SURGEON COULD NOT "FIRE" LOAD. OPENED SECOND LOAD, SAME ISSUE.
21-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT SURGICAL SCRUB NURSE REPORTS PROBABLE INCOMPLETE STAPLING OF 01 LOAD USED IN A SURGICAL PROCEDURE. HE DOES NOT KNOW WHICH OF THE 03 LOADS USED, SO I AM SENDING THE TWO BATCHES OF THE USED LOADS. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON WAS USING TO DEVICE FOR ANASTOMOSIS IN A LAR PROCEDURE. THE SURGEON USUALLY WILL LEAVE A LONG PURSE STRING BEFORE FIRING THE DEVICE, AS INDICATION THAT THE BLADE CUT THROUGH THE COLON PERFECTLY. HOWEVER, IN THIS CASE, THE SUTURE DID NOT BREAK DURING ANASTOMOSIS WHERE THIS INDICATING THE BLADE UNABLE TO CUT THE SUTURE. FOR THE SURGEON, HE THOUGHT THE ANASTOMOSIS WAS INCOMPLETE AND ORDER FOR COLONOSCOPY TO INSPECT THE STAPLE LINE AT THE ANASTOMOSIS SITE. STAPLE LINE WAS FOUND INTACT, AND NO LEAK DETECTED DURING AIR LEAK TEST. THE PROCEDURE WAS DELAYED BY 10-MINUTES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO FRAGMENTS GENERATED. THERE WERE NO PATIENT CONSEQUENCES.
18-12-2022 | Injury | ethicon endo-surgery | 10-01-2023 | CIRCULAR MECH STD 29MM | Failure to Cut | A050702 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE CIRCULAR STAPLER MISFIRED. STAPLE LINE WAS NOT FORMED ALSO KNIFE WAS NOT FULLY EXTENDED TO CUT THE TISSUE. THE SURGERY WAS DELAYED 15 MINUTES. THERE WAS A CHANGE IN THE PROCEDURE, AN APR WAS DONE INSTEAD OF A LAR.
18-12-2022 | Injury | ethicon endo-surgery | 10-01-2023 | CIRCULAR MECH STD 29MM | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE CIRCULAR STAPLER MISFIRED. STAPLE LINE WAS NOT FORMED ALSO KNIFE WAS NOT FULLY EXTENDED TO CUT THE TISSUE. THE SURGERY WAS DELAYED 15 MINUTES. THERE WAS A CHANGE IN THE PROCEDURE, AN APR WAS DONE INSTEAD OF A LAR.
18-12-2022 | Injury | ethicon endo-surgery | 10-01-2023 | CIRCULAR MECH STD 29MM | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE CIRCULAR STAPLER MISFIRED. STAPLE LINE WAS NOT FORMED ALSO KNIFE WAS NOT FULLY EXTENDED TO CUT THE TISSUE. THE SURGERY WAS DELAYED 15 MINUTES. THERE WAS A CHANGE IN THE PROCEDURE, AN APR WAS DONE INSTEAD OF A LAR.
18-12-2022 | Injury | ethicon endo-surgery | 10-01-2023 | CIRCULAR MECH STD 29MM | Firing Problem | A0505 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE CIRCULAR STAPLER MISFIRED. STAPLE LINE WAS NOT FORMED ALSO KNIFE WAS NOT FULLY EXTENDED TO CUT THE TISSUE. THE SURGERY WAS DELAYED 15 MINUTES. THERE WAS A CHANGE IN THE PROCEDURE, AN APR WAS DONE INSTEAD OF A LAR.
19-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A WHIPPLE PROCEDURE THAT THE DEVICE LOCKED OUT. STAPLES WERE DEPLOYED BUT THE STAPLE LINE WAS NOT COMPLETE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
21-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PROBABLE INCOMPLETE STAPLING OF 01 LOAD USED IN A SURGICAL PROCEDURE. HE DOES NOT KNOW WHICH OF THE 03 LOADS USED, SO I AM SENDING THE TWO BATCHES OF THE USED LOADS. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
07-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ABDOMINAL RETOSSIGMOIDECTOMY IN ONCOLOGY THE STAPLER DID NOT STAPLE, REMOVED THE LOAD AGAIN, AND STILL DID NOT MAKE THE STAPLE LINE AFTER SHOOTING. DOCTOR PREFERRED TO CONTINUE WITHOUT ENERGY USE.
21-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WAS PROBABLE CAUSE INCOMPLETE STAPLING OF 01 LOAD USED IN A SURGICAL PROCEDURE. HE DOES NOT KNOW WHICH OF THE 03 LOADS USED, SO I AM SENDING THE TWO BATCHES OF THE USED LOADS. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
07-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY THE STAPLER DID NOT PERFORM STAPLING. THE LOAD WAS REMOVED AND RE-FITTED AND EVEN THEN THE STAPLING LINE WAS NOT MADE AFTER FIRING. USE OF ANOTHER SIMILAR PRODUCT.
28-11-2022 | Injury | covidien lp llc north haven | 10-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. HAND SUTURING OVER THE STAPLE LINE WAS DONE TO COMPLETE THE CASE.
01-01-2022 | Injury | ethicon endo-surgery | 10-01-2023 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS USED TO RESECT THE DISTAL COLON, AND THE DEVICE SEEMED TO WORK AS EXPECTED. HOWEVER LATE IN THE CASE AS THE END TO END ANASTOMOSIS WAS BEING MADE, THE CIRCULAR STAPLER WAS APPLIED INTRALUMINALLY TO THE DISTAL STUMP AND WITH MINIMAL PRESSURE THE STAPLE LINE CAME APART. THIS RESULTED IN THE NEED TO SUTURE THE DISTAL STUMP CLOSED. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y PROCEDURE, WHILE AT POUCH, THE SHAFT CAME OFF OF THE WHITE PART OF THE HANDLE, THE RELOAD WAS STILL ATTACHED TO THE HANDLE, RELOAD COULD NOT OPEN OR REMOVE AND GRASPERS FROM THE LAPAROSCOPIC SET WAS USED TO PRY OPEN. IT WAS NOTED THAT THE TISSUE WAS NORMAL TISSUE AS ROUTINELY ENCOUNTERED. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y PROCEDURE, WHILE AT POUCH, THE SHAFT CAME OFF OF THE WHITE PART OF THE HANDLE, THE RELOAD WAS STILL ATTACHED TO THE HANDLE, RELOAD COULD NOT OPEN OR REMOVE AND GRASPERS FROM THE LAPAROSCOPIC SET WAS USED TO PRY OPEN. IT WAS NOTED THAT THE TISSUE WAS NORMAL TISSUE AS ROUTINELY ENCOUNTERED. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-11-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS, THE SUTURE WAS FOUND TO STILL BE ATTACHED TO THE DEVICE; THE SURGEON NOTED THAT THE STAPLER FIRED NORMALLY. THE DEVICE WAS DIFFICULT TO REMOVE; THE USER NOTED THAT THE RETENTION SAFETY STITCH WAS NOT CUT WHEN THE STAPLER WAS DEPLOYED; IT WAS CUT WITH ENDOSCOPIC SCISSORS. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y PROCEDURE, WHILE AT POUCH, THE SHAFT CAME OFF OF THE WHITE PART OF THE HANDLE, THE RELOAD WAS STILL ATTACHED TO THE HANDLE, RELOAD COULD NOT OPEN OR REMOVE AND GRASPERS FROM THE LAPAROSCOPIC SET WAS USED TO PRY OPEN. IT WAS NOTED THAT THE TISSUE WAS NORMAL TISSUE AS ROUTINELY ENCOUNTERED. THE HANDLE AND RELOAD WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-06-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | GIA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLES DID NOT DEPLOY BUT CUT THE TISSUE.
28-11-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | EEA ORVIL | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS, THE SUTURE WAS FOUND TO STILL BE ATTACHED TO THE DEVICE; THE SURGEON NOTED THAT THE STAPLER FIRED NORMALLY. THE DEVICE WAS DIFFICULT TO REMOVE; THE USER NOTED THAT THE RETENTION SAFETY STITCH WAS NOT CUT WHEN THE STAPLER WAS DEPLOYED; IT WAS CUT WITH ENDOSCOPIC SCISSORS. THERE WAS NO PATIENT INJURY.
22-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 10-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACIC LUNG VOLUME REDUCTION SURGERY, WHILE DETACHING LUNG PARENCHYMA, THE DEVICE DID NOT FIRE, THE GREEN BUTTON WAS PRESSED AND UNABLE TO PULL THE HANDLE AFTER PRESSING THE GREEN FIRING INSURANCE. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 10-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACIC LUNG VOLUME REDUCTION SURGERY, WHILE DETACHING LUNG PARENCHYMA, THE DEVICE DID NOT FIRE, THE GREEN BUTTON WAS PRESSED AND UNABLE TO PULL THE HANDLE AFTER PRESSING THE GREEN FIRING INSURANCE. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-11-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE, WITH THE SECOND RELOAD, THE SURGEON WAS UNABLE TO FIRE THE DEVICE. A SECOND HANDLE WAS OPENED WHICH ALSO DID NOT FIRE. A THIRD STAPLER FROM THE SAME BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-11-2022 | Injury | covidien lp llc north haven | 10-01-2023 | UNKNOWN REINFORCED RELOAD | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2017 TO APRIL 2021 REGARDING A STUDY TO EVALUATE THE IMPACT OF THE USE OF A REINFORCED STAPLER WITH THAT OF A STANDARD STAPLER FOR CLOSURE OF THE PANCREATIC PARENCHYMA DURING DISTAL PANCREATECTOMY ON 199 PATIENTS. PATIENTS WERE RANDOMIZED IN 2 GROUPS FOR THE USE OF A STANDARD STAPLER OR A REINFORCED STAPLER TO CLOSE REMNANT PANCREATIC PARENCHYMA. THERE WERE 99 PATIENTS IN THE STANDARD STAPLER GROUP AND 100 PATIENTS IN THE REINFORCED STAPLER GROUP. POSTOPERATIVE COMPLICATIONS INCLUDED: INTRAOPERATIVE PANCREATIC STUMP BLEEDING AFTER STAPLING, PANCREATIC FISTULA IN 83 PATIENTS, AND POSTOPERATIVE HEMORRHAGE IN 3 PATIENTS IN THE REINFORCED STAPLER GROUP. INTRAOPERATIVE HEMOSTASIS WAS REQUIRED IN 16 PATIENTS AFTER STAPLING OF PANCREAS IN THE REINFORCED STAPLER GROUP. DIAGNOSTIC POSTOPERATIVE COMPUTED TOMOGRAPHY SCAN WAS PERFORMED IN SYMPTOMATIC PATIENTS. READMISSIONS DUE TO PANCREATIC FISTULA WERE REPORTED IN 8 PATIENTS IN THE REINFORCED STAPLER GROUP. READMISSIONS DUE TO HEMORRHAGE WAS ALSO REPORTED. A PROBLEM DURING THE STAPLING PROCEDURE WAS REPORTED IN 7 PATIENTS IN THE REINFORCED STAPLER GROUP, THOUGH PROBLEM WAS NOT SPECIFIED.
01-11-2022 | Injury | covidien lp llc north haven | 10-01-2023 | UNKNOWN REINFORCED RELOAD | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2017 TO APRIL 2021 REGARDING A STUDY TO EVALUATE THE IMPACT OF THE USE OF A REINFORCED STAPLER WITH THAT OF A STANDARD STAPLER FOR CLOSURE OF THE PANCREATIC PARENCHYMA DURING DISTAL PANCREATECTOMY ON 199 PATIENTS. PATIENTS WERE RANDOMIZED IN 2 GROUPS FOR THE USE OF A STANDARD STAPLER OR A REINFORCED STAPLER TO CLOSE REMNANT PANCREATIC PARENCHYMA. THERE WERE 99 PATIENTS IN THE STANDARD STAPLER GROUP AND 100 PATIENTS IN THE REINFORCED STAPLER GROUP. POSTOPERATIVE COMPLICATIONS INCLUDED: INTRAOPERATIVE PANCREATIC STUMP BLEEDING AFTER STAPLING, PANCREATIC FISTULA IN 83 PATIENTS, AND POSTOPERATIVE HEMORRHAGE IN 3 PATIENTS IN THE REINFORCED STAPLER GROUP. INTRAOPERATIVE HEMOSTASIS WAS REQUIRED IN 16 PATIENTS AFTER STAPLING OF PANCREAS IN THE REINFORCED STAPLER GROUP. DIAGNOSTIC POSTOPERATIVE COMPUTED TOMOGRAPHY SCAN WAS PERFORMED IN SYMPTOMATIC PATIENTS. READMISSIONS DUE TO PANCREATIC FISTULA WERE REPORTED IN 8 PATIENTS IN THE REINFORCED STAPLER GROUP. READMISSIONS DUE TO HEMORRHAGE WAS ALSO REPORTED. A PROBLEM DURING THE STAPLING PROCEDURE WAS REPORTED IN 7 PATIENTS IN THE REINFORCED STAPLER GROUP, THOUGH PROBLEM WAS NOT SPECIFIED.
01-11-2022 | Injury | covidien lp llc north haven | 10-01-2023 | UNKNOWN REINFORCED RELOAD | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2017 TO APRIL 2021 REGARDING A STUDY TO EVALUATE THE IMPACT OF THE USE OF A REINFORCED STAPLER WITH THAT OF A STANDARD STAPLER FOR CLOSURE OF THE PANCREATIC PARENCHYMA DURING DISTAL PANCREATECTOMY ON 199 PATIENTS. PATIENTS WERE RANDOMIZED IN 2 GROUPS FOR THE USE OF A STANDARD STAPLER OR A REINFORCED STAPLER TO CLOSE REMNANT PANCREATIC PARENCHYMA. THERE WERE 99 PATIENTS IN THE STANDARD STAPLER GROUP AND 100 PATIENTS IN THE REINFORCED STAPLER GROUP. POSTOPERATIVE COMPLICATIONS INCLUDED: INTRAOPERATIVE PANCREATIC STUMP BLEEDING AFTER STAPLING, PANCREATIC FISTULA IN 83 PATIENTS, AND POSTOPERATIVE HEMORRHAGE IN 3 PATIENTS IN THE REINFORCED STAPLER GROUP. INTRAOPERATIVE HEMOSTASIS WAS REQUIRED IN 16 PATIENTS AFTER STAPLING OF PANCREAS IN THE REINFORCED STAPLER GROUP. DIAGNOSTIC POSTOPERATIVE COMPUTED TOMOGRAPHY SCAN WAS PERFORMED IN SYMPTOMATIC PATIENTS. READMISSIONS DUE TO PANCREATIC FISTULA WERE REPORTED IN 8 PATIENTS IN THE REINFORCED STAPLER GROUP. READMISSIONS DUE TO HEMORRHAGE WAS ALSO REPORTED. A PROBLEM DURING THE STAPLING PROCEDURE WAS REPORTED IN 7 PATIENTS IN THE REINFORCED STAPLER GROUP, THOUGH PROBLEM WAS NOT SPECIFIED.
24-08-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE DIDN´T CLAMPED AND ALSO DIDN´T CUT. A NEW DEVICE WAS OPENED. THERE WERE NO CONSEQUENCES FOR THE PATIENT.
21-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOTOMY/ OPEN LEFT LOWER LOBECTOMY THE STAPLER MALFUNCTIONED. THE STAPLES NEVER RELEASED WHEN FIRED CREATING OPEN ARTERY INTO PLURAL SPACE. THIS REQUIRED MORE HOURS FOR REPAIR.
15-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLECTOMY SIGMOID WITH END TO END ANASTOMOSIS PROCEDURE, THE TWO STAPLERS DID NOT FIRE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON FELT THAT THERE'S SOMETHING WRONG WITH THE STAPLER HE CAN'T FIRE IT PROPERLY AS USUAL DESPITE FOLLOWING ALL THE STEPS (AND IT WAS VERY HEAVY DURING FIRING) THEN THE FIRING KNOB SEPARATED FROM THE STAPLER THERE WERE NO PATIENT CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | 75MM SELECTABLE NEW TLC | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON FELT THAT THERE'S SOMETHING WRONG WITH THE STAPLER HE CAN'T FIRE IT PROPERLY AS USUAL DESPITE FOLLOWING ALL THE STEPS (AND IT WAS VERY HEAVY DURING FIRING) THEN THE FIRING KNOB SEPARATED FROM THE STAPLER THERE WERE NO PATIENT CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | 75MM SELECTABLE NEW TLC | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON FELT THAT THERE'S SOMETHING WRONG WITH THE STAPLER HE CAN'T FIRE IT PROPERLY AS USUAL DESPITE FOLLOWING ALL THE STEPS (AND IT WAS VERY HEAVY DURING FIRING) THEN THE FIRING KNOB SEPARATED FROM THE STAPLER THERE WERE NO PATIENT CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | 75MM SELECTABLE NEW TLC | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON FELT THAT THERE'S SOMETHING WRONG WITH THE STAPLER HE CAN'T FIRE IT PROPERLY AS USUAL DESPITE FOLLOWING ALL THE STEPS (AND IT WAS VERY HEAVY DURING FIRING) THEN THE FIRING KNOB SEPARATED FROM THE STAPLER THERE WERE NO PATIENT CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE ECHELON LOCKS AND COULD NOT BE OPENED. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
23-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVERSEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVERSEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVERSEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVERSEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
25-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC COLORECTAL SURGERY, WHEN CUTTING THE RECTUM, THE STAPLER STOPPED FIRING HALFWAY AND THERE WERE STAPLES DID NOT FORM INTO B SHAPE. TO FIX THE STAPLE LINE, THE HANDLE AND RELOAD WERE REPLACED AND RE-CUTTING WAS PERFORMED.
25-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC COLORECTAL SURGERY, WHEN CUTTING THE RECTUM, THE STAPLER STOPPED FIRING HALFWAY AND THERE WERE STAPLES DID NOT FORM INTO B SHAPE. TO FIX THE STAPLE LINE, THE HANDLE AND RELOAD WERE REPLACED AND RE-CUTTING WAS PERFORMED.
23-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVERSEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVERSEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVERSEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVERSEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
12-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC BYPASS, WHEN DIVIDING THE JEJUNUM, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-11-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE, WITH THE SECOND RELOAD, THE SURGEON WAS UNABLE TO FIRE THE DEVICE. A SECOND HANDLE WAS OPENED WHICH ALSO DID NOT FIRE. A THIRD STAPLER FROM THE SAME BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC HEAD AND DUODENAL RESECTION, WHILE RESECTIONING THE STOMACH, THE KNIFE CUT THROUGH, BUT THERE WAS A PART OF THE STAPLER THAT WAS NOT APPLIED. THIS WAS OBSERVED ON TWO RELOADS. REINFORCEMENT SUTURES WERE USED TO RESOLVE THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC HEAD AND DUODENAL RESECTION, WHILE RESECTIONING THE STOMACH, THE KNIFE CUT THROUGH, BUT THERE WAS A PART OF THE STAPLER THAT WAS NOT APPLIED. THIS WAS OBSERVED ON TWO RELOADS. REINFORCEMENT SUTURES WERE USED TO RESOLVE THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC HEAD AND DUODENAL RESECTION, WHILE RESECTIONING THE STOMACH, THE KNIFE CUT THROUGH, BUT THERE WAS A PART OF THE STAPLER THAT WAS NOT APPLIED. THIS WAS OBSERVED ON TWO RELOADS. REINFORCEMENT SUTURES WERE USED TO RESOLVE THE ISSUE.
12-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | ENDO GIA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC HEAD AND DUODENAL RESECTION, WHILE RESECTIONING THE STOMACH, THE KNIFE CUT THROUGH, BUT THERE WAS A PART OF THE STAPLER THAT WAS NOT APPLIED. THIS WAS OBSERVED ON TWO RELOADS. REINFORCEMENT SUTURES WERE USED TO RESOLVE THE ISSUE.
21-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Premature Activation | A150103 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE TRANSECTING THE STOMACH TISSUE, THE BUTTRESS STRIPS CAME OFF THE ANVIL PRIOR TO FIRING. THE STAPLE LINE WAS VERY SUCCESSFUL, DESPITE ONE SIDE OF THE STAPLE LINE MISSING THE BUTTRESS. THE SURGEON THEN NOTICED THE STAND ALONE BUTTRESS STRIP WHILE CLOSING THE JAWS OF THE STAPLER IN ORDER TO REMOVE IT FROM THE PATIENT. A GRASPER WAS USED TO REMOVE THE STAND ALONE BUTTRESS STRIP; AND ANOTHER RELOAD WAS USED TO TRANSECT THE STOMACH A LITTLE CLOSER TO THE BOUGIE. THE COMPLETE STAPLE LINE WAS THEN TESTED OUT SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
21-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Premature Activation | A150103 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE TRANSECTING THE STOMACH TISSUE, THE BUTTRESS STRIPS CAME OFF THE ANVIL PRIOR TO FIRING. THE STAPLE LINE WAS VERY SUCCESSFUL, DESPITE ONE SIDE OF THE STAPLE LINE MISSING THE BUTTRESS. THE SURGEON THEN NOTICED THE STAND ALONE BUTTRESS STRIP WHILE CLOSING THE JAWS OF THE STAPLER IN ORDER TO REMOVE IT FROM THE PATIENT. A GRASPER WAS USED TO REMOVE THE STAND ALONE BUTTRESS STRIP; AND ANOTHER RELOAD WAS USED TO TRANSECT THE STOMACH A LITTLE CLOSER TO THE BOUGIE. THE COMPLETE STAPLE LINE WAS THEN TESTED OUT SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
21-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE TRANSECTING THE STOMACH TISSUE, THE BUTTRESS STRIPS CAME OFF THE ANVIL PRIOR TO FIRING. THE STAPLE LINE WAS VERY SUCCESSFUL, DESPITE ONE SIDE OF THE STAPLE LINE MISSING THE BUTTRESS. THE SURGEON THEN NOTICED THE STAND ALONE BUTTRESS STRIP WHILE CLOSING THE JAWS OF THE STAPLER IN ORDER TO REMOVE IT FROM THE PATIENT. A GRASPER WAS USED TO REMOVE THE STAND ALONE BUTTRESS STRIP; AND ANOTHER RELOAD WAS USED TO TRANSECT THE STOMACH A LITTLE CLOSER TO THE BOUGIE. THE COMPLETE STAPLE LINE WAS THEN TESTED OUT SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
21-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE TRANSECTING THE STOMACH TISSUE, THE BUTTRESS STRIPS CAME OFF THE ANVIL PRIOR TO FIRING. THE STAPLE LINE WAS VERY SUCCESSFUL, DESPITE ONE SIDE OF THE STAPLE LINE MISSING THE BUTTRESS. THE SURGEON THEN NOTICED THE STAND ALONE BUTTRESS STRIP WHILE CLOSING THE JAWS OF THE STAPLER IN ORDER TO REMOVE IT FROM THE PATIENT. A GRASPER WAS USED TO REMOVE THE STAND ALONE BUTTRESS STRIP; AND ANOTHER RELOAD WAS USED TO TRANSECT THE STOMACH A LITTLE CLOSER TO THE BOUGIE. THE COMPLETE STAPLE LINE WAS THEN TESTED OUT SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
21-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE TRANSECTING THE STOMACH TISSUE, THE BUTTRESS STRIPS CAME OFF THE ANVIL PRIOR TO FIRING. THE STAPLE LINE WAS VERY SUCCESSFUL, DESPITE ONE SIDE OF THE STAPLE LINE MISSING THE BUTTRESS. THE SURGEON THEN NOTICED THE STAND ALONE BUTTRESS STRIP WHILE CLOSING THE JAWS OF THE STAPLER IN ORDER TO REMOVE IT FROM THE PATIENT. A GRASPER WAS USED TO REMOVE THE STAND ALONE BUTTRESS STRIP; AND ANOTHER RELOAD WAS USED TO TRANSECT THE STOMACH A LITTLE CLOSER TO THE BOUGIE. THE COMPLETE STAPLE LINE WAS THEN TESTED OUT SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
21-12-2022 | Injury | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHILE TRANSECTING THE STOMACH TISSUE, THE BUTTRESS STRIPS CAME OFF THE ANVIL PRIOR TO FIRING. THE STAPLE LINE WAS VERY SUCCESSFUL, DESPITE ONE SIDE OF THE STAPLE LINE MISSING THE BUTTRESS. THE SURGEON THEN NOTICED THE STAND ALONE BUTTRESS STRIP WHILE CLOSING THE JAWS OF THE STAPLER IN ORDER TO REMOVE IT FROM THE PATIENT. A GRASPER WAS USED TO REMOVE THE STAND ALONE BUTTRESS STRIP; AND ANOTHER RELOAD WAS USED TO TRANSECT THE STOMACH A LITTLE CLOSER TO THE BOUGIE. THE COMPLETE STAPLE LINE WAS THEN TESTED OUT SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
22-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | TA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION FOR CLOSING STUMP OF RECTUM, AFTER FIRING THE STAPLER GOT OPEN AND THERE WERE NO ANY STAPLES. A RE-OPERATION WAS NECESSARY AND THE SURGICAL TIME WAS EXTENDED DUE TO THE DEVICE PROBLEM. THE SURGEON MADE A HARTMAN POUCH AND COLOSTOMY FOR THE PATIENT.
22-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION FOR CLOSING STUMP OF RECTUM, AFTER FIRING THE STAPLER GOT OPEN AND THERE WERE NO ANY STAPLES. A RE-OPERATION WAS NECESSARY AND THE SURGICAL TIME WAS EXTENDED DUE TO THE DEVICE PROBLEM. THE SURGEON MADE A HARTMAN POUCH AND COLOSTOMY FOR THE PATIENT.
22-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | UNKNOWN TA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION FOR CLOSING STUMP OF RECTUM, AFTER FIRING THE STAPLER GOT OPEN AND THERE WERE NO ANY STAPLES. A RE-OPERATION WAS NECESSARY AND THE SURGICAL TIME WAS EXTENDED DUE TO THE DEVICE PROBLEM. THE SURGEON MADE A HARTMAN POUCH AND COLOSTOMY FOR THE PATIENT.
22-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | UNKNOWN TA | Mechanics Altered | A0507 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION FOR CLOSING STUMP OF RECTUM, AFTER FIRING THE STAPLER GOT OPEN AND THERE WERE NO ANY STAPLES. A RE-OPERATION WAS NECESSARY AND THE SURGICAL TIME WAS EXTENDED DUE TO THE DEVICE PROBLEM. THE SURGEON MADE A HARTMAN POUCH AND COLOSTOMY FOR THE PATIENT.
22-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | UNKNOWN TA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION FOR CLOSING STUMP OF RECTUM, AFTER FIRING THE STAPLER GOT OPEN AND THERE WERE NO ANY STAPLES. A RE-OPERATION WAS NECESSARY AND THE SURGICAL TIME WAS EXTENDED DUE TO THE DEVICE PROBLEM. THE SURGEON MADE A HARTMAN POUCH AND COLOSTOMY FOR THE PATIENT.
22-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | UNKNOWN TA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION FOR CLOSING STUMP OF RECTUM, AFTER FIRING THE STAPLER GOT OPEN AND THERE WERE NO ANY STAPLES. A RE-OPERATION WAS NECESSARY AND THE SURGICAL TIME WAS EXTENDED DUE TO THE DEVICE PROBLEM. THE SURGEON MADE A HARTMAN POUCH AND COLOSTOMY FOR THE PATIENT.
22-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | UNKNOWN TA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION FOR CLOSING STUMP OF RECTUM, AFTER FIRING THE STAPLER GOT OPEN AND THERE WERE NO ANY STAPLES. A RE-OPERATION WAS NECESSARY AND THE SURGICAL TIME WAS EXTENDED DUE TO THE DEVICE PROBLEM. THE SURGEON MADE A HARTMAN POUCH AND COLOSTOMY FOR THE PATIENT.
22-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | UNKNOWN TA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION FOR CLOSING STUMP OF RECTUM, AFTER FIRING THE STAPLER GOT OPEN AND THERE WERE NO ANY STAPLES. A RE-OPERATION WAS NECESSARY AND THE SURGICAL TIME WAS EXTENDED DUE TO THE DEVICE PROBLEM. THE SURGEON MADE A HARTMAN POUCH AND COLOSTOMY FOR THE PATIENT.
21-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Separation Failure | A150301 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IN THE FIRST FIRING OF THE STAPLE ON THE STOMACH, THE STAPLER FULLY FIRED BUT THE DISTAL STAPLES OF THE STAPLE LINE WERE MALFORMED. THE KNIFE BLADE ALSO DID NOT CUT THROUGH THE SMALL DISTAL PORTION OF THE REINFORCEMENT MATERIAL. A NEW RELOAD WAS PLACED ON THE HANDLE, AND WAS FIRED OVER THE AREA WHERE THE STAPLES DID NOT FORM. NO LEAKS SHOWN DURING LEAK TEST OF STAPLE LINE AT THE END OF THE CASE.
21-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IN THE FIRST FIRING OF THE STAPLE ON THE STOMACH, THE STAPLER FULLY FIRED BUT THE DISTAL STAPLES OF THE STAPLE LINE WERE MALFORMED. THE KNIFE BLADE ALSO DID NOT CUT THROUGH THE SMALL DISTAL PORTION OF THE REINFORCEMENT MATERIAL. A NEW RELOAD WAS PLACED ON THE HANDLE, AND WAS FIRED OVER THE AREA WHERE THE STAPLES DID NOT FORM. NO LEAKS SHOWN DURING LEAK TEST OF STAPLE LINE AT THE END OF THE CASE.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ANTRECTOMY THEY FIRED THREE TIMES BUT WERE NOT GETTING THROUGH TISSUE AND THE STAPLES WERE NOT FORMED AT ALL. A NEW DEVICE WAS OPENED. THERE WERE NO PATIENT CONSEQUENCES.
08-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPY PULMONARY RESECTION, THE SURGEON WAS BALE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZE ON THE HALF WAY, THE DEVICE WAS NOT ABLE TO BE FIRED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
16-11-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SLEEVE, WITH THE SECOND RELOAD, THE SURGEON WAS UNABLE TO FIRE THE DEVICE. A SECOND HANDLE WAS OPENED WHICH ALSO DID NOT FIRE. A THIRD STAPLER FROM THE SAME BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPY PULMONARY RESECTION, THE SURGEON WAS BALE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZE ON THE HALF WAY, THE DEVICE WAS NOT ABLE TO BE FIRED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
16-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD ECH FLEX 60MM 34CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD ECH FLEX 60MM 34CM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD ECH FLEX 60MM 34CM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD ECH FLEX 60MM 34CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD ECH FLEX 60MM 34CM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD ECH FLEX 60MM 34CM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE ON THE UPPER RIGHT LOBECTOMY, THE STAPLER DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. IT WAS ALSO NOTED THAT THE JAWS OF THE DEVICE WERE LOCKED IN TISSUE AND WAS OPENED TO REMOVED IT FROM THE TISSUE. IT WAS STIFF WHILE BLACK RETURN KNOB WAS BEING PULLED BACK. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE ON THE UPPER RIGHT LOBECTOMY, THE STAPLER DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND UNABLE TO SQUEEZE THE HANDLE. IT WAS ALSO NOTED THAT THE JAWS OF THE DEVICE WERE LOCKED IN TISSUE AND WAS OPENED TO REMOVED IT FROM THE TISSUE. IT WAS STIFF WHILE BLACK RETURN KNOB WAS BEING PULLED BACK. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-11-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC PSEUDOCYST RESECTION SURGERY, IT WAS FOUND THAT THE DEVICE COULD NOT BE COMPLETELY CLOSED. THE DOCTOR IMMEDIATELY USED MANUAL SUTURE TO CLOSE THE SMALL INTESTINE, AND THE SURGERY WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED FOR SOME TIME, AND THE EXACT TIME WAS UNKNOWN. ANASTOMOSIS BY HAND SUTURING WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-11-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC PSEUDOCYST RESECTION SURGERY, IT WAS FOUND THAT THE DEVICE COULD NOT BE COMPLETELY CLOSED. THE DOCTOR IMMEDIATELY USED MANUAL SUTURE TO CLOSE THE SMALL INTESTINE, AND THE SURGERY WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED FOR SOME TIME, AND THE EXACT TIME WAS UNKNOWN. ANASTOMOSIS BY HAND SUTURING WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-11-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC PSEUDOCYST RESECTION SURGERY, IT WAS FOUND THAT THE DEVICE COULD NOT BE COMPLETELY CLOSED. THE DOCTOR IMMEDIATELY USED MANUAL SUTURE TO CLOSE THE SMALL INTESTINE, AND THE SURGERY WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED FOR SOME TIME, AND THE EXACT TIME WAS UNKNOWN. ANASTOMOSIS BY HAND SUTURING WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-11-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC PSEUDOCYST RESECTION SURGERY, IT WAS FOUND THAT THE DEVICE COULD NOT BE COMPLETELY CLOSED. THE DOCTOR IMMEDIATELY USED MANUAL SUTURE TO CLOSE THE SMALL INTESTINE, AND THE SURGERY WAS COMPLETED. THE SURGICAL TIME WAS EXTENDED FOR SOME TIME, AND THE EXACT TIME WAS UNKNOWN. ANASTOMOSIS BY HAND SUTURING WAS DONE TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE ATTACHED YELLOW TAB WAS HARD, SO IT BROKE WHEN IT WAS PULLED FORCIBLY. ANOTHER RELOAD WAS TAKEN OUT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
12-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE ATTACHED YELLOW TAB WAS HARD, SO IT BROKE WHEN IT WAS PULLED FORCIBLY. ANOTHER RELOAD WAS TAKEN OUT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
08-12-2022 | Injury | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Cut | A050702 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE THE SURGEON PERFORMED AN OPEN COLON RESECTION FOR DIVERTICULITIS USING A CDH29P FOR THE ANASTOMOSIS. UPON INSPECTION OF THE TISSUE DONUTS IN THE CIRCULAR STAPLER, IT WAS FOUND THAT THE DISTAL DONUT WAS THIN ON ONE SIDE. A RIGOROUS LEAK TEST WAS PERFORMED USING AIR AND NO INDICATIONS OF LEAKS WERE FOUND. ON (B)(6), THE PATIENT EXPERIENCED SEVERE ABDOMINAL PAIN AND SLIGHTLY ELEVATED WBC. CONTRAST WAS INJECTED IN THE PATIENT AND A LEAK WAS IDENTIFIED. THE SURGEON RE-OPERATED ON THE PATIENT AND FOUND A HOLE IN THE ANASTOMOSIS ON THE PATIENT'S RIGHT SIDE. A DIVERTING LOOP ILEOSTOMY WAS CREATED AND THE PATIENT IS RECOVERING. THE PROCEDURE WAS NOT DELAYED AND WAS COMPLETED SUCCESSFULLY. THERE WERE PATIENT CONSEQUENCES.
08-12-2022 | Injury | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Cut | A050702 | Abdominal Pain | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE THE SURGEON PERFORMED AN OPEN COLON RESECTION FOR DIVERTICULITIS USING A CDH29P FOR THE ANASTOMOSIS. UPON INSPECTION OF THE TISSUE DONUTS IN THE CIRCULAR STAPLER, IT WAS FOUND THAT THE DISTAL DONUT WAS THIN ON ONE SIDE. A RIGOROUS LEAK TEST WAS PERFORMED USING AIR AND NO INDICATIONS OF LEAKS WERE FOUND. ON (B)(6), THE PATIENT EXPERIENCED SEVERE ABDOMINAL PAIN AND SLIGHTLY ELEVATED WBC. CONTRAST WAS INJECTED IN THE PATIENT AND A LEAK WAS IDENTIFIED. THE SURGEON RE-OPERATED ON THE PATIENT AND FOUND A HOLE IN THE ANASTOMOSIS ON THE PATIENT'S RIGHT SIDE. A DIVERTING LOOP ILEOSTOMY WAS CREATED AND THE PATIENT IS RECOVERING. THE PROCEDURE WAS NOT DELAYED AND WAS COMPLETED SUCCESSFULLY. THERE WERE PATIENT CONSEQUENCES.
08-12-2022 | Injury | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE THE SURGEON PERFORMED AN OPEN COLON RESECTION FOR DIVERTICULITIS USING A CDH29P FOR THE ANASTOMOSIS. UPON INSPECTION OF THE TISSUE DONUTS IN THE CIRCULAR STAPLER, IT WAS FOUND THAT THE DISTAL DONUT WAS THIN ON ONE SIDE. A RIGOROUS LEAK TEST WAS PERFORMED USING AIR AND NO INDICATIONS OF LEAKS WERE FOUND. ON (B)(6), THE PATIENT EXPERIENCED SEVERE ABDOMINAL PAIN AND SLIGHTLY ELEVATED WBC. CONTRAST WAS INJECTED IN THE PATIENT AND A LEAK WAS IDENTIFIED. THE SURGEON RE-OPERATED ON THE PATIENT AND FOUND A HOLE IN THE ANASTOMOSIS ON THE PATIENT'S RIGHT SIDE. A DIVERTING LOOP ILEOSTOMY WAS CREATED AND THE PATIENT IS RECOVERING. THE PROCEDURE WAS NOT DELAYED AND WAS COMPLETED SUCCESSFULLY. THERE WERE PATIENT CONSEQUENCES.
08-12-2022 | Injury | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE THE SURGEON PERFORMED AN OPEN COLON RESECTION FOR DIVERTICULITIS USING A CDH29P FOR THE ANASTOMOSIS. UPON INSPECTION OF THE TISSUE DONUTS IN THE CIRCULAR STAPLER, IT WAS FOUND THAT THE DISTAL DONUT WAS THIN ON ONE SIDE. A RIGOROUS LEAK TEST WAS PERFORMED USING AIR AND NO INDICATIONS OF LEAKS WERE FOUND. ON (B)(6), THE PATIENT EXPERIENCED SEVERE ABDOMINAL PAIN AND SLIGHTLY ELEVATED WBC. CONTRAST WAS INJECTED IN THE PATIENT AND A LEAK WAS IDENTIFIED. THE SURGEON RE-OPERATED ON THE PATIENT AND FOUND A HOLE IN THE ANASTOMOSIS ON THE PATIENT'S RIGHT SIDE. A DIVERTING LOOP ILEOSTOMY WAS CREATED AND THE PATIENT IS RECOVERING. THE PROCEDURE WAS NOT DELAYED AND WAS COMPLETED SUCCESSFULLY. THERE WERE PATIENT CONSEQUENCES.
01-11-2022 | Injury | covidien lp llc north haven | 10-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2017 TO APRIL 2021 REGARDING A STUDY TO EVALUATE THE IMPACT OF THE USE OF A REINFORCED STAPLER WITH THAT OF A STANDARD STAPLER FOR CLOSURE OF THE PANCREATIC PARENCHYMA DURING DISTAL PANCREATECTOMY ON 199 PATIENTS. PATIENTS WERE RANDOMIZED IN 2 GROUPS FOR THE USE OF A STANDARD STAPLER OR A REINFORCED STAPLER TO CLOSE REMNANT PANCREATIC PARENCHYMA. THERE WERE 99 PATIENTS IN THE STANDARD STAPLER GROUP AND 100 PATIENTS IN THE REINFORCED STAPLER GROUP. POSTOPERATIVE COMPLICATIONS INCLUDED: INTRAOPERATIVE PANCREATIC STUMP BLEEDING AFTER STAPLING, PANCREATIC FISTULA IN 67 PATIENTS, AND POSTOPERATIVE HEMORRHAGE IN 5 PATIENTS IN THE STANDARD STAPLER GROUP. INTRAOPERATIVE HEMOSTASIS WAS REQUIRED IN 24 PATIENTS AFTER STAPLING OF PANCREAS. DIAGNOSTIC POSTOPERATIVE COMPUTED TOMOGRAPHY SCAN WAS PERFORMED IN SYMPTOMATIC PATIENTS. READMISSIONS DUE TO PANCREATIC FISTULA AND HEMORRHAGE WERE REPORTED IN 8 PATIENTS IN THE STANDARD STAPLER GROUP. READMISSIONS DUE TO HEMORRHAGE WAS ALSO REPORTED. A PROBLEM DURING THE STAPLING PROCEDURE WAS REPORTED IN 5 PATIENTS IN THE STANDARD STAPLER GROUP, THOUGH PROBLEM WAS NOT SPECIFIED.
01-11-2022 | Injury | covidien lp llc north haven | 10-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2017 TO APRIL 2021 REGARDING A STUDY TO EVALUATE THE IMPACT OF THE USE OF A REINFORCED STAPLER WITH THAT OF A STANDARD STAPLER FOR CLOSURE OF THE PANCREATIC PARENCHYMA DURING DISTAL PANCREATECTOMY ON 199 PATIENTS. PATIENTS WERE RANDOMIZED IN 2 GROUPS FOR THE USE OF A STANDARD STAPLER OR A REINFORCED STAPLER TO CLOSE REMNANT PANCREATIC PARENCHYMA. THERE WERE 99 PATIENTS IN THE STANDARD STAPLER GROUP AND 100 PATIENTS IN THE REINFORCED STAPLER GROUP. POSTOPERATIVE COMPLICATIONS INCLUDED: INTRAOPERATIVE PANCREATIC STUMP BLEEDING AFTER STAPLING, PANCREATIC FISTULA IN 67 PATIENTS, AND POSTOPERATIVE HEMORRHAGE IN 5 PATIENTS IN THE STANDARD STAPLER GROUP. INTRAOPERATIVE HEMOSTASIS WAS REQUIRED IN 24 PATIENTS AFTER STAPLING OF PANCREAS. DIAGNOSTIC POSTOPERATIVE COMPUTED TOMOGRAPHY SCAN WAS PERFORMED IN SYMPTOMATIC PATIENTS. READMISSIONS DUE TO PANCREATIC FISTULA AND HEMORRHAGE WERE REPORTED IN 8 PATIENTS IN THE STANDARD STAPLER GROUP. READMISSIONS DUE TO HEMORRHAGE WAS ALSO REPORTED. A PROBLEM DURING THE STAPLING PROCEDURE WAS REPORTED IN 5 PATIENTS IN THE STANDARD STAPLER GROUP, THOUGH PROBLEM WAS NOT SPECIFIED.
01-11-2022 | Injury | covidien lp llc north haven | 10-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2017 TO APRIL 2021 REGARDING A STUDY TO EVALUATE THE IMPACT OF THE USE OF A REINFORCED STAPLER WITH THAT OF A STANDARD STAPLER FOR CLOSURE OF THE PANCREATIC PARENCHYMA DURING DISTAL PANCREATECTOMY ON 199 PATIENTS. PATIENTS WERE RANDOMIZED IN 2 GROUPS FOR THE USE OF A STANDARD STAPLER OR A REINFORCED STAPLER TO CLOSE REMNANT PANCREATIC PARENCHYMA. THERE WERE 99 PATIENTS IN THE STANDARD STAPLER GROUP AND 100 PATIENTS IN THE REINFORCED STAPLER GROUP. POSTOPERATIVE COMPLICATIONS INCLUDED: INTRAOPERATIVE PANCREATIC STUMP BLEEDING AFTER STAPLING, PANCREATIC FISTULA IN 67 PATIENTS, AND POSTOPERATIVE HEMORRHAGE IN 5 PATIENTS IN THE STANDARD STAPLER GROUP. INTRAOPERATIVE HEMOSTASIS WAS REQUIRED IN 24 PATIENTS AFTER STAPLING OF PANCREAS. DIAGNOSTIC POSTOPERATIVE COMPUTED TOMOGRAPHY SCAN WAS PERFORMED IN SYMPTOMATIC PATIENTS. READMISSIONS DUE TO PANCREATIC FISTULA AND HEMORRHAGE WERE REPORTED IN 8 PATIENTS IN THE STANDARD STAPLER GROUP. READMISSIONS DUE TO HEMORRHAGE WAS ALSO REPORTED. A PROBLEM DURING THE STAPLING PROCEDURE WAS REPORTED IN 5 PATIENTS IN THE STANDARD STAPLER GROUP, THOUGH PROBLEM WAS NOT SPECIFIED.
26-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, WHEN USING THE SECOND RELOAD, THE RELOAD WAS INSTALLED ON THE HANDLE BUT THE JAWS WILL NOT CLOSE COMPLETELY. WHEN A NEW HANDLE WAS TAKEN OUT, THE SAME RELOAD COULD BE USED NORMALLY. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE TWO OF THE STAPLES DID NOT FORM CORRECTLY PRESENTING, THE SURGEON HAND-SUTURED CLOSED. THERE WERE NO PATIENT CONSEQUENCES.
18-10-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Component Missing | A020602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN ENHANCED RECOVERY AFTER SURGERY (ERAS) LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS A MANUFACTURING DEFECT WITH THE RELOAD USED. THE SURGEON HAD COMPLETED ONE OF THE TOTAL FIVE FIRINGS. ON THE SECOND FIRING WHEN THE PURPLE RELOAD WAS REMOVED FROM THE PACKAGING THE SUTURE THAT WAS TO RETAIN THE BUTTRESS MATERIAL ON THE UPPER JAW OF THE RELOAD WAS MISSING. THE BUTTRESS MATERIAL WAS NOT RETAINED ON THE ANTERIOR SIDE OF THE RELOAD, HENCE THE SURGEON DID NOT USE THE RELOAD AND REMOVED THE DEVICE FROM THE PATIENT AND OBTAINED A NEW PURPLE RELOAD FOR THE FIRING. THERE WAS NO PATIENT INJURY. THE SURGEON COMMENTED THAT THIS HAD HAPPENED TO THEM IN A CASE THE PREVIOUS WEEK AND WAS WONDERING IF IT WAS ALL THE SAME LOT OF RELOAD BUT DID NOT HAVE ANY FURTHER INFORMATION.
18-10-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN ENHANCED RECOVERY AFTER SURGERY (ERAS) LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS A MANUFACTURING DEFECT WITH THE RELOAD USED. THE SURGEON HAD COMPLETED ONE OF THE TOTAL FIVE FIRINGS. ON THE SECOND FIRING WHEN THE PURPLE RELOAD WAS REMOVED FROM THE PACKAGING THE SUTURE THAT WAS TO RETAIN THE BUTTRESS MATERIAL ON THE UPPER JAW OF THE RELOAD WAS MISSING. THE BUTTRESS MATERIAL WAS NOT RETAINED ON THE ANTERIOR SIDE OF THE RELOAD, HENCE THE SURGEON DID NOT USE THE RELOAD AND REMOVED THE DEVICE FROM THE PATIENT AND OBTAINED A NEW PURPLE RELOAD FOR THE FIRING. THERE WAS NO PATIENT INJURY. THE SURGEON COMMENTED THAT THIS HAD HAPPENED TO THEM IN A CASE THE PREVIOUS WEEK AND WAS WONDERING IF IT WAS ALL THE SAME LOT OF RELOAD BUT DID NOT HAVE ANY FURTHER INFORMATION.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SEMI-OPEN COLECTOMY, THE STAPLE WAS MALFORMED AT 2ND FIRING ON THE FEEA. THE DEVICE WAS USED ON THE COLON. THE DEVICE IN THE QUESTION WAS USED, AND ANOTHER CARTRIDGE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
16-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. THE DEVICE OPERATED UNINTENDEDLY. WHEN THE RIGHT ARTICULATION WAS ATTEMPTED, THE JAWS CLOSED. THE RELOAD ALSO COULD NOT BE REMOVED FROM THE ADAPTER EASILY. THIS HAPPENED ON TWO RELOADS. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
16-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. THE DEVICE OPERATED UNINTENDEDLY. WHEN THE RIGHT ARTICULATION WAS ATTEMPTED, THE JAWS CLOSED. THE RELOAD ALSO COULD NOT BE REMOVED FROM THE ADAPTER EASILY. THIS HAPPENED ON TWO RELOADS. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
16-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Unintended Movement | A0512 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. THE DEVICE OPERATED UNINTENDEDLY. WHEN THE RIGHT ARTICULATION WAS ATTEMPTED, THE JAWS CLOSED. THE RELOAD ALSO COULD NOT BE REMOVED FROM THE ADAPTER EASILY. THIS HAPPENED ON TWO RELOADS. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION PROCEDURE THAT THE DOCTOR AT (B)(6) HOSPITAL DESCRIBED CONTINUED ISSUES WITH THE REMOVAL OF THE CDH29P. THE STAPLER WAS NOT RELEASING FROM STAPLE LINE AS INTENDED AND DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION PROCEDURE THAT THE DOCTOR AT (B)(6) HOSPITAL DESCRIBED CONTINUED ISSUES WITH THE REMOVAL OF THE CDH29P. THE STAPLER WAS NOT RELEASING FROM STAPLE LINE AS INTENDED AND DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION PROCEDURE THAT THE DOCTOR AT (B)(6) HOSPITAL DESCRIBED CONTINUED ISSUES WITH THE REMOVAL OF THE CDH29P. THE STAPLER WAS NOT RELEASING FROM STAPLE LINE AS INTENDED AND DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION PROCEDURE THAT THE DOCTOR AT (B)(6) HOSPITAL DESCRIBED CONTINUED ISSUES WITH THE REMOVAL OF THE CDH29P. THE STAPLER WAS NOT RELEASING FROM STAPLE LINE AS INTENDED AND DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION PROCEDURE THAT THE DOCTOR AT (B)(6) HOSPITAL DESCRIBED CONTINUED ISSUES WITH THE REMOVAL OF THE CDH29P. THE STAPLER WAS NOT RELEASING FROM STAPLE LINE AS INTENDED AND DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION PROCEDURE THAT THE DOCTOR AT (B)(6) HOSPITAL DESCRIBED CONTINUED ISSUES WITH THE REMOVAL OF THE CDH29P. THE STAPLER WAS NOT RELEASING FROM STAPLE LINE AS INTENDED AND DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
16-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. THE DEVICE OPERATED UNINTENDEDLY. WHEN THE RIGHT ARTICULATION WAS ATTEMPTED, THE JAWS CLOSED. THE RELOAD ALSO COULD NOT BE REMOVED FROM THE ADAPTER EASILY. THIS HAPPENED ON 2 RELOADS. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
16-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. THE DEVICE OPERATED UNINTENDEDLY. WHEN THE RIGHT ARTICULATION WAS ATTEMPTED, THE JAWS CLOSED. THE RELOAD ALSO COULD NOT BE REMOVED FROM THE ADAPTER EASILY. THIS HAPPENED ON 2 RELOADS. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
16-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Unintended Movement | A0512 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE PRODUCT WAS ATTACHED TO THE POWERED STAPLER BUT THE JAWS COULD NOT BE CLOSED WHEN TISSUE WAS APPROACHED AFTER CHECKING THE OPENING AND CLOSING. THE DEVICE OPERATED UNINTENDEDLY. WHEN THE RIGHT ARTICULATION WAS ATTEMPTED, THE JAWS CLOSED. THE RELOAD ALSO COULD NOT BE REMOVED FROM THE ADAPTER EASILY. THIS HAPPENED ON 2 RELOADS. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PROXIMATE*LINEAR CUTTER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AFTER THE DEVICE WAS FIRED, THE STAPLE LEG DIDN'T MAKE A FULL B SHAPE. RETRIEVED RELOAD AND AS MUCH AS STAPLES AS POSSIBLE, HAND SUTURED CUT SITE AND FINISHED SURGERY. THERE WERE NO PATIENT CONSEQUENCES.
18-10-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE OPEN LOBECTOMY, WHEN STAPLING, THE SURGEON COULD NOT SQUEEZE THE HANDLE AND STAPLER DID NOT FIRE. GOT ANOTHER HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-10-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE OPEN LOBECTOMY, WHEN STAPLING, THE SURGEON COULD NOT SQUEEZE THE HANDLE AND STAPLER DID NOT FIRE. GOT ANOTHER HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY THE STAPLER FIRED, NO STAPLES CAME OUT. A NEW STAPLER WAS OPENED. THE SURGERY WAS DELAYED. THERE WERE NO PATIENT CONSEQUENCES.
12-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PULMONARY RESECTION PROCEDURE, WHILE ATTEMPTING TO RESECTION THE INFERIOR PULMONARY VEIN, THE STAPLE DID NOT GO ALL THE WAY AND STIFFENED. THEN JAWS COULD NOT BE OPENED ANYMORE; THE BLACK KNOB COULD NOT BE PULLED UNTIL THE END. THE SURGEON USED FORCEPS TO FORCIBLY OPEN THE JAWS; AND A NEW HANDLE AND RELOAD WERE USED TO PERFORM THE RESECTION. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PULMONARY RESECTION PROCEDURE, WHILE ATTEMPTING TO RESECTION THE INFERIOR PULMONARY VEIN, THE STAPLE DID NOT GO ALL THE WAY AND STIFFENED. THEN JAWS COULD NOT BE OPENED ANYMORE; THE BLACK KNOB COULD NOT BE PULLED UNTIL THE END. THE SURGEON USED FORCEPS TO FORCIBLY OPEN THE JAWS; AND A NEW HANDLE AND RELOAD WERE USED TO PERFORM THE RESECTION. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE STAPLER DID NOT RELEASE FROM THE STAPLE LINE AS INTENDED AND THE DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE STAPLER DID NOT RELEASE FROM THE STAPLE LINE AS INTENDED AND THE DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE STAPLER DID NOT RELEASE FROM THE STAPLE LINE AS INTENDED AND THE DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE STAPLER DID NOT RELEASE FROM THE STAPLE LINE AS INTENDED AND THE DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE STAPLER DID NOT RELEASE FROM THE STAPLE LINE AS INTENDED AND THE DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE STAPLER DID NOT RELEASE FROM THE STAPLE LINE AS INTENDED AND THE DIFFICULT REMOVALS HAVE CAUSED TRAUMA AND DAMAGE TO THE STAPLE LINE. IT WAS VERY DIFFICULT TO RELEASE THE STAPLER FROM THE ANASTOMOSIS WHILE FOLLOWING PROPER STEPS FOR REMOVAL. CAUSED VERY BLOODY INTERNAL LUMEN AT THE STAPLE LINE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. IT IS UNKNOWN IF THERE WAS ANY CONSEQUENCES.
21-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | CIRCULAR MECH XL SEAL 25MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID COLECTOMY THAT THE DOCTOR HAD A STAPLER MISFIRE. THE DEVICE CUT BUT THERE WEREN¿T ANY STAPLES AT ALL IN THE DEVICE. UNKNOWN OF ANY PATIENT CONSEQUENCES.
12-07-2020 | Injury | us surgical puerto rico | 10-01-2023 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: THE PATIENTS ATTORNEY ALLEGED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE SURGEON USED AN ENDO GIA STAPLER TO FORM A STAPLE LINE. FOLLOWING THE TRANSECTION, THE SURGEON IDENTIFIED A SMALL BLEEDING VESSEL ABOVE THE STAPLE LINE. A CLIP WAS USED TO STOP THE BLEEDING. POST-OPERATIVELY, THE PATIENT EXPERIENCED DECREASE IN BLOOD PRESSURE, INTRAABDOMINAL BLEEDING, INCREASE IN HEART RATE, LOW HEMOGLOBIN COUNT, INTRAPERITONEAL HEMATOMA, LOCALIZED PAIN TO POST-OP, ABDOMINAL PAIN, FLUID COLLECTION/ABSCESS. THE PATIENTS TREATMENTS INCLUDED EXTENDED HOSPITAL STAY, CT SCAN, ANTIBIOTICS, BLOOD TRANSFUSION.
12-07-2020 | Injury | us surgical puerto rico | 10-01-2023 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: THE PATIENTS ATTORNEY ALLEGED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE SURGEON USED AN ENDO GIA STAPLER TO FORM A STAPLE LINE. FOLLOWING THE TRANSECTION, THE SURGEON IDENTIFIED A SMALL BLEEDING VESSEL ABOVE THE STAPLE LINE. A CLIP WAS USED TO STOP THE BLEEDING. POST-OPERATIVELY, THE PATIENT EXPERIENCED DECREASE IN BLOOD PRESSURE, INTRAABDOMINAL BLEEDING, INCREASE IN HEART RATE, LOW HEMOGLOBIN COUNT, INTRAPERITONEAL HEMATOMA, LOCALIZED PAIN TO POST-OP, ABDOMINAL PAIN, FLUID COLLECTION/ABSCESS. THE PATIENTS TREATMENTS INCLUDED EXTENDED HOSPITAL STAY, CT SCAN, ANTIBIOTICS, BLOOD TRANSFUSION.
12-07-2020 | Injury | us surgical puerto rico | 10-01-2023 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Anemia | Event Description: THE PATIENTS ATTORNEY ALLEGED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE SURGEON USED AN ENDO GIA STAPLER TO FORM A STAPLE LINE. FOLLOWING THE TRANSECTION, THE SURGEON IDENTIFIED A SMALL BLEEDING VESSEL ABOVE THE STAPLE LINE. A CLIP WAS USED TO STOP THE BLEEDING. POST-OPERATIVELY, THE PATIENT EXPERIENCED DECREASE IN BLOOD PRESSURE, INTRAABDOMINAL BLEEDING, INCREASE IN HEART RATE, LOW HEMOGLOBIN COUNT, INTRAPERITONEAL HEMATOMA, LOCALIZED PAIN TO POST-OP, ABDOMINAL PAIN, FLUID COLLECTION/ABSCESS. THE PATIENTS TREATMENTS INCLUDED EXTENDED HOSPITAL STAY, CT SCAN, ANTIBIOTICS, BLOOD TRANSFUSION.
12-07-2020 | Injury | us surgical puerto rico | 10-01-2023 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: THE PATIENTS ATTORNEY ALLEGED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE SURGEON USED AN ENDO GIA STAPLER TO FORM A STAPLE LINE. FOLLOWING THE TRANSECTION, THE SURGEON IDENTIFIED A SMALL BLEEDING VESSEL ABOVE THE STAPLE LINE. A CLIP WAS USED TO STOP THE BLEEDING. POST-OPERATIVELY, THE PATIENT EXPERIENCED DECREASE IN BLOOD PRESSURE, INTRAABDOMINAL BLEEDING, INCREASE IN HEART RATE, LOW HEMOGLOBIN COUNT, INTRAPERITONEAL HEMATOMA, LOCALIZED PAIN TO POST-OP, ABDOMINAL PAIN, FLUID COLLECTION/ABSCESS. THE PATIENTS TREATMENTS INCLUDED EXTENDED HOSPITAL STAY, CT SCAN, ANTIBIOTICS, BLOOD TRANSFUSION.
12-07-2020 | Injury | us surgical puerto rico | 10-01-2023 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: THE PATIENTS ATTORNEY ALLEGED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE SURGEON USED AN ENDO GIA STAPLER TO FORM A STAPLE LINE. FOLLOWING THE TRANSECTION, THE SURGEON IDENTIFIED A SMALL BLEEDING VESSEL ABOVE THE STAPLE LINE. A CLIP WAS USED TO STOP THE BLEEDING. POST-OPERATIVELY, THE PATIENT EXPERIENCED DECREASE IN BLOOD PRESSURE, INTRAABDOMINAL BLEEDING, INCREASE IN HEART RATE, LOW HEMOGLOBIN COUNT, INTRAPERITONEAL HEMATOMA, LOCALIZED PAIN TO POST-OP, ABDOMINAL PAIN, FLUID COLLECTION/ABSCESS. THE PATIENTS TREATMENTS INCLUDED EXTENDED HOSPITAL STAY, CT SCAN, ANTIBIOTICS, BLOOD TRANSFUSION.
12-07-2020 | Injury | us surgical puerto rico | 10-01-2023 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Low Blood Pressure/ Hypotension | Event Description: THE PATIENTS ATTORNEY ALLEGED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE SURGEON USED AN ENDO GIA STAPLER TO FORM A STAPLE LINE. FOLLOWING THE TRANSECTION, THE SURGEON IDENTIFIED A SMALL BLEEDING VESSEL ABOVE THE STAPLE LINE. A CLIP WAS USED TO STOP THE BLEEDING. POST-OPERATIVELY, THE PATIENT EXPERIENCED DECREASE IN BLOOD PRESSURE, INTRAABDOMINAL BLEEDING, INCREASE IN HEART RATE, LOW HEMOGLOBIN COUNT, INTRAPERITONEAL HEMATOMA, LOCALIZED PAIN TO POST-OP, ABDOMINAL PAIN, FLUID COLLECTION/ABSCESS. THE PATIENTS TREATMENTS INCLUDED EXTENDED HOSPITAL STAY, CT SCAN, ANTIBIOTICS, BLOOD TRANSFUSION.
12-07-2020 | Injury | us surgical puerto rico | 10-01-2023 | nan | Adverse Event Without Identified Device or Use Problem | A24 | Tachycardia | Event Description: THE PATIENTS ATTORNEY ALLEGED THAT DURING A LAPAROSCOPIC APPENDECTOMY, THE SURGEON USED AN ENDO GIA STAPLER TO FORM A STAPLE LINE. FOLLOWING THE TRANSECTION, THE SURGEON IDENTIFIED A SMALL BLEEDING VESSEL ABOVE THE STAPLE LINE. A CLIP WAS USED TO STOP THE BLEEDING. POST-OPERATIVELY, THE PATIENT EXPERIENCED DECREASE IN BLOOD PRESSURE, INTRAABDOMINAL BLEEDING, INCREASE IN HEART RATE, LOW HEMOGLOBIN COUNT, INTRAPERITONEAL HEMATOMA, LOCALIZED PAIN TO POST-OP, ABDOMINAL PAIN, FLUID COLLECTION/ABSCESS. THE PATIENTS TREATMENTS INCLUDED EXTENDED HOSPITAL STAY, CT SCAN, ANTIBIOTICS, BLOOD TRANSFUSION.
23-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | 55 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE RELOAD DID NOT DELIVER ANY STAPLES WHEN FIRING. THE SURGERY WAS DELAYED 20 MINUTES. HAND SUTURING TOOK PLACE TO SUCCESSFULLY COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES.
16-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, PANCREATICODUODENECTOMY, AND SMALL BOWEL RESECTION, DURING THE FOURTH FIRING, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, PANCREATICODUODENECTOMY, AND SMALL BOWEL RESECTION, DURING THE FOURTH FIRING, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, WHILE THE DEVICE WAS USED ON THE SIGMOID, THE TISSUE COMPRESSION SCALE BACKLIGHT WAS ILLUMINATED, BUT THE DEVICE COULD NOT BE FIRED ALTHOUGH THE TRIGGER WAS GRASPED. ADDITIONAL CUT WAS PERFORMED ON INTESTINAL TRACT. THE ANVIL OF THE DEVICE IN THE QUESTION WAS PLACED AS IT WAS, AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PROBLEM IN TIGHTENING THE ADJUSTING KNOB, AND RELEASING THE RED SAFETY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, WHILE THE DEVICE WAS USED ON THE SIGMOID, THE TISSUE COMPRESSION SCALE BACKLIGHT WAS ILLUMINATED, BUT THE DEVICE COULD NOT BE FIRED ALTHOUGH THE TRIGGER WAS GRASPED. ADDITIONAL CUT WAS PERFORMED ON INTESTINAL TRACT. THE ANVIL OF THE DEVICE IN THE QUESTION WAS PLACED AS IT WAS, AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PROBLEM IN TIGHTENING THE ADJUSTING KNOB, AND RELEASING THE RED SAFETY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS LACK OF STAPLING AND THE METAL PART OF THE LOADING UNIT CAME APART AND REMAINED IN THE PATIENT'S ABDOMEN. THE SURGEON FOUND IT IN THE COLON.
20-12-2022 | Injury | us surgical puerto rico | 10-01-2023 | GIA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS LACK OF STAPLING AND THE METAL PART OF THE LOADING UNIT CAME APART AND REMAINED IN THE PATIENT'S ABDOMEN. THE SURGEON FOUND IT IN THE COLON.
25-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | CIRCULAR MECH STD 25MM | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ANVIL COULD NOT BE REMOVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RECTUM PROCEDURE, WHEN THE DISTAL RECTUM WAS SECTIONED/STAPLED WITH THE CONTOUR STAPLER, THE DISTAL RECTUM WAS POSITIONED IN THE STAPLER AND THE CORRECT MOVEMENTS WERE FOLLOWED. STAPLES WERE NOT FIXED IN THE RECTUM OR WERE LOOSE IN THE PELVIS. THEREFORE, A NEW STAPLING WITH ANOTHER CONTOUR WAS CHOSEN, WHICH WORKED PERFECTLY.
25-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BOWEL PROCEDURE, AFTER FIRED, NOTED THE STAPLES WERE MALFORMED AND THE DONUTS WERE INCOMPLETE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-11-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BOWEL PROCEDURE, AFTER FIRED, NOTED THE STAPLES WERE MALFORMED AND THE DONUTS WERE INCOMPLETE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2022 | Malfunction | ethicon endo-surgery | 10-01-2023 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGEAL SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, WHEN THE USER PRESSED THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT WAS INVOLVEMENT.
22-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECONSTRUCTION AFTER AN ESOPHAGECTOMY, WHEN THE SURGEON USED THE DEVICE TO MAKE THE GASTRIC TUBE, THE DEVICE DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, IN CHOLECYSTIC ARTERY RESECTION, AN ATTEMPT WAS MADE TO FIRE BUT THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WERE ABLE TO BE REMOVED IT NORMALLY. WHEN THE RELOAD WAS CHECKED, IT HAD FEW SIGNS OF HAVING BEEN FIRED, BUT PRE-FIRE MAY NOT HAVE BEEN PERFORMED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
23-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, IN CHOLECYSTIC ARTERY RESECTION, AN ATTEMPT WAS MADE TO FIRE BUT THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WERE ABLE TO BE REMOVED IT NORMALLY. WHEN THE RELOAD WAS CHECKED, IT HAD FEW SIGNS OF HAVING BEEN FIRED, BUT PRE-FIRE MAY NOT HAVE BEEN PERFORMED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
23-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, IN CHOLECYSTIC ARTERY RESECTION, AN ATTEMPT WAS MADE TO FIRE, BUT THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, THE JAWS OF THE DEVICE LOCKED ON TISSUE BUT WERE ABLE TO BE REMOVED IT NORMALLY. WHEN THE RELOAD WAS CHECKED, IT HAD FEW SIGNS OF HAVING BEEN FIRED, BUT PRE-FIRE MAY NOT HAVE BEEN PERFORMED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, IN RECTUM RESECTION, THE FIRST FIRING WITH A REINFORCED FORTY FIVE MILLIMETER (45MM) RELOAD WAS COMPLETED, BUT THE POWERED HANDLE DISPLAY SHOWED AN IMPEDANCE VALUE HIGHEST FORCE ZONE WHICH WAS 3 AT ONE TIME. THE FIRST FIRING WAS PERFORMED WITHOUT ARTICULATION. THE SECOND FIRING WAS STARTED WITH THE SAME TYPE OF RELOAD IN MAXIMUM ARTICULATION, BUT IN APPROXIMATELY TWO CENTIMETERS (2CM) OF FIRING, THE LIMIT OF ADAPTATION APPEARED ON THE HANDLE DISPLAY; AN EXCESSIVE FORCE INDICATION WAS DISPLAYED. THE FIRING BUTTON WAS PRESSED TWICE, BUT THE FIRING COULD NOT BE COMPLETED, SO THE JAWS WERE RELEASED. ADDITIONALLY, THE TIP SIDE OF THE STAPLE FROM THE SECOND RELOAD HAD A U-SHAPED. THE SHEET WAS RESECTED WITH LAPAROSCOPY SCISSORS AND WAS RELEASED FROM THE TISSUE. A SIXTY MILLIMETER (60MM) RELOAD WAS THEN USED TO COMPLETE THE RESECTION OF THE REMAINING RECTUM. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, IN RECTUM RESECTION, THE FIRST FIRING WITH A REINFORCED FORTY FIVE MILLIMETER (45MM) RELOAD WAS COMPLETED, BUT THE POWERED HANDLE DISPLAY SHOWED AN IMPEDANCE VALUE HIGHEST FORCE ZONE WHICH WAS 3 AT ONE TIME. THE FIRST FIRING WAS PERFORMED WITHOUT ARTICULATION. THE SECOND FIRING WAS STARTED WITH THE SAME TYPE OF RELOAD IN MAXIMUM ARTICULATION, BUT IN APPROXIMATELY TWO CENTIMETERS (2CM) OF FIRING, THE LIMIT OF ADAPTATION APPEARED ON THE HANDLE DISPLAY;AN EXCESSIVE FORCE INDICATION WAS DISPLAYED. THE FIRING BUTTON WAS PRESSED TWICE, BUT THE FIRING COULD NOT BE COMPLETED, SO THE JAWS WERE RELEASED. ADDITIONALLY, THE TIP SIDE OF THE STAPLE FROM THE SECOND RELOAD HAD A U-SHAPED. THE SHEET WAS RESECTED WITH LAPAROSCOPY SCISSORS AND WAS RELEASED FROM THE TISSUE. A SIXTY MILLIMETER (60MM) RELOAD WAS THEN USED TO COMPLETE THE RESECTION OF THE REMAINING RECTUM. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, IN RECTUM RESECTION, THE FIRST FIRING WITH A REINFORCED FORTY FIVE MILLIMETER (45MM) RELOAD WAS COMPLETED, BUT THE POWERED HANDLE DISPLAY SHOWED AN IMPEDANCE VALUE HIGHEST FORCE ZONE WHICH WAS 3 AT ONE TIME. THE FIRST FIRING WAS PERFORMED WITHOUT ARTICULATION. THE SECOND FIRING WAS STARTED WITH THE SAME TYPE OF RELOAD IN MAXIMUM ARTICULATION, BUT IN APPROXIMATELY TWO CENTIMETERS (2CM) OF FIRING, THE LIMIT OF ADAPTATION APPEARED ON THE HANDLE DISPLAY;AN EXCESSIVE FORCE INDICATION WAS DISPLAYED. THE FIRING BUTTON WAS PRESSED TWICE, BUT THE FIRING COULD NOT BE COMPLETED, SO THE JAWS WERE RELEASED. ADDITIONALLY, THE TIP SIDE OF THE STAPLE FROM THE SECOND RELOAD HAD A U-SHAPED. THE SHEET WAS RESECTED WITH LAPAROSCOPY SCISSORS AND WAS RELEASED FROM THE TISSUE. A SIXTY MILLIMETER (60MM) RELOAD WAS THEN USED TO COMPLETE THE RESECTION OF THE REMAINING RECTUM. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, IN RECTUM RESECTION, THE FIRST FIRING WITH A REINFORCED FORTY FIVE MILLIMETER (45MM) RELOAD WAS COMPLETED, BUT THE POWERED HANDLE DISPLAY SHOWED AN IMPEDANCE VALUE HIGHEST FORCE ZONE WHICH WAS 3 AT ONE TIME. THE FIRST FIRING WAS PERFORMED WITHOUT ARTICULATION. THE SECOND FIRING WAS STARTED WITH THE SAME TYPE OF RELOAD IN MAXIMUM ARTICULATION, BUT IN APPROXIMATELY TWO CENTIMETERS (2CM) OF FIRING, THE LIMIT OF ADAPTATION APPEARED ON THE HANDLE DISPLAY;AN EXCESSIVE FORCE INDICATION WAS DISPLAYED. THE FIRING BUTTON WAS PRESSED TWICE, BUT THE FIRING COULD NOT BE COMPLETED, SO THE JAWS WERE RELEASED. ADDITIONALLY, THE TIP SIDE OF THE STAPLE FROM THE SECOND RELOAD HAD A U-SHAPED. THE SHEET WAS RESECTED WITH LAPAROSCOPY SCISSORS AND WAS RELEASED FROM THE TISSUE. A SIXTY MILLIMETER (60MM) RELOAD WAS THEN USED TO COMPLETE THE RESECTION OF THE REMAINING RECTUM. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | covidien lp llc north haven | 10-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE BLACK REINFORCED RELOAD WAS ATTEMPTED TO BE USED BUT FIRING COULD NOT BE DONE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. OPENED A NEW PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE BLACK REINFORCED RELOAD WAS ATTEMPTED TO BE USED BUT FIRING COULD NOT BE DONE. THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. OPENED A NEW PRODUCT TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | us surgical puerto rico | 10-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, WHILE AT INTESTINAL RESECTION, UPON USING THE PRODUCT, THE HANDLE COULD NOT BE SQUEEZED EVEN WHEN THE GREEN BUTTON WAS PRESSED AND THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED, BUT THE SITUATION DID NOT CHANGE, THE STAPLER STILL DID NOT FIRE. THE DEVICE WAS FIXATED TO TISSUE AND WAS REMOVED WITHOUT ANY PROBLEM. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. IT WAS NOTED THAT THE SITUATION BEFORE PRESSING THE GREEN BUTTON HAS NOT BEEN CONFIRMED AT THIS POINT, AND WHEN THE COLLECTED CARTRIDGE WAS CHECKED, THE STAPLER TIP WAS SLIGHTLY PROTRUDING INSIDE THE JAWS.
09-12-2022 | Injury | johnson & johnson / ethicon endo-surgery | 10-01-2023 | ETHICON VASCULAR STAPLER | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: PRE-OP DX: LEFT RENAL MASS. PROCEDURE: ROBOTICALLY-ASSISTED LAPAROSCOPIC LEFT RADICAL NEPHRECTOMY (DAVINCI). SURGEON: UROLOGIST, ANESTHESIA. STAFF: FIRST ASSISTANT, SCRUB, CIRCULATOR. IN ROOM 0732, OUT OF ROOM 1421. UPON OPENING THE STAPLER, IT WAS NOTICED THAT IT MISFIRED. THE MARYLAND BIPOLAR WAS THEN USED TO GRASP THE RENAL ARTERY AS IT WAS BLEEDING. THEN ATTEMPTED TO PLACE ROBOTIC HEM-O-LOK WHERE THE VESSEL SEALER WAS, HOWEVER, THIS PORT WAS INADVERTENTLY PULLED OUT SO THE ASSISTANT PORT WAS THEN USED TO PUT THE KIDNEY ON TRACTION. THE FOURTH ARM WAS THEN USED TO GRASP THE BLEEDER. NEXT, THE MARYLAND WAS REMOVED AND HEM-O-LOK CLIP APPLIER WAS THEN PLACED AND A HEM O-LOK WAS PLACED AROUND THE RENAL ARTERY. ONCE THIS WAS DONE, THERE WAS NO MORE BLEEDING. THE ARTERY BLED FOR APPROXIMATELY 45-60 SECONDS AND THERE WAS 1 L OF BLOOD LOSS AT THIS TIME. NEXT, THERE WAS A NEW VASCULAR STAPLER WITH NEW VASCULAR LOADS ON THE FIELD AND THE RENAL VEIN WAS TRANSECTED WITHOUT COMPLICATION WITH THE STAPLER AND HEM-O-LOK WAS THEN PLACED ALONG THE STAPLE LINE. THERE WAS NO ACTIVE BLEEDING AND SURGICEL POWDER, AS WELL AS TISSEEL WAS PLACED IN THE .SURGICAL AREA. ONCE THE SPECIMEN WAS REMOVED AND THE FASCIA WAS CLOSED, THE ABDOMEN WAS REINSUFFLATED. THERE WAS NO BOWEL INJURY AS A RESULT OF THE CLOSURE. A CARTER-THOMASON WAS THEN USED TO C:LOSE THE TROCAR SITES. (B)(6) 2022 17:14 POST OP HGB 9.8 AND HCT 33.3; (B)(6) 2022 PT. RECEIVED PRBC 1 UNIT (250CC) AND ALBUMIN 500CC. (B)(6) 2022 0620 HGB 5.7 AND HCT 29.8, POD#1 (B)(6) COURSE. PAIN CONTROLLED, AMBULATED POD #1 (B)(6) 2022 HGB 6.7 AND HCT 29.8, DISCHARGED HOME SELF CARE. (B)(6) 2022 HGB 9.1. RETURNED TO ED ALTERED MENTAL STATUS., UNABLE TO KEEP ANTIBIOTICS POST OP SURGERY, HOSPITAL COURSE: UROLOGY AND NEPHROLOGY WERE CONSULTED ON THE PATIENT'S CASE. PATIENT RECEIVED IVF HYDRATION, 3 DAYS OF ROCEPHIN THERAPY, AND REPLACEMENT OF HER ELECTROLYTES. HER MENTAL STATUS RETURNED TO BASELINE AND. SHE WAS EATING WELL WITHOUT ANY FURTHER VOMITING. SHE WORKED WITH PT/OT AND THEY RECOMMENDED PLACEMENT; HOWEVER, PATIENT REFUSED AND WAS AGREEABLE WITH HOME PT/OT. FOLEY KEPT IN PLACE UNTIL FOLLOW UP WITH UROLOGY OUTPATIENT. (B)(6) 2022 PATIENT DISCHARGED IN A HEMODYNAMICALLY STABLE CONDITION.(B)(6) 2022 PT RECEIVED PRBC 1 UNIT (250CC) AND ALBUMIN 500CC.
09-12-2022 | Injury | johnson & johnson / ethicon endo-surgery | 10-01-2023 | ETHICON VASCULAR STAPLER | Misfire | A050502 | Vomiting | Event Description: PRE-OP DX: LEFT RENAL MASS. PROCEDURE: ROBOTICALLY-ASSISTED LAPAROSCOPIC LEFT RADICAL NEPHRECTOMY (DAVINCI). SURGEON: UROLOGIST, ANESTHESIA. STAFF: FIRST ASSISTANT, SCRUB, CIRCULATOR. IN ROOM 0732, OUT OF ROOM 1421. UPON OPENING THE STAPLER, IT WAS NOTICED THAT IT MISFIRED. THE MARYLAND BIPOLAR WAS THEN USED TO GRASP THE RENAL ARTERY AS IT WAS BLEEDING. THEN ATTEMPTED TO PLACE ROBOTIC HEM-O-LOK WHERE THE VESSEL SEALER WAS, HOWEVER, THIS PORT WAS INADVERTENTLY PULLED OUT SO THE ASSISTANT PORT WAS THEN USED TO PUT THE KIDNEY ON TRACTION. THE FOURTH ARM WAS THEN USED TO GRASP THE BLEEDER. NEXT, THE MARYLAND WAS REMOVED AND HEM-O-LOK CLIP APPLIER WAS THEN PLACED AND A HEM O-LOK WAS PLACED AROUND THE RENAL ARTERY. ONCE THIS WAS DONE, THERE WAS NO MORE BLEEDING. THE ARTERY BLED FOR APPROXIMATELY 45-60 SECONDS AND THERE WAS 1 L OF BLOOD LOSS AT THIS TIME. NEXT, THERE WAS A NEW VASCULAR STAPLER WITH NEW VASCULAR LOADS ON THE FIELD AND THE RENAL VEIN WAS TRANSECTED WITHOUT COMPLICATION WITH THE STAPLER AND HEM-O-LOK WAS THEN PLACED ALONG THE STAPLE LINE. THERE WAS NO ACTIVE BLEEDING AND SURGICEL POWDER, AS WELL AS TISSEEL WAS PLACED IN THE .SURGICAL AREA. ONCE THE SPECIMEN WAS REMOVED AND THE FASCIA WAS CLOSED, THE ABDOMEN WAS REINSUFFLATED. THERE WAS NO BOWEL INJURY AS A RESULT OF THE CLOSURE. A CARTER-THOMASON WAS THEN USED TO C:LOSE THE TROCAR SITES. (B)(6) 2022 17:14 POST OP HGB 9.8 AND HCT 33.3; (B)(6) 2022 PT. RECEIVED PRBC 1 UNIT (250CC) AND ALBUMIN 500CC. (B)(6) 2022 0620 HGB 5.7 AND HCT 29.8, POD#1 (B)(6) COURSE. PAIN CONTROLLED, AMBULATED POD #1 (B)(6) 2022 HGB 6.7 AND HCT 29.8, DISCHARGED HOME SELF CARE. (B)(6) 2022 HGB 9.1. RETURNED TO ED ALTERED MENTAL STATUS., UNABLE TO KEEP ANTIBIOTICS POST OP SURGERY, HOSPITAL COURSE: UROLOGY AND NEPHROLOGY WERE CONSULTED ON THE PATIENT'S CASE. PATIENT RECEIVED IVF HYDRATION, 3 DAYS OF ROCEPHIN THERAPY, AND REPLACEMENT OF HER ELECTROLYTES. HER MENTAL STATUS RETURNED TO BASELINE AND. SHE WAS EATING WELL WITHOUT ANY FURTHER VOMITING. SHE WORKED WITH PT/OT AND THEY RECOMMENDED PLACEMENT; HOWEVER, PATIENT REFUSED AND WAS AGREEABLE WITH HOME PT/OT. FOLEY KEPT IN PLACE UNTIL FOLLOW UP WITH UROLOGY OUTPATIENT. (B)(6) 2022 PATIENT DISCHARGED IN A HEMODYNAMICALLY STABLE CONDITION.(B)(6) 2022 PT RECEIVED PRBC 1 UNIT (250CC) AND ALBUMIN 500CC.
09-12-2022 | Injury | johnson & johnson / ethicon endo-surgery | 10-01-2023 | ETHICON VASCULAR STAPLER | Misfire | A050502 | Cognitive Changes | Event Description: PRE-OP DX: LEFT RENAL MASS. PROCEDURE: ROBOTICALLY-ASSISTED LAPAROSCOPIC LEFT RADICAL NEPHRECTOMY (DAVINCI). SURGEON: UROLOGIST, ANESTHESIA. STAFF: FIRST ASSISTANT, SCRUB, CIRCULATOR. IN ROOM 0732, OUT OF ROOM 1421. UPON OPENING THE STAPLER, IT WAS NOTICED THAT IT MISFIRED. THE MARYLAND BIPOLAR WAS THEN USED TO GRASP THE RENAL ARTERY AS IT WAS BLEEDING. THEN ATTEMPTED TO PLACE ROBOTIC HEM-O-LOK WHERE THE VESSEL SEALER WAS, HOWEVER, THIS PORT WAS INADVERTENTLY PULLED OUT SO THE ASSISTANT PORT WAS THEN USED TO PUT THE KIDNEY ON TRACTION. THE FOURTH ARM WAS THEN USED TO GRASP THE BLEEDER. NEXT, THE MARYLAND WAS REMOVED AND HEM-O-LOK CLIP APPLIER WAS THEN PLACED AND A HEM O-LOK WAS PLACED AROUND THE RENAL ARTERY. ONCE THIS WAS DONE, THERE WAS NO MORE BLEEDING. THE ARTERY BLED FOR APPROXIMATELY 45-60 SECONDS AND THERE WAS 1 L OF BLOOD LOSS AT THIS TIME. NEXT, THERE WAS A NEW VASCULAR STAPLER WITH NEW VASCULAR LOADS ON THE FIELD AND THE RENAL VEIN WAS TRANSECTED WITHOUT COMPLICATION WITH THE STAPLER AND HEM-O-LOK WAS THEN PLACED ALONG THE STAPLE LINE. THERE WAS NO ACTIVE BLEEDING AND SURGICEL POWDER, AS WELL AS TISSEEL WAS PLACED IN THE .SURGICAL AREA. ONCE THE SPECIMEN WAS REMOVED AND THE FASCIA WAS CLOSED, THE ABDOMEN WAS REINSUFFLATED. THERE WAS NO BOWEL INJURY AS A RESULT OF THE CLOSURE. A CARTER-THOMASON WAS THEN USED TO C:LOSE THE TROCAR SITES. (B)(6) 2022 17:14 POST OP HGB 9.8 AND HCT 33.3; (B)(6) 2022 PT. RECEIVED PRBC 1 UNIT (250CC) AND ALBUMIN 500CC. (B)(6) 2022 0620 HGB 5.7 AND HCT 29.8, POD#1 (B)(6) COURSE. PAIN CONTROLLED, AMBULATED POD #1 (B)(6) 2022 HGB 6.7 AND HCT 29.8, DISCHARGED HOME SELF CARE. (B)(6) 2022 HGB 9.1. RETURNED TO ED ALTERED MENTAL STATUS., UNABLE TO KEEP ANTIBIOTICS POST OP SURGERY, HOSPITAL COURSE: UROLOGY AND NEPHROLOGY WERE CONSULTED ON THE PATIENT'S CASE. PATIENT RECEIVED IVF HYDRATION, 3 DAYS OF ROCEPHIN THERAPY, AND REPLACEMENT OF HER ELECTROLYTES. HER MENTAL STATUS RETURNED TO BASELINE AND. SHE WAS EATING WELL WITHOUT ANY FURTHER VOMITING. SHE WORKED WITH PT/OT AND THEY RECOMMENDED PLACEMENT; HOWEVER, PATIENT REFUSED AND WAS AGREEABLE WITH HOME PT/OT. FOLEY KEPT IN PLACE UNTIL FOLLOW UP WITH UROLOGY OUTPATIENT. (B)(6) 2022 PATIENT DISCHARGED IN A HEMODYNAMICALLY STABLE CONDITION.(B)(6) 2022 PT RECEIVED PRBC 1 UNIT (250CC) AND ALBUMIN 500CC.
26-12-2022 | Malfunction | us surgical puerto rico | 11-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, WHILE AT INTESTINAL RESECTION, UPON USING THE PRODUCT, THE HANDLE COULD NOT BE SQUEEZED EVEN WHEN THE GREEN BUTTON WAS PRESSED AND THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED, BUT THE SITUATION DID NOT CHANGE, THE STAPLER STILL DID NOT FIRE. THE DEVICE WAS FIXATED TO TISSUE AND WAS REMOVED WITHOUT ANY PROBLEM. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. IT WAS NOTED THAT THE SITUATION BEFORE PRESSING THE GREEN BUTTON HAS NOT BEEN CONFIRMED AT THIS POINT, AND WHEN THE COLLECTED CARTRIDGE WAS CHECKED, THE STAPLER TIP WAS SLIGHTLY PROTRUDING INSIDE THE JAWS.
26-12-2022 | Malfunction | us surgical puerto rico | 11-01-2023 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, WHILE AT INTESTINAL RESECTION, UPON USING THE PRODUCT, THE HANDLE COULD NOT BE SQUEEZED EVEN WHEN THE GREEN BUTTON WAS PRESSED AND THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED, BUT THE SITUATION DID NOT CHANGE, THE STAPLER STILL DID NOT FIRE. THE DEVICE WAS FIXATED TO TISSUE AND WAS REMOVED WITHOUT ANY PROBLEM. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. IT WAS NOTED THAT THE SITUATION BEFORE PRESSING THE GREEN BUTTON HAS NOT BEEN CONFIRMED AT THIS POINT, AND WHEN THE COLLECTED CARTRIDGE WAS CHECKED, THE STAPLER TIP WAS SLIGHTLY PROTRUDING INSIDE THE JAWS.
26-12-2022 | Malfunction | us surgical puerto rico | 11-01-2023 | EEA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, AFTER THE BILLROTH I SURGERY, THE STAPLE WAS MALFORMED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2022 | Malfunction | us surgical puerto rico | 11-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY), WHILE RESECTING THE PULMONARY VEIN, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED, BUT THE DEVICE DID NOT FIRE. A NEW DEVICE OF THE SAME KIND WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2022 | Malfunction | us surgical puerto rico | 11-01-2023 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY), WHILE RESECTING THE PULMONARY VEIN, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED, BUT THE DEVICE DID NOT FIRE. A NEW DEVICE OF THE SAME KIND WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2022 | Malfunction | us surgical puerto rico | 11-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL LUNG RESECTION, THREE FIRINGS WERE DONE WITHOUT ANY PROBLEMS BUT THE FOURTH FIRING COULD NOT BE FIRED. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2022 | Malfunction | us surgical puerto rico | 11-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL LUNG RESECTION, THREE FIRINGS WERE DONE WITHOUT ANY PROBLEMS BUT THE FOURTH FIRING COULD NOT BE FIRED. THE SURGEON WAS NOT ABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2022 | Injury | covidien lp llc north haven | 11-01-2023 | GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVER SEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 11-01-2023 | GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVER SEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 11-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVER SEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 11-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVER SEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
07-12-2022 | Injury | ethicon endo-surgery | 11-01-2023 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE THE SURGEON USED THE DEVICE AND CREATED AN ANASTOMOSIS AFTER THE STAPLES WERE FORMED. THE ANASTOMOSIS LEAKED AND THE SURGEON HAD TO STAPLE WITH AN ECHELON DEVICE TO COMPLETE THE PROCEDURE.
26-12-2022 | Malfunction | covidien lp llc north haven | 11-01-2023 | EEA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, AFTER TAKING THE PURSE STRING SUTURE, THE SURGEON PLACED THE ANVIL, TIED SUTURE IN ANCHOR HOLES THEN ATTACHED STAPLER. THE SURGEON ROTATED THE KNOB BUT IT GOT FREE (LOOSE) AND THE DEVICE WOULD NOT CLOSE (GREEN INDICATOR WAS NOT VISIBLE). THE SURGEON TOOK THE DEVICE OUT AND USED ANOTHER HEMORRHOID STAPLER. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 11-01-2023 | EEA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING OPEN MINIMALLY INVASIVE PROCEDURE FOR HEMORRHOID, AFTER TAKING THE PURSE STRING SUTURE, THE SURGEON PLACED THE ANVIL, TIED SUTURE IN ANCHOR HOLES THEN ATTACHED STAPLER. THE SURGEON ROTATED THE KNOB BUT IT GOT FREE (LOOSE) AND THE DEVICE WOULD NOT CLOSE (GREEN INDICATOR WAS NOT VISIBLE). THE SURGEON TOOK THE DEVICE OUT AND USED ANOTHER HEMORRHOID STAPLER. THERE WAS NO PATIENT INJURY.
03-11-2022 | Injury | ethicon endo-surgery | 11-01-2023 | ILS 25MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT 1WEEK AFTER LAPAROSCOPIC ANTERIOR RESECTION, ANASTOMOTIC LEAKAGE OCCURRED SOON AFTER THE LIQUID FOOD WAS STARTED. IT WAS MINOR LEAK GRADE 2, AND IT DID NOT NEED COLOSTOMY. THE HYPERALIMENTATION WAS USED, AND THE PATIENT HAS BEEN FOLLOWED UP. IN THE INITIAL OPERATION, THE RECTUM RA RB WAS CUT WITH PSEE60A AND GOLD CARTRIDGE FOR 2 TIMES, AND CDH25A WAS USED TO ANASTOMOSE. DST RECONSTRUCTION WAS PERFORMED. THERE WAS NO PROBLEM TO THE DEVICE OR THE STAPLING DURING THE INITIAL OPERATION. THE DEVICE WAS USED BETWEEN COLON AND RECTUM. NO FURTHER INFORMATION IS AVAILABLE.
26-12-2022 | Malfunction | covidien lp llc north haven | 11-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, THERE WAS A PURPLE SUTURE THAT CAME OUT OF THE STAPLE AFTER FIRING THE DEVICE. THE GROOVE WHERE THE KNIFE ADVANCES HAS BEEN SCRAPPED. THE DEVICE LOCKED ON THE TISSUE, BUT THE SURGEON WAS ABLE TO REMOVE IT IN A NORMAL WAY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 11-01-2023 | ENDO GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN THORACOTOMY PLUS LOBECTOMY DUE TO PNEUMONIA, THE DEVICE STAPLED BUT DID NOT CUT. TO RESOLVE THE ISSUE, A COLD CUT USING SCISSORS WAS MADE. THERE WAS NO PATIENT INJURY.
24-04-2023 | Malfunction | covidien lp llc north haven | 11-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, WHEN THE SURGEON PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-12-2022 | Death | intuitive surgical | 11-01-2023 | ENDOWRIST | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY PROCEDURE, THE CUSTOMER UNDERWENT A REVISION SURGERY ON POST-OPERATIVE DAY (POD) #3 VIA OPEN TECHNIQUE BY THE SAME SURGEON. THE PATIENT DIED ON POD #10 DUE TO PERITONITIS. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE SURGEON AND OBTAINED THE FOLLOWING INFORMATION: THE PROCEDURE WAS PERFORMED DUE TO ADENOCARCINOMA OF THE "ASCENDING COLON T2N0MO, STAGE 2." THE PROCEDURE LASTED ABOUT 3 HOURS WITH NO INTRAOPERATIVE COMPLICATIONS, MINOR BLEEDING ABOUT 10-15 ML, AND ENDED WITH AN INTRACORPOREAL ANASTOMOSIS. THE TERMINAL ILEUM WAS DIVIDED BY AN ENDOWRIST STAPLER 45 INSTRUMENT, USING ONE BLUE STAPLER 45 RELOAD. THE TRANSVERSE COLON WAS DIVIDED BY A ENDOWRIST STAPLER 45 INSTRUMENT USING TWO BLUE RELOADS. THE ANASTOMOSIS BETWEEN THE ILEUM AND TRANSVERSE COLON WAS DONE USING ONE BLUE RELOAD. ON POD #3, THE PATIENT WAS DIAGNOSED WITH POST-OPERATIVE PERITONITIS AND AN EXPLORATORY LAPAROTOMY WAS DONE. TOTAL PERITONITIS WAS FOUND, AND THE SOURCE OF INFECTION WAS LEAKAGE FROM THE STUMP OF THE SMALL INTESTINE WITH A DIAMETER OF 3 MM, AND VITAL MUCOSA WAS SEEN. THE INTESTINAL CONTENTS LEAKED THROUGH THE STAPLES. A RESECTION OF THE STUMP WAS PERFORMED, AN ILEOSTOMY PROXIMAL TO THE ANASTOMOSIS WAS DONE, AND THE ABDOMEN WAS LEFT OPEN FOR FURTHER REVISION. AFTER SEVERAL REVISIONS OF THE ABDOMEN AND ICU TREATMENT, SEPSIS LED TO MULTIORGAN FAILURE AND THE DEATH OF THE PATIENT. THE SURGEON BELIEVES THAT THE CAUSE OF THE COMPLICATION WAS DUE TO A STAPLER LINE FAILURE. NO IMAGES/VIDEO ARE AVAILABLE FOR VIEWING.
02-12-2022 | Death | intuitive surgical | 11-01-2023 | ENDOWRIST | Failure to Form Staple | A050704 | Sepsis | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY PROCEDURE, THE CUSTOMER UNDERWENT A REVISION SURGERY ON POST-OPERATIVE DAY (POD) #3 VIA OPEN TECHNIQUE BY THE SAME SURGEON. THE PATIENT DIED ON POD #10 DUE TO PERITONITIS. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE SURGEON AND OBTAINED THE FOLLOWING INFORMATION: THE PROCEDURE WAS PERFORMED DUE TO ADENOCARCINOMA OF THE "ASCENDING COLON T2N0MO, STAGE 2." THE PROCEDURE LASTED ABOUT 3 HOURS WITH NO INTRAOPERATIVE COMPLICATIONS, MINOR BLEEDING ABOUT 10-15 ML, AND ENDED WITH AN INTRACORPOREAL ANASTOMOSIS. THE TERMINAL ILEUM WAS DIVIDED BY AN ENDOWRIST STAPLER 45 INSTRUMENT, USING ONE BLUE STAPLER 45 RELOAD. THE TRANSVERSE COLON WAS DIVIDED BY A ENDOWRIST STAPLER 45 INSTRUMENT USING TWO BLUE RELOADS. THE ANASTOMOSIS BETWEEN THE ILEUM AND TRANSVERSE COLON WAS DONE USING ONE BLUE RELOAD. ON POD #3, THE PATIENT WAS DIAGNOSED WITH POST-OPERATIVE PERITONITIS AND AN EXPLORATORY LAPAROTOMY WAS DONE. TOTAL PERITONITIS WAS FOUND, AND THE SOURCE OF INFECTION WAS LEAKAGE FROM THE STUMP OF THE SMALL INTESTINE WITH A DIAMETER OF 3 MM, AND VITAL MUCOSA WAS SEEN. THE INTESTINAL CONTENTS LEAKED THROUGH THE STAPLES. A RESECTION OF THE STUMP WAS PERFORMED, AN ILEOSTOMY PROXIMAL TO THE ANASTOMOSIS WAS DONE, AND THE ABDOMEN WAS LEFT OPEN FOR FURTHER REVISION. AFTER SEVERAL REVISIONS OF THE ABDOMEN AND ICU TREATMENT, SEPSIS LED TO MULTIORGAN FAILURE AND THE DEATH OF THE PATIENT. THE SURGEON BELIEVES THAT THE CAUSE OF THE COMPLICATION WAS DUE TO A STAPLER LINE FAILURE. NO IMAGES/VIDEO ARE AVAILABLE FOR VIEWING.
02-12-2022 | Death | intuitive surgical | 11-01-2023 | ENDOWRIST | Failure to Form Staple | A050704 | Peritonitis | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY PROCEDURE, THE CUSTOMER UNDERWENT A REVISION SURGERY ON POST-OPERATIVE DAY (POD) #3 VIA OPEN TECHNIQUE BY THE SAME SURGEON. THE PATIENT DIED ON POD #10 DUE TO PERITONITIS. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE SURGEON AND OBTAINED THE FOLLOWING INFORMATION: THE PROCEDURE WAS PERFORMED DUE TO ADENOCARCINOMA OF THE "ASCENDING COLON T2N0MO, STAGE 2." THE PROCEDURE LASTED ABOUT 3 HOURS WITH NO INTRAOPERATIVE COMPLICATIONS, MINOR BLEEDING ABOUT 10-15 ML, AND ENDED WITH AN INTRACORPOREAL ANASTOMOSIS. THE TERMINAL ILEUM WAS DIVIDED BY AN ENDOWRIST STAPLER 45 INSTRUMENT, USING ONE BLUE STAPLER 45 RELOAD. THE TRANSVERSE COLON WAS DIVIDED BY A ENDOWRIST STAPLER 45 INSTRUMENT USING TWO BLUE RELOADS. THE ANASTOMOSIS BETWEEN THE ILEUM AND TRANSVERSE COLON WAS DONE USING ONE BLUE RELOAD. ON POD #3, THE PATIENT WAS DIAGNOSED WITH POST-OPERATIVE PERITONITIS AND AN EXPLORATORY LAPAROTOMY WAS DONE. TOTAL PERITONITIS WAS FOUND, AND THE SOURCE OF INFECTION WAS LEAKAGE FROM THE STUMP OF THE SMALL INTESTINE WITH A DIAMETER OF 3 MM, AND VITAL MUCOSA WAS SEEN. THE INTESTINAL CONTENTS LEAKED THROUGH THE STAPLES. A RESECTION OF THE STUMP WAS PERFORMED, AN ILEOSTOMY PROXIMAL TO THE ANASTOMOSIS WAS DONE, AND THE ABDOMEN WAS LEFT OPEN FOR FURTHER REVISION. AFTER SEVERAL REVISIONS OF THE ABDOMEN AND ICU TREATMENT, SEPSIS LED TO MULTIORGAN FAILURE AND THE DEATH OF THE PATIENT. THE SURGEON BELIEVES THAT THE CAUSE OF THE COMPLICATION WAS DUE TO A STAPLER LINE FAILURE. NO IMAGES/VIDEO ARE AVAILABLE FOR VIEWING.
23-12-2022 | Injury | covidien lp llc north haven | 11-01-2023 | GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVER SEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 11-01-2023 | GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVER SEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 11-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVER SEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
23-12-2022 | Injury | covidien lp llc north haven | 11-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY OF AN OPEN LAPAROTOMY BOWEL RESECTION WHERE THE DEVICE WAS USED FOR RESECTION AND ANASTOMOSIS, THE RESECTED BOWEL STARTED OOZING. THE PATIENT WAS RETURNED TO THEATRE AND THE SURGEON OVER SEWED THE STAPLE LINE WITH SUTURES. THE PATIENT¿S HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
13-12-2022 | Malfunction | covidien lp llc north haven | 11-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC NEPHRECTOMY, THE SURGEON APPROACHED THE BLOOD VESSEL, CLOSED THE JAWS, AND PRESSED THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 11-01-2023 | ENDO GIA | Failure to Cycle | A050703 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, THE RELOAD WAS FIRED COMPLETELY, BUT THE STAPLES ON THE STOMACH TISSUE PILED UP, THE BLADE DID NOT COMPLETE THE CUT, PUSHING THE TISSUE FORWARD. THE USER RELEASED THE BLACK KNOB AND A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
20-12-2022 | Injury | us surgical puerto rico | 11-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC (VATS) WEDGE AT LOWER LEFT LOBE, THE DEVICE PARTIALLY FIRED AND STOPPED AT 30MM, THE DEVICE LOCKED ON TISSUE AND CANNOT START THE SECOND FIRING. STAPLE LINE WAS FORMED. THE SURGEON HAD TO TEAR IT OFF TO REMOVE THE DEVICE. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE.
20-12-2022 | Injury | us surgical puerto rico | 11-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC (VATS) WEDGE AT LOWER LEFT LOBE, THE DEVICE PARTIALLY FIRED AND STOPPED AT 30MM, THE DEVICE LOCKED ON TISSUE AND CANNOT START THE SECOND FIRING. STAPLE LINE WAS FORMED. THE SURGEON HAD TO TEAR IT OFF TO REMOVE THE DEVICE. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE.
09-03-2023 | Injury | ethicon endo-surgery | 11-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST-OP A GASTRIC SLEEVE PROCEDURE THAT TWO PATIENTS PRESENTED WITH STAPLE LINE BREAK DOWN. ONE PATIENT WAS FIVE WEEKS POST-OP AND THE SECOND PATIENT WAS TWO WEEKS POST-OP. REPAIR OF THE STAPLE LINE BREAK DOWN WAS NOT PROVIDED TO THE SALES REP. PATIENTS¿ STATUS WAS NOT PROVIDED.
09-03-2023 | Injury | ethicon endo-surgery | 11-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST-OP A GASTRIC SLEEVE PROCEDURE THAT TWO PATIENTS PRESENTED WITH STAPLE LINE BREAK DOWN. ONE PATIENT WAS FIVE WEEKS POST-OP AND THE SECOND PATIENT WAS TWO WEEKS POST-OP. REPAIR OF THE STAPLE LINE BREAK DOWN WAS NOT PROVIDED TO THE SALES REP. PATIENTS¿ STATUS WAS NOT PROVIDED.
11-01-2023 | Injury | ethicon endo-surgery | 11-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST-OP A GASTRIC SLEEVE PROCEDURE THAT TWO PATIENTS PRESENTED WITH STAPLE LINE BREAK DOWN. ONE PATIENT WAS FIVE WEEKS POST-OP AND THE SECOND PATIENT WAS TWO WEEKS POST-OP. REPAIR OF THE STAPLE LINE BREAK DOWN WAS NOT PROVIDED TO THE SALES REP. PATIENTS CURRENT STATUS WAS NOT PROVIDED.
11-01-2023 | Injury | ethicon endo-surgery | 11-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST-OP A GASTRIC SLEEVE PROCEDURE THAT TWO PATIENTS PRESENTED WITH STAPLE LINE BREAK DOWN. ONE PATIENT WAS FIVE WEEKS POST-OP AND THE SECOND PATIENT WAS TWO WEEKS POST-OP. REPAIR OF THE STAPLE LINE BREAK DOWN WAS NOT PROVIDED TO THE SALES REP. PATIENTS CURRENT STATUS WAS NOT PROVIDED.
01-12-2022 | Injury | ethicon endo-surgery | 11-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Patient Device Interaction Problem | A01 | Insufficient Information | Event Description: EVENT DAY WAS IN (B)(6) 2022. SURGEON UNSURE WHETHER IT WAS OUR PRODUCT OR PATIENT FACTORS BUT HIS PATIENT REQUIRED READMISSION SOME DAYS POST OP. SURGEON REFUSED TO PROVIDE ANY MORE DETAILS SUCH AS DETAILS OF EVENT OR DATE DUE TO PATIENT CONFIDENTIALITY. DID THE PATIENT EXPERIENCE AN ADVERSE EVENT SUCH AS INFECTION, NON-UNION, ALLERGIC REACTION, OSTEOPOROSIS, OVERLOADING, PAIN, DEGENERATIVE DISEASES, BLEEDING OR OOZING? UNKNOWN, DID THE PATIENT REQUIRE REVISION SURGERY OR HARDWARE REMOVAL? UNKNOWN, IF NO, WAS THERE ANY ADDITIONAL MEDICAL INTERVENTION REQUIRED SUCH AS X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS? UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES YES, PATIENT CONSEQUENCE DESCRIPTION/WAS THERE A CLINICAL OUTCOME EXPERIENCED BY THE PATIENT (INFECTION, INFLAMMATION, ETC.)? SURGEON WOULD NOT PROVIDE ADDITIONAL INFORMATION, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN, IF YES, DESCRIBE UNKNOWN EXCEPT READMISSION WAS REQUIRED, IS THE PATIENT PART OF A CLINICAL STUDY UNKNOWN, (B)(4). DEVICE PROPERTY OF HOSPITAL, DEVICE IN POSSESSION OF NONE, (B)(4). DEVICE PROPERTY OF NONE, DEVICE IN POSSESSION OF NONE BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST. TRUE
14-12-2022 | Injury | intuitive surgical | 11-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT FIRED A BLUE SUREFORM 60 RELOAD AND PUSHED/BUNCHED THE TARGET TISSUE INSTEAD OF CUTTING IT. THE CUSTOMER REPLACED THIS STAPLER INSTRUMENT WITH A SECOND SUREFORM 60 STAPLER INSTRUMENT WHICH EXPERIENCED AN ENGAGEMENT ISSUE. THE CUSTOMER THEN USED A THIRD SUREFORM 60 STAPLER INSTRUMENT WHICH REPORTEDLY WORKED AS EXPECTED AND THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE SURGEON REPORTED THAT THE INDICATION FOR THIS PROCEDURE WAS SEVERE OBESITY. THE SUREFORM 60 STAPLER INSTRUMENT WAS FIRING A BLUE SUREFORM 60 RELOAD TO CREATE THE SLEEVE ON THE STOMACH; THIS WAS THE STAPLER INSTRUMENT¿S FOURTH FIRE OF THE PROCEDURE. THE SURGEON REPORTED THERE WERE NO CLAMPING ISSUES, ERRORS, OR UNEXPECTED BEHAVIORS FROM THE STAPLER RELATED TO THIS FIRE UNTIL IT PUSHED THE TARGET TISSUE DURING THE FIRE SEQUENCE. THE SURGEON REPORTED THAT THIS TISSUE PUSH EVENT CAUSED A HEMATOMA AND ABOUT 5CC OF BLOOD LOSS. THE SURGEON SAID THIS CAUSED ABOUT A 30 MINUTE DELAY DURING A CRITICAL PART OF THE PROCEDURE. TO RESOLVE THIS EVENT, THE SURGEON MADE A TIGHTER SLEEVE AND RESECTED THE HEMATOMA. THE SURGEON SAID THIS SLEEVE WAS ¿TOO TIGHT¿ AS THEY NORMALLY DO NOT GO THIS TIGHT. HOWEVER, THE SURGEON ADDED THAT THE PATIENT DID NOT HAVE ANY POST-OPERATIVE COMPLICATIONS AND WAS DISCHARGED IN GOOD HEALTH. THE PATIENT CARE PLAN WAS NOT CHANGED DUE TO THIS EVENT. THE SURGEON IS CONCERNED THE PATIENT MAY DEVELOP FOOD INTOLERANCE OR ISSUES SWALLOWING DUE TO THE TIGHT SLEEVE. THE SURGEON ADDED THAT THEY DO NOT KNOW WHY THIS TISSUE PUSH EVENT WOULD HAVE OCCURRED.
14-12-2022 | Injury | intuitive surgical | 11-01-2023 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT FIRED A BLUE SUREFORM 60 RELOAD AND PUSHED/BUNCHED THE TARGET TISSUE INSTEAD OF CUTTING IT. THE CUSTOMER REPLACED THIS STAPLER INSTRUMENT WITH A SECOND SUREFORM 60 STAPLER INSTRUMENT WHICH EXPERIENCED AN ENGAGEMENT ISSUE. THE CUSTOMER THEN USED A THIRD SUREFORM 60 STAPLER INSTRUMENT WHICH REPORTEDLY WORKED AS EXPECTED AND THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE SURGEON REPORTED THAT THE INDICATION FOR THIS PROCEDURE WAS SEVERE OBESITY. THE SUREFORM 60 STAPLER INSTRUMENT WAS FIRING A BLUE SUREFORM 60 RELOAD TO CREATE THE SLEEVE ON THE STOMACH; THIS WAS THE STAPLER INSTRUMENT¿S FOURTH FIRE OF THE PROCEDURE. THE SURGEON REPORTED THERE WERE NO CLAMPING ISSUES, ERRORS, OR UNEXPECTED BEHAVIORS FROM THE STAPLER RELATED TO THIS FIRE UNTIL IT PUSHED THE TARGET TISSUE DURING THE FIRE SEQUENCE. THE SURGEON REPORTED THAT THIS TISSUE PUSH EVENT CAUSED A HEMATOMA AND ABOUT 5CC OF BLOOD LOSS. THE SURGEON SAID THIS CAUSED ABOUT A 30 MINUTE DELAY DURING A CRITICAL PART OF THE PROCEDURE. TO RESOLVE THIS EVENT, THE SURGEON MADE A TIGHTER SLEEVE AND RESECTED THE HEMATOMA. THE SURGEON SAID THIS SLEEVE WAS ¿TOO TIGHT¿ AS THEY NORMALLY DO NOT GO THIS TIGHT. HOWEVER, THE SURGEON ADDED THAT THE PATIENT DID NOT HAVE ANY POST-OPERATIVE COMPLICATIONS AND WAS DISCHARGED IN GOOD HEALTH. THE PATIENT CARE PLAN WAS NOT CHANGED DUE TO THIS EVENT. THE SURGEON IS CONCERNED THE PATIENT MAY DEVELOP FOOD INTOLERANCE OR ISSUES SWALLOWING DUE TO THE TIGHT SLEEVE. THE SURGEON ADDED THAT THEY DO NOT KNOW WHY THIS TISSUE PUSH EVENT WOULD HAVE OCCURRED.
27-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOBECTOMY, THE STAPLER WAS USED TO STAPLE AND TRANSECT A VESSEL, HOWEVER, WHEN FIRING, THE HANDLE SNAPPED, WHICH WAS VERY STRANGE SINCE THE TISSUE WAS REALLY THIN AND THE APPROPRIATE RELOAD WAS BEING USED. THIS LED TO BREAKING OF SOME KIND OF INTERNAL COMPONENT. A NEW HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | UNKNOWN GIA DST PRODUCT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. A DIFFERENT PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PARTIAL PULMONARY RESECTION, DURING A PULMONARY PARENCHYMA RESECTION, THE FIRING STOPPED MIDWAY ON BOTH RELOADS, AND THE TIP OF THE STAPLE THAT STOPPED WAS IN U-SHAPE. A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER WHERE THE SURGEON CONTINUED THE SECOND FIRING WHERE THE FIRST FIRING STOPPED. NO PATIENT INJURY.
23-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PARTIAL PULMONARY RESECTION, DURING A PULMONARY PARENCHYMA RESECTION, THE FIRING STOPPED MIDWAY ON BOTH RELOADS, AND THE TIP OF THE STAPLE THAT STOPPED WAS IN U-SHAPE. A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER WHERE THE SURGEON CONTINUED THE SECOND FIRING WHERE THE FIRST FIRING STOPPED. NO PATIENT INJURY.
23-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PARTIAL PULMONARY RESECTION, DURING A PULMONARY PARENCHYMA RESECTION, THE FIRING STOPPED MIDWAY ON BOTH RELOADS, AND THE TIP OF THE STAPLE THAT STOPPED WAS IN U-SHAPE. A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER WHERE THE SURGEON CONTINUED THE SECOND FIRING WHERE THE FIRST FIRING STOPPED. NO PATIENT INJURY.
23-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PARTIAL PULMONARY RESECTION, DURING A PULMONARY PARENCHYMA RESECTION, THE FIRING STOPPED MIDWAY ON BOTH RELOADS, AND THE TIP OF THE STAPLE THAT STOPPED WAS IN U-SHAPE. A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER WHERE THE SURGEON CONTINUED THE SECOND FIRING WHERE THE FIRST FIRING STOPPED. NO PATIENT INJURY.
21-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION AND SEGMENTECTOMY SURGERY, AT THE LUNG LOBE BOUNDARIES, THREE RELOADS DID NOT FIRE WHEN PERFORMING SEGMENTECTOMY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. IT WAS ALSO NOTED THAT WHEN THE SURGEON WAS FIRING THE FIRST SHOT, THERE WAS A CLICK SOUND. WEDGE RESECTION WAS A SUCCESS BUT THERE WAS AN AIR LEAK. SWITCHED TO ANOTHER COMPANY'S PRODUCT FOR THE THIRD TIME TO RESOLVE THE ISSUE.
21-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION AND SEGMENTECTOMY SURGERY, AT THE LUNG LOBE BOUNDARIES, THREE RELOADS DID NOT FIRE WHEN PERFORMING SEGMENTECTOMY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. IT WAS ALSO NOTED THAT WHEN THE SURGEON WAS FIRING THE FIRST SHOT, THERE WAS A CLICK SOUND. WEDGE RESECTION WAS A SUCCESS BUT THERE WAS AN AIR LEAK. SWITCHED TO ANOTHER COMPANY'S PRODUCT FOR THE THIRD TIME TO RESOLVE THE ISSUE.
21-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION AND SEGMENTECTOMY SURGERY, AT THE LUNG LOBE BOUNDARIES, THREE RELOADS DID NOT FIRE WHEN PERFORMING SEGMENTECTOMY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. IT WAS ALSO NOTED THAT WHEN THE SURGEON WAS FIRING THE FIRST SHOT, THERE WAS A CLICK SOUND. WEDGE RESECTION WAS A SUCCESS BUT THERE WAS AN AIR LEAK. SWITCHED TO ANOTHER COMPANY'S PRODUCT FOR THE THIRD TIME TO RESOLVE THE ISSUE.
27-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, WHILE RESECTING GASTRIC BODY, THE FIRST RELOAD WAS LOADED INTO THE STAPLER, JAWS OF THE RELOAD WERE SET INTO THE GASTRIC BODY, THEN GREEN BUTTON WAS PRESSED, AND HANDLE WAS SQUEEZED, HOWEVER, THE HANDLE COULD NO LONGER BE SQUEEZED HALFWAY THROUGH. RELOAD WAS REPLACED WITH BLACK RELOAD. JAWS OF THE BLACK RELOAD WERE SET INTO THE GASTRIC BODY, THE GREEN BUTTON WAS THEN PRESSED, AND THE HANDLE WAS ATTEMPTED TO BE SQUEEZED, BUT SINCE THE HANDLE COULD NOT BE SQUEEZED, IT WAS REPLACED WITH A NEW PRODUCT THAT WAS IN THE HOSPITAL AND FIRING WAS DONE WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
27-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, WHILE RESECTING GASTRIC BODY, THE FIRST RELOAD WAS LOADED INTO THE STAPLER, JAWS OF THE RELOAD WERE SET INTO THE GASTRIC BODY, THEN GREEN BUTTON WAS PRESSED, AND HANDLE WAS SQUEEZED, HOWEVER, THE HANDLE COULD NO LONGER BE SQUEEZED HALFWAY THROUGH. RELOAD WAS REPLACED WITH BLACK RELOAD. JAWS OF THE BLACK RELOAD WERE SET INTO THE GASTRIC BODY, THE GREEN BUTTON WAS THEN PRESSED, AND THE HANDLE WAS ATTEMPTED TO BE SQUEEZED, BUT SINCE THE HANDLE COULD NOT BE SQUEEZED, IT WAS REPLACED WITH A NEW PRODUCT THAT WAS IN THE HOSPITAL AND FIRING WAS DONE WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
27-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, WHILE RESECTING GASTRIC BODY, THE FIRST RELOAD WAS LOADED INTO THE STAPLER, JAWS OF THE RELOAD WERE SET INTO THE GASTRIC BODY, THEN GREEN BUTTON WAS PRESSED, AND HANDLE WAS SQUEEZED, HOWEVER, THE HANDLE COULD NO LONGER BE SQUEEZED HALFWAY THROUGH. RELOAD WAS REPLACED WITH BLACK RELOAD. JAWS OF THE BLACK RELOAD WERE SET INTO THE GASTRIC BODY, THE GREEN BUTTON WAS THEN PRESSED, AND THE HANDLE WAS ATTEMPTED TO BE SQUEEZED, BUT SINCE THE HANDLE COULD NOT BE SQUEEZED, IT WAS REPLACED WITH A NEW PRODUCT THAT WAS IN THE HOSPITAL AND FIRING WAS DONE WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
26-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS, WHEN MAKING THE POUCH OF THE STOMACH, THE STAPLE LINE WAS INCOMPLETE DISTALLY. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE.
26-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS, WHEN MAKING THE POUCH OF THE STOMACH, THE STAPLE LINE WAS INCOMPLETE DISTALLY. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE.
21-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION AND SEGMENTECTOMY SURGERY, AT THE LUNG LOBE BOUNDARIES, THREE RELOADS DID NOT FIRE WHEN PERFORMING SEGMENTECTOMY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. IT WAS ALSO NOTED THAT WHEN THE SURGEON WAS FIRING THE FIRST SHOT, THERE WAS A CLICK SOUND. WEDGE RESECTION WAS A SUCCESS BUT THERE WAS AN AIR LEAK. SWITCHED TO ANOTHER COMPANY'S PRODUCT FOR THE THIRD TIME TO RESOLVE THE ISSUE.
21-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION AND SEGMENTECTOMY SURGERY, AT THE LUNG LOBE BOUNDARIES, THREE RELOADS DID NOT FIRE WHEN PERFORMING SEGMENTECTOMY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. IT WAS ALSO NOTED THAT WHEN THE SURGEON WAS FIRING THE FIRST SHOT, THERE WAS A CLICK SOUND. WEDGE RESECTION WAS A SUCCESS BUT THERE WAS AN AIR LEAK. SWITCHED TO ANOTHER COMPANY'S PRODUCT FOR THE THIRD TIME TO RESOLVE THE ISSUE.
14-11-2022 | Malfunction | ethicon endo-surgery | 12-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE HAD NO STAPLES IN IT. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
27-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE DEVICE DID NOT FIRE AND THE GREEN BUTTON WAS NOT PRESSED. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
28-09-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW RECTAL RESECTION, WHEN DETACHING THE RECTUM, AFTER THE HANDLE WAS FIRED, THE STAPLE COULD NOT BE PUSHED. THE SURGEON REPLACED ONE HANDLE AND ONE RELOAD, AND STILL COULD NOT PUSH THE STAPLE. IN ORDER TO RESOLVE THE ISSUE, THE SURGEON REPLACED AN ELECTRIC STAPLER FROM ANOTHER BRAND. IT WAS NOTED THAT THE SURGERY TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE REPLICATION OF THE SHEARED TEETH ON THE FIRING RACK MAY OCCUR.
28-09-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW RECTAL RESECTION, WHEN DETACHING THE RECTUM, AFTER THE HANDLE WAS FIRED, THE STAPLE COULD NOT BE PUSHED. THE SURGEON REPLACED ONE HANDLE AND ONE RELOAD, AND STILL COULD NOT PUSH THE STAPLE. IN ORDER TO RESOLVE THE ISSUE, THE SURGEON REPLACED AN ELECTRIC STAPLER FROM ANOTHER BRAND. IT WAS NOTED THAT THE SURGERY TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE REPLICATION OF THE SHEARED TEETH ON THE FIRING RACK MAY OCCUR.
28-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC INTESTINAL TUMOR RESECTION, WHEN RECONSTRUCTING DIGESTIVE TRACK, THE DEVICE DID NOT FIRE. WHEN THE INSTRUMENT WAS WITHDRAWN, THE ANVIL DID NOT TILT. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THE EVENT LED TO PROLONGED ANESTHESIA AND SURGICAL TIME (LESS THAN 30 MINUTES). THERE WAS NO PATIENT INJURY.
28-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC INTESTINAL TUMOR RESECTION, WHEN RECONSTRUCTING DIGESTIVE TRACK, THE DEVICE DID NOT FIRE. WHEN THE INSTRUMENT WAS WITHDRAWN, THE ANVIL DID NOT TILT. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THE EVENT LED TO PROLONGED ANESTHESIA AND SURGICAL TIME (LESS THAN 30 MINUTES). THERE WAS NO PATIENT INJURY.
28-09-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW RECTAL RESECTION, WHEN DETACHING THE RECTUM, AFTER THE HANDLE WAS FIRED, THE STAPLE COULD NOT BE PUSHED. THE SURGEON REPLACED ONE HANDLE AND ONE RELOAD, AND STILL COULD NOT PUSH THE STAPLE. IN ORDER TO RESOLVE THE ISSUE, THE SURGEON REPLACED AN ELECTRIC STAPLER FROM ANOTHER BRAND. IT WAS NOTED THAT THE SURGERY TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE REPLICATION OF THE SHEARED TEETH ON THE FIRING RACK MAY OCCUR.
28-09-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW RECTAL RESECTION, WHEN DETACHING THE RECTUM, AFTER THE HANDLE WAS FIRED, THE STAPLE COULD NOT BE PUSHED. THE SURGEON REPLACED ONE HANDLE AND ONE RELOAD, AND STILL COULD NOT PUSH THE STAPLE. IN ORDER TO RESOLVE THE ISSUE, THE SURGEON REPLACED AN ELECTRIC STAPLER FROM ANOTHER BRAND. IT WAS NOTED THAT THE SURGERY TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE REPLICATION OF THE SHEARED TEETH ON THE FIRING RACK MAY OCCUR.
25-08-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY PROCEDURE, WHILE STAPLING THE SIGMOID COLON, THE SURGEON WAS PLACING A CIRCULAR STAPLER ON THE PATIENT, WHEN IT WAS OBSERVED THAT THE STAPLES IN THE DISTAL PORTION OF THE STAPLE LINE WERE NOT DEPLOYED; THE STAPLE LINE WAS DISTALLY INCOMPLETE EVEN WHEN THE FIRING CYCLE OF THE RELOAD WAS COMPLETED. THE SURGEON SUTURED THE DISTAL PORTION TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
25-08-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY PROCEDURE, WHILE STAPLING THE SIGMOID COLON, THE SURGEON WAS PLACING A CIRCULAR STAPLER ON THE PATIENT, WHEN IT WAS OBSERVED THAT THE STAPLES IN THE DISTAL PORTION OF THE STAPLE LINE WERE NOT DEPLOYED; THE STAPLE LINE WAS DISTALLY INCOMPLETE EVEN WHEN THE FIRING CYCLE OF THE RELOAD WAS COMPLETED. THE SURGEON SUTURED THE DISTAL PORTION TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
23-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. A DIFFERENT PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESECTION, WHILE AT RESECTION FOR COLECTOMY, THE BODY OF THE STOMACH WAS PLACED INTO THE JAWS OF THIS CARTRIDGE, THE HANDLE WAS SQUEEZED, JAWS WERE CLOSED, GREEN BUTTON WAS PRESSED, BUT WHEN AN ATTEMPT WAS MADE TO SQUEEZE IT, THE HANDLE COULD NO LONGER BE SQUEEZED AFTER ABOUT 1CM OF FIRING. AS SUCH, IT WAS SQUEEZED STRONGLY, BUT THE SITUATION DID NOT IMPROVE; SO THE RELOAD WAS REPLACED AND WAS USED WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
01-05-2023 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY, AFTER FIRING THE FIRST RELOAD, THE HANDLE LOCKED UP AND DID NOT FIRE FULLY. THE HANDLE AND RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2023 | Malfunction | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY, AFTER FIRING THE FIRST RELOAD, THE HANDLE LOCKED UP AND DID NOT FIRE FULLY. THE HANDLE AND RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2022 | Injury | us surgical puerto rico | 12-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), DURING STAPLING OF THE VESSELS, THERE WAS BLEEDING FROM BOTH STAPLE LINES. THE SURGEON CLIPPED THE STAPLE LINES TO STOP THE BLEEDING.
27-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), DURING STAPLING OF THE VESSELS, THERE WAS BLEEDING FROM BOTH STAPLE LINES. THE SURGEON CLIPPED THE STAPLE LINES TO STOP THE BLEEDING.
26-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PEDIATRIC SURGERY, THE SURGEON WAS ABLE TO SQUEEZE THE HANDLE TO CLOSE THE JAWS AND THEN PRESSED THE GREEN BUTTON. AFTER PRESSING THE GREEN BUTTON, THE SURGEON FOUND OUT THAT THE JAWS WOULD NOT CLOSED COMPLETELY. A NEW HANDLE WITH A DIFFERENT LOT NUMBER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2022 | Injury | ethicon endo-surgery | 12-01-2023 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A KIDNEY TRANSPLANT DONOR PROCEDURE THAT THE FIRST FIRE ON THE RENAL ARTERY, IT BLED IMMEDIATELY AND THEY HAD TO OPEN THE PATIENT. CURRENTLY THE PATIENT IS STABLE ON THE TABLE. REP COULD NOT PROVIDE MORE INFORMATION AT THIS TIME.
27-12-2022 | Injury | ethicon endo-surgery | 12-01-2023 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A KIDNEY TRANSPLANT DONOR PROCEDURE THAT THE FIRST FIRE ON THE RENAL ARTERY, IT BLED IMMEDIATELY AND THEY HAD TO OPEN THE PATIENT. CURRENTLY THE PATIENT IS STABLE ON THE TABLE. REP COULD NOT PROVIDE MORE INFORMATION AT THIS TIME.
27-12-2022 | Injury | ethicon endo-surgery | 12-01-2023 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A KIDNEY TRANSPLANT DONOR PROCEDURE THAT THE FIRST FIRE ON THE RENAL ARTERY, IT BLED IMMEDIATELY AND THEY HAD TO OPEN THE PATIENT. CURRENTLY THE PATIENT IS STABLE ON THE TABLE. REP COULD NOT PROVIDE MORE INFORMATION AT THIS TIME.
27-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, WHILE RESECTING GASTRIC BODY, THE FIRST RELOAD WAS LOADED INTO THE STAPLER, JAWS OF THE RELOAD WERE SET INTO THE GASTRIC BODY, THEN GREEN BUTTON WAS PRESSED, AND HANDLE WAS SQUEEZED, HOWEVER, THE HANDLE COULD NO LONGER BE SQUEEZED HALFWAY THROUGH. RELOAD WAS REPLACED WITH BLACK RELOAD. JAWS OF THE BLACK RELOAD WERE SET INTO THE GASTRIC BODY, THE GREEN BUTTON WAS THEN PRESSED, AND THE HANDLE WAS ATTEMPTED TO BE SQUEEZED, BUT SINCE THE HANDLE COULD NOT BE SQUEEZED, IT WAS REPLACED WITH A NEW PRODUCT THAT WAS IN THE HOSPITAL AND FIRING WAS DONE WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
15-12-2022 | Injury | ethicon endo-surgery | 12-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT 1WEEK AFTER A LAPAROSCOPIC ANTERIOR RESECTION, ANASTOMOTIC LEAKAGE OCCURRED SOON AFTER THE LIQUID FOOD WAS STARTED. THE HYPERALIMENTATION WAS USED, AND THE PATIENT HAS BEEN FOLLOWED UP. IT DEPENDS ON THE PATIENT CONDITION, BUT THERE IS A POSSIBILITY THAT THE COLOSTOMY WILL BE PERFORMED. IN THE INITIAL OPERATION, THE RECTUM RA RB WAS CUT WITH PSEE60A AND GOLD CARTRIDGE FOR 2 TIMES, AND CDH25B WAS USED TO ANASTOMOSE. DST RECONSTRUCTION WAS PERFORMED. THE STAPLING HEIGHT WAS 1.8MM. THERE WAS NO PROBLEM TO THE DEVICE OR THE STAPLING DURING THE INITIAL OPERATION. THE DEVICE WAS USED BETWEEN COLON AND RECTUM. THE PATIENT IS 64 YEARS-OLD, MALE. 4:00PM 19TH DEC, THE PATIENT IS STILL HOSPITALIZED, AND HYPERALIMENTATION WAS USED AND HAS BEEN FOLLOWED UP. THE SURGEON¿S COMMENT: I DON¿T BELIEVE THAT THE PROBLEM WAS 100% RELATED TO THE INSTRUMENT. THE POSSIBLE CAUSES OTHER THAN PRODUCT IS THAT BECAUSE OF DR. O. WHO IS COMING FOR GUIDANCE, THEY CHOSE GOLD CARTRIDGE FOR THE RECTUM, AND USED PSEE60A TO CUT AS MUCH TISSUE AS THE DEVICE CAN CUT, AND THEN ADDITIONAL CUT WILL BE PERFORMED ON THE REST OF THE TARGET TISSUE. SO, THE STAPLES ARE EASILY STACKED ON THE LEFT SIDE. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2022 | Injury | ethicon endo-surgery | 12-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT 1WEEK AFTER A LAPAROSCOPIC ANTERIOR RESECTION, ANASTOMOTIC LEAKAGE OCCURRED SOON AFTER THE LIQUID FOOD WAS STARTED. THE HYPERALIMENTATION WAS USED, AND THE PATIENT HAS BEEN FOLLOWED UP. IT DEPENDS ON THE PATIENT CONDITION, BUT THERE IS A POSSIBILITY THAT THE COLOSTOMY WILL BE PERFORMED. IN THE INITIAL OPERATION, THE RECTUM RA RB WAS CUT WITH PSEE60A AND GOLD CARTRIDGE FOR 2 TIMES, AND CDH25B WAS USED TO ANASTOMOSE. DST RECONSTRUCTION WAS PERFORMED. THE STAPLING HEIGHT WAS 1.8MM. THERE WAS NO PROBLEM TO THE DEVICE OR THE STAPLING DURING THE INITIAL OPERATION. THE DEVICE WAS USED BETWEEN COLON AND RECTUM. THE PATIENT IS 64 YEARS-OLD, MALE. 4:00PM 19TH DEC, THE PATIENT IS STILL HOSPITALIZED, AND HYPERALIMENTATION WAS USED AND HAS BEEN FOLLOWED UP. THE SURGEON¿S COMMENT: I DON¿T BELIEVE THAT THE PROBLEM WAS 100% RELATED TO THE INSTRUMENT. THE POSSIBLE CAUSES OTHER THAN PRODUCT IS THAT BECAUSE OF DR. O. WHO IS COMING FOR GUIDANCE, THEY CHOSE GOLD CARTRIDGE FOR THE RECTUM, AND USED PSEE60A TO CUT AS MUCH TISSUE AS THE DEVICE CAN CUT, AND THEN ADDITIONAL CUT WILL BE PERFORMED ON THE REST OF THE TARGET TISSUE. SO, THE STAPLES ARE EASILY STACKED ON THE LEFT SIDE. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2022 | Malfunction | ethicon endo-surgery | 12-01-2023 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE RELOAD NO SHAPE. NO OTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCE REPORTED.
15-12-2022 | Injury | ethicon endo-surgery | 12-01-2023 | CIRCULAR MECH STD 25MM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT 1WEEK AFTER A LAPAROSCOPIC ANTERIOR RESECTION, ANASTOMOTIC LEAKAGE OCCURRED SOON AFTER THE LIQUID FOOD WAS STARTED. THE HYPERALIMENTATION WAS USED, AND THE PATIENT HAS BEEN FOLLOWED UP. IT DEPENDS ON THE PATIENT CONDITION, BUT THERE IS A POSSIBILITY THAT THE COLOSTOMY WILL BE PERFORMED. IN THE INITIAL OPERATION, THE RECTUM RA RB WAS CUT WITH PSEE60A AND GOLD CARTRIDGE FOR 2 TIMES, AND CDH25B WAS USED TO ANASTOMOSE. DST RECONSTRUCTION WAS PERFORMED. THE STAPLING HEIGHT WAS 1.8MM. THERE WAS NO PROBLEM TO THE DEVICE OR THE STAPLING DURING THE INITIAL OPERATION. THE DEVICE WAS USED BETWEEN COLON AND RECTUM. THE PATIENT IS 64 YEARS-OLD, MALE. 4:00PM 19TH DEC, THE PATIENT IS STILL HOSPITALIZED, AND HYPERALIMENTATION WAS USED AND HAS BEEN FOLLOWED UP. THE SURGEON¿S COMMENT: I DON¿T BELIEVE THAT THE PROBLEM WAS 100% RELATED TO THE INSTRUMENT. THE POSSIBLE CAUSES OTHER THAN PRODUCT IS THAT BECAUSE OF DR. O. WHO IS COMING FOR GUIDANCE, THEY CHOSE GOLD CARTRIDGE FOR THE RECTUM, AND USED PSEE60A TO CUT AS MUCH TISSUE AS THE DEVICE CAN CUT, AND THEN ADDITIONAL CUT WILL BE PERFORMED ON THE REST OF THE TARGET TISSUE. SO, THE STAPLES ARE EASILY STACKED ON THE LEFT SIDE. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2022 | Injury | ethicon endo-surgery | 12-01-2023 | CIRCULAR MECH STD 25MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT 1WEEK AFTER A LAPAROSCOPIC ANTERIOR RESECTION, ANASTOMOTIC LEAKAGE OCCURRED SOON AFTER THE LIQUID FOOD WAS STARTED. THE HYPERALIMENTATION WAS USED, AND THE PATIENT HAS BEEN FOLLOWED UP. IT DEPENDS ON THE PATIENT CONDITION, BUT THERE IS A POSSIBILITY THAT THE COLOSTOMY WILL BE PERFORMED. IN THE INITIAL OPERATION, THE RECTUM RA RB WAS CUT WITH PSEE60A AND GOLD CARTRIDGE FOR 2 TIMES, AND CDH25B WAS USED TO ANASTOMOSE. DST RECONSTRUCTION WAS PERFORMED. THE STAPLING HEIGHT WAS 1.8MM. THERE WAS NO PROBLEM TO THE DEVICE OR THE STAPLING DURING THE INITIAL OPERATION. THE DEVICE WAS USED BETWEEN COLON AND RECTUM. THE PATIENT IS 64 YEARS-OLD, MALE. 4:00PM 19TH DEC, THE PATIENT IS STILL HOSPITALIZED, AND HYPERALIMENTATION WAS USED AND HAS BEEN FOLLOWED UP. THE SURGEON¿S COMMENT: I DON¿T BELIEVE THAT THE PROBLEM WAS 100% RELATED TO THE INSTRUMENT. THE POSSIBLE CAUSES OTHER THAN PRODUCT IS THAT BECAUSE OF DR. O. WHO IS COMING FOR GUIDANCE, THEY CHOSE GOLD CARTRIDGE FOR THE RECTUM, AND USED PSEE60A TO CUT AS MUCH TISSUE AS THE DEVICE CAN CUT, AND THEN ADDITIONAL CUT WILL BE PERFORMED ON THE REST OF THE TARGET TISSUE. SO, THE STAPLES ARE EASILY STACKED ON THE LEFT SIDE. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
27-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOBECTOMY, THE STAPLER WAS BEING USED TO STAPLE AND TRANSECT A VESSEL, WHEN FIRING, THE HANDLE SNAPPED EVEN THOUGH THE TISSUE WAS REALLY THIN AND AN APPROPRIATE RELOAD WAS BEING USED. THEN, WHEN THE LUNG WAS STAPLED USING AN APPROPRIATE RELOAD FOR THE TISSUE, THE HANDLE SNAPPED AGAIN WHICH LED TO COMPLETELY BREAKING OF SOME KIND OF INTERNAL COMPONENT. USED ANOTHER HANDLE WHICH WORKED FINE FOR THE REMAINING FIRES. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
21-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC WEDGE RESECTION AND SEGMENTECTOMY SURGERY, AT THE LUNG LOBE BOUNDARIES, THREE RELOADS DID NOT FIRE WHEN PERFORMING SEGMENTECTOMY. THE SURGEON ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. IT WAS ALSO NOTED THAT WHEN THE SURGEON WAS FIRING THE FIRST SHOT, THERE WAS A CLICK SOUND. WEDGE RESECTION WAS A SUCCESS BUT THERE WAS AN AIR LEAK. SWITCHED TO ANOTHER COMPANY'S PRODUCT FOR THE THIRD TIME TO RESOLVE THE ISSUE.
23-11-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ILEOSTOMY REVERSAL, WHEN ANASTOMOSIS (SMALL BOWEL TO SMALL BOWEL) WAS PERFORMED, THE DEVICE FIRED, STAPLES DID NOT COMPLETELY CLOSED AND STAPLE LINE WAS INCOMPLETE DISTALLY. IT WAS NOTED THAT THE JAWS OPENED WHEN THE OPENING LEVER WAS OPENED AS NORMAL PROCEDURE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO HAND SUTURE.
23-11-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ILEOSTOMY REVERSAL, WHEN ANASTOMOSIS (SMALL BOWEL TO SMALL BOWEL) WAS PERFORMED, THE DEVICE FIRED, STAPLES DID NOT COMPLETELY CLOSED AND STAPLE LINE WAS INCOMPLETE DISTALLY. IT WAS NOTED THAT THE JAWS OPENED WHEN THE OPENING LEVER WAS OPENED AS NORMAL PROCEDURE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO HAND SUTURE.
07-12-2022 | Death | ethicon endo-surgery | 12-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED AN OPEN EXPLORATORY LAPAROTOMY THAT INCLUDED A COLON RESECTION AND END-TO-END ANASTOMOSIS WITH AN ETHICON POWERED CIRCULAR DEVICE. THE ANASTOMOSIS PASSED AN AIR LEAK TEST. THE PATIENT WENT HOME FROM THE HOSPITAL ON (B)(6). THE PATIENT WAS READMITTED TO THE HOSPITAL ON (B)(6) AND THE DECISION WAS MADE TO RE-OPERATE DUE TO THE LIKELIHOOD OF AN ANASTOMOTIC LEAK. DURING THE RE-OPERATION, IT WAS NOTED THAT THE ANASTOMOSIS HAD "FAILED" AND HAD FALLEN APART. DURING THE RE-OPERATION, THE PATIENT DIED.
07-12-2022 | Death | ethicon endo-surgery | 12-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED AN OPEN EXPLORATORY LAPAROTOMY THAT INCLUDED A COLON RESECTION AND END-TO-END ANASTOMOSIS WITH AN ETHICON POWERED CIRCULAR DEVICE. THE ANASTOMOSIS PASSED AN AIR LEAK TEST. THE PATIENT WENT HOME FROM THE HOSPITAL ON (B)(6). THE PATIENT WAS READMITTED TO THE HOSPITAL ON (B)(6) AND THE DECISION WAS MADE TO RE-OPERATE DUE TO THE LIKELIHOOD OF AN ANASTOMOTIC LEAK. DURING THE RE-OPERATION, IT WAS NOTED THAT THE ANASTOMOSIS HAD "FAILED" AND HAD FALLEN APART. DURING THE RE-OPERATION, THE PATIENT DIED.
07-12-2022 | Death | ethicon endo-surgery | 12-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED AN OPEN EXPLORATORY LAPAROTOMY THAT INCLUDED A COLON RESECTION AND END-TO-END ANASTOMOSIS WITH AN ETHICON POWERED CIRCULAR DEVICE. THE ANASTOMOSIS PASSED AN AIR LEAK TEST. THE PATIENT WENT HOME FROM THE HOSPITAL ON (B)(6). THE PATIENT WAS READMITTED TO THE HOSPITAL ON (B)(6) AND THE DECISION WAS MADE TO RE-OPERATE DUE TO THE LIKELIHOOD OF AN ANASTOMOTIC LEAK. DURING THE RE-OPERATION, IT WAS NOTED THAT THE ANASTOMOSIS HAD "FAILED" AND HAD FALLEN APART. DURING THE RE-OPERATION, THE PATIENT DIED.
07-12-2022 | Death | ethicon endo-surgery | 12-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Septic Shock | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED AN OPEN EXPLORATORY LAPAROTOMY THAT INCLUDED A COLON RESECTION AND END-TO-END ANASTOMOSIS WITH AN ETHICON POWERED CIRCULAR DEVICE. THE ANASTOMOSIS PASSED AN AIR LEAK TEST. THE PATIENT WENT HOME FROM THE HOSPITAL ON (B)(6). THE PATIENT WAS READMITTED TO THE HOSPITAL ON (B)(6) AND THE DECISION WAS MADE TO RE-OPERATE DUE TO THE LIKELIHOOD OF AN ANASTOMOTIC LEAK. DURING THE RE-OPERATION, IT WAS NOTED THAT THE ANASTOMOSIS HAD "FAILED" AND HAD FALLEN APART. DURING THE RE-OPERATION, THE PATIENT DIED.
07-12-2022 | Death | ethicon endo-surgery | 12-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Multiple Organ Dysfunction Syndrome | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED AN OPEN EXPLORATORY LAPAROTOMY THAT INCLUDED A COLON RESECTION AND END-TO-END ANASTOMOSIS WITH AN ETHICON POWERED CIRCULAR DEVICE. THE ANASTOMOSIS PASSED AN AIR LEAK TEST. THE PATIENT WENT HOME FROM THE HOSPITAL ON (B)(6). THE PATIENT WAS READMITTED TO THE HOSPITAL ON (B)(6) AND THE DECISION WAS MADE TO RE-OPERATE DUE TO THE LIKELIHOOD OF AN ANASTOMOTIC LEAK. DURING THE RE-OPERATION, IT WAS NOTED THAT THE ANASTOMOSIS HAD "FAILED" AND HAD FALLEN APART. DURING THE RE-OPERATION, THE PATIENT DIED.
20-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC (VATS) WEDGE AT LOWER LEFT LOBE, THE DEVICE PARTIALLY FIRED AND STOPPED AT 30MM, THE DEVICE LOCKED ON TISSUE AND CANNOT START THE SECOND FIRING. STAPLE LINE WAS FORMED. THE SURGEON HAD TO TEAR IT OFF TO REMOVE THE DEVICE. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE.
20-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC (VATS) WEDGE AT LOWER LEFT LOBE, THE DEVICE PARTIALLY FIRED AND STOPPED AT 30MM, THE DEVICE LOCKED ON TISSUE AND CANNOT START THE SECOND FIRING. STAPLE LINE WAS FORMED. THE SURGEON HAD TO TEAR IT OFF TO REMOVE THE DEVICE. A NEW HANDLE AND RELOAD WAS USED TO RESOLVE THE ISSUE.
27-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ESOPHAGEAL PROCEDURE, WHEN MAKING TUBE FROM STOMACH, THERE WAS A POOR STAPLE FORMATION AFTER FIRING. A NEW DEVICE WAS USED TO OVER SEW THE STAPLE LINE AND REMOVE MALFORMED STAPLE PINS.
27-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | GIA | Failure to Form Staple | A050704 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ESOPHAGEAL PROCEDURE, WHEN MAKING TUBE FROM STOMACH, THERE WAS A POOR STAPLE FORMATION AFTER FIRING. A NEW DEVICE WAS USED TO OVER SEW THE STAPLE LINE AND REMOVE MALFORMED STAPLE PINS.
21-11-2022 | Injury | intuitive surgical | 12-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
21-11-2022 | Injury | intuitive surgical | 12-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED RIGHT HEMICOLECTOMY PROCEDURE, THE SUREFORM 60 STAPLE LINE HAD A LEAK. BEFORE THIS STAPLE RELOAD WAS FIRED, THERE WAS AN ENGAGEMENT ERROR AND THE SURGEON MANUALLY SELECTED THE COLOR RELOAD BEING FIRED (THE COLOR OF THE RELOAD WAS NOT PROVIDED). THE PATIENT RETURNED TO THE HOSPITAL POST-OPERATIVELY DUE TO THE STAPLE LINE LEAK AND THE PATIENT WAS HOSPITALIZED. AN INTUITIVE SURGICAL INC. (ISI) CLINICAL SALES ASSOCIATE (CSA) WAS PRESENT FOR THIS RIGHT HEMICOLECTOMY PROCEDURE, BUT THERE WERE NO INTRA-OPERATIVE COMPLICATIONS. THE CSA WAS INFORMED OF THE PATIENT'S POST-OPERATIVE STAPLE LINE LEAK BY THE SURGEON ON A LATER DATE. THE PATIENT RECEIVED A SECONDARY PROCEDURE TO RESOLVE THE LEAKS. THE SURGEON REPORTED THAT THEY PERFORMED A THREE STAPLE LINE ANASTOMOSIS WITH ALL BLUE SUREFORM 60 RELOADS, AND THIS STAPLE ANASTOMOSIS HAD COME UNDONE POST-OPERATIVELY. THE PATIENT'S CURRENT STATUS IS UNKNOWN. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION. HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
27-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ESOPHAGEAL PROCEDURE, WHEN MAKING TUBE FROM STOMACH, THERE WAS A POOR STAPLE FORMATION AFTER FIRING. A NEW DEVICE WAS USED TO OVER SEW THE STAPLE LINE AND REMOVE MALFORMED STAPLE PINS.
27-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | GIA | Failure to Form Staple | A050704 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN ESOPHAGEAL PROCEDURE, WHEN MAKING TUBE FROM STOMACH, THERE WAS A POOR STAPLE FORMATION AFTER FIRING. A NEW DEVICE WAS USED TO OVER SEW THE STAPLE LINE AND REMOVE MALFORMED STAPLE PINS.
21-11-2022 | Injury | intuitive surgical | 12-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED RIGHT HEMICOLECTOMY PROCEDURE, THE SUREFORM 60 STAPLE LINE HAD A LEAK. BEFORE THIS STAPLE RELOAD WAS FIRED, THERE WAS AN ENGAGEMENT ERROR AND THE SURGEON MANUALLY SELECTED THE COLOR RELOAD BEING FIRED (THE COLOR OF THE RELOAD WAS NOT PROVIDED). THE PATIENT RETURNED TO THE HOSPITAL POST-OPERATIVELY DUE TO THE STAPLE LINE LEAK AND THE PATIENT WAS HOSPITALIZED. AN INTUITIVE SURGICAL INC. (ISI) CLINICAL SALES ASSOCIATE (CSA) WAS PRESENT FOR THIS RIGHT HEMICOLECTOMY PROCEDURE, BUT THERE WERE NO INTRA-OPERATIVE COMPLICATIONS. THE CSA WAS INFORMED OF THE PATIENT'S POST-OPERATIVE STAPLE LINE LEAK BY THE SURGEON ON A LATER DATE. THE PATIENT RECEIVED A SECONDARY PROCEDURE TO RESOLVE THE LEAKS. THE SURGEON REPORTED THAT THEY PERFORMED A THREE STAPLE LINE ANASTOMOSIS WITH ALL BLUE SUREFORM 60 RELOADS, AND THIS STAPLE ANASTOMOSIS HAD COME UNDONE POST-OPERATIVELY. THE PATIENT'S CURRENT STATUS IS UNKNOWN. ISI HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION. HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
26-10-2022 | Malfunction | intuitive surgical | 12-01-2023 | SUREFORM | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
13-09-2022 | Injury | intuitive surgical | 12-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT AFTER COMPLETION OF A DA VINCI-ASSISTED LOW ANTERIOR RESECTION (LAR) PROCEDURE, A STAPLE LINE LEAK OCCURRED. IT IS UNKNOWN HOW THE POST-OPERATIVE LEAK WAS DISCOVERED. THE STAPLE LINE WAS REPORTEDLY FORMED WITH THE SUREFORM 60 AND AN UNKNOWN COLORED RELOAD. IT IS ALSO UNKNOWN HOW THE LEAK WAS RESOLVED. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION. HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
13-09-2022 | Injury | intuitive surgical | 12-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT AFTER COMPLETION OF A DA VINCI-ASSISTED LOW ANTERIOR RESECTION (LAR) PROCEDURE, A STAPLE LINE LEAK OCCURRED. IT IS UNKNOWN HOW THE POST-OPERATIVE LEAK WAS DISCOVERED. THE STAPLE LINE WAS REPORTEDLY FORMED WITH THE SUREFORM 60 AND AN UNKNOWN COLORED RELOAD. IT IS ALSO UNKNOWN HOW THE LEAK WAS RESOLVED. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION. HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
01-09-2022 | Injury | intuitive surgical | 12-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY PROCEDURE, A STAPLE LINE LEAK OCCURRED IN RELATION TO AN ANASTOMOSIS. THE STAPLE LINE ALLEGEDLY FELL APART POST-OPERATIVELY. THE STAPLE LINE WAS REPORTEDLY FORMED WITH THE SUREFORM 60 AND AN UNKNOWN COLORED SUREFORM 60 RELOAD. IT IS UNKNOWN HOW THE LEAK WAS RESOLVED. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION. HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
13-12-2022 | Injury | intuitive surgical | 12-01-2023 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS (ROUX-EN-Y) SURGICAL PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT, LOADED WITH A BLUE SUREFORM 60 RELOAD, DID NOT CUT THE TARGET TISSUE DURING THE FIRING SEQUENCE. WHILE THE STAPLES WERE LAID DOWN AS INTENDED, BECAUSE THE BLADE DID NOT CUT AS INTENDED, THE TISSUE WAS PUSHED FORWARD, PULLING OUT STAPLES AND CAUSING TEARS IN THE SURROUNDING TISSUE. THE SURGEON CLOSED THE TORN AREAS WITH SUTURES AND RAN AN ADDITIONAL STAPLE LINE, RESULTING IN THE CREATION OF A GASTRIC POUCH THAT WAS 20% SMALLER THAN INTENDED. THE PROCEDURE WAS COMPLETED, WITH A REPORT OF PATIENT INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE SURGEON AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: HE CONFIRMED THAT THE ISSUE OCCURRED DURING THE THIRD BLUE RELOAD FIRE (THE FIFTH FIRE OF THE PROCEDURE). THE FIRST FOUR FIRES WERE COMPLETED NORMALLY. PRIOR TO THE FIFTH FIRE, THE STAPLER CLAMPED NORMALLY, AND "THE FIRST HALF OF THE FIRING WAS FINE." BUT, THEN HE COULD SEE THE KNIFE PUSHING THE TISSUE, RATHER THAN CUTTING IT, AS IF "THE BLADE GOT OUT OF POSITION." THE PUSHED TISSUE "PULLED THE STAPLES OUT OF THE JAWS," AND WHEN HE OPENED THE JAWS, HE SAW THAT "THE DISTAL HALF OF THE STAPLE LINE, BETWEEN THE STOMACH AND THE POUCH, WASN'T CUT, IT WAS RIPPED." THE MALFORMED STAPLES WERE SUCCESSFULLY REMOVED FROM THE PATIENT, AND NO FRAGMENTS FELL. THERE WAS NO UNEXPECTED BLEEDING DUE TO THE INCIDENT. HE CONFIRMED THAT THIS OCCURRED ON THE FIRST STAPLE LINE, WHILE SEPARATING THE STOMACH FROM THE GASTRIC POUCH, AND THAT HE UTILIZED SUTURES TO FIX THE LATERAL EDGE OF THE POUCH BEFORE FIRING ANOTHER LINE OF STAPLES MEDIALLY TO THE FAILED STAPLE LINE. PER THE SURGEON, THE PATIENT "IS DOING FINE," AND SHE DID NOT EXPERIENCE ANY COMPLICATIONS, EITHER DURING THE SURGERY OR AFTER. THE SURGEON DOES NOT EXPECT HER TO HAVE FUTURE COMPLICATIONS DUE TO THE SMALLER-THAN-INTENDED SIZE OF THE GASTRIC POUCH. THE SURGEON DID NOT KNOW IF THE STAPLER INSTRUMENT WAS INSPECTED PRIOR TO THE PROCEDURE, BUT NO ISSUES WERE REPORTED BY THE STAFF. THE PROCEDURE WAS ROUTINE; THE TISSUE WAS ALL NORMAL, WITH NO CALCIFICATIONS, BUNCHING, TENSION, OR EXPOSURE TO RADIATION OR CHEMOTHERAPY. PER THE SURGEON, IT IS RELATIVELY COMMON IN GASTROINTESTINAL SURGERIES TO "SOMETIMES CROSS OVER THE STAPLE LINE." HE DID NOT REMEMBER IF THE STAPLER FIRED OVER PART OF THE PRECEDING STAPLE LINE AT THE TIME OF THE EVENT.
28-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, BEFORE USE ON THE LUNGS, THE DISTAL BLACK TIP OF THE TWO HANDLES BROKE OFF. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 12-01-2023 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, BEFORE USE ON THE LUNGS, THE DISTAL BLACK TIP OF THE TWO HANDLES BROKE OFF. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
24-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDICITIS SURGERY, WHEN THE FIRST STAPLE WAS DRIVEN OUT, THE STAPLE LINE DID NOT FORM INTO COMPLETE B-SHAPE, STAPLE LINE WAS INCOMPLETE PROXIMALLY, AND HANDLE WAS NO LONGER USABLE. THE DOCTOR NEEDED TO PULL OUT ONE BY ONE. THE STAPLE LINE WAS FIXED BY SUTURING. IT WAS NOTED THAT THE FEMALE PATIENT WAS THIN, SO THE TISSUE WAS NOT THICK. AFTER REINSTALLING THE NEW HANDLE AND RELOAD, ISSUE WAS RESOLVE WITH NO PROBLEM.
28-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, FOLLOWING AN OPEN COLON RESECTION WITH COLOSTOMY, FOR RECTAL STUMP TRANSECTION, THE PATIENT WAS NOT DOING WELL. THE PATIENT WAS BROUGHT BACK FOR A SECOND OPEN COLON RESECTION WITH COLOSTOMY AND IT WAS FOUND THAT THE RECTAL STAPLE LINE OPENED UP. THE STAPLE LINE LOOKED FINE DURING THE PROCEDURE, BUT WHEN REITERATING, IT HAD OPENED UP FROM THE MIDDLE OUT TO THE ENDS. THE SURGICAL TIME FOR SECOND SURGERY WAS EXTENDED FOR 30 MINUTES OR MORE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED DUE TO THE PRODUCT PROBLEM. THE PATIENT HAD TO BE WASHED OUT AND THE RECTAL STUMP WAS CLOSED WHEN OVERSEWN.
28-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, FOLLOWING AN OPEN COLON RESECTION WITH COLOSTOMY, FOR RECTAL STUMP TRANSECTION, THE PATIENT WAS NOT DOING WELL. THE PATIENT WAS BROUGHT BACK FOR A SECOND OPEN COLON RESECTION WITH COLOSTOMY AND IT WAS FOUND THAT THE RECTAL STAPLE LINE OPENED UP. THE STAPLE LINE LOOKED FINE DURING THE PROCEDURE, BUT WHEN REITERATING, IT HAD OPENED UP FROM THE MIDDLE OUT TO THE ENDS. THE SURGICAL TIME FOR SECOND SURGERY WAS EXTENDED FOR 30 MINUTES OR MORE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED DUE TO THE PRODUCT PROBLEM. THE PATIENT HAD TO BE WASHED OUT AND THE RECTAL STUMP WAS CLOSED WHEN OVERSEWN.
28-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, FOLLOWING AN OPEN COLON RESECTION WITH COLOSTOMY, FOR RECTAL STUMP TRANSECTION, THE PATIENT WAS NOT DOING WELL. THE PATIENT WAS BROUGHT BACK FOR A SECOND OPEN COLON RESECTION WITH COLOSTOMY AND IT WAS FOUND THAT THE RECTAL STAPLE LINE OPENED UP. THE STAPLE LINE LOOKED FINE DURING THE PROCEDURE, BUT WHEN REITERATING, IT HAD OPENED UP FROM THE MIDDLE OUT TO THE ENDS. THE SURGICAL TIME FOR SECOND SURGERY WAS EXTENDED FOR 30 MINUTES OR MORE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED DUE TO THE PRODUCT PROBLEM. THE PATIENT HAD TO BE WASHED OUT AND THE RECTAL STUMP WAS CLOSED WHEN OVERSEWN.
28-12-2022 | Injury | covidien lp llc north haven | 12-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, FOLLOWING AN OPEN COLON RESECTION WITH COLOSTOMY, FOR RECTAL STUMP TRANSECTION, THE PATIENT WAS NOT DOING WELL. THE PATIENT WAS BROUGHT BACK FOR A SECOND OPEN COLON RESECTION WITH COLOSTOMY AND IT WAS FOUND THAT THE RECTAL STAPLE LINE OPENED UP. THE STAPLE LINE LOOKED FINE DURING THE PROCEDURE, BUT WHEN REITERATING, IT HAD OPENED UP FROM THE MIDDLE OUT TO THE ENDS. THE SURGICAL TIME FOR SECOND SURGERY WAS EXTENDED FOR 30 MINUTES OR MORE AND THE PATIENT'S HOSPITALIZATION WAS EXTENDED DUE TO THE PRODUCT PROBLEM. THE PATIENT HAD TO BE WASHED OUT AND THE RECTAL STUMP WAS CLOSED WHEN OVERSEWN.
15-11-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | SIGNIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG SURGERY, THE ADAPTER WAS CONNECTED TO THE HANDLE, BUT THE ADAPTER WAS NOT ABLE TO RECOGNIZE THE RELOAD; THE RELOAD COULD BE CLOSED BUT NOT RE-OPENED. THE SURGEON DECIDED TO USE THE MANUAL RETRACTOR TO RE-OPEN THE RELOAD AND UNLOAD. THE DOCTOR USED THE SAME ADAPTOR WITH ANOTHER RELOAD, BUT THERE WAS A CONNECTION PROBLEM WITH THE NEW RELOAD AND THE ADAPTER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | SIGNIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG SURGERY, THE ADAPTER WAS CONNECTED TO THE HANDLE, BUT THE ADAPTER WAS NOT ABLE TO RECOGNIZE THE RELOAD; THE RELOAD COULD BE CLOSED BUT NOT RE-OPENED. THE SURGEON DECIDED TO USE THE MANUAL RETRACTOR TO RE-OPEN THE RELOAD AND UNLOAD. THE DOCTOR USED THE SAME ADAPTOR WITH ANOTHER RELOAD, BUT THERE WAS A CONNECTION PROBLEM WITH THE NEW RELOAD AND THE ADAPTER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG SURGERY, THE ADAPTER WAS CONNECTED TO THE HANDLE, BUT THE ADAPTER WAS NOT ABLE TO RECOGNIZE THE RELOAD; THE RELOAD COULD BE CLOSED BUT NOT RE-OPENED. THE SURGEON DECIDED TO USE THE MANUAL RETRACTOR TO RE-OPEN THE RELOAD AND UNLOAD. THE DOCTOR USED THE SAME ADAPTOR WITH ANOTHER RELOAD, BUT THERE WAS A CONNECTION PROBLEM WITH THE NEW RELOAD AND THE ADAPTER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | SIGNIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG SURGERY, THE ADAPTER WAS CONNECTED TO THE HANDLE, BUT THE ADAPTER WAS NOT ABLE TO RECOGNIZE THE RELOAD; THE RELOAD COULD BE CLOSED BUT NOT RE-OPENED. THE SURGEON DECIDED TO USE THE MANUAL RETRACTOR TO RE-OPEN THE RELOAD AND UNLOAD. THE DOCTOR USED THE SAME ADAPTOR WITH ANOTHER RELOAD, BUT THERE WAS A CONNECTION PROBLEM WITH THE NEW RELOAD AND THE ADAPTER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG SURGERY, THE ADAPTER WAS CONNECTED TO THE HANDLE, BUT THE ADAPTER WAS NOT ABLE TO RECOGNIZE THE RELOAD; THE RELOAD COULD BE CLOSED BUT NOT RE-OPENED. THE SURGEON DECIDED TO USE THE MANUAL RETRACTOR TO RE-OPEN THE RELOAD AND UNLOAD. THE DOCTOR USED THE SAME ADAPTOR WITH ANOTHER RELOAD, BUT THERE WAS A CONNECTION PROBLEM WITH THE NEW RELOAD AND THE ADAPTER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG SURGERY, THE ADAPTER WAS CONNECTED TO THE HANDLE, BUT THE ADAPTER WAS NOT ABLE TO RECOGNIZE THE RELOAD; THE RELOAD COULD BE CLOSED BUT NOT RE-OPENED. THE SURGEON DECIDED TO USE THE MANUAL RETRACTOR TO RE-OPEN THE RELOAD AND UNLOAD. THE DOCTOR USED THE SAME ADAPTOR WITH ANOTHER RELOAD, BUT THERE WAS A CONNECTION PROBLEM WITH THE NEW RELOAD AND THE ADAPTER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2022 | Malfunction | covidien lp llc north haven | 12-01-2023 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG SURGERY, THE ADAPTER WAS CONNECTED TO THE HANDLE, BUT THE ADAPTER WAS NOT ABLE TO RECOGNIZE THE RELOAD; THE RELOAD COULD BE CLOSED BUT NOT RE-OPENED. THE SURGEON DECIDED TO USE THE MANUAL RETRACTOR TO RE-OPEN THE RELOAD AND UNLOAD. THE DOCTOR USED THE SAME ADAPTOR WITH ANOTHER RELOAD, BUT THERE WAS A CONNECTION PROBLEM WITH THE NEW RELOAD AND THE ADAPTER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-11-2022 | Injury | ethicon endo-surgery | 13-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT ON NOV 28TH, THE SURGEON PERFORMED A ROBOTICALLY ASSISTED SIGMOID COLECTOMY. THE SURGEON USED AN ETHICON POWERED CIRCULAR STAPLER FOR THE ANASTOMOSIS. THE ANASTOMOSIS PASSED AN AIR-LEAK TEST. ON DEC 5TH, THE SURGEON DECIDED TO RE-OPERATE ON THE PATIENT SUSPECTING A LEAK IN THE COLON. THE SURGEON FOUND A HOLE IN THE SIDE OF THE ANASTOMOSIS. A COLOSTOMY WAS PERFORMED, AND THE PATIENT IS RECOVERING. THE SURGERY WAS NOT DELAYED AND WAS COMPLETED SUCCESSFULLY. THERE WERE NO FRAGMENTS GENERATED.
28-11-2022 | Injury | ethicon endo-surgery | 13-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: IT WAS REPORTED THAT ON NOV 28TH, THE SURGEON PERFORMED A ROBOTICALLY ASSISTED SIGMOID COLECTOMY. THE SURGEON USED AN ETHICON POWERED CIRCULAR STAPLER FOR THE ANASTOMOSIS. THE ANASTOMOSIS PASSED AN AIR-LEAK TEST. ON DEC 5TH, THE SURGEON DECIDED TO RE-OPERATE ON THE PATIENT SUSPECTING A LEAK IN THE COLON. THE SURGEON FOUND A HOLE IN THE SIDE OF THE ANASTOMOSIS. A COLOSTOMY WAS PERFORMED, AND THE PATIENT IS RECOVERING. THE SURGERY WAS NOT DELAYED AND WAS COMPLETED SUCCESSFULLY. THERE WERE NO FRAGMENTS GENERATED.
13-12-2022 | Death | ethicon endo-surgery | 13-01-2023 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE THAT THE SURGEON PERFORMED AN OPEN COLON RESECTION. THE ANASTOMOSIS WAS A SIDE-TO-SIDE, PERFORMED WITH A TLC75 AND WAS OVER-SEWN WITH SUTURE. THE PATIENT HAD BEEN RECOVERING WELL WITH SUCCESSFUL BOWEL MOVEMENTS. ON DEC 21ST, THE PATIENT STARTED TO EXPERIENCE ABDOMINAL PAIN AND OTHER INDICATIONS OF A LEAK IN THE COLON. ON DECEMBER 21ST, THE SURGEON RE-OPERATED ON THE PATIENT. THE ABDOMEN WAS COMPLETELY FILLED WITH STOOL, INCLUDING IN THE ABDOMINAL WALL CLOSURE FROM THE INITIAL OPERATION. AFTER REMOVAL OF STOOL, A HOLE WAS FOUND ON THE SIDE OF THE ANASTOMOSIS. THE ANASTOMOSIS WAS RESECTED USING A TLC75. A COLOSTOMY WAS CREATED, AND THE ABDOMEN WAS PARTIALLY CLOSED. STRATAFIX SYMMETRIC WAS USED TO CLOSE THE FASCIA AND IMPROVED BRIDGING SUTURES WERE MADE WITH A RED-RUBBER CATHETER AND SUTURED. AN ANTIBACTERIAL DRESSING WAS PACKED INTO THE WOUND. THE PATIENT IS CURRENTLY RECOVERING FROM THE RE-OPERATION.
13-12-2022 | Death | ethicon endo-surgery | 13-01-2023 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE THAT THE SURGEON PERFORMED AN OPEN COLON RESECTION. THE ANASTOMOSIS WAS A SIDE-TO-SIDE, PERFORMED WITH A TLC75 AND WAS OVER-SEWN WITH SUTURE. THE PATIENT HAD BEEN RECOVERING WELL WITH SUCCESSFUL BOWEL MOVEMENTS. ON DEC 21ST, THE PATIENT STARTED TO EXPERIENCE ABDOMINAL PAIN AND OTHER INDICATIONS OF A LEAK IN THE COLON. ON DECEMBER 21ST, THE SURGEON RE-OPERATED ON THE PATIENT. THE ABDOMEN WAS COMPLETELY FILLED WITH STOOL, INCLUDING IN THE ABDOMINAL WALL CLOSURE FROM THE INITIAL OPERATION. AFTER REMOVAL OF STOOL, A HOLE WAS FOUND ON THE SIDE OF THE ANASTOMOSIS. THE ANASTOMOSIS WAS RESECTED USING A TLC75. A COLOSTOMY WAS CREATED, AND THE ABDOMEN WAS PARTIALLY CLOSED. STRATAFIX SYMMETRIC WAS USED TO CLOSE THE FASCIA AND IMPROVED BRIDGING SUTURES WERE MADE WITH A RED-RUBBER CATHETER AND SUTURED. AN ANTIBACTERIAL DRESSING WAS PACKED INTO THE WOUND. THE PATIENT IS CURRENTLY RECOVERING FROM THE RE-OPERATION.
25-11-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | 55MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SMALL BOWEL RESECTION PROCEDURE THE STAPLER WOULDN¿T FIRE. ANOTHER STAPLER WAS OPENED AND USED SUCCESSFULLY. NO FRAGMENTS WERE GENERATED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
23-12-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 13-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC RIGHT HEMICOINTESTINAL RESECTION, IN DISSOCIATING THE INTESTINE, THE SIXTY MILLIMETER (60MM) RELOAD DID NOT FIRE. EVEN THOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE STAPLE BURST ON THE SECOND RELOAD. TO RESOLVE THE ISSUE, A NEW 60MM RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
27-10-2022 | Malfunction | covidien lp llc north haven | 13-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE THREE HANDLES FAILED TO CHARGE. THE HANDLES WERE REMOVED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS OPENED AND BOTH BATTERY CELLS WERE VENTED.
27-10-2022 | Malfunction | covidien lp llc north haven | 13-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE THREE HANDLES FAILED TO CHARGE. THE HANDLES WERE REMOVED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS OPENED AND BOTH BATTERY CELLS WERE VENTED.
20-12-2022 | Malfunction | covidien lp llc north haven | 13-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PULMONARY RESECTION, A BLACK RELOAD WAS USED IN PULMONARY PARENCHYMA, HOWEVER, IT STOPPED MIDWAY. FIRED A NEW RELOAD THAT WAS OPENED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | RELOAD- CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE STAPLER MADE A CUTTING SOUND WITHOUT STAPLING.
28-12-2022 | Injury | covidien lp llc north haven | 13-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC (VATS) LOBECTOMY, DURING LIGATION OF LINGULAR ARTERIAL BRANCH. ALL THE STAPLE LINES BROKE DUE TO COMPRESSION. ARTERIAL BRANCH HAD TO BE HAND SUTURED. THE SURGERY WAS CONVERTED INTO AN OPEN PROCEDURE AND SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES. SURGICAL INCISION WAS EXTENDED FOR MORE THAN ONE INCH AND THE PATIENT NEEDED TO EXTEND HOSPITALIZATION FOR 3-4 DAYS DUE TO MALFUNCTION. UNANTICIPATED TISSUE LOSS WAS ALSO NOTED.
27-12-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGECTOMY IT WAS DIFFICULT TO PLACE ANVIL ON THE DEVICE TROCAR. DEVICE TROCAR PULLED BACKWARDS WITHOUT PUTTING TOO MUCH PRESSURE IN IT. THE PROCEDURE WAS DELAYED BY 5-MINUTES. THE PROCEDURE WAS COMPLETED WITH NO FRAGMENTS GENERATED. THERE WERE NO PATIENT CONSEQUENCES.
01-01-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE DIDN¿T FIRE DURING THE CASE. THE PROCEDURE WAS DELAYED BUT COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED.
10-11-2022 | Malfunction | us surgical puerto rico | 13-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOBECTOMY, WHEN 60MILLIEMETER PURPLE STAPLE WAS USED TO CUT THE LOBE, THE STAPLES WERE POORLY FORMED. SUTURES WERE ADDED TO STOP THE BLEEDING. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
10-11-2022 | Malfunction | us surgical puerto rico | 13-01-2023 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOBECTOMY, WHEN 60MILLIEMETER PURPLE STAPLE WAS USED TO CUT THE LOBE, THE STAPLES WERE POORLY FORMED. SUTURES WERE ADDED TO STOP THE BLEEDING. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
15-12-2022 | Malfunction | us surgical puerto rico | 13-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-12-2022 | Malfunction | us surgical puerto rico | 13-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, AFTER COMPLETION OF STAPLING, THE SURGEON FOUND ONE MALFORMED STAPLE ON THE OUTER ROW OF THE STAPLE LINE. CASE WAS COMPLETED WITHOUT ISSUES OR ANY IMPACT TO THE PATIENT.
29-12-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN PROCEDURE THE STAPLE DID NOT COME OUT FULLY.
03-05-2021 | Malfunction | ethicon endo-surgery | 13-01-2023 | POWERED PLUS LONG 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN LAR THERE WAS ONLY A PARTIAL STAPLE LINE. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
29-12-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RECTUM SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-12-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE PATIENT SUBMITTED TO THE VIDEO PROCEDURE IN AN EMERGENCY SECONDARY OBSTRUCTIVE TUMOR OF THE SIGMOID COLON. DURING ANASTOMOSIS OF THE RECTUM WITH DESCENDING COLON, AN OPENING AREA WAS OBSERVED ON THE POSTERIOR WALL OF THE ANASTOMOSIS DUE TO THE FAILURE OF THE STAPLER NOT STAPLING THE AREA TO BE ANASTOMOSED, BEING NECESSARY TO REDO THE ANASTOMOSIS USING A NEW CIRCULAR STAPLER AND 02 MORE LOADS FOR THE ECHELON ENDO-STAPLER OF 60MM. THERE WERE NO PATIENT CONSEQUENCES.
01-11-2022 | Death | covidien lp llc north haven | 13-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN AUGUST 2017 TO APRIL 2021 REGARDING A STUDY TO EVALUATE THE IMPACT OF THE USE OF A REINFORCED STAPLER WITH THAT OF A STANDARD STAPLER FOR CLOSURE OF THE PANCREATIC PARENCHYMA DURING DISTAL PANCREATECTOMY ON 199 PATIENTS. PATIENTS WERE RANDOMIZED IN 2 GROUPS FOR THE USE OF A STANDARD STAPLER OR A REINFORCED STAPLER TO CLOSE REMNANT PANCREATIC PARENCHYMA. THERE WERE 99 PATIENTS IN THE STANDARD STAPLER GROUP AND 100 PATIENTS IN THE REINFORCED STAPLER GROUP. THERE WAS ONE DEATH IN THE SIMPLE STAPLER GROUP DUE TO UNSPECIFIED CAUSE. EFFECT OF THE USE OF REINFORCED STAPLING ON THE OCCURRENCE OF PANCREATIC FISTULA AFTER DISTAL PANCREATECTOMY AUDE MERDRIGNAC, MD, PHD, JONATHAN GARNIER, MD, SAFI DOKMAK, MD, NICOLAS REGENET, MD, MICKAËL LESURTEL, MD, PHD, JEAN YVES MABRUT, MD, PHD, ANTONIO SA CUNHA, MD, PHD, DAVID FUKS, MD, PHD, DAMIEN BERGEAT, MD, FABIEN ROBIN, MD, ESTELLE LE PABIC, MSC, KARIM BOUDJEMA, MD, PHD, OLIVIER TURRINI, MD, PHD,§ BRUNO LAVIOLLE, MD, PHD, AND LAURENT SULPICE, MD, PHD 2022
29-12-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | ILS 29MM, CURVED | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THE DURING THE UNKNOWN PROCEDURE THE ANVIL AND CDH BODY TROCAR WERE NOT COMBINED.
28-11-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT DEVICE WOULD NOT FIRE. UNKNOWN IF THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-09-2022 | Malfunction | ethicon endo-surgery | 13-01-2023 | GOLD RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-12-2022 | Malfunction | us surgical puerto rico | 13-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC DUODENECTOMY, IN THE RESECTION OF THE GASTRODUODENAL ARTERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Injury | us surgical puerto rico | 13-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, ON PULMONARY PARENCHYMA RESECTION, THE FIRING COULD BE DONE TILL MIDWAY BUT THE HANDLE COULD NOT BE SQUEEZED SINCE HANDLE MADE ABNORMAL NOISE. THE HANDLE AND RELOAD WERE REPLACED AND ONCE AGAIN ADDITIONAL CUTTING WAS PERFORMED BUT THE SAME EVENT OCCURRED AGAIN THEN WHEN THE HANDLE AND THE RELOAD WERE REPLACED AGAIN. BOTH THE THIRD HANDLE AND RELOAD WERE ABLE TO BE USED WITH NO PROBLEM. AFTER THAT, RESECTION WAS DONE WITHOUT ANY PROBLEMS.
15-12-2022 | Injury | us surgical puerto rico | 13-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, ON PULMONARY PARENCHYMA RESECTION, THE FIRING COULD BE DONE TILL MIDWAY BUT THE HANDLE COULD NOT BE SQUEEZED SINCE HANDLE MADE ABNORMAL NOISE. THE HANDLE AND RELOAD WERE REPLACED AND ONCE AGAIN ADDITIONAL CUTTING WAS PERFORMED BUT THE SAME EVENT OCCURRED AGAIN THEN WHEN THE HANDLE AND THE RELOAD WERE REPLACED AGAIN. BOTH THE THIRD HANDLE AND RELOAD WERE ABLE TO BE USED WITH NO PROBLEM. AFTER THAT, RESECTION WAS DONE WITHOUT ANY PROBLEMS.
15-12-2022 | Injury | covidien lp llc north haven | 13-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, ON PULMONARY PARENCHYMA RESECTION, THE FIRING COULD BE DONE TILL MIDWAY BUT THE HANDLE COULD NOT BE SQUEEZED SINCE HANDLE MADE ABNORMAL NOISE. THE HANDLE AND RELOAD WERE REPLACED AND ONCE AGAIN ADDITIONAL CUTTING WAS PERFORMED BUT THE SAME EVENT OCCURRED AGAIN THEN WHEN THE HANDLE AND THE RELOAD WERE REPLACED AGAIN. BOTH THE THIRD HANDLE AND RELOAD WERE ABLE TO BE USED WITH NO PROBLEM. AFTER THAT, RESECTION WAS DONE WITHOUT ANY PROBLEMS.
29-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 13-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOENTERECTOMY, IN DISSOCIATING THE BOWEL, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2022 | Injury | us surgical puerto rico | 13-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, ON PULMONARY PARENCHYMA RESECTION, THE FIRING COULD BE DONE TILL MIDWAY BUT THE HANDLE COULD NOT BE SQUEEZED SINCE HANDLE MADE ABNORMAL NOISE. THE HANDLE AND RELOAD WERE REPLACED AND ONCE AGAIN ADDITIONAL CUTTING WAS PERFORMED BUT THE SAME EVENT OCCURRED AGAIN THEN WHEN THE HANDLE AND THE RELOAD WERE REPLACED AGAIN. BOTH THE THIRD HANDLE AND RELOAD WERE ABLE TO BE USED WITH NO PROBLEM. AFTER THAT, RESECTION WAS DONE WITHOUT ANY PROBLEMS.
13-12-2022 | Injury | intuitive surgical | 13-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GENERAL GASTRIC BYPASS SURGICAL PROCEDURE, STAPLES WERE MISSING FROM THE STAPLE LINE AFTER A WHITE SUREFORM 60 RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. INTUITIVE SURGICAL, INC. (ISI) CONTACTED A NURSE AT THE SITE AND OBTAINED THE FOLLOWING INFORMATION: THE EVENT OCCURRED DURING THE 4TH OR 5TH STAPLER FIRE WHILE THE SURGEON WAS STAPLING THE INTESTINE AND DIVIDING THE ROUX LIMB. THE STAPLER INSTRUMENT FIRED, BUT THE STAPLE LINE WAS NOT INTACT. THE SURGEON REPAIRED THE STAPLE LINE DEFECT BY FIRING ANOTHER STAPLER RELOAD AND APPLYING A SUTURE. THIS FIRING FAILURE RESULTED IN UNPLANNED TISSUE REMOVAL. THE NURSE CONFIRMED THAT THERE WAS NO ISSUE CLAMPING DOWN, NO BUTTRESS MATERIAL USED, NO SYSTEM-GENERATED ERROR MESSAGES, NO CROSSING OF STAPLE LINES, AND NO HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. THE TISSUE WAS NOT CALCIFIED AND EXPOSED TO RADIATION OR CHEMOTHERAPY BEFORE THE PROCEDURE. THERE WAS NO TISSUE TENSION OR TISSUE BUNCHING. THE SURGEON USED A BACKUP SUREFORM 60 STAPLER INSTRUMENT TO CONTINUE THE PROCEDURE WITHOUT FURTHER ISSUES. THE PATIENT IS REPORTED TO BE RECOVERING WELL WITH NO POST-OPERATIVE COMPLICATIONS.
25-10-2022 | Injury | intuitive surgical | 13-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
15-12-2022 | Injury | intuitive surgical | 14-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
15-12-2022 | Injury | intuitive surgical | 14-01-2023 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
12-12-2022 | Injury | intuitive surgical | 14-01-2023 | SUREFORM | Misfire | A050502 | Insufficient Information | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
24-11-2022 | Malfunction | intuitive surgical | 15-01-2023 | SUREFORM | No Device Output | A0907 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
01-11-2022 | Malfunction | intuitive surgical | 15-01-2023 | SUREFORM | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY WEDGE RESECTION SURGICAL PROCEDURE, AN ERROR MESSAGE ¿TURN 40 TIMES MANUALLY¿ APPEARED WHEN THE CUSTOMER WAS USING THE GREEN RELOAD WITH THE SUREFORM 60 STAPLER INSTRUMENT. A BACKUP RELOAD OF SAME KIND WAS USED TO CONTINUE THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITH NO REPORT OF PATIENT INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION: THE RELOAD WAS INSPECTED PRIOR TO USE WITH NO ABNORMALITY. WHILE USING THE INSTRUMENT WITH THE GREEN RELOAD, AN ERROR ¿TURN 40 TIMES MANUALLY¿ MESSAGE OCCURRED AFTER THE 5TH FIRING. THE 1ST TO 4TH FIRES WERE SUCCESSFUL WITH THE BLUE AND WHITE RELOADS. THE INSTRUMENT JAWS WERE STUCK ON THE LUNG TISSUE WHEN THE ISSUE OCCURRED. THEY TURNED THE MANUAL RELEASE KNOB AND SUCCESSFULLY OPENED THE JAWS INTRAOPERATIVELY AND RELEASED THE TISSUE. THERE WAS NO ADVERSE EFFECT TO THE GRASPED TISSUE. PORT INCISIONS WERE NOT INCREASED. THE SURGEON DID NOT EXPERIENCE ANY CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD. THERE WAS NO REPORT OF UNEXPECTED TISSUE REMOVAL AND BLEEDING.
16-12-2022 | Injury | intuitive surgical | 16-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA REPAIR PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT FIRED AN UNSPECIFIED COLORED RELOAD AND THE TARGET TISSUE WAS PUSHED INSTEAD OF CUT. IT IS UNKNOWN WHAT MEDICAL INTERVENTION WAS RENDERED DUE TO THE STAPLING ISSUE. THE PROCEDURE WAS REPORTEDLY COMPLETED ROBOTICALLY. INTUITIVE SURGICAL, INC. (ISI) MADE MULTIPLE FOLLOW-UP ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO FURTHER DETAILS HAVE BEEN RECEIVED AS OF THE DATE OF THIS REPORT.
16-12-2022 | Injury | intuitive surgical | 16-01-2023 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA REPAIR PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT FIRED AN UNSPECIFIED COLORED RELOAD AND THE TARGET TISSUE WAS PUSHED INSTEAD OF CUT. IT IS UNKNOWN WHAT MEDICAL INTERVENTION WAS RENDERED DUE TO THE STAPLING ISSUE. THE PROCEDURE WAS REPORTEDLY COMPLETED ROBOTICALLY. INTUITIVE SURGICAL, INC. (ISI) MADE MULTIPLE FOLLOW-UP ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO FURTHER DETAILS HAVE BEEN RECEIVED AS OF THE DATE OF THIS REPORT.
16-12-2022 | Malfunction | intuitive surgical | 16-01-2023 | SUREFORM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI- ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, THE SUREFORM 45 STAPLER INSTRUMENT WAS WORKING NORMALLY, THEN A "PAUSING FOR COMPRESSION, TISSUE TOO THICK TO CONTINUE" MESSAGE DISPLAYED. THE INSTRUMENT STOPPED WORKING WHEN USING THE WHITE RELOAD ON THE 6TH FIRE. THE CUSTOMER REMOVED THE INSTRUMENT TO CHECK AND FOUND THAT THE BLADE WAS BURIED IN THE MIDDLE. IT WAS CONFIRMED THAT THE INSTRUMENT TIP WAS WASHED AND USED FOR EACH FIRE. AT THE TIME OF FIRING ON THE TARGET BLOOD VESSEL (A3), THE INSTRUMENT STOPPED AT 27MM. ONE STAPLE WAS APPLIED TO A3, BUT NO TISSUE OR OTHERS WERE CAUGHT UP TO A POINT OF 27 MM. THERE APPEARED TO BE NO OBSTRUCTIVE TISSUE. IT WAS UNKNOWN IF A FRAGMENT FELL INTO THE PATIENT. THE PROCEDURE WAS DELAYED FOR GREATER THAN 30 MINUTES. THE PROCEDURE WAS COMPLETED WITH A THIRD-PARTY PRODUCT. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION: THE INSTRUMENT AND RELOAD WERE INSPECTED PRIOR TO USE WITH NO ABNORMALITY. THIS EVENT INVOLVED A PARTIAL FIRE. THE STAPLER STOPPED IN THE MIDDLE OF FIRING WHEN RESECTING THE A3 VESSEL. THE CUSTOMER ELECTED THE WHITE (2.5MM) RELOAD BECAUSE IT WAS THE MOST APPROPRIATE FOR THE SUPERIOR LOBAR ARTERY (A3 VESSEL OF PULMONARY ARTERY) FROM THE PAST EXPERIENCE. THE TARGET TISSUE WAS NOT CALCIFIED NOR EXPOSED TO ANY RADIATION OR CHEMOTHERAPY PRIOR TO THE PROCEDURE. THERE WAS NO TISSUE TENSION NOR BUNCHING. NO BLEEDING WAS OBSERVED ON THE STAPLE LINE DUE TO THE STAPLER ISSUE. IN ADDITION, THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS BETWEEN THE INSTRUMENT JAWS AND DID NOT FIRE OVER AN EXISTING STAPLE LINE. BUTTRESS MATERIAL WAS NOT USED. THE SURGEON DID NOT EXPERIENCE ANY CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD. THE CUSTOMER CONFIRMED THAT NO FRAGMENT FELL INSIDE OF THE PATIENT. THE PROCEDURE WAS DELAYED FOR MORE THAN 30 MINUTES BUT THERE WERE NO INTRA OR POST OPERATIVE COMPLICATIONS IDENTIFIED DUE TO THIS DELAY. THE PATIENT HAD NO INJURY/HARM. INFORMATION REGARDING RELEVANT TESTING, AND MEDICAL HISTORY WERE REQUESTED HOWEVER, THE REPORTER WAS NOT ABLE TO PROVIDE THAT INFORMATION.
16-12-2022 | Malfunction | intuitive surgical | 16-01-2023 | SUREFORM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
15-12-2022 | Injury | covidien lp llc north haven | 16-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE VIDEO ASSISTED THORACOSCOPIC LOBECTOMY SURGERY, ON PULMONARY PARENCHYMA RESECTION, THE FIRING COULD BE DONE TILL MIDWAY BUT THE HANDLE COULD NOT BE SQUEEZED SINCE HANDLE MADE ABNORMAL NOISE. THE HANDLE AND RELOAD WERE REPLACED AND ONCE AGAIN ADDITIONAL CUTTING WAS PERFORMED BUT THE SAME EVENT OCCURRED AGAIN THEN WHEN THE HANDLE AND THE RELOAD WERE REPLACED AGAIN. BOTH THE THIRD HANDLE AND RELOAD WERE ABLE TO BE USED WITH NO PROBLEM. AFTER THAT, RESECTION WAS DONE WITHOUT ANY PROBLEMS.
23-11-2022 | Malfunction | us surgical puerto rico | 16-01-2023 | GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE EDGE OF THE DEVICE BROKE. THERE WAS NO PATIENT INVOLVEMENT.
01-01-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE ANVIL WAS NOT ATTACHING TO THE TROCAR TIP OF THE DEVICE AND KEPT POPPING OFF SEVERAL ATTEMPTS. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
21-12-2022 | Malfunction | davol inc sub c.r bard inc -1213643 | 17-01-2023 | CAPSURE STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED DURING A LAPAROSCOPIC UMBILICAL HERNIA REPAIR PROCEDURE ON (B)(6) 2022, A CAPSURE FIXATION DEVICE WAS USED TO FIXATE THE MESH. AS REPORTED THE FIRST FOUR FASTENERS SUCCESSFULLY FIXATED THE MESH, HOWEVER THE NEXT FOUR FASTENERS DEPLOYED FROM THE DEVICE ASKEW/TWISTED AND DID NOT FIXATE THE MESH. THESE LOOSE FASTENERS WERE REMOVED FROM THE PATIENT'S BODY. A SECOND DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO INJURY/HARM TO THE PATIENT.
03-01-2023 | Malfunction | ethicon endo-surgery | 17-01-2023 | ECHELON FLEX | Battery Problem | A0705 | Insufficient Information | Event Description: SURGEON WAS USING STAPLER TO CUT BOWEL DURING A BOWEL RESECTION PROCEDURE. BATTERY ON STAPLER DIED WHILE STAPLING. SURGEON WAS ABLE TO REMOVE STAPLER FROM PATIENT. CIRCULATOR REPLACED DEFECTIVE STAPLER AND PROCEDURE CONTINUED WITHOUT FURTHER ISSUES.
29-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY PROCEDURE, IN LIGATION OF A PULMONARY VESSEL, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON USED A LOT OF FORCE TO RE-OPEN THE BLOCKED JAW. THE SURGERY WAS EXTENDED FOR OVER THIRTY MINUTES. IT WAS NOTED THAT THE HANDLE WORKED WITH THE OTHER LOAD, WHEN USE OF THE HANDLE WAS CONTINUED AFTER THE PROBLEM WITH THE LOAD. THERE WAS NO PATIENT INJURY.
29-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY PROCEDURE, IN LIGATION OF A PULMONARY VESSEL, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON USED A LOT OF FORCE TO RE-OPEN THE BLOCKED JAW. THE SURGERY WAS EXTENDED FOR OVER THIRTY MINUTES. IT WAS NOTED THAT THE HANDLE WORKED WITH THE OTHER LOAD, WHEN USE OF THE HANDLE WAS CONTINUED AFTER THE PROBLEM WITH THE LOAD. THERE WAS NO PATIENT INJURY.
30-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATECTOMY, AFTER PRESSING THE GREEN BUTTON, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
30-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEPATECTOMY, AFTER PRESSING THE GREEN BUTTON, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
30-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 17-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, IN DETACHING THE RECTUM, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED, AND THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
30-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL SURGERY PROCEDURE TO THE PATIENT THAT RECEIVED CHEMOTHERAPY OR RADIATION THERAPY, IN RECTAL RESECTION, WITH THE FIRST RELOAD, THE FIRING STOPPED HALFWAY. WITH THE SECOND RELOAD, THE SURGEON WAS UNABLE TO CONTINUE FIRING MIDWAY. BOTH STAPLES THAT HAVE BEEN FIRED DID NOT FORM AS WELL. TO RESOLVE THE ISSUE, A RECUT WAS MADE AND A NEW RELOAD WAS USED.
30-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL SURGERY PROCEDURE TO THE PATIENT THAT RECEIVED CHEMOTHERAPY OR RADIATION THERAPY, IN RECTAL RESECTION, WITH THE FIRST RELOAD, THE FIRING STOPPED HALFWAY. WITH THE SECOND RELOAD, THE SURGEON WAS UNABLE TO CONTINUE FIRING MIDWAY. BOTH STAPLES THAT HAVE BEEN FIRED DID NOT FORM AS WELL. TO RESOLVE THE ISSUE, A RECUT WAS MADE AND A NEW RELOAD WAS USED.
30-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL SURGERY PROCEDURE TO THE PATIENT THAT RECEIVED CHEMOTHERAPY OR RADIATION THE RAPY, IN RECTAL RESECTION, WITH THE FIRST RELOAD, THE FIRING STOPPED HALFWAY. WITH THE SECOND RELOAD, THE SURGEON WAS UNABLE TO CONTINUE FIRING MIDWAY. BOTH STAPLES THAT HAVE BEEN FIRED DID NOT FORM AS WELL. TO RESOLVE THE ISSUE, A RECUT WAS MADE AND A NEW RELOAD WAS USED.
30-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL SURGERY PROCEDURE TO THE PATIENT THAT RECEIVED CHEMOTHERAPY OR RADIATION THE RAPY, IN RECTAL RESECTION, WITH THE FIRST RELOAD, THE FIRING STOPPED HALFWAY. WITH THE SECOND RELOAD, THE SURGEON WAS UNABLE TO CONTINUE FIRING MIDWAY. BOTH STAPLES THAT HAVE BEEN FIRED DID NOT FORM AS WELL. TO RESOLVE THE ISSUE, A RECUT WAS MADE AND A NEW RELOAD WAS USED.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Misfire | A050502 | Feeding Problem | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Misfire | A050502 | Sepsis | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Failure to Form Staple | A050704 | Feeding Problem | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Failure to Form Staple | A050704 | Sepsis | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Material Twisted/Bent | A040609 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Material Twisted/Bent | A040609 | Feeding Problem | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Material Twisted/Bent | A040609 | Sepsis | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Material Twisted/Bent | A040609 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Feeding Problem | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
21-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS NO STAPLE FORMATION; THE DONUT WAS INCOMPLETE; THE ANVIL WAS BENT. AFTER THE CORRECT CLOSURE OF THE CIRCULAR STAPLER, WHICH WAS VERIFIED BY THE GREEN LINE, THE STAPLER WAS ACTIVATED BY MEANS OF THE APPROPRIATE LEVER. FROM ACTIVATION, HOWEVER, THERE WAS ONLY THE SECTION OF THE CIRCULAR RINGS WITHOUT THE CORRECT CLOSURE OF THE METAL STAPLES THAT ALLOWED THE ANASTOMOTIC SUTURE. THEREFORE, THE STAPLES WERE FOUND TO BE OPEN WITH CONSEQUENT NON-SUTURING OF THE ANASTOMOTIC TISSUE IN THE PATIENT WITH LOW RECTAL RESECTION. AFTER THE DOCTOR HAD CUT WITHOUT APPLYING THE STITCHES, THE SURGEON HAD TO REPACKAGE THE POUCH AND THEN HAD TO REDO THE STEP OF INSERTING THE NEW STAPLER AND RESECTED AND RE-STAPLED TO FIX THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. IN THE POST-OPERATIVE PERIOD, THE STAPLES DID NOT HOLD AND REOPENED. WITH AN ILEOSTOMY FOR TEMPORARY PROTECTION, THE PATIENT WAS STILL HOSPITALIZED IN THE WARD IN A SEPTIC STATE AND WAS TREATED WITH MEDICAL THERAPY WITH ENDOLUMINAL VACUUM THERAPY, ANTIBIOTICS, PARENTERAL SUPPORT AND PAINKILLERS.
29-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | EEA | Positioning Problem | A1502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN HEMORRHOIDECTOMY, WHEN THE SURGEON INSERTED THE DEVICE INTO THE ANUS TO CUT HEMORRHOIDS, THE ANVIL HEAD AND THE DEVICE DID NOT MATCH. THE DEVICE WAS UNABLE TO BE CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | EEA | Misassembly by Users | A230501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN OPEN HEMORRHOIDECTOMY, WHEN THE SURGEON INSERTED THE DEVICE INTO THE ANUS TO CUT HEMORRHOIDS, THE ANVIL HEAD AND THE DEVICE DID NOT MATCH. THE DEVICE WAS UNABLE TO BE CLOSED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC RIGHT HEMICOINTESTINAL RESECTION, IN DISSOCIATING THE INTESTINE, THE SIXTY MILLIMETER (60MM) RELOAD DID NOT FIRE. EVEN THOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE STAPLE BURST ON THE SECOND RELOAD. TO RESOLVE THE ISSUE, A NEW 60MM RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
25-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC COLORECTAL SURGERY, WHEN CUTTING THE RECTUM, THE STAPLER STOPPED FIRING HALFWAY AND THERE WERE STAPLES DID NOT FORM INTO B SHAPE. TO FIX THE STAPLE LINE, THE HANDLE AND RELOAD WERE REPLACED AND RE-CUTTING WAS PERFORMED.
18-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHEN THE GASTRIC PARENCHYMA WAS RESECTED, ALTHOUGH THE TISSUE WAS NOT THAT THICK, THE HANDLE MADE A CRACKING SOUND. THE HANDLE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, WHEN THE GASTRIC PARENCHYMA WAS RESECTED, ALTHOUGH THE TISSUE WAS NOT THAT THICK, THE HANDLE MADE A CRACKING SOUND. THE HANDLE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2022 | Injury | us surgical puerto rico | 17-01-2023 | GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN DISTAL GASTRECTOMY, AT THE DISSOCIATED STOMACH, THE STAPLE LINE WAS INCOMPLETE AND HAS POOR FORMATION. THE SURGEON OVER SEW THE TISSUE TO FIX THE STAPLE LINE.
30-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC TOTAL GASTRECTOMY, WHEN DISSOCIATING THE ESOPHAGUS, THERE WAS A BROKEN SOUND INSIDE THE HANDLE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
30-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOSCOPIC TOTAL GASTRECTOMY, WHEN DISSOCIATING THE ESOPHAGUS, THERE WAS A BROKEN SOUND INSIDE THE HANDLE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
22-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | GIA | Failure to Align | A050701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY TO THE VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, THE JAWS WOULD NOT OPEN SO USER WOULD NOT BE ABLE TO PROCEED TO FIRING. THE FIRING KNOB WAS UNABLE TO SWIPE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
29-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC HEPATECTOMY, ON LIVER RESECTION, THE RELOAD DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED AS REPLACEMENT BUT STILL THE RELOAD DID NOT FIRE. ANOTHER RELOAD WAS USED WITH THE NEW HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, IN STAPLING OF THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE TWO DEVICES DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
29-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION, IN STAPLING OF THE LUNG TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE TWO DEVICES DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, LITTLE PIECE FROM THE YELLOW SHIPPING WEDGE BROKE OFF FROM THE RELOAD WHILE IN THE PATIENT'S ABDOMEN. THE SURGEON HAD TO REMOVE IT INTRA OPERATIVELY WHEN IT WAS SEEN. THERE WAS NO PATIENT INJURY.
13-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, LITTLE PIECE FROM THE YELLOW SHIPPING WEDGE BROKE OFF FROM THE RELOAD WHILE IN THE PATIENT'S ABDOMEN. THE SURGEON HAD TO REMOVE IT INTRA OPERATIVELY WHEN IT WAS SEEN. THERE WAS NO PATIENT INJURY.
02-01-2023 | Malfunction | covidien lp llc north haven | 17-01-2023 | ABSORBATACK | Positioning Failure | A150201 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN IPOM (INTRAPERITONEAL ONLAY MESH) FIXATION, THE TRIGGER WAS HARD TO FIRE AFTER THE TEN FIRING. TACKS DISENGAGE INTO THE PATIENTS CAVITY AND WERE RETRIEVED. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
02-01-2023 | Malfunction | covidien lp llc north haven | 17-01-2023 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN IPOM (INTRAPERITONEAL ONLAY MESH) FIXATION, THE TRIGGER WAS HARD TO FIRE AFTER THE TEN FIRING. TACKS DISENGAGE INTO THE PATIENTS CAVITY AND WERE RETRIEVED. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
02-01-2023 | Malfunction | covidien lp llc north haven | 17-01-2023 | ABSORBATACK | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN IPOM (INTRAPERITONEAL ONLAY MESH) FIXATION, THE TRIGGER WAS HARD TO FIRE AFTER THE TEN FIRING. TACKS DISENGAGE INTO THE PATIENTS CAVITY AND WERE RETRIEVED. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
02-01-2023 | Malfunction | covidien lp llc north haven | 17-01-2023 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN IPOM (INTRAPERITONEAL ONLAY MESH) FIXATION, THE TRIGGER WAS HARD TO FIRE AFTER THE TEN FIRING. TACKS DISENGAGE INTO THE PATIENTS CAVITY AND WERE RETRIEVED. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
02-01-2023 | Malfunction | covidien lp llc north haven | 17-01-2023 | ABSORBATACK | Firing Problem | A0505 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN IPOM (INTRAPERITONEAL ONLAY MESH) FIXATION, THE TRIGGER WAS HARD TO FIRE AFTER THE TEN FIRING. TACKS DISENGAGE INTO THE PATIENTS CAVITY AND WERE RETRIEVED. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
02-01-2023 | Malfunction | covidien lp llc north haven | 17-01-2023 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN IPOM (INTRAPERITONEAL ONLAY MESH) FIXATION, THE TRIGGER WAS HARD TO FIRE AFTER THE TEN FIRING. TACKS DISENGAGE INTO THE PATIENTS CAVITY AND WERE RETRIEVED. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE.
26-10-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, ON AN EXPLORATORY LAPAROTOMY FOR TRAUMA WITH SMALL BOWEL RESECTION, PARTIAL SIGMOID COLECTOMY, CONTROL OF HEMORRHAGE FROM MESENTERY, AND TEMPORARY CLOSURE OF THE ABDOMEN, THE DEVICE MISFIRED.
26-10-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON AN EXPLORATORY LAPAROTOMY FOR TRAUMA WITH SMALL BOWEL RESECTION, PARTIAL SIGMOID COLECTOMY, CONTROL OF HEMORRHAGE FROM MESENTERY, AND TEMPORARY CLOSURE OF THE ABDOMEN, THE DEVICE MISFIRED.
28-12-2022 | Malfunction | covidien lp llc north haven | 17-01-2023 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN THE DEVICE WAS INSERTED IN THE TROCAR INTO THE PATIENT'S CAVITY, THE SUTURE AT THE DISTAL END OF THE CARTRIDGE'S SIDE FELL INTO THE THE PATIENT'S CAVITY. THE SURGEON WAS ABLE TO LOCATE AND USED A GRASPER TO SAFELY RETRIEVED THE SUTURE NEAR THE FUNDUS OF THE STOMACH. THE RELOAD WAS NOT FIRED AND WAS REMOVED FROM THE CAVITY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-01-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: THIS CASE IS FROM HEALTH AUTHORITY. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED . CHANGED THE NEW ONE TO COMPLETE SURGERY. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-01-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | PWRD ECH FLEX 60MM 34CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: THIS CASE IS FROM HEALTH AUTHORITY. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. AND THE STAPLES WAS NORMAL AFTER DISASSEMBLY . NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 17-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, WHILE SEVERING THE BLOOD VESSEL, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD WAS USED TO RESOLVE HE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 17-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHEN AT THE STOMACH CURVATURE, THE DEVICE PARTIALLY FIRED INITIAL TISSUE WHERE AN ATTEMPT WAS MADE. SOME OF THE STAPLES HAVE A POOR FORMATION. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
26-12-2022 | Malfunction | celestica de monterrey s.a de c.v | 17-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, WHEN AT THE STOMACH CURVATURE, THE DEVICE PARTIALLY FIRED INITIAL TISSUE WHERE AN ATTEMPT WAS MADE. SOME OF THE STAPLES HAVE A POOR FORMATION. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
27-12-2022 | Injury | covidien lp llc north haven | 17-01-2023 | EEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING OPEN HEMORRHOIDS SURGERY, THE TISSUE CUT BUT THE DEVICE FAILED TO FIRE STAPLES. AFTER FIRING ATTEMPT, THE ANVIL COULD NOT BE TILTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
27-12-2022 | Injury | covidien lp llc north haven | 17-01-2023 | EEA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING OPEN HEMORRHOIDS SURGERY, THE TISSUE CUT BUT THE DEVICE FAILED TO FIRE STAPLES. AFTER FIRING ATTEMPT, THE ANVIL COULD NOT BE TILTED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE.
02-01-2023 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION PROCEDURE, WHEN RELOAD WAS LOADED AND SURGEON WAS TRYING TO ARTICULATE, THE ARTICULATING LEVER BROKE, SO DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2023 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Mechanics Altered | A0507 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, AT THE TIME OF BOWEL ANASTOMOSIS, THE KNIFE BLADE STILL CUT, BUT NO STAPLES WERE FIRED. THE SURGICAL TIME WAS EXTENDED FOR 45 MINUTES DUE TO THE PRODUCT PROBLEM. THE SURGEON HAD TO HAND SEW TO RESOLVE THE ISSUE.
01-01-2023 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, AT THE TIME OF BOWEL ANASTOMOSIS, THE KNIFE BLADE STILL CUT, BUT NO STAPLES WERE FIRED. THE SURGICAL TIME WAS EXTENDED FOR 45 MINUTES DUE TO THE PRODUCT PROBLEM. THE SURGEON HAD TO HAND SEW TO RESOLVE THE ISSUE.
01-01-2023 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, AT THE TIME OF BOWEL ANASTOMOSIS, THE KNIFE BLADE STILL CUT, BUT NO STAPLES WERE FIRED. THE SURGICAL TIME WAS EXTENDED FOR 45 MINUTES DUE TO THE PRODUCT PROBLEM. THE SURGEON HAD TO HAND SEW TO RESOLVE THE ISSUE.
01-01-2023 | Injury | us surgical puerto rico | 17-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, AT THE TIME OF BOWEL ANASTOMOSIS, THE KNIFE BLADE STILL CUT, BUT NO STAPLES WERE FIRED. THE SURGICAL TIME WAS EXTENDED FOR 45 MINUTES DUE TO THE PRODUCT PROBLEM. THE SURGEON HAD TO HAND SEW TO RESOLVE THE ISSUE.
17-01-2023 | Malfunction | covidien lp llc north haven | 17-01-2023 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: PRODUCT RECEIVED WITHOUT ACCOMPANYING INFORMATION. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT THE ANVIL CLAMPING MECHANISM WAS DEFORMED.
17-01-2023 | Malfunction | covidien lp llc north haven | 17-01-2023 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: PRODUCT RECEIVED WITHOUT ACCOMPANYING INFORMATION. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT THE ANVIL CLAMPING MECHANISM WAS DEFORMED.
29-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC HEPATECTOMY, ON LIVER RESECTION, THE RELOAD DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED AS REPLACEMENT BUT STILL THE RELOAD DID NOT FIRE. ANOTHER RELOAD WAS USED WITH THE NEW HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-12-2022 | Malfunction | davol inc sub c.r bard inc -1213643 | 17-01-2023 | OPTIFIX AT | Device Contamination with Chemical or Other Material | A180104 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, ON 22-DEC-2022 THE OPTIFIX AT FIXATION DEVICE WAS NOTED TO HAVE "PLASTIC STRINGS" COMING OUT OF THE FIXATION DEVICE HANDLE. AS REPORTED THE DEVICE WAS NOT USED ON THE PATIENT.
01-11-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: WE HAVE RECEIVED AN E-MAIL FROM SR GBP MANAGER FOR EM STATING: "AS PART OF OUR INVESTIGATIONS, WE HAVE OBTAINED SAMPLES FROM A SUSPECTED SELLER OF COUNTERFEIT GST RELOADS IN TURKEY (TREND MEDIKAL). PLEASE FIND ATTACHED SAMPLES FOR GST60B & GST60G ¿ I WILL CONFIRM TOTAL NUMBER OF SAMPLES FOR EACH SKU" .
01-01-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | GST GREEN RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: WHILE ATTENDING THE OT WITH DR. (B)(6) AT (B)(6) HOSPITAL, THE FOLLOWING RELOADS HAVE BEEN FOUND AND USED IN THE CASES. THERE WERE NOT LESS THAN 50 RELOADS/CODE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
26-12-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS BECAME NOT TO OPEN WHEN THE ECH60R WAS LOAD AT 2ND FIRING. THE NURSE COMMENTED THAT, IT MIGHT BE NOT CLEANED ENOUGH. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
01-11-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | GST GREEN RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: WE HAVE RECEIVED AN E-MAIL FROM SR GBP MANAGER FOR EM STATING: "AS PART OF OUR INVESTIGATIONS, WE HAVE OBTAINED SAMPLES FROM A SUSPECTED SELLER OF COUNTERFEIT GST RELOADS IN TURKEY ((B)(6) MEDICAL).
01-01-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: WHILE ATTENDING THE OT WITH DR. (B)(6), THE FOLLOWING RELOADS HAVE BEEN FOUND AND USED IN THE CASES. THERE WERE NOT LESS THAN 50 RELOADS/ CODE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
29-12-2022 | Malfunction | us surgical puerto rico | 17-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GENERAL SURGERY, THE DEVICE DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | GST GOLD RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: WHILE ATTENDING THE OT WITH (B)(6), THE FOLLOWING RELOADS HAVE BEEN FOUND AND USED IN THE CASES. THERE WERE NOT LESS THAN 50 RELOADS/ CODE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | ECHELON FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER | Failure to Fire | A050501 | Device Embedded In Tissue or Plaque | Event Description: IN MAIN OR DURING SURGERY, DOCTOR WAS USING ECHELON FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER STAPLER AND THE PRODUCT HAD A MALFUNCTION. WHILE TRYING TO FIRE THE LOAD IT WAS STUCK ON THE TISSUE AND WOULD NOT FIRE OR OPEN THEN WHEN THE DOCTOR WAS ABLE TO GET THE STAPLER WAS BENT INSIDE. NOTIFIED VENDOR.
15-12-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | ECHELON FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER | Difficult to Open or Close | A051104 | Device Embedded In Tissue or Plaque | Event Description: IN MAIN OR DURING SURGERY, DOCTOR WAS USING ECHELON FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER STAPLER AND THE PRODUCT HAD A MALFUNCTION. WHILE TRYING TO FIRE THE LOAD IT WAS STUCK ON THE TISSUE AND WOULD NOT FIRE OR OPEN THEN WHEN THE DOCTOR WAS ABLE TO GET THE STAPLER WAS BENT INSIDE. NOTIFIED VENDOR.
15-12-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | ECHELON FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER | Material Twisted/Bent | A040609 | Device Embedded In Tissue or Plaque | Event Description: IN MAIN OR DURING SURGERY, DOCTOR WAS USING ECHELON FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER STAPLER AND THE PRODUCT HAD A MALFUNCTION. WHILE TRYING TO FIRE THE LOAD IT WAS STUCK ON THE TISSUE AND WOULD NOT FIRE OR OPEN THEN WHEN THE DOCTOR WAS ABLE TO GET THE STAPLER WAS BENT INSIDE. NOTIFIED VENDOR.
15-12-2022 | Malfunction | ethicon endo-surgery | 17-01-2023 | ECHELON FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER | Mechanical Jam | A0506 | Device Embedded In Tissue or Plaque | Event Description: IN MAIN OR DURING SURGERY, DOCTOR WAS USING ECHELON FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER STAPLER AND THE PRODUCT HAD A MALFUNCTION. WHILE TRYING TO FIRE THE LOAD IT WAS STUCK ON THE TISSUE AND WOULD NOT FIRE OR OPEN THEN WHEN THE DOCTOR WAS ABLE TO GET THE STAPLER WAS BENT INSIDE. NOTIFIED VENDOR.
30-12-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVE OF A LAPAROSCOPIC SIGMOID RESECTION WHERE THE ANVIL USED IN THE COLON DETACHED AND DISENGAGED DURING DEVICE EXTRACTION, FELL INTO THE CAVITY, AND RETRIEVED WITH A GRASPER AND SCOPE, THE PATIENT HAD A HEART ISSUE. THE PATIENT DIED. THE DEATH OF THE PATIENT WAS UNRELATED TO THE PROCEDURE. THERE WAS NO INTRA-OPERATIVE PATIENT INJURY. THE ANASTOMOSIS, LEAK TEST, AND DOUGHNUTS SENT TO PATHOLOGY HAD NO ISSUES.
30-12-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | EEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVE OF A LAPAROSCOPIC SIGMOID RESECTION WHERE THE ANVIL USED IN THE COLON DETACHED AND DISENGAGED DURING DEVICE EXTRACTION, FELL INTO THE CAVITY, AND RETRIEVED WITH A GRASPER AND SCOPE, THE PATIENT HAD A HEART ISSUE. THE PATIENT DIED. THE DEATH OF THE PATIENT WAS UNRELATED TO THE PROCEDURE. THERE WAS NO INTRA-OPERATIVE PATIENT INJURY. THE ANASTOMOSIS, LEAK TEST, AND DOUGHNUTS SENT TO PATHOLOGY HAD NO ISSUES.
30-03-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE RELOAD DID NOT CLOSE FIRMLY WHEN STAPLING THE TISSUE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-04-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, THE RELOAD DID NOT CLOSE FIRMLY FOR STAPLING THE TISSUE. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
15-12-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | UNKNOWN GIA DST PRODUCT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INVOLVEMENT.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE RELOAD WAS DIFFICULT TO LOAD AND THE DEVICE DID NOT FIRE; THE RELOAD WAS PRE-SHOT, THUS IT DID NOT ADVANCE THE STAPLE LINE. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUES OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE RELOAD WAS DIFFICULT TO LOAD AND THE DEVICE DID NOT FIRE; THE RELOAD WAS PRE-SHOT, THUS IT DID NOT ADVANCE THE STAPLE LINE. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUES OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE RELOAD WAS DIFFICULT TO LOAD AND THE DEVICE DID NOT FIRE; THE RELOAD WAS PRE-SHOT, THUS IT DID NOT ADVANCE THE STAPLE LINE. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUES OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE DEVICE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE DEVICE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE DEVICE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE DEVICE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-10-2022 | Malfunction | ethicon endo-surgery | 18-01-2023 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE BLADDER PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, AFTER OPENING THE DEVICE, THE CARTRIDGE WAS STUCK AND NOT COMING OUT ON TWO RELOADS. A RELOAD DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
04-12-2022 | Malfunction | ethicon endo-surgery | 18-01-2023 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RECTUM SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-09-2022 | Malfunction | covidien lp llc north haven | 18-01-2023 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE IS SHOWING THE RED CIRCLE WHICH INDICATES A PROBLEM WITH THE HANDLE. TO RESOLVE THE ISSUE THE SALE REPRESENTATIVE FIRST TRIED TO RESET IT BY HOLDING DOWN BUTTONS, RESET TO SAME ERROR. THEN PULLED THE HANDLE INTO VLEX TO SEE THE LOGS. THE HANDLE IMMEDIATELY STATED IT NEEDED TO BE UPGRADED TO NEWEST SOFTWARE, WHEN REPRESENTATIVE AGREED TO INSTALL, IT TRIED BUT TIMED OUT AT 35%. THERE WAS NO PATIENT INVOLVEMENT. PLI-10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THE CAUSE OF THE REPORTED CONDITION OF HANDLE ERROR SCREEN IS DUE TO FIELD PROGRAMMABLE GATE ARRAY (FPGA) BEING UNABLE TO RESET CAUSING THE MOTOR TEST TO FAIL. THE CAUSE OF THE FPGA REPROGRAM/RESETS CAN BE TRACED TO HANDLE SOFTWARE.
08-09-2022 | Malfunction | covidien lp llc north haven | 18-01-2023 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE IS SHOWING THE RED CIRCLE WHICH INDICATES A PROBLEM WITH THE HANDLE. TO RESOLVE THE ISSUE THE SALE REPRESENTATIVE FIRST TRIED TO RESET IT BY HOLDING DOWN BUTTONS, RESET TO SAME ERROR. THEN PULLED THE HANDLE INTO VLEX TO SEE THE LOGS. THE HANDLE IMMEDIATELY STATED IT NEEDED TO BE UPGRADED TO NEWEST SOFTWARE, WHEN REPRESENTATIVE AGREED TO INSTALL, IT TRIED BUT TIMED OUT AT 35%. THERE WAS NO PATIENT INVOLVEMENT. PLI-10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THE CAUSE OF THE REPORTED CONDITION OF HANDLE ERROR SCREEN IS DUE TO FIELD PROGRAMMABLE GATE ARRAY (FPGA) BEING UNABLE TO RESET CAUSING THE MOTOR TEST TO FAIL. THE CAUSE OF THE FPGA REPROGRAM/RESETS CAN BE TRACED TO HANDLE SOFTWARE.
08-09-2022 | Malfunction | covidien lp llc north haven | 18-01-2023 | SIGNIA | Application Program Version or Upgrade Problem | A110206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE IS SHOWING THE RED CIRCLE WHICH INDICATES A PROBLEM WITH THE HANDLE. TO RESOLVE THE ISSUE THE SALE REPRESENTATIVE FIRST TRIED TO RESET IT BY HOLDING DOWN BUTTONS, RESET TO SAME ERROR. THEN PULLED THE HANDLE INTO VLEX TO SEE THE LOGS. THE HANDLE IMMEDIATELY STATED IT NEEDED TO BE UPGRADED TO NEWEST SOFTWARE, WHEN REPRESENTATIVE AGREED TO INSTALL, IT TRIED BUT TIMED OUT AT 35%. THERE WAS NO PATIENT INVOLVEMENT. PLI-10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THE CAUSE OF THE REPORTED CONDITION OF HANDLE ERROR SCREEN IS DUE TO FIELD PROGRAMMABLE GATE ARRAY (FPGA) BEING UNABLE TO RESET CAUSING THE MOTOR TEST TO FAIL. THE CAUSE OF THE FPGA REPROGRAM/RESETS CAN BE TRACED TO HANDLE SOFTWARE.
03-01-2023 | Injury | us surgical puerto rico | 18-01-2023 | EEA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOWER ANTERIOR RESECTION, DURING END-TO-END ANASTOMOSIS, THE HANDLE WAS FULLY SQUEEZED BUT THERE WAS POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE. THE DONUTS WERE INCOMPLETE. THE STAPLE LINE WAS RESECTED AND RE-STAPLED WITH A NEW STAPLER TO FIX THE ISSUE.
03-01-2023 | Injury | us surgical puerto rico | 18-01-2023 | EEA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOWER ANTERIOR RESECTION, DURING END-TO-END ANASTOMOSIS, THE HANDLE WAS FULLY SQUEEZED BUT THERE WAS POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE. THE DONUTS WERE INCOMPLETE. THE STAPLE LINE WAS RESECTED AND RE-STAPLED WITH A NEW STAPLER TO FIX THE ISSUE.
03-01-2023 | Injury | us surgical puerto rico | 18-01-2023 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOWER ANTERIOR RESECTION, DURING END-TO-END ANASTOMOSIS, THE HANDLE WAS FULLY SQUEEZED BUT THERE WAS POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE. THE DONUTS WERE INCOMPLETE. THE STAPLE LINE WAS RESECTED AND RE-STAPLED WITH A NEW STAPLER TO FIX THE ISSUE.
03-01-2023 | Injury | us surgical puerto rico | 18-01-2023 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC LOWER ANTERIOR RESECTION, DURING END-TO-END ANASTOMOSIS, THE HANDLE WAS FULLY SQUEEZED BUT THERE WAS POOR STAPLE FORMATION AND INCOMPLETE STAPLE LINE. THE DONUTS WERE INCOMPLETE. THE STAPLE LINE WAS RESECTED AND RE-STAPLED WITH A NEW STAPLER TO FIX THE ISSUE.
01-01-2023 | Malfunction | ethicon endo-surgery | 18-01-2023 | PROXIMATE*RL LINEAR STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED BY THE SALES REP THAT DURING A LOW ANTERIOR PROCEDURE THE SURGEON REMOVED THE BLUE RELOAD AND RELOADED WITH A GREEN RELOAD. THE SURGEON FIRED ONCE AND THEN RELOADED AND FIRED AGAIN. THE SURGEON FIRED A THIRD TIME AND THAT'S WHEN THE SURGEON NOTICED NONE OF THE STAPLES WERE FORMED. THEY ENDED UP OPENING A CONTOUR DEVICE FOR THE LAST FIRING. THEY COMPLETED THE ANASTOMOSIS IN A NORMAL FASHION. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE POWERED HANDLE WAS NOT FIRING. IT WAS STATED THAT ONCE THE STAPLER WAS ENGAGED IT TURNED GREEN, BUT NEVER BLINKED OR FIRED. THE SALES REPRESENTATIVE TESTED THE HANDLE AND IT FIRED. THERE WERE NO FURTHER ISSUES. THERE WAS NO PATIENT INVOLVEMENT.
23-12-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL PULMONARY RESECTION, WHILE AT PULMONARY RESECTION, WHEN RELOAD WAS CONNECTED TO HANDLE, A CRACK SOUND CAME FROM THE STAPLER HANDLE, AND THE HANDLE COULD NOT BE GRIPPED ANYMORE. IT WAS NOTED THAT THE INTERNAL GEAR WAS BROKEN, BUT THE DEVICE WAS ABLE TO FIRE. THE HANDLE AND RELOAD WERE REPLACED WITH A NEW ONE, AND THE SURGERY WAS COMPLETED AFTERWARDS. THERE WAS NO PATIENT INJURY.
23-12-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL PULMONARY RESECTION, WHILE AT PULMONARY RESECTION, WHEN RELOAD WAS CONNECTED TO HANDLE, A CRACK SOUND CAME FROM THE STAPLER HANDLE, AND THE HANDLE COULD NOT BE GRIPPED ANYMORE. IT WAS NOTED THAT THE INTERNAL GEAR WAS BROKEN, BUT THE DEVICE WAS ABLE TO FIRE. THE HANDLE AND RELOAD WERE REPLACED WITH A NEW ONE, AND THE SURGERY WAS COMPLETED AFTERWARDS. THERE WAS NO PATIENT INJURY.
17-08-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY PROCEDURE, WHILE BEING APPLIED AT THE RECTUM, AFTER THE RELOAD WAS USED, THE SURGEON COULD NOT UNLOAD IT FROM THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WAS SHEARED TEETH ON THE FIRING RACK.
17-08-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY PROCEDURE, WHILE BEING APPLIED AT THE RECTUM, AFTER THE RELOAD WAS USED, THE SURGEON COULD NOT UNLOAD IT FROM THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WAS SHEARED TEETH ON THE FIRING RACK.
21-12-2022 | Malfunction | covidien lp llc north haven | 18-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY), WHILE RESECTING THE PULMONARY VEIN, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED, BUT THE DEVICE DID NOT FIRE. A NEW RELOAD OF THE SAME KIND WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2022 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE TESTING THE DEVICE, THE DEVICE WAS BLOCKED AND WAS UNABLE TO BE USED.
09-12-2022 | Malfunction | covidien lp llc north haven | 18-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE TESTING THE DEVICE, THE DEVICE WAS BLOCKED AND WAS UNABLE TO BE USED.
02-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR) PROCEDURE, WHILE TRANSECTING THE RECTUM, THE SURGEON PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE, BUT THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2022 | Malfunction | ethicon endo-surgery | 18-01-2023 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOWER ANTERIOR BOWEL RESECTION, THE DEVICE MISFIRED. THE SURGICAL PROCEDURE WAS DELAYED. NO FRAGMENTS WERE GENERATED, AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY.
18-12-2022 | Malfunction | ethicon endo-surgery | 18-01-2023 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOWER ANTERIOR BOWEL RESECTION, THE DEVICE MISFIRED. THE SURGICAL PROCEDURE WAS DELAYED. NO FRAGMENTS WERE GENERATED, AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, WHILE RESECTIONING THE BODY OF THE STOMACH, THE SURGEON WAS ABLE TO FIRE THE FIRST TWO RELOADS. WHEN THE THIRD RELOAD WAS FIRED, THE HANDLE MADE A NOISE LIKE IT WAS BROKEN AND THE DEVICE DID NOT FIRE. A SECOND HANDLE WAS USED WITH THE THIRD RELOAD, BUT THE SAME ISSUES OCCURRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE ON BOTH HANDLES. A THIRD HANDLE AND A NEW RELOAD WERE USED TO COMPLETE THE PROCEDURE. THE SURGEON NOTED THAT THE THIRD HANDLE'S ROTATION WAS VERY STIFF. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, WHILE RESECTIONING THE BODY OF THE STOMACH, THE SURGEON WAS ABLE TO FIRE THE FIRST TWO RELOADS. WHEN THE THIRD RELOAD WAS FIRED, THE HANDLE MADE A NOISE LIKE IT WAS BROKEN AND THE DEVICE DID NOT FIRE. A SECOND HANDLE WAS USED WITH THE THIRD RELOAD, BUT THE SAME ISSUES OCCURRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE ON BOTH HANDLES. A THIRD HANDLE AND A NEW RELOAD WERE USED TO COMPLETE THE PROCEDURE. THE SURGEON NOTED THAT THE THIRD HANDLE'S ROTATION WAS VERY STIFF. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, WHILE RESECTIONING THE BODY OF THE STOMACH, THE SURGEON WAS ABLE TO FIRE THE FIRST TWO RELOADS. WHEN THE THIRD RELOAD WAS FIRED, THE HANDLE MADE A NOISE LIKE IT WAS BROKEN AND THE DEVICE DID NOT FIRE. A SECOND HANDLE WAS USED WITH THE THIRD RELOAD, BUT THE SAME ISSUES OCCURRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE ON BOTH HANDLES. A THIRD HANDLE AND A NEW RELOAD WERE USED TO COMPLETE THE PROCEDURE. THE SURGEON NOTED THAT THE THIRD HANDLE'S ROTATION WAS VERY STIFF. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, WHILE RESECTIONING THE BODY OF THE STOMACH, THE SURGEON WAS ABLE TO FIRE THE FIRST TWO RELOADS. WHEN THE THIRD RELOAD WAS FIRED, THE HANDLE MADE A NOISE LIKE IT WAS BROKEN AND THE DEVICE DID NOT FIRE. A SECOND HANDLE WAS USED WITH THE THIRD RELOAD, BUT THE SAME ISSUES OCCURRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE ON BOTH HANDLES. A THIRD HANDLE AND A NEW RELOAD WERE USED TO COMPLETE THE PROCEDURE. THE SURGEON NOTED THAT THE THIRD HANDLE'S ROTATION WAS VERY STIFF. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SUBTOTAL GASTRECTOMY, WHILE RESECTIONING THE BODY OF THE STOMACH, THE SURGEON WAS ABLE TO FIRE THE FIRST TWO RELOADS. WHEN THE THIRD RELOAD WAS FIRED, THE HANDLE MADE A NOISE LIKE IT WAS BROKEN AND THE DEVICE DID NOT FIRE. A SECOND HANDLE WAS USED WITH THE THIRD RELOAD, BUT THE SAME ISSUES OCCURRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE ON BOTH HANDLES. A THIRD HANDLE AND A NEW RELOAD WERE USED TO COMPLETE THE PROCEDURE. THE SURGEON NOTED THAT THE THIRD HANDLE'S ROTATION WAS VERY STIFF. THERE WAS NO PATIENT INJURY.
19-12-2022 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE KNIFE DID NOT GO THROUGH AND STOPPED THE CARTRIDGE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE DEVICE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE DEVICE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE DEVICE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | covidien lp llc north haven | 18-01-2023 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BARIATRIC PROCEDURE, WHILE MAKING THE NEW STOMACH, THE DEVICE STOPPED FIRING MIDWAY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY PROCEDURE, WHILE DETACHING THE LUNG PARENCHYMA, THE DEVICE WAS UNABLE TO BE FIRED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y PROCEDURE, WHILE STAPLING OF THE FUNDUS OF STOMACH, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. A NEW LOAD WAS USED AFTER LOAD DID NOT WORK ON NEW HAND PIECE TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 18-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROUX-EN-Y PROCEDURE, WHILE STAPLING OF THE FUNDUS OF STOMACH, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. A NEW LOAD WAS USED AFTER LOAD DID NOT WORK ON NEW HAND PIECE TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 18-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SLEEVE GASTRECTOMY PROCEDURE THAT DELIVERY BOX AND THE PRODUCT BOX HAD WET DRY MARKS ON IT. THE DOCTOR FELT THE DEVICE WAS UNSTERILE SO HE DID NOT USE THE PRODUCT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. ONE DEVICE WILL BE RETURNED.
22-08-2022 | Malfunction | ethicon endo-surgery | 18-01-2023 | PWRD ECH FLEX 60MM 28CM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: THIS CASE IS FROM HEALTH AUTHORITY. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW, AND A SMALL SPRING FELL FROM THE GAP OF THE RED EXCITATION KEY . CHANGED THE NEW ONE TO COMPLETE SURGERY. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-08-2022 | Malfunction | ethicon endo-surgery | 18-01-2023 | PWRD ECH FLEX 60MM 28CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: THIS CASE IS FROM HEALTH AUTHORITY. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW, AND A SMALL SPRING FELL FROM THE GAP OF THE RED EXCITATION KEY . CHANGED THE NEW ONE TO COMPLETE SURGERY. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2022 | Injury | intuitive surgical | 18-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, STAPLE LINE BLEEDING ON STOMACH TISSUE WAS OBSERVED WHERE A BLUE SUREFORM 60 RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. ACCORDING TO THE SURGEON, THERE WAS NOTHING WRONG WITH THE STAPLE LINE, AS THERE WERE NO MALFORMED STAPLES, NO HOLES WITHIN THE STAPLE LINE, AND NO TISSUE PUSHOUT OBSERVED. THE BLEEDING WAS MORE THAN USUAL AND CLIPS WERE APPLIED TO STOP THE BLEEDING. NO OTHER MEDICAL INTERVENTION WAS NEEDED. THE PATIENT HAS BEEN RECOVERING WELL WITH NO POST-OPERATIVE COMPLICATIONS.
19-12-2022 | Malfunction | intuitive surgical | 19-01-2023 | SUREFORM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LUNG SURGICAL PROCEDURE, THE SURGEON FIRED A SUREFORM 45 STAPLER INSTRUMENT, AND THE STAPLER INSTRUMENT WOULD NOT RELEASE. AND THE RELOAD WAS JAMMED IN DEVICE. THE CUSTOMER WAS EVENTUALLY ABLE TO GET IT OUT. THERE WAS A PROCEDURE DELAY IN SURGEON TO HANDLE MALFUNCTION DEVICE. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION: THE INSTRUMENT WAS BRAND NEW TAKEN OUT OF THE BOX INSERTED IN DA VINCI AND USED WHEN IT MISS FIRED AND JAMMED. UNSURE IT WAS THE APPROPRIATE RELOAD FOR DA VINCI LUNG CASES AS THE SURGEON IS QUALIFIED AND DOES MANY OF THESE PROCEDURES WITH THIS DEVICE. STAPLING OF LUNG TISSUE. UNABLE TO PROVIDE STAPLER INFORMATION. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. NO BUTTRESS MATERIAL WAS USED. THERE WAS NO INJURY OR ISSUE TO THE PATIENT OR SURGICAL FIELD. THE PROCEDURE WAS COMPLETED WITH NO INCIDENT. AND NO PATIENT INJURY. ISI FOLLOWED UP WITH THE INITIAL REPORTER AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION: THE COLOR OF THE RELOAD JAMMED WHEN THE EVENT OCCURRED WAS BLACK. NO THE RELOAD WAS NOT ABLE TO COMPLETE FIRING. YES THERE WERE MALFORMED STAPLES. THERE WAS UNSUCCESSFUL FIRE AND COULD NOT RELOAD. WHEN THE USER REMOVED THE INSTRUMENT, YES WITH DIFFICULTY, THE RELOAD WAS ALSO REMOVED FROM THE INSTRUMENT. THE SURGICAL TASK WAS COMPLETED WITH A NEW STAPLER SINCE THE UNABLE TO RELOAD.
04-01-2023 | Malfunction | covidien lp llc north haven | 19-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON LOW ANTERIOR RESECTION PROCEDURE, WHILE TRANSECTING RECTUM, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND HANDLE WAS NOT SQUEEZED. THE TWO HANDLES AND TWO RELOADS THAT DID NOT FIRE WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VERTICAL SLEEVE GASTRECTOMY, DURING THE SECOND SHOT, THE TRIGGER WAS SQUEEZED COMPL ETELY, AND THE GREEN BUTTON WAS PRESSED TO PREPARE THE SHOT, BUT THE DEVICE DID NOT FIRE. THE THICKNESS OF THE TISSUE WAS VERIFIED. THE RELOAD WAS EXCHANGED FROM LARGER TO SMALLER STAPLES, BUT THE SAME ISSUE OCCURRED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, WHILE CLOSING THE STAPLER, THE STAPLER DID NOT FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS SQUEEZED. A SECOND HANDLE WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2022 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE RELOAD JAWS WOULD NOT FULLY CLOSED. A DIFFERENT TYPE AND DIFFERENT SIZES OF RELOADS WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-10-2022 | Malfunction | covidien lp llc north haven | 19-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEVICE FOLLOW UP, THE DEMO HANDLE FAILED TO CHARGE. IT NO LONGER LIGHTS UP AND THE SCREEN REMAINED BLACK. THE CHARGER WORKED FINE ON OTHER HANDLES. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
11-10-2022 | Malfunction | covidien lp llc north haven | 19-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DEVICE FOLLOW UP, THE DEMO HANDLE FAILED TO CHARGE. IT NO LONGER LIGHTS UP AND THE SCREEN REMAINED BLACK. THE CHARGER WORKED FINE ON OTHER HANDLES. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
06-12-2022 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE PROCEDURE, THE RELOAD JAWS WOULD NOT FULLY CLOSED. A DIFFERENT TYPE AND DIFFERENT SIZES OF RELOADS WERE USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-10-2022 | Malfunction | covidien lp llc north haven | 19-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE USER MOVED ALL THE HANDLES AND CHARGERS TO A NEW LOCATION. WHEN MOVED, ALL THE FOUR HANDLES FLASHED A RED LIGHT ERROR AND WOULD NOT TURN ON. TROUBLESHOOTING WAS DONE BUT STILL DID NOT RESOLVE THE ISSUE. THE SALES REP TRIED TO CHARGE A PERSONAL HANDLE AND IT WORKED FINE WITH ALL THE CHARGERS. NO PATIENT INVOLVED.
25-10-2022 | Malfunction | covidien lp llc north haven | 19-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE USER MOVED ALL THE HANDLES AND CHARGERS TO A NEW LOCATION. WHEN MOVED, ALL THE FOUR HANDLES FLASHED A RED LIGHT ERROR AND WOULD NOT TURN ON. TROUBLESHOOTING WAS DONE BUT STILL DID NOT RESOLVE THE ISSUE. THE SALES REP TRIED TO CHARGE A PERSONAL HANDLE AND IT WORKED FINE WITH ALL THE CHARGERS. NO PATIENT INVOLVED.
19-12-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE KNIFE WAS NOT RETRIEVED AFTER FIRING. NO OTHER DETAILS PROVIDED. NO PATIENT CONSEQUENCES REPORTED.
03-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE.
03-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE.
03-10-2022 | Malfunction | us surgical puerto rico | 19-01-2023 | EEA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A ROBOTIC LAPAROSCOPIC SIGMOID RESECTION, AFTER THE INTESTINAL MEMBRANE TREATMENT OF THE ANAL COLON WITH A SEALING DEVICE, AND AFTER A CLOSED DISSECTION WITH AN ENDOSCOPIC STAPLER RELOAD, THE ANVIL OF A CIRCULAR STAPLER WAS INSERTED TRANSANALLY. WHEN THE WING NUT WAS ROTATED FOR DOCKING AND EXTENDING THE CENTER ROD, IT WAS DIFFICULT TO ROTATE. THE DEVICE WAS FORCED AND THE CENTER ROD EXTENDED BUT THE HANDLE WAS STOPPED FROM BEING USED. THE HANDLE WAS REPLACED WITH A NEW ONE AND THE SAME ANVIL WAS USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE WING NUT WAS BINDING WHEN THE CENTER ROD WAS ALMOST EXTENDED AND WHEN IT RETRACTED.
03-10-2022 | Malfunction | us surgical puerto rico | 19-01-2023 | EEA | Positioning Problem | A1502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A ROBOTIC LAPAROSCOPIC SIGMOID RESECTION, AFTER THE INTESTINAL MEMBRANE TREATMENT OF THE ANAL COLON WITH A SEALING DEVICE, AND AFTER A CLOSED DISSECTION WITH AN ENDOSCOPIC STAPLER RELOAD, THE ANVIL OF A CIRCULAR STAPLER WAS INSERTED TRANSANALLY. WHEN THE WING NUT WAS ROTATED FOR DOCKING AND EXTENDING THE CENTER ROD, IT WAS DIFFICULT TO ROTATE. THE DEVICE WAS FORCED AND THE CENTER ROD EXTENDED BUT THE HANDLE WAS STOPPED FROM BEING USED. THE HANDLE WAS REPLACED WITH A NEW ONE AND THE SAME ANVIL WAS USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE WING NUT WAS BINDING WHEN THE CENTER ROD WAS ALMOST EXTENDED AND WHEN IT RETRACTED.
05-01-2023 | Injury | us surgical puerto rico | 19-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHILE STAPLING, THERE WAS A PROBLEM LOADING THEN DEVICE WHEN THE LOAD DID NOT FIT THE STAPLER, THE DEVICE DID FIRE AND THE STAPLE LINE WAS INCOMPLETE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
05-01-2023 | Injury | us surgical puerto rico | 19-01-2023 | ENDO GIA | Component or Accessory Incompatibility | A1701 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHILE STAPLING, THERE WAS A PROBLEM LOADING THEN DEVICE WHEN THE LOAD DID NOT FIT THE STAPLER, THE DEVICE DID FIRE AND THE STAPLE LINE WAS INCOMPLETE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
04-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE-PORT THORACOSCOPIC RESECTION OF THE LEFT UPPER LOBE OF THE LUNG, WHILE DETACHING THE UPPER LOBE VEIN, THE BLACK KNOB FELL OFF. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-02-2022 | Malfunction | covidien lp llc north haven | 19-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 TO 2021 REGARDING A STUDY TO COMPARE TUMOR DISTANCE TO MARGIN ON 56 PATIENTS AFTER ROBOTIC VS. OPEN HEPATECTOMY FOR COLORECTAL LIVER METASTASES (CLM) AND TO DETERMINE THE RELATIONSHIP BETWEEN PERIOPERATIVE VARIABLES, SURGICAL APPROACH AND TUMOR DISTANCE TO MARGIN WITH OVERALL SURVIVAL PARENCHYMAL. PATIENTS WERE RANDOMIZED IN 2 GROUPS, WITH 42 PATIENTS FOR ROBOTIC AND 14 PATIENTS FOR OPEN PROCEDURE. THE ASSISTANT USED SEALER TO ACHIEVE HEMOSTASIS FROM THE LIVER SURFACE AS THE PARENCHYMAL TRANSECTION WAS BEING CARRIED OUT. SMALL CROSSING VESSELS LESS THAN 7 MM WERE DIVIDED WITH SEALING DEVICE, WHILE LARGE CROSSING VESSELS MORE THAN 7 MM WERE HANDLED WITH 3¿0 SILK SEALING DEVICE OR VASCULAR STAPLERS. THERE WAS BLEEDING COMPLICATION AND ESTIMATED BLOOD LOSS WAS 200 ML.
07-10-2022 | Malfunction | covidien lp llc north haven | 19-01-2023 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREAS TAIL RESECTION, THE ADAPTER WAS DIFFICULT TO CONNECT INTO THE CLAMSHELL. A NEW MANUAL STAPLER WAS USED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INJURY.
07-10-2022 | Malfunction | covidien lp llc north haven | 19-01-2023 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREAS TAIL RESECTION, THE ADAPTER WAS DIFFICULT TO CONNECT INTO THE CLAMSHELL. A NEW MANUAL STAPLER WAS USED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THE STAPLE LINE WAS INTACT ON ONE SIDE OF THE DIVIDED ARTERIES (KIDNEY SPECIMEN SIDE), WHEREAS IT WAS DEFICIENT OF THE PATIENT RENAL ARTERY STUMP SIDE. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED;THE TOP OF THE CLIP WAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THE STAPLE LINE WAS INTACT ON ONE SIDE OF THE DIVIDED ARTERIES (KIDNEY SPECIMEN SIDE), WHEREAS IT WAS DEFICIENT OF THE PATIENT RENAL ARTERY STUMP SIDE. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED;THE TOP OF THE CLIP WAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Product Quality Problem | A0201 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THE STAPLE LINE WAS INTACT ON ONE SIDE OF THE DIVIDED ARTERIES (KIDNEY SPECIMEN SIDE), WHEREAS IT WAS DEFICIENT OF THE PATIENT RENAL ARTERY STUMP SIDE. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED;THE TOP OF THE CLIP WAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Product Quality Problem | A0201 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THE STAPLE LINE WAS INTACT ON ONE SIDE OF THE DIVIDED ARTERIES (KIDNEY SPECIMEN SIDE), WHEREAS IT WAS DEFICIENT OF THE PATIENT RENAL ARTERY STUMP SIDE. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED;THE TOP OF THE CLIP WAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THE STAPLE LINE WAS INTACT ON ONE SIDE OF THE DIVIDED ARTERIES (KIDNEY SPECIMEN SIDE), WHEREAS IT WAS DEFICIENT OF THE PATIENT RENAL ARTERY STUMP SIDE. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED;THE TOP OF THE CLIP WAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THE STAPLE LINE WAS INTACT ON ONE SIDE OF THE DIVIDED ARTERIES (KIDNEY SPECIMEN SIDE), WHEREAS IT WAS DEFICIENT OF THE PATIENT RENAL ARTERY STUMP SIDE. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED;THE TOP OF THE CLIP WAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THE STAPLE LINE WAS INTACT ON ONE SIDE OF THE DIVIDED ARTERIES (KIDNEY SPECIMEN SIDE), WHEREAS IT WAS DEFICIENT OF THE PATIENT RENAL ARTERY STUMP SIDE. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED;THE TOP OF THE CLIP WAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THE STAPLE LINE WAS INTACT ON ONE SIDE OF THE DIVIDED ARTERIES (KIDNEY SPECIMEN SIDE), WHEREAS IT WAS DEFICIENT OF THE PATIENT RENAL ARTERY STUMP SIDE. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED;THE TOP OF THE CLIP WAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Insufficient Information | E2401 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-01-2023 | Injury | covidien lp llc north haven | 19-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LIVER DONOR NEPHRECTOMY, UPON THE STAPLING OF THE ARTERY FROM THE KIDNEY, THE STAPLE LINE WAS INCOMPLETE AND HAD CAUGHT ONTO CLIP. THERE WAS AN ARTERIAL BLEEDING AND THE PATIENT HAD TO BE OPENED FROM A LAPAROSCOPIC CASE TO MANAGE THE BLEEDING AND REMOVING THE KIDNEY OUT. THERE WAS A CLIP NEXT TO THE AREA THAT BEING STAPLED; THE TOP OF THE CLIPWAS VISIBLE ON THE SCREEN PRIOR TO STAPLING BUT THE BACK WAS NOT. WHEN TRYING TO REMOVE THE STAPLER, THE CLIP WAS ALSO BEING REMOVED IN THE JAWS OF THE RELOAD. BASED ON THE FIRST IMAGE, THE KNIFE BLADE WAS SKEWED DUE TO CATCHING ONTO SOMETHING AND A DENT WAS SEEN IN THE STAPLER. ON THE SECOND IMAGE, WHEN THE KNIFE BLADE WAS REMOVED WITH TISSUE, A SMALL PLASTIC PART WAS JAMMED BETWEEN THE JAWS, CAUSING THE KNIFE NOT ABLE TO GO FURTHER. BLOOD LOSS WAS MORE THAN 500CC AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
28-12-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY AFTER THE FIRST FIRING THE ANVIL WAS BENT/DEFLECTED UPWARD SO THE DEVICE WOULD NOT FIT THROUGH THE TROCAR. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
28-12-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, A PSEE60A WAS USED. GOOD SUCCESS OF THE FIRST SHOT ON TERMINAL ILEUS. DURING THE SECOND SHOT ON THE TRANSVERSE COLON AT THE TIME OF THE OPERATION OF THE GRID, THE OPERATORS HEARD A VERY SHORT START OF THE MACHINE WHICH THEN WENT INTO LOCKDOWN. SUBSEQUENTLY, AN ATTEMPT WAS MADE TO CONTACT THE SPECIALISTS TO COMMUNICATE THE BLOCK AND UNDERSTAND THE PROCEDURE TO BE ACTIVATED, WITHOUT SUCCESS. OPENED THE UPPER DOOR FOR MANUAL UNLOCKING ALSO WITHOUT SUCCESS. FINALLY HEARD BY PHONE THE SPECIALIST BY REMOVING AND REINSERTING THE BATTERY IT WAS POSSIBLE TO REOPEN THE STAPLER. THE INTERVENTION ENDED USING A NEW PSEE60A. NO CONSEQUENCE FOR THE PATIENT.
28-12-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, A PSEE60A WAS USED. GOOD SUCCESS OF THE FIRST SHOT ON TERMINAL ILEUS. DURING THE SECOND SHOT ON THE TRANSVERSE COLON AT THE TIME OF THE OPERATION OF THE GRID, THE OPERATORS HEARD A VERY SHORT START OF THE MACHINE WHICH THEN WENT INTO LOCKDOWN. SUBSEQUENTLY, AN ATTEMPT WAS MADE TO CONTACT THE SPECIALISTS TO COMMUNICATE THE BLOCK AND UNDERSTAND THE PROCEDURE TO BE ACTIVATED, WITHOUT SUCCESS. OPENED THE UPPER DOOR FOR MANUAL UNLOCKING ALSO WITHOUT SUCCESS. FINALLY HEARD BY PHONE THE SPECIALIST BY REMOVING AND REINSERTING THE BATTERY IT WAS POSSIBLE TO REOPEN THE STAPLER. THE INTERVENTION ENDED USING A NEW PSEE60A. NO CONSEQUENCE FOR THE PATIENT.
28-12-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, A PSEE60A WAS USED. GOOD SUCCESS OF THE FIRST SHOT ON TERMINAL ILEUS. DURING THE SECOND SHOT ON THE TRANSVERSE COLON AT THE TIME OF THE OPERATION OF THE GRID, THE OPERATORS HEARD A VERY SHORT START OF THE MACHINE WHICH THEN WENT INTO LOCKDOWN. SUBSEQUENTLY, AN ATTEMPT WAS MADE TO CONTACT THE SPECIALISTS TO COMMUNICATE THE BLOCK AND UNDERSTAND THE PROCEDURE TO BE ACTIVATED, WITHOUT SUCCESS. OPENED THE UPPER DOOR FOR MANUAL UNLOCKING ALSO WITHOUT SUCCESS. FINALLY HEARD BY PHONE THE SPECIALIST BY REMOVING AND REINSERTING THE BATTERY IT WAS POSSIBLE TO REOPEN THE STAPLER. THE INTERVENTION ENDED USING A NEW PSEE60A. NO CONSEQUENCE FOR THE PATIENT.
05-10-2022 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOW ANTERIOR RESECTION, WHILE BEING APPLIED AT THE RECTUM, THE STAPLER DID NOT FIRE. THE RELOAD WAS LOADED ONTO THE HANDLE. THE SURGEON PLACED THE JAWS OF THE RELOAD OVER THE RECTUM TO PROCEED TO RESECT. SURGEON WAS ABLE TO CLOSE THE JAW OF THE RELOAD, PRESSED THE GREEN BUTTON AND READY TO FIRE. THE HANDLE ADVANCED FORWARD TO THE READY TO FIRE POSITION, BUT ONCE THE SURGEON ATTEMPTED TO FIRE, THE HANDLE HAD NO RESISTANCE AND THE STAPLES DID NOT DEPLOY FROM THE RELOAD. IT WAS ALSO STATED THAT THE JAWS OF THE DEVICE LOCK ON TISSUE BUT IT WAS REMOVED BY USING THE BLACK RESET BUTTON. ANOTHER HANDLE WAS USED BUT HAD THE SAME ISSUE AS THE FIRST. A NEW DEVICE FROM OTHER MANUFACTURER WAS THEN USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-07-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
04-12-2022 | Malfunction | us surgical puerto rico | 19-01-2023 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A SECOND LOOK LAPAROTOMY WITH APPENDECTOMY, SMALL BOWEL ANASTOMOSIS, WASHOUT, VACUUM-ASSISTED CLOSURE (VAC) DEVICE PLACEMENT, AND DAMAGE CONTROL ABDOMEN, THE DEVICE JAMMED AND WAS NOT DISPENSING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, THE HANDLE WAS SQUEEZED, AND THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
04-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, THE HANDLE WAS SQUEEZED, AND THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
04-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, THE HANDLE WAS SQUEEZED, AND THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
22-12-2022 | Malfunction | us surgical puerto rico | 19-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE THERE WAS NO PATIENT INVOLVEMENT.
04-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON LOW ANTERIOR RESECTION PROCEDURE, WHILE TRANSECTING RECTUM, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND HANDLE WAS NOT SQUEEZED. THE 2 HANDLES AND 2 RELOADS THAT DID NOT FIRE WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Malfunction | us surgical puerto rico | 19-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON LOW ANTERIOR RESECTION PROCEDURE, WHILE TRANSECTING RECTUM, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND HANDLE WAS NOT SQUEEZED. THE 2 HANDLES AND 2 RELOADS THAT DID NOT FIRE WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Malfunction | covidien lp llc north haven | 19-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON LOW ANTERIOR RESECTION PROCEDURE, WHILE TRANSECTING RECTUM, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND HANDLE WAS NOT SQUEEZED. THE 2 HANDLES AND 2 RELOADS THAT DID NOT FIRE WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-11-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE FARTHER AFTER FIRED HALF. CHANGED TO ANOTHER ONE TO USE, STILL COULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-12-2022 | Malfunction | johnson & johnson medical (suzhou | 19-01-2023 | CARTRIDGE FINAL ASSY 60MM WHITE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
01-01-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING FIRING WITH A POWERED ECHELON STAPLER WITH GREEN RELOAD, THE STAPLER MALFUNCTIONED BY STAPLING ONE SIDE WITHOUT THE OTHER. NO PATIENT CONSEQUENCES.
30-12-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2022 | Malfunction | ethicon endo-surgery | 19-01-2023 | GST GOLD RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-11-2022 | Malfunction | ethicon | 19-01-2023 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2022 AND ABSORBABLE STRAPS WERE USED. THE FIRST TWO STRAPS FIRED NORMALLY WHEN THE TRIGGER WAS DEPRESSED, AND THE LAST CLIP FAILED TO FIRE NORMALLY AFTER MULTIPLE ATTEMPTS. THE STRAPS DID NOT FIX THE MESH. CHANGED TO ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED.
20-12-2022 | Injury | intuitive surgical | 19-01-2023 | SUREFORM | Difficult to Remove | A150207 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, TISSUE WAS STUCK TO A WHITE SUREFORM 60 RELOAD AFTER THE STAPLER RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. THE STAPLER INSTRUMENT ALLEGEDLY DID NOT CUT THROUGH TISSUE AND CAUSED MORE BLEEDING THEN NORMAL. INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: THE CSR STATED SHE WAS PRESENT DURING THE PROCEDURE AND DISCUSSED THE EVENT WITH THE SURGEON AFTER AS WELL. DURING THE SECOND TO LAST FIRE, WHILE THE SURGEON FIRED ON GASTRIC TISSUE WITH A WHITE SUREFORM 60 RELOAD, THERE WAS NO ISSUE WITH THE FIRING SEQUENCE. HOWEVER, WHEN THE SURGEON UNCLAMPED THE STAPLER INSTRUMENT, A LITTLE BIT OF GASTRIC TISSUE WAS STUCK TO THE TIP OF THE STAPLER RELOAD. THE SURGEON WAS ABLE TO MANIPULATE THE TISSUE WITHOUT HAVING TO CUT. A MINIMAL AMOUNT OF OOZING WAS PRESENT, AND THE SURGEON ADDRESSED THE MINOR OOZE WITH BIPOLAR ENERGY. THE CSR CONFIRMED THAT THERE WERE NO SYSTEM GENERATED MESSAGES, BUTTRESS MATERIAL USAGE, OR VIDEO RECORDING AVAILABLE. THE SURGEON SAID HE USED THE WHITE RELOAD ON GASTRIC TISSUE TO AVOID BUTTRESS MATERIAL USAGE AND OOZING. THE PATIENT IS REPORTED TO BE RECOVERING WITH NO ISSUES.
20-12-2022 | Injury | intuitive surgical | 19-01-2023 | SUREFORM | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, TISSUE WAS STUCK TO A WHITE SUREFORM 60 RELOAD AFTER THE STAPLER RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. THE STAPLER INSTRUMENT ALLEGEDLY DID NOT CUT THROUGH TISSUE AND CAUSED MORE BLEEDING THEN NORMAL. INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: THE CSR STATED SHE WAS PRESENT DURING THE PROCEDURE AND DISCUSSED THE EVENT WITH THE SURGEON AFTER AS WELL. DURING THE SECOND TO LAST FIRE, WHILE THE SURGEON FIRED ON GASTRIC TISSUE WITH A WHITE SUREFORM 60 RELOAD, THERE WAS NO ISSUE WITH THE FIRING SEQUENCE. HOWEVER, WHEN THE SURGEON UNCLAMPED THE STAPLER INSTRUMENT, A LITTLE BIT OF GASTRIC TISSUE WAS STUCK TO THE TIP OF THE STAPLER RELOAD. THE SURGEON WAS ABLE TO MANIPULATE THE TISSUE WITHOUT HAVING TO CUT. A MINIMAL AMOUNT OF OOZING WAS PRESENT, AND THE SURGEON ADDRESSED THE MINOR OOZE WITH BIPOLAR ENERGY. THE CSR CONFIRMED THAT THERE WERE NO SYSTEM GENERATED MESSAGES, BUTTRESS MATERIAL USAGE, OR VIDEO RECORDING AVAILABLE. THE SURGEON SAID HE USED THE WHITE RELOAD ON GASTRIC TISSUE TO AVOID BUTTRESS MATERIAL USAGE AND OOZING. THE PATIENT IS REPORTED TO BE RECOVERING WITH NO ISSUES.
20-12-2022 | Injury | intuitive surgical | 19-01-2023 | SUREFORM | Failure to Cut | A050702 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, TISSUE WAS STUCK TO A WHITE SUREFORM 60 RELOAD AFTER THE STAPLER RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. THE STAPLER INSTRUMENT ALLEGEDLY DID NOT CUT THROUGH TISSUE AND CAUSED MORE BLEEDING THEN NORMAL. INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: THE CSR STATED SHE WAS PRESENT DURING THE PROCEDURE AND DISCUSSED THE EVENT WITH THE SURGEON AFTER AS WELL. DURING THE SECOND TO LAST FIRE, WHILE THE SURGEON FIRED ON GASTRIC TISSUE WITH A WHITE SUREFORM 60 RELOAD, THERE WAS NO ISSUE WITH THE FIRING SEQUENCE. HOWEVER, WHEN THE SURGEON UNCLAMPED THE STAPLER INSTRUMENT, A LITTLE BIT OF GASTRIC TISSUE WAS STUCK TO THE TIP OF THE STAPLER RELOAD. THE SURGEON WAS ABLE TO MANIPULATE THE TISSUE WITHOUT HAVING TO CUT. A MINIMAL AMOUNT OF OOZING WAS PRESENT, AND THE SURGEON ADDRESSED THE MINOR OOZE WITH BIPOLAR ENERGY. THE CSR CONFIRMED THAT THERE WERE NO SYSTEM GENERATED MESSAGES, BUTTRESS MATERIAL USAGE, OR VIDEO RECORDING AVAILABLE. THE SURGEON SAID HE USED THE WHITE RELOAD ON GASTRIC TISSUE TO AVOID BUTTRESS MATERIAL USAGE AND OOZING. THE PATIENT IS REPORTED TO BE RECOVERING WITH NO ISSUES.
20-12-2022 | Injury | intuitive surgical | 19-01-2023 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, TISSUE WAS STUCK TO A WHITE SUREFORM 60 RELOAD AFTER THE STAPLER RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. THE STAPLER INSTRUMENT ALLEGEDLY DID NOT CUT THROUGH TISSUE AND CAUSED MORE BLEEDING THEN NORMAL. INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT: THE CSR STATED SHE WAS PRESENT DURING THE PROCEDURE AND DISCUSSED THE EVENT WITH THE SURGEON AFTER AS WELL. DURING THE SECOND TO LAST FIRE, WHILE THE SURGEON FIRED ON GASTRIC TISSUE WITH A WHITE SUREFORM 60 RELOAD, THERE WAS NO ISSUE WITH THE FIRING SEQUENCE. HOWEVER, WHEN THE SURGEON UNCLAMPED THE STAPLER INSTRUMENT, A LITTLE BIT OF GASTRIC TISSUE WAS STUCK TO THE TIP OF THE STAPLER RELOAD. THE SURGEON WAS ABLE TO MANIPULATE THE TISSUE WITHOUT HAVING TO CUT. A MINIMAL AMOUNT OF OOZING WAS PRESENT, AND THE SURGEON ADDRESSED THE MINOR OOZE WITH BIPOLAR ENERGY. THE CSR CONFIRMED THAT THERE WERE NO SYSTEM GENERATED MESSAGES, BUTTRESS MATERIAL USAGE, OR VIDEO RECORDING AVAILABLE. THE SURGEON SAID HE USED THE WHITE RELOAD ON GASTRIC TISSUE TO AVOID BUTTRESS MATERIAL USAGE AND OOZING. THE PATIENT IS REPORTED TO BE RECOVERING WITH NO ISSUES.
06-12-2022 | Malfunction | intuitive surgical | 20-01-2023 | SUREFORM | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, THE SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT WAS MOVING WHEN THE SURGEON HAD NOT MOVED AT THE CONSOLE. THE PROCEDURE WAS COMPLETED AS PLANNED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) HAS MADE MULTIPLE FOLLOW-UP ATTEMPTS TO OBTAIN ADDITIONAL PATIENT/EVENT INFORMATION. HOWEVER, DESPITE THE FOLLOW-UP GOOD FAITH EFFORTS NO FURTHER DETAILS HAVE BEEN RECEIVED AS OF THE DATE OF THIS REPORT.
05-01-2023 | Malfunction | covidien lp llc north haven | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
03-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, IN STOMACH RESECTION, ON TWO DEVICES, THE SURGEON COULD NOT PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
23-12-2022 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2023 | SORBAFIX ENHANCED | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNSPECIFIED PROCEDURE, THE SORBAFIX ENHANCED FIXATION DEVICE SUCCESSFULLY DEPLOYED FIRST FASTENER AND RELEASED SECOND FASTENER AS A BROKEN FASTENER. IT WAS REPORTED THAT THE REMNANTS WERE REMOVED FROM THE PATIENT¿S BODY. THERE WAS NO REPORTED PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | 75MM SELECTABLE NEW TLC | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN PROCEDURE, THE CLOSURE RETAINING KNOB GOT DETACHED ON THE OPERATING FIELD AT THE TIME OF USE.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | High Blood Pressure/ Hypertension | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | Septic Shock | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | High Blood Pressure/ Hypertension | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | Septic Shock | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Material Integrity Problem | A04 | High Blood Pressure/ Hypertension | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Material Integrity Problem | A04 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Material Integrity Problem | A04 | Septic Shock | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Patient Device Interaction Problem | A01 | High Blood Pressure/ Hypertension | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Patient Device Interaction Problem | A01 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
06-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Patient Device Interaction Problem | A01 | Septic Shock | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRIC SLEEVE PROCEDURE ON THE FOURTH OR FIFTH RELOAD FIRING, THE PATIENT SIDE OF THE STAPLE LINE (LEFT SIDE AS LOOKING DOWN THE RELOAD PROXIMAL TO DISTAL WITH THE STAPLE DECK UP) WAS NOT STAPLED FOR "MOST IF NOT ALL OF THE STAPLE LINE". THE OR ASSIST AND SURGEON NOTED THAT UNFORMED STAPLES LAY ALONG THE CUT LINE. THE RIGHT SIDE STAPLE LINE APPEARED INTACT WITHOUT ISSUE. THE UNSTAPLED SECTION WAS STAPLED OVER BY A NEW LIKE DEVICE. IT IS NOTED THAT A GORE SEAM GUARD WAS USED ON THE FIRING. THE OR ASSIST ALSO STATED THAT HE OBSERVED A "CRACK" OF THE PLASTIC ALONG THE LENGTH OF THE LEFT SIDE OF THE RELOAD. THE PATIENT RETURNED TO THE OR ON (B)(6) 2023. THE SURGERY WAS DELAYED. THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WERE PATIENT CONSEQUENCES.
05-01-2023 | Malfunction | covidien lp llc north haven | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
03-10-2022 | Malfunction | covidien lp llc north haven | 20-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, WHILE CHARGING, THE HANDLE FAILED TO CHARGE; THE HANDLE WAS SITTING BLACK IN THE CHARGER. ANOTHER POWERED HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
03-10-2022 | Malfunction | covidien lp llc north haven | 20-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, WHILE CHARGING, THE HANDLE FAILED TO CHARGE; THE HANDLE WAS SITTING BLACK IN THE CHARGER. ANOTHER POWERED HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
23-02-2022 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL RESECTION OF ESOPHAGEAL CANCER, ON THE STOMACH TISSUE, THE STAPLES OPENED AFTER THE DOCTOR WAS ABLE TO FIRE THE DEVICE. THE JAWS WERE UNABLE TO OPEN WITHOUT INVOLVING TISSUE. THE DEVICE WAS REPLACED WITH A NEW STAPLER AND STAPLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-02-2022 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL RESECTION OF ESOPHAGEAL CANCER, ON THE STOMACH TISSUE, THE STAPLES OPENED AFTER THE DOCTOR WAS ABLE TO FIRE THE DEVICE. THE JAWS WERE UNABLE TO OPEN WITHOUT INVOLVING TISSUE. THE DEVICE WAS REPLACED WITH A NEW STAPLER AND STAPLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-02-2022 | Malfunction | us surgical puerto rico | 20-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL RESECTION OF ESOPHAGEAL CANCER, ON THE STOMACH TISSUE, THE STAPLES OPENED AFTER THE DOCTOR WAS ABLE TO FIRE THE DEVICE. THE JAWS WERE UNABLE TO OPEN WITHOUT INVOLVING TISSUE. THE DEVICE WAS REPLACED WITH A NEW STAPLER AND STAPLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Malfunction | covidien lp llc north haven | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN TESTING THE DEVICE, THE STAPLER DID NOT FIRE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
05-01-2023 | Malfunction | covidien lp llc north haven | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN TESTING THE DEVICE, THE STAPLER DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
05-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN TESTING THE DEVICE, THE STAPLER DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
05-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN TESTING THE DEVICE, THE STAPLER DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
01-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | ILS 29MM, CURVED | Failure to Cut | A050702 | Pain | Event Description: IT WAS REPORTED THAT THERE WAS A STAPLING DEFECT OF THE ANTERIOR HEMI-CIRCUMFERENCE. COMPLETE STAPLING OF THE POSTERIOR HEMI-CIRCUMFERENCE. COMPLETE RECTAL COLLAR. NO COLIC COLLAR FIND. THIS ISSUE NECESSITATED THE CONVERSION TO LAPAROTOMY OF THE INITIAL LAPAROSCOPY AND INCREASED THE DURATION OF THE INTERVENTION BY APPROXIMATELY 1H30 AND THE ANASTOMOSIS WAS PERFORMED MANUALLY. PROCEDURE: SIGMOID COLON RESECTION.
01-01-2023 | Injury | ethicon endo-surgery | 20-01-2023 | ILS 29MM, CURVED | Failure to Fire | A050501 | Pain | Event Description: IT WAS REPORTED THAT THERE WAS A STAPLING DEFECT OF THE ANTERIOR HEMI-CIRCUMFERENCE. COMPLETE STAPLING OF THE POSTERIOR HEMI-CIRCUMFERENCE. COMPLETE RECTAL COLLAR. NO COLIC COLLAR FIND. THIS ISSUE NECESSITATED THE CONVERSION TO LAPAROTOMY OF THE INITIAL LAPAROSCOPY AND INCREASED THE DURATION OF THE INTERVENTION BY APPROXIMATELY 1H30 AND THE ANASTOMOSIS WAS PERFORMED MANUALLY. PROCEDURE: SIGMOID COLON RESECTION.
18-10-2022 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE STAPLER GOT STUCK, AND COULD NOT ACTIVATE THE RELOAD TO COMPLETE THE DETACHMENT OPERATION. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-10-2022 | Malfunction | covidien lp llc north haven | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE STAPLER GOT STUCK, AND COULD NOT ACTIVATE THE RELOAD TO COMPLETE THE DETACHMENT OPERATION. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2023 | CAPSURE STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC HERNIA REPAIR PROCEDURE, THE CAPSURE FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE DEVICE FAILED TO FIXATE THE MESH WHEN DEPLOYING THE FIRST FASTENER. AS REPORTED, NO BROKEN FASTENER WAS FOUND AND ANOTHER CAPSURE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY.
05-01-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2023 | CAPSURE STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC HERNIA REPAIR PROCEDURE, THE CAPSURE FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE DEVICE FAILED TO FIXATE THE MESH WHEN DEPLOYING THE FIRST FASTENER. AS REPORTED, NO BROKEN FASTENER WAS FOUND AND ANOTHER CAPSURE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY.
24-08-2022 | Malfunction | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE BLOCKED WHILE OPENED. NO PATIENT CONSEQUENCES.
24-08-2022 | Malfunction | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE BLOCKED WHILE OPENED. NO PATIENT CONSEQUENCES.
25-12-2022 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, IN TRANSECTING THE APPENDIX, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
03-01-2022 | Death | intuitive surgical | 20-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Fistula | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED ESOPHAGECTOMY WITH INTRATHORACIC ANASTOMOSIS PROCEDURE, A LEAK WAS DISCOVERED ON THE 19TH POST-OPERATIVE DAY (POD) AT THE VERTICAL STAPLE LINE OF THE GASTRIC CONDUIT. THE LEAK WAS DISCOVERED AT THE CROSS POINT OF THE 1ST AND 2ND STAPLE LINE. THE LEAK CAUSED AN AORTO-ESOPHAGEAL FISTULA WHICH RESULTED IN ACUTE BLEEDING TO THE GASTROINTESTINAL TRACT (GIT). A BALLOON TAMPONADE OF THE GASTRIC TUBE AND PLACEMENT OF A TEVAR (THORACIC ENDOVASCULAR AORTIC REPAIR) WAS DONE TO TREAT THE BLEEDING. THE PATIENT DIED OF MASSIVE BLEEDING BEFORE A RIGHT THORACOTOMY COULD BE PERFORMED. DURING THE PRIMARY PROCEDURE, THERE WERE NO INTRA-OPERATIVE COMPLICATIONS OR STAPLING ISSUES. THERE WERE NO ERROR MESSAGES, OBSTRUCTION IN THE JAWS OF THE STAPLER, BUTTRESS MATERIAL USED, NO CLAMPING ISSUES AND NO MALFORMED STAPLES PRODUCED. THE SURGICAL STAFF STATED THAT THEY WASH THE STAPLERS BETWEEN EACH STAPLER FIRES. THE TARGET TISSUE (GASTRIC TISSUE) DID NOT HAVE CALCIFICATIONS OR TISSUE TENSION AND WAS NOT EXPOSED TO CHEMOTHERAPY/RADIATION. THERE WAS NO TISSUE BUNCHING DURING STAPLING. THE STAPLER AND RELOADS HAD BEEN DISCARDED BEFORE THE CREATION OF THIS COMPLAINT. ISI MADE MULTIPLE FOLLOW-UP ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE SURGEON. HOWEVER, NO FURTHER DETAILS WERE RECEIVED AS OF THE DATE OF THIS REPORT.
03-01-2022 | Death | intuitive surgical | 20-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED ESOPHAGECTOMY WITH INTRATHORACIC ANASTOMOSIS PROCEDURE, A LEAK WAS DISCOVERED ON THE 19TH POST-OPERATIVE DAY (POD) AT THE VERTICAL STAPLE LINE OF THE GASTRIC CONDUIT. THE LEAK WAS DISCOVERED AT THE CROSS POINT OF THE 1ST AND 2ND STAPLE LINE. THE LEAK CAUSED AN AORTO-ESOPHAGEAL FISTULA WHICH RESULTED IN ACUTE BLEEDING TO THE GASTROINTESTINAL TRACT (GIT). A BALLOON TAMPONADE OF THE GASTRIC TUBE AND PLACEMENT OF A TEVAR (THORACIC ENDOVASCULAR AORTIC REPAIR) WAS DONE TO TREAT THE BLEEDING. THE PATIENT DIED OF MASSIVE BLEEDING BEFORE A RIGHT THORACOTOMY COULD BE PERFORMED. DURING THE PRIMARY PROCEDURE, THERE WERE NO INTRA-OPERATIVE COMPLICATIONS OR STAPLING ISSUES. THERE WERE NO ERROR MESSAGES, OBSTRUCTION IN THE JAWS OF THE STAPLER, BUTTRESS MATERIAL USED, NO CLAMPING ISSUES AND NO MALFORMED STAPLES PRODUCED. THE SURGICAL STAFF STATED THAT THEY WASH THE STAPLERS BETWEEN EACH STAPLER FIRES. THE TARGET TISSUE (GASTRIC TISSUE) DID NOT HAVE CALCIFICATIONS OR TISSUE TENSION AND WAS NOT EXPOSED TO CHEMOTHERAPY/RADIATION. THERE WAS NO TISSUE BUNCHING DURING STAPLING. THE STAPLER AND RELOADS HAD BEEN DISCARDED BEFORE THE CREATION OF THIS COMPLAINT. ISI MADE MULTIPLE FOLLOW-UP ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE SURGEON. HOWEVER, NO FURTHER DETAILS WERE RECEIVED AS OF THE DATE OF THIS REPORT.
05-01-2023 | Injury | us surgical puerto rico | 20-01-2023 | EEA | Mechanics Altered | A0507 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHEN THE SURGEON COMPLETELY FIRED THE DEVICE AND REMOVED IT OUT, THE CIRCULAR STAPLER CUT THE RECTUM TISSUE (COMPLETE TWO DONUT RING) BUT NO STAPLER IN ANASTOMOSIS SIDE THAT MEANS NO STAPLER LINE IN ANASTOMOSIS SIDE. WHEN THE SURGEON USE FINGER TO PALPATE AND CHECK ON ANASTOMOSIS SIDE, THERE WAS NO ANY STAPLER WIRE PRESENT. THE SURGEON NEED TO OVERSEW TO CLOSE THE STUMP AND USE A NEW CIRCULAR STAPLER TO DO END TO END ANASTOMOSIS AGAIN. THE SURGICAL TIME EXTENDED AN HOUR DUE TO THE PRODUCT PROBLEM.
05-01-2023 | Injury | us surgical puerto rico | 20-01-2023 | EEA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHEN THE SURGEON COMPLETELY FIRED THE DEVICE AND REMOVED IT OUT, THE CIRCULAR STAPLER CUT THE RECTUM TISSUE (COMPLETE TWO DONUT RING) BUT NO STAPLER IN ANASTOMOSIS SIDE THAT MEANS NO STAPLER LINE IN ANASTOMOSIS SIDE. WHEN THE SURGEON USE FINGER TO PALPATE AND CHECK ON ANASTOMOSIS SIDE, THERE WAS NO ANY STAPLER WIRE PRESENT. THE SURGEON NEED TO OVERSEW TO CLOSE THE STUMP AND USE A NEW CIRCULAR STAPLER TO DO END TO END ANASTOMOSIS AGAIN. THE SURGICAL TIME EXTENDED AN HOUR DUE TO THE PRODUCT PROBLEM.
05-01-2023 | Injury | us surgical puerto rico | 20-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHEN THE SURGEON COMPLETELY FIRED THE DEVICE AND REMOVED IT OUT, THE CIRCULAR STAPLER CUT THE RECTUM TISSUE (COMPLETE TWO DONUT RING) BUT NO STAPLER IN ANASTOMOSIS SIDE THAT MEANS NO STAPLER LINE IN ANASTOMOSIS SIDE. WHEN THE SURGEON USE FINGER TO PALPATE AND CHECK ON ANASTOMOSIS SIDE, THERE WAS NO ANY STAPLER WIRE PRESENT. THE SURGEON NEED TO OVERSEW TO CLOSE THE STUMP AND USE A NEW CIRCULAR STAPLER TO DO END TO END ANASTOMOSIS AGAIN. THE SURGICAL TIME EXTENDED AN HOUR DUE TO THE PRODUCT PROBLEM.
05-01-2023 | Injury | us surgical puerto rico | 20-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHEN THE SURGEON COMPLETELY FIRED THE DEVICE AND REMOVED IT OUT, THE CIRCULAR STAPLER CUT THE RECTUM TISSUE (COMPLETE TWO DONUT RING) BUT NO STAPLER IN ANASTOMOSIS SIDE THAT MEANS NO STAPLER LINE IN ANASTOMOSIS SIDE. WHEN THE SURGEON USE FINGER TO PALPATE AND CHECK ON ANASTOMOSIS SIDE, THERE WAS NO ANY STAPLER WIRE PRESENT. THE SURGEON NEED TO OVERSEW TO CLOSE THE STUMP AND USE A NEW CIRCULAR STAPLER TO DO END TO END ANASTOMOSIS AGAIN. THE SURGICAL TIME EXTENDED AN HOUR DUE TO THE PRODUCT PROBLEM.
23-09-2022 | Malfunction | celestica de monterrey s.a de c.v | 20-01-2023 | ENDO GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC SEGMENTAL RESECTION OF LUNG, IN PLANE CUTTING BETWEEN THE LUNG SEGMENTS, THE RELOAD WAS DEPLOYED BUT COULD NOT CUT, AND THE BUTTON COULD NOT BE PRESSED. ANOTHER RELOAD WAS USED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INJURY.
23-09-2022 | Malfunction | celestica de monterrey s.a de c.v | 20-01-2023 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC SEGMENTAL RESECTION OF LUNG, IN PLANE CUTTING BETWEEN THE LUNG SEGMENTS, THE RELOAD WAS DEPLOYED BUT COULD NOT CUT, AND THE BUTTON COULD NOT BE PRESSED. ANOTHER RELOAD WAS USED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT LIVER TUMOR RESECTION, AT THE TIME OF DISSOCIATED LIVER PARENCHYMA, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
05-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HARTMANN'S REVERSAL THE SURGEON PUT DEVICE IN AND FIRED THEN WENT TO REMOVE IT AND THE ANVIL GOT STUCK IN RECTUM. HE RETRIEVED IT WITH A GRASPER. HE TURNED THE KNOB 2.5 TURNS PER USUAL. A LEAK TEST WAS NEGATIVE.
05-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMI-HEPATECTOMY, DURING A DISSOCIATED RIGHT HEPATIC CAUDLE AND BLOOD VESSEL, THE TWO DEVICES DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW DEVICE FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
20-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | EEA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID/RECTUM END-TO-END ANASTOMOSIS, THE TISSUE ENCAPSULATED IN HOUSING AND DISTAL DONUT WAS THIN AND FRAGILE. THE SURGEON COMPLETED THE PROCEDURE AS NORMAL TO RESOLVE THE ISSUE.
20-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | EEA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID/RECTUM END-TO-END ANASTOMOSIS, THE TISSUE ENCAPSULATED IN HOUSING AND DISTAL DONUT WAS THIN AND FRAGILE. THE SURGEON COMPLETED THE PROCEDURE AS NORMAL TO RESOLVE THE ISSUE.
05-01-2023 | Malfunction | covidien lp llc north haven | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, DURING A DISSOCIATED RIGHT HEPATIC CAUDLE AND BLOOD VESSEL, THE TWO DEVICES DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW DEVICE FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
05-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, DURING A DISSOCIATED RIGHT HEPATIC CAUDLE AND BLOOD VESSEL, THE TWO DEVICES DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW DEVICE FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
05-01-2023 | Malfunction | covidien lp llc north haven | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, DURING A DISSOCIATED RIGHT HEPATIC CAUDLE AND BLOOD VESSEL, THE TWO DEVICES DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. A NEW DEVICE FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
20-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | EEA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID/RECTUM END-TO-END ANASTOMOSIS, THE TISSUE ENCAPSULATED IN HOUSING AND DISTAL DONUT WAS THIN AND FRAGILE. THE SURGEON COMPLETED THE PROCEDURE AS NORMAL TO RESOLVE THE ISSUE.
20-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | EEA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID/RECTUM END-TO-END ANASTOMOSIS, THE TISSUE ENCAPSULATED IN HOUSING AND DISTAL DONUT WAS THIN AND FRAGILE. THE SURGEON COMPLETED THE PROCEDURE AS NORMAL TO RESOLVE THE ISSUE.
05-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL MESORECTAL RESECTION, WHILE AT RECTAL CLOSED DISSECTION, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS SQUEEZED. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Injury | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE ON RESECTION THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
05-01-2023 | Injury | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE ON RESECTION THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
05-01-2023 | Injury | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE ON RESECTION THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
05-01-2023 | Injury | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE ON RESECTION THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
05-01-2023 | Injury | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE ON RESECTION THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
20-08-2022 | Injury | ethicon endo-surgery | 20-01-2023 | ILS 29MM, CURVED | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN PROCEDURE, THE STEP TO USE WAS DONE CORRECTLY BUT DURING FIRING THE STAPLER, THE FIRING TRIGGER DID NOT CLOSE. THE SURGEON PUT A LOT OF PRESSURE TO FIRE THE TRIGGER THEN IT WAS CLOSED BUT THE STAPLER HAD AN UNUSUAL SOUND, AND THE DONUT DID NOT HAVE A COMPLETE FORM AND THE STAPLE LINE WAS NOT A COMPLETE CIRCLE. AT THE END THE SURGEON FINISHED THE ANASTOMOSIS BY SUTURING.
20-08-2022 | Injury | ethicon endo-surgery | 20-01-2023 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN PROCEDURE, THE STEP TO USE WAS DONE CORRECTLY BUT DURING FIRING THE STAPLER, THE FIRING TRIGGER DID NOT CLOSE. THE SURGEON PUT A LOT OF PRESSURE TO FIRE THE TRIGGER THEN IT WAS CLOSED BUT THE STAPLER HAD AN UNUSUAL SOUND, AND THE DONUT DID NOT HAVE A COMPLETE FORM AND THE STAPLE LINE WAS NOT A COMPLETE CIRCLE. AT THE END THE SURGEON FINISHED THE ANASTOMOSIS BY SUTURING.
20-08-2022 | Injury | ethicon endo-surgery | 20-01-2023 | ILS 29MM, CURVED | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN PROCEDURE, THE STEP TO USE WAS DONE CORRECTLY BUT DURING FIRING THE STAPLER, THE FIRING TRIGGER DID NOT CLOSE. THE SURGEON PUT A LOT OF PRESSURE TO FIRE THE TRIGGER THEN IT WAS CLOSED BUT THE STAPLER HAD AN UNUSUAL SOUND, AND THE DONUT DID NOT HAVE A COMPLETE FORM AND THE STAPLE LINE WAS NOT A COMPLETE CIRCLE. AT THE END THE SURGEON FINISHED THE ANASTOMOSIS BY SUTURING.
20-08-2022 | Injury | ethicon endo-surgery | 20-01-2023 | ILS 29MM, CURVED | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN PROCEDURE, THE STEP TO USE WAS DONE CORRECTLY BUT DURING FIRING THE STAPLER, THE FIRING TRIGGER DID NOT CLOSE. THE SURGEON PUT A LOT OF PRESSURE TO FIRE THE TRIGGER THEN IT WAS CLOSED BUT THE STAPLER HAD AN UNUSUAL SOUND, AND THE DONUT DID NOT HAVE A COMPLETE FORM AND THE STAPLE LINE WAS NOT A COMPLETE CIRCLE. AT THE END THE SURGEON FINISHED THE ANASTOMOSIS BY SUTURING.
03-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, IN STOMACH RESECTION, ON TWO DEVICES, THE SURGEON COULD NOT PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
21-12-2022 | Malfunction | covidien lp llc north haven | 20-01-2023 | EEA | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A OPEN HEMORRHOIDECTOMY, WHEN THE DEVICE WAS INSERTED INTO THE ANUS TO CUT HEMORRHOIDS, THE ANVIL HEAD AND THE HANDLE DID NOT MATCH. THE HANDLE WAS DIFFICULT OR NOT ABLE TO CLOSE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 20-01-2023 | BLUE RELOAD FOR ECHELON 60 | Structural Problem | A0511 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE CHANGE OF LOAD (BLUE) OF THE STAPLER, THE PROTECTIVE PART (ORANGE COLOR) CAME LOOSE CAUSING INSTABILITY OF THE STAPLES, SO IT WAS NECESSARY TO DISCARD IT AND REPLACE IT WITH A NEW LOAD. THERE WAS NO HARM TO THE PATIENT.
18-11-2022 | Malfunction | covidien lp llc north haven | 20-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE. A NEW PRODUCT WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-11-2022 | Malfunction | us surgical puerto rico | 20-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-11-2022 | Malfunction | us surgical puerto rico | 20-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOW ANTERIOR RESECTION, POWERED CIRCULAR STAPLER DID NOT WORK. THEY PREPARED THE DEVICE, INSERTED THE BATTERY AND AT THE ANASTOMOSIS CREATION STAGE IT DID NOT FIRE. INDICATOR WAS IN A GREEN ZONE ACCORDING TO THEM, BUT AFTER TRIGGER WAS PUSHED DEVICE DID NOT WORKED. THERE WAS A 15-MINUTE DELAY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO PATIENT CONSEQUENCES OR CHANGE IN THE POST-OPERATIVE CARE.
05-01-2023 | Malfunction | us surgical puerto rico | 20-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN CHANGING THE RELOAD, THE STEM BROKE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | ECHELON 3000 60MM LONG | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE THAT ON THE FIRST FIRING OF THE GASTRIC POUCH WITH A LITTLE BIT OF ARTICULATION, THE DEVICE WAS FIRED BUT THEY DID NOT HEAR THE KNIFE RETURN HOME. SURGEON PRESSED ON THE FIRING TRIGGER MORE TO SEE IF KNIFE WOULD RETURN HOME. SURGEON ATTEMPTED TO OPEN THE DEVICE, BUT DEVICE WOULD NOT OPEN. HOME BUTTON WAS PRESSED THEN DEVICE WAS ATTEMPTED TO OPEN BUT DEVICE WOULD NOT OPEN. BATTERY WAS THEN REMOVED AND REPLACE WITH ANOTHER ATTEMPT TO OPEN THE DEVICE, BUT THE DEVICE WOULD NOT OPEN. MANUAL OVERRIDE WAS THEN USED TO RETURN THE KNIFE HOME AND OPEN THE DEVICE. THERE WERE A FEW LOOSE STAPLES AT THE END OF THE STAPLE LINE. SURGEON OVER-SEWED THE STAPLE LINE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
04-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE THAT ON THE FIRST FIRING OF THE GASTRIC POUCH WITH A LITTLE BIT OF ARTICULATION, THE DEVICE WAS FIRED BUT THEY DID NOT HEAR THE KNIFE RETURN HOME. SURGEON PRESSED ON THE FIRING TRIGGER MORE TO SEE IF KNIFE WOULD RETURN HOME. SURGEON ATTEMPTED TO OPEN THE DEVICE, BUT DEVICE WOULD NOT OPEN. HOME BUTTON WAS PRESSED THEN DEVICE WAS ATTEMPTED TO OPEN BUT DEVICE WOULD NOT OPEN. BATTERY WAS THEN REMOVED AND REPLACE WITH ANOTHER ATTEMPT TO OPEN THE DEVICE, BUT THE DEVICE WOULD NOT OPEN. MANUAL OVERRIDE WAS THEN USED TO RETURN THE KNIFE HOME AND OPEN THE DEVICE. THERE WERE A FEW LOOSE STAPLES AT THE END OF THE STAPLE LINE. SURGEON OVER-SEWED THE STAPLE LINE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
04-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | ECHELON 3000 60MM LONG | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE THAT ON THE FIRST FIRING OF THE GASTRIC POUCH WITH A LITTLE BIT OF ARTICULATION, THE DEVICE WAS FIRED BUT THEY DID NOT HEAR THE KNIFE RETURN HOME. SURGEON PRESSED ON THE FIRING TRIGGER MORE TO SEE IF KNIFE WOULD RETURN HOME. SURGEON ATTEMPTED TO OPEN THE DEVICE, BUT DEVICE WOULD NOT OPEN. HOME BUTTON WAS PRESSED THEN DEVICE WAS ATTEMPTED TO OPEN BUT DEVICE WOULD NOT OPEN. BATTERY WAS THEN REMOVED AND REPLACE WITH ANOTHER ATTEMPT TO OPEN THE DEVICE, BUT THE DEVICE WOULD NOT OPEN. MANUAL OVERRIDE WAS THEN USED TO RETURN THE KNIFE HOME AND OPEN THE DEVICE. THERE WERE A FEW LOOSE STAPLES AT THE END OF THE STAPLE LINE. SURGEON OVER-SEWED THE STAPLE LINE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
04-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | ECHELON 3000 60MM LONG | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE THAT ON THE FIRST FIRING OF THE GASTRIC POUCH WITH A LITTLE BIT OF ARTICULATION, THE DEVICE WAS FIRED BUT THEY DID NOT HEAR THE KNIFE RETURN HOME. SURGEON PRESSED ON THE FIRING TRIGGER MORE TO SEE IF KNIFE WOULD RETURN HOME. SURGEON ATTEMPTED TO OPEN THE DEVICE, BUT DEVICE WOULD NOT OPEN. HOME BUTTON WAS PRESSED THEN DEVICE WAS ATTEMPTED TO OPEN BUT DEVICE WOULD NOT OPEN. BATTERY WAS THEN REMOVED AND REPLACE WITH ANOTHER ATTEMPT TO OPEN THE DEVICE, BUT THE DEVICE WOULD NOT OPEN. MANUAL OVERRIDE WAS THEN USED TO RETURN THE KNIFE HOME AND OPEN THE DEVICE. THERE WERE A FEW LOOSE STAPLES AT THE END OF THE STAPLE LINE. SURGEON OVER-SEWED THE STAPLE LINE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
04-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | ECHELON 3000 60MM LONG | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE THAT ON THE FIRST FIRING OF THE GASTRIC POUCH WITH A LITTLE BIT OF ARTICULATION, THE DEVICE WAS FIRED BUT THEY DID NOT HEAR THE KNIFE RETURN HOME. SURGEON PRESSED ON THE FIRING TRIGGER MORE TO SEE IF KNIFE WOULD RETURN HOME. SURGEON ATTEMPTED TO OPEN THE DEVICE, BUT DEVICE WOULD NOT OPEN. HOME BUTTON WAS PRESSED THEN DEVICE WAS ATTEMPTED TO OPEN BUT DEVICE WOULD NOT OPEN. BATTERY WAS THEN REMOVED AND REPLACE WITH ANOTHER ATTEMPT TO OPEN THE DEVICE, BUT THE DEVICE WOULD NOT OPEN. MANUAL OVERRIDE WAS THEN USED TO RETURN THE KNIFE HOME AND OPEN THE DEVICE. THERE WERE A FEW LOOSE STAPLES AT THE END OF THE STAPLE LINE. SURGEON OVER-SEWED THE STAPLE LINE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
03-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. USED SUTURE TO OVERSEW.THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-12-2022 | Malfunction | ethicon endo-surgery | 20-01-2023 | POWERED PLUS COMPACT 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THORACOSCOPIC PNEUMONECTOMY, THE DEVICE LOCKED OUT WHEN PULSE-FIRING AT 1ST FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
04-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TOTAL ABDOMINAL COLECTOMY PROCEDURE THAT AFTER FIRING ON THE BOWEL THE JAWS WOULD NOT OPEN. REVERSE WAS ATTEMPTED BUT DID NOT OPEN THE DEVICE. MANUAL OVERRIDE WAS ALSO ATTEMPTED BUT THE DEVICE STILL WOULD NOT OPEN. DEVICE WAS REMOVED BY UNKNOWN MEANS. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
04-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TOTAL ABDOMINAL COLECTOMY PROCEDURE THAT AFTER FIRING ON THE BOWEL THE JAWS WOULD NOT OPEN. REVERSE WAS ATTEMPTED BUT DID NOT OPEN THE DEVICE. MANUAL OVERRIDE WAS ALSO ATTEMPTED BUT THE DEVICE STILL WOULD NOT OPEN. DEVICE WAS REMOVED BY UNKNOWN MEANS. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
04-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TOTAL ABDOMINAL COLECTOMY PROCEDURE THAT AFTER FIRING ON THE BOWEL THE JAWS WOULD NOT OPEN. REVERSE WAS ATTEMPTED BUT DID NOT OPEN THE DEVICE. MANUAL OVERRIDE WAS ALSO ATTEMPTED BUT THE DEVICE STILL WOULD NOT OPEN. DEVICE WAS REMOVED BY UNKNOWN MEANS. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE GUN IS NOT OPENING. NO OTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCE REPORTED.
01-01-2022 | Malfunction | ethicon endo-surgery | 20-01-2023 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE DEVICE DID NOT FUNCTION, AND THAT THE BATTERY APPEARED DEAD. UNKNOWN AS TO HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
21-12-2022 | Injury | intuitive surgical | 20-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED TRANSTHORACIC ESOPHAGECTOMY WITH CHEST ANASTOMOSIS PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT FIRED A GREEN RELOAD AND THE TARGET TISSUE WAS PUSHED AND BUNCHED INSTEAD OF CUT. AN INTUITIVE SURGICAL INC. (ISI) CLINICAL SALES REPRESENTATIVE (CSR) WAS NOT PRESENT DURING THIS EVENT, BUT WAS IN THE HOSPITAL AND WAS CALLED INTO THE ROOM AFTER THE EVENT OCCURRED. THIS EVENT OCCURRED WHILE THE SURGEON WAS STAPLING THE STOMACH TO CREATE THE GASTRIC CONDUIT. THE SURGEON REPORTEDLY UNCLAMPED THE STAPLER JAWS, AND THE STOMACH TISSUE RELAXED AND WAS NO LONGER BUNCHED. THE SURGEON USED SUCTION TO REMOVE THE LOOSE STAPLES. THE SURGEON NEXT STAPLED LATERAL (CLOSER TO CONDUIT SIDE) TO THIS PREVIOUS LINE THAT BUNCHED. THE SURGEON PLACED THE STAPLER WERE THERE WOULD NOT BE CROSSING STAPLE LINES, AND SUTURED THE LOCATION WHERE THERE WAS A CUT WITHOUT STAPLES DUE TO NO CROSSING STAPLE LINES. ISI HAS ATTEMPTED TO CONTACT THE SURGEON OF THIS PROCEDURE TO OBTAIN ADDITIONAL INFORMATION, HOWEVER NO RESPONSE HAS BEEN RECEIVED AT THIS TIME.
17-01-2023 | Malfunction | covidien | 20-01-2023 | TRI-STAPLE 2.0 | Component Missing | A020602 | Insufficient Information | Event Description: THE PURPLE STRAP ON THE TRI-STAPLE 2.0 RELOAD (X2 WITH SAME LOT) WAS MISSING FROM THE DEVICE.
13-12-2022 | Malfunction | intuitive surgical | 21-01-2023 | ENDOWRIST | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
13-12-2022 | Malfunction | intuitive surgical | 21-01-2023 | ENDOWRIST | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
22-12-2022 | Malfunction | intuitive surgical | 21-01-2023 | SUREFORM | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
06-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC BYPASS RESECTION, THE NURSE OPENED THE ASEPTIC PACKAGE OF THE HANDLE AND TESTED THE RELOAD. AFTER THE INSTALLATION WAS SUCCESSFUL, THE DOCTOR NORMALLY INSERTED THE STAPLES INTO THE PATIENT'S ABDOMINAL CAVITY THROUGH THE TROCAR TO PREPARE FOR GASTRECTOMY. THE DOCTOR SUCCESSFULLY CLAMPED THE GASTRIC TISSUE, PRESSED THE GREEN SAFETY BUTTON, AND WAS READY TO FIRE THE BLUE RELOAD. THE BLUE RELOAD COULD BE FIRED NORMALLY, BUT AFTER THE FIRING, THE STAPLE FAILED TO CLOSE AND FORM. DOCTOR AND NURSE PREDICTED THAT IT COULD NOT BE USED NORMALLY, SO IT WAS REPLACED WITH THE SAME MODEL, AND THE GASTRIC TISSUE WAS SUCCESSFULLY ANASTOMOSED. WHEN THE DOCTOR PERFORMED INTESTINAL ANASTOMOSIS, PRESSED THE GREEN SAFETY BUTTON TO PREPARE TO FIRE THE WHITE RELOAD, WHICH COULD BE FIRED NORMALLY, BUT AFTER FIRING, THE STAPLE COULD NOT BE CLOSED AND COULD NOT BE FORMED SUCCESSFULLY. THE WHITE RELOAD WAS REPLACED WITH THE SAME MODEL AND THE GASTRIC TISSUE WAS SUCCESSFULLY ANASTOMOSED. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC BYPASS RESECTION, THE NURSE OPENED THE ASEPTIC PACKAGE OF THE HANDLE AND TESTED THE RELOAD. AFTER THE INSTALLATION WAS SUCCESSFUL, THE DOCTOR NORMALLY INSERTED THE STAPLES INTO THE PATIENT'S ABDOMINAL CAVITY THROUGH THE TROCAR TO PREPARE FOR GASTRECTOMY. THE DOCTOR SUCCESSFULLY CLAMPED THE GASTRIC TISSUE, PRESSED THE GREEN SAFETY BUTTON, AND WAS READY TO FIRE THE BLUE RELOAD. THE BLUE RELOAD COULD BE FIRED NORMALLY, BUT AFTER THE FIRING, THE STAPLE FAILED TO CLOSE AND FORM. DOCTOR AND NURSE PREDICTED THAT IT COULD NOT BE USED NORMALLY, SO IT WAS REPLACED WITH THE SAME MODEL, AND THE GASTRIC TISSUE WAS SUCCESSFULLY ANASTOMOSED. WHEN THE DOCTOR PERFORMED INTESTINAL ANASTOMOSIS, PRESSED THE GREEN SAFETY BUTTON TO PREPARE TO FIRE THE WHITE RELOAD, WHICH COULD BE FIRED NORMALLY, BUT AFTER FIRING, THE STAPLE COULD NOT BE CLOSED AND COULD NOT BE FORMED SUCCESSFULLY. THE WHITE RELOAD WAS REPLACED WITH THE SAME MODEL AND THE GASTRIC TISSUE WAS SUCCESSFULLY ANASTOMOSED. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | covidien lp llc north haven | 23-01-2023 | ENDO GIA | Fail-Safe Problem | A1602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE TWO RELOADS WERE FIRED WITHOUT ANY PROBLEMS. AFTER THE PROCEDURE, THE SAME RELOADS WERE CONNECTED AGAIN TO THE POWERED STAPLER FOR TRAINING PURPOSES. WHEN THE FIRST RELOAD (45MM) WAS CONNECTED AGAIN, THE HANDLE DID NOT RECOGNIZE THE RELOAD AS USED; ON THE DISPLAY, THE RELOAD WAS HIGHLIGHTED IN GREEN AND MARKED WITH ONE. IT WAS ALSO POSSIBLE TO SHOOT AND FIRE; THE BLADE MOVED ONLY A FEW MILLIMETERS AND THEN STOPPED. WHEN THE SECOND RELOAD (60MM) WAS CONNECTED AGAIN, IT HAD THE SAME ISSUE; THE HANDLE DID NOT RECOGNIZE THE RELOAD AS USED. FIRING THE SECOND RELOAD AGAIN WAS NOT ATTEMPTED. THE ISSUE CONTINUED TO EXIST. THERE WAS NO PATIENT INVOLVED.
06-01-2023 | Malfunction | covidien lp llc north haven | 23-01-2023 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE TWO RELOADS WERE FIRED WITHOUT ANY PROBLEMS. AFTER THE PROCEDURE, THE SAME RELOADS WERE CONNECTED AGAIN TO THE POWERED STAPLER FOR TRAINING PURPOSES. WHEN THE FIRST RELOAD (45MM) WAS CONNECTED AGAIN, THE HANDLE DID NOT RECOGNIZE THE RELOAD AS USED; ON THE DISPLAY, THE RELOAD WAS HIGHLIGHTED IN GREEN AND MARKED WITH ONE. IT WAS ALSO POSSIBLE TO SHOOT AND FIRE; THE BLADE MOVED ONLY A FEW MILLIMETERS AND THEN STOPPED. WHEN THE SECOND RELOAD (60MM) WAS CONNECTED AGAIN, IT HAD THE SAME ISSUE; THE HANDLE DID NOT RECOGNIZE THE RELOAD AS USED. FIRING THE SECOND RELOAD AGAIN WAS NOT ATTEMPTED. THE ISSUE CONTINUED TO EXIST. THERE WAS NO PATIENT INVOLVED.
08-12-2022 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED, POST-OP A SIGMOID COLECTOMY PROCEDURE EVERYTHING WENT AS IT SHOULD AND THE ANASTOMOSIS PASSED AN AIR LEAK TEST. SEVERAL DAYS LATER, THE PATIENT PRESENTED COMPLICATIONS OF A POSSIBLE LEAK. THE SURGEON PERFORMED AN EXPLORATORY LAPAROSCOPY BUT DID NOT FIND ANYTHING BUT PLACED A DRAIN IN THE PATIENT. THE PATIENT DID NOT IMPROVE SEVERAL DAYS AFTER THE FIRST REOPERATION. THE SECOND REOPERATION WAS CONDUCTED, AND A ONE-CENTIMETER HOLE WAS FOUND ON THE ANASTOMOSIS AND THE SURGEON PERFORMED A COLOSTOMY. TOMORROW ON JAN 6, THE SURGEON WILL NEED TO PERFORM A COLOSTOMY REVISION.
08-12-2022 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: IT WAS REPORTED, POST-OP A SIGMOID COLECTOMY PROCEDURE EVERYTHING WENT AS IT SHOULD AND THE ANASTOMOSIS PASSED AN AIR LEAK TEST. SEVERAL DAYS LATER, THE PATIENT PRESENTED COMPLICATIONS OF A POSSIBLE LEAK. THE SURGEON PERFORMED AN EXPLORATORY LAPAROSCOPY BUT DID NOT FIND ANYTHING BUT PLACED A DRAIN IN THE PATIENT. THE PATIENT DID NOT IMPROVE SEVERAL DAYS AFTER THE FIRST REOPERATION. THE SECOND REOPERATION WAS CONDUCTED, AND A ONE-CENTIMETER HOLE WAS FOUND ON THE ANASTOMOSIS AND THE SURGEON PERFORMED A COLOSTOMY. TOMORROW ON JAN 6, THE SURGEON WILL NEED TO PERFORM A COLOSTOMY REVISION.
08-12-2022 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Adhesion(s) | Event Description: IT WAS REPORTED, POST-OP A SIGMOID COLECTOMY PROCEDURE EVERYTHING WENT AS IT SHOULD AND THE ANASTOMOSIS PASSED AN AIR LEAK TEST. SEVERAL DAYS LATER, THE PATIENT PRESENTED COMPLICATIONS OF A POSSIBLE LEAK. THE SURGEON PERFORMED AN EXPLORATORY LAPAROSCOPY BUT DID NOT FIND ANYTHING BUT PLACED A DRAIN IN THE PATIENT. THE PATIENT DID NOT IMPROVE SEVERAL DAYS AFTER THE FIRST REOPERATION. THE SECOND REOPERATION WAS CONDUCTED, AND A ONE-CENTIMETER HOLE WAS FOUND ON THE ANASTOMOSIS AND THE SURGEON PERFORMED A COLOSTOMY. TOMORROW ON JAN 6, THE SURGEON WILL NEED TO PERFORM A COLOSTOMY REVISION.
08-12-2022 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED, POST-OP A SIGMOID COLECTOMY PROCEDURE EVERYTHING WENT AS IT SHOULD AND THE ANASTOMOSIS PASSED AN AIR LEAK TEST. SEVERAL DAYS LATER, THE PATIENT PRESENTED COMPLICATIONS OF A POSSIBLE LEAK. THE SURGEON PERFORMED AN EXPLORATORY LAPAROSCOPY BUT DID NOT FIND ANYTHING BUT PLACED A DRAIN IN THE PATIENT. THE PATIENT DID NOT IMPROVE SEVERAL DAYS AFTER THE FIRST REOPERATION. THE SECOND REOPERATION WAS CONDUCTED, AND A ONE-CENTIMETER HOLE WAS FOUND ON THE ANASTOMOSIS AND THE SURGEON PERFORMED A COLOSTOMY. TOMORROW ON JAN 6, THE SURGEON WILL NEED TO PERFORM A COLOSTOMY REVISION.
08-12-2022 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Inflammation | Event Description: IT WAS REPORTED, POST-OP A SIGMOID COLECTOMY PROCEDURE EVERYTHING WENT AS IT SHOULD AND THE ANASTOMOSIS PASSED AN AIR LEAK TEST. SEVERAL DAYS LATER, THE PATIENT PRESENTED COMPLICATIONS OF A POSSIBLE LEAK. THE SURGEON PERFORMED AN EXPLORATORY LAPAROSCOPY BUT DID NOT FIND ANYTHING BUT PLACED A DRAIN IN THE PATIENT. THE PATIENT DID NOT IMPROVE SEVERAL DAYS AFTER THE FIRST REOPERATION. THE SECOND REOPERATION WAS CONDUCTED, AND A ONE-CENTIMETER HOLE WAS FOUND ON THE ANASTOMOSIS AND THE SURGEON PERFORMED A COLOSTOMY. TOMORROW ON JAN 6, THE SURGEON WILL NEED TO PERFORM A COLOSTOMY REVISION.
08-12-2022 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: IT WAS REPORTED, POST-OP A SIGMOID COLECTOMY PROCEDURE EVERYTHING WENT AS IT SHOULD AND THE ANASTOMOSIS PASSED AN AIR LEAK TEST. SEVERAL DAYS LATER, THE PATIENT PRESENTED COMPLICATIONS OF A POSSIBLE LEAK. THE SURGEON PERFORMED AN EXPLORATORY LAPAROSCOPY BUT DID NOT FIND ANYTHING BUT PLACED A DRAIN IN THE PATIENT. THE PATIENT DID NOT IMPROVE SEVERAL DAYS AFTER THE FIRST REOPERATION. THE SECOND REOPERATION WAS CONDUCTED, AND A ONE-CENTIMETER HOLE WAS FOUND ON THE ANASTOMOSIS AND THE SURGEON PERFORMED A COLOSTOMY. TOMORROW ON JAN 6, THE SURGEON WILL NEED TO PERFORM A COLOSTOMY REVISION.
08-12-2022 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Urinary Tract Infection | Event Description: IT WAS REPORTED, POST-OP A SIGMOID COLECTOMY PROCEDURE EVERYTHING WENT AS IT SHOULD AND THE ANASTOMOSIS PASSED AN AIR LEAK TEST. SEVERAL DAYS LATER, THE PATIENT PRESENTED COMPLICATIONS OF A POSSIBLE LEAK. THE SURGEON PERFORMED AN EXPLORATORY LAPAROSCOPY BUT DID NOT FIND ANYTHING BUT PLACED A DRAIN IN THE PATIENT. THE PATIENT DID NOT IMPROVE SEVERAL DAYS AFTER THE FIRST REOPERATION. THE SECOND REOPERATION WAS CONDUCTED, AND A ONE-CENTIMETER HOLE WAS FOUND ON THE ANASTOMOSIS AND THE SURGEON PERFORMED A COLOSTOMY. TOMORROW ON JAN 6, THE SURGEON WILL NEED TO PERFORM A COLOSTOMY REVISION.
08-12-2022 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: IT WAS REPORTED, POST-OP A SIGMOID COLECTOMY PROCEDURE EVERYTHING WENT AS IT SHOULD AND THE ANASTOMOSIS PASSED AN AIR LEAK TEST. SEVERAL DAYS LATER, THE PATIENT PRESENTED COMPLICATIONS OF A POSSIBLE LEAK. THE SURGEON PERFORMED AN EXPLORATORY LAPAROSCOPY BUT DID NOT FIND ANYTHING BUT PLACED A DRAIN IN THE PATIENT. THE PATIENT DID NOT IMPROVE SEVERAL DAYS AFTER THE FIRST REOPERATION. THE SECOND REOPERATION WAS CONDUCTED, AND A ONE-CENTIMETER HOLE WAS FOUND ON THE ANASTOMOSIS AND THE SURGEON PERFORMED A COLOSTOMY. TOMORROW ON JAN 6, THE SURGEON WILL NEED TO PERFORM A COLOSTOMY REVISION.
26-12-2022 | Malfunction | covidien lp llc north haven | 23-01-2023 | TRI-STAPLE 2.0 | Premature Activation | A150103 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC BYPASS, THE BUTTRESS STRIP CAME OFF THE RELOAD AND FELL INTO THE PATIENT'S CAVITY WHEN POSITIONING THE STAPLER ON THE TISSUE. SURGEON USED A GRASPER TO RETRIEVE THE LOOSE BUTTRESS STRIP. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 23-01-2023 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC BYPASS, THE BUTTRESS STRIP CAME OFF THE RELOAD AND FELL INTO THE PATIENT'S CAVITY WHEN POSITIONING THE STAPLER ON THE TISSUE. SURGEON USED A GRASPER TO RETRIEVE THE LOOSE BUTTRESS STRIP. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A HARTMANN'S REVERSAL THE SURGEON WENT TO REVERSE THE COLOSTOMY, FIRED THE DEVICE, AND REMOVED IT. A LEAK TEST FAILED. HE ATTEMPTED TO SUTURE BUT IT WOULD NOT STOP LEAKING. A COLOSTOMY WAS PERFORMED.
03-01-2023 | Malfunction | ethicon endo-surgery | 23-01-2023 | ECHELON FLEX 60 LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: THE PRODUCTS DID NOT STAPLE AT THE TIME OF USE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED AN LAR FOR DIVERTICULITIS. BUTTRESSING WAS ATTACHED TO THE CIRCULAR STAPLER. THE ANASTOMOSIS PASSED AN AIR LEAK TEST AND THERE WERE TWO COMPLETE DONUTS IN THE STAPLER. ON (B)(6)2023, A RE-OPERATION WAS PERFORMED FOR A SUSPECTED LEAK. THE SURGEON DRAINED AN ABSCESS ON THE ANASTOMOSIS AND FOUND A SMALL HOLE. THE HOLE WAS REPAIRED, AND A DIVERTING LOOP ILEOSTOMY WAS PERFORMED. THE PATIENT IS RECOVERING.
01-01-2023 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: IT WAS REPORTED THAT THE SURGEON PERFORMED AN LAR FOR DIVERTICULITIS. BUTTRESSING WAS ATTACHED TO THE CIRCULAR STAPLER. THE ANASTOMOSIS PASSED AN AIR LEAK TEST AND THERE WERE TWO COMPLETE DONUTS IN THE STAPLER. ON (B)(6)2023, A RE-OPERATION WAS PERFORMED FOR A SUSPECTED LEAK. THE SURGEON DRAINED AN ABSCESS ON THE ANASTOMOSIS AND FOUND A SMALL HOLE. THE HOLE WAS REPAIRED, AND A DIVERTING LOOP ILEOSTOMY WAS PERFORMED. THE PATIENT IS RECOVERING.
08-12-2022 | Malfunction | davol inc sub c.r bard inc -1213643 | 23-01-2023 | CAPSURE STRAIGHT | Component Misassembled | A020601 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING USE OF THE CAPSURE STRAIGHT FIXATION DEVICE ON (B)(6) 2022, THE FIRST FASTENER DEPLOYED BACKWARDS AND DID NOT PROVIDE FIXATION. AS REPORTED, THE LOOSE FASTENER WAS REMOVED FROM THE PATIENT AND THE PROCEDURE WAS COMPLETED USING A NEW CAPSURE STRAIGHT DEVICE. THERE WAS NO REPORTED PATIENT INJURY.
24-02-2023 | Malfunction | covidien lp llc north haven | 23-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION. ALL THE PIECES WERE RETRIEVED FROM THE PATIENT BODY. THERE WAS NO PATIENT INJURY.
24-02-2023 | Malfunction | covidien lp llc north haven | 23-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN WHICH SIX RELOADS WERE USED. THERE WAS A DIGESTIVE BLOCK OF YELLOW PLASTIC TIP THAT WAS FOUND IN THE PATIENT AT THE END OF THE INTERVENTION. ALL THE PIECES WERE RETRIEVED FROM THE PATIENT BODY. THERE WAS NO PATIENT INJURY.
27-10-2022 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, THE FORCEPS HANDLE BROKE IN THE HAND OF THE SURGEON WHILE DEVICE WAS USED FOR FIRING. A SECOND ONE WAS USED AND IT WAS BROKEN AGAIN. THE CLAMP WAS REMOVED AND A THIRD HANDLE WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WERE NO CLINICAL CONSEQUENCES. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
08-01-2023 | Malfunction | covidien lp llc north haven | 23-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, WHILE STAPLING THE STOMACH, AT THE FIRST RELOAD OF STAPLING THE STOMACH, THE DOCTOR PRESSED THE GREEN BUTTON, BUT WAS UNABLE TO SQUEEZE THE STAPLER (GUN), AND THE DEVICE DID NOT FIRE. THE SURGEON TOOK THE STAPLER OUT OF THE PATIENT ABDOMEN. THE RELOAD WAS CHANGED AND SAME GUN WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE, AND AFTER THAT EVERYTHING WENT OKAY. THERE WAS NO PATIENT INJURY.
02-08-2022 | Malfunction | ethicon endo-surgery | 23-01-2023 | PROXIMATE*PPH PROCEDURE SET | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE HEMORRHOIDECTOMY THE PPH MISFIRED WHILE OPERATING. THE PROCEDURE WAS DELAYED BY 1 HOUR.
06-01-2023 | Malfunction | covidien lp llc north haven | 23-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY, THE FIRST TWO RELOADS WAS USED NORMALLY. WITH THE THIRD BACK RELOAD, THE DEVICE FIRED ON FIRST ACTIVATION BUT THE HANDLE COULD NOT BE SQUEEZED ON THE SECOND ACTIVATION. THE USER WAS ABLE TO OPEN THE CLAMP AND WAS REPLACED WITH A NEW EXTRA LONG HANDLE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE COMPONENT DISENGAGED. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
05-01-2023 | Injury | ethicon endo-surgery | 23-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A HARTMANN'S REVERSAL THE SURGEON WENT TO REVERSE THE COLOSTOMY, FIRED THE DEVICE, AND REMOVED IT. A LEAK TEST FAILED, AND IT WAS DISCOVERED THERE WAS A HOLE IN THE STAPLE LINE. THE PATIENT WAS OPENED BACK UP AND SUTURED AND THE SURGEON DID A DIVERTING LOOP COLOSTOMY.
09-01-2023 | Injury | covidien lp llc north haven | 23-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE DISSOCIATING BLOOD VESSELS, THE DEVICE WAS FIRED UNTIL THE END. THE DISTAL STAPLES FIRED, BUT SOME NAILS WERE NOT NAILED/FORMED. THE SURGEON SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE.
08-01-2023 | Injury | us surgical puerto rico | 23-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF GASTRIC CANCER, WHILE DISSOCIATING THE SMALL INTESTINE, THE DEVICE FIRED, SOME STAPLES WERE POORLY FORMED AS SOME WERE NOT NAILED, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. TO FIX THE STAPLE LINE, HAND STITCHING WAS DONE BY THE DOCTOR, THAT RESOLVE THE ISSUE AND COMPLETE THE CASE.
09-01-2023 | Injury | covidien lp llc north haven | 23-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL PROCEDURE, WHILE DISSOCIATING BLOOD VESSELS, THE DEVICE FIRED UNTIL THE END. THE DISTAL STAPLES FIRE, BUT THE DISTAL STAPLES HAD POOR STAPLE FORMATION; SOME NAILS WERE POORLY NAILED/FORMED. THE SURGEON SUTURED THE STAPLE LINE TO RESOLVE THE ISSUE AND USED A NEW RELOAD TO CUT AND ANASTOMOSE OTHER PARTS.
25-11-2022 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE BLACK RETRACTING KNOB WAS DIFFICULT TO RETRACT. THE SURGEON HAD TO USE EXCESSIVE FORCE TO RETRACT THE BLADE (AND THUS TO OPEN THE JAWS). A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. THE INITIAL RELOAD WAS TRIED ON THE NEW HANDLE BUT THE RELOAD STILL WOULD NOT FIRE. THE SURGERY TOOK A LITTLE LONGER DUE TO THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | covidien lp llc north haven | 23-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD AND HANDLE WERE USED TO COMPLETE THE CASE. THE INITIAL RELOAD WAS TRIED ON THE NEW HANDLE BUT THE RELOAD STILL WOULD NOT FIRE. THE SURGERY TOOK A LITTLE LONGER DUE TO THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ASSISTED CYSTOPROSTATECTOMY WITH CONSTRUCTION OF INTESTINAL CONDUIT OF BLADDER, WHILE STAPLING THE BOWEL, DURING THE THIRD FIRING, THE DEVICE DID NOT FIRE WHEN IT WAS LOCKED ON BOWEL TISSUE. THE SURGEON PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. WHEN THE STAPLER WAS REMOVED BY DOING A STANDARD ACTION TO OPEN THE STAPLER, THERE WAS NO VISIBLE DAMAGED TO THE BOWELS. BOTH HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Malfunction | covidien lp llc north haven | 23-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ASSISTED CYSTOPROSTATECTOMY WITH CONSTRUCTION OF INTESTINAL CONDUIT OF BLADDER, WHILE STAPLING THE BOWEL, DURING THE THIRD FIRING, THE DEVICE DID NOT FIRE WHEN IT WAS LOCKED ON BOWEL TISSUE. THE SURGEON PRESSED THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. WHEN THE STAPLER WAS REMOVED BY DOING A STANDARD ACTION TO OPEN THE STAPLER, THERE WAS NO VISIBLE DAMAGED TO THE BOWELS. BOTH HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROLYSE PROCEDURE, BOTH HANDLES ONLY TRIGGERED TWICE AND DID NOT WORK WHEN TRIGGERED THE THIRD TIME. THE HANDLES HAD A CRACKING NOISE WHEN TRYING TO FIRE THE THIRD TIME USING A NEW RELOAD. THE DEVICES WERE BLOCKED. FOR THE SECOND HANDLE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, AND THE GREEN BUTTON WAS NOT PRESSED. ALL MAGAZINES COULD BE RELEASED WITH A NEW HANDLE.
09-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROLYSE PROCEDURE, BOTH HANDLES ONLY TRIGGERED TWICE AND DID NOT WORK WHEN TRIGGERED THE THIRD TIME. THE HANDLES HAD A CRACKING NOISE WHEN TRYING TO FIRE THE THIRD TIME USING A NEW RELOAD. THE DEVICES WERE BLOCKED. FOR THE SECOND HANDLE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, AND THE GREEN BUTTON WAS NOT PRESSED. ALL MAGAZINES COULD BE RELEASED WITH A NEW HANDLE.
09-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROLYSE PROCEDURE, BOTH HANDLES ONLY TRIGGERED TWICE AND DID NOT WORK WHEN TRIGGERED THE THIRD TIME. THE HANDLES HAD A CRACKING NOISE WHEN TRYING TO FIRE THE THIRD TIME USING A NEW RELOAD. THE DEVICES WERE BLOCKED. FOR THE SECOND HANDLE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, AND THE GREEN BUTTON WAS NOT PRESSED. ALL MAGAZINES COULD BE RELEASED WITH A NEW HANDLE.
09-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROLYSE PROCEDURE, BOTH HANDLES ONLY TRIGGERED TWICE AND DID NOT WORK WHEN TRIGGERED THE THIRD TIME. THE HANDLES HAD A CRACKING NOISE WHEN TRYING TO FIRE THE THIRD TIME USING A NEW RELOAD. THE DEVICES WERE BLOCKED. FOR THE SECOND HANDLE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, AND THE GREEN BUTTON WAS NOT PRESSED. ALL MAGAZINES COULD BE RELEASED WITH A NEW HANDLE.
07-01-2023 | Malfunction | celestica de monterrey s.a de c.v | 23-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, ILEOSTOMY AND LIVER MASTECTOMY, DURING RESECTION, AFTER PRESSING THE GREEN BUTTON; THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. THE RELOAD WAS DIFFICULT TO UNLOAD AFTER THE FIRING ATTEMPT. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES SINCE THE SURGEON HAD TO WAIT FOR THE NEW PRODUCT TO COME. THE RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2023 | Malfunction | celestica de monterrey s.a de c.v | 23-01-2023 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, ILEOSTOMY AND LIVER MASTECTOMY, DURING RESECTION, AFTER PRESSING THE GREEN BUTTON; THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. THE RELOAD WAS DIFFICULT TO UNLOAD AFTER THE FIRING ATTEMPT. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES SINCE THE SURGEON HAD TO WAIT FOR THE NEW PRODUCT TO COME. THE RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2023 | Malfunction | celestica de monterrey s.a de c.v | 23-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, ILEOSTOMY AND LIVER MASTECTOMY, DURING RESECTION, AFTER PRESSING THE GREEN BUTTON; THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. THE RELOAD WAS DIFFICULT TO UNLOAD AFTER THE FIRING ATTEMPT. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES SINCE THE SURGEON HAD TO WAIT FOR THE NEW PRODUCT TO COME. THE RELOAD WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY, DURING THE FIFTH STAPLING, THE STAPLER LOCKED AND LOST THE RELOAD. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE; THE DEVICE DID NOT FIRE. BOTH HANDLE AND RELOAD WERE REPLACED IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-01-2023 | Malfunction | covidien lp llc north haven | 23-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY, DURING THE FIFTH STAPLING, THE STAPLER LOCKED AND LOST THE RELOAD. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT UNABLE TO SQUEEZE THE HANDLE; THE DEVICE DID NOT FIRE. BOTH HANDLE AND RELOAD WERE REPLACED IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-11-2022 | Injury | us surgical puerto rico | 23-01-2023 | EEA | Detachment of Device or Device Component | A0501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE OF THE LEFT COLON, EXTRACTION OF THE CIRCULAR STAPLER, THE HEAD DISENGAGED AFTER FIRING. THERE WAS NO COMPONENT THAT FELL INTO PATIENT'S CAVITY. IT WAS NOTED THAT DID NOT UNDERGONE A FURTHER RESECTION OF THE RECTUM. ANOTHER CIRCULAR STAPLER WAS USED TO RESOLVE THE ISSUE.
06-01-2023 | Malfunction | us surgical puerto rico | 23-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE GASTRECTOMY, THE FIRST TWO RELOADS WAS USED NORMALLY. WITH THE THIRD BACK RELOAD, THE DEVICE FIRED ON FIRST ACTIVATION BUT THE HANDLE COULD NOT BE SQUEEZED ON THE SECOND ACTIVATION. THE USER WAS ABLE TO OPEN THE CLAMP AND WAS REPLACED WITH A NEW EXTRA LONG HANDLE. THERE WAS NO PATIENT INJURY.
01-01-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS THAT DURING AN UNKNOWN PROCEDURE TWO DEVICES MISFIRED. THE SURGEON COULD NOT SQUEEZE THE TRIGGER AT ALL ON THE FIRST DEVICE AND ONLY PARTIALLY ON THE SECOND. A THIRD DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
22-11-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-01-2023 | Injury | covidien lp llc north haven | 24-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE FIRST FIRING FROM THE GREATER CURVATURE SIDE, THE FIRING STOPPED WHEN THE STAPLING ADVANCED FOR ABOUT 30MM. AT THAT TIME, THE DISPLAY SHOWED AN INDICATION MARK OF 60 DEGREES. BASED ON THE IMAGE, THERE WAS AN EXCESSIVE LOAD DETECTED DURING USE. THE RELOAD WAS RELEASED AND THE RECOVERY SHOT WAS CONTINUED USING A DIFFERENT RELOAD. A PORTION OF THE STAPLE LINE WAS NOT IN A NORMAL CONDITION, THE FIRED STAPLES OVERLAPPED, AND AN ADDITIONAL SUTURING WAS MADE ON THE AREA WHERE FIRING WAS PERFORMED HALFWAY. A REINFORCED RELOAD WAS SUCCESSFULLY USED AGAIN IN THE LAST FIRING.
09-01-2023 | Injury | covidien lp llc north haven | 24-01-2023 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE FIRST FIRING FROM THE GREATER CURVATURE SIDE, THE FIRING STOPPED WHEN THE STAPLING ADVANCED FOR ABOUT 30MM. AT THAT TIME, THE DISPLAY SHOWED AN INDICATION MARK OF 60 DEGREES. BASED ON THE IMAGE, THERE WAS AN EXCESSIVE LOAD DETECTED DURING USE. THE RELOAD WAS RELEASED AND THE RECOVERY SHOT WAS CONTINUED USING A DIFFERENT RELOAD. A PORTION OF THE STAPLE LINE WAS NOT IN A NORMAL CONDITION, THE FIRED STAPLES OVERLAPPED, AND AN ADDITIONAL SUTURING WAS MADE ON THE AREA WHERE FIRING WAS PERFORMED HALFWAY. A REINFORCED RELOAD WAS SUCCESSFULLY USED AGAIN IN THE LAST FIRING.
03-01-2023 | Injury | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PARTIAL GASTRECTOMY, WHILE TRANSACTIONING THE STOMACH, THE FIRST HANDLE CLOSED DOWN ON THE TISSUE DURING FIRING AND WOULD NOT OPEN BACK UP. THE SURGEON RESULTED IN OPENING TWO MORE HANDLES TO REMOVE THE SPECIMEN (IN WHICH THE HANDLE WAS STILL STUCK ONTO THE TISSUE WHEN THE SPECIMEN CAME OUT). SOMEHOW OPENING IT WAS EVENTUALLY MANAGED. THE SECOND HANDLE DID NOT PROVIDE GREAT FEEDBACK AND MADE AN UNUSUAL NOISE. THE THIRD HANDLE DID NOT PROVIDE GOOD FEEDBACK, BUT THE DOCTOR STILL CONTINUED TO USE THE LAST HANDLE TO COMPLETE THE PROCEDURE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Injury | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PARTIAL GASTRECTOMY, WHILE TRANSACTIONING THE STOMACH, THE FIRST HANDLE CLOSED DOWN ON THE TISSUE DURING FIRING AND WOULD NOT OPEN BACK UP. THE SURGEON RESULTED IN OPENING TWO MORE HANDLES TO REMOVE THE SPECIMEN (IN WHICH THE HANDLE WAS STILL STUCK ONTO THE TISSUE WHEN THE SPECIMEN CAME OUT). SOMEHOW OPENING IT WAS EVENTUALLY MANAGED. THE SECOND HANDLE DID NOT PROVIDE GREAT FEEDBACK AND MADE AN UNUSUAL NOISE. THE THIRD HANDLE DID NOT PROVIDE GOOD FEEDBACK, BUT THE DOCTOR STILL CONTINUED TO USE THE LAST HANDLE TO COMPLETE THE PROCEDURE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Injury | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A PARTIAL GASTRECTOMY, WHILE TRANSACTIONING THE STOMACH, THE FIRST HANDLE CLOSED DOWN ON THE TISSUE DURING FIRING AND WOULD NOT OPEN BACK UP. THE SURGEON RESULTED IN OPENING TWO MORE HANDLES TO REMOVE THE SPECIMEN (IN WHICH THE HANDLE WAS STILL STUCK ONTO THE TISSUE WHEN THE SPECIMEN CAME OUT). SOMEHOW OPENING IT WAS EVENTUALLY MANAGED. THE SECOND HANDLE DID NOT PROVIDE GREAT FEEDBACK AND MADE AN UNUSUAL NOISE. THE THIRD HANDLE DID NOT PROVIDE GOOD FEEDBACK, BUT THE DOCTOR STILL CONTINUED TO USE THE LAST HANDLE TO COMPLETE THE PROCEDURE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Injury | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PARTIAL GASTRECTOMY, WHILE TRANSACTIONING THE STOMACH, THE FIRST HANDLE CLOSED DOWN ON THE TISSUE DURING FIRING AND WOULD NOT OPEN BACK UP. THE SURGEON RESULTED IN OPENING TWO MORE HANDLES TO REMOVE THE SPECIMEN (IN WHICH THE HANDLE WAS STILL STUCK ONTO THE TISSUE WHEN THE SPECIMEN CAME OUT). SOMEHOW OPENING IT WAS EVENTUALLY MANAGED. THE SECOND HANDLE DID NOT PROVIDE GREAT FEEDBACK AND MADE AN UNUSUAL NOISE. THE THIRD HANDLE DID NOT PROVIDE GOOD FEEDBACK, BUT THE DOCTOR STILL CONTINUED TO USE THE LAST HANDLE TO COMPLETE THE PROCEDURE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | ECHELON FLEX | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: DURING THE CASE, THE ECHELON POWERED STAPLER WAS HARD TO CLOSE ON THE STOMACH. UPON CLOSING, IT MADE A LOUD CLICK AS IF SOMETHING BROKE AND THE STAPLER WOULD NOT FIRE THE STAPLE LOAD WHEN IT WAS PLACED ON THE PATIENT'S STOMACH. STAPLER WAS ABLE TO BE REMOVED FROM THE STOMACH AND NO INJURY WAS NOTED.
21-12-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | ECHELON FLEX | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: DURING THE CASE, THE ECHELON POWERED STAPLER WAS HARD TO CLOSE ON THE STOMACH. UPON CLOSING, IT MADE A LOUD CLICK AS IF SOMETHING BROKE AND THE STAPLER WOULD NOT FIRE THE STAPLE LOAD WHEN IT WAS PLACED ON THE PATIENT'S STOMACH. STAPLER WAS ABLE TO BE REMOVED FROM THE STOMACH AND NO INJURY WAS NOTED.
06-01-2023 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, AFTER FIRING THE OTHER RELOADS WHILE STAPLING SOME THICK AND TOUGH TISSUE, A PURPLE 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE HANDLE. WHEN FIRING WAS ATTEMPTED, THE DEVICE DID NOT FIRE AND THE HANDLE WAS VISUALLY BROKEN. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A SECOND HANDLE WAS USED WITH THE SAME RELOAD, BUT THE SAME ISSUES OCCURRED. A THIRD HANDLE AND A NEW RELOAD WAS THEN USED WITHOUT ANY ISSUES. LATER IN THE PROCEDURE, THE SAME 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE THIRD HANDLE. WHEN FIRED, THE SAME ISSUES OCCURRED. DUE TO THICK TISSUE, A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, AFTER FIRING THE OTHER RELOADS WHILE STAPLING SOME THICK AND TOUGH TISSUE, A PURPLE 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE HANDLE. WHEN FIRING WAS ATTEMPTED, THE DEVICE DID NOT FIRE AND THE HANDLE WAS VISUALLY BROKEN. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A SECOND HANDLE WAS USED WITH THE SAME RELOAD, BUT THE SAME ISSUES OCCURRED. A THIRD HANDLE AND A NEW RELOAD WAS THEN USED WITHOUT ANY ISSUES. LATER IN THE PROCEDURE, THE SAME 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE THIRD HANDLE. WHEN FIRED, THE SAME ISSUES OCCURRED. DUE TO THICK TISSUE, A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, AFTER FIRING THE OTHER RELOADS WHILE STAPLING SOME THICK AND TOUGH TISSUE, A PURPLE 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE HANDLE. WHEN FIRING WAS ATTEMPTED, THE DEVICE DID NOT FIRE AND THE HANDLE WAS VISUALLY BROKEN. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A SECOND HANDLE WAS USED WITH THE SAME RELOAD, BUT THE SAME ISSUES OCCURRED. A THIRD HANDLE AND A NEW RELOAD WAS THEN USED WITHOUT ANY ISSUES. LATER IN THE PROCEDURE, THE SAME 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE THIRD HANDLE. WHEN FIRED, THE SAME ISSUES OCCURRED. DUE TO THICK TISSUE, A NEWRELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, AFTER FIRING THE OTHER RELOADS WHILE STAPLING SOME THICK AND TOUGH TISSUE, A PURPLE 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE HANDLE. WHEN FIRING WAS ATTEMPTED, THE DEVICE DID NOT FIRE AND THE HANDLE WAS VISUALLY BROKEN. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A SECOND HANDLE WAS USED WITH THE SAME RELOAD, BUT THE SAME ISSUES OCCURRED. A THIRD HANDLE AND A NEW RELOAD WAS THEN USED WITHOUT ANY ISSUES. LATER IN THE PROCEDURE, THE SAME 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE THIRD HANDLE. WHEN FIRED, THE SAME ISSUES OCCURRED. DUE TO THICK TISSUE, A NEWRELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2022 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ATRIAL APPENDAGE, WHEN TRYING TO STAPLE THE HEART EAR, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. NO PATIENT INJURY.
28-12-2022 | Malfunction | covidien lp llc north haven | 24-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ATRIAL APPENDAGE, WHEN TRYING TO STAPLE THE HEART EAR, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. NO PATIENT INJURY.
12-12-2022 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) RESECTION PROCEDURE, WHILE STAPLING THE LUNG, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | covidien lp llc north haven | 24-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) RESECTION PROCEDURE, WHILE STAPLING THE LUNG, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-12-2022 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC CECAL RESECTION, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
29-12-2022 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC CECAL RESECTION, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, AFTER FIRING THE OTHER RELOADS WHILE STAPLING SOME THICK AND TOUGH TISSUE, A PURPLE 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE HANDLE. WHEN FIRING WAS ATTEMPTED, THE DEVICE DID NOT FIRE AND THE HANDLE WAS VISUALLY BROKEN; IT CRACKED. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A SECOND HANDLE WAS USED WITH THE SAME RELOAD, BUT THE SAME ISSUES OCCURRED. A THIRD HANDLE AND A NEW RELOAD WAS THEN USED WITHOUT ANY ISSUES. LATER IN THE PROCEDURE, THE SAME 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE THIRD HANDLE AND WHEN IT WAS FIRED ON THE PATIENT, THE SAME ISSUES OCCURRED. DUE TO THICK TISSUE, A NEW RELOAD AND A FOURTH HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, AFTER FIRING THE OTHER RELOADS WHILE STAPLING SOME THICK AND TOUGH TISSUE, A PURPLE 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE HANDLE. WHEN FIRING WAS ATTEMPTED, THE DEVICE DID NOT FIRE AND THE HANDLE WAS VISUALLY BROKEN; IT CRACKED. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A SECOND HANDLE WAS USED WITH THE SAME RELOAD, BUT THE SAME ISSUES OCCURRED. A THIRD HANDLE AND A NEW RELOAD WAS THEN USED WITHOUT ANY ISSUES. LATER IN THE PROCEDURE, THE SAME 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE THIRD HANDLE AND WHEN IT WAS FIRED ON THE PATIENT, THE SAME ISSUES OCCURRED. DUE TO THICK TISSUE, A NEW RELOAD AND A FOURTH HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | covidien lp llc north haven | 24-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, AFTER FIRING THE OTHER RELOADS WHILE STAPLING SOME THICK AND TOUGH TISSUE, A PURPLE 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE HANDLE. WHEN FIRING WAS ATTEMPTED, THE DEVICE DID NOT FIRE AND THE HANDLE WAS VISUALLY BROKEN. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A SECOND HANDLE WAS USED WITH THE SAME RELOAD, BUT THE SAME ISSUES OCCURRED. A THIRD HANDLE AND A NEW RELOAD WAS THEN USED WITHOUT ANY ISSUES. LATER IN THE PROCEDURE, THE SAME 60MM RELOAD WITHOUT TIP WAS LOADED INTO THE THIRD HANDLE. WHEN FIRED, THE SAME ISSUES OCCURRED. DUE TO THICK TISSUE, A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 24-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS IMPOSSIBLE TO STAPLE AFTER THE SURGEON SQUEEZED THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE.
30-11-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2023 | Malfunction | covidien lp llc north haven | 24-01-2023 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, DURING GASTRIC SECRETION, THE DEVICE MAKES A STRANGE SOUND, OUT OF ORDINARY. THE SURGEON ALSO REPORTED THAT THE CHARGE MAKES A ZIGZAG-LIKE MOVEMENT DURING THE SHOT. THE CASE WAS COMPLETED WITH THE SAME DEVICE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | covidien lp llc north haven | 24-01-2023 | SIGNIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, DURING GASTRIC SECRETION, THE DEVICE MAKES A STRANGE SOUND, OUT OF ORDINARY. THE SURGEON ALSO REPORTED THAT THE CHARGE MAKES A ZIGZAG-LIKE MOVEMENT DURING THE SHOT. THE CASE WAS COMPLETED WITH THE SAME DEVICE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | covidien lp llc north haven | 24-01-2023 | ENDO GIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, DURING GASTRIC SECRETION, THE DEVICE MAKES A STRANGE SOUND, OUT OF ORDINARY. THE SURGEON ALSO REPORTED THAT THE LOAD MOVES DURING SHOOTING. THE CASE WAS COMPLETED WITH THE SAME DEVICE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | covidien lp llc north haven | 24-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, DURING GASTRIC SECRETION, THE DEVICE MAKES A STRANGE SOUND, OUT OF ORDINARY. THE SURGEON ALSO REPORTED THAT THE LOAD MOVES DURING SHOOTING. THE CASE WAS COMPLETED WITH THE SAME DEVICE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | ethicon endo-surgery | 24-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP RIGHT HEMI, ON THE ILEO-COLIC VASCULAR PEDICLE. IT WAS SLIGHTLY BULKY/FATTY, BUT THE STAPLER CLOSED NORMALLY, HOWEVER IT DID NOT FIRE COMPLETELY. WE WERE ABLE TO OPEN AND REMOVE IT, AND FOUND SOME ATTACHED AND SOME LOOSE STAPLES, BUT THE VESSELS HAD NOT BEEN CUT. WE THEN USED A WHOLE NEW STAPLER AND CARTRIDGE (WHITE, AS BEFORE) AT THE SAME PLACE, AND THIS CLOSED AND FIRED WITHOUT ANY PROBLEMS.
03-01-2023 | Malfunction | ethicon endo-surgery | 24-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP RIGHT HEMI, ON THE ILEO-COLIC VASCULAR PEDICLE. IT WAS SLIGHTLY BULKY/FATTY, BUT THE STAPLER CLOSED NORMALLY, HOWEVER IT DID NOT FIRE COMPLETELY. WE WERE ABLE TO OPEN AND REMOVE IT, AND FOUND SOME ATTACHED AND SOME LOOSE STAPLES, BUT THE VESSELS HAD NOT BEEN CUT. WE THEN USED A WHOLE NEW STAPLER AND CARTRIDGE (WHITE, AS BEFORE) AT THE SAME PLACE, AND THIS CLOSED AND FIRED WITHOUT ANY PROBLEMS.
27-12-2022 | Injury | intuitive surgical | 24-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DAVINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, STAPLES DID NOT FORM CORRECTLY ON TISSUE WHEN A BLUE SUREFORM 45 STAPLER RELOAD WAS FIRED WITH A SUREFORM 45 STAPLER INSTRUMENT. AS A RESULT, THE TISSUE WAS PUSHED OUT AND TORN, AND ADDITIONAL UNPLANNED STAPLING WAS REQUIRED. THE SURGEON WAS ABLE TO RETRIEVE THE UNFORMED STAPLES. INTUITIVE SURGICAL, INC. (ISI) OBTAINED THE FOLLOWING ADDITIONAL INFORMATION FROM THE SURGEON REGARDING THE REPORTED EVENT: PER THE SURGEON, THE STAPLE WAS INSPECTED PRIOR TO THE PROCEDURE, AND NO ISSUES WERE FOUND. IT WAS STATED THAT THE 1ST FIRE WAS NORMAL. DURING THE 2ND FIRE, "A VERY SMALL PIECE OF THE RELOAD (BLUE) BROKE OFF AND FELL INTO THE PATIENT," BUT THE FRAGMENT WAS RETRIEVED DURING THE PROCEDURE. IT WAS CONFIRMED THAT THE STAPLE LINE WAS COMPLETED. THE SURGEON CONFIRMED THAT THE EVENT WHEN THE TISSUE WAS PUSHED OUT AND TORN OCCURRED DURING THE 3RD FIRE WHILE USING A BLUE SUREFORM 45 RELOAD, WHICH IS HIS STANDARD CHOICE FOR THAT STEP OF THE PROCEDURE. AT THE TIME OF THE EVENT, THE SURGEON WAS CREATING THE GASTRIC POUCH AND STAPLING GASTRIC TISSUE. THE TISSUE WAS NOT CALCIFIED OR EXPOSED TO RADIATION OR CHEMOTHERAPY, AND THERE WAS NO TISSUE TENSION. THERE WAS ALSO NO TISSUE BUNCHING "UNTIL THE BLADE PUSHED THE TISSUE FORWARD DURING FIRING AND AT THAT TIME, THE TISSUE BECAME BUNCHED." PER THE SURGEON, THIS ISSUE RESULTED IN HOLES IN THE STAPLE LINE AND A LEAK FROM THE GASTRIC REMNANT, WHICH WAS ADDRESSED VIA ADDITIONAL RESECTION OF GASTRIC TISSUE AND RE-STAPLING THE AFFECTED AREA UTILIZING A NEW STAPLER. THE EVENT DID NOT INVOLVE A FAILURE TO FIRE NOR A PARTIAL FIRE, AS THE STAPLER "TRIED TO FIRE ALL THE WAY." THE STAPLER JAWS DID NOT OPEN/RELEASE DURING THE FIRING SEQUENCE, AND NO STAPLES WERE DEPLOYED FROM THE RELOAD UNEXPECTEDLY. THERE WERE NO ERRORS OR MESSAGES GENERATED WHEN THE ISSUE OCCURRED. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS, NOR DID HE FIRE OVER AN EXISTING STAPLE LINE. NO BUTTRESS MATERIAL WAS USED. THE SURGEON DID NOT EXPERIENCE ANY CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD.
27-12-2022 | Injury | intuitive surgical | 24-01-2023 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DAVINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, STAPLES DID NOT FORM CORRECTLY ON TISSUE WHEN A BLUE SUREFORM 45 STAPLER RELOAD WAS FIRED WITH A SUREFORM 45 STAPLER INSTRUMENT. AS A RESULT, THE TISSUE WAS PUSHED OUT AND TORN, AND ADDITIONAL UNPLANNED STAPLING WAS REQUIRED. THE SURGEON WAS ABLE TO RETRIEVE THE UNFORMED STAPLES. INTUITIVE SURGICAL, INC. (ISI) OBTAINED THE FOLLOWING ADDITIONAL INFORMATION FROM THE SURGEON REGARDING THE REPORTED EVENT: PER THE SURGEON, THE STAPLE WAS INSPECTED PRIOR TO THE PROCEDURE, AND NO ISSUES WERE FOUND. IT WAS STATED THAT THE 1ST FIRE WAS NORMAL. DURING THE 2ND FIRE, "A VERY SMALL PIECE OF THE RELOAD (BLUE) BROKE OFF AND FELL INTO THE PATIENT," BUT THE FRAGMENT WAS RETRIEVED DURING THE PROCEDURE. IT WAS CONFIRMED THAT THE STAPLE LINE WAS COMPLETED. THE SURGEON CONFIRMED THAT THE EVENT WHEN THE TISSUE WAS PUSHED OUT AND TORN OCCURRED DURING THE 3RD FIRE WHILE USING A BLUE SUREFORM 45 RELOAD, WHICH IS HIS STANDARD CHOICE FOR THAT STEP OF THE PROCEDURE. AT THE TIME OF THE EVENT, THE SURGEON WAS CREATING THE GASTRIC POUCH AND STAPLING GASTRIC TISSUE. THE TISSUE WAS NOT CALCIFIED OR EXPOSED TO RADIATION OR CHEMOTHERAPY, AND THERE WAS NO TISSUE TENSION. THERE WAS ALSO NO TISSUE BUNCHING "UNTIL THE BLADE PUSHED THE TISSUE FORWARD DURING FIRING AND AT THAT TIME, THE TISSUE BECAME BUNCHED." PER THE SURGEON, THIS ISSUE RESULTED IN HOLES IN THE STAPLE LINE AND A LEAK FROM THE GASTRIC REMNANT, WHICH WAS ADDRESSED VIA ADDITIONAL RESECTION OF GASTRIC TISSUE AND RE-STAPLING THE AFFECTED AREA UTILIZING A NEW STAPLER. THE EVENT DID NOT INVOLVE A FAILURE TO FIRE NOR A PARTIAL FIRE, AS THE STAPLER "TRIED TO FIRE ALL THE WAY." THE STAPLER JAWS DID NOT OPEN/RELEASE DURING THE FIRING SEQUENCE, AND NO STAPLES WERE DEPLOYED FROM THE RELOAD UNEXPECTEDLY. THERE WERE NO ERRORS OR MESSAGES GENERATED WHEN THE ISSUE OCCURRED. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS, NOR DID HE FIRE OVER AN EXISTING STAPLE LINE. NO BUTTRESS MATERIAL WAS USED. THE SURGEON DID NOT EXPERIENCE ANY CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD.
09-01-2023 | Injury | covidien lp llc north haven | 24-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, WHILE DISSOCIATING LUNG LOBES, THE DEVICE FIRED, BUT NOT ALL STAPLES WERE B-SHAPED AND STAPLE LINE WAS INCOMPLETE CENTRALLY. THE STAPLE LINE WAS FIXED WHEN SUTURED. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
09-01-2023 | Malfunction | covidien lp llc north haven | 24-01-2023 | ABSORBATACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC HERNIA PROCEDURE, WHILE ON FIXATION OF THE MESH, THE SHAFT BROKE WHEN APPROACHING THE POINT TO FIRE THE TACK. A NEW TACKER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE DOCTOR WENT TO FIRE, IT WOULD NOT FIRE AND THE JAWS WOULD NOT OPEN. HE PULLED THE EMERGENCY RELEASE AND IT STILL DID NOT OPEN. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
27-12-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE DOCTOR WENT TO FIRE, IT WOULD NOT FIRE AND THE JAWS WOULD NOT OPEN. HE PULLED THE EMERGENCY RELEASE AND IT STILL DID NOT OPEN. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
27-12-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE DOCTOR WENT TO FIRE, IT WOULD NOT FIRE AND THE JAWS WOULD NOT OPEN. HE PULLED THE EMERGENCY RELEASE AND IT STILL DID NOT OPEN. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
02-01-2023 | Malfunction | ethicon endo-surgery | 24-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE STAPLER MISFIRED AND SHOT STAPLES OUT. STATED THAT THE TISSUE WAS NOT STAPLED. UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. UNKNOWN IF THERE WERE ANY ADVERSE CONSEQUENCES FOR THE PATIENT.
04-11-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | ECHELON FLEX 60 LONG | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-01-2023 | Malfunction | ethicon endo-surgery | 24-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED REP THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, SURGEON FIRED A PLEE60A ON A GREEN RELOAD, STAPLES WERE ONLY DEPLOYED ON THE LEFT HALF OF THE STAPLE LINE AND THUS DID NOT FORM A PROPER B SHAPE ON THE RIGHT SIDE. STAPLER CONTINUED TO CUT WITHOUT BEING ABLE TO SEAL DUE TO THE LACK OF ABILITY TO STAPLE WITH THIS MALFUNCTION. NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCES.
03-01-2023 | Malfunction | ethicon endo-surgery | 24-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED REP THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, SURGEON FIRED A PLEE60A ON A GREEN RELOAD, STAPLES WERE ONLY DEPLOYED ON THE LEFT HALF OF THE STAPLE LINE AND THUS DID NOT FORM A PROPER B SHAPE ON THE RIGHT SIDE. STAPLER CONTINUED TO CUT WITHOUT BEING ABLE TO SEAL DUE TO THE LACK OF ABILITY TO STAPLE WITH THIS MALFUNCTION. NEW STAPLER AND RELOAD WERE USED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCES.
04-01-2023 | Malfunction | ethicon endo-surgery | 24-01-2023 | GST BLUE RELOAD, 60MM, 6 ROW | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE RELOAD SEPARATED FROM THE SILVER SLED UNDER THE RELOAD. UNKNOWN AS TO HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. PICTURES AVAILABLE.
16-11-2022 | Malfunction | ethicon endo-surgery | 24-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: PLEASE SEE ATTACHED (B)(4).
20-01-2023 | Malfunction | ethicon endo-surgery llc / ethicon endo-surgery | 24-01-2023 | ENDOPATH ETS-FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: DURING THE SURGICAL PROCEDURE THE SURGEON HAD NOT FIRED THE STAPLER REALIZING HE STILL HAD ADDITIONAL WORK TO DO. WHEN HE HANDED THE DEVICE BACK TO THE SCRUB, THE JAWS WERE CLOSED. THE SCRUB PUSHED THE RELEASE BUTTON AND HEARD THE NORMAL CLICKING NOISE, BUT THE JAWS WOULD NOT OPEN. WHEN THE SURGEON WAS ABLE TO OPEN THE JAWS, FOLLOWING SEVERAL TRIES, HE WAS THEN NOT ABLE TO CLOSE THE JAWS.
19-12-2022 | Malfunction | covidien lp | 25-01-2023 | GIA | Device Damaged Prior to Use | A0204 | Insufficient Information | Event Description: COVIDIEN 2/3 COVIDIEN 80 BLUE RELOADS ACCORDING TO DR. [REDACTED NAME], "APPEARED TO BE PRE-FIRED OUT OF PACKAGE". BOTH RELOADS ARE FROM SAME LOT#.
19-12-2022 | Malfunction | covidien lp | 25-01-2023 | GIA | Firing Problem | A0505 | Insufficient Information | Event Description: COVIDIEN 2/3 COVIDIEN 80 BLUE RELOADS ACCORDING TO DR. [REDACTED NAME], "APPEARED TO BE PRE-FIRED OUT OF PACKAGE". BOTH RELOADS ARE FROM SAME LOT#.
10-01-2023 | Injury | covidien lp llc north haven | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, AT THE TIME OF LUNG SEGMENT RESECTION, THE DEVICE FIRED UNTIL THE END, BUT THE DISTAL STAPLES DID NOT FIRE AT ALL. THE SURGEON PERFORMED RE-ANASTOMOSIS USING A NEW RELOAD TO RESOLVE THE ISSUE.
10-01-2023 | Injury | covidien lp llc north haven | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, AT THE TIME OF LUNG SEGMENT RESECTION, THE DEVICE FIRED UNTIL THE END, BUT THE DISTAL STAPLES DID NOT FIRE AT ALL. THE SURGEON PERFORMED RE-ANASTOMOSIS USING A NEW RELOAD TO RESOLVE THE ISSUE.
06-01-2023 | Malfunction | us surgical puerto rico | 25-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LAPAROTOMY RECTAL RESECTION, WHEN APPROACHING THE ARTICULATION IN THE RECTUM, THE JAW WAS CLOSED BUT IT DID NOT CLOSE PROPERLY. THE CONNECTION BETWEEN THE HANDLE AND THE RELOAD WAS BAD. THE JAW WAS OPENED AND REPLACED WITH A NEW DEVICE. THE NEW HANDLE AND RELOAD WORKED WITHOUT A PROBLEM WITH THE SAME TISSUE THICKNESS. THE TISSUE WAS NEITHER HARD NOR THICK. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | covidien lp llc north haven | 25-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AT THE TIME OF STAPLING THE TISSUE, THE TWO RELOADS FIRED HALFWAY AND THE RELOADS WERE STUCK IN THE HANDLE. A NEW RELOAD FROM A DIFFERENT BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | covidien lp llc north haven | 25-01-2023 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AT THE TIME OF STAPLING THE TISSUE, THE TWO RELOADS FIRED HALFWAY AND THE RELOADS WERE STUCK IN THE HANDLE. A NEW RELOAD FROM A DIFFERENT BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | covidien lp llc north haven | 25-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AT THE TIME OF STAPLING THE TISSUE, THE TWO RELOADS FIRED HALFWAY AND THE RELOADS WERE STUCK IN THE HANDLE. A NEW RELOAD FROM A DIFFERENT BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | covidien lp llc north haven | 25-01-2023 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, AT THE TIME OF STAPLING THE TISSUE, THE TWO RELOADS FIRED HALFWAY AND THE RELOADS WERE STUCK IN THE HANDLE. A NEW RELOAD FROM A DIFFERENT BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | ethicon endo-surgery | 25-01-2023 | ECHELON CIRCULAR | Defective Device | A0203 | No Clinical Signs, Symptoms or Conditions | Event Description: DEFECTIVE STAPLER, STAPLER DID NOT FIRE.
20-12-2022 | Malfunction | ethicon endo-surgery | 25-01-2023 | ECHELON CIRCULAR | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: DEFECTIVE STAPLER, STAPLER DID NOT FIRE.
09-01-2023 | Malfunction | us surgical puerto rico | 25-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AFTER THE FIRST FIRING, THE RELOAD WAS LOADED ONTO THE LINEAR STAPLER APPROPRIATELY WHEN THE SURGEON GRASPED THE APPENDIX AND PLACED THE STAPLER AROUND IT, CLOSED THE JAWS, AND ATTEMPTED TO FIRE WHEN A POP WAS HEARD. THE JAWS COULD NOT BE OPENED. THE TWO SURGEONS WERE ABLE TO DISENGAGE THE RELOAD FROM THE HANDLE AND REMOVE THE RELOAD AND THE APPENDIX TOGETHER. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES AND MORE TO REMOVE THE RELOAD AND CONTINUING WITH OTHER OPTION TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 25-01-2023 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AFTER THE FIRST FIRING, THE RELOAD WAS LOADED ONTO THE LINEAR STAPLER APPROPRIATELY WHEN THE SURGEON GRASPED THE APPENDIX AND PLACED THE STAPLER AROUND IT, CLOSED THE JAWS, AND ATTEMPTED TO FIRE WHEN A POP WAS HEARD. THE JAWS COULD NOT BE OPENED. THE TWO SURGEONS WERE ABLE TO DISENGAGE THE RELOAD FROM THE HANDLE AND REMOVE THE RELOAD AND THE APPENDIX TOGETHER. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES AND MORE TO REMOVE THE RELOAD AND CONTINUING WITH OTHER OPTION TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 25-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AFTER THE FIRST FIRING, THE RELOAD WAS LOADED ONTO THE LINEAR STAPLER APPROPRIATELY WHEN THE SURGEON GRASPED THE APPENDIX AND PLACED THE STAPLER AROUND IT, CLOSED THE JAWS, AND ATTEMPTED TO FIRE WHEN A POP WAS HEARD. THE JAWS COULD NOT BE OPENED. THE TWO SURGEONS WERE ABLE TO DISENGAGE THE RELOAD FROM THE HANDLE AND REMOVE THE RELOAD AND THE APPENDIX TOGETHER. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES AND MORE TO REMOVE THE RELOAD AND CONTINUING WITH OTHER OPTION TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, IN THE ANTRUM OF THE STOMACH, THE RELOAD MISFIRE. THERE WAS A INCOMPLETE DISTAL STAPLE LINE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED.
10-12-2022 | Malfunction | us surgical puerto rico | 25-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, WHEN THE SURGEON WAS STARTING TO STAPLE THE GASTRIC TISSUE, THE STAPLER DID NOT FIRE. THIS WAS TRIED TWO TIMES. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. ANOTHER KIND OF DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2023 | Injury | davol inc sub c.r bard inc -1213643 | 25-01-2023 | CAPSURE STRAIGHT | Patient-Device Incompatibility | A0101 | Abscess | Event Description: AS REPORTED, DURING TAPP LAPAROSCOPIC PROCEDURE, CAPSURE STRAIGHT FIXATION WAS USED TO FIXATE THE 3DMAX LIGHT. IT WAS REPORTED THAT THE PATIENT DEVELOPED ABSCESS POST IMPLANTATION AND UNDERWENT ADDITIONAL SURGERY FOR REMOVAL OF THE MESH AND FASTENERS ON (B)(6) 2023.
03-01-2023 | Malfunction | us surgical puerto rico | 25-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHEN TRYING TO TRANSECT APPENDIX, THE SURGEON COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
05-01-2023 | Malfunction | ethicon endo-surgery | 25-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE KNIFE MOVED TO THE TIP, BUT DID NOT RETURN TO THE HOME POSITION AT 1ST FIRING WHEN CUTTING THE SIGMOID. KNIFE REVERSE SWITCH DID NOT WORK, AND MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE STAPLE WAS STAPLED PROPERLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE DEVICE WAS USED ON THE COLON. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THE TARGET TISSUE WAS NOT THICK. THE FORCEPS WAS NOT USED, SO THERE IS NO POSSIBILITY THAT THE DEVICE CLAMPED THE HARD THINGS. AFTER THE OPERATION, THE PULSE FIRING WAS PERFORMED, AND IT CONFIRMED THAT THE KNIFE COULD NOT BE RETURNED AFTER THE KNIFE STOPPED. KNIFE REVERSE SWITCH COULD NOT BE MOVED. MANUAL OVERRIDE LEVER WAS USED, BUT LEVER STOPPED AT 1.5 ROUND TRIP. AT THAT TIME, KNIFE REVERSE SWITCH COULD BE USED, AND THE KNIFE RETURNED.
05-01-2023 | Malfunction | ethicon endo-surgery | 25-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE KNIFE MOVED TO THE TIP, BUT DID NOT RETURN TO THE HOME POSITION AT 1ST FIRING WHEN CUTTING THE SIGMOID. KNIFE REVERSE SWITCH DID NOT WORK, AND MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE STAPLE WAS STAPLED PROPERLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE DEVICE WAS USED ON THE COLON. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THE TARGET TISSUE WAS NOT THICK. THE FORCEPS WAS NOT USED, SO THERE IS NO POSSIBILITY THAT THE DEVICE CLAMPED THE HARD THINGS. AFTER THE OPERATION, THE PULSE FIRING WAS PERFORMED, AND IT CONFIRMED THAT THE KNIFE COULD NOT BE RETURNED AFTER THE KNIFE STOPPED. KNIFE REVERSE SWITCH COULD NOT BE MOVED. MANUAL OVERRIDE LEVER WAS USED, BUT LEVER STOPPED AT 1.5 ROUND TRIP. AT THAT TIME, KNIFE REVERSE SWITCH COULD BE USED, AND THE KNIFE RETURNED.
05-01-2023 | Malfunction | ethicon endo-surgery | 25-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE KNIFE MOVED TO THE TIP, BUT DID NOT RETURN TO THE HOME POSITION AT 1ST FIRING WHEN CUTTING THE SIGMOID. KNIFE REVERSE SWITCH DID NOT WORK, AND MANUAL OVERRIDE LEVER WAS USED TO RETURN THE KNIFE. THE STAPLE WAS STAPLED PROPERLY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE DEVICE WAS USED ON THE COLON. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. THE TARGET TISSUE WAS NOT THICK. THE FORCEPS WAS NOT USED, SO THERE IS NO POSSIBILITY THAT THE DEVICE CLAMPED THE HARD THINGS. AFTER THE OPERATION, THE PULSE FIRING WAS PERFORMED, AND IT CONFIRMED THAT THE KNIFE COULD NOT BE RETURNED AFTER THE KNIFE STOPPED. KNIFE REVERSE SWITCH COULD NOT BE MOVED. MANUAL OVERRIDE LEVER WAS USED, BUT LEVER STOPPED AT 1.5 ROUND TRIP. AT THAT TIME, KNIFE REVERSE SWITCH COULD BE USED, AND THE KNIFE RETURNED.
29-12-2022 | Malfunction | covidien lp llc north haven | 25-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN APPENDECTOMY, WHILE AT THE APPENDIX, THE DEVICE DID NOT FIRE. A SECOND RELOAD WAS OPENED AND USED BUT SAME ISSUE OCCURRED.
30-06-2020 | Injury | covidien lp llc north haven | 25-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE REPORT, A STUDY COMPARED LAPAROSCOPIC SPLENECTOMY COMBINED WITH OESOPHAGOGASTRIC DEVASCULARISATION VERSUS OPEN SPLENECTOMY COMBINED WITH OESOPHAGOGASTRIC DEVASCULARISATION IN PATIENTS WITH PORTAL HYPERTENSION SECONDARY TO LIVER CIRRHOSIS. THIS STUDY INCLUDED 192 PATIENTS DIAGNOSED WITH PORTAL HYPERTENSION AND SEVERE GASTROESOPHAGEAL VARICES BETWEEN JANUARY 2002 AND DECEMBER 2018; 62 PATIENTS UNDERWENT LAPAROSCOPIC SPLENECTOMY COMBINED WITH OESOPHAGOGASTRIC DEVASCULARISATION AND 130 PATIENTS UNDERWENT THE OPEN PROCEDURE (OPEN GROUP). THE COMPLICATIONS INCLUDE: IN PATIENTS WITH SPLENIC HILAR BLEEDING, HEMORRHAGE USUALLY OCCURRED FROM THE TIP OF THE STAPLE LINE WHEN THE FIRST STAPLER CANNOT CROSS THE ENTIRE SPLENIC HILAR PEDICLE, AND THE HIGH SPLENIC BLOOD FLOW MAY ACCUMULATE IN THE PARTIALLY RESECTED SPLENIC VEIN. USE OF ENDOSCOPIC VASCULAR STAPLERS MORE THAN TWICE FOR SPLENIC HILAR DIVISION RESULTS IN OVERLAPPING OF STAPLE LINES. OOZING FROM THE RESECTED STUMP OF THE SPLENIC HILAR PEDICLES WAS TREATED WITH SUTURING OR CLIPPING THE STAPLE LINES. BLOOD TRANSFUSIONS IN 12 PATIENTS IN LAPAROSCOPIC GROUP, WHILE 3 PATIENTS UNDERWENT CONVERSION FR OM LAPAROSCOPIC SURGERY TO OPEN SURGERY, NECESSITATED BY UNCONTROLLED INTRAOPERATIVE BLEEDING. LAPAROSCOPIC GENERAL COMPLICATIONS INCLUDED PULMONARY COMPLICATIONS AND ILEUS.
30-06-2020 | Injury | covidien lp llc north haven | 25-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Obstruction/Occlusion | Event Description: ACCORDING TO THE LITERATURE REPORT, A STUDY COMPARED LAPAROSCOPIC SPLENECTOMY COMBINED WITH OESOPHAGOGASTRIC DEVASCULARISATION VERSUS OPEN SPLENECTOMY COMBINED WITH OESOPHAGOGASTRIC DEVASCULARISATION IN PATIENTS WITH PORTAL HYPERTENSION SECONDARY TO LIVER CIRRHOSIS. THIS STUDY INCLUDED 192 PATIENTS DIAGNOSED WITH PORTAL HYPERTENSION AND SEVERE GASTROESOPHAGEAL VARICES BETWEEN JANUARY 2002 AND DECEMBER 2018; 62 PATIENTS UNDERWENT LAPAROSCOPIC SPLENECTOMY COMBINED WITH OESOPHAGOGASTRIC DEVASCULARISATION AND 130 PATIENTS UNDERWENT THE OPEN PROCEDURE (OPEN GROUP). THE COMPLICATIONS INCLUDE: IN PATIENTS WITH SPLENIC HILAR BLEEDING, HEMORRHAGE USUALLY OCCURRED FROM THE TIP OF THE STAPLE LINE WHEN THE FIRST STAPLER CANNOT CROSS THE ENTIRE SPLENIC HILAR PEDICLE, AND THE HIGH SPLENIC BLOOD FLOW MAY ACCUMULATE IN THE PARTIALLY RESECTED SPLENIC VEIN. USE OF ENDOSCOPIC VASCULAR STAPLERS MORE THAN TWICE FOR SPLENIC HILAR DIVISION RESULTS IN OVERLAPPING OF STAPLE LINES. OOZING FROM THE RESECTED STUMP OF THE SPLENIC HILAR PEDICLES WAS TREATED WITH SUTURING OR CLIPPING THE STAPLE LINES. BLOOD TRANSFUSIONS IN 12 PATIENTS IN LAPAROSCOPIC GROUP, WHILE 3 PATIENTS UNDERWENT CONVERSION FR OM LAPAROSCOPIC SURGERY TO OPEN SURGERY, NECESSITATED BY UNCONTROLLED INTRAOPERATIVE BLEEDING. LAPAROSCOPIC GENERAL COMPLICATIONS INCLUDED PULMONARY COMPLICATIONS AND ILEUS.
25-01-2023 | Malfunction | covidien lp llc north haven | 25-01-2023 | UNKNOWN ENDO GIA SULU | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE WAS BROKEN. THE ARTICULATION PART OF JOINT TO THE RELOAD WAS BROKEN.
29-12-2022 | Malfunction | covidien lp llc north haven | 25-01-2023 | UNKNOWN STAPLING DEVICE | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN APPENDECTOMY, WHILE AT THE APPENDIX, THE DEVICE DID NOT FIRE. A SECOND RELOAD WAS OPENED AND USED BUT SAME ISSUE OCCURRED.
19-12-2022 | Malfunction | us surgical puerto rico | 25-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, PRIOR TO PATIENT USE, THE JAWS WOULD NOT CLOSE. A NEW RELOAD WAS USED BUT THE JAWS STILL WOULD NOT CLOSE. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED WITH THE TWO RELOADS. THE ISSUE CAUSED DELAY IN THE PROCEDURE. THERE WAS NO PATIENT INJURY.
29-12-2022 | Malfunction | covidien lp llc north haven | 25-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN APPENDECTOMY, WHILE AT THE APPENDIX, THE DEVICE DID NOT FIRE. A SECOND RELOAD WAS OPENED AND USED BUT SAME ISSUE OCCURRED.
01-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON HAD A HARD TIME INSERTING THE CIRCULAR STAPLER TRANSANALLY DUE TO THE PATIENT'S ANATOMY. THE SURGEON DEFINED ANATOMY TO HAVE "KINKS." THE SURGEON HAD A HARD TIME CONNECTING THE ANVIL TO THE TROCAR DUE TO THE TISSUE THICKNESS OF THE RECTAL STUMP. THE SURGEON FELT HIGH TENSION WHILE RETRACTING TROCAR/ANVIL CONFIGURATION. FIRING EXECUTED, THE WASHER BROKE, AND A GREEN CHECK MARK APPEARED. THE SURGEON TURNED COUNTERCLOCKWISE TWO FULL REVOLUTIONS - HAD A HARD TIME RELEASING FROM THE ANASTOMOSIS. THE SURGEON ROTATED THE DEVICE ANOTHER REVOLUTION TO RELEASE. THE SURGEON PULLED THE DEVICE, AND THE ANVIL WAS LEFT IN THE RECTUM. THE SURGEON PULLED THE ANVIL FROM THE RECTUM AND THERE WAS A LEAK AS A RESULT. DONUTS LOOKED VERY LONG (LONGITUDINALLY). THERE WAS A LEAK.
01-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON HAD A HARD TIME INSERTING THE CIRCULAR STAPLER TRANSANALLY DUE TO THE PATIENT'S ANATOMY. THE SURGEON DEFINED ANATOMY TO HAVE "KINKS." THE SURGEON HAD A HARD TIME CONNECTING THE ANVIL TO THE TROCAR DUE TO THE TISSUE THICKNESS OF THE RECTAL STUMP. THE SURGEON FELT HIGH TENSION WHILE RETRACTING TROCAR/ANVIL CONFIGURATION. FIRING EXECUTED, THE WASHER BROKE, AND A GREEN CHECK MARK APPEARED. THE SURGEON TURNED COUNTERCLOCKWISE TWO FULL REVOLUTIONS - HAD A HARD TIME RELEASING FROM THE ANASTOMOSIS. THE SURGEON ROTATED THE DEVICE ANOTHER REVOLUTION TO RELEASE. THE SURGEON PULLED THE DEVICE, AND THE ANVIL WAS LEFT IN THE RECTUM. THE SURGEON PULLED THE ANVIL FROM THE RECTUM AND THERE WAS A LEAK AS A RESULT. DONUTS LOOKED VERY LONG (LONGITUDINALLY). THERE WAS A LEAK.
01-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON HAD A HARD TIME INSERTING THE CIRCULAR STAPLER TRANSANALLY DUE TO THE PATIENT'S ANATOMY. THE SURGEON DEFINED ANATOMY TO HAVE "KINKS." THE SURGEON HAD A HARD TIME CONNECTING THE ANVIL TO THE TROCAR DUE TO THE TISSUE THICKNESS OF THE RECTAL STUMP. THE SURGEON FELT HIGH TENSION WHILE RETRACTING TROCAR/ANVIL CONFIGURATION. FIRING EXECUTED, THE WASHER BROKE, AND A GREEN CHECK MARK APPEARED. THE SURGEON TURNED COUNTERCLOCKWISE TWO FULL REVOLUTIONS - HAD A HARD TIME RELEASING FROM THE ANASTOMOSIS. THE SURGEON ROTATED THE DEVICE ANOTHER REVOLUTION TO RELEASE. THE SURGEON PULLED THE DEVICE, AND THE ANVIL WAS LEFT IN THE RECTUM. THE SURGEON PULLED THE ANVIL FROM THE RECTUM AND THERE WAS A LEAK AS A RESULT. DONUTS LOOKED VERY LONG (LONGITUDINALLY). THERE WAS A LEAK.
01-01-2022 | Malfunction | ethicon endo-surgery | 25-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON HAD A HARD TIME INSERTING THE CIRCULAR STAPLER TRANSANALLY DUE TO THE PATIENT'S ANATOMY. THE SURGEON DEFINED ANATOMY TO HAVE "KINKS." THE SURGEON HAD A HARD TIME CONNECTING THE ANVIL TO THE TROCAR DUE TO THE TISSUE THICKNESS OF THE RECTAL STUMP. THE SURGEON FELT HIGH TENSION WHILE RETRACTING TROCAR/ANVIL CONFIGURATION. FIRING EXECUTED, THE WASHER BROKE, AND A GREEN CHECK MARK APPEARED. THE SURGEON TURNED COUNTERCLOCKWISE TWO FULL REVOLUTIONS - HAD A HARD TIME RELEASING FROM THE ANASTOMOSIS. THE SURGEON ROTATED THE DEVICE ANOTHER REVOLUTION TO RELEASE. THE SURGEON PULLED THE DEVICE, AND THE ANVIL WAS LEFT IN THE RECTUM. THE SURGEON PULLED THE ANVIL FROM THE RECTUM AND THERE WAS A LEAK AS A RESULT. DONUTS LOOKED VERY LONG (LONGITUDINALLY). THERE WAS A LEAK.
01-01-2023 | Malfunction | ethicon endo-surgery | 25-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER DID JUST CLIP AND CUT ONE HALF, ALSO THE CLIPS WHERE MISSHAPED. NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 25-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER DID JUST CLIP AND CUT ONE HALF, ALSO THE CLIPS WHERE MISSHAPED. NO PATIENT HARM.
14-11-2022 | Malfunction | intuitive surgical | 25-01-2023 | SUREFORM | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
26-10-2022 | Malfunction | covidien lp llc north haven | 25-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LIGATION AND CUTTING OF LUNG TISSUE, IN THE MIDDLE OF FIRING, THE HANDLE FLICKERED AND DID NOT FIRE. THE TISSUE WAS NOT CUT. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
09-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, IN DISSOCIATING THE LUNG, THE DEVICE FULLY FIRED BUT THE STAPLE LINE WAS DISTALLY INCOMPLETE. THE DEPLOYED STAPLES HAD POOR FORMATION AS WELL. TO RESOLVE THE ISSUE, THE SURGEON HAD TO HAND SEW.
09-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMIGUT RESECTION, IN THE ANASTOMOSIS, THERE WERE UNPERFECT B-SHAPED STAPLES. THE DEVICE FIRED ALL THE WAY BUT NO STAPLES DEPLOYED IN THE DISTAL PART OF THE STAPLE LINE. RE-ANASTOMOSIS WERE DONE USING A NEW RELOAD TO RESOLVE THE ISSUE.
09-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMIGUT RESECTION, IN THE ANASTOMOSIS, THERE WERE UNPERFECT B-SHAPED STAPLES. THE DEVICE FIRED ALL THE WAY BUT NO STAPLES DEPLOYED IN THE DISTAL PART OF THE STAPLE LINE. RE-ANASTOMOSIS WERE DONE USING A NEW RELOAD TO RESOLVE THE ISSUE.
11-11-2022 | Malfunction | ethicon endo-surgery | 25-01-2023 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE POWERED ECHELON GUN GOT STUCK DURING FIRING IN THE SURGERY. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
26-10-2022 | Malfunction | covidien lp llc north haven | 25-01-2023 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LIGATION AND CUTTING OF LUNG TISSUE, IN THE MIDDLE OF FIRING, THE HANDLE FLICKERED AND DID NOT FIRE. THE TISSUE WAS NOT CUT. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
26-10-2022 | Malfunction | covidien lp llc north haven | 25-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LIGATION AND CUTTING OF LUNG TISSUE, IN THE MIDDLE OF FIRING, THE HANDLE FLICKERED AND DID NOT FIRE. THE TISSUE WAS NOT CUT. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
25-10-2022 | Malfunction | covidien lp llc north haven | 25-01-2023 | ENDO GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER STAPLING, IT WAS OBSERVE THAT THERE WAS NO INCISION MADE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Malfunction | ethicon endo-surgery | 25-01-2023 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE ECHELON FLEX¿ POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP COULD ONLY BE RELEASED ONCE, ON THE SECOND MAGAZINE THE MAGAZINE DID NOT RELEASE. THE STAPLER ALSO FAILED TO CLOSE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
10-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBECTOMY, WHEN CUTTING THE LUNG, THERE WAS A POOR STAPLE FORMATION. THE DISTAL END OF THE STAPLE FORMED INTO THE SHAPE OF "DOOR". THERE WAS A STUCK DURING THE FIRING PROCESS OF THE STAPLER. AFTER CONTINUING TO FIRE, IT WAS FULLY FIRED. THE STAPLE LINE WAS PARTIALLY COMPLETE. A NEW RELOAD WAS USED AND MANUAL STITCHING WAS DONE TO COMPLETE THE CASE.
10-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | ENDO GIA | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBECTOMY, WHEN CUTTING THE LUNG, THERE WAS A POOR STAPLE FORMATION. THE DISTAL END OF THE STAPLE FORMED INTO THE SHAPE OF "DOOR". THERE WAS A STUCK DURING THE FIRING PROCESS OF THE STAPLER. AFTER CONTINUING TO FIRE, IT WAS FULLY FIRED. THE STAPLE LINE WAS PARTIALLY COMPLETE. A NEW RELOAD WAS USED AND MANUAL STITCHING WAS DONE TO COMPLETE THE CASE.
09-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AFTER DISSECTING THE MESORECTUM, THE ANVIL WAS POSITIONED AND JOINED TO THE STAPLER FOR FIRING. WHEN THE HANDLE OF THE STAPLER WAS FULLY SQUEEZED, SOME STAPLES DEPLOYED WITH POOR FORMATION AND SOME DID NOT DEPLOY. THE STAPLER WAS TAKEN OUT AND THE SURGEON PERFORMED OVER SEWING TO FIX THE ANASTOMOSIS.
09-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AFTER DISSECTING THE MESORECTUM, THE ANVIL WAS POSITIONED AND JOINED TO THE STAPLER FOR FIRING. WHEN THE HANDLE OF THE STAPLER WAS FULLY SQUEEZED, SOME STAPLES DEPLOYED WITH POOR FORMATION AND SOME DID NOT DEPLOY. THE STAPLER WAS TAKEN OUT AND THE SURGEON PERFORMED OVER SEWING TO FIX THE ANASTOMOSIS.
09-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | EEA | Failure to Fire | A050501 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AFTER DISSECTING THE MESORECTUM, THE ANVIL WAS POSITIONED AND JOINED TO THE STAPLER FOR FIRING. WHEN THE HANDLE OF THE STAPLER WAS FULLY SQUEEZED, SOME STAPLES DEPLOYED WITH POOR FORMATION AND SOME DID NOT DEPLOY. THE STAPLER WAS TAKEN OUT AND THE SURGEON PERFORMED OVER SEWING TO FIX THE ANASTOMOSIS.
09-01-2023 | Injury | us surgical puerto rico | 25-01-2023 | EEA | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION, AFTER DISSECTING THE MESORECTUM, THE ANVIL WAS POSITIONED AND JOINED TO THE STAPLER FOR FIRING. WHEN THE HANDLE OF THE STAPLER WAS FULLY SQUEEZED, SOME STAPLES DEPLOYED WITH POOR FORMATION AND SOME DID NOT DEPLOY. THE STAPLER WAS TAKEN OUT AND THE SURGEON PERFORMED OVER SEWING TO FIX THE ANASTOMOSIS.
19-12-2022 | Malfunction | intuitive surgical | 25-01-2023 | SUREFORM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: INTUITIVE SURGICAL INC (ISI) RECEIVED USER FACILITY REPORT 4600060000-2022-8006 STATING: "DOCTOR WAS USING THE SUREFORM STAPLER 60 ON THE STOMACH RESECTION. DURING THREE SEPARATE TIMES, THE STAPLER FAILED TO REACH THE COMPLETE DISTANCE THAT WAS REQUIRED TO COMPLETE THE STAPLER LINE. NEW STAPLE LOADS WERE OPENED EACH TIME, AND EVENTUALLY A NEW STAPLER." ISI FOLLOWED UP WITH THE SITE'S CLINICAL SALES REP AND THE FOLLOWING ADDITIONAL INFORMATION WAS OBTAINED: THE SUREFORM 60 BLUE RELOADS WERE THE RELOADS INVOLVED WITH THE REPORTED EVENT. THE RELOAD WAS SELECTED BECAUSE IT IS THE RELOAD THE SURGEON TYPICALLY USES FOR 1 RELOAD ON SLEEVES. STAPLER FIRES 1-3 WERE THE STAPLER FIRES INVOLVES WITH THE REPORTED EVENT. A STOMACH TRANSACTION WAS THE SURGICAL TASK BEING PERFORMED AT THE TIME OF THE EVENT. THE STOMACH WAS THE TARGET TISSUE. THE TARGET TISSUE WAS NOT CALCIFIED AND WAS NOT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY, THERE WAS NO TISSUE TENSION OR BUNCHING. THE EVENT DID NOT INVOLVE A FAILURE TO FIRE. THE STAPLER SEEMED TO STOP IN THE SAME AREA X3 SEPARATE TIMES, WITH THE TISSUE BEING TOO THICK TO CONTINUE. THERE WAS NO TISSUE BEING PUSHED FORWARD/DRAGGED DURING THE FIRING PROCESS. THE JAWS DID NOT OPEN/RELEASE DURING THE FIRING SEQUENCE. THE RELOAD STAPLES DID NOT DEPLOY UNEXPECTEDLY. THERE WERE NO MALFORMED STAPLES. THE RELOAD DID HAVE AN EXPOSED BLADE. THERE WERE NO STAPLES MISSING FROM THE STAPLE LINE. THERE WERE NO HOLES/GAPS IN THE STAPLE LINE. A "TISSUE TOO THICK" ERROR MESSAGE OCCURRED. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS BETWEEN THE INSTRUMENT JAWS. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. NO BUTTRESS MATERIAL WAS USED. NO BLEEDING ON THE STAPLE LINE WAS OBSERVED. THERE WAS NO UNPLANNED TISSUE REMOVAL DUE TO THE FIRING FAILURE. THE SURGEON RESOLVED THE ISSUE BY OPENING A NEW STAPLER. THE PROCEDURE WAS COMPLETED ROBOTICALLY AS PLANNED WITH NO REPORTS OF PATIENT INJURY. THE RELOAD IS NOT AVAILABLE FOR RETURN.
10-01-2023 | Injury | covidien lp llc north haven | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE-HOLE THORACOSCOPIC LOBECTOMY, WHILE THE LOBE WAS DISSOCIATED, THE DEVICE FIRED, STAPLES WERE NOT B-SHAPED AND STAPLE LINE WAS INCOMPLETE CENTRALLY. THE STAPLE LINE WAS FIXED BY STITCHING. THE RELOAD WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
10-01-2023 | Malfunction | us surgical puerto rico | 25-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION IN THE UPPER LOBE OF THE RIGHT LUNG, WHILE WEDGE-CUTTING OF LOBES, THE DEVICE DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZED. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2023 | Injury | celestica de monterrey s.a de c.v | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE, WHILE RESECTING THE SIGMOID COLON, TWO RELOADS FIRED. HOWEVER, THERE WAS POOR STAPLE FORMATION, AFTER FIRST SQUEEZING THE STAPLER, INITIALS STAPLES MOVE A LITTLE AND NO MOVEMENT AFTER THAT. THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. THE RELOADS WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
10-01-2023 | Injury | celestica de monterrey s.a de c.v | 25-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE, WHILE RESECTING THE SIGMOID COLON, TWO RELOADS FIRED. HOWEVER, THERE WAS POOR STAPLE FORMATION, AFTER FIRST SQUEEZING THE STAPLER, INITIAL STAPLES MOVE A LITTLE AND NO MOVEMENT AFTER THAT. THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. THE RELOADS WERE REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
07-11-2022 | Malfunction | ethicon endo-surgery | 25-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IN THE CLINICIANS WORDS IT WAS REPORTED THE DEVICE WAS STUCK WHILE BEING USED INSIDE THE PATIENT'S ABDOMEN. THE DR. HAD TO MANUALLY PUT THE DEVICE'S HEAD PART TO ITS NORMAL POSITION FOR IT TO BE REMOVED THROUGH THE PORT¿ THEY USED ANOTHER AND IT DID THE SAME. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
07-11-2022 | Malfunction | ethicon endo-surgery | 25-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE GOT STUCK WHILE BEING USED INSIDE THE PATIENT'S ABDOMEN. THE DR. HAD TO MANUALLY PUT THE DEVICE'S HEAD PART TO ITS NORMAL POSITION FOR IT TO BE REMOVED THROUGH THE PORT¿ THEY USED ANOTHER AND IT DID THE SAME. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2022 | Injury | intuitive surgical | 25-01-2023 | SUREFORM | Material Fragmentation | A040103 | Foreign Body In Patient | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS PROCEDURE, STAPLES DID NOT FORM CORRECTLY ON TISSUE WHEN A BLUE SUREFORM 45 STAPLER RELOAD WAS FIRED WITH A SUREFORM 45 STAPLER INSTRUMENT. AS A RESULT, THE TISSUE WAS PUSHED OUT AND TORN, AND ADDITIONAL UNPLANNED STAPLING WAS REQUIRED. THE SURGEON WAS ABLE TO RETRIEVE THE UNFORMED STAPLES. INTUITIVE SURGICAL, INC. (ISI) OBTAINED THE FOLLOWING ADDITIONAL INFORMATION FROM THE SURGEON REGARDING THE REPORTED EVENT: PER THE SURGEON, THE STAPLE WAS INSPECTED PRIOR TO THE PROCEDURE, AND NO ISSUES WERE FOUND. IT WAS STATED THAT THE 1ST FIRE WAS NORMAL. DURING THE 2ND FIRE, "A VERY SMALL PIECE OF THE RELOAD (BLUE) BROKE OFF AND FELL INTO THE PATIENT," BUT THE FRAGMENT WAS RETRIEVED DURING THE PROCEDURE. IT WAS CONFIRMED THAT THE STAPLE LINE WAS COMPLETED. THE SURGEON CONFIRMED THAT THE EVENT WHEN THE TISSUE WAS PUSHED OUT AND TORN OCCURRED DURING THE 3RD FIRE WHILE USING A BLUE SUREFORM 45 RELOAD, WHICH IS HIS STANDARD CHOICE FOR THAT STEP OF THE PROCEDURE. AT THE TIME OF THE EVENT, THE SURGEON WAS CREATING THE GASTRIC POUCH AND STAPLING GASTRIC TISSUE. THE TISSUE WAS NOT CALCIFIED OR EXPOSED TO RADIATION OR CHEMOTHERAPY, AND THERE WAS NO TISSUE TENSION. THERE WAS ALSO NO TISSUE BUNCHING "UNTIL THE BLADE PUSHED THE TISSUE FORWARD DURING FIRING AND AT THAT TIME, THE TISSUE BECAME BUNCHED." PER THE SURGEON, THIS ISSUE RESULTED IN HOLES IN THE STAPLE LINE AND A LEAK FROM THE GASTRIC REMNANT, WHICH WAS ADDRESSED VIA ADDITIONAL RESECTION OF GASTRIC TISSUE AND RE-STAPLING THE AFFECTED AREA UTILIZING A NEW STAPLER. THE EVENT DID NOT INVOLVE A FAILURE TO FIRE NOR A PARTIAL FIRE, AS THE STAPLER "TRIED TO FIRE ALL THE WAY." THE STAPLER JAWS DID NOT OPEN/RELEASE DURING THE FIRING SEQUENCE, AND NO STAPLES WERE DEPLOYED FROM THE RELOAD UNEXPECTEDLY. THERE WERE NO ERRORS OR MESSAGES GENERATED WHEN THE ISSUE OCCURRED. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS, NOR DID HE FIRE OVER AN EXISTING STAPLE LINE. NO BUTTRESS MATERIAL WAS USED. THE SURGEON DID NOT EXPERIENCE ANY CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD.
12-01-2023 | Malfunction | intuitive surgical | 26-01-2023 | SUREFORM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: INSTRUMENT ENGAGEMENT FAILED. THE SUREFORM 45 STAPLER WAS USED, THE LOAD FIRED CORRECTLY THE FIRST TIME, THE SECOND LOAD DID NOT FIRE PROPERLY, AND THE SUREFORM 45 INSTRUMENT COULD NO LONGER BE USED. HEALTHCARE PROVIDERS WERE NOTIFIED OF IFU UPDATE.
12-01-2023 | Malfunction | intuitive surgical | 26-01-2023 | SUREFORM | Intermittent Energy Output | A090403 | No Clinical Signs, Symptoms or Conditions | Event Description: INSTRUMENT ENGAGEMENT FAILED. THE SUREFORM 45 STAPLER WAS USED, THE LOAD FIRED CORRECTLY THE FIRST TIME, THE SECOND LOAD DID NOT FIRE PROPERLY, AND THE SUREFORM 45 INSTRUMENT COULD NO LONGER BE USED. HEALTHCARE PROVIDERS WERE NOTIFIED OF IFU UPDATE.
06-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | ECHELON FLEX 60 LONG | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE AND DURING A SIGMOID COLON RESECTION, BEFORE THE PROCEDURE THEY CLOSED THE STAPLER BEFORE PUTTING THE RELOAD IN. THEY WERE UNABLE TO OPEN THE STAPLER ONCE CLOSED. IT IS UNKNOWN IF THE REVERSE WAS USED OR IF THE MANUAL OVERRIDE WAS USED. THE STAPLER WOULD NOT OPEN. ANOTHER STAPLER WAS OPENED FOR THE PROCEDURE. THEY LOADED THE STAPLER. UPON BEING READY TO FIRE THE STAPLER THE SURGEON WANTED TO RE-APPROXIMATE AND WAS UNABLE TO OPEN THE CLOSED STAPLER. THE REVERSE WAS ATTEMPTED AND THEN THE MANUAL OVERRIDE WAS ATTEMPTED, BOTH DID NOT WORK. THE STAPLER WAS REMOVED FROM TISSUE AFTER TWENTY MINUTES. THE TRANSECTION WAS COMPLETED FURTHER DOWN AND WITH A THIRD STAPLER. NO PATIENT HARM WAS REPORTED.
06-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | ECHELON FLEX 60 LONG | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE AND DURING A SIGMOID COLON RESECTION, BEFORE THE PROCEDURE THEY CLOSED THE STAPLER BEFORE PUTTING THE RELOAD IN. THEY WERE UNABLE TO OPEN THE STAPLER ONCE CLOSED. IT IS UNKNOWN IF THE REVERSE WAS USED OR IF THE MANUAL OVERRIDE WAS USED. THE STAPLER WOULD NOT OPEN. ANOTHER STAPLER WAS OPENED FOR THE PROCEDURE. THEY LOADED THE STAPLER. UPON BEING READY TO FIRE THE STAPLER THE SURGEON WANTED TO RE-APPROXIMATE AND WAS UNABLE TO OPEN THE CLOSED STAPLER. THE REVERSE WAS ATTEMPTED AND THEN THE MANUAL OVERRIDE WAS ATTEMPTED, BOTH DID NOT WORK. THE STAPLER WAS REMOVED FROM TISSUE AFTER TWENTY MINUTES. THE TRANSECTION WAS COMPLETED FURTHER DOWN AND WITH A THIRD STAPLER. NO PATIENT HARM WAS REPORTED.
10-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, IN THE GASTRIC, THE SURGEON COULD NOT PRESS THE GREEN BUTTON. THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, THE ARTICULATING KNOB BROKE OFF. TO RESOLVE THE ISSUE, SUTURES WERE USED. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, IN THE GASTRIC, THE SURGEON COULD NOT PRESS THE GREEN BUTTON. THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, THE ARTICULATING KNOB BROKE OFF. TO RESOLVE THE ISSUE, SUTURES WERE USED. THERE WAS NO PATIENT INJURY.
11-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | EEA | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHILE ON ANASTOMOSIS OF RECTUM STUMP TO THE COLON, THE DEVICE PARTIALLY FIRED, HAD INCOMPLETE STAPLE LINE, AND THE ANVIL COULD NOT BE TILTED AFTER FIRING. MANUAL OVER SEWING WAS DONE TO TIE THE ANASTOMOTIC SITE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE TO ALLOW MORE TIME TO SUTURE.
11-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | EEA | Difficult to Remove | A150207 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHILE ON ANASTOMOSIS OF RECTUM STUMP TO THE COLON, THE DEVICE PARTIALLY FIRED, HAD INCOMPLETE STAPLE LINE, AND THE ANVIL COULD NOT BE TILTED AFTER FIRING. MANUAL OVER SEWING WAS DONE TO TIE THE ANASTOMOTIC SITE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE TO ALLOW MORE TIME TO SUTURE.
11-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHILE ON ANASTOMOSIS OF RECTUM STUMP TO THE COLON, THE DEVICE PARTIALLY FIRED, HAD INCOMPLETE STAPLE LINE, AND THE ANVIL COULD NOT BE TILTED AFTER FIRING. MANUAL OVER SEWING WAS DONE TO TIE THE ANASTOMOTIC SITE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE TO ALLOW MORE TIME TO SUTURE.
11-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHILE ON ANASTOMOSIS OF RECTUM STUMP TO THE COLON, THE DEVICE PARTIALLY FIRED, HAD INCOMPLETE STAPLE LINE, AND THE ANVIL COULD NOT BE TILTED AFTER FIRING. MANUAL OVER SEWING WAS DONE TO TIE THE ANASTOMOTIC SITE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE TO ALLOW MORE TIME TO SUTURE.
11-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHILE ON ANASTOMOSIS OF RECTUM STUMP TO THE COLON, THE DEVICE PARTIALLY FIRED, HAD INCOMPLETE STAPLE LINE, AND THE ANVIL COULD NOT BE TILTED AFTER FIRING. MANUAL OVER SEWING WAS DONE TO TIE THE ANASTOMOTIC SITE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE TO ALLOW MORE TIME TO SUTURE.
11-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOW ANTERIOR RESECTION, WHILE ON ANASTOMOSIS OF RECTUM STUMP TO THE COLON, THE DEVICE PARTIALLY FIRED, HAD INCOMPLETE STAPLE LINE, AND THE ANVIL COULD NOT BE TILTED AFTER FIRING. MANUAL OVER SEWING WAS DONE TO TIE THE ANASTOMOTIC SITE. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE TO ALLOW MORE TIME TO SUTURE.
30-12-2022 | Malfunction | us surgical puerto rico | 26-01-2023 | EEA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL CANCER PROCEDURE, THE CIRCULAR STAPLER FIRED BUT IRREGULAR AND UNSAFE STAPLING. IT WAS ALSO REPORTED THAT THE ANVIL BENT. A NEW RADIAL RELOAD STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2022 | Malfunction | us surgical puerto rico | 26-01-2023 | EEA | Material Twisted/Bent | A040609 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL CANCER PROCEDURE, THE CIRCULAR STAPLER FIRED BUT IRREGULAR AND UNSAFE STAPLING. IT WAS ALSO REPORTED THAT THE ANVIL BENT. A NEW RADIAL RELOAD STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE USING THE HANDLE, THE FIRST RELOAD WAS LOADED WITH A PIN WHEN DISENGAGING THE PATIENT¿S BRONCHUS, AND DURING FIRST FIRING, THE STAPLE WAS STUCK AND THE DISTAL END OF THE SUTURE WAS INCOMPLETE. AFTER REPLACING THE RELOAD, IT WAS FOUND THAT THE GRIP WAS SQUEEZED, BUT DEVICE COULD NOT BE FIRED. THE GREEN BUTTON WAS PRESSED. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
04-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, WHILE CLOSING THE ENTRY HOLE FOR FEEA (FUNCTIONAL END-TO-END ANAST OMOSIS) RECONSTRUCTION OF THE COLON OUTSIDE THE BODY, THE DEVICE WAS DIFFICULT TO FIRE. IT INITIALLY STOPPED FIRING MIDWAY, BUT THE SURGEON REAPPLIED FORCE AND WAS ABLE TO COMPLETE FIRING THE DEVICE UNTIL THE END. LESS THAN 500CC BLEEDING FROM THE PROXIMAL SIDE OF THE STAPLE LINE OCCURRED; AND SINCE THE DEGREE OF STAPLE FORMATION COULD NOT BE CONFIRMED, SUTURE REINFORCEMENT WAS PERFORMED AND FURTHER IMPLANTATION IN THE SEROUS MUSCLE LAYER (SEROMUSCULAR). NO POSTOPERATIVE LEAK WAS NOTED, BUT ILEUS WAS OCCURRING.
04-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Obstruction/Occlusion | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, WHILE CLOSING THE ENTRY HOLE FOR FEEA (FUNCTIONAL END-TO-END ANAST OMOSIS) RECONSTRUCTION OF THE COLON OUTSIDE THE BODY, THE DEVICE WAS DIFFICULT TO FIRE. IT INITIALLY STOPPED FIRING MIDWAY, BUT THE SURGEON REAPPLIED FORCE AND WAS ABLE TO COMPLETE FIRING THE DEVICE UNTIL THE END. LESS THAN 500CC BLEEDING FROM THE PROXIMAL SIDE OF THE STAPLE LINE OCCURRED; AND SINCE THE DEGREE OF STAPLE FORMATION COULD NOT BE CONFIRMED, SUTURE REINFORCEMENT WAS PERFORMED AND FURTHER IMPLANTATION IN THE SEROUS MUSCLE LAYER (SEROMUSCULAR). NO POSTOPERATIVE LEAK WAS NOTED, BUT ILEUS WAS OCCURRING.
04-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, WHILE CLOSING THE ENTRY HOLE FOR FEEA (FUNCTIONAL END-TO-END ANAST OMOSIS) RECONSTRUCTION OF THE COLON OUTSIDE THE BODY, THE DEVICE WAS DIFFICULT TO FIRE. IT INITIALLY STOPPED FIRING MIDWAY, BUT THE SURGEON REAPPLIED FORCE AND WAS ABLE TO COMPLETE FIRING THE DEVICE UNTIL THE END. LESS THAN 500CC BLEEDING FROM THE PROXIMAL SIDE OF THE STAPLE LINE OCCURRED; AND SINCE THE DEGREE OF STAPLE FORMATION COULD NOT BE CONFIRMED, SUTURE REINFORCEMENT WAS PERFORMED AND FURTHER IMPLANTATION IN THE SEROUS MUSCLE LAYER (SEROMUSCULAR). NO POSTOPERATIVE LEAK WAS NOTED, BUT ILEUS WAS OCCURRING.
04-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | GIA | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, WHILE CLOSING THE ENTRY HOLE FOR FEEA (FUNCTIONAL END-TO-END ANAST OMOSIS) RECONSTRUCTION OF THE COLON OUTSIDE THE BODY, THE DEVICE WAS DIFFICULT TO FIRE. IT INITIALLY STOPPED FIRING MIDWAY, BUT THE SURGEON REAPPLIED FORCE AND WAS ABLE TO COMPLETE FIRING THE DEVICE UNTIL THE END. LESS THAN 500CC BLEEDING FROM THE PROXIMAL SIDE OF THE STAPLE LINE OCCURRED; AND SINCE THE DEGREE OF STAPLE FORMATION COULD NOT BE CONFIRMED, SUTURE REINFORCEMENT WAS PERFORMED AND FURTHER IMPLANTATION IN THE SEROUS MUSCLE LAYER (SEROMUSCULAR). NO POSTOPERATIVE LEAK WAS NOTED, BUT ILEUS WAS OCCURRING.
04-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | GIA | Firing Problem | A0505 | Obstruction/Occlusion | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, WHILE CLOSING THE ENTRY HOLE FOR FEEA (FUNCTIONAL END-TO-END ANAST OMOSIS) RECONSTRUCTION OF THE COLON OUTSIDE THE BODY, THE DEVICE WAS DIFFICULT TO FIRE. IT INITIALLY STOPPED FIRING MIDWAY, BUT THE SURGEON REAPPLIED FORCE AND WAS ABLE TO COMPLETE FIRING THE DEVICE UNTIL THE END. LESS THAN 500CC BLEEDING FROM THE PROXIMAL SIDE OF THE STAPLE LINE OCCURRED; AND SINCE THE DEGREE OF STAPLE FORMATION COULD NOT BE CONFIRMED, SUTURE REINFORCEMENT WAS PERFORMED AND FURTHER IMPLANTATION IN THE SEROUS MUSCLE LAYER (SEROMUSCULAR). NO POSTOPERATIVE LEAK WAS NOTED, BUT ILEUS WAS OCCURRING.
04-01-2023 | Injury | us surgical puerto rico | 26-01-2023 | GIA | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, WHILE CLOSING THE ENTRY HOLE FOR FEEA (FUNCTIONAL END-TO-END ANAST OMOSIS) RECONSTRUCTION OF THE COLON OUTSIDE THE BODY, THE DEVICE WAS DIFFICULT TO FIRE. IT INITIALLY STOPPED FIRING MIDWAY, BUT THE SURGEON REAPPLIED FORCE AND WAS ABLE TO COMPLETE FIRING THE DEVICE UNTIL THE END. LESS THAN 500CC BLEEDING FROM THE PROXIMAL SIDE OF THE STAPLE LINE OCCURRED; AND SINCE THE DEGREE OF STAPLE FORMATION COULD NOT BE CONFIRMED, SUTURE REINFORCEMENT WAS PERFORMED AND FURTHER IMPLANTATION IN THE SEROUS MUSCLE LAYER (SEROMUSCULAR). NO POSTOPERATIVE LEAK WAS NOTED, BUT ILEUS WAS OCCURRING.
04-10-2021 | Malfunction | covidien lp llc north haven | 26-01-2023 | UNKNOWN EGIA TRI STAPLE | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE IN THE STOMACH, THE DEVICE MISFIRED, GOT STUCK ON THE PATIENT'S TISSUE, AND WAS UNABLE TO OPEN THE RELOAD TO COMPLETE THE FIRING. THE SURGEON CAREFULLY REMOVED THE MISFIRED DEVICE WITHOUT HARMING THE BOWEL. THERE WAS NO PATIENT INJURY.
04-10-2021 | Malfunction | covidien lp llc north haven | 26-01-2023 | UNKNOWN EGIA TRI STAPLE | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE IN THE STOMACH, THE DEVICE MISFIRED, GOT STUCK ON THE PATIENT'S TISSUE, AND WAS UNABLE TO OPEN THE RELOAD TO COMPLETE THE FIRING. THE SURGEON CAREFULLY REMOVED THE MISFIRED DEVICE WITHOUT HARMING THE BOWEL. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | covidien lp llc north haven | 26-01-2023 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH REPAIR (IPOM), WHILE ON FIXATION OF THE MESH, THE TACKER WAS JAMMED AND THE SURGEON WAS UNABLE TO FIRE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE SINCE THE SURGEON HAD TO USE ANOTHER TACKER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | covidien lp llc north haven | 26-01-2023 | ABSORBATACK | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH REPAIR (IPOM), WHILE ON FIXATION OF THE MESH, THE TACKER WAS JAMMED AND THE SURGEON WAS UNABLE TO FIRE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE SINCE THE SURGEON HAD TO USE ANOTHER TACKER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2022 | Malfunction | ethicon endo-surgery | 26-01-2023 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2022 | Malfunction | ethicon endo-surgery | 26-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE "STAPLER WAS LOADED WITH GREEN LOAD, BATTERY INSERTED. WHEN PLACED INSIDE THE PATIENT, THE STAPLER WOULD NOT FIRE NO MATTER WHAT WE DID. WE TRIED A NEW STAPLER BATTERY FROM A NEWLY OPENED STAPLER AND THIS ONE STILL WOULDN'T FIRE." THERE WERE NO PATIENT CONSEQUENCES.
16-12-2022 | Malfunction | ethicon endo-surgery | 26-01-2023 | PROXIMATE*75MM LINEAR CUTTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN EXPLORATORY LAPAROTOMY WITH A BOWEL RESECTION THAT ALL THREE STAPLE LINES DID NOT FIRE. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
05-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | PROXIMATE*RL LINEAR STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON WAS PERFORMING SIDE TO SIDE ANASTOMOSIS AND DEVICE MISFIRED. THE STAPLES DID NOT FULLY FORM ON ONE SIDE. THERE WERE NO PATIENT CONSEQUENCES.
06-11-2022 | Malfunction | ethicon endo-surgery | 26-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE MISS FIRING OCCURRED IN THE THIRD RELOAD (GST60B) ON THE STAPLE LINE AND THERE WAS LEAK AND BLEEDING DURING THE CASE, AND THE SURGEON MADE STITCHES WITH SUTURE IN THE AREA TO AVOID BLEEDING AND LEAKAGE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
06-11-2022 | Malfunction | ethicon endo-surgery | 26-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE MISS FIRING OCCURRED IN THE THIRD RELOAD (GST60B) ON THE STAPLE LINE AND THERE WAS LEAK AND BLEEDING DURING THE CASE, AND THE SURGEON MADE STITCHES WITH SUTURE IN THE AREA TO AVOID BLEEDING AND LEAKAGE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
11-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMIHEPATECTOMY, WHILE DISSOCIATING LEFT HEPATIC PEDICLE, THERE WERE 2 RELOADS, THE DEVICE FIRED AND FAILED TO STAPLE ALL B-LINE. THE DEVICE WAS REPLACED WITH ANOTHER MANUFACTURER¿S PRODUCT TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, DURING A GASTROJEJUNAL ANASTOMOSIS, AFTER SETTING THE DEVICE ON THE TISSUE AND PRESSING THE GREEN BUTTON, THE DEVICE DID NOT ADVANCE HALFWAY THROUGH AFTER FIRING WAS STARTED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2022 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, WHEN THE SURGEON WAS STARTING TO STAPLE THE GASTRIC TISSUE, THE STAPLER DID NOT FIRE. THIS WAS TRIED TWO TIMES. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. ANOTHER KIND OF DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-11-2022 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHILE BEING APPLIED ACROSS A FISSURE ON THE LUNG, THE HANDLE WENT LOOSE WHEN GREEN BUTTON WAS PRESSED, AND SURGEON COULD NOT FIRE THE RELOAD. THE HANDLE WAS THEN REMOVED AND WAS REPLACED WITH A NEW HANDLE WITH THE SAME BLACK RELOAD AND WORKED WITH NO ISSUES TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
07-11-2022 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHILE BEING APPLIED ACROSS A FISSURE ON THE LUNG, THE HANDLE WENT LOOSE WHEN GREEN BUTTON WAS PRESSED, AND SURGEON COULD NOT FIRE THE RELOAD. THE HANDLE WAS THEN REMOVED AND WAS REPLACED WITH A NEW HANDLE WITH THE SAME BLACK RELOAD AND WORKED WITH NO ISSUES TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
01-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON STATES THAT THIN TISSUE GETS STUCK ON THE TROCAR PAST ORANGE BAND. IT FIRES AND THERE ARE NO LEAKS AND NO ISSUES, BUT THE TROCAR DOESN'T ALWAYS SIT ON ANVIL PROPERLY. PROCEDURE: UNKNOWN
01-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON STATES THAT THIN TISSUE GETS STUCK ON THE TROCAR PAST ORANGE BAND. IT FIRES AND THERE ARE NO LEAKS AND NO ISSUES, BUT THE TROCAR DOESN'T ALWAYS SIT ON ANVIL PROPERLY. PROCEDURE: UNKNOWN
11-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, WHEN TRYING TO TRANSECT DISTAL PART OF STOMACH, THE HANDLE MADE A CRACKING SOUND AS IF SOMETHING BROKE IN THE INSTRUMENT. THE DEVICE DID NOT FIRE, THE HANDLE WAS NOT SQUEEZED, BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. IT WAS NOTED THAT OPENING IT WAS POSSIBLE, ALTHOUGH MAYBE A BIT HARD TO OPEN. TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THE RELOAD WAS REMOVED FROM THE "CRACKING STAPLER", LOADED IT TO A NEW STAPLER AND EVERYTHING WORKED OUT. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, WHEN TRYING TO TRANSECT DISTAL PART OF STOMACH, THE HANDLE MADE A CRACKING SOUND AS IF SOMETHING BROKE IN THE INSTRUMENT. THE DEVICE DID NOT FIRE, THE HANDLE WAS NOT SQUEEZED, BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. IT WAS NOTED THAT OPENING IT WAS POSSIBLE, ALTHOUGH MAYBE A BIT HARD TO OPEN. TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THE RELOAD WAS REMOVED FROM THE "CRACKING STAPLER", LOADED IT TO A NEW STAPLER AND EVERYTHING WORKED OUT. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN WHIPPLE PROCEDURE, WHEN TRYING TO TRANSECT DISTAL PART OF STOMACH, THE HANDLE MADE A CRACKING SOUND AS IF SOMETHING BROKE IN THE INSTRUMENT. THE DEVICE DID NOT FIRE, THE HANDLE WAS NOT SQUEEZED, BUT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. IT WAS NOTED THAT OPENING IT WAS POSSIBLE, ALTHOUGH MAYBE A BIT HARD TO OPEN. TO RESOLVE THE ISSUE AND COMPLETE THE CASE, THE RELOAD WAS REMOVED FROM THE "CRACKING STAPLER", LOADED IT TO A NEW STAPLER AND EVERYTHING WORKED OUT. THERE WAS NO PATIENT INJURY.
15-11-2022 | Malfunction | ethicon endo-surgery | 26-01-2023 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE COULD NOT BE TRIGGERED. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED.
11-01-2023 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AT THE TIME OF RESECTION OF THE LUNG LOBES, WHEN THE HANDLE WAS LOADED WITH STAPLES, THE HANDLE EXPLODED AND THE STAPLES WERE BROKEN AND HAD A POOR FORMATION. SAME ISSUE OCCURRED ON THE SECOND RELOAD WITH THE SAME HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2023 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AT THE TIME OF RESECTION OF THE LUNG LOBES, WHEN THE HANDLE WAS LOADED WITH STAPLES, THE HANDLE EXPLODED AND THE STAPLES WERE BROKEN AND HAD A POOR FORMATION. SAME ISSUE OCCURRED ON THE SECOND RELOAD WITH THE SAME HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2022 | Malfunction | ethicon endo-surgery | 26-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE; BLADE DID NOT RETRACT FOLLOWING FIRING. BLADE MANUALLY RETRACTED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. DEVICE IS AVAILABLE FOR RETURN.
20-12-2022 | Malfunction | ethicon endo-surgery | 26-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE; BLADE DID NOT RETRACT FOLLOWING FIRING. BLADE MANUALLY RETRACTED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. DEVICE IS AVAILABLE FOR RETURN.
11-01-2023 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AT THE TIME OF RESECTION OF THE LUNG LOBES, WHEN THE HANDLE WAS LOADED WITH STAPLES, THE HANDLE EXPLODED AND THE STAPLES WERE BROKEN AND HAD A POOR FORMATION. SAME ISSUE OCCURRED ON THE SECOND RELOAD WITH THE SAME HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | celestica de monterrey s.a de c.v | 26-01-2023 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AT THE TIME OF RESECTION OF THE LUNG LOBES, WHEN THE HANDLE WAS LOADED WITH STAPLES, THE HANDLE EXPLODED AND THE STAPLES WERE BROKEN AND HAD A POOR FORMATION. SAME ISSUE OCCURRED ON THE SECOND RELOAD WITH THE SAME HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-11-2022 | Malfunction | ethicon endo-surgery | 26-01-2023 | 55MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAP GASTRECTOMY SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-01-2023 | Malfunction | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, IN THE GASTRIC, THE SURGEON COULD NOT PRESS THE GREEN BUTTON. THE HANDLE COULD NOT BE SQUEEZE. ADDITIONALLY, THE ARTICULATING KNOB BROKE OFF. TO RESOLVE THE ISSUE, SUTURES WERE USED. THERE WAS NO PATIENT INJURY.
02-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2022 | Malfunction | us surgical puerto rico | 26-01-2023 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTOSIGMOIDECTOMY, THE DEVICE DID NOT WORK AND THE ANVIL COULD NOT BE TITLED AFTER FIRING. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2022 | Malfunction | us surgical puerto rico | 26-01-2023 | EEA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTOSIGMOIDECTOMY, THE DEVICE DID NOT WORK AND THE ANVIL COULD NOT BE TITLED AFTER FIRING. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AFTER THE FIRST FIRING, THE RELOAD WAS LOADED ONTO THE LINEAR STAPLER APPROPRIATELY, WHEN THE SURGEON GRASPED THE APPENDIX AND PLACED THE STAPLER AROUND IT, CLOSED THE JAWS, AND ATTEMPTED TO FIRE WHEN A POP WAS HEARD. THE JAWS COULD NOT BE OPENED. THE SURGEON WAS ABLE TO DISENGAGE THE RELOAD FROM THE HANDLE AND USE A DIFFERENT TECHNIQUE TO REMOVE THE RELOAD AND THE APPENDIX TOGETHER. THE TWO SURGEON USED A DIFFERENT CLINICAL APPROACH TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AFTER THE FIRST FIRING, THE RELOAD WAS LOADED ONTO THE LINEAR STAPLER APPROPRIATELY, WHEN THE SURGEON GRASPED THE APPENDIX AND PLACED THE STAPLER AROUND IT, CLOSED THE JAWS, AND ATTEMPTED TO FIRE WHEN A POP WAS HEARD. THE JAWS COULD NOT BE OPENED. THE SURGEON WAS ABLE TO DISENGAGE THE RELOAD FROM THE HANDLE AND USE A DIFFERENT TECHNIQUE TO REMOVE THE RELOAD AND THE APPENDIX TOGETHER. THE TWO SURGEON USED A DIFFERENT CLINICAL APPROACH TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AFTER THE FIRST FIRING, THE RELOAD WAS LOADED ONTO THE LINEAR STAPLER APPROPRIATELY, WHEN THE SURGEON GRASPED THE APPENDIX AND PLACED THE STAPLER AROUND IT, CLOSED THE JAWS, AND ATTEMPTED TO FIRE WHEN A POP WAS HEARD. THE JAWS COULD NOT BE OPENED. THE SURGEON WAS ABLE TO DISENGAGE THE RELOAD FROM THE HANDLE AND USE A DIFFERENT TECHNIQUE TO REMOVE THE RELOAD AND THE APPENDIX TOGETHER. THE TWO SURGEON USED A DIFFERENT CLINICAL APPROACH TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | UNKNOWN GIA DST PRODUCT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-11-2021 | Injury | us surgical puerto rico | 26-01-2023 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY EVALUATED 21 PATIENTS WHO UNDERWENT AN ANASTOMOTIC PROCEDURE BETWEEN 2016 AND 2017. THERE WERE POSTOPERATIVE COMPLICATIONS OF ANASTOMOTIC FAILURE IN ONE PATIENT USING THE CIRCULAR STAPLER. A 50-YEAR OLD PATIENT WITH A DIAGNOSIS OF CANCER OF THE MIDDLE AMPULLARY RECTUM, STAGE 4, UNDERWENT LAPAROSCOPIC TOTAL MESORECTUMECTOMY WITH ULTRA-LOW RECTAL AND LIVER RESECTION. A HARDWARE CIRCULAR ANASTOMOSIS WITH AUTO-SUTURE 28 WAS FORMED. ON DAY SIX, THE COLORECTAL ANASTOMOSIS FAILED. REVISION OF THE SMALL PELVIS, ILIAC AREAS, WHERE THERE WAS A SMALL AMOUNT OF CLOUDY DISCHARGE, AREAS OF PERITONEUM WITH FIBRIN, NO CHYME. THE ABDOMINAL CAVITY WAS WASHED WITH CLEAN WATER. IN ORDER TO PREVENT THE FAILURE OF LOW COLORECTAL AND COLOANAL ANASTOMOSES, AN ILEOSTOMY WAS NOT REMOVED.
12-11-2021 | Injury | us surgical puerto rico | 26-01-2023 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY EVALUATED 21 PATIENTS WHO UNDERWENT AN ANASTOMOTIC PROCEDURE BETWEEN 2016 AND 2017. THERE WERE POSTOPERATIVE COMPLICATIONS OF ANASTOMOTIC FAILURE IN ONE PATIENT USING THE CIRCULAR STAPLER. A 50-YEAR OLD PATIENT WITH A DIAGNOSIS OF CANCER OF THE MIDDLE AMPULLARY RECTUM, STAGE 4, UNDERWENT LAPAROSCOPIC TOTAL MESORECTUMECTOMY WITH ULTRA-LOW RECTAL AND LIVER RESECTION. A HARDWARE CIRCULAR ANASTOMOSIS WITH AUTO-SUTURE 28 WAS FORMED. ON DAY SIX, THE COLORECTAL ANASTOMOSIS FAILED. REVISION OF THE SMALL PELVIS, ILIAC AREAS, WHERE THERE WAS A SMALL AMOUNT OF CLOUDY DISCHARGE, AREAS OF PERITONEUM WITH FIBRIN, NO CHYME. THE ABDOMINAL CAVITY WAS WASHED WITH CLEAN WATER. IN ORDER TO PREVENT THE FAILURE OF LOW COLORECTAL AND COLOANAL ANASTOMOSES, AN ILEOSTOMY WAS NOT REMOVED.
10-01-2023 | Malfunction | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE, DUE TO THE RELOADS BEING FIRED AHEAD OF TIME. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE, DUE TO THE RELOADS BEING FIRED AHEAD OF TIME. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE, DUE TO THE RELOADS BEING FIRED AHEAD OF TIME. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE, DUE TO THE RELOADS BEING FIRED AHEAD OF TIME. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE, DUE TO THE RELOADS BEING FIRED AHEAD OF TIME. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-11-2022 | Malfunction | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE SURGEON PREPARED THE DEVICE WITH AN INSTALLED SEAMGUARD (REINFORCEMENT MATERIAL) ON THE RELOAD. THE SURGEON WAS UNABLE TO CLOSE THE JAWS COMPLETELY DUE TO THE THICKNESS OF THE TISSUE AND SEAMGUARD. THEREFORE THE SURGEON COULD NOT PROCEED WITH PRESSING THE GREEN BUTTON. JAWS WERE UNABLE TO CLOSE COMPLETELY ON THE THREE RELOADS THAT WERE USED ON THE TWO HANDLES. THE SURGEON ALSO USED A BIPOLAR INSTRUMENT FOR THE HEMOSTASIS. THERE WAS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOTOMY/LOBECTOMY PROCEDURE, THE SURGEON WAS USING THE DEVICE ON THE PULMONARY VEIN. THIS WAS THE THIRD RELOAD FIRE. NO ISSUE WAS NOTED ON THE FIRST TWO FIRINGS. HE REPORTED TO ME THAT THE DEVICE PARTIALLY FIRED, "JUST STOPPED" AND HAD "LOCKED OUT". THE REVERSE KNIFE SWITCH BROUGHT THE BLADE BACK AND THE DEVICE WAS OPENED WITHOUT DIFFICULTY. UPON OPENING, THE VESSEL WAS BLEEDING BRISKLY. THE RELOAD APPEARED TO BE FULLY DEPLOYED (ALL THE DRIVERS APPEARED UP) WITH A COUPLE OF LOOSE, FORMED STAPLES NOTED. A VESSEL CLAMP WAS PLACED AND THE SURGEON USED THE SAME DEVICE WITH A NEW RELOAD ON THE SAME VESSEL. THE DEVICE FIRED AS EXPECTED WITHOUT ISSUE. UPON REMOVING THE DEVICE, THE VESSEL CONTINUED TO BLEED BRISKLY AND WAS OVERSEWN WITH PROLENE. THE SURGEON NOTED THAT ALL THE STAPLES APPEARED WELL FORMED. THE SURGEON SHARED THAT THERE WAS AN INITIAL "TEAR" IN THE VEIN PRIOR TO THE THIRD AND FOURTH FIRING OF THE DEVICE. THE DEVICE AND THE LAST TWO FIRED RELOADS WERE REMOVED FROM FIELD AND A LIKE DEVICE WAS USED TO COMPLETE THE CASE WITHOUT DIFFICULTY. THE SURGERY WAS DELAYED BY FIVE-TEN MINUTES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
01-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOTOMY/LOBECTOMY PROCEDURE, THE SURGEON WAS USING THE DEVICE ON THE PULMONARY VEIN. THIS WAS THE THIRD RELOAD FIRE. NO ISSUE WAS NOTED ON THE FIRST TWO FIRINGS. HE REPORTED TO ME THAT THE DEVICE PARTIALLY FIRED, "JUST STOPPED" AND HAD "LOCKED OUT". THE REVERSE KNIFE SWITCH BROUGHT THE BLADE BACK AND THE DEVICE WAS OPENED WITHOUT DIFFICULTY. UPON OPENING, THE VESSEL WAS BLEEDING BRISKLY. THE RELOAD APPEARED TO BE FULLY DEPLOYED (ALL THE DRIVERS APPEARED UP) WITH A COUPLE OF LOOSE, FORMED STAPLES NOTED. A VESSEL CLAMP WAS PLACED AND THE SURGEON USED THE SAME DEVICE WITH A NEW RELOAD ON THE SAME VESSEL. THE DEVICE FIRED AS EXPECTED WITHOUT ISSUE. UPON REMOVING THE DEVICE, THE VESSEL CONTINUED TO BLEED BRISKLY AND WAS OVERSEWN WITH PROLENE. THE SURGEON NOTED THAT ALL THE STAPLES APPEARED WELL FORMED. THE SURGEON SHARED THAT THERE WAS AN INITIAL "TEAR" IN THE VEIN PRIOR TO THE THIRD AND FOURTH FIRING OF THE DEVICE. THE DEVICE AND THE LAST TWO FIRED RELOADS WERE REMOVED FROM FIELD AND A LIKE DEVICE WAS USED TO COMPLETE THE CASE WITHOUT DIFFICULTY. THE SURGERY WAS DELAYED BY FIVE-TEN MINUTES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES.
20-12-2022 | Injury | covidien lp llc north haven | 26-01-2023 | TRI-STAPLE 2.0 | Material Rupture | A0412 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, IT WAS NOTED THAT THE PURPLE REINFORCED RELOAD WAS USED IN GASTROJEJUNOSTOMY. FUNCTIONAL END TO END (FEE) ANASTOMOSIS WAS DONE WITH ANOTHER TYPE OF REINFORCED PURPLE RELOAD FOR CLOSURE OF THE INSERTION HOLE. THE NEXT DAY AFTER, THE EDGE OF THE REINFORCED SHEET WAS OPEN LIKE IT WAS TORN APART FROM A MAJOR LEAK IN THE GASTRO-JEJUNAL FEE ANASTOMOSIS. DUE TO THE MAJOR LEAK, THERE WAS TISSUE DAMAGE, EXTENDED INCISION AS THE PATIENT HAD REOPERATION WITH REANASTOMOSIS.
20-12-2022 | Injury | covidien lp llc north haven | 26-01-2023 | TRI-STAPLE 2.0 | Material Rupture | A0412 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, IT WAS NOTED THAT THE PURPLE REINFORCED RELOAD WAS USED IN GASTROJEJUNOSTOMY. FUNCTIONAL END TO END (FEE) ANASTOMOSIS WAS DONE WITH ANOTHER TYPE OF REINFORCED PURPLE RELOAD FOR CLOSURE OF THE INSERTION HOLE. THE NEXT DAY AFTER, THE EDGE OF THE REINFORCED SHEET WAS OPEN LIKE IT WAS TORN APART FROM A MAJOR LEAK IN THE GASTRO-JEJUNAL FEE ANASTOMOSIS. DUE TO THE MAJOR LEAK, THERE WAS TISSUE DAMAGE, EXTENDED INCISION AS THE PATIENT HAD REOPERATION WITH REANASTOMOSIS.
20-12-2022 | Injury | covidien lp llc north haven | 26-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, IT WAS NOTED THAT THE PURPLE REINFORCED RELOAD WAS USED IN GASTROJEJUNOSTOMY. FUNCTIONAL END TO END (FEE) ANASTOMOSIS WAS DONE WITH ANOTHER TYPE OF REINFORCED PURPLE RELOAD FOR CLOSURE OF THE INSERTION HOLE. THE NEXT DAY AFTER, THE EDGE OF THE REINFORCED SHEET WAS OPEN LIKE IT WAS TORN APART FROM A MAJOR LEAK IN THE GASTRO-JEJUNAL FEE ANASTOMOSIS. DUE TO THE MAJOR LEAK, THERE WAS TISSUE DAMAGE, EXTENDED INCISION AS THE PATIENT HAD REOPERATION WITH REANASTOMOSIS.
20-12-2022 | Injury | covidien lp llc north haven | 26-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, IT WAS NOTED THAT THE PURPLE REINFORCED RELOAD WAS USED IN GASTROJEJUNOSTOMY. FUNCTIONAL END TO END (FEE) ANASTOMOSIS WAS DONE WITH ANOTHER TYPE OF REINFORCED PURPLE RELOAD FOR CLOSURE OF THE INSERTION HOLE. THE NEXT DAY AFTER, THE EDGE OF THE REINFORCED SHEET WAS OPEN LIKE IT WAS TORN APART FROM A MAJOR LEAK IN THE GASTRO-JEJUNAL FEE ANASTOMOSIS. DUE TO THE MAJOR LEAK, THERE WAS TISSUE DAMAGE, EXTENDED INCISION AS THE PATIENT HAD REOPERATION WITH REANASTOMOSIS.
10-01-2023 | Injury | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PNEUMOTHORAX UPPER LEFT PROCEDURE, FOR THE FIRST RELOAD, IT GOT STUCK AFTER ONLY FIRING 1 CM. THE STAPLE DID NOT FORM AND THE KNIFE DID NOT COME OUT. THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. THERE WAS AN UNANTICIPATED TISSUE LOSS, TISSUE WAS DAMAGED, AND INCISION WAS EXTENDED NOT MORE THAN 1 INCH DUE PRODUCT PROBLEM. THE GRIP WAS VERY TIGHT WHEN FIRED, DOCTOR WAS ABLE TO SUCCESSFULLY RETURN IT BACK, HOWEVER DOCTOR USED NEW RELOAD NEXT TO THE TISSUE SUCCESSFULLY. AFTER THAT, 3 PIECES RELOAD SUCCESSFULLY FIRED WITH THE SAME HANDLE.
10-01-2023 | Injury | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PNEUMOTHORAX UPPER LEFT PROCEDURE, FOR THE FIRST RELOAD, IT GOT STUCK AFTER ONLY FIRING 1 CM. THE STAPLE DID NOT FORM AND THE KNIFE DID NOT COME OUT. THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. THERE WAS AN UNANTICIPATED TISSUE LOSS, TISSUE WAS DAMAGED, AND INCISION WAS EXTENDED NOT MORE THAN 1 INCH DUE PRODUCT PROBLEM. THE GRIP WAS VERY TIGHT WHEN FIRED, DOCTOR WAS ABLE TO SUCCESSFULLY RETURN IT BACK, HOWEVER DOCTOR USED NEW RELOAD NEXT TO THE TISSUE SUCCESSFULLY. AFTER THAT, 3 PIECES RELOAD SUCCESSFULLY FIRED WITH THE SAME HANDLE.
20-12-2022 | Injury | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Material Rupture | A0412 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, IT WAS NOTED THAT THE PURPLE REINFORCED RELOAD WAS USED IN GASTROJEJUNOSTOMY. FUNCTIONAL END TO END (FEE) ANASTOMOSIS WAS DONE WITH ANOTHER TYPE OF REINFORCED PURPLE RELOAD FOR CLOSURE OF THE INSERTION HOLE. THE NEXT DAY AFTER, THE EDGE OF THE REINFORCED SHEET WAS OPEN LIKE IT WAS TORN APART FROM A MAJOR LEAK IN THE GASTRO-JEJUNAL FEE ANASTOMOSIS. DUE TO THE MAJOR LEAK, THERE WAS TISSUE DAMAGE, EXTENDED INCISION AS THE PATIENT HAD REOPERATION WITH REANASTOMOSIS.
20-12-2022 | Injury | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Material Rupture | A0412 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, IT WAS NOTED THAT THE PURPLE REINFORCED RELOAD WAS USED IN GASTROJEJUNOSTOMY. FUNCTIONAL END TO END (FEE) ANASTOMOSIS WAS DONE WITH ANOTHER TYPE OF REINFORCED PURPLE RELOAD FOR CLOSURE OF THE INSERTION HOLE. THE NEXT DAY AFTER, THE EDGE OF THE REINFORCED SHEET WAS OPEN LIKE IT WAS TORN APART FROM A MAJOR LEAK IN THE GASTRO-JEJUNAL FEE ANASTOMOSIS. DUE TO THE MAJOR LEAK, THERE WAS TISSUE DAMAGE, EXTENDED INCISION AS THE PATIENT HAD REOPERATION WITH REANASTOMOSIS.
20-12-2022 | Injury | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, IT WAS NOTED THAT THE PURPLE REINFORCED RELOAD WAS USED IN GASTROJEJUNOSTOMY. FUNCTIONAL END TO END (FEE) ANASTOMOSIS WAS DONE WITH ANOTHER TYPE OF REINFORCED PURPLE RELOAD FOR CLOSURE OF THE INSERTION HOLE. THE NEXT DAY AFTER, THE EDGE OF THE REINFORCED SHEET WAS OPEN LIKE IT WAS TORN APART FROM A MAJOR LEAK IN THE GASTRO-JEJUNAL FEE ANASTOMOSIS. DUE TO THE MAJOR LEAK, THERE WAS TISSUE DAMAGE, EXTENDED INCISION AS THE PATIENT HAD REOPERATION WITH REANASTOMOSIS.
20-12-2022 | Injury | covidien lp llc north haven | 26-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, AT THE TIME OF LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY, IT WAS NOTED THAT THE PURPLE REINFORCED RELOAD WAS USED IN GASTROJEJUNOSTOMY. FUNCTIONAL END TO END (FEE) ANASTOMOSIS WAS DONE WITH ANOTHER TYPE OF REINFORCED PURPLE RELOAD FOR CLOSURE OF THE INSERTION HOLE. THE NEXT DAY AFTER, THE EDGE OF THE REINFORCED SHEET WAS OPEN LIKE IT WAS TORN APART FROM A MAJOR LEAK IN THE GASTRO-JEJUNAL FEE ANASTOMOSIS. DUE TO THE MAJOR LEAK, THERE WAS TISSUE DAMAGE, EXTENDED INCISION AS THE PATIENT HAD REOPERATION WITH REANASTOMOSIS.
12-11-2021 | Injury | covidien lp llc north haven | 26-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Peritonitis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY EVALUATED 21 PATIENTS WHO UNDERWENT AN ANASTOMOTIC PROCEDURE BETWEEN 2016 AND 2017. THERE WERE POSTOPERATIVE COMPLICATIONS OF ANASTOMOTIC FAILURE IN ONE PATIENT USING A LINEAR STAPLER. A 77-YEAR-OLD PATIENT WITH GASTRIC CANCER UNDERWENT LAPAROSCOPIC GASTRECTOMY WITH LYMPH DISSECTION. THE ESOPHAGOJUNOANASTOMOSIS AND THE JEJUNOJEJUNAL ANASTOMOSIS WERE PERFORMED WITH ENDO GIA. ON DAY 5, A PARTIAL FAILURE OF THE ESOPHAGOJUNOANASTOMOSIS OCCURRED. REOPERATION WAS PERFORMED TO REPAIR THE DEFECT AND TO REVISE THE JEJUNOJEJUNAL ANASTOMOSIS. ON THE 2ND DAY AFTER LAPAROTOMY, THE FAILURE OF ESOPHAGOJUNOANASTOMOSIS DEVELOPED, WHICH LED TO PERITONITIS. A SECOND REOPERATION WAS PERFORMED WHERE IT WAS NOTED THAT THERE WAS NO FUSION OF THE ANASTOMOSIS. IT IS UNCLEAR WHICH 60 MM LINEAR RELOAD DEVICE WAS USED FOR THE ANASTOMOSIS. A LEAK TEST WAS NOT PERFORMED.
12-11-2021 | Injury | covidien lp llc north haven | 26-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY EVALUATED 21 PATIENTS WHO UNDERWENT AN ANASTOMOTIC PROCEDURE BETWEEN 2016 AND 2017. THERE WERE POSTOPERATIVE COMPLICATIONS OF ANASTOMOTIC FAILURE IN ONE PATIENT USING A LINEAR STAPLER. A 77-YEAR-OLD PATIENT WITH GASTRIC CANCER UNDERWENT LAPAROSCOPIC GASTRECTOMY WITH LYMPH DISSECTION. THE ESOPHAGOJUNOANASTOMOSIS AND THE JEJUNOJEJUNAL ANASTOMOSIS WERE PERFORMED WITH ENDO GIA. ON DAY 5, A PARTIAL FAILURE OF THE ESOPHAGOJUNOANASTOMOSIS OCCURRED. REOPERATION WAS PERFORMED TO REPAIR THE DEFECT AND TO REVISE THE JEJUNOJEJUNAL ANASTOMOSIS. ON THE 2ND DAY AFTER LAPAROTOMY, THE FAILURE OF ESOPHAGOJUNOANASTOMOSIS DEVELOPED, WHICH LED TO PERITONITIS. A SECOND REOPERATION WAS PERFORMED WHERE IT WAS NOTED THAT THERE WAS NO FUSION OF THE ANASTOMOSIS. IT IS UNCLEAR WHICH 60 MM LINEAR RELOAD DEVICE WAS USED FOR THE ANASTOMOSIS. A LEAK TEST WAS NOT PERFORMED.
08-12-2022 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY, WHILE AT THE LUNG TISSUE, THE FIRST THREE DEVICES HAVE AN INCOMPLETE STAPLE LINE CENTRALLY WHILE THE FOURTH HANDLE HAD AN INCOMPLETE STAPLE LINE DISTALLY. ALL FOUR HANDLES WERE DIFFICULT TO FIRE. THE TISSUE WAS CALCIFIED AND HARD. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON EACH FIRING RACK. BASED ON THE EVIDENCE AVAILABLE THE REPORTED CONDITION OF THE INSTRUMENT WAS DIFFICULT TO FIRE WAS CONFIRMED.
08-12-2022 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Firing Problem | A0505 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY, WHILE AT THE LUNG TISSUE, THE FIRST THREE DEVICES HAVE AN INCOMPLETE STAPLE LINE CENTRALLY WHILE THE FOURTH HANDLE HAD AN INCOMPLETE STAPLE LINE DISTALLY. ALL FOUR HANDLES WERE DIFFICULT TO FIRE. THE TISSUE WAS CALCIFIED AND HARD. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON EACH FIRING RACK. BASED ON THE EVIDENCE AVAILABLE THE REPORTED CONDITION OF THE INSTRUMENT WAS DIFFICULT TO FIRE WAS CONFIRMED.
11-01-2023 | Malfunction | us surgical puerto rico | 26-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE USING THE HANDLE, THE FIRST RELOAD WAS LOADED WITH A PIN WHEN DISENGAGING THE PATIENT¿S BRONCHUS, AND DURING FIRST FIRING, THE STAPLE WAS STUCK AND THE DISTAL END OF THE SUTURE WAS INCOMPLETE. AFTER REPLACING THE RELOAD, IT WAS FOUND THAT THE GRIP WAS SQUEEZED, BUT DEVICE COULD NOT BE FIRED. THE GREEN BUTTON WAS PRESSED. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
16-12-2022 | Malfunction | intuitive surgical | 26-01-2023 | SUREFORM | Failure to Deliver Energy | A090402 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GENERAL SURGICAL PROCEDURE, THE SUREFORM STAPLER 45 HAD AN UNSPECIFIED ISSUE. THE REPORTED INFORMATION INDICATED IT WAS UNKNOWN IF A FRAGMENT FELL INSIDE THE PATIENT AS A RESULT OF THE ISSUE. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER WHO STATED THAT THEY HAD NO FURTHER INFORMATION TO PROVIDE.
06-01-2023 | Malfunction | ethicon endo-surgery | 26-01-2023 | GST GREEN RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
10-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE PROCEDURE, THERE WERE FIVE RELOADS USED. THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND THE HANDLE WAS NOT SQUEEZED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE PROCEDURE, THERE WERE FIVE RELOADS USED. THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND THE HANDLE WAS NOT SQUEEZED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | LINEAR CUTTER/STAPLER-STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY OPENED THE PACKING, AND IT WAS NOTED THAT THE JJ55B RELOADS DO NOT HAVE SAFETY-LOCK-OUT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
10-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE PROCEDURE, THERE WERE FIVE RELOADS USED. THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, AND THE HANDLE WAS NOT SQUEEZED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE MISFIRED USING THE RELOAD. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 27-01-2023 | UNKNOWN EGIA TRI STAPLE | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH RESECTION BY DA VINCI, DELTA RECONSTRUCTION PROCEDURE, PROTRUDING PART ON THE YELLOW TAB OF THE CARTRIDGE THAT PREVENTS PRE-FIRE WAS DAMAGED, AND THE BROKEN PIECE FELL IN THE ABDOMINAL CAVITY WHICH WAS THEN COLLECTED. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 27-01-2023 | UNKNOWN EGIA TRI STAPLE | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH RESECTION BY DA VINCI, DELTA RECONSTRUCTION PROCEDURE, PROTRUDING PART ON THE YELLOW TAB OF THE CARTRIDGE THAT PREVENTS PRE-FIRE WAS DAMAGED, AND THE BROKEN PIECE FELL IN THE ABDOMINAL CAVITY WHICH WAS THEN COLLECTED. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 27-01-2023 | UNKNOWN EGIA TRI STAPLE | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH RESECTION BY DA VINCI, DELTA RECONSTRUCTION PROCEDURE, PROTRUDING PART ON THE YELLOW TAB OF THE CARTRIDGE THAT PREVENTS PRE-FIRE WAS DAMAGED, AND THE BROKEN PIECE FELL IN THE ABDOMINAL CAVITY WHICH WAS THEN COLLECTED. THERE WAS NO PATIENT INJURY.
26-12-2022 | Malfunction | covidien lp llc north haven | 27-01-2023 | UNKNOWN EGIA TRI STAPLE | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH RESECTION BY DA VINCI, DELTA RECONSTRUCTION PROCEDURE, PROTRUDING PART ON THE YELLOW TAB OF THE CARTRIDGE THAT PREVENTS PRE-FIRE WAS DAMAGED, AND THE BROKEN PIECE FELL IN THE ABDOMINAL CAVITY WHICH WAS THEN COLLECTED. THERE WAS NO PATIENT INJURY.
22-11-2022 | Malfunction | ethicon endo-surgery | 27-01-2023 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AFTER LOADING FOR THE SECOND TIME, FIRING WAS NOT POSSIBLE ANYMORE. NO PATIENT HARM REPORTED.
21-11-2022 | Malfunction | ethicon endo-surgery | 27-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER COULD NOT BE FIRED AND WAS BLOCKED. NO PATIENT CONSEQUENCES REPORTED.
09-11-2022 | Malfunction | us surgical puerto rico | 27-01-2023 | TA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
12-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THERE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
09-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED, IT RETURNED AFTER FIRING HALF. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED, IT RETURNED AFTER FIRING HALF. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREATIC DUODENECTOMY, WHILE DISSOCIATING THE GASTRIC ANTRUM, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. ANOTHER SET WITH THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED, IT RETURNED AFTER FIRING HALF.
09-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE PARTIALLY FIRED, IT RETURNED AFTER FIRING HALF.
19-12-2022 | Malfunction | intuitive surgical | 27-01-2023 | SUREFORM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ON (B)(6)2023, INTUITIVE SURGICAL INC (ISI) RECEIVED USER FACILITY REPORT 4600060000-2022-8006 STATING: DOCTOR WAS USING THE SUREFORM STAPLER 60 ON THE STOMACH RESECTION. DURING THREE SEPARATE TIMES, THE STAPLER FAILED TO REACH THE COMPLETE DISTANCE THAT WAS REQUIRED TO COMPLETE THE STAPLER LINE. NEW STAPLE LOADS WERE OPENED EACH TIME, AND EVENTUALLY A NEW STAPLER. ISI FOLLOWED UP WITH THE SITE'S CLINICAL SALES REP AND THE FOLLOWING ADDITIONAL INFORMATION WAS OBTAINED: THE SUREFORM 60 BLUE RELOADS WERE THE RELOADS INVOLVED WITH THE REPORTED EVENT. THE RELOAD WAS SELECTED BECAUSE IT IS THE RELOAD THE SURGEON TYPICALLY USES FOR 1 RELOAD ON SLEEVES. STAPLER FIRES 1-3 WHERE THE STAPLER FIRES INVOLVE WITH THE REPORTED EVENT. A STOMACH TRANSACTION WAS THE SURGICAL TASK BEING PERFORMED AT THE TIME OF THE EVENT. THE STOMACH WAS THE TARGET TISSUE. THE TARGET TISSUE WAS NOT CALCIFIED AND WAS NOT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY, THERE WAS NO TISSUE TENSION OR BUNCHING. THE EVENT DID NOT INVOLVE A FAILURE TO FIRE. THE STAPLER SEEMED TO STOP IN THE SAME AREA X3 SEPARATE TIMES, WITH THE TISSUE BEING TOO THICK TO CONTINUE. THERE WAS NO TISSUE BEING PUSHED FORWARD/DRAGGED DURING THE FIRING PROCESS. THE JAWS DID NOT OPEN/RELEASE DURING THE FIRING SEQUENCE. THE RELOAD STAPLES DID NOT DEPLOY UNEXPECTEDLY. THERE WERE NO MALFORMED STAPLES. THE RELOAD DID HAVE AN EXPOSED BLADE. THERE WERE NO STAPLES MISSING FROM THE STAPLE LINE. THERE WERE NO HOLES/GAPS IN THE STAPLE LINE. A "TISSUE TOO THICK" ERROR MESSAGE OCCURRED. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS BETWEEN THE INSTRUMENT JAWS. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. NO BUTTRESS MATERIAL WAS USED. NO BLEEDING ON THE STAPLE LINE WAS OBSERVED. THERE WAS NO UNPLANNED TISSUE REMOVAL DUE TO THE FIRING FAILURE. THE SURGEON RESOLVED THE ISSUE BY OPENING A NEW STAPLER. THE PROCEDURE WAS COMPLETED ROBOTICALLY AS PLANNED WITH NO REPORTS OF PATIENT INJURY. THE RELOAD IS NOT AVAILABLE FOR RETURN.
10-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREATIC DUODENECTOMY, WHILE DISSOCIATING THE GASTRIC ANTRUM, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. ANOTHER SET WITH THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE-HOLE THORACOSCOPIC PNEUMOTOMY, WHILE PREPARING TO DISSOCIATE THE VESSEL, THE DEVICE DID NOT FIRE, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS NOT SQUEEZED. A NEW HANDLE WAS USED, AND THE PROCEDURE WENT SMOOTHLY. THERE WAS NO PATIENT INJURY.
12-01-2023 | Injury | covidien lp llc north haven | 27-01-2023 | TRI-STAPLE 2.0 | Failure to Cycle | A050703 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE FIRST TWO FIRINGS OF THE POWERED STAPLER WERE SUCCESSFUL. THERE WERE NO OBVIOUS ISSUES IN LOADING OR SETTING UP THE DEVICE. ON THE NEXT FIRING, IT APPEARED NORMAL AT FIRST. APPROXIMATELY 2/3 OF THE WAY INTO FIRING, IT SEEMED TO START PUSHING TISSUE WHILE PLACING STAPLES. UPON RETRACTING STAPLER, THE DEVICE HAD NOT APPEARED TO CUT TISSUE PROPERLY, AND CONSEQUENTLY STAPLES BEGAN TO BUNCH AND MALFORMED AROUND THE DISTAL END OF THE STAPLE LINE. THERE WAS ALSO AN INCOMPLETE STAPLE LINE DISTALLY EVEN THOUGH HANDLE WAS COMPLETELY FIRED. A MANUAL STAPLER AND A NEW RELOAD WERE USED TO FIRE ANOTHER LINE, MEDIAL (TOWARDS THE BOUGIE) TO THE AFFECTED AREA, THIS RESECTED DAMAGED TISSUE TOWARDS THE SIDE OF THE GASTRIC SPECIMEN. SOME TISSUE WAS LOST AND DAMAGED AS A RESULT, TIGHTER THAN USUAL SLEEVE AND CAUSED A DEVIATION FROM THE SURGEONS' USUAL TECHNIQUE.
12-01-2023 | Injury | covidien lp llc north haven | 27-01-2023 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE FIRST TWO FIRINGS OF THE POWERED STAPLER WERE SUCCESSFUL. THERE WERE NO OBVIOUS ISSUES IN LOADING OR SETTING UP THE DEVICE. ON THE NEXT FIRING, IT APPEARED NORMAL AT FIRST. APPROXIMATELY 2/3 OF THE WAY INTO FIRING, IT SEEMED TO START PUSHING TISSUE WHILE PLACING STAPLES. UPON RETRACTING STAPLER, THE DEVICE HAD NOT APPEARED TO CUT TISSUE PROPERLY, AND CONSEQUENTLY STAPLES BEGAN TO BUNCH AND MALFORMED AROUND THE DISTAL END OF THE STAPLE LINE. THERE WAS ALSO AN INCOMPLETE STAPLE LINE DISTALLY EVEN THOUGH HANDLE WAS COMPLETELY FIRED. A MANUAL STAPLER AND A NEW RELOAD WERE USED TO FIRE ANOTHER LINE, MEDIAL (TOWARDS THE BOUGIE) TO THE AFFECTED AREA, THIS RESECTED DAMAGED TISSUE TOWARDS THE SIDE OF THE GASTRIC SPECIMEN. SOME TISSUE WAS LOST AND DAMAGED AS A RESULT, TIGHTER THAN USUAL SLEEVE AND CAUSED A DEVIATION FROM THE SURGEONS' USUAL TECHNIQUE.
12-01-2023 | Injury | covidien lp llc north haven | 27-01-2023 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE FIRST TWO FIRINGS OF THE POWERED STAPLER WERE SUCCESSFUL. THERE WERE NO OBVIOUS ISSUES IN LOADING OR SETTING UP THE DEVICE. ON THE NEXT FIRING, IT APPEARED NORMAL AT FIRST. APPROXIMATELY 2/3 OF THE WAY INTO FIRING, IT SEEMED TO START PUSHING TISSUE WHILE PLACING STAPLES. UPON RETRACTING STAPLER, THE DEVICE HAD NOT APPEARED TO CUT TISSUE PROPERLY, AND CONSEQUENTLY STAPLES BEGAN TO BUNCH AND MALFORMED AROUND THE DISTAL END OF THE STAPLE LINE. THERE WAS ALSO AN INCOMPLETE STAPLE LINE DISTALLY EVEN THOUGH HANDLE WAS COMPLETELY FIRED. A MANUAL STAPLER AND A NEW RELOAD WERE USED TO FIRE ANOTHER LINE, MEDIAL (TOWARDS THE BOUGIE) TO THE AFFECTED AREA, THIS RESECTED DAMAGED TISSUE TOWARDS THE SIDE OF THE GASTRIC SPECIMEN. SOME TISSUE WAS LOST AND DAMAGED AS A RESULT, TIGHTER THAN USUAL SLEEVE AND CAUSED A DEVIATION FROM THE SURGEONS' USUAL TECHNIQUE.
11-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT HEMIHEPATECTOMY, WHILE DISSOCIATING LEFT HEPATIC PEDICLE, THERE WERE 2 RELOADS, THE DEVICE FIRED AND FAILED TO STAPLE ALL B-LINE. THE DEVICE WAS REPLACED WITH ANOTHER MANUFACTURER¿S PRODUCT TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | GST GOLD RELOAD, 60MM, 6 ROW | Structural Problem | A0511 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHEN THEY OPENED THE RELOAD IT LOOKED BENT AND HAD LOOSE DRIVERS AND STAPLES WITHIN IT. THERE WAS NO PATIENT INVOLVEMENT.
01-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | GST GOLD RELOAD, 60MM, 6 ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHEN THEY OPENED THE RELOAD IT LOOKED BENT AND HAD LOOSE DRIVERS AND STAPLES WITHIN IT. THERE WAS NO PATIENT INVOLVEMENT.
07-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE MISFIRED AND DIDN'T STAPLE THE APPENDIX. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
27-12-2022 | Injury | covidien lp llc north haven | 27-01-2023 | TRI-STAPLE 2.0 | Material Rupture | A0412 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, A DAY FOLLOWING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, WHEREIN THE DEVICE WAS USED FOR FEE (FUNCTIONAL END-END) ANASTOMOSIS FOR CLOSURE OF THE INSERTION HOLE, THE EDGE OF THE REINFORCEMENT MATERIAL WAS OPENED LIKE IT WAS TORN APART FROM A MAJOR LEAK. THIS LED TO REOPERATION WITH REANASTOMOSIS.
27-12-2022 | Injury | covidien lp llc north haven | 27-01-2023 | TRI-STAPLE 2.0 | Material Rupture | A0412 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, A DAY FOLLOWING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, WHEREIN THE DEVICE WAS USED FOR FEE (FUNCTIONAL END-END) ANASTOMOSIS FOR CLOSURE OF THE INSERTION HOLE, THE EDGE OF THE REINFORCEMENT MATERIAL WAS OPENED LIKE IT WAS TORN APART FROM A MAJOR LEAK. THIS LED TO REOPERATION WITH REANASTOMOSIS.
27-12-2022 | Injury | covidien lp llc north haven | 27-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, A DAY FOLLOWING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, WHEREIN THE DEVICE WAS USED FOR FEE (FUNCTIONAL END-END) ANASTOMOSIS FOR CLOSURE OF THE INSERTION HOLE, THE EDGE OF THE REINFORCEMENT MATERIAL WAS OPENED LIKE IT WAS TORN APART FROM A MAJOR LEAK. THIS LED TO REOPERATION WITH REANASTOMOSIS.
27-12-2022 | Injury | covidien lp llc north haven | 27-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, A DAY FOLLOWING A LAPAROSCOPIC ASSISTED DISTAL GASTRECTOMY PROCEDURE, WHEREIN THE DEVICE WAS USED FOR FEE (FUNCTIONAL END-END) ANASTOMOSIS FOR CLOSURE OF THE INSERTION HOLE, THE EDGE OF THE REINFORCEMENT MATERIAL WAS OPENED LIKE IT WAS TORN APART FROM A MAJOR LEAK. THIS LED TO REOPERATION WITH REANASTOMOSIS.
06-12-2022 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE STOMACH, THE 45MM RELOAD WAS DIFFICULT TO FIRE AND HAD POOR STAPLE FORMATION. THE 60MM RELOAD HAD AN UNKNOWN FAILURE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2022 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE AT THE STOMACH, THE 45MM RELOAD WAS DIFFICULT TO FIRE AND HAD POOR STAPLE FORMATION. THE 60MM RELOAD HAD AN UNKNOWN FAILURE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST-OP TO THE LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING, NAUSEOUS AND DIZZINESS. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWEL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION. NO OTHER INFORMATION COULD BE PROVIDED AT THIS TIME.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: IT WAS REPORTED THAT POST-OP TO THE LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING, NAUSEOUS AND DIZZINESS. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWEL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION. NO OTHER INFORMATION COULD BE PROVIDED AT THIS TIME.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Dizziness | Event Description: IT WAS REPORTED THAT POST-OP TO THE LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING, NAUSEOUS AND DIZZINESS. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWEL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION. NO OTHER INFORMATION COULD BE PROVIDED AT THIS TIME.
01-01-2022 | Malfunction | ethicon endo-surgery | 27-01-2023 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNKNOWN PROCEDURE THE DEVICE WOULD NOT FIRE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST-OP TO THE LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING, NAUSEOUS AND DIZZINESS. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWEL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION. NO OTHER INFORMATION COULD BE PROVIDED AT THIS TIME.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: IT WAS REPORTED THAT POST-OP TO THE LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING, NAUSEOUS AND DIZZINESS. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWEL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION. NO OTHER INFORMATION COULD BE PROVIDED AT THIS TIME.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Dizziness | Event Description: IT WAS REPORTED THAT POST-OP TO THE LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING, NAUSEOUS AND DIZZINESS. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWEL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION. NO OTHER INFORMATION COULD BE PROVIDED AT THIS TIME.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST-OP TO THE LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING, NAUSEOUS AND DIZZINESS. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWEL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION. NO OTHER INFORMATION COULD BE PROVIDED AT THIS TIME.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: IT WAS REPORTED THAT POST-OP TO THE LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING, NAUSEOUS AND DIZZINESS. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWEL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION. NO OTHER INFORMATION COULD BE PROVIDED AT THIS TIME.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Dizziness | Event Description: IT WAS REPORTED THAT POST-OP TO THE LAPAROSCOPIC GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING, NAUSEOUS AND DIZZINESS. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWEL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION. NO OTHER INFORMATION COULD BE PROVIDED AT THIS TIME.
10-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE RELOAD WAS CHANGED AND THE PROBLEM WAS SOLVED. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE RELOAD WAS CHANGED AND THE PROBLEM WAS SOLVED. THERE WAS NO PATIENT INJURY.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST-OP TO A LAP GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THERE WAS DELAYED BLEEDING WHERE PATIENT HAD NAUSEOUS AND DIZZINESS.THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: IT WAS REPORTED THAT POST-OP TO A LAP GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THERE WAS DELAYED BLEEDING WHERE PATIENT HAD NAUSEOUS AND DIZZINESS.THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION.
01-03-2022 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Dizziness | Event Description: IT WAS REPORTED THAT POST-OP TO A LAP GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THERE WAS DELAYED BLEEDING WHERE PATIENT HAD NAUSEOUS AND DIZZINESS.THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWL MOVEMENTS UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION.
01-01-2023 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST-OP TO A LAP GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THERE WAS DELAYED BLEEDING WHERE PATIENT HAD NAUSEOUS AND DIZZINESS. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWL MOVEMENTS. UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION.
01-01-2023 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Nausea | Event Description: IT WAS REPORTED THAT POST-OP TO A LAP GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THERE WAS DELAYED BLEEDING WHERE PATIENT HAD NAUSEOUS AND DIZZINESS. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWL MOVEMENTS. UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION.
01-01-2023 | Injury | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Dizziness | Event Description: IT WAS REPORTED THAT POST-OP TO A LAP GASTRIC BYPASS PROCEDURE THAT PATIENT RETURNED WITH DELAYED BLEEDING. THE SURGEON BELIEVED THAT THREE TIMES IT CAME FROM THE GASTRIC REMNANT. THERE WAS DELAYED BLEEDING WHERE PATIENT HAD NAUSEOUS AND DIZZINESS. THE SURGEON REFERRED TO IT AS INTRA LUMINAL BLEEDING THAT EVENTUALLY PASSED A BLOOD CLOT THROUGH HIS BOWL MOVEMENTS. UNKNOWN HOW IT WAS TREATED. THE PATIENT REQUIRED A BLOOD TRANSFUSION.
02-12-2022 | Malfunction | ethicon endo-surgery | 27-01-2023 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE WHEN SEVERING LUNG TISSUE ON THE 3TH FIRING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
04-11-2022 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ESOPHAGEAL RESECTION, IN RECONSTRUCTION, THE PROTRUDING PART ON THE YELLOW TAB OF THE RELOAD THAT PREVENTS PRE-FIRE WAS DAMAGED. THE BROKEN PIECE FELL IN THE ABDOMINAL CAVITY WHICH WAS THEN COLLECTED. THE CASE WAS CONTINUED AS IT IS. THERE WAS NO PATIENT INJURY.
04-11-2022 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOTIC ESOPHAGEAL RESECTION, IN RECONSTRUCTION, THE PROTRUDING PART ON THE YELLOW TAB OF THE RELOAD THAT PREVENTS PRE-FIRE WAS DAMAGED. THE BROKEN PIECE FELL IN THE ABDOMINAL CAVITY WHICH WAS THEN COLLECTED. THE CASE WAS CONTINUED AS IT IS. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE THAT ON THE LAST FIRING WHILE FIRING ACROSS THE STOMACH ONLY ONE SIDE OF STAPLES WERE DEPLOYED. THE SPECIMEN SIDE DID NOT DEPLOY. PATIENT SIDE STAPLE LINE WAS INTACT WITH NO ISSUES. ANOTHER DEVICE WAS NOT NEEDED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
19-12-2022 | Malfunction | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ROBOTIC CYSTECTOMY ON THE SIXTH FIRE THEY HAD A LOCK OUT. THE MANUAL OVERRIDE WAS USED TO REMOVE THE DEVICE AND A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
19-12-2022 | Malfunction | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ROBOTIC CYSTECTOMY ON THE SIXTH FIRE THEY HAD A LOCK OUT. THE MANUAL OVERRIDE WAS USED TO REMOVE THE DEVICE AND A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
19-12-2022 | Malfunction | intuitive surgical | 27-01-2023 | SUREFORM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ON (B)(6) 2023, INTUITIVE SURGICAL, INC. (ISI) RECEIVED USER FACILITY REPORT (B)(4) STATING: DOCTOR WAS USING THE SUREFORM STAPLER 60 INSTRUMENT ON THE STOMACH RESECTION. DURING THREE SEPARATE TIMES, THE STAPLER FAILED TO REACH THE COMPLETE DISTANCE THAT WAS REQUIRED TO COMPLETE THE STAPLER LINE. NEW STAPLE LOADS WERE OPENED EACH TIME, AND EVENTUALLY A NEW STAPLER. ISI FOLLOWED UP WITH THE ISI CLINICAL SALES REPRESENTATIVE (CSR) AND OBTAINED THE FOLLOWING INFORMATION: THE SUREFORM 60 BLUE RELOADS WERE THE RELOADS INVOLVED WITH THE REPORTED EVENT. THE RELOAD WAS SELECTED BECAUSE IT IS THE RELOAD THE SURGEON TYPICALLY USES FOR 1 RELOAD ON SLEEVES. THE STAPLER FIRE 1-3 WERE THE ONES INVOLVED WITH THE REPORTED EVENT. A STOMACH TRANSACTION WAS THE SURGICAL TASK BEING PERFORMED AT THE TIME OF THE EVENT. THE STOMACH WAS THE TARGET TISSUE. THE TARGET TISSUE WAS NOT CALCIFIED AND WAS NOT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY. THERE WAS NO TISSUE TENSION OR BUNCHING. THE EVENT DID NOT INVOLVE A FAILURE TO FIRE. THE STAPLER SEEMED TO STOP IN THE SAME AREA THREE SEPARATE TIMES, WITH THE TISSUE BEING TOO THICK TO CONTINUE. THERE WAS NO TISSUE BEING PUSHED FORWARD/DRAGGED DURING THE FIRING PROCESS. THE JAWS DID NOT OPEN/RELEASE DURING THE FIRING SEQUENCE. THE RELOAD STAPLES DID NOT DEPLOY UNEXPECTEDLY. THERE WERE NO MALFORMED STAPLES. THE RELOAD DID HAVE AN EXPOSED BLADE. THERE WERE NO STAPLES MISSING FROM THE STAPLE LINE. THERE WERE NO HOLES/GAPS IN THE STAPLE LINE. A "TISSUE TOO THICK" ERROR MESSAGE OCCURRED. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS BETWEEN THE INSTRUMENT JAWS. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. NO BUTTRESS MATERIAL WAS USED. NO BLEEDING ON THE STAPLE LINE WAS OBSERVED. THERE WAS NO UNPLANNED TISSUE REMOVAL DUE TO THE FIRING FAILURE. THE SURGEON RESOLVED THE ISSUE BY OPENING A NEW STAPLER. THE PROCEDURE WAS COMPLETED ROBOTICALLY AS PLANNED WITH NO REPORTS OF PATIENT INJURY. THE RELOAD IS NOT AVAILABLE FOR RETURN.
12-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO PROCEDURE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
21-12-2022 | Malfunction | intuitive surgical | 27-01-2023 | SUREFORM | No Device Output | A0907 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED CARDIAC SURGICAL PROCEDURE, THE SUREFORM STAPLER HAD CLAMPING ISSUES. THE CUSTOMER ALSO INDICATED THERE WERE MALFORMED STAPLES. THE PROCEDURE COMPLETED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION: THESE ARE 3 DIFFERENT STAPLERS WHICH HAD THE SAME MATERIAL AND BATCH/LOT # AND THE ISSUE WAS THE EXACT SAME FOR ALL 3 STAPLERS. IT WAS INADEQUATE CLAMPING WITH MALFORMED STAPLES NOTED ON ALL 3. THE PROCEDURE WAS COMPLETED WITH A BACKUP DV STAPLER. THE CUSTOMER CLARIFIED THAT THE FIELD ABORTED/NO PATIENT INVOLVED WAS ENTERED INCORRECTLY AND THAT THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY. THE PROCEDURE FOR ALL WAS WEDGE RESECTION. NO PATIENT DEMOGRAPHIC AVAILABLE. THIS IS ALL THE INFORMATION THEY GIVE US AND NOTHING MORE WAS PROVIDED SO I AM UNABLE TO ANSWER ANY ADDITIONAL QUESTIONS ON THESE.
11-01-2023 | Injury | celestica de monterrey s.a de c.v | 27-01-2023 | ENDO GIA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A BYPASS PROCEDURE, IN FIRING, THE DEVICE FIRED AND CUT THE TISSUE, BUT THERE WAS NO STAPLE LINE. TO RESOLVE THE ISSUE, A NEW RELOAD FROM ANOTHER BATCH WAS USED.
14-07-2022 | Malfunction | davis & geck caribe | 27-01-2023 | UNKNOWN ENDO STITCH SULU | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL-ILEAL BYPASS, WHEN SUTURING PART OF THE NEEDLE FELL INTO THE PATIENT CAVITY. THE PIECE WAS RETRIEVED USING GRASPERS. THERE WAS NO PATIENT INJURY.
14-07-2022 | Malfunction | davis & geck caribe | 27-01-2023 | UNKNOWN ENDO STITCH SULU | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL-ILEAL BYPASS, WHEN SUTURING PART OF THE NEEDLE FELL INTO THE PATIENT CAVITY. THE PIECE WAS RETRIEVED USING GRASPERS. THERE WAS NO PATIENT INJURY.
14-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2023 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, THE HANDLE TURNED OFF WHILE IT WAS INSERTED IN THE SHELL. IN ADDITION, SOMETIMES THE STEM WAS NOT RECOGNIZED, AND IT MUST BE DISASSEMBLED IN ORDER TO BE RECOGNIZED. THE HANDLE WAS RESET TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE THAT THE ONE BATTERY CELL WAS FOUND TO BE VENTED.
14-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2023 | SIGNIA | Calibration Problem | A08 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, THE HANDLE TURNED OFF WHILE IT WAS INSERTED IN THE SHELL. IN ADDITION, SOMETIMES THE STEM WAS NOT RECOGNIZED, AND IT MUST BE DISASSEMBLED IN ORDER TO BE RECOGNIZED. THE HANDLE WAS RESET TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE THAT THE ONE BATTERY CELL WAS FOUND TO BE VENTED.
14-09-2021 | Malfunction | covidien lp llc north haven | 27-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PREPARATION, THE HANDLE TURNED OFF WHILE IT WAS INSERTED IN THE SHELL. IN ADDITION, SOMETIMES THE STEM WAS NOT RECOGNIZED, AND IT MUST BE DISASSEMBLED IN ORDER TO BE RECOGNIZED. THE HANDLE WAS RESET TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE THAT THE ONE BATTERY CELL WAS FOUND TO BE VENTED.
10-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE USAGE OF THE THIRD LOADING UNIT, THE FIRST TWO HANDLES DID NOT FIRE EVEN WHEN THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED CORRECTLY. THE HANDLE WAS REPLACED BUT THE SAME ISSUE OCCURRED. ANOTHER HANDLE WAS USED, BUT IT HAD THE SAME ISSUE. THE SURGEON USED A FIFTH HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | EEA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID/RECTUM END-TO-END ANASTOMOSIS, THE TISSUE ENCAPSULATED IN HOUSING AND DISTAL DONUT WAS THIN AND FRAGILE. THE SURGEON COMPLETED THE PROCEDURE AS NORMAL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | EEA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOID/RECTUM END-TO-END ANASTOMOSIS, THE TISSUE ENCAPSULATED IN HOUSING AND DISTAL DONUT WAS THIN AND FRAGILE. THE SURGEON COMPLETED THE PROCEDURE AS NORMAL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE, ON STOMACH, THE STAPLER WAS PROPERLY LOADED, HOWEVER, WHEN THE SURGEON TRIED TO CLOSE THE STAPLER, THE JAWS OF STAPLER WOULD NOT FULLY CLOSE ON TISSUE OR WOULD REOPEN SLIGHTLY WHEN THE SURGEON TRIED CLOSING THE JAWS. NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE, ON STOMACH, THE STAPLER WAS PROPERLY LOADED, HOWEVER, WHEN THE SURGEON TRIED TO CLOSE THE STAPLER, THE JAWS OF STAPLER WOULD NOT FULLY CLOSE ON TISSUE OR WOULD REOPEN SLIGHTLY WHEN THE SURGEON TRIED CLOSING THE JAWS. NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE, ON STOMACH, THE STAPLER WAS PROPERLY LOADED, HOWEVER, WHEN THE SURGEON TRIED TO CLOSE THE STAPLER, THE JAWS OF STAPLER WOULD NOT FULLY CLOSE ON TISSUE OR WOULD REOPEN SLIGHTLY WHEN THE SURGEON TRIED CLOSING THE JAWS. NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE, ON STOMACH, THE STAPLER WAS PROPERLY LOADED, HOWEVER, WHEN THE SURGEON TRIED TO CLOSE THE STAPLER, THE JAWS OF STAPLER WOULD NOT FULLY CLOSE ON TISSUE OR WOULD REOPEN SLIGHTLY WHEN THE SURGEON TRIED CLOSING THE JAWS. NEW DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2020 | Malfunction | covidien lp llc north haven | 27-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEMONSTRATION DEVICE WOULD NOT TURN ON. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES TO BE VENTED.
04-12-2020 | Malfunction | covidien lp llc north haven | 27-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEMONSTRATION DEVICE WOULD NOT TURN ON. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND BOTH BATTERIES TO BE VENTED.
21-12-2022 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATICODUODENECTOMY PROCEDURE, WHILE CONNECTING THE RELOAD AND THE HANDLE, THE USER NOTICED THAT THE CONDITION OF THE HANDLE WAS NOT GOOD; THE HANDLE COULD NOT BE SQUEEZED. THE USE OF THE HANDLE WAS IMMEDIATELY STOPPED AND A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY , DURING THE USAGE OF THE THIRD LOADING UNIT, THE FIRST TWO HANDLES DID NOT FIRE EVEN WHEN THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED CORRECTLY. THE HANDLE WAS REPLACED BUT THE SAME ISSUE OCCURRED. ANOTHER HANDLE WAS USED, BUT IT HAD THE SAME ISSUE. THE SURGEON USED A FIFTH HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-11-2022 | Malfunction | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS DIFFICULT TO CLOSE. NO PATIENT CONSEQUENCES. USED OF ANOTHER DEVICE.
10-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE USAGE OF THE THIRD LOADING UNIT, THE FIRST TWO HANDLES DID NOT FIRE EVEN WHEN THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED CORRECTLY. THE HANDLE WAS REPLACED BUT THE SAME ISSUE OCCURRED. ANOTHER HANDLE WAS USED, BUT IT HAD THE SAME ISSUE. THE SURGEON USED A FIFTH HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING THE USAGE OF THE THIRD LOADING UNIT, THE FIRST TWO HANDLES DID NOT FIRE EVEN WHEN THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED CORRECTLY. THE HANDLE WAS REPLACED BUT THE SAME ISSUE OCCURRED. ANOTHER HANDLE WAS USED, BUT IT HAD THE SAME ISSUE. THE SURGEON USED A FIFTH HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | celestica de monterrey s.a de c.v | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A BYPASS PROCEDURE, THE SURGEON PRESSED THE GREEN BUTTON WAS AS ABLE TO SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. A NEW RELOAD FROM A DIFFERENT BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-12-2022 | Malfunction | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS DID NOT OPEN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE. AFFILIATE REFERENCE NUMBER: (B)(4).
10-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE PROCEDURE, THERE WERE 5 RELOADS USED, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2022 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SEGMENTECTOMY, THE STAPLER DID NOT FIRE. AFTER REPLACING THE HANDLE AND RELOAD, THE OPERATION WAS SUCCESSFULLY COMPLETED.
08-01-2023 | Malfunction | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. AS THE TISSUE WAS CUT OFF, REMOVED THE DEVICE FROM THE PATIENT THROUGH TISSUE GAP. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
27-12-2022 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SEGMENTECTOMY, THE STAPLER DID NOT FIRE. AFTER REPLACING THE HANDLE AND RELOAD, THE OPERATION WAS SUCCESSFULLY COMPLETED.
10-01-2023 | Malfunction | covidien lp llc north haven | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING GASTRIC SLEEVE PROCEDURE, THERE WERE FIVE RELOADS USED, THE STAPLER WAS UNABLE TO FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE HANDLE WAS NOT SQUEEZE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 27-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. THE RELOAD WAS CHANGED AND THE PROBLEM WAS SOLVED. THERE WAS NO PATIENT INJURY.
04-01-2023 | Injury | us surgical puerto rico | 27-01-2023 | EEA | Failure to Cut | A050702 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, AND DURING THE TRANS-ANAL ANASTOMOSIS, AN INCOMPLETE CUT, LEADS TO THE TEARING OF THE ANASTOMOSED TISSUE AND THE BLOCKAGE OF THE INCUS IN THE PATIENT'S RECTUM. THE ANASTOMOSIS WAS REDONE, AND WHEN USING A NEW CIRCULAR STAPLER FROM THE SAME LOT, AN IMPERFECT CUT WITH INCUS BLOCKAGE AGAIN OCCURRED, BUT THE ANASTOMOSIS WAS NOT IMPAIRED. VERY SIGNIFICANT LENGTHENING OF THE DURATION OF THE INTERVENTION AND RECOURSE TO A SECOND SURGEON TO ENSURE RECOVERY AT THE RECTAL LEVEL.
04-01-2023 | Injury | us surgical puerto rico | 27-01-2023 | EEA | Failure to Cut | A050702 | Obstruction/Occlusion | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, AND DURING THE TRANS-ANAL ANASTOMOSIS, AN INCOMPLETE CUT, LEADS TO THE TEARING OF THE ANASTOMOSED TISSUE AND THE BLOCKAGE OF THE INCUS IN THE PATIENT'S RECTUM. THE ANASTOMOSIS WAS REDONE, AND WHEN USING A NEW CIRCULAR STAPLER FROM THE SAME LOT, AN IMPERFECT CUT WITH INCUS BLOCKAGE AGAIN OCCURRED, BUT THE ANASTOMOSIS WAS NOT IMPAIRED. VERY SIGNIFICANT LENGTHENING OF THE DURATION OF THE INTERVENTION AND RECOURSE TO A SECOND SURGEON TO ENSURE RECOVERY AT THE RECTAL LEVEL.
04-01-2023 | Injury | us surgical puerto rico | 27-01-2023 | EEA | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, AND DURING THE TRANS-ANAL ANASTOMOSIS, AN INCOMPLETE CUT, LEADS TO THE TEARING OF THE ANASTOMOSED TISSUE AND THE BLOCKAGE OF THE INCUS IN THE PATIENT'S RECTUM. THE ANASTOMOSIS WAS REDONE, AND WHEN USING A NEW CIRCULAR STAPLER FROM THE SAME LOT, AN IMPERFECT CUT WITH INCUS BLOCKAGE AGAIN OCCURRED, BUT THE ANASTOMOSIS WAS NOT IMPAIRED. VERY SIGNIFICANT LENGTHENING OF THE DURATION OF THE INTERVENTION AND RECOURSE TO A SECOND SURGEON TO ENSURE RECOVERY AT THE RECTAL LEVEL.
13-09-2022 | Malfunction | ethicon endo-surgery | 27-01-2023 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS ROUX-EN-Y THE JAWS WOULD NOT OPEN WITHIN THE TROCAR. A NEW DEVICE WAS USED TO COMPLETE THE CASE WITH NO PATIENT CONSEQUENCES.
15-11-2022 | Malfunction | intuitive surgical | 27-01-2023 | SUREFORM | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
05-01-2023 | Injury | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE RESECTING THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
05-01-2023 | Injury | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE RESECTING THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
05-01-2023 | Injury | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Component or Accessory Incompatibility | A1701 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE RESECTING THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
05-01-2023 | Injury | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE RESECTING THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
05-01-2023 | Injury | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PROCEDURE, WHILE RESECTING THE RIGHT PULMONARY ARTERY, THE FIRING STOPPED MIDWAY AND THE DEVICE LOCKED ON TISSUE; THE BLACK RETURN KNOB CANNOT BE PULLED TO THE END. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO LOAD AND UNLOAD. THE AREA AROUND THE BLOOD VESSEL WAS DETACHED AS MUCH AS POSSIBLE USING A PROXIMAL LIGATION AND SMALL DIAMETER RELOAD. ADDITIONAL RESECTION WAS PERFORMED TO RESOLVE THE ISSUE. THIS LED TO LESS THAN 1-INCH EXTENDED INCISION, TISSUE DEFECT, AND 30 MINUTE OR MORE EXTENDED SURGICAL TIME.
02-12-2022 | Malfunction | ethicon endo-surgery | 30-01-2023 | ECHELON FLEX | Mechanical Jam | A0506 | Insufficient Information | Event Description: ECHELON ENDO-GIA 60 WAS USED ON PT, FIRED ONE WHITE LOAD, WAS REMOVED FROM PT, AND THEN WAS JAMMED AND WOULD NOT OPEN TO REMOVE USED STAPLE LOAD. IT WAS REMOVED FROM FIELD AND REPLACEMENT OPENED TO COMPLETE CASE.
15-11-2022 | Malfunction | ethicon endo-surgery | 30-01-2023 | ECHELON | Defective Device | A0203 | Insufficient Information | Event Description: DEVICE IS SUPPOSED TO STAPLE AND CUT SIMULTANEOUSLY, AND THIS DEVICE FAILED BOTH ACTIONS. ECHELON CONTOUR CUTTER DID NOT CUT OR WORK AS INTENDED AND IT FAILED.
02-12-2022 | Malfunction | ethicon endo-surgery | 30-01-2023 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Manufacturing, Packaging or Shipping Problem | A02 | No Clinical Signs, Symptoms or Conditions | Event Description: PROCEDURE: IT WAS REPORTED THAT DURING A LUNG PROCEDURE. THE DEVICE COULD NOT BE USED DUE TO ARTICULATING JAWS BEING BROKEN. DEVICE WAS NOT USED IN PATIENT, BECAUSE IT WAS BROKEN WHEN TAKEN OUT OF PACKAGING.
02-12-2022 | Malfunction | ethicon endo-surgery | 30-01-2023 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE DID NOT WORK. THERE WERE NO PATIENT CONSEQUENCES.
15-09-2020 | Malfunction | covidien lp llc north haven | 30-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, DURING THE COLON STAPLING, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
13-01-2023 | Injury | celestica de monterrey s.a de c.v | 30-01-2023 | ENDO GIA | Material Separation | A0413 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PULMONARY SEGMENTAL RESECTION, ON LUNG PARENCHYMA, THE STAPLES BURST OUT AND THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. ADDED SUTURING TO FIX THE STAPLE LINE AND THE DEVICE WAS REPLACED TO COMPLETE THE PROCEDURE.
13-01-2023 | Injury | celestica de monterrey s.a de c.v | 30-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PULMONARY SEGMENTAL RESECTION, ON LUNG PARENCHYMA, THE STAPLES BURST OUT AND THERE WAS AN INCOMPLETE STAPLE LINE DISTALLY. ADDED SUTURING TO FIX THE STAPLE LINE AND THE DEVICE WAS REPLACED TO COMPLETE THE PROCEDURE.
13-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE THIRD FIRING OF THE STOMACH, THE STAPLER FIRED, HOWEVER, THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE AND THE HANDLE LOCKED. EVENTUALLY ABLE TO WIGGLE IT AND WITH A LITTLE FORCE THE JAWS WERE ABLE TO OPEN BUT THE HANDLE STILL LOCKED IN HALF FIRE POSITION. THE HANDLE WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE THIRD FIRING OF THE STOMACH, THE STAPLER FIRED, HOWEVER, THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE AND THE HANDLE LOCKED. EVENTUALLY ABLE TO WIGGLE IT AND WITH A LITTLE FORCE THE JAWS WERE ABLE TO OPEN BUT THE HANDLE STILL LOCKED IN HALF FIRE POSITION. THE HANDLE WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE THIRD FIRING OF THE STOMACH, THE STAPLER FULLY FIRED, HOWEVER, THE JAWS OF THE DEVICE WERE LOCKED ON TISSUE AND HANDLE LOCKED. EVENTUALLY ABLE TO WIGGLE IT AND WITH A LITTLE FORCE THE JAWS WERE ABLE TO OPEN BUT THE HANDLE STILL LOCKED IN HALF FIRE POSITION. THE HANDLE WAS REPLACED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | ethicon endo-surgery | 30-01-2023 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT HEMICOLECTOMY, AFTER FIRED, NOTED THE STAPLES MALFORMED. USED SUTURE TO OVER SEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-01-2023 | Malfunction | ethicon endo-surgery | 30-01-2023 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-01-2023 | Malfunction | ethicon endo-surgery | 30-01-2023 | PWRD 23MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE ON THE CERVICAL ESOPHAGEAL CANCER, THE ANVIL CAME OFF FREQUENTLY WHEN TIGHTENING THE ANVIL. WHEN THE ANVIL CAME OFF, THE SPRING EXPANDED. THE DEVICE WAS USED BETWEEN FREE JEJUNUM AND ESOPHAGEAL. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT.
10-01-2023 | Malfunction | ethicon endo-surgery | 30-01-2023 | ENDO ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION, THE DEVICE COULD NOT FIRE. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-01-2023 | Injury | ethicon endo-surgery | 30-01-2023 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLON DESCENT PROCEDURE THE DOCTOR REPORTS THAT THE DEVICE MADE THE CUT BUT DID NOT STAPLE, THAT IS, IT DID NOT CLOSE. THE COMPLICATION WERE DEHISCENCE AND ILEOSTOMY. THE POSTOPERATIVE PROCESS IS THE PATIENT MUST ATTEND THE HOSPITAL EVERY 4 DAYS TO RECEIVE WOUND CARE. IT SHOULD BE NOTED THAT THE PATIENT IS STABLE WITHIN HIS CONDITION.
10-01-2023 | Injury | ethicon endo-surgery | 30-01-2023 | 75MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLON DESCENT PROCEDURE THE DOCTOR REPORTS THAT THE DEVICE MADE THE CUT BUT DID NOT STAPLE, THAT IS, IT DID NOT CLOSE. THE COMPLICATION WERE DEHISCENCE AND ILEOSTOMY. THE POSTOPERATIVE PROCESS IS THE PATIENT MUST ATTEND THE HOSPITAL EVERY 4 DAYS TO RECEIVE WOUND CARE. IT SHOULD BE NOTED THAT THE PATIENT IS STABLE WITHIN HIS CONDITION.
11-01-2023 | Malfunction | ethicon endo-surgery | 30-01-2023 | PROXIMATE*75MM LINEAR CUTTER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ABDOMINAL ILECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-09-2022 | Malfunction | ethicon endo-surgery | 30-01-2023 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-11-2022 | Injury | intuitive surgical | 30-01-2023 | SUREFORM | Insufficient Information | E2401 | Abdominal Pain | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED SLEEVE GASTRECTOMY, THE PATIENT HAD A ¿LEAK.¿ THROUGH FOLLOW-UP WITH THE SURGEON, IT WAS LEARNED THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH ABDOMINAL PAIN FIVE WEEKS STATUS POST SLEEVE GASTRECTOMY. AFTER A CT SCAN SHOWED A LEAK AND ABSCESS AT THE GASTRIC STAPLE LINE, THE PATIENT WAS REFERRED TO A SPECIALIST AND UNDERWENT AN ENDOSCOPY TO HAVE THE ABSCESS DRAINED. DURING THE ENDOSCOPY, THE SPECIALIST ALSO REMOVED "ABOUT 5 LOOSE STAPLES." THE PATIENT IS ON A CLEAR LIQUID DIET AND TOTAL PARENTERAL NUTRITION. THE SPECIALIST PLANS TO SUTURE THE STAPLE LINE VIA ENDOSCOPE. BECAUSE THERE WERE NO ISSUES INTRAOPERATIVELY AND THE LENGTH OF TIME POST-OP THIS OCCURRED, THE SURGEON BELIEVES THE STAPLE LINE LEAK IS A RESULT OF A VASCULAR ISSUE AND NOT THE SUREFORM 60 STAPLER INSTRUMENT ¿ THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE SURGEON SAID BECAUSE THE LEAK WAS AT THE TOP OF THE STOMACH, A BLUE STAPLER 60 RELOAD, LIKELY THE LAST FIRE OF THE PROCEDURE, WAS USED AT THE AREA OF THE LEAK.
15-11-2022 | Injury | intuitive surgical | 30-01-2023 | SUREFORM | Insufficient Information | E2401 | Abscess | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED SLEEVE GASTRECTOMY, THE PATIENT HAD A ¿LEAK.¿ THROUGH FOLLOW-UP WITH THE SURGEON, IT WAS LEARNED THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH ABDOMINAL PAIN FIVE WEEKS STATUS POST SLEEVE GASTRECTOMY. AFTER A CT SCAN SHOWED A LEAK AND ABSCESS AT THE GASTRIC STAPLE LINE, THE PATIENT WAS REFERRED TO A SPECIALIST AND UNDERWENT AN ENDOSCOPY TO HAVE THE ABSCESS DRAINED. DURING THE ENDOSCOPY, THE SPECIALIST ALSO REMOVED "ABOUT 5 LOOSE STAPLES." THE PATIENT IS ON A CLEAR LIQUID DIET AND TOTAL PARENTERAL NUTRITION. THE SPECIALIST PLANS TO SUTURE THE STAPLE LINE VIA ENDOSCOPE. BECAUSE THERE WERE NO ISSUES INTRAOPERATIVELY AND THE LENGTH OF TIME POST-OP THIS OCCURRED, THE SURGEON BELIEVES THE STAPLE LINE LEAK IS A RESULT OF A VASCULAR ISSUE AND NOT THE SUREFORM 60 STAPLER INSTRUMENT ¿ THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE SURGEON SAID BECAUSE THE LEAK WAS AT THE TOP OF THE STOMACH, A BLUE STAPLER 60 RELOAD, LIKELY THE LAST FIRE OF THE PROCEDURE, WAS USED AT THE AREA OF THE LEAK.
15-11-2022 | Injury | intuitive surgical | 30-01-2023 | SUREFORM | Insufficient Information | E2401 | Feeding Problem | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED SLEEVE GASTRECTOMY, THE PATIENT HAD A ¿LEAK.¿ THROUGH FOLLOW-UP WITH THE SURGEON, IT WAS LEARNED THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH ABDOMINAL PAIN FIVE WEEKS STATUS POST SLEEVE GASTRECTOMY. AFTER A CT SCAN SHOWED A LEAK AND ABSCESS AT THE GASTRIC STAPLE LINE, THE PATIENT WAS REFERRED TO A SPECIALIST AND UNDERWENT AN ENDOSCOPY TO HAVE THE ABSCESS DRAINED. DURING THE ENDOSCOPY, THE SPECIALIST ALSO REMOVED "ABOUT 5 LOOSE STAPLES." THE PATIENT IS ON A CLEAR LIQUID DIET AND TOTAL PARENTERAL NUTRITION. THE SPECIALIST PLANS TO SUTURE THE STAPLE LINE VIA ENDOSCOPE. BECAUSE THERE WERE NO ISSUES INTRAOPERATIVELY AND THE LENGTH OF TIME POST-OP THIS OCCURRED, THE SURGEON BELIEVES THE STAPLE LINE LEAK IS A RESULT OF A VASCULAR ISSUE AND NOT THE SUREFORM 60 STAPLER INSTRUMENT ¿ THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE SURGEON SAID BECAUSE THE LEAK WAS AT THE TOP OF THE STOMACH, A BLUE STAPLER 60 RELOAD, LIKELY THE LAST FIRE OF THE PROCEDURE, WAS USED AT THE AREA OF THE LEAK.
12-11-2021 | Death | covidien lp llc north haven | 30-01-2023 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY EVALUATED 21 PATIENTS WHO UNDERWENT AN ANASTOMOTIC PROCEDURE BETWEEN 2016 AND 2017. THERE WERE POSTOPERATIVE COMPLICATIONS OF ANASTOMOTIC FAILURE IN ONE PATIENT USING A LINEAR STAPLER WHICH LEAD TO DEATH. THE MORTALITY RATE FOR ESOPHAGOGASTRO AND ESOPHAGOENTERO ANASTOMOSES WAS 1 CASE.
12-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | EEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, ON LAPAROSCOPIC SIGMOID COLON RESECTION, WHEN REMOVING THE DEVICE AFTER THE ANASTOMOSIS WAS DONE, THE ANVIL SEPARATED FROM THE STAPLER AND REMAINED IN THE PATIENT¿S CAVITY. IT WAS NECESSARY TO INSERT A CLAMP AND REMOVE THE ANVIL SEPARATELY. THE ANASTOMOSIS WORKED WELL WITHOUT LEAKS. THERE WAS NO PATIENT INJURY.
13-01-2022 | Injury | davis & geck caribe | 30-01-2023 | UNKNOWN ENDO STITCH SULU | Detachment of Device or Device Component | A0501 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HYSTERECTOMY, WHEN THE SURGEON WAS SEWING THE CUFF, THE NEEDLE AND SUTURE COMPLETELY BROKE OFF THE HANDLE. THE NEEDED AND THE SUTURE FELL INTO THE PATIENT AND WERE NOT ABLE TO BE RETRIEVED. X-RAY RESULT STATED THAT THE NEEDLE WOULD BE TOO SMALL TO IDENTIFY THROUGH. AS A RESULT, THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE TRYING TO FIND THE NEEDLE AND SUTURE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
13-01-2022 | Injury | davis & geck caribe | 30-01-2023 | UNKNOWN ENDO STITCH SULU | Detachment of Device or Device Component | A0501 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HYSTERECTOMY, WHEN THE SURGEON WAS SEWING THE CUFF, THE NEEDLE AND SUTURE COMPLETELY BROKE OFF THE HANDLE. THE NEEDED AND THE SUTURE FELL INTO THE PATIENT AND WERE NOT ABLE TO BE RETRIEVED. X-RAY RESULT STATED THAT THE NEEDLE WOULD BE TOO SMALL TO IDENTIFY THROUGH. AS A RESULT, THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE TRYING TO FIND THE NEEDLE AND SUTURE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
13-01-2022 | Injury | davis & geck caribe | 30-01-2023 | UNKNOWN ENDO STITCH SULU | Adverse Event Without Identified Device or Use Problem | A24 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HYSTERECTOMY, WHEN THE SURGEON WAS SEWING THE CUFF, THE NEEDLE AND SUTURE COMPLETELY BROKE OFF THE HANDLE. THE NEEDED AND THE SUTURE FELL INTO THE PATIENT AND WERE NOT ABLE TO BE RETRIEVED. X-RAY RESULT STATED THAT THE NEEDLE WOULD BE TOO SMALL TO IDENTIFY THROUGH. AS A RESULT, THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE TRYING TO FIND THE NEEDLE AND SUTURE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
13-01-2022 | Injury | davis & geck caribe | 30-01-2023 | UNKNOWN ENDO STITCH SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HYSTERECTOMY, WHEN THE SURGEON WAS SEWING THE CUFF, THE NEEDLE AND SUTURE COMPLETELY BROKE OFF THE HANDLE. THE NEEDED AND THE SUTURE FELL INTO THE PATIENT AND WERE NOT ABLE TO BE RETRIEVED. X-RAY RESULT STATED THAT THE NEEDLE WOULD BE TOO SMALL TO IDENTIFY THROUGH. AS A RESULT, THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE TRYING TO FIND THE NEEDLE AND SUTURE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
11-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC RESECTION, WHEN THE CARTRIDGE WAS ATTACHED TO THE HANDLE AND THE TISSUE WAS CLAMPED, THE GREEN BUTTON DID NOT LIGHT UP AND THE FIRING COULD NOT BE PERFORMED. THE CARTRIDGE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZE. THE DEVICE DID NOT FIRE. ADDITIONALLY, THE HANDLE OF THE DEVICE BROKE WHEN THE RELOAD WAS USED. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZE. THE DEVICE DID NOT FIRE. ADDITIONALLY, THE HANDLE OF THE DEVICE BROKE WHEN THE RELOAD WAS USED. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, AFTER USING THE HANDLE SEVERAL TIMES AT THE TIME OF PULMONARY PARENCHYMA RESECTION, A NEW RELOAD WAS ATTACHED. WHEN FIRING WAS ATTEMPTED, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN WHEN THE GREEN BUTTON WAS PRESSED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Injury | ethicon endo-surgery | 30-01-2023 | 75MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER AN OPEN PROCEDURE ON THE DESCENDING COLON, LEAKAGE OCCURRED. THE PATIENT FEVERED AND THE LEAKAGE WAS FOUND. THE DRAINAGE WAS PLACED, FOLLOW UP, AND DISCHARGED FROM THE HOSPITAL. AT THE INITIAL OPERATION, THE DEVICE WAS USED ON FEE, AND THE STAPLE HEIGHT OF 1ST FIRING ON SIDE-TO-SIDE ANASTOMOSIS WAS GOLD, 2ND FIRING ON CLOSURE WAS GREEN. THE DEVICE WAS USED ON THE HILUM LIENIS, AND IT WAS A DIFFICULT CASE. THE BIFURCATION PART, CLOSED END, AND THE OVERLAPPING STAPLES WERE REINFORCED WITH THE SUTURE. NO FURTHER INFORMATION IS AVAILABLE.
11-01-2023 | Injury | ethicon endo-surgery | 30-01-2023 | 75MM SELECTABLE NEW TLC | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: IT WAS REPORTED THAT AFTER AN OPEN PROCEDURE ON THE DESCENDING COLON, LEAKAGE OCCURRED. THE PATIENT FEVERED AND THE LEAKAGE WAS FOUND. THE DRAINAGE WAS PLACED, FOLLOW UP, AND DISCHARGED FROM THE HOSPITAL. AT THE INITIAL OPERATION, THE DEVICE WAS USED ON FEE, AND THE STAPLE HEIGHT OF 1ST FIRING ON SIDE-TO-SIDE ANASTOMOSIS WAS GOLD, 2ND FIRING ON CLOSURE WAS GREEN. THE DEVICE WAS USED ON THE HILUM LIENIS, AND IT WAS A DIFFICULT CASE. THE BIFURCATION PART, CLOSED END, AND THE OVERLAPPING STAPLES WERE REINFORCED WITH THE SUTURE. NO FURTHER INFORMATION IS AVAILABLE.
06-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, AFTER USING THE HANDLE SEVERAL TIMES AT THE TIME OF PULMONARY PARENCHYMA RESECTION, A NEW RELOAD WAS ATTACHED. WHEN FIRING WAS ATTEMPTED, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE EVEN WHEN THE GREEN BUTTON WAS PRESSED. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LAPAROTOMY RECTAL RESECTION, WHEN APPROACHING THE ARTICULATION IN THE RECTUM, THE JAW WAS CLOSED BUT IT DID NOT CLOSE PROPERLY. THE CONNECTION BETWEEN THE HANDLE AND THE RELOAD WAS BAD. THE JAW WAS OPENED AND REPLACED WITH A NEW DEVICE. THE NEW HANDLE AND RELOAD WORKED WITHOUT A PROBLEM WITH THE SAME TISSUE THICKNESS. THE TISSUE WAS NEITHER HARD NOR THICK. THERE WAS NO PATIENT INJURY.
02-11-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE CHARGER SHOWED A NORMAL BEHAVIOR WHEN THE HANDLE WAS PUT IN. AFTER ABOUT A MINUTE, THE FLASHING GREEN LIGHT TURNED TO PERMANENT GREEN. WHEN THE HANDLE WAS TAKEN OUT OF THE CHARGER, IT WAS NOT ABLE TO START. ANOTHER TYPE OF STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS CELL VENTING.
02-11-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE CHARGER SHOWED A NORMAL BEHAVIOR WHEN THE HANDLE WAS PUT IN. AFTER ABOUT A MINUTE, THE FLASHING GREEN LIGHT TURNED TO PERMANENT GREEN. WHEN THE HANDLE WAS TAKEN OUT OF THE CHARGER, IT WAS NOT ABLE TO START. ANOTHER TYPE OF STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS CELL VENTING.
12-01-2023 | Malfunction | ethicon endo-surgery | 30-01-2023 | ENDO ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE THE DEVICE WOULD NOT FIRE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-01-2023 | Injury | celestica de monterrey s.a de c.v | 30-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY SEGMENTAL RESECTION, WHILE STAPLING THE LUNG PARENCHYMA, THE STAPLES HAD POOR FORMATION AND BURST OUT AND THE STAPLE LINE WAS CENTRALLY INCOMPLETE. THE SURGEON SUTURED THE STAPLE LINE AND REPLACED THE RELOAD TO RESOLVE THE ISSUE.
12-01-2023 | Injury | celestica de monterrey s.a de c.v | 30-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY SEGMENTAL RESECTION, WHILE STAPLING THE LUNG PARENCHYMA, THE STAPLES HAD POOR FORMATION AND BURST OUT AND THE STAPLE LINE WAS CENTRALLY INCOMPLETE. THE SURGEON SUTURED THE STAPLE LINE AND REPLACED THE RELOAD TO RESOLVE THE ISSUE.
12-01-2023 | Injury | celestica de monterrey s.a de c.v | 30-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY SEGMENTAL RESECTION, WHILE STAPLING THE PULMONARY ARTERY, THE STAPLES HAD POOR FORMATION WHICH LED TO BLEEDING AND THE STAPLE LINE WAS CENTRALLY INCOMPLETE. THE SURGEON SUTURED THE STAPLE LINE AND REPLACED THE RELOAD TO RESOLVE THE ISSUE.
12-01-2023 | Injury | celestica de monterrey s.a de c.v | 30-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY SEGMENTAL RESECTION, WHILE STAPLING THE PULMONARY ARTERY, THE STAPLES HAD POOR FORMATION WHICH LED TO BLEEDING AND THE STAPLE LINE WAS CENTRALLY INCOMPLETE. THE SURGEON SUTURED THE STAPLE LINE AND REPLACED THE RELOAD TO RESOLVE THE ISSUE.
12-01-2023 | Injury | celestica de monterrey s.a de c.v | 30-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY SEGMENTAL RESECTION, WHILE STAPLING THE PULMONARY ARTERY, THE STAPLES HAD POOR FORMATION WHICH LED TO BLEEDING AND THE STAPLE LINE WAS CENTRALLY INCOMPLETE. THE SURGEON SUTURED THE STAPLE LINE AND REPLACED THE RELOAD TO RESOLVE THE ISSUE.
12-01-2023 | Injury | celestica de monterrey s.a de c.v | 30-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY SEGMENTAL RESECTION, WHILE STAPLING THE PULMONARY ARTERY, THE STAPLES HAD POOR FORMATION WHICH LED TO BLEEDING AND THE STAPLE LINE WAS CENTRALLY INCOMPLETE. THE SURGEON SUTURED THE STAPLE LINE AND REPLACED THE RELOAD TO RESOLVE THE ISSUE.
31-10-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE HAD INITIAL FILE ERROR, THEN STATED A TRIANGLE WITH AN EXCLAMATION POINT. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
31-10-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE HAD INITIAL FILE ERROR, THEN STATED A TRIANGLE WITH AN EXCLAMATION POINT. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
31-10-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Power Problem | A0708 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE HAD INITIAL FILE ERROR, THEN STATED A TRIANGLE WITH AN EXCLAMATION POINT. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
11-01-2023 | Malfunction | ethicon endo-surgery | 30-01-2023 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE POWERED CIRCULAR STAPLER DID NOT FIRE. THE BATTERY WAS ON, EVERY STEP WAS TAKEN ACCORDINGLY, AND DEVICE DID NOT FIRE. IT WAS END TO END ANASTOMOSES OF THE OESOPHAGUS. THE SURGERY WAS DELAYED 15 MINUTES. MULTIPLE TRIALS WERE DONE, A NEW PRODUCT WAS OPEN TO USE BATTERY, BUT PRODUCT STILL DID NOT FIRE.
11-01-2023 | Malfunction | ethicon endo-surgery | 30-01-2023 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE POWERED CIRCULAR STAPLER DID NOT FIRE. THE BATTERY WAS ON, EVERY STEP WAS TAKEN ACCORDINGLY, AND DEVICE DID NOT FIRE. IT WAS END TO END ANASTOMOSES OF THE OESOPHAGUS. THE SURGERY WAS DELAYED 15 MINUTES. MULTIPLE TRIALS WERE DONE, A NEW PRODUCT WAS OPEN TO USE BATTERY, BUT PRODUCT STILL DID NOT FIRE.
16-12-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
14-11-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE LIGHT-EMITTING DIODE (LED) DISPLAY AT THE CHARGER SHOWED A TRIANGLE ERROR MASSAGE AND FAILED TO CHARGE. THE HANDLE WAS REBOOTED BY PRESSING THE GREEN BUTTONS ON THE SIDE BUT NOTHING HAPPENED FOR IT DID NOT CHARGE AT ALL AND WOULD NOT TURN ON. A PERSONAL, SINGLE-BAY CHARGER WAS USED BUT THE HANDLE SHOWED AN ERROR MESSAGE PERTAINING TO A BATTERY ERROR AND THE CHARGER FLASHED RED LIGHT. A NEW HANDLE WAS REQUESTED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS CELL VENTING.
14-11-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE LIGHT-EMITTING DIODE (LED) DISPLAY AT THE CHARGER SHOWED A TRIANGLE ERROR MASSAGE AND FAILED TO CHARGE. THE HANDLE WAS REBOOTED BY PRESSING THE GREEN BUTTONS ON THE SIDE BUT NOTHING HAPPENED FOR IT DID NOT CHARGE AT ALL AND WOULD NOT TURN ON. A PERSONAL, SINGLE-BAY CHARGER WAS USED BUT THE HANDLE SHOWED AN ERROR MESSAGE PERTAINING TO A BATTERY ERROR AND THE CHARGER FLASHED RED LIGHT. A NEW HANDLE WAS REQUESTED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS CELL VENTING.
16-12-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | UNKNOWN SIGNIA EGIA ADAPTER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW DEVICE OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
04-01-2023 | Injury | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE PHYSICIAN COULD NOT STAPLE UNTIL THE END OF THE RELOAD, IT WAS TOO DIFFICULT. THE JAWS WERE LOCKED ON TISSUE. THE PHYSICIAN HAD TO USE EXCESSIVE FORCE TO RETRACT THE BLACK KNOB AND REMOVE THE DEVICE. AFTER THE PHYSICIAN MANAGED TO OPEN THE JAWS, THE TISSUE HAD BEEN A LITTLE BIT DAMAGED. SOME STAPLES HAVE NOT BEEN FORMED PROPERLY OR AT ALL. THERE WAS ALSO AN INCOMPLETE STAPLE LINE PROXIMALLY OR CENTRALLY. THE PHYSICIAN DECIDED TO PLACE THE NEW STAPLER A LITTLE BIT HIGHER TO STAPLE A LITTLE BIT ABOVE THIS DAMAGED TISSUE PART. THIS TISSUE PART HAD BEEN REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
04-01-2023 | Injury | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Retraction Problem | A0510 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE PHYSICIAN COULD NOT STAPLE UNTIL THE END OF THE RELOAD, IT WAS TOO DIFFICULT. THE JAWS WERE LOCKED ON TISSUE. THE PHYSICIAN HAD TO USE EXCESSIVE FORCE TO RETRACT THE BLACK KNOB AND REMOVE THE DEVICE. AFTER THE PHYSICIAN MANAGED TO OPEN THE JAWS, THE TISSUE HAD BEEN A LITTLE BIT DAMAGED. SOME STAPLES HAVE NOT BEEN FORMED PROPERLY OR AT ALL. THERE WAS ALSO AN INCOMPLETE STAPLE LINE PROXIMALLY OR CENTRALLY. THE PHYSICIAN DECIDED TO PLACE THE NEW STAPLER A LITTLE BIT HIGHER TO STAPLE A LITTLE BIT ABOVE THIS DAMAGED TISSUE PART. THIS TISSUE PART HAD BEEN REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
04-01-2023 | Injury | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE PHYSICIAN COULD NOT STAPLE UNTIL THE END OF THE RELOAD, IT WAS TOO DIFFICULT. THE JAWS WERE LOCKED ON TISSUE. THE PHYSICIAN HAD TO USE EXCESSIVE FORCE TO RETRACT THE BLACK KNOB AND REMOVE THE DEVICE. AFTER THE PHYSICIAN MANAGED TO OPEN THE JAWS, THE TISSUE HAD BEEN A LITTLE BIT DAMAGED. SOME STAPLES HAVE NOT BEEN FORMED PROPERLY OR AT ALL. THERE WAS ALSO AN INCOMPLETE STAPLE LINE PROXIMALLY OR CENTRALLY. THE PHYSICIAN DECIDED TO PLACE THE NEW STAPLER A LITTLE BIT HIGHER TO STAPLE A LITTLE BIT ABOVE THIS DAMAGED TISSUE PART. THIS TISSUE PART HAD BEEN REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
04-01-2023 | Injury | covidien lp llc north haven | 30-01-2023 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE PHYSICIAN COULD NOT STAPLE UNTIL THE END OF THE RELOAD, IT WAS TOO DIFFICULT. THE JAWS WERE LOCKED ON TISSUE. THE PHYSICIAN HAD TO USE EXCESSIVE FORCE TO RETRACT THE BLACK KNOB AND REMOVE THE DEVICE. AFTER THE PHYSICIAN MANAGED TO OPEN THE JAWS, THE TISSUE HAD BEEN A LITTLE BIT DAMAGED. SOME STAPLES HAVE NOT BEEN FORMED PROPERLY OR AT ALL. THERE WAS ALSO AN INCOMPLETE STAPLE LINE PROXIMALLY OR CENTRALLY. THE PHYSICIAN DECIDED TO PLACE THE NEW STAPLER A LITTLE BIT HIGHER TO STAPLE A LITTLE BIT ABOVE THIS DAMAGED TISSUE PART. THIS TISSUE PART HAD BEEN REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
04-01-2023 | Injury | covidien lp llc north haven | 30-01-2023 | UNKNOWN ENDO GIA SULU | Retraction Problem | A0510 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE PHYSICIAN COULD NOT STAPLE UNTIL THE END OF THE RELOAD, IT WAS TOO DIFFICULT. THE JAWS WERE LOCKED ON TISSUE. THE PHYSICIAN HAD TO USE EXCESSIVE FORCE TO RETRACT THE BLACK KNOB AND REMOVE THE DEVICE. AFTER THE PHYSICIAN MANAGED TO OPEN THE JAWS, THE TISSUE HAD BEEN A LITTLE BIT DAMAGED. SOME STAPLES HAVE NOT BEEN FORMED PROPERLY OR AT ALL. THERE WAS ALSO AN INCOMPLETE STAPLE LINE PROXIMALLY OR CENTRALLY. THE PHYSICIAN DECIDED TO PLACE THE NEW STAPLER A LITTLE BIT HIGHER TO STAPLE A LITTLE BIT ABOVE THIS DAMAGED TISSUE PART. THIS TISSUE PART HAD BEEN REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
04-01-2023 | Injury | covidien lp llc north haven | 30-01-2023 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE PHYSICIAN COULD NOT STAPLE UNTIL THE END OF THE RELOAD, IT WAS TOO DIFFICULT. THE JAWS WERE LOCKED ON TISSUE. THE PHYSICIAN HAD TO USE EXCESSIVE FORCE TO RETRACT THE BLACK KNOB AND REMOVE THE DEVICE. AFTER THE PHYSICIAN MANAGED TO OPEN THE JAWS, THE TISSUE HAD BEEN A LITTLE BIT DAMAGED. SOME STAPLES HAVE NOT BEEN FORMED PROPERLY OR AT ALL. THERE WAS ALSO AN INCOMPLETE STAPLE LINE PROXIMALLY OR CENTRALLY. THE PHYSICIAN DECIDED TO PLACE THE NEW STAPLER A LITTLE BIT HIGHER TO STAPLE A LITTLE BIT ABOVE THIS DAMAGED TISSUE PART. THIS TISSUE PART HAD BEEN REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
04-01-2023 | Injury | covidien lp llc north haven | 30-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE STOMACH, THE PHYSICIAN COULD NOT STAPLE UNTIL THE END OF THE RELOAD, IT WAS TOO DIFFICULT. THE JAWS WERE LOCKED ON TISSUE. THE PHYSICIAN HAD TO USE EXCESSIVE FORCE TO RETRACT THE BLACK KNOB AND REMOVE THE DEVICE. AFTER THE PHYSICIAN MANAGED TO OPEN THE JAWS, THE TISSUE HAD BEEN A LITTLE BIT DAMAGED. SOME STAPLES HAVE NOT BEEN FORMED PROPERLY OR AT ALL. THERE WAS ALSO AN INCOMPLETE STAPLE LINE PROXIMALLY OR CENTRALLY. THE PHYSICIAN DECIDED TO PLACE THE NEW STAPLER A LITTLE BIT HIGHER TO STAPLE A LITTLE BIT ABOVE THIS DAMAGED TISSUE PART. THIS TISSUE PART HAD BEEN REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
22-12-2022 | Malfunction | intuitive surgical | 30-01-2023 | SUREFORM | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE CUSTOMER HAD ISSUES REMOVING THE SUREFORM STAPLER 60 INSTRUMENT FROM THE ARM. AFTER STAPLER INSTRUMENT WAS FIRED, IT RELEASED FROM THE TISSUE, BUT MESSAGE APPEARED ON MONITOR TO TURN THE KNOB FOR THE EMERGENCY RELEASE TO RELEASE THE TISSUE. THE CUSTOMER TURNED THE EMERGENCY RELEASE BUT HAD DIFFICULTY REMOVING THE STAPLER FROM THE NEW ARM. THE SUREFORM STAPLER 60 WHITE RELOAD WAS RETURNED WITH THE STAPLER INSTRUMENT. THERE WAS NO ALLEGED MALFUNCTION REPORTED AGAINST THE RELOAD. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY. AN INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER AND FOLLOWING ADDITIONAL INFORMATION WAS OBTAINED ABOUT THE COMPLAINT: THE STAPLER INSTRUMENT WAS INSPECTED PRIOR TO USE AND NO DAMAGE WAS NOTED. WHITE STAPLER RELOAD WAS INVOLVED WITH THE REPORTED EVENT. THE STAPLER INSTRUMENT WORKED THE FIRST TWO TIMES. THE REPORTED ISSUE OCCURRED WHILE STAPLING THE STOMACH TISSUE. THE TISSUE WAS NOT CALCIFIED OR EXPOSED TO ANY RADIATION OR CHEMOTHERAPY PRIOR TO THE PROCEDURE. THERE WAS NO TISSUE TENSION OR TISSUE BUNCHING. THE STAPLER HAD AN ERROR MESSAGE STATING "STAPLER WAS NOT RECOGNIZED", IT WAS REMOVED AND REINSERTED MULTIPLE TIMES, THE STAPLE RELOAD WAS REPLACED BEFORE OPENING A NEW STAPLER. THE STAPLING ISSUE DID NOT RESULT IN ANY MALFORMED OR UNFORMED STAPLES. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. BUTTRESS MATERIAL WAS NOT USED. THE SURGEON DID NOT EXPERIENCE ANY CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD. THERE WAS NO UNPLANNED TISSUE REMOVAL DUE TO THE STAPLER FIRING FAILURE. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER INSTRUMENT. THE PROCEDURE WAS COMPLETED ROBOTICALLY WITH NO REPORTS OF PATIENT INJURY.
14-11-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE TURN OFF AFTER THE SOFTWARE WAS UPLOADED AND WOULD NOT TURN BACK ON. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE). DEVICE BATTERY MANAGEMENT ENHANCEMENTS HAVE BEEN IMPLEMENTED TO EXTEND BATTERY LIFE AND ENHANCE BATTERY HEALTH MONITORING.
14-11-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Application Program Version or Upgrade Problem | A110206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE TURN OFF AFTER THE SOFTWARE WAS UPLOADED AND WOULD NOT TURN BACK ON. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE). DEVICE BATTERY MANAGEMENT ENHANCEMENTS HAVE BEEN IMPLEMENTED TO EXTEND BATTERY LIFE AND ENHANCE BATTERY HEALTH MONITORING.
14-11-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE TURN OFF AFTER THE SOFTWARE WAS UPLOADED AND WOULD NOT TURN BACK ON. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. PLI 10- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE). DEVICE BATTERY MANAGEMENT ENHANCEMENTS HAVE BEEN IMPLEMENTED TO EXTEND BATTERY LIFE AND ENHANCE BATTERY HEALTH MONITORING.
11-01-2023 | Injury | us surgical puerto rico | 30-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE USING THE HANDLE, THE FIRST RELOAD WAS LOADED WITH A PIN WHEN DISENGAGING THE PATIENT¿S BRONCHUS, AND DURING FIRST FIRING, THE STAPLE WAS STUCK AND THE DISTAL END OF THE SUTURE WAS INCOMPLETE. AFTER REPLACING THE RELOAD, IT WAS FOUND THAT THE GRIP WAS SQUEEZED, BUT DEVICE COULD NOT BE FIRED. THE GREEN BUTTON WAS PRESSED. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE CASE.
30-01-2023 | Malfunction | ethicon | 30-01-2023 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNILATERAL INGUINAL HERNIOPLASTY PROCEDURE ON AN UNKNOWN DATE AND ABSORBABLE STRAPS WERE USED. THE HOOKS WERE NOT FIXED BY THE DEVICE AFTER IT WAS TRIGGERED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES.
31-10-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE CHARGER LAMP TURNED RED DURING CHARGING, AND THE DISPLAY SCREEN DID NOT APPEAR (BLACKOUT) WHEN THE HANDLE WAS CHECKED. WHEN THE GREEN BUTTON WAS PRESSED AND HELD DOWN FROM BOTH SIDES FOR 10 SECONDS, IT WAS NOT WORKING AT ALL. THE HANDLE WAS UNABLE TO BE USED SO IT WAS HANDLED WITH A DISPOSABLE HANDLE. THE CHARGER WAS ABLE TO WORK WITH OTHER HANDLE WITH NO PROBLEM. THERE WAS NO PATIENT INVOLVED. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
31-10-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, THE CHARGER LAMP TURNED RED DURING CHARGING, AND THE DISPLAY SCREEN DID NOT APPEAR (BLACKOUT) WHEN THE HANDLE WAS CHECKED. WHEN THE GREEN BUTTON WAS PRESSED AND HELD DOWN FROM BOTH SIDES FOR 10 SECONDS, IT WAS NOT WORKING AT ALL. THE HANDLE WAS UNABLE TO BE USED SO IT WAS HANDLED WITH A DISPOSABLE HANDLE. THE CHARGER WAS ABLE TO WORK WITH OTHER HANDLE WITH NO PROBLEM. THERE WAS NO PATIENT INVOLVED. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ESOPHAGECTOMY, AFTER CLAMPING THE ESOPHAGUS AT THE TIME OF ESOPHAGEAL RESECTION, THE SURGEON PRESSED THE GREEN BUTTON AND SQUEEZED THE HANDLE, BUT THE FIRING DID NOT ADVANCE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. IT WAS ALSO NOTED THAT THE DEVICE LOCKED ON TISSUE, BUT WAS ABLE TO BE REMOVED. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREAS TAIL RESECTION, WHILE AT PANCREAS, THE RESTART INDICATOR APPEARED IN THE MIDDLE OF FIRING, AND FIRING COULD NOT BE PERFORMED TILL THE END. THE INCISION RANGE EXPANDED LESS THAN ONE INCH DUE TO A PROBLEM WITH THIS PRODUCT. TO RESOLVE THE ISSUE AND COMPLETE THE CASE, A MANUAL ADAPTER TOOL WITH THE SAME RELOAD WERE USED TO PERFORM FIRING TILL THE END, AND CENTRAL SIDE OF THE PANCREAS WAS RESECTED WITH THE HANDLE AND RELOAD.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREAS TAIL RESECTION, WHILE AT PANCREAS, THE RESTART INDICATOR APPEARED IN THE MIDDLE OF FIRING, AND FIRING COULD NOT BE PERFORMED TILL THE END. THE INCISION RANGE EXPANDED LESS THAN ONE INCH DUE TO A PROBLEM WITH THIS PRODUCT. TO RESOLVE THE ISSUE AND COMPLETE THE CASE, A MANUAL ADAPTER TOOL WITH THE SAME RELOAD WERE USED TO PERFORM FIRING TILL THE END, AND CENTRAL SIDE OF THE PANCREAS WAS RESECTED WITH THE HANDLE AND RELOAD.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREAS TAIL RESECTION, WHILE AT PANCREAS, THE RESTART INDICATOR APPEARED IN THE M IDDLE OF FIRING, AND FIRING COULD NOT BE PERFORMED TILL THE END. THERE WAS A SLIGHT OOZING FROM THE TIP OF THE STAPLE LINE (THE PART THAT THE FIRING HAD ADVANCED MANUALLY). VISUAL CHECKING WAS PERFORMED AND POOR STAPLE FORMATION WAS FOUND. THE INCISION RANGE EXPAN DED LESS THAN 1 INCH DUE TO A PROBLEM WITH THIS PRODUCT. TO COMPLETE THE CASE AND RESOLVE THE ISSUE, THE MANUAL ADAPTER TOOL WAS USED AND THE RELOAD USED AT FIRST WAS ABLE TO BE FIRED TO THE END. AFTER THAT, THE PROXIMAL SIDE (INNER SIDE) WAS RESECTED WITH MANUAL HANDLE AND RELOAD.
30-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: TWO MANUAL STAPLER HANDLE AND TWO RELOADS WERE RETURNED WITHOUT ANY ACCOMPANYING INFORMATION. PLI20: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
30-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: TWO MANUAL STAPLER HANDLE AND TWO RELOADS WERE RETURNED WITHOUT ANY ACCOMPANYING INFORMATION. PLI20: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
23-12-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBE RESECTION, WHILE RESECTING THE LUNG, THE DEVICE WAS FIRED UNTIL THE END, BUT THE ANCHORING SUTURE AT THE TIP WAS NOT CUT DURING FIRING. THE SURGEON USED FORCEPS TO DIRECTLY CUT THE SUTURE. THERE WAS NO PATIENT INJURY. PLI 10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE RELOAD WAS PARTIALLY FIRED AND THE INTERLOCK WAS ENGAGED.
23-12-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | TRI-STAPLE 2.0 | Separation Failure | A150301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBE RESECTION, WHILE RESECTING THE LUNG, THE DEVICE WAS FIRED UNTIL THE END, BUT THE ANCHORING SUTURE AT THE TIP WAS NOT CUT DURING FIRING. THE SURGEON USED FORCEPS TO DIRECTLY CUT THE SUTURE. THERE WAS NO PATIENT INJURY. PLI 10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE RELOAD WAS PARTIALLY FIRED AND THE INTERLOCK WAS ENGAGED.
04-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE BROKE. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-11-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, BOTH POWER HANDLES WERE NOT TURNING ON OR HOLDING A CHARGE AND THE CHARGING BAY SEEMED TO BE MALFUNCTIONING. TROUBLESHOOTING WAS PERFORMED TO FIX THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT, BOTH OF THE BATTERY CELLS WERE FOUND TO BE VENTED.
14-11-2022 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, BOTH POWER HANDLES WERE NOT TURNING ON OR HOLDING A CHARGE AND THE CHARGING BAY SEEMED TO BE MALFUNCTIONING. TROUBLESHOOTING WAS PERFORMED TO FIX THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT, BOTH OF THE BATTERY CELLS WERE FOUND TO BE VENTED.
13-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, IN THE MIDDLE OF A PANCREATIC TAIL RESECTION, THE FIRING WAS COMPLETED AND THE KNIFE STOPPED IN THE MIDWAY OF RETURNING. THE DISPLAY SHOWED A HANDLE WITH RED CIRCLE AND A SLASH, AND PRESSING THE BUTTON AGAIN TO RETURN THE KNIFE DID NOT RESUME THE OPERATION, SO THE ADAPTER WAS FORCIBLY REMOVED AND THE KNIFE WAS MANUALLY RETURNED USING THE MANUAL ADAPTER TOOL. THE SURGERY EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, IN THE MIDDLE OF A PANCREATIC TAIL RESECTION, THE FIRING WAS COMPLETED AND THE KNIFE STOPPED IN THE MIDWAY OF RETURNING. THE DISPLAY SHOWED A HANDLE WITH RED CIRCLE AND A SLASH, AND PRESSING THE BUTTON AGAIN TO RETURN THE KNIFE DID NOT RESUME THE OPERATION, SO THE ADAPTER WAS FORCIBLY REMOVED AND THE KNIFE WAS MANUALLY RETURNED USING THE MANUAL ADAPTER TOOL. THE SURGERY EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, IN THE MIDDLE OF A PANCREATIC TAIL RESECTION, THE FIRING WAS COMPLETED AND THE KNIFE STOPPED IN THE MIDWAY OF RETURNING. THE DISPLAY SHOWED A HANDLE WITH RED CIRCLE AND A SLASH, AND PRESSING THE BUTTON AGAIN TO RETURN THE KNIFE DID NOT RESUME THE OPERATION, SO THE ADAPTER WAS FORCIBLY REMOVED AND THE KNIFE WAS MANUALLY RETURNED USING THE MANUAL ADAPTER TOOL. THE SURGERY EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, IN THE MIDDLE OF A PANCREATIC TAIL RESECTION, THE FIRING WAS COMPLETED AND THE KNIFE STOPPED IN THE MIDWAY OF RETURNING. THE DISPLAY SHOWED A HANDLE WITH RED CIRCLE AND A SLASH, AND PRESSING THE BUTTON AGAIN TO RETURN THE KNIFE DID NOT RESUME THE OPERATION, SO THE ADAPTER WAS FORCIBLY REMOVED AND THE KNIFE WAS MANUALLY RETURNED USING THE MANUAL ADAPTER TOOL. THE SURGERY EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, IN THE MIDDLE OF A PANCREATIC TAIL RESECTION, THE FIRING WAS COMPLETED AND THE KNIFE STOPPED IN THE MIDWAY OF RETURNING. THE DISPLAY SHOWED A HANDLE WITH RED CIRCLE AND A SLASH, AND PRESSING THE BUTTON AGAIN TO RETURN THE KNIFE DID NOT RESUME THE OPERATION, SO THE ADAPTER WAS FORCIBLY REMOVED AND THE KNIFE WAS MANUALLY RETURNED USING THE MANUAL ADAPTER TOOL. THE SURGERY EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, IN THE MIDDLE OF A PANCREATIC TAIL RESECTION, THE FIRING WAS COMPLETED AND THE KNIFE STOPPED IN THE MIDWAY OF RETURNING. THE DISPLAY SHOWED A HANDLE WITH RED CIRCLE AND A SLASH, AND PRESSING THE BUTTON AGAIN TO RETURN THE KNIFE DID NOT RESUME THE OPERATION, SO THE ADAPTER WAS FORCIBLY REMOVED AND THE KNIFE WAS MANUALLY RETURNED USING THE MANUAL ADAPTER TOOL. THE SURGERY EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | covidien lp llc north haven | 30-01-2023 | SIGNIA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY, IN THE MIDDLE OF A PANCREATIC TAIL RESECTION, THE FIRING WAS COMPLETED AND THE KNIFE STOPPED IN THE MIDWAY OF RETURNING. THE DISPLAY SHOWED A HANDLE WITH RED CIRCLE AND A SLASH, AND PRESSING THE BUTTON AGAIN TO RETURN THE KNIFE DID NOT RESUME THE OPERATION, SO THE ADAPTER WAS FORCIBLY REMOVED AND THE KNIFE WAS MANUALLY RETURNED USING THE MANUAL ADAPTER TOOL. THE SURGERY EXTENDED FOR MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
14-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC SURGERY, WHILE AT UPPER LOBE OF LEFT LUNG, THE STAPLES WERE POORLY FORMED, AND THE FIRING WAS NOT SMOOTH FOR THE 3 RELOADS. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2023 | Malfunction | us surgical puerto rico | 30-01-2023 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC SURGERY, WHILE AT UPPER LOBE OF LEFT LUNG, THE STAPLES WERE POORLY FORMED, AND THE FIRING WAS NOT SMOOTH FOR THE 3 RELOADS. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING PROCEDURE, THE TISSUE COULD NOT BE GRASPED, AND THE DEVICE COULD NOT BE USED. A NEW DIFFERENT RELOAD WAS TAKEN OUT AND WAS USED WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIORRHAPHY, DURING MESH FIXATION INTO THE ABDOMINAL WALL, AFTER THE FIRING SEQUENCE, THE TACK GOT STUCK IN THE SHAFT OF THE TACKER. THE TACKER STUCK TO THE ABDOMINAL WALL. THE SURGEON UNSCREWED THE TACK FROM THE MESH TO BE ABLE TO RELEASE THE TACKER. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | SIGNIA | Imprecision | A0803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE WAS NOT FUNCTIONAL EVEN WITH CASES LEFT. THE HANDLE WOULD NOT COMPLETE THE SO FTWARE UPDATE. THE SCREEN ONLY READS ¿VER¿, NOT THE FULL WORD OF VERSION AND THE SOFTWARE LEVEL. THE DEVICE WAS ALSO INITIALLY UNABLE TO ENGAGE WITH A SHELL AND ADAPTER. OTHER SHELLS AND ADAPTERS WERE TRIED WITH NO SUCCESS. THE DEVICE TOOK SEVEN HARD RESETS TO GET THE DEVICE FUNCTIONAL FOR TWO FIRINGS UNTIL THE DEVICE SHUT DOWN AGAIN. THERE WAS NO PATIENT INVOLVEMENT.
13-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE WAS NOT FUNCTIONAL EVEN WITH CASES LEFT. THE HANDLE WOULD NOT COMPLETE THE SO FTWARE UPDATE. THE SCREEN ONLY READS ¿VER¿, NOT THE FULL WORD OF VERSION AND THE SOFTWARE LEVEL. THE DEVICE WAS ALSO INITIALLY UNABLE TO ENGAGE WITH A SHELL AND ADAPTER. OTHER SHELLS AND ADAPTERS WERE TRIED WITH NO SUCCESS. THE DEVICE TOOK SEVEN HARD RESETS TO GET THE DEVICE FUNCTIONAL FOR TWO FIRINGS UNTIL THE DEVICE SHUT DOWN AGAIN. THERE WAS NO PATIENT INVOLVEMENT.
13-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | SIGNIA | Application Program Version or Upgrade Problem | A110206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE WAS NOT FUNCTIONAL EVEN WITH CASES LEFT. THE HANDLE WOULD NOT COMPLETE THE SO FTWARE UPDATE. THE SCREEN ONLY READS ¿VER¿, NOT THE FULL WORD OF VERSION AND THE SOFTWARE LEVEL. THE DEVICE WAS ALSO INITIALLY UNABLE TO ENGAGE WITH A SHELL AND ADAPTER. OTHER SHELLS AND ADAPTERS WERE TRIED WITH NO SUCCESS. THE DEVICE TOOK SEVEN HARD RESETS TO GET THE DEVICE FUNCTIONAL FOR TWO FIRINGS UNTIL THE DEVICE SHUT DOWN AGAIN. THERE WAS NO PATIENT INVOLVEMENT.
13-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | SIGNIA | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE WAS NOT FUNCTIONAL EVEN WITH CASES LEFT. THE HANDLE WOULD NOT COMPLETE THE SO FTWARE UPDATE. THE SCREEN ONLY READS ¿VER¿, NOT THE FULL WORD OF VERSION AND THE SOFTWARE LEVEL. THE DEVICE WAS ALSO INITIALLY UNABLE TO ENGAGE WITH A SHELL AND ADAPTER. OTHER SHELLS AND ADAPTERS WERE TRIED WITH NO SUCCESS. THE DEVICE TOOK SEVEN HARD RESETS TO GET THE DEVICE FUNCTIONAL FOR TWO FIRINGS UNTIL THE DEVICE SHUT DOWN AGAIN. THERE WAS NO PATIENT INVOLVEMENT.
13-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE WAS NOT FUNCTIONAL EVEN WITH CASES LEFT. THE HANDLE WOULD NOT COMPLETE THE SO FTWARE UPDATE. THE SCREEN ONLY READS ¿VER¿, NOT THE FULL WORD OF VERSION AND THE SOFTWARE LEVEL. THE DEVICE WAS ALSO INITIALLY UNABLE TO ENGAGE WITH A SHELL AND ADAPTER. OTHER SHELLS AND ADAPTERS WERE TRIED WITH NO SUCCESS. THE DEVICE TOOK SEVEN HARD RESETS TO GET THE DEVICE FUNCTIONAL FOR TWO FIRINGS UNTIL THE DEVICE SHUT DOWN AGAIN. THERE WAS NO PATIENT INVOLVEMENT.
16-11-2022 | Malfunction | intuitive surgical | 31-01-2023 | SUREFORM | Connection Problem | A12 | Insufficient Information | Event Description: SUREFORM 45 STAPLER NOT ENGAGING PROPERLY DURING PROCEDURE.
22-12-2022 | Malfunction | ethicon endo-surgery | 31-01-2023 | PROXIMATE*RELOAD THICK TISSUE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, WHEN FIRED THE DISTAL PORTION OF THE STAPLE DIDN¿T MAKE B SHAPE COMPLETELY. A NEW RELOAD WAS USED TO FINISH THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES.
16-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. THIS HAPPENED ON THREE RELOADS.
21-11-2022 | Malfunction | ethicon endo-surgery | 31-01-2023 | 75MM SELECTABLE NEW TLC | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC UNKNOWN PROCEDURE, THE FIRING KNOB BECAME NOT TO MOVE AND COULD NOT BE FIRED AT 4TH FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
16-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. THIS HAPPENED ON THREE RELOADS.
16-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. THIS HAPPENED ON THREE RELOADS.
16-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. THIS HAPPENED ON THREE RELOADS.
16-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE, THE STAPLER WAS NOT ABLE TO FIRE THE RELOAD. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE, THE WHOLE DEVICE WAS REPLACED FOR A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRIC SLEEVE, THE STAPLER WAS NOT ABLE TO FIRE THE RELOAD. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE, THE WHOLE DEVICE WAS REPLACED FOR A NEW ONE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. THE ISSUE OCCURRED ON TWO HANDLES AND ONE RELOAD. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE.
10-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER PRESSING THE GREEN BUTTON, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE THREE DEVICES DID NOT FIRE.
17-11-2022 | Malfunction | ethicon endo-surgery | 31-01-2023 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT OPEN HEMICOLECTOMY FOR APPENDICEAL CA, THE MENTIONED STAPLER WITH THE CORRESPONDING RELOADS WAS TO BE USED. THE STAPLER WAS APPLIED, BUT THE CLIPS WERE NOT RELEASED, NOR WAS CUTTING POSSIBLE. THE PROCEDURE WAS SUCCESSFULLY CONTINUED WITH A NEW STAPLER AND MAGAZINE. THERE WERE NO PATIENT CONSEQUENCES.
17-11-2022 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SMALL INTESTINAL RESECTION PROCEDURE, AFTER REMOVING THE PRODUCT FROM THE BLISTER PACK, THE SURGEON WAS WORRIED ABOUT THE OPENING OF THE JAW. THE JAWS WERE HALF-OPENED COMPARED TO THE NORMAL SITUATION. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE INTERLOCK WAS OVERRIDDEN AND THE RELOAD WAS CYCLED WITHOUT HESITATION OR BINDING AND THE RELOAD INTERLOCK WAS TESTED AND FOUND TO FUNCTION PROPERLY.
17-11-2022 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SMALL INTESTINAL RESECTION PROCEDURE, AFTER REMOVING THE PRODUCT FROM THE BLISTER PACK, THE SURGEON WAS WORRIED ABOUT THE OPENING OF THE JAW. THE JAWS WERE HALF-OPENED COMPARED TO THE NORMAL SITUATION. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE INTERLOCK WAS OVERRIDDEN AND THE RELOAD WAS CYCLED WITHOUT HESITATION OR BINDING AND THE RELOAD INTERLOCK WAS TESTED AND FOUND TO FUNCTION PROPERLY.
08-01-2023 | Malfunction | ethicon endo-surgery | 31-01-2023 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHEN THE SURGEON TRIED TO FIRE THE STAPLER, IT DID NOT RELEASE THE STAPLES, THEY GOT STUCK, FOR WHICH HE REQUESTED A NEW ONE. WITH THIS, SHE DID NOT HAVE ANY PROBLEMS AND THE SURGERY WAS COMPLETED SUCCESSFULLY WITHOUT ANY DAMAGE TO THE PATIENT OR TISSUE. PROCEDURE: UNKNOWN.
13-12-2022 | Malfunction | ethicon endo-surgery | 31-01-2023 | GST GOLD RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: THIS CASE IS FROM HEALTH AUTHORITY. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED . IMMEDIATELY STOP USING. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-11-2022 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE SURGEON PREPARED THE DEVICE WITH AN INSTALLED SEAM GUARD (REINFORCEMENT MATERIAL) ON THE RELOAD. THE SURGEON WAS UNABLE TO CLOSE THE JAWS COMPLETELY DUE TO THE THICKNESS OF THE TISSUE AND SEAM GUARD. THEREFORE THE SURGEON COULD NOT PROCEED WITH PRESSING THE GREEN BUTTON. JAWS WERE UNABLE TO CLOSE COMPLETELY ON THE THREE RELOADS THAT WERE USED ON THE TWO HANDLES. THE SURGEON ALSO USED A BIPOLAR INSTRUMENT FOR THE HEMOSTASIS. THERE WAS NO PATIENT INJURY.
02-11-2022 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE SURGEON PREPARED THE DEVICE WITH AN INSTALLED SEAM GUARD (REINFORCEMENT MATERIAL) ON THE RELOAD. THE SURGEON WAS UNABLE TO CLOSE THE JAWS COMPLETELY DUE TO THE THICKNESS OF THE TISSUE AND SEAM GUARD. THEREFORE THE SURGEON COULD NOT PROCEED WITH PRESSING THE GREEN BUTTON. JAWS WERE UNABLE TO CLOSE COMPLETELY ON THE THREE RELOADS THAT WERE USED ON THE TWO HANDLES. THE SURGEON ALSO USED A BIPOLAR INSTRUMENT FOR THE HEMOSTASIS. THERE WAS NO PATIENT INJURY.
14-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC SURGERY, WHILE AT UPPER LOBE OF LEFT LUNG, THE STAPLES WERE POORLY FORMED, AND THE FIRING WAS NOT SMOOTH FOR THE 3 RELOADS. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC SURGERY, WHILE AT UPPER LOBE OF LEFT LUNG, THE STAPLES WERE POORLY FORMED, AND THE FIRING WAS NOT SMOOTH FOR THE 3 RELOADS. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-11-2022 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, THE SURGEON PREPARED THE DEVICE WITH AN INSTALLED SEAM GUARD (REINFORCEMENT MATERIAL) ON THE RELOAD. THE SURGEON WAS UNABLE TO CLOSE THE JAWS COMPLETELY DUE TO THE THICKNESS OF THE TISSUE AND SEAM GUARD. THEREFORE THE SURGEON COULD NOT PROCEED WITH PRESSING THE GREEN BUTTON. JAWS WERE UNABLE TO CLOSE COMPLETELY ON THE THREE RELOADS THAT WERE USED ON THE TWO HANDLES. THE SURGEON ALSO USED A BIPOLAR INSTRUMENT FOR THE HEMOSTASIS. THERE WAS NO PATIENT INJURY.
14-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC SURGERY, WHILE AT UPPER LOBE OF LEFT LUNG, THE STAPLES WERE POORLY FORMED, AND THE FIRING WAS NOT SMOOTH FOR THE 3 RELOADS. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC SURGERY, WHILE AT UPPER LOBE OF LEFT LUNG, THE STAPLES WERE POORLY FORMED, AND THE FIRING WAS NOT SMOOTH FOR THE 3 RELOADS. THE DEVICE WAS REPLACED TO COMPLETE THE CASE AND RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN COELIO APPENDECTOMY, WHILE AT THE BASE OF APPENDIX, THE JAWS OF THE FORCEPS REMAINED LOCKED IN THE CLOSED POSITION INSIDE THE PATIENT'S TISSUE. THE SURGEON WAS ABLE TO REMOVE THE JAWS BY MANEUVERING THE HANDLE SEVERAL TIMES. A NEW HANDLE AND RELOAD WERE OPENED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER PRESSING THE GREEN BUTTON, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE THREE DEVICES DID NOT FIRE.
30-12-2022 | Injury | intuitive surgical | 31-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Abscess | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LOW ANTERIOR RESECTION AND ABDOMINOPERINEAL RESECTION SURGICAL PROCEDURE, THE STAPLE LINE OPENED UP AFTER USING A SUREFORM BLUE RELOAD. THE STAPLES DID NOT PROPERLY FORM. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE INDICATION FOR THIS PROCEDURE WAS RECTAL CANCER. THE SURGEON REPORTED THAT THIS COMPLICATION OCCURRED DURING THE STAPLER INSTRUMENT¿S SECOND FIRE OF THE PROCEDURE, AND OCCURRED WHILE TRANSECTING THE DISTAL RECTUM. THE SURGEON REPORTED THAT THE STAPLER¿S FIRST FIRE OCCURRED WITHOUT ISSUE. HOWEVER, DURING THE CLAMPING PHASE OF THIS SECOND FIRE, A MESSAGE WAS RECEIVED SAYING THE TISSUE WAS TOO THICK. THE SURGEON DISSECTED THE TISSUE AND CLAMPED AGAIN WITHOUT ERROR. THE SUREFORM 45 FIRED THE BLUE RELOAD WITHOUT ISSUE, BUT UPON UNCLAMPING, THE SURGEON NOTICED THAT THE STAPLE LINE DID NOT STAY TOGETHER. THE SURGEON REPORTED BEING ABLE TO SEE INTO THE LUMEN OF THE RECTUM. THE STAPLES ALL DEPLOYED, BUT THE STAPLE LINE DID NOT HOLD. THE SURGEON SAID THERE WERE NO MALFORMED STAPLES, AND NO BLEEDING. THE SURGEON REPORTED THAT THERE WERE NO OBSTRUCTIONS IN THE WAY OF THIS STAPLE LINE, BUT THAT THEY DID FIRE OVER AN EXISTING STAPLE LINE. DUE TO THE STAPLE LINE COMPLICATION, THE SURGEON REPORTED LEAVING IN A DRAIN FOR HEALING. THE SURGEON FIRED ANOTHER STAPLE LINE TO REMOVE THE INCOMPLETE STAPLE LINE WITH THE SPECIMEN. THE SURGEON SAID IT IS UNCLEAR HOW THIS EVENT EFFECTED THE PROCEDURE AND PATIENT OUTCOME, BUT ADDED THAT IT MAY HAVE CONTRIBUTED TO THE PATIENT ACQUIRING A POST-OPERATION SURGICAL SITE INFECTION AT THE RECTAL STUMP WITH STAPLE LINE DEHISCENCE. THE SURGEON SAID THIS EVENT COULD HAVE SPILLED MICROSCOPIC TUMOR PARTICLES AND CAUSE A LONG-TERM COMPLICATION. THE SURGEON ADDED THAT IT IS THE LONG-TERM RELEVANCE OF THE RECTAL CUFF IS UNDETERMINED; THE SURGEON SAID IT SHOULD HEAL WITH A DRAIN PLACED. THE PATIENT EXPERIENCED POST-OPERATIVE STAPLE LINE DEHISCENCE WITH A SMALL PELVIC ABSCESS THAT REQUIRED ORAL ANTIBIOTICS, PROLONGED INPATIENT STAY, AND TRANSRECTAL DRAINAGE. THE PATIENT WAS READMITTED TO THE HOSPITAL DUE TO A FALL AT REHAB THAT RESULTED IN A BROKEN FEMUR.
30-12-2022 | Injury | intuitive surgical | 31-01-2023 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LOW ANTERIOR RESECTION AND ABDOMINOPERINEAL RESECTION SURGICAL PROCEDURE, THE STAPLE LINE OPENED UP AFTER USING A SUREFORM BLUE RELOAD. THE STAPLES DID NOT PROPERLY FORM. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE INDICATION FOR THIS PROCEDURE WAS RECTAL CANCER. THE SURGEON REPORTED THAT THIS COMPLICATION OCCURRED DURING THE STAPLER INSTRUMENT¿S SECOND FIRE OF THE PROCEDURE, AND OCCURRED WHILE TRANSECTING THE DISTAL RECTUM. THE SURGEON REPORTED THAT THE STAPLER¿S FIRST FIRE OCCURRED WITHOUT ISSUE. HOWEVER, DURING THE CLAMPING PHASE OF THIS SECOND FIRE, A MESSAGE WAS RECEIVED SAYING THE TISSUE WAS TOO THICK. THE SURGEON DISSECTED THE TISSUE AND CLAMPED AGAIN WITHOUT ERROR. THE SUREFORM 45 FIRED THE BLUE RELOAD WITHOUT ISSUE, BUT UPON UNCLAMPING, THE SURGEON NOTICED THAT THE STAPLE LINE DID NOT STAY TOGETHER. THE SURGEON REPORTED BEING ABLE TO SEE INTO THE LUMEN OF THE RECTUM. THE STAPLES ALL DEPLOYED, BUT THE STAPLE LINE DID NOT HOLD. THE SURGEON SAID THERE WERE NO MALFORMED STAPLES, AND NO BLEEDING. THE SURGEON REPORTED THAT THERE WERE NO OBSTRUCTIONS IN THE WAY OF THIS STAPLE LINE, BUT THAT THEY DID FIRE OVER AN EXISTING STAPLE LINE. DUE TO THE STAPLE LINE COMPLICATION, THE SURGEON REPORTED LEAVING IN A DRAIN FOR HEALING. THE SURGEON FIRED ANOTHER STAPLE LINE TO REMOVE THE INCOMPLETE STAPLE LINE WITH THE SPECIMEN. THE SURGEON SAID IT IS UNCLEAR HOW THIS EVENT EFFECTED THE PROCEDURE AND PATIENT OUTCOME, BUT ADDED THAT IT MAY HAVE CONTRIBUTED TO THE PATIENT ACQUIRING A POST-OPERATION SURGICAL SITE INFECTION AT THE RECTAL STUMP WITH STAPLE LINE DEHISCENCE. THE SURGEON SAID THIS EVENT COULD HAVE SPILLED MICROSCOPIC TUMOR PARTICLES AND CAUSE A LONG-TERM COMPLICATION. THE SURGEON ADDED THAT IT IS THE LONG-TERM RELEVANCE OF THE RECTAL CUFF IS UNDETERMINED; THE SURGEON SAID IT SHOULD HEAL WITH A DRAIN PLACED. THE PATIENT EXPERIENCED POST-OPERATIVE STAPLE LINE DEHISCENCE WITH A SMALL PELVIC ABSCESS THAT REQUIRED ORAL ANTIBIOTICS, PROLONGED INPATIENT STAY, AND TRANSRECTAL DRAINAGE. THE PATIENT WAS READMITTED TO THE HOSPITAL DUE TO A FALL AT REHAB THAT RESULTED IN A BROKEN FEMUR.
30-12-2022 | Injury | intuitive surgical | 31-01-2023 | SUREFORM | Failure to Fire | A050501 | Abscess | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LOW ANTERIOR RESECTION AND ABDOMINOPERINEAL RESECTION SURGICAL PROCEDURE, THE STAPLE LINE OPENED UP AFTER USING A SUREFORM BLUE RELOAD. THE STAPLES DID NOT PROPERLY FORM. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE INDICATION FOR THIS PROCEDURE WAS RECTAL CANCER. THE SURGEON REPORTED THAT THIS COMPLICATION OCCURRED DURING THE STAPLER INSTRUMENT¿S SECOND FIRE OF THE PROCEDURE, AND OCCURRED WHILE TRANSECTING THE DISTAL RECTUM. THE SURGEON REPORTED THAT THE STAPLER¿S FIRST FIRE OCCURRED WITHOUT ISSUE. HOWEVER, DURING THE CLAMPING PHASE OF THIS SECOND FIRE, A MESSAGE WAS RECEIVED SAYING THE TISSUE WAS TOO THICK. THE SURGEON DISSECTED THE TISSUE AND CLAMPED AGAIN WITHOUT ERROR. THE SUREFORM 45 FIRED THE BLUE RELOAD WITHOUT ISSUE, BUT UPON UNCLAMPING, THE SURGEON NOTICED THAT THE STAPLE LINE DID NOT STAY TOGETHER. THE SURGEON REPORTED BEING ABLE TO SEE INTO THE LUMEN OF THE RECTUM. THE STAPLES ALL DEPLOYED, BUT THE STAPLE LINE DID NOT HOLD. THE SURGEON SAID THERE WERE NO MALFORMED STAPLES, AND NO BLEEDING. THE SURGEON REPORTED THAT THERE WERE NO OBSTRUCTIONS IN THE WAY OF THIS STAPLE LINE, BUT THAT THEY DID FIRE OVER AN EXISTING STAPLE LINE. DUE TO THE STAPLE LINE COMPLICATION, THE SURGEON REPORTED LEAVING IN A DRAIN FOR HEALING. THE SURGEON FIRED ANOTHER STAPLE LINE TO REMOVE THE INCOMPLETE STAPLE LINE WITH THE SPECIMEN. THE SURGEON SAID IT IS UNCLEAR HOW THIS EVENT EFFECTED THE PROCEDURE AND PATIENT OUTCOME, BUT ADDED THAT IT MAY HAVE CONTRIBUTED TO THE PATIENT ACQUIRING A POST-OPERATION SURGICAL SITE INFECTION AT THE RECTAL STUMP WITH STAPLE LINE DEHISCENCE. THE SURGEON SAID THIS EVENT COULD HAVE SPILLED MICROSCOPIC TUMOR PARTICLES AND CAUSE A LONG-TERM COMPLICATION. THE SURGEON ADDED THAT IT IS THE LONG-TERM RELEVANCE OF THE RECTAL CUFF IS UNDETERMINED; THE SURGEON SAID IT SHOULD HEAL WITH A DRAIN PLACED. THE PATIENT EXPERIENCED POST-OPERATIVE STAPLE LINE DEHISCENCE WITH A SMALL PELVIC ABSCESS THAT REQUIRED ORAL ANTIBIOTICS, PROLONGED INPATIENT STAY, AND TRANSRECTAL DRAINAGE. THE PATIENT WAS READMITTED TO THE HOSPITAL DUE TO A FALL AT REHAB THAT RESULTED IN A BROKEN FEMUR.
30-12-2022 | Injury | intuitive surgical | 31-01-2023 | SUREFORM | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LOW ANTERIOR RESECTION AND ABDOMINOPERINEAL RESECTION SURGICAL PROCEDURE, THE STAPLE LINE OPENED UP AFTER USING A SUREFORM BLUE RELOAD. THE STAPLES DID NOT PROPERLY FORM. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE INDICATION FOR THIS PROCEDURE WAS RECTAL CANCER. THE SURGEON REPORTED THAT THIS COMPLICATION OCCURRED DURING THE STAPLER INSTRUMENT¿S SECOND FIRE OF THE PROCEDURE, AND OCCURRED WHILE TRANSECTING THE DISTAL RECTUM. THE SURGEON REPORTED THAT THE STAPLER¿S FIRST FIRE OCCURRED WITHOUT ISSUE. HOWEVER, DURING THE CLAMPING PHASE OF THIS SECOND FIRE, A MESSAGE WAS RECEIVED SAYING THE TISSUE WAS TOO THICK. THE SURGEON DISSECTED THE TISSUE AND CLAMPED AGAIN WITHOUT ERROR. THE SUREFORM 45 FIRED THE BLUE RELOAD WITHOUT ISSUE, BUT UPON UNCLAMPING, THE SURGEON NOTICED THAT THE STAPLE LINE DID NOT STAY TOGETHER. THE SURGEON REPORTED BEING ABLE TO SEE INTO THE LUMEN OF THE RECTUM. THE STAPLES ALL DEPLOYED, BUT THE STAPLE LINE DID NOT HOLD. THE SURGEON SAID THERE WERE NO MALFORMED STAPLES, AND NO BLEEDING. THE SURGEON REPORTED THAT THERE WERE NO OBSTRUCTIONS IN THE WAY OF THIS STAPLE LINE, BUT THAT THEY DID FIRE OVER AN EXISTING STAPLE LINE. DUE TO THE STAPLE LINE COMPLICATION, THE SURGEON REPORTED LEAVING IN A DRAIN FOR HEALING. THE SURGEON FIRED ANOTHER STAPLE LINE TO REMOVE THE INCOMPLETE STAPLE LINE WITH THE SPECIMEN. THE SURGEON SAID IT IS UNCLEAR HOW THIS EVENT EFFECTED THE PROCEDURE AND PATIENT OUTCOME, BUT ADDED THAT IT MAY HAVE CONTRIBUTED TO THE PATIENT ACQUIRING A POST-OPERATION SURGICAL SITE INFECTION AT THE RECTAL STUMP WITH STAPLE LINE DEHISCENCE. THE SURGEON SAID THIS EVENT COULD HAVE SPILLED MICROSCOPIC TUMOR PARTICLES AND CAUSE A LONG-TERM COMPLICATION. THE SURGEON ADDED THAT IT IS THE LONG-TERM RELEVANCE OF THE RECTAL CUFF IS UNDETERMINED; THE SURGEON SAID IT SHOULD HEAL WITH A DRAIN PLACED. THE PATIENT EXPERIENCED POST-OPERATIVE STAPLE LINE DEHISCENCE WITH A SMALL PELVIC ABSCESS THAT REQUIRED ORAL ANTIBIOTICS, PROLONGED INPATIENT STAY, AND TRANSRECTAL DRAINAGE. THE PATIENT WAS READMITTED TO THE HOSPITAL DUE TO A FALL AT REHAB THAT RESULTED IN A BROKEN FEMUR.
10-01-2023 | Injury | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY, POOR STAPLE FORMATION AND MISFIRING ON FOUR RELOADS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
03-01-2023 | Injury | intuitive surgical | 31-01-2023 | ENDOWRIST | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RIGHT PULMONARY LOBECTOMY, WHILE STAPLING THE PULMONARY ARTERY WITH AN ENDOWRIST 30 CURVED-TIP STAPLER INSTRUMENT, THE PATIENT EXPERIENCED ABNORMAL BLEEDING. THE ESTIMATED BLOOD LOSS VOLUME ASSOCIATED WITH THE EVENT IS UNKNOWN. DURING FOLLOW UP, THE INTUITIVE SURGICAL, INC. (ISI) CLINICAL TERRITORY ASSOCIATE (CTA), WHO WAS PRESENT DURING THE EVENT, INDICATED THE BLEEDING OCCURRED AFTER THE USE OF A WHITE STAPLER 30 RELOAD AND WAS CONTROLLED WITH THE USE OF MULTIPLE HEMOSTATIC AGENTS: SURGICEL, FLOSEAL AND PROGEL. THE PROCEDURE WAS COMPLETED ROBOTICALLY. PER CTA, THE SURGEON CONCLUDED THE PATIENT'S TISSUE CONDITION WAS PROBABLY THE REASON FOR THE BLEEDING AND NOT NECESSARILY THE INSTRUMENT. ISI HAS ATTEMPTED TO FOLLOW-UP WITH THE SURGEON TO GATHER ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT. HOWEVER, AS OF THE DATE OF THIS REPORT, NO NEW INFORMATION HAS BEEN OBTAINED.
10-01-2023 | Injury | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY, POOR STAPLE FORMATION AND MISFIRING ON FOUR RELOADS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
28-09-2022 | Malfunction | us surgical puerto rico | 31-01-2023 | TA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING BOWEL RESECTION, WHEN TRYING TO FIRE THE DEVICE, STAPLER DID NOT FIRE COMPLETELY. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-01-2023 | Injury | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY, POOR STAPLE FORMATION AND MISFIRING ON FOUR RELOADS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
10-01-2023 | Injury | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS AN INCOMPLETE STAPLE LINE CENTRALLY, POOR STAPLE FORMATION AND MISFIRING ON FOUR RELOADS. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
03-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN COELIO APPENDECTOMY, WHILE AT THE BASE OF APPENDIX, THE JAWS OF THE FORCEPS REMAINED LOCKED IN THE CLOSED POSITION INSIDE THE PATIENT'S TISSUE. THE SURGEON WAS ABLE TO REMOVE THE JAWS BY MANEUVERING THE HANDLE SEVERAL TIMES. A NEW HANDLE AND RELOAD WERE OPENED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOTOMY, DURING THE SECOND STAPLING, WHEN THE GREEN BUTTON WAS PRESSED, THE STAPLING DID NOT OCCUR. MULTIPLE ATTEMPTS WERE MADE BUT WERE NOT SUCCESSFUL. THERE WAS DIFFICULTY ON LOADING AND UNLOADING THE DEVICE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOTOMY, DURING THE SECOND STAPLING, WHEN THE GREEN BUTTON WAS PRESSED, THE STAPLING DID NOT OCCUR. MULTIPLE ATTEMPTS WERE MADE BUT WERE NOT SUCCESSFUL. THERE WAS DIFFICULTY ON LOADING AND UNLOADING THE DEVICE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOTOMY, DURING THE SECOND STAPLING, WHEN THE GREEN BUTTON WAS PRESSED, THE STAPLING DID NOT OCCUR. MULTIPLE ATTEMPTS WERE MADE BUT WERE NOT SUCCESSFUL. THERE WAS DIFFICULTY ON LOADING AND UNLOADING THE DEVICE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DEROOFING OF HEPATIC CYSTS WITH CHOLECYSTECTOMY PROCEDURE, WHILE STAPLING THE LIVER, THE DEVICE DID NOT FIRE. WHEN HE SURGEON TRIED FIRING THE RELOAD, THE HANDLE MADE A CRACKING SOUND WHEN HANDLE WAS SQUEEZED AND WAS NO LONGER WORKING. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO UNLOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DEROOFING OF HEPATIC CYSTS WITH CHOLECYSTECTOMY PROCEDURE, WHILE STAPLING THE LIVER, THE DEVICE DID NOT FIRE. WHEN HE SURGEON TRIED FIRING THE RELOAD, THE HANDLE MADE A CRACKING SOUND WHEN HANDLE WAS SQUEEZED AND WAS NO LONGER WORKING. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO UNLOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DEROOFING OF HEPATIC CYSTS WITH CHOLECYSTECTOMY PROCEDURE, WHILE STAPLING THE LIVER, THE DEVICE DID NOT FIRE. WHEN HE SURGEON TRIED FIRING THE RELOAD, THE HANDLE MADE A CRACKING SOUND WHEN HANDLE WAS SQUEEZED AND WAS NO LONGER WORKING. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO UNLOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DEROOFING OF HEPATIC CYSTS WITH CHOLECYSTECTOMY PROCEDURE, WHILE STAPLING THE LIVER, THE DEVICE DID NOT FIRE. WHEN HE SURGEON TRIED FIRING THE RELOAD, THE HANDLE MADE A CRACKING SOUND WHEN HANDLE WAS SQUEEZED AND WAS NO LONGER WORKING. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO UNLOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING, THE STAPLER DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
16-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING, THE STAPLER DID NOT FIRE. THERE WAS NO PATIENT INVOLVED.
11-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DEROOFING OF HEPATIC CYSTS WITH CHOLECYSTECTOMY PROCEDURE, WHILE STAPLING THE LIVER, THE DEVICE DID NOT FIRE. WHEN HE SURGEON TRIED FIRING THE RELOAD, THE HANDLE MADE A CRACKING SOUND WHEN HANDLE WAS SQUEEZED AND WAS NO LONGER WORKING. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO UNLOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-01-2023 | Malfunction | covidien lp llc north haven | 31-01-2023 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DEROOFING OF HEPATIC CYSTS WITH CHOLECYSTECTOMY PROCEDURE, WHILE STAPLING THE LIVER, THE DEVICE DID NOT FIRE. WHEN HE SURGEON TRIED FIRING THE RELOAD, THE HANDLE MADE A CRACKING SOUND WHEN HANDLE WAS SQUEEZED AND WAS NO LONGER WORKING. IT WAS ALSO NOTED THAT THE DEVICE WAS DIFFICULT TO UNLOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2023 | Injury | us surgical puerto rico | 31-01-2023 | ENDO GIA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE OPERATOR FIRED THE STAPLER ACROSS THE CECUM, WHEN BLADE WAS CUTTING STAPLED TISSUE, A LOUD CRUNCHING SOUND WAS HEARD. WHEN STAPLER WAS REMOVED, THE STAPLE LINE WAS NOT FULLY CLOSED AND BOWEL WAS STILL OPENED. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THERE WAS AN UNANTICIPATED TISSUE LOSS, TISSUE DAMAGED, SURGICAL TIME EXTENDED 30 MINUTES OR MORE AND INCISION WAS EXTENDED FOR MORE THAN 1 INCH DUE TO THE PRODUCT PROBLEM. STAPLE LINE WAS FIXED WHEN PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE LAPAROSCOPIC PROCEDURE WAS UNABLE TO BE COMPLETED/CONVERTED TO OPEN PROCEDURE TO REPAIR BOWEL INJURY. THE HOSPITALIZATION HAS BEEN EXTENDED FOR 5 DAYS.
12-01-2023 | Injury | us surgical puerto rico | 31-01-2023 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE OPERATOR FIRED THE STAPLER ACROSS THE CECUM, WHEN BLADE WAS CUTTING STAPLED TISSUE, A LOUD CRUNCHING SOUND WAS HEARD. WHEN STAPLER WAS REMOVED, THE STAPLE LINE WAS NOT FULLY CLOSED AND BOWEL WAS STILL OPENED. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THERE WAS AN UNANTICIPATED TISSUE LOSS, TISSUE DAMAGED, SURGICAL TIME EXTENDED 30 MINUTES OR MORE AND INCISION WAS EXTENDED FOR MORE THAN 1 INCH DUE TO THE PRODUCT PROBLEM. STAPLE LINE WAS FIXED WHEN PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE LAPAROSCOPIC PROCEDURE WAS UNABLE TO BE COMPLETED/CONVERTED TO OPEN PROCEDURE TO REPAIR BOWEL INJURY. THE HOSPITALIZATION HAS BEEN EXTENDED FOR 5 DAYS.
12-01-2023 | Injury | us surgical puerto rico | 31-01-2023 | ENDO GIA | Noise, Audible | A0508 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE OPERATOR FIRED THE STAPLER ACROSS THE CECUM, WHEN BLADE WAS CUTTING STAPLED TISSUE, A LOUD CRUNCHING SOUND WAS HEARD. WHEN STAPLER WAS REMOVED, THE STAPLE LINE WAS NOT FULLY CLOSED AND BOWEL WAS STILL OPENED. THE STAPLE LINE WAS ALSO INCOMPLETE DISTALLY. THERE WAS AN UNANTICIPATED TISSUE LOSS, TISSUE DAMAGED, SURGICAL TIME EXTENDED 30 MINUTES OR MORE AND INCISION WAS EXTENDED FOR MORE THAN 1 INCH DUE TO THE PRODUCT PROBLEM. STAPLE LINE WAS FIXED WHEN PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE. THE LAPAROSCOPIC PROCEDURE WAS UNABLE TO BE COMPLETED/CONVERTED TO OPEN PROCEDURE TO REPAIR BOWEL INJURY. THE HOSPITALIZATION HAS BEEN EXTENDED FOR 5 DAYS.
13-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PULMONARY VEIN RESECTION ON THE RIGHT UPPER LOBE, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC MALIGNANT TUMOR SURGERY, WHILE UNLOADING THE DEVICE AFTER COMPLETION OF THE STAPLING, THE PROXIMAL SIDE OF THE RELOAD WAS DAMAGED. THE DEVICE COULD NOT BE REMOVED FROM THE HANDLE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
12-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC MALIGNANT TUMOR SURGERY, WHILE UNLOADING THE DEVICE AFTER COMPLETION OF THE STAPLING, THE PROXIMAL SIDE OF THE RELOAD WAS DAMAGED. THE DEVICE COULD NOT BE REMOVED FROM THE HANDLE. TO RESOLVE THE ISSUE, A NEW DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
14-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | EEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, DURING RECTAL ANASTOMOSIS, AT THE TIME OF THE DEVICE REMOVAL, THE ANVIL AND THE CENTER SHAFT WERE DISCONNECTED, AND ONLY THE ANVIL REMAINED IN THE INTESTINE. THE ANVIL WAS TAKEN OUT USING AN ANOSCOPE. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A NEPHROLITHOTOMY, THE INSIDE OF THE TWO DEVICES BROKE OFF. TO RESOLVE THE ISSUE, A NEW STAPLER WAS USED. THERE WAS NO PATIENT INJURY.
10-01-2023 | Malfunction | us surgical puerto rico | 31-01-2023 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A NEPHROLITHOTOMY, THE INSIDE OF THE TWO DEVICES BROKE OFF. TO RESOLVE THE ISSUE, A NEW STAPLER WAS USED. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 01-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC RECTAL CANCER SURGERY, IN DETACHING THE TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE SAME RELOAD. THERE WAS NO PATIENT INJURY.
04-12-2023 | Injury | covidien lp llc north haven | 01-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC PULMONARY SEGMENTECTOMY, WHEN THE PULMONARY ARTERY WAS BEING PROCESSED, FIRING WAS PERFORMED AND AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT IT COULD NO LONGER BE OPENED HALFWAY THROUGH; A BLEEDING OCCURRED. THORACOTOMY WAS PERFORMED, THE BLOOD VESSEL WAS CLAMPED, AND IT WAS PROCESSED WITH LIGATION. THE PULMONARY ARTERY WAS DAMAGED BY MORE THAN ONE INCH AND THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE PATIENT WAS CURRENTLY UNDER OBSERVATION.
04-12-2023 | Injury | covidien lp llc north haven | 01-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC PULMONARY SEGMENTECTOMY, WHEN THE PULMONARY ARTERY WAS BEING PROCESSED, FIRING WAS PERFORMED AND AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT IT COULD NO LONGER BE OPENED HALFWAY THROUGH; A BLEEDING OCCURRED. THORACOTOMY WAS PERFORMED, THE BLOOD VESSEL WAS CLAMPED, AND IT WAS PROCESSED WITH LIGATION. THE PULMONARY ARTERY WAS DAMAGED BY MORE THAN ONE INCH AND THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE PATIENT WAS CURRENTLY UNDER OBSERVATION.
04-12-2023 | Injury | covidien lp llc north haven | 01-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC PULMONARY SEGMENTECTOMY, WHEN THE PULMONARY ARTERY WAS BEING PROCESSED, FIRING WAS PERFORMED AND AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT IT COULD NO LONGER BE OPENED HALFWAY THROUGH; A BLEEDING OCCURRED. THORACOTOMY WAS PERFORMED, THE BLOOD VESSEL WAS CLAMPED, AND IT WAS PROCESSED WITH LIGATION. THE PULMONARY ARTERY WAS DAMAGED BY MORE THAN ONE INCH AND THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE PATIENT WAS CURRENTLY UNDER OBSERVATION.
04-12-2023 | Injury | covidien lp llc north haven | 01-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC PULMONARY SEGMENTECTOMY, WHEN THE PULMONARY ARTERY WAS BEING PROCESSED, FIRING WAS PERFORMED AND AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT IT COULD NO LONGER BE OPENED HALFWAY THROUGH; A BLEEDING OCCURRED. THORACOTOMY WAS PERFORMED, THE BLOOD VESSEL WAS CLAMPED, AND IT WAS PROCESSED WITH LIGATION. THE PULMONARY ARTERY WAS DAMAGED BY MORE THAN ONE INCH AND THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE PATIENT WAS CURRENTLY UNDER OBSERVATION.
04-12-2023 | Injury | us surgical puerto rico | 01-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC PULMONARY SEGMENTECTOMY, WHEN THE PULMONARY ARTERY WAS BEING PROCESSED, FIRING WAS PERFORMED AND AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT IT COULD NO LONGER BE OPENED HALFWAY THROUGH; A BLEEDING OCCURRED. THORACOTOMY WAS PERFORMED, THE BLOOD VESSEL WAS CLAMPED, AND IT WAS PROCESSED WITH LIGATION. THE PULMONARY ARTERY WAS DAMAGED BY MORE THAN ONE INCH AND THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE PATIENT WAS CURRENTLY UNDER OBSERVATION.
04-12-2023 | Injury | us surgical puerto rico | 01-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC PULMONARY SEGMENTECTOMY, WHEN THE PULMONARY ARTERY WAS BEING PROCESSED, FIRING WAS PERFORMED AND AN ATTEMPT WAS MADE TO OPEN THE CARTRIDGE, BUT IT COULD NO LONGER BE OPENED HALFWAY THROUGH; A BLEEDING OCCURRED. THORACOTOMY WAS PERFORMED, THE BLOOD VESSEL WAS CLAMPED, AND IT WAS PROCESSED WITH LIGATION. THE PULMONARY ARTERY WAS DAMAGED BY MORE THAN ONE INCH AND THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE PATIENT WAS CURRENTLY UNDER OBSERVATION.
01-04-2023 | Malfunction | us surgical puerto rico | 01-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WAS SQUEEZED, BUT THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-04-2023 | Malfunction | us surgical puerto rico | 01-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE HANDLE WAS SQUEEZED, BUT THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-01-2024 | Malfunction | covidien lp llc north haven | 01-01-2024 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE FIRING BAR STOPPED UP/DOWN MOTION.
19-01-2024 | Malfunction | covidien lp llc north haven | 01-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE FIRING BAR STOPPED UP/DOWN MOTION.
04-12-2023 | Malfunction | covidien lp llc north haven | 01-01-2024 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE DETACHING THE INTERLOBAR FISSURE, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED, BUT THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | covidien lp llc north haven | 01-01-2024 | UNKNOWN EGIA TRI STAPLE | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE DETACHING THE INTERLOBAR FISSURE, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED, BUT THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | us surgical puerto rico | 01-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE DETACHING THE INTERLOBAR FISSURE, THE GREEN BUTTON WAS PRESSED AND THE HANDLE WAS SQUEEZED, BUT THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2023 | Malfunction | covidien lp llc north haven | 01-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LUNG WEDGE RESECTION SURGERY, DURING THE REMOVAL OF LUNG TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD HAD THE ISSUE. THE DEVICE WAS REPLACED WITH A SET OF HANDLE AND STAPLES WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | covidien lp llc north haven | 01-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, BRONCHIAL RESECTION, ONE STAPLE NEEDLE DID NOT GET SEPARATED FROM THE CARTRIDGE AND HAS POOR STAPLE FORMATION. THE STAPLE LINE WAS ALSO INCOMPLETE CENTRALLY. THEREFORE, THE CARTRIDGE COULD NOT BE RELEASED FROM THE TISSUE. THE MALFORMED STAPLES WERE DETACHED, AND THE TISSUE WAS RELEASED FROM THE CARTRIDGE. THERE WAS NO ISSUE WITH THE LEAK TEST AND NO ADDITIONAL FIXATION NEEDED.
05-12-2023 | Malfunction | covidien lp llc north haven | 01-01-2024 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, BRONCHIAL RESECTION, ONE STAPLE NEEDLE DID NOT GET SEPARATED FROM THE CARTRIDGE AND HAS POOR STAPLE FORMATION. THE STAPLE LINE WAS ALSO INCOMPLETE CENTRALLY. THEREFORE, THE CARTRIDGE COULD NOT BE RELEASED FROM THE TISSUE. THE MALFORMED STAPLES WERE DETACHED, AND THE TISSUE WAS RELEASED FROM THE CARTRIDGE. THERE WAS NO ISSUE WITH THE LEAK TEST AND NO ADDITIONAL FIXATION NEEDED.
01-01-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE FAULTY STAPLER DIDN¿T STAPLE PROPERLY AND HE NEEDED TO USE EXTRA CLIPS TO SEAL THE LINE. NO PATIENT HARM REPORTED.
20-11-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | CARTRIDGE FINAL ASSY 45MM GREEN | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE THE SURGERY, NOTED THERE ARE FOREIGN MATTERS IN THE PACKING. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
07-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ANTERIOR RESECTION, THE JAWS DID NOT OPEN AFTER THE DEVICE WAS INSERTED THROUGH THE TROCAR WITH THE JAWS CLOSED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
02-01-2024 | Malfunction | covidien lp llc north haven | 02-01-2024 | UNKNOWN STAPLING DEVICE | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE STAPLER DID NOT APPLY ANY STAPLES.
02-05-2024 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE STAPLER DID NOT APPLY ANY STAPLES.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RECTAL RESECTION, AT THE TIME OF RECTAL TRANSECTION, WHEN STARTED THE INTRAPERITONEAL APPROACH WITH THE JAWS NOT ARTICULATING, IN A STRAIGHT STATE, THE STATE OF THE RELOAD OPENING WAS WEAK. THE DOCTOR THEN TOOK THE DEVICE OUTSIDE THE SURGICAL FIELD THEN CHECKED THE OPENING AND CLOSING WITH FINGER, IT WAS NOTICED THAT THE SPRING OF OPENING AND CLOSING SEEMED TO BE STRONGER THAN USUAL. THERE WAS NO ISSUE WITH THE FIRING EVEN THOUGH THERE WAS SOME DEVIATION IN JAW OPENING AND CLOSING. ANOTHER OBJECT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RECTAL RESECTION, AT THE TIME OF RECTAL TRANSECTION, WHEN STARTED THE INTRAPERITONEAL APPROACH WITH THE JAWS NOT ARTICULATING, IN A STRAIGHT STATE, THE STATE OF THE RELOAD OPENING WAS WEAK. THE DOCTOR THEN TOOK THE DEVICE OUTSIDE THE SURGICAL FIELD THEN CHECKED THE OPENING AND CLOSING WITH FINGER, IT WAS NOTICED THAT THE SPRING OF OPENING AND CLOSING SEEMED TO BE STRONGER THAN USUAL. THERE WAS NO ISSUE WITH THE FIRING EVEN THOUGH THERE WAS SOME DEVIATION IN JAW OPENING AND CLOSING. ANOTHER OBJECT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RECTAL RESECTION, AT THE TIME OF RECTAL TRANSECTION, WHEN STARTED THE INTRAPERITONEAL APPROACH WITH THE JAWS NOT ARTICULATING, IN A STRAIGHT STATE, THE STATE OF THE RELOAD OPENING WAS WEAK. THE DOCTOR THEN TOOK THE DEVICE OUTSIDE THE SURGICAL FIELD THEN CHECKED THE OPENING AND CLOSING WITH FINGER, IT WAS NOTICED THAT THE SPRING OF OPENING AND CLOSING SEEMED TO BE STRONGER THAN USUAL. THERE WAS NO ISSUE WITH THE FIRING EVEN THOUGH THERE WAS SOME DEVIATION IN JAW OPENING AND CLOSING. ANOTHER OBJECT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RECTAL RESECTION, AT THE TIME OF RECTAL TRANSECTION, WHEN STARTED THE INTRAPERITONEAL APPROACH WITH THE JAWS NOT ARTICULATING, IN A STRAIGHT STATE, THE STATE OF THE RELOAD OPENING WAS WEAK. THE DOCTOR THEN TOOK THE DEVICE OUTSIDE THE SURGICAL FIELD THEN CHECKED THE OPENING AND CLOSING WITH FINGER, IT WAS NOTICED THAT THE SPRING OF OPENING AND CLOSING SEEMED TO BE STRONGER THAN USUAL. THERE WAS NO ISSUE WITH THE FIRING EVEN THOUGH THERE WAS SOME DEVIATION IN JAW OPENING AND CLOSING. ANOTHER OBJECT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Physical Resistance/Sticking | A0509 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RECTAL RESECTION, AT THE TIME OF RECTAL TRANSECTION, WHEN STARTED THE INTRAPERITONEAL APPROACH WITH THE JAWS NOT ARTICULATING, IN A STRAIGHT STATE, THE STATE OF THE RELOAD OPENING WAS WEAK. THE DOCTOR THEN TOOK THE DEVICE OUTSIDE THE SURGICAL FIELD THEN CHECKED THE OPENING AND CLOSING WITH FINGER, IT WAS NOTICED THAT THE SPRING OF OPENING AND CLOSING SEEMED TO BE STRONGER THAN USUAL. THERE WAS NO ISSUE WITH THE FIRING EVEN THOUGH THERE WAS SOME DEVIATION IN JAW OPENING AND CLOSING. ANOTHER OBJECT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG VOLUME REDUCTION SURGERY, IN LUNG TISSUE ANASTOMOSIS, THE STAPLES BURST OUT AND COULD NOT BE FORMED. THE STAPLE LINE WAS DISTALLY INCOMPLETE. THE ISSUE OCCURRED ON TWO RELOADS. REINFORCEMENT SUTURE WAS DONE TO COMPLETE THE CASE.
05-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG VOLUME REDUCTION SURGERY, IN LUNG TISSUE ANASTOMOSIS, THE STAPLES BURST OUT AND COULD NOT BE FORMED. THE STAPLE LINE WAS DISTALLY INCOMPLETE. THE ISSUE OCCURRED ON TWO RELOADS. REINFORCEMENT SUTURE WAS DONE TO COMPLETE THE CASE.
05-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG VOLUME REDUCTION SURGERY, IN LUNG TISSUE ANASTOMOSIS, THE STAPLES BURST OUT AND COULD NOT BE FORMED. THE STAPLE LINE WAS DISTALLY INCOMPLETE. THE ISSUE OCCURRED ON TWO RELOADS. REINFORCEMENT SUTURE WAS DONE TO COMPLETE THE CASE.
05-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG VOLUME REDUCTION SURGERY, IN LUNG TISSUE ANASTOMOSIS, THE STAPLES BURST OUT AND COULD NOT BE FORMED. THE STAPLE LINE WAS DISTALLY INCOMPLETE. THE ISSUE OCCURRED ON TWO RELOADS. REINFORCEMENT SUTURE WAS DONE TO COMPLETE THE CASE.
10-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 02-01-2024 | OPTIFIX STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN INGUINAL HERNIA REPAIR THE OPTIFIX STRAIGHT WAS USED TO FIXATE THE 3DMAX LIGHT MESH. IT WAS REPORTED THAT THE DEVICE STOPPED DEPLOYING AFTER THE THIRD AND FOURTH FASTENER. IT WAS ALSO REPORTED THAT THE FIRST TWO FASTENERS WERE ABLE TO FIXATE THE MESH AND IT IS UNKNOWN IF LOOSE FASTENERS WERE REMOVED FROM THE PATIENT'S BODY. IT WAS REPORTED THAT THE SURGEON DRY FIRED THE DEVICE TO CHECK IF THE PROBLEM RE-OCCURRED AND USED ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROTOMY WITH PULMONARY RESECTION, WHILE ON PARENCHYMA RESECTION, THERE WAS A DAMAGED TO THE YELLOW TAB PROTRUSION. IT WAS NOTED THAT A YELLOW FOREIGN OBJECT WAS VISUALLY CONFIRMED ON THE PATIENT'S ABDOMINAL CAVITY WHEN PULMONARY RESECTION WAS DONE. BASED ON THE MISSING YELLOW TAB PROTRUSION AND THE SIZE OF THE FOREIGN BODY, IT APPEARS TO BE A FRAGMENT OF THE YELLOW TAB PROTRUSION. THE FRAGMENT OF THE YELLOW TAB WAS THEN COLLECTED WITH FORCEPS. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, WHEN THE SURGEON TRIED TO FIRE THE DEVICE, IT DID NOT WORK. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC COLORECTAL SURGERY, THE STAPLER WAS UNABLE TO FIRE AFTER INSTALLING THE RELOAD. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-08-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY ESOPHAGECTOMY, DURING SETTING UP THE DEVICE, A POWER HANDLE ERROR WITH A RED CIRCLE AND A SLASH APPEARED DURING WHEN IT WAS INSERTED TO THE SHELL, AND THE SHELL ALSO HAVE A CONNECTION ERROR WHICH WAS SHELL NOT RECOGNIZED. A NEW POWERED HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE WAS LEFT IN A LOW VOLTAGE STATE FOR A SHORT PERIOD OF TIME AND ENTERED GRACEFUL SHUTDOWN AS A RESULT OF THE LOW VOLTAGE.
10-08-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Imprecision | A0803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY ESOPHAGECTOMY, DURING SETTING UP THE DEVICE, A POWER HANDLE ERROR WITH A RED CIRCLE AND A SLASH APPEARED DURING WHEN IT WAS INSERTED TO THE SHELL, AND THE SHELL ALSO HAVE A CONNECTION ERROR WHICH WAS SHELL NOT RECOGNIZED. A NEW POWERED HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE WAS LEFT IN A LOW VOLTAGE STATE FOR A SHORT PERIOD OF TIME AND ENTERED GRACEFUL SHUTDOWN AS A RESULT OF THE LOW VOLTAGE.
10-08-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY ESOPHAGECTOMY, DURING SETTING UP THE DEVICE, A POWER HANDLE ERROR WITH A RED CIRCLE AND A SLASH APPEARED DURING WHEN IT WAS INSERTED TO THE SHELL, AND THE SHELL ALSO HAVE A CONNECTION ERROR WHICH WAS SHELL NOT RECOGNIZED. A NEW POWERED HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE WAS LEFT IN A LOW VOLTAGE STATE FOR A SHORT PERIOD OF TIME AND ENTERED GRACEFUL SHUTDOWN AS A RESULT OF THE LOW VOLTAGE.
10-10-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE OF THE DEVICE, THE HANDLE IS NOT TURNING ON. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE WAS LEFT IN A LOW VOLTAGE STATE FOR A SHORT PERIOD OF TIME AND ENTERED GRACEFUL SHUTDOWN AS A RESULT OF THE LOW VOLTAGE.
10-10-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE OF THE DEVICE, THE HANDLE IS NOT TURNING ON. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE HANDLE WAS LEFT IN A LOW VOLTAGE STATE FOR A SHORT PERIOD OF TIME AND ENTERED GRACEFUL SHUTDOWN AS A RESULT OF THE LOW VOLTAGE.
28-11-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC DISTAL GASTRECTOMY, WHEN ATTEMPTING TO CLAMP THE TISSUE, THE USER FELT THAT THE RELOAD WAS NOT CLAMPING PROPERLY AND THE USE WAS DISCONTINUED. A NEW RELOAD WAS TAKEN OUT AND IT COULD BE USED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, AT THE TIME OF STAPLING THE SLEEVE, THE DEVICE DID NOT CLOSE PROPERLY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. THE SURGEON USED A NEW RELOAD WITH THE SAME POWERED DEVICES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE-PORT THORACOSCOPIC LEFT LOWER LOBECTOMY PROCEDURE, THE STAPLES FIRED ALL THE WAY BUT THE DISTAL STAPLES WERE INCOMPLETE AND SOME WERE POORLY FORMED. TO RESOLVE THE ISSUE, THE SURGEON USED SUTURE TO REINFORCE THE STAPLE LINE.
26-10-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LEFT HEMICOLECTOMY, AT THE TIME OF FIRING, THE RELOAD DOES NOT FINISH CLOSING, AS THE SIGNAL INDICATOR MARKED TISSUE SUITABLE FOR THAT LOAD, IT IS REOPENED, AND WHEN ATTEMPTING TO TAKE TISSUE AGAIN, IT NO LONGER CLOSES. THE RELOAD WAS REPLACED TO FIX THE ISSUE. THERE WAS NO PATIENT INJURY.
26-10-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LEFT HEMICOLECTOMY, AT THE TIME OF FIRING, THE RELOAD DOES NOT FINISH CLOSING, AS THE SIGNAL INDICATOR MARKED TISSUE SUITABLE FOR THAT LOAD, IT WAS REOPENED, AND WHEN ATTEMPTING TO TAKE TISSUE AGAIN, IT NO LONGER CLOSES. THE RELOAD WAS REPLACED TO FIX THE ISSUE. THERE WAS NO PATIENT INJURY.
26-10-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LEFT HEMICOLECTOMY, AT THE TIME OF FIRING, THE RELOAD DOES NOT FINISH CLOSING, AS THE SIGNAL INDICATOR MARKED TISSUE SUITABLE FOR THAT LOAD, IT WAS REOPENED, AND WHEN ATTEMPTING TO TAKE TISSUE AGAIN, IT NO LONGER CLOSES. THE RELOAD WAS REPLACED TO FIX THE ISSUE. THERE WAS NO PATIENT INJURY.
01-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PARTIAL LUNG RESECTION PROCEDURE, AT THE TIME OF THE 5TH TO 6TH FIRING OF LUNG PARENCHYMA RESECTION, THE SURGEON ATTEMPTED TO RETURN THE BLACK RETURNED KNOB WHILE THE RELOAD WAS ARTICULATED AT THE MAXIMUM. HOWEVER, THE BLACK RETURN KNOB WAS STIFF AND DID NOT RETURN. WHEN THE SURGEON RETURNED THE ARTICULATION TO STRAIGHT, THE BLACK RETURN KNOB WAS ABLE TO BE PULLED LITTLE BY LITTLE, AND FINALLY, THE JAWS WERE OPENED AND RELEASED. THE STAPLE LINE WAS THEN CHECKED AND NO PROBLEM WITH THE FORMATION AND NO LEAK WERE NOTED. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT THE VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
01-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) PARTIAL LUNG RESECTION PROCEDURE, AT THE TIME OF THE 5TH TO 6TH FIRING OF LUNG PARENCHYMA RESECTION, THE SURGEON ATTEMPTED TO RETURN THE BLACK RETURNED KNOB WHILE THE RELOAD WAS ARTICULATED AT THE MAXIMUM. HOWEVER, THE BLACK RETURN KNOB WAS STIFF AND DID NOT RETURN. WHEN THE SURGEON RETURNED THE ARTICULATION TO STRAIGHT, THE BLACK RETURN KNOB WAS ABLE TO BE PULLED LITTLE BY LITTLE, AND FINALLY, THE JAWS WERE OPENED AND RELEASED. THE STAPLE LINE WAS THEN CHECKED AND NO PROBLEM WITH THE FORMATION AND NO LEAK WERE NOTED. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION OF THE DEVICE FOUND THAT THE VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
12-11-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A REPAYMENT SURGERY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | CIRCULAR MECH STD 29MM | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION OPERATION IN LV, AFTER THE CIRCULAR SUTURE HAD BEEN USED CORRECTLY, THE INSTRUMENTALIST WAS NOT ABLE TO EXTRACT THE DEVICE.AFTER SEVERAL ATTEMPTS, WE WERE ABLE TO REMOVE IT ABOVE SUTURE RHYME, BUT ON THE PROXIMAL STUMP, THERE WAS A LACERATION. THE DAMAGE HAS BEEN REPAIRED, AND THE NECESSARY POST-OPERATIVE CHECKS HAVE BEEN CARRIED OUT.
14-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | CIRCULAR MECH STD 29MM | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION OPERATION IN LV, AFTER THE CIRCULAR SUTURE HAD BEEN USED CORRECTLY, THE INSTRUMENTALIST WAS NOT ABLE TO EXTRACT THE DEVICE.AFTER SEVERAL ATTEMPTS, WE WERE ABLE TO REMOVE IT ABOVE SUTURE RHYME, BUT ON THE PROXIMAL STUMP, THERE WAS A LACERATION. THE DAMAGE HAS BEEN REPAIRED, AND THE NECESSARY POST-OPERATIVE CHECKS HAVE BEEN CARRIED OUT.
14-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | CIRCULAR MECH STD 29MM | Ejection Problem | A0503 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION OPERATION IN LV, AFTER THE CIRCULAR SUTURE HAD BEEN USED CORRECTLY, THE INSTRUMENTALIST WAS NOT ABLE TO EXTRACT THE DEVICE.AFTER SEVERAL ATTEMPTS, WE WERE ABLE TO REMOVE IT ABOVE SUTURE RHYME, BUT ON THE PROXIMAL STUMP, THERE WAS A LACERATION. THE DAMAGE HAS BEEN REPAIRED, AND THE NECESSARY POST-OPERATIVE CHECKS HAVE BEEN CARRIED OUT.
14-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | CIRCULAR MECH STD 29MM | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION OPERATION IN LV, AFTER THE CIRCULAR SUTURE HAD BEEN USED CORRECTLY, THE INSTRUMENTALIST WAS NOT ABLE TO EXTRACT THE DEVICE.AFTER SEVERAL ATTEMPTS, WE WERE ABLE TO REMOVE IT ABOVE SUTURE RHYME, BUT ON THE PROXIMAL STUMP, THERE WAS A LACERATION. THE DAMAGE HAS BEEN REPAIRED, AND THE NECESSARY POST-OPERATIVE CHECKS HAVE BEEN CARRIED OUT.
13-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | 75 MM NTLC SELECTABLE RELOAD | Product Quality Problem | A0201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN STOMA CLOSURE, BEFORE OPENING THE PACKAGE, IT WAS FOUND THAT THERE WAS A BLACK DREGS-LIKE SUBSTANCE INSIDE THE PACKAGE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
13-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | 75 MM NTLC SELECTABLE RELOAD | Product Quality Problem | A0201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN STOMA CLOSURE, BEFORE OPENING THE PACKAGE, IT WAS FOUND THAT THERE WAS A BLACK DREGS-LIKE SUBSTANCE INSIDE THE PACKAGE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
05-01-2024 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THERE WAS A SIZEABLE GAP BETWEEN THE STAPLE CARTRIDGE AND ANVIL.
05-01-2024 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THERE WAS A SIZEABLE GAP BETWEEN THE STAPLE CARTRIDGE AND ANVIL.
12-01-2024 | Malfunction | covidien lp llc north haven | 02-01-2024 | UNKNOWN STAPLING DEVICE | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THERE WAS A SIZEABLE GAP BETWEEN THE STAPLE CARTRIDGE AND ANVIL.
12-01-2024 | Malfunction | covidien lp llc north haven | 02-01-2024 | UNKNOWN STAPLING DEVICE | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THERE WAS A SIZEABLE GAP BETWEEN THE STAPLE CARTRIDGE AND ANVIL.
15-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE IS NOT WORKING TO FIRE.
14-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN ANTERIOR RESECTION, THE STAPLE WAS STAPLED ON THE FRONT WALL SIDE, BUT IT COULD NOT BE STAPLED ON THE BACK WALL SIDE AT 1ST FIRING. THE DEVICE WAS USED ON THE RECTUM. ADDITIONAL CUT WAS PERFORMED, AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN ANTERIOR RESECTION, THE STAPLE WAS STAPLED ON THE FRONT WALL SIDE, BUT IT COULD NOT BE STAPLED ON THE BACK WALL SIDE AT 1ST FIRING. THE DEVICE WAS USED ON THE RECTUM. ADDITIONAL CUT WAS PERFORMED, AND ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-01-2024 | Malfunction | covidien lp llc north haven | 02-01-2024 | UNKNOWN EGIA ADAPTER | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THERE WAS A SIZEABLE GAP BETWEEN THE STAPLE CARTRIDGE AND ANVIL.
12-01-2024 | Malfunction | covidien lp llc north haven | 02-01-2024 | UNKNOWN EGIA ADAPTER | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THERE WAS A SIZEABLE GAP BETWEEN THE STAPLE CARTRIDGE AND ANVIL.
15-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | PROXIMATE*ILS CURVED CIR STAPL | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DISTAL STAPLES WERE POORLY FORMED AND INCOMPLETE. THE STAPLE LINE WAS REINFORCED WITH SUTURE TO FIX THE ISSUE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE.
05-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRECTOMY, THE SURGEON ATTEMPTED TO FIRE THE DEVICE, BUT THE DEVICE DID NOT WORK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE WAS UNFIRED.
16-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | PROXIMATE*PPH PROCEDURE SET | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY THE PPH03 THE STAPLE PINS WERE NOT LOGGED IN AFTER THE KNIFE CUT THE HEMORRHOID.
04-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACIC SURGERY, ON THE 6TH STAPLE FIRING WITH THE RELOAD, IT CRACKED AND THE SHAFT WAS SEPARATED FROM THE HANDLE. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACIC SURGERY, ON THE 6TH STAPLE FIRING WITH THE RELOAD, IT CRACKED AND THE SHAFT WAS SEPARATED FROM THE HANDLE. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING AN UNKNOWN PROCEDURE A CDH29P DID NOT FIRE. THERE WAS NO SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
15-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DISTAL GASTRECTOMY THE DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2023 | Malfunction | ethicon endo-surgery | 02-01-2024 | PROXIMATE*PPH PROCEDURE SET | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY THE PPH03 THE STAPLE PINS WERE NOT LOGGED IN AFTER THE KNIFE CUT THE HEMORRHOID.
01-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | GIA | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A CHOLECYSTECTOMY OPEN PROCEDURE, THE STAPLE CUT IN THE SHOT BUT DID NOT PERFORM THE STAPLES. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE.
01-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A CHOLECYSTECTOMY OPEN PROCEDURE, THE STAPLE CUT IN THE SHOT BUT DID NOT PERFORM THE STAPLES. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE.
28-11-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, WHILE IN THE PULMONARY ARTERY AND VEIN, THE DEVICE FIRED BUT THERE WAS AVERAGE BLEEDING ALONG THE ENTIRE STAPLE LINE AFTER SHOOTING BOTH VEIN AND ARTERY. THE BLEEDING WAS STOPPED BY COMPRESSING THE VESSELS FOR 15 TO 20 MINUTES. THE ISSUE RESULTED TO TISSUE DAMAGE.
28-11-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, WHILE IN THE PULMONARY ARTERY AND VEIN, THE DEVICE FIRED BUT THERE WAS AVERAGE BLEEDING ALONG THE ENTIRE STAPLE LINE AFTER SHOOTING BOTH VEIN AND ARTERY. THE BLEEDING WAS STOPPED BY COMPRESSING THE VESSELS FOR 15 TO 20 MINUTES. THE ISSUE RESULTED TO TISSUE DAMAGE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC AND OPEN NEPHROURETERECTOMY WITH RESECTION OF BLADDER CUFF AND REGIONAL LYMPH NODE DISSECTION, DURING ARTERY LIGATION, THERE WAS STAPLE LINE BLEEDING. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE TO CONTROL THE BLEEDING BOTH DURING THE LAPAROSCOPIC PORTION AND OPEN PROCEDURE PORTION. THE ISSUE OCCURRED ON A TOTAL OF SIX RELOADS. CLIPS AND SEALANT FROM ANOTHER MANUFACTURER WERE USED TO STOP THE BLEEDING.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC AND OPEN NEPHROURETERECTOMY WITH RESECTION OF BLADDER CUFF AND REGIONAL LYMPH NODE DISSECTION, DURING ARTERY LIGATION, THERE WAS STAPLE LINE BLEEDING. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE TO CONTROL THE BLEEDING BOTH DURING THE LAPAROSCOPIC PORTION AND OPEN PROCEDURE PORTION. THE ISSUE OCCURRED ON A TOTAL OF SIX RELOADS. CLIPS AND SEALANT FROM ANOTHER MANUFACTURER WERE USED TO STOP THE BLEEDING.
29-11-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC AND OPEN NEPHROURETERECTOMY WITH RESECTION OF BLADDER CUFF AND REGIONAL LYMPH NODE DISSECTION, DURING ARTERY LIGATION, THERE WAS STAPLE LINE BLEEDING. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE TO CONTROL THE BLEEDING BOTH DURING THE LAPAROSCOPIC PORTION AND OPEN PROCEDURE PORTION. THE ISSUE OCCURRED ON A TOTAL OF SIX RELOADS. CLIPS AND SEALANT FROM ANOTHER MANUFACTURER WERE USED TO STOP THE BLEEDING.
29-11-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | UNKNOWN STAPLING DEVICE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC AND OPEN NEPHROURETERECTOMY WITH RESECTION OF BLADDER CUFF AND REGIONAL LYMPH NODE DISSECTION, DURING ARTERY LIGATION, THERE WAS STAPLE LINE BLEEDING. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE TO CONTROL THE BLEEDING BOTH DURING THE LAPAROSCOPIC PORTION AND OPEN PROCEDURE PORTION. THE ISSUE OCCURRED ON A TOTAL OF SIX RELOADS. CLIPS AND SEALANT FROM ANOTHER MANUFACTURER WERE USED TO STOP THE BLEEDING.
05-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC TAIL RESECTION, THE REINFORCED RELOAD WAS CONNECTED TO THE POWER HANDLE AND ADAPTER, AND FIRING WAS PERFORMED UNTIL THE END. HOWEVER, EVEN AFTER ATTEMPTING TO OPEN THE JAWS OF THE CARTRIDGE, THE JAWS DIDNOT OPEN. A MANUAL ADAPTER TOOL WAS USED, AND THE JAWS WERE RELEASED, AND WAS REMOVED FROM THE PANCREATIC PARENCHYMA. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC AND OPEN NEPHROURETERECTOMY WITH RESECTION OF BLADDER CUFF AND REGIONAL LYMPH NODE DISSECTION, DURING ARTERY LIGATION, THERE WAS STAPLE LINE BLEEDING. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE TO CONTROL THE BLEEDING BOTH DURING THE LAPAROSCOPIC PORTION AND OPEN PROCEDURE PORTION. THE ISSUE OCCURRED ON A TOTAL OF SIX RELOADS. CLIPS AND SEALANT FROM ANOTHER MANUFACTURER WERE USED TO STOP THE BLEEDING.
04-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE WAS ABLE TO CLOSE. HOWEVER, IT DID NOT OPEN. THERE WAS NO PATIENT INVOLVED.
05-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC TAIL RESECTION, THE REINFORCED RELOAD WAS CONNECTED TO THE POWER HANDLE AND ADAPTER, AND FIRING WAS PERFORMED UNTIL THE END. HOWEVER, EVEN AFTER ATTEMPTING TO OPEN THE JAWS OF THE CARTRIDGE, THE JAWS DIDNOT OPEN. A MANUAL ADAPTER TOOL WAS USED, AND THE JAWS WERE RELEASED, AND WAS REMOVED FROM THE PANCREATIC PARENCHYMA. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC AND OPEN NEPHROURETERECTOMY WITH RESECTION OF BLADDER CUFF AND REGIONAL LYMPH NODE DISSECTION, DURING ARTERY LIGATION, THERE WAS STAPLE LINE BLEEDING. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES OR MORE TO CONTROL THE BLEEDING BOTH DURING THE LAPAROSCOPIC PORTION AND OPEN PROCEDURE PORTION. THE ISSUE OCCURRED ON A TOTAL OF SIX RELOADS. CLIPS AND SEALANT FROM ANOTHER MANUFACTURER WERE USED TO STOP THE BLEEDING.
04-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE WAS ABLE TO CLOSE. HOWEVER, IT DID NOT OPEN. THERE WAS NO PATIENT INVOLVED.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, AT THE TIME OF STAPLING THE SLEEVE, THE DEVICE DID NOT CLOSE PROPERLY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. THE SURGEON USED A NEW RELOAD WITH THE SAME POWERED DEVICES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, AT THE TIME OF STAPLING THE SLEEVE, THE DEVICE DID NOT CLOSE PROPERLY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. THE SURGEON USED A NEW RELOAD WITH THE SAME POWERED DEVICES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE STAPLER SEAM WAS NOT MOVING FORWARD. THERE WAS NO PATIENT INVOLVEMENT.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SLEEVE PROCEDURE, AT THE TIME OF STAPLING THE SLEEVE, THE DEVICE DID NOT CLOSE PROPERLY. ANOTHER RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. THE SURGEON USED A NEW RELOAD WITH THE SAME POWERED DEVICES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRECTOMY, WHILE IN THE GASTRIC TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL TRANSIT RECONSTRUCTION PROCEDURE, AT THE TIME OF TESTING, EVERYTHING WORKED; HO WEVER, WHEN OPENING THE LOAD TO TAKE THE TISSUE, THE USER HEARD A NOISE BETWEEN THE RELOAD AND THE SHAFT. THE RELOAD DID NOT OPEN. WHEN TRYING TO REMOVE THE LOAD, IT WAS OBSERVED THAT THE UNION OF BOTH WAS LOOSE, AND THE LOAD COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW RELOAD AND ADAPTER WERE USED WITH THE SAME HANDLE AND SHELL. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC TAIL RESECTION, THE REINFORCED RELOAD WAS CONNECTED TO THE POWER HANDLE AND ADAPTER, AND FIRING WAS PERFORMED UNTIL THE END. HOWEVER, EVEN AFTER ATTEMPTING TO OPEN THE JAWS OF THE CARTRIDGE, THE JAWS DID NOT OPEN. A MANUAL ADAPTER TOOL WAS USED, AND THE JAWS WERE RELEASED, AND WAS REMOVED FROM THE PANCREATIC PARENCHYMA. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRECTOMY, THE SURGEON ATTEMPTED TO FIRE THE DEVICE, BUT THE DEVICE DID NOT WORK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRECTOMY, THE SURGEON ATTEMPTED TO FIRE THE DEVICE, BUT THE DEVICE DID NOT WORK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL TRANSIT RECONSTRUCTION PROCEDURE, AT THE TIME OF TESTING, EVERYTHING WORKED; HOWEVER, WHEN OPENING THE LOAD TO TAKE THE TISSUE, THE USER HEARD A NOISE BETWEEN THE RELOAD AND THE SHAFT. THE RELOAD DID NOT OPEN. WHEN TRYING TO REMOVE THE LOAD, IT WAS OBSERVED THAT THE UNION OF BOTH WAS LOOSE, AND THE LOAD COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW RELOAD AND ADAPTER WERE USED WITH THE SAME HANDLE AND SHELL. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL TRANSIT RECONSTRUCTION PROCEDURE, AT THE TIME OF TESTING, EVERYTHING WORKED; HOWEVER, WHEN OPENING THE LOAD TO TAKE THE TISSUE, THE USER HEARD A NOISE BETWEEN THE RELOAD AND THE SHAFT. THE RELOAD DID NOT OPEN. WHEN TRYING TO REMOVE THE LOAD, IT WAS OBSERVED THAT THE UNION OF BOTH WAS LOOSE, AND THE LOAD COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW RELOAD AND ADAPTER WERE USED WITH THE SAME HANDLE AND SHELL. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL TRANSIT RECONSTRUCTION PROCEDURE, AT THE TIME OF TESTING, EVERYTHING WORKED; HOWEVER, WHEN OPENING THE LOAD TO TAKE THE TISSUE, THE USER HEARD A NOISE BETWEEN THE RELOAD AND THE SHAFT. THE RELOAD DID NOT OPEN. WHEN TRYING TO REMOVE THE LOAD, IT WAS OBSERVED THAT THE UNION OF BOTH WAS LOOSE, AND THE LOAD COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW RELOAD AND ADAPTER WERE USED WITH THE SAME HANDLE AND SHELL. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL TRANSIT RECONSTRUCTION PROCEDURE, AT THE TIME OF TESTING, EVERYTHING WORKED; HOWEVER, WHEN OPENING THE LOAD TO TAKE THE TISSUE, THE USER HEARD A NOISE BETWEEN THE RELOAD AND THE SHAFT. THE RELOAD DID NOT OPEN. WHEN TRYING TO REMOVE THE LOAD, IT WAS OBSERVED THAT THE UNION OF BOTH WAS LOOSE, AND THE LOAD COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW RELOAD AND ADAPTER WERE USED WITH THE SAME HANDLE AND SHELL. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL TRANSIT RECONSTRUCTION PROCEDURE, AT THE TIME OF TESTING, EVERYTHING WORKED; HO WEVER, WHEN OPENING THE LOAD TO TAKE THE TISSUE, THE USER HEARD A NOISE BETWEEN THE RELOAD AND THE SHAFT. THE RELOAD DID NOT OPEN. WHEN TRYING TO REMOVE THE LOAD, IT WAS OBSERVED THAT THE UNION OF BOTH WAS LOOSE, AND THE LOAD COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW RELOAD AND ADAPTER WERE USED WITH THE SAME HANDLE AND SHELL. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL TRANSIT RECONSTRUCTION PROCEDURE, AT THE TIME OF TESTING, EVERYTHING WORKED; HO WEVER, WHEN OPENING THE LOAD TO TAKE THE TISSUE, THE USER HEARD A NOISE BETWEEN THE RELOAD AND THE SHAFT. THE RELOAD DID NOT OPEN. WHEN TRYING TO REMOVE THE LOAD, IT WAS OBSERVED THAT THE UNION OF BOTH WAS LOOSE, AND THE LOAD COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW RELOAD AND ADAPTER WERE USED WITH THE SAME HANDLE AND SHELL. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL TRANSIT RECONSTRUCTION PROCEDURE, AT THE TIME OF TESTING, EVERYTHING WORKED; HO WEVER, WHEN OPENING THE LOAD TO TAKE THE TISSUE, THE USER HEARD A NOISE BETWEEN THE RELOAD AND THE SHAFT. THE RELOAD DID NOT OPEN. WHEN TRYING TO REMOVE THE LOAD, IT WAS OBSERVED THAT THE UNION OF BOTH WAS LOOSE, AND THE LOAD COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW RELOAD AND ADAPTER WERE USED WITH THE SAME HANDLE AND SHELL. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL TRANSIT RECONSTRUCTION PROCEDURE, AT THE TIME OF TESTING, EVERYTHING WORKED; HO WEVER, WHEN OPENING THE LOAD TO TAKE THE TISSUE, THE USER HEARD A NOISE BETWEEN THE RELOAD AND THE SHAFT. THE RELOAD DID NOT OPEN. WHEN TRYING TO REMOVE THE LOAD, IT WAS OBSERVED THAT THE UNION OF BOTH WAS LOOSE, AND THE LOAD COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW RELOAD AND ADAPTER WERE USED WITH THE SAME HANDLE AND SHELL. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE DEVICE DID NOT FIRE. THE ISSUE OCCURRED IN TWO RELOADS. THERE WAS NO PATIENT INVOLVEMENT.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE PERFORMING ANASTOMOSIS, THE DEVICE LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON PHYSICALLY PULLED THE RELOAD OFF THE ADAPTER AND USED A COMPETITOR'S DEVICE TO STAPLE AND CUT AROUND THE RELOAD. THIS LED TO MORE TISSUE BEING TAKEN IN ORDER TO PERFORM THE ANASTOMOSIS.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | UNKNOWN ENDO GIA SULU | Human-Device Interface Problem | A22 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE PERFORMING ANASTOMOSIS, THE DEVICE LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON PHYSICALLY PULLED THE RELOAD OFF THE ADAPTER AND USED A COMPETITOR'S DEVICE TO STAPLE AND CUT AROUND THE RELOAD. THIS LED TO MORE TISSUE BEING TAKEN IN ORDER TO PERFORM THE ANASTOMOSIS.
28-11-2023 | Injury | intuitive surgical | 02-01-2024 | SUREFORM | Insufficient Information | E2401 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED RIGHT HEMICOLECTOMY, THE PATIENT EXPERIENCED A STAPLE LINE BLEED. THE SURGEON STATED THEIR OVERALL TECHNIQUE FOR THE PROCEDURE HAS NOT CHANGED. THE SURGEON SAID THEY PERFORM AN INTRACORPOREAL ANASTOMOSIS USING THE SUREFORM 60 STAPLER INSTRUMENT. THEN, TWO SUREFORM 60 BLUE RELOADS ARE USED TO DIVIDE THE BOWEL. OF RECENT, A SUREFORM 60 WHITE RELOAD IS USED FOR THE SIDE-TO-SIDE ANASTOMOSIS; HOWEVER, DUE TO ASSOCIATED CASE ISSUES, THE SURGEON OPTED TO USE A BLUE 60 RELOAD. AFTER THE COMPLETION OF THE FIRE, THE STAPLE LINE IS CONFIRMED TO BE IN GOOD CONDITION. THE SURGEON CHECKS THE ENTIRE ABDOMEN FOR ANY BLEEDING; THE PATIENT IS THEN CLOSED, AND THE PROCEDURE COMPLETED. HOWEVER, POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERAL SELF-LIMITED DARK BLOODY BOWEL MOVEMENTS. THE SURGEON BELIEVED THIS WAS A RESULT OF A STAPLE LINE LEAK FROM THE SIDE-TO-SIDE ANASTOMOSIS. FURTHER INFORMATION REGARDING ANY INTERVENTIONS PERFORMED OR THE PATIENT¿S CURRENT STATUS, IS UNKNOWN AT THIS TIME.
06-10-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC SURGERY, THE DEVICE DID NOT CLOSE COMPLETELY LEADING TO PARTIAL STAPLING. THE SURGEON USED ANOTHER HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-10-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC SURGERY, THE DEVICE DID NOT CLOSE COMPLETELY LEADING TO PARTIAL STAPLING. THE SURGEON USED ANOTHER HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-11-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT UPPER LOBECTOMY, THE CUTTING STAPLER WAS BLOCKED, RESULTING IN A FAILURE TO FIRE NORMALLY. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE PERFORMING ANASTOMOSIS, THE DEVICE LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON PHYSICALLY PULLED THE RELOAD OFF THE ADAPTER AND USED A COMPETITOR'S DEVICE TO STAPLE AND CUT AROUND THE RELOAD. THIS LED TO MORE TISSUE BEING TAKEN IN ORDER TO PERFORM THE ANASTOMOSIS.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE PERFORMING ANASTOMOSIS, THE DEVICE LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON PHYSICALLY PULLED THE RELOAD OFF THE ADAPTER AND USED A COMPETITOR'S DEVICE TO STAPLE AND CUT AROUND THE RELOAD. THIS LED TO MORE TISSUE BEING TAKEN IN ORDER TO PERFORM THE ANASTOMOSIS.
06-12-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | SIGNIA | Human-Device Interface Problem | A22 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE PERFORMING ANASTOMOSIS, THE DEVICE LOCKED ON TISSUE. TO RESOLVE THE ISSUE, THE SURGEON PHYSICALLY PULLED THE RELOAD OFF THE ADAPTER AND USED A COMPETITOR'S DEVICE TO STAPLE AND CUT AROUND THE RELOAD. THIS LED TO MORE TISSUE BEING TAKEN IN ORDER TO PERFORM THE ANASTOMOSIS.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLORECTAL SURGERY, AFTER REPLACING THE RELOAD, THE STAPLES BURST OUT WHEN FIRED. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ESOPHAGEAL CANCER RESECTION PROCEDURE, DURING THE DISSOCIATION AND FABRICATION OF THE TUBULAR STOMACH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER PURPLE RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 02-01-2024 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE INITIALLY FIRED BUT THEN FAILED TO COMPLETE THE FIRING CYCLE. ANOTHER OPEN STAPLER WAS USED BUT HAD THE SAME ISSUE. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE INITIALLY FIRED BUT THEN FAILED TO COMPLETE THE FIRING CYCLE. ANOTHER OPEN STAPLER WAS USED BUT HAD THE SAME ISSUE. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-11-2023 | Malfunction | covidien lp llc north haven | 02-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEPATECTOMY, DURING CUTTING OF THE ANATOMICAL PIECE, THE DEVICE FIRED UNTIL THE END BUT STAPLE LINE WAS INCOMPLETE DISTALLY. A SECOND RELOAD WAS FIRED TO RESOLVE THE ISSUE.
05-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LUNG CANCER SURGERY, WHEN AN ATTEMPT WAS MADE TO PERFORM FIRING, FIRING COULD NOT BE PERFORMED. IT WAS ALSO MENTIONED THAT THE STAPLE FROM THE PROXIMAL BASE WAS COMING OUT, SO THERE WAS A POSSIBILITY THAT IT HAD BEEN PRE-FIRED AT SOME POINT. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 02-01-2024 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LUNG CANCER SURGERY, WHEN AN ATTEMPT WAS MADE TO PERFORM FIRING, FIRING COULD NOT BE PERFORMED. IT WAS ALSO MENTIONED THAT THE STAPLE FROM THE PROXIMAL BASE WAS COMING OUT, SO THERE WAS A POSSIBILITY THAT IT HAD BEEN PRE-FIRED AT SOME POINT. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2023 | Malfunction | ethicon endo-surgery | 03-01-2024 | PROXIMATE*PPH PROCEDURE SET | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY THE PPH03 THE STAPLE PINS WERE NOT LOGGED IN AFTER THE KNIFE CUT THE HEMORRHOID.
13-10-2023 | Malfunction | ethicon endo-surgery | 03-01-2024 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-12-2023 | Malfunction | ethicon endo-surgery | 03-01-2024 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE GST RELOAD GOT STUCK IN-BETWEEN WHILE FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
07-12-2023 | Malfunction | ethicon endo-surgery | 03-01-2024 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE GST RELOAD GOT STUCK IN-BETWEEN WHILE FIRING. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Malfunction | ethicon endo-surgery | 03-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP APPENDECTOMY, THE DEVICE WAS NOT FIRING. THE ATTENDING PHYSICIAN WAS HAVING THE RESIDENT FIRE THE DEVICE. THEY WERE UNABLE TO FIRE THE DEVICE. CASE COMPLETED WITH A LIKE DEVICE. NO PATIENT HARM WAS REPORTED.
10-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | SIGNIA | Entrapment of Device | A150208 | Fistula | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACT THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE.
10-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | SIGNIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACT THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE.
10-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACT THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE.
26-09-2023 | Injury | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , AT THE TIME OF STAPLING OF THE STOMACH, THE STAPLER WAS OK ON THE FIRST FIRING BUT THEN IT COULD NOT SHOT WITH PURPLE RELOADS. IT WAS MENTIONED THAT AT LEAST 5 RELOADS HAD THE SAME ISSUE. IT SEEMS THAT THERE WAS SOMETHING INSIDE THE STAPLERS THAT GOT BROKEN SO ANOTHER STAPLER WAS TAKEN BUT THE SAME THING OCCURRED. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY AND THE PHYSICIAN USED ANOTHER STAPLER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THE PATIENT HOSPITALIZATION WAS EXTENDED FOR TWO DAYS AND THE PATIENT WAS HOSPITALIZED AGAIN FOR PAIN ISSUE.
26-09-2023 | Injury | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | Pain | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , AT THE TIME OF STAPLING OF THE STOMACH, THE STAPLER WAS OK ON THE FIRST FIRING BUT THEN IT COULD NOT SHOT WITH PURPLE RELOADS. IT WAS MENTIONED THAT AT LEAST 5 RELOADS HAD THE SAME ISSUE. IT SEEMS THAT THERE WAS SOMETHING INSIDE THE STAPLERS THAT GOT BROKEN SO ANOTHER STAPLER WAS TAKEN BUT THE SAME THING OCCURRED. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY AND THE PHYSICIAN USED ANOTHER STAPLER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THE PATIENT HOSPITALIZATION WAS EXTENDED FOR TWO DAYS AND THE PATIENT WAS HOSPITALIZED AGAIN FOR PAIN ISSUE.
26-09-2023 | Injury | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , AT THE TIME OF STAPLING OF THE STOMACH, THE STAPLER WAS OK ON THE FIRST FIRING BUT THEN IT COULD NOT SHOT WITH PURPLE RELOADS. IT WAS MENTIONED THAT AT LEAST 5 RELOADS HAD THE SAME ISSUE. IT SEEMS THAT THERE WAS SOMETHING INSIDE THE STAPLERS THAT GOT BROKEN SO ANOTHER STAPLER WAS TAKEN BUT THE SAME THING OCCURRED. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY AND THE PHYSICIAN USED ANOTHER STAPLER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THE PATIENT HOSPITALIZATION WAS EXTENDED FOR TWO DAYS AND THE PATIENT WAS HOSPITALIZED AGAIN FOR PAIN ISSUE.
26-09-2023 | Injury | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | Pain | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , AT THE TIME OF STAPLING OF THE STOMACH, THE STAPLER WAS OK ON THE FIRST FIRING BUT THEN IT COULD NOT SHOT WITH PURPLE RELOADS. IT WAS MENTIONED THAT AT LEAST 5 RELOADS HAD THE SAME ISSUE. IT SEEMS THAT THERE WAS SOMETHING INSIDE THE STAPLERS THAT GOT BROKEN SO ANOTHER STAPLER WAS TAKEN BUT THE SAME THING OCCURRED. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY AND THE PHYSICIAN USED ANOTHER STAPLER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THE PATIENT HOSPITALIZATION WAS EXTENDED FOR TWO DAYS AND THE PATIENT WAS HOSPITALIZED AGAIN FOR PAIN ISSUE.
26-09-2023 | Injury | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , AT THE TIME OF STAPLING OF THE STOMACH, THE STAPLER WAS OK ON THE FIRST FIRING BUT THEN IT COULD NOT SHOT WITH PURPLE RELOADS. IT WAS MENTIONED THAT AT LEAST 5 RELOADS HAD THE SAME ISSUE. IT SEEMS THAT THERE WAS SOMETHING INSIDE THE STAPLERS THAT GOT BROKEN SO ANOTHER STAPLER WAS TAKEN BUT THE SAME THING OCCURRED. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY AND THE PHYSICIAN USED ANOTHER STAPLER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THE PATIENT HOSPITALIZATION WAS EXTENDED FOR TWO DAYS AND THE PATIENT WAS HOSPITALIZED AGAIN FOR PAIN ISSUE.
26-09-2023 | Injury | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | Pain | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , AT THE TIME OF STAPLING OF THE STOMACH, THE STAPLER WAS OK ON THE FIRST FIRING BUT THEN IT COULD NOT SHOT WITH PURPLE RELOADS. IT WAS MENTIONED THAT AT LEAST 5 RELOADS HAD THE SAME ISSUE. IT SEEMS THAT THERE WAS SOMETHING INSIDE THE STAPLERS THAT GOT BROKEN SO ANOTHER STAPLER WAS TAKEN BUT THE SAME THING OCCURRED. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY AND THE PHYSICIAN USED ANOTHER STAPLER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THE PATIENT HOSPITALIZATION WAS EXTENDED FOR TWO DAYS AND THE PATIENT WAS HOSPITALIZED AGAIN FOR PAIN ISSUE.
26-09-2023 | Injury | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , AT THE TIME OF STAPLING OF THE STOMACH, THE STAPLER WAS OK ON THE FIRST FIRING BUT THEN IT COULD NOT SHOT WITH PURPLE RELOADS. IT WAS MENTIONED THAT AT LEAST 5 RELOADS HAD THE SAME ISSUE. IT SEEMS THAT THERE WAS SOMETHING INSIDE THE STAPLERS THAT GOT BROKEN SO ANOTHER STAPLER WAS TAKEN BUT THE SAME THING OCCURRED. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY AND THE PHYSICIAN USED ANOTHER STAPLER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THE PATIENT HOSPITALIZATION WAS EXTENDED FOR TWO DAYS AND THE PATIENT WAS HOSPITALIZED AGAIN FOR PAIN ISSUE.
26-09-2023 | Injury | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | Pain | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , AT THE TIME OF STAPLING OF THE STOMACH, THE STAPLER WAS OK ON THE FIRST FIRING BUT THEN IT COULD NOT SHOT WITH PURPLE RELOADS. IT WAS MENTIONED THAT AT LEAST 5 RELOADS HAD THE SAME ISSUE. IT SEEMS THAT THERE WAS SOMETHING INSIDE THE STAPLERS THAT GOT BROKEN SO ANOTHER STAPLER WAS TAKEN BUT THE SAME THING OCCURRED. THE PROCEDURE WAS CONVERTED TO OPEN SURGERY AND THE PHYSICIAN USED ANOTHER STAPLER FROM DIFFERENT COMPANY TO RESOLVE THE ISSUE. THE PATIENT HOSPITALIZATION WAS EXTENDED FOR TWO DAYS AND THE PATIENT WAS HOSPITALIZED AGAIN FOR PAIN ISSUE.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STOMACH STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE HANDLE CRACKED. THE SURGEON NO LONGER WANTED TO STAPLE FOLLOWING THE CRACK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STOMACH STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE HANDLE CRACKED. THE SURGEON NO LONGER WANTED TO STAPLE FOLLOWING THE CRACK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STOMACH STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE HANDLE CRACKED. THE SURGEON NO LONGER WANTED TO STAPLE FOLLOWING THE CRACK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-12-2023 | Malfunction | ethicon endo-surgery | 03-01-2024 | BLUE RELOAD FOR ECHELON 60 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SURGERY, THE STAPLER MALFUNCTIONED, CAUSING THE CARTRIDGE DAMAGED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
16-12-2023 | Malfunction | ethicon endo-surgery | 03-01-2024 | BLUE RELOAD FOR ECHELON 60 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SURGERY, THE STAPLER MALFUNCTIONED, CAUSING THE CARTRIDGE DAMAGED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
04-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORACOSCOPIC PNEUMONECTOMY, WHEN THE KNOB WAS PULLED BACK IN ORDER TO RETRACT THE KNIFE AFTER BULLA RESECTION, THE RETRACTION WAS HEAVY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORACOSCOPIC PNEUMONECTOMY, WHEN THE KNOB WAS PULLED BACK IN ORDER TO RETRACT THE KNIFE AFTER BULLA RESECTION, THE RETRACTION WAS HEAVY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-11-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROTOMY TOTAL GASTRECTOMY, DURING RESECTION, THE DEVICE WAS ATTACHED BUT THE DEVICE DID NOT FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHEN DETACHING THE PULMONARY VEINS, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE PRODUCT WAS REPLACED WITH THE SAME MODEL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHEN DETACHING THE PULMONARY VEINS, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE PRODUCT WAS REPLACED WITH THE SAME MODEL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHEN DETACHING THE PULMONARY VEINS, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE PRODUCT WAS REPLACED WITH THE SAME MODEL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-10-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANTERIOR RESECTION, BEFORE USING IN PROXIMAL RECTUM, AFTER PLACING THE SHELL, A ERROR MESSAGE OF A HANDLE AND A RED CIRCLE WITH A LINE GOING THROUGH IT WAS DISPLAYED AND THE HANDLE COULD NOT BE USED. TO RESOLVE THE ISSUE, A NEW HANDLE AND SHELL WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION FOUND THAT THE ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) R ESET/REPROGRAM FAILURES.
23-10-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANTERIOR RESECTION, BEFORE USING IN PROXIMAL RECTUM, AFTER PLACING THE SHELL, A ERROR MESSAGE OF A HANDLE AND A RED CIRCLE WITH A LINE GOING THROUGH IT WAS DISPLAYED AND THE HANDLE COULD NOT BE USED. TO RESOLVE THE ISSUE, A NEW HANDLE AND SHELL WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION FOUND THAT THE ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) R ESET/REPROGRAM FAILURES.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC COLORECTAL RESECTION, WHILE DETACHING THE COLON TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER 60 MM PURPLE RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STOMACH STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE HANDLE CRACKED. THE SURGEON NO LONGER WANTED TO STAPLE FOLLOWING THE CRACK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STOMACH STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE HANDLE CRACKED. THE SURGEON NO LONGER WANTED TO STAPLE FOLLOWING THE CRACK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | UNKNOWN EGIA TRI STAPLE | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORACOSCOPIC PNEUMONECTOMY, WHEN THE KNOB WAS PULLED BACK IN ORDER TO RETRACT THE KNIFE AFTER BULLA RESECTION, THE RETRACTION WAS HEAVY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | UNKNOWN EGIA TRI STAPLE | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORACOSCOPIC PNEUMONECTOMY, WHEN THE KNOB WAS PULLED BACK IN ORDER TO RETRACT THE KNIFE AFTER BULLA RESECTION, THE RETRACTION WAS HEAVY. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | SIGNIA | Entrapment of Device | A150208 | Fistula | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACTS THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE. THERE WAS A BLOOD LOSS AND THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
10-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | SIGNIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACTS THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE. THERE WAS A BLOOD LOSS AND THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
10-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACTS THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE. THERE WAS A BLOOD LOSS AND THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
10-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Entrapment of Device | A150208 | Fistula | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACT THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE.
10-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACT THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE.
10-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACT THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE.
10-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Fistula | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACT THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE.
10-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACT THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE.
10-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENOPANCREATECTOMY, AFTER FIRING, THE JAWS OF THE RELOAD DID NOT OPEN. MANUAL METHODS WERE USED TO RETRACT THE BLADE AND OPEN THE JAWS, BUT IT DID NOT WORK. AN ATTEMPT WAS ALSO MADE TO DO SO BY RESTARTING THE DEVICE, BUT IT DID NOT RESOLVE THE ISSUE. THE SURGEON HAD TO CUT ALL THE SURROUNDING TISSUE TO REMOVE THE RELOAD, WHICH CAUSED A TEAR IN THE PANCREATIC TISSUE AND A FISTULA. THE STAPLE LINE WAS ALSO INCOMPLETE.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STOMACH STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE HANDLE CRACKED. THE SURGEON NO LONGER WANTED TO STAPLE FOLLOWING THE CRACK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING STOMACH STAPLING, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE HANDLE CRACKED. THE SURGEON NO LONGER WANTED TO STAPLE FOLLOWING THE CRACK. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
24-10-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN CUTTING TO CLOSE THE GASTRIC BODY, THERE WAS POOR STAPLE FORMATION. THE BODY OF THE STOMACH WAS NOT COMPLETELY CLOSED, AND BLEEDING NOT MORE THAN 500CC OCCURRED. HAND SUTURING WAS PERFORMED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP DUE TO THICK TISSUE.
24-10-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN CUTTING TO CLOSE THE GASTRIC BODY, THERE WAS POOR STAPLE FORMATION. THE BODY OF THE STOMACH WAS NOT COMPLETELY CLOSED, AND BLEEDING NOT MORE THAN 500CC OCCURRED. HAND SUTURING WAS PERFORMED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP DUE TO THICK TISSUE.
05-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LAPAROTOMY, PRIOR TO RECTAL RESECTION, WHEN THE 60MM RELOAD WAS CONNECTED TO THE HANDLE WITH NO ISSUE, THE FIRING COULD NOT BE PERFORMED EVEN WHEN THE HANDLE WAS SQUEEZED. IT WAS ATTEMPTED SEVERAL TIMES, BUT THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-09-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE UNDER THE LEFT LUNG LOBE, WHEN THE SURGEON SQUEEZED THE HANDLE ONCE, THERE WAS A LOUD CRACK, THE HANDLE COULD NOT BE SQUEEZED AND THE THE FIRING STOPPED. THE STAPLER WAS STUCK AND THE JAWS LOCKED ON TISSUE. THE STAPLER HAD TO BE REMOVED BY TEARING AND RETRACTING FROM THE TISSUE. SAME ISSUE OCCURRED WITH THE SECOND HANDLE. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. PRODUCT RECEIVED WITHOUT ACCOMPANYING INFORMATION MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
19-09-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY, WHILE UNDER THE LEFT LUNG LOBE, WHEN THE SURGEON SQUEEZED THE HANDLE ONCE, THERE WAS A LOUD CRACK, THE HANDLE COULD NOT BE SQUEEZED AND THE THE FIRING STOPPED. THE STAPLER WAS STUCK AND THE JAWS LOCKED ON TISSUE. THE STAPLER HAD TO BE REMOVED BY TEARING AND RETRACTING FROM THE TISSUE. SAME ISSUE OCCURRED WITH THE SECOND HANDLE. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. PRODUCT RECEIVED WITHOUT ACCOMPANYING INFORMATION MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
07-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE GREEN BUTTON WAS UNABLE TO BE PRESSED AND THE DEVICE DID NOT FIRE FOR THE THREE RELOADS. THERE WAS NO PATIENT INVOLVED.
07-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE GREEN BUTTON WAS UNABLE TO BE PRESSED AND THE DEVICE DID NOT FIRE FOR THE THREE RELOADS. THERE WAS NO PATIENT INVOLVED.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE GREEN BUTTON WAS UNABLE TO BE PRESSED AND THE DEVICE DID NOT FIRE FOR THE THREE RELOADS. THERE WAS NO PATIENT INVOLVED.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE GREEN BUTTON WAS UNABLE TO BE PRESSED AND THE DEVICE DID NOT FIRE FOR THE THREE RELOADS. THERE WAS NO PATIENT INVOLVED.
01-01-2023 | Injury | ethicon endo-surgery | 03-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE SURGEON REPORTED HE HAD HEMOSTASIS RELATED PATIENT ISSUES. NO FURTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCES REPORTED.
01-01-2023 | Injury | ethicon endo-surgery | 03-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE SURGEON REPORTED HE HAD HEMOSTASIS RELATED PATIENT ISSUES. NO FURTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCES REPORTED.
01-01-2023 | Injury | ethicon endo-surgery | 03-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE SURGEON REPORTED HE HAD HEMOSTASIS RELATED PATIENT ISSUES. NO FURTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCES REPORTED.
01-01-2023 | Injury | ethicon endo-surgery | 03-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE SURGEON REPORTED HE HAD HEMOSTASIS RELATED PATIENT ISSUES. NO FURTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCES REPORTED.
01-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A TRANSPLANT DONOR PROCEDURE, IN CROSS-CLAMPING THE VESSEL, RIGHT AFTER FIRING THE DEVICE, A BRISK ARTERIAL BLEEDING WAS NOTICED FROM THE MID PORTION OF THE ARTERIAL STUMP. THE DEVICE SEALED ON BOTH ENDS, BUT THERE WAS AN INCOMPLETE CLOSURE IN THE MID-SECTION.
03-11-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG TRANSPLANT PROCEDURE, AFTER FIRING THE STAPLER, THE DEVICE WAS UNABLE TO DETACH FROM THE PATIENT'S TISSUE. THE DEVICE WOULD NOT UNCLAMP. ANOTHER HANDLE AND RELOAD WERE USED TO STAPLE A NEW SITE TO THE PATIENT'S BRONCHUS. THE SURGEON WAS THEN ABLE TO REMOVE THE STAPLER FROM THE PATIENT'S BODY.
08-12-2023 | Injury | covidien lp llc north haven | 03-01-2024 | EEA | Difficult to Remove | A150207 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOIDECTOMY, AFTER THE STAPLER WAS FIRED ON THE HEMORRHOID TISSUE, THERE WAS POOR STAPLE FORMATION. ADDITIONALLY, THE STAPLE LINE AND DONUT WERE INCOMPLETE. THERE WAS ALSO DIFFICULTY WITHDRAWING THE CIRCULAR STAPLER OUT OF THE PATIENT'S BODY, EVEN WHEN THE STAPLER WAS FULLY SQUEEZED. THE PATIENT HAD A BLOOD LOSS OF MORE THAN 500 CC. THE SURGEON USED A MONOPOLAR DEVICE TO STOP THE BLEEDING AND ALSO USED MANUAL SUTURING. THE PROCEDURE TIME WAS ALSO EXTENDED BY MORE THAN 30 MINUTES. THE PATIENT WAS THEN TRANSFERRED TO A HIGH DEPENDENCY UNIT DUE TO SEVERE PAIN, AND HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
08-12-2023 | Injury | covidien lp llc north haven | 03-01-2024 | EEA | Difficult to Remove | A150207 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOIDECTOMY, AFTER THE STAPLER WAS FIRED ON THE HEMORRHOID TISSUE, THERE WAS POOR STAPLE FORMATION. ADDITIONALLY, THE STAPLE LINE AND DONUT WERE INCOMPLETE. THERE WAS ALSO DIFFICULTY WITHDRAWING THE CIRCULAR STAPLER OUT OF THE PATIENT'S BODY, EVEN WHEN THE STAPLER WAS FULLY SQUEEZED. THE PATIENT HAD A BLOOD LOSS OF MORE THAN 500 CC. THE SURGEON USED A MONOPOLAR DEVICE TO STOP THE BLEEDING AND ALSO USED MANUAL SUTURING. THE PROCEDURE TIME WAS ALSO EXTENDED BY MORE THAN 30 MINUTES. THE PATIENT WAS THEN TRANSFERRED TO A HIGH DEPENDENCY UNIT DUE TO SEVERE PAIN, AND HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
08-12-2023 | Injury | covidien lp llc north haven | 03-01-2024 | EEA | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOIDECTOMY, AFTER THE STAPLER WAS FIRED ON THE HEMORRHOID TISSUE, THERE WAS POOR STAPLE FORMATION. ADDITIONALLY, THE STAPLE LINE AND DONUT WERE INCOMPLETE. THERE WAS ALSO DIFFICULTY WITHDRAWING THE CIRCULAR STAPLER OUT OF THE PATIENT'S BODY, EVEN WHEN THE STAPLER WAS FULLY SQUEEZED. THE PATIENT HAD A BLOOD LOSS OF MORE THAN 500 CC. THE SURGEON USED A MONOPOLAR DEVICE TO STOP THE BLEEDING AND ALSO USED MANUAL SUTURING. THE PROCEDURE TIME WAS ALSO EXTENDED BY MORE THAN 30 MINUTES. THE PATIENT WAS THEN TRANSFERRED TO A HIGH DEPENDENCY UNIT DUE TO SEVERE PAIN, AND HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
08-12-2023 | Injury | covidien lp llc north haven | 03-01-2024 | EEA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOIDECTOMY, AFTER THE STAPLER WAS FIRED ON THE HEMORRHOID TISSUE, THERE WAS POOR STAPLE FORMATION. ADDITIONALLY, THE STAPLE LINE AND DONUT WERE INCOMPLETE. THERE WAS ALSO DIFFICULTY WITHDRAWING THE CIRCULAR STAPLER OUT OF THE PATIENT'S BODY, EVEN WHEN THE STAPLER WAS FULLY SQUEEZED. THE PATIENT HAD A BLOOD LOSS OF MORE THAN 500 CC. THE SURGEON USED A MONOPOLAR DEVICE TO STOP THE BLEEDING AND ALSO USED MANUAL SUTURING. THE PROCEDURE TIME WAS ALSO EXTENDED BY MORE THAN 30 MINUTES. THE PATIENT WAS THEN TRANSFERRED TO A HIGH DEPENDENCY UNIT DUE TO SEVERE PAIN, AND HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
08-12-2023 | Injury | covidien lp llc north haven | 03-01-2024 | EEA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOIDECTOMY, AFTER THE STAPLER WAS FIRED ON THE HEMORRHOID TISSUE, THERE WAS POOR STAPLE FORMATION. ADDITIONALLY, THE STAPLE LINE AND DONUT WERE INCOMPLETE. THERE WAS ALSO DIFFICULTY WITHDRAWING THE CIRCULAR STAPLER OUT OF THE PATIENT'S BODY, EVEN WHEN THE STAPLER WAS FULLY SQUEEZED. THE PATIENT HAD A BLOOD LOSS OF MORE THAN 500 CC. THE SURGEON USED A MONOPOLAR DEVICE TO STOP THE BLEEDING AND ALSO USED MANUAL SUTURING. THE PROCEDURE TIME WAS ALSO EXTENDED BY MORE THAN 30 MINUTES. THE PATIENT WAS THEN TRANSFERRED TO A HIGH DEPENDENCY UNIT DUE TO SEVERE PAIN, AND HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
08-12-2023 | Injury | covidien lp llc north haven | 03-01-2024 | EEA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN HEMORRHOIDECTOMY, AFTER THE STAPLER WAS FIRED ON THE HEMORRHOID TISSUE, THERE WAS POOR STAPLE FORMATION. ADDITIONALLY, THE STAPLE LINE AND DONUT WERE INCOMPLETE. THERE WAS ALSO DIFFICULTY WITHDRAWING THE CIRCULAR STAPLER OUT OF THE PATIENT'S BODY, EVEN WHEN THE STAPLER WAS FULLY SQUEEZED. THE PATIENT HAD A BLOOD LOSS OF MORE THAN 500 CC. THE SURGEON USED A MONOPOLAR DEVICE TO STOP THE BLEEDING AND ALSO USED MANUAL SUTURING. THE PROCEDURE TIME WAS ALSO EXTENDED BY MORE THAN 30 MINUTES. THE PATIENT WAS THEN TRANSFERRED TO A HIGH DEPENDENCY UNIT DUE TO SEVERE PAIN, AND HOSPITALIZATION WAS EXTENDED FOR THREE DAYS.
03-11-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG TRANSPLANT PROCEDURE, AFTER FIRING THE STAPLER, THE DEVICE WAS UNABLE TO DETACH FROM THE PATIENT'S TISSUE. THE DEVICE WOULD NOT UNCLAMP. ANOTHER HANDLE AND RELOAD WERE USED TO STAPLE A NEW SITE TO THE PATIENT'S BRONCHUS. THE SURGEON WAS THEN ABLE TO REMOVE THE STAPLER FROM THE PATIENT'S BODY.
04-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Great Vessel Perforation | Event Description: ACCORDING TO THE REPORTER, IN AN OPEN LOBECTOMY, WHEN STAPLING THE LUNG VESSEL, THE DEVICE FIRED, BUT THE STAPLES WERE POORLY FORMED AND INCOMPLETE. THE STAPLES IN THE MIDDLE OF THE VESSEL STAPLE LINE WERE PARTIALLY OPEN. BOTH SIDES OF THE VESSEL, THE RESECTION SIDE AND THE PATIENT SIDE, WERE OPEN. SOME STAPLES AT THE BEGINNING AND AT THE END OF THE STAPLE LINE WERE FORMED. TO RESOLVE THE ISSUE, BOTH SIDES OF THE OPEN LOBE ARTERIES (RESECTION AND PATIENT SIDE) WERE CLAMPED, COMPRESSED, AND OVER SEWED. THERE WAS BLOOD LOSS OF APPROXIMATELY 1.700 CC. NO BLOOD TRANSFUSION WAS NECESSARY. THE PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES.
04-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, IN AN OPEN LOBECTOMY, WHEN STAPLING THE LUNG VESSEL, THE DEVICE FIRED, BUT THE STAPLES WERE POORLY FORMED AND INCOMPLETE. THE STAPLES IN THE MIDDLE OF THE VESSEL STAPLE LINE WERE PARTIALLY OPEN. BOTH SIDES OF THE VESSEL, THE RESECTION SIDE AND THE PATIENT SIDE, WERE OPEN. SOME STAPLES AT THE BEGINNING AND AT THE END OF THE STAPLE LINE WERE FORMED. TO RESOLVE THE ISSUE, BOTH SIDES OF THE OPEN LOBE ARTERIES (RESECTION AND PATIENT SIDE) WERE CLAMPED, COMPRESSED, AND OVER SEWED. THERE WAS BLOOD LOSS OF APPROXIMATELY 1.700 CC. NO BLOOD TRANSFUSION WAS NECESSARY. THE PROCEDURE WAS EXTENDED FOR MORE THAN 30 MINUTES.
08-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AT THE TIME OF DISSOCIATING BLOOD VESSELS, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS A POOR STAPLE FORMATION AFTER FIRING AND THE TISSUE WAS NOT CLOSED DURING CUTTING. THE STAPLE LINE WAS ALSO INCOMPLETE. THIS ISSUE OCCURRED WITH THREE RELOADS. A NEW STAPLER WAS USED TO RE-DO THE ANASTOMOSIS.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC COLORECTAL SURGERY, THE STAPLER WAS UNABLE TO FIRE AFTER LOADING THE STAPLER. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS A POOR STAPLE FORMATION AFTER FIRING AND THE TISSUE WAS NOT CLOSED DURING CUTTING. THE STAPLE LINE WAS ALSO INCOMPLETE. THIS ISSUE OCCURRED WITH THREE RELOADS. A NEW STAPLER WAS USED TO RE-DO THE ANASTOMOSIS.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, THERE WAS A POOR STAPLE FORMATION AFTER FIRING AND THE TISSUE WAS NOT CLOSED DURING CUTTING. THE STAPLE LINE WAS ALSO INCOMPLETE. THIS ISSUE OCCURRED WITH THREE RELOADS. A NEW STAPLER WAS USED TO RE-DO THE ANASTOMOSIS.
03-10-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION, WHILE TRANSECTING THE PARENCHYMA, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE OTHER RELOAD WAS PARTIALLY FIRED. THE FULLY FIRED RELOAD HAD A DEFORMED CLAMPING MECHANISM.
03-10-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION, WHILE TRANSECTING THE PARENCHYMA, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE OTHER RELOAD WAS PARTIALLY FIRED. THE FULLY FIRED RELOAD HAD A DEFORMED CLAMPING MECHANISM.
03-10-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION, WHILE TRANSECTING THE PARENCHYMA, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE OTHER RELOAD WAS PARTIALLY FIRED. THE FULLY FIRED RELOAD HAD A DEFORMED CLAMPING MECHANISM.
03-10-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION, WHILE TRANSECTING THE PARENCHYMA, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE OTHER RELOAD WAS PARTIALLY FIRED. THE FULLY FIRED RELOAD HAD A DEFORMED CLAMPING MECHANISM.
03-10-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION, WHILE TRANSECTING THE PARENCHYMA, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE OTHER RELOAD WAS PARTIALLY FIRED. THE FULLY FIRED RELOAD HAD A DEFORMED CLAMPING MECHANISM.
03-10-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION, WHILE TRANSECTING THE PARENCHYMA, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE OTHER RELOAD WAS PARTIALLY FIRED. THE FULLY FIRED RELOAD HAD A DEFORMED CLAMPING MECHANISM.
06-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOTOMY PROCEDURE, THE DEVICE WAS USED WITH NO ISSUE WITH SIX RELOADS. HOWEVER, ON THE SEVENTH RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE; BUT, THE DEVICE WAS MAKING A "CLICKING SOUND" AND WOULD NOT FIRE THE STAPLE. SEVERAL ATTEMPTS WERE MADE TO CORRECT THE ISSUE WITHOUT RESOLUTION. THE SURGEON REQUESTED ANOTHER 45-MILLIMETER RELOAD AND HANDLE BOTH OF THE SAME TYPE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE PORT SEGMENTECTOMY, DURING CUTTING AND CLOSING THE LUNG SEGMENT TISSUE, THE RELOAD BURST THE STAPLES OUT AND HAD POOR FORMATION AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE.
04-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE FIRING THE SIXTH RELOAD, THE HANDLE MADE A LOUD CRACKING NOISE AND WOULD NOT WORK OR OPEN. ANOTHER HANDLE WAS OPENED AND USED TO RESOLVE THE ISSUE AND THE PROCEDURE WAS COMPLETED WITHOUT FURTHER INCIDENT. THERE WAS NO PATIENT INJURY.
08-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RESECTION OF MALIGNANT TUMOR OF DIGESTIVE PROCEDURE, THE JAWS DID NOT CLOSE. THE DEVICE WAS NOT USED ON THE PATIENT. ANOTHER RELOAD WAS USED WITH THE SAME POWERED DEVICES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, AS THE SURGEON BEGAN TO CUT AND STAPLE THE STOMACH AT THE LEVEL OF THE PYLORUS, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON COULD NOT BE PRESSED. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. AT THE END OF THE PROCEDURE, THE RELOAD WAS TESTED ON A DIFFERENT HANDLE AND THE DEVICE COULD NOT BE TRIGGERED.
06-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, AS THE SURGEON BEGAN TO CUT AND STAPLE THE STOMACH AT THE LEVEL OF THE PYLORUS, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON COULD NOT BE PRESSED. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. AT THE END OF THE PROCEDURE, THE RELOAD WAS TESTED ON A DIFFERENT HANDLE AND THE DEVICE COULD NOT BE TRIGGERED.
06-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, AS THE SURGEON BEGAN TO CUT AND STAPLE THE STOMACH AT THE LEVEL OF THE PYLORUS, THE FIRST RELOAD DID NOT FIRE. THE GREEN BUTTON COULD NOT BE PRESSED. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. AT THE END OF THE PROCEDURE, THE RELOAD WAS TESTED ON A DIFFERENT HANDLE AND THE DEVICE COULD NOT BE TRIGGERED.
08-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RESECTION OF MALIGNANT TUMOR OF DIGESTIVE PROCEDURE, THE JAWS DID NOT CLOSE. THE DEVICE WAS NOT USED ON THE PATIENT. ANOTHER RELOAD WAS USED WITH THE SAME POWERED DEVICES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RESECTION OF MALIGNANT TUMOR OF DIGESTIVE PROCEDURE, THE JAWS DID NOT CLOSE. THE DEVICE WAS NOT USED ON THE PATIENT. ANOTHER RELOAD WAS USED WITH THE SAME POWERED DEVICES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO GASTROPLASTY PROCEDURE, THE RELOAD PRESENTED A PROBLEM AND DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO GASTROPLASTY PROCEDURE, THE RELOAD PRESENTED A PROBLEM AND DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-11-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL GASTRECTOMY WITH GASTROJEJUNAL ANASTOMOSIS, WHEN THE STAPLER WAS USED FOR CLOSURE, IT WAS FOUND THAT THE HANDLE COULD NOT BE SQUEEZE AND DID NOT FIRE AFTER PRESSING THE GREEN BUTTON, AND STILL FAILED TO FIRE AFTER REPLACING THE RELOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-11-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL GASTRECTOMY WITH GASTROJEJUNAL ANASTOMOSIS, WHEN THE STAPLER WAS USED FOR CLOSURE, IT WAS FOUND THAT THE HANDLE COULD NOT BE SQUEEZE AND DID NOT FIRE AFTER PRESSING THE GREEN BUTTON, AND STILL FAILED TO FIRE AFTER REPLACING THE RELOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-11-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL GASTRECTOMY WITH GASTROJEJUNAL ANASTOMOSIS, WHEN THE STAPLER WAS USED FOR CLOSURE, IT WAS FOUND THAT THE HANDLE COULD NOT BE SQUEEZE AND DID NOT FIRE AFTER PRESSING THE GREEN BUTTON, AND STILL FAILED TO FIRE AFTER REPLACING THE RELOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-11-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL GASTRECTOMY WITH GASTROJEJUNAL ANASTOMOSIS, WHEN THE STAPLER WAS USED FOR CLOSURE, IT WAS FOUND THAT THE HANDLE COULD NOT BE SQUEEZE AND DID NOT FIRE AFTER PRESSING THE GREEN BUTTON, AND STILL FAILED TO FIRE AFTER REPLACING THE RELOAD. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER THE RELOAD HAS BEEN CONNECTED, THE DEVICE DID NOT ACTIVATE THE SHOOT BUTTON, EVEN THOUGH REPEATING THE MANEUVER SEVERAL TIMES. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE.
07-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE GREEN BUTTON WAS UNABLE TO BE PRESSED AND THE DEVICE DID NOT FIRE FOR THE THREE RELOADS. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
04-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE FIRING THE SIXTH RELOAD, THE HANDLE MADE A LOUD CRACKING NOISE AND WOULD NOT WORK OR OPEN. ANOTHER HANDLE WAS OPENED AND USED TO RESOLVE THE ISSUE AND THE PROCEDURE WAS COMPLETED WITHOUT FURTHER INCIDENT. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | us surgical puerto rico | 03-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE FIRING THE SIXTH RELOAD, THE HANDLE MADE A LOUD CRACKING NOISE AND WOULD NOT WORK OR OPEN. ANOTHER HANDLE WAS OPENED AND USED TO RESOLVE THE ISSUE AND THE PROCEDURE WAS COMPLETED WITHOUT FURTHER INCIDENT. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESECTION, DURING THE THIRD FIRING OF COLON RECONSTRUCTION (FUNCTIONAL END-TO-END A NASTOMOSIS), THE DEPLOYED STAPLES ON THE TIP SIDE WERE NOT FORMED PROPERLY. THE STAPLE LINE WAS DISTALLY INCOMPLETE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO REINFORCE IT WITH A SUTURE.
04-12-2023 | Malfunction | covidien lp llc north haven | 03-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RESECTION, DURING THE THIRD FIRING OF COLON RECONSTRUCTION (FUNCTIONAL END-TO-END A NASTOMOSIS), THE DEPLOYED STAPLES ON THE TIP SIDE WERE NOT FORMED PROPERLY. THE STAPLE LINE WAS DISTALLY INCOMPLETE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO REINFORCE IT WITH A SUTURE.
20-11-2023 | Malfunction | intuitive surgical | 04-01-2024 | ENDOWRIST | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, THE CURVED-TIP STAPLER 30 INSTRUMENT STOPPED WHILE FIRING ACROSS A VESSEL. USE OF THE INSTRUMENT WAS DISCONTINUED, AND IT WAS REPLACED TO THE BACKUP. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY.
20-11-2023 | Malfunction | intuitive surgical | 04-01-2024 | ENDOWRIST | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, THE CURVED-TIP STAPLER 30 INSTRUMENT STOPPED WHILE FIRING ACROSS A VESSEL. USE OF THE INSTRUMENT WAS DISCONTINUED, AND IT WAS REPLACED TO THE BACKUP. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY.
06-12-2023 | Malfunction | covidien | 04-01-2024 | GIA |  |  | Insufficient Information | Event Description: DEVICE WOULD NOT FULLY STAPLE.
06-12-2023 | Malfunction | covidien | 04-01-2024 | ENDO GIA |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: WHILE PERFORMING THE VENTRAL INCISIONAL HERNIA REPAIR WITH MESH, THE COVIDIEN STAPLER FAILED TO RELEASE THE LOAD WHEN CLOSING ON A PORTION OF SMALL BOWEL. MANUFACTURER RESPONSE FOR RELOAD ENDO GIA INTNL TAN 60, RELOAD ENDO GIA INTNL TAN 60 (PER SITE REPORTER). WOULD LIKE TO PICK THE DEVICE UP FROM THE HOSPITAL.
26-09-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE HANDLE WOULD NOT TURN ON. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED BATTERY CELLS WERE FOUND TO BE VENTED.
26-09-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE HANDLE WOULD NOT TURN ON. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED BATTERY CELLS WERE FOUND TO BE VENTED.
12-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT LOWER LOBECTOMY PROCEDURE, IN DETACHING THE INTERLOBAR FISSURE IN THE LOWER LEFT LUNG, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON; HOWEVER, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A CEPHALIC DUODENOPANCREATECTOMY AND CHOLECYSTECTOMY PROCEDURE, THE DEVICE GRIPPER BROKE DESPITE PROPER USE. ANOTHER HANDLE OF THE SAME TYPE AND LOT NUMBER WAS THEN USED, BUT IT ALSO BROKE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE RELOADS. IT WAS NOTED THAT THERE WAS A LOSS OF PROCEDURE TIME.
07-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A CEPHALIC DUODENOPANCREATECTOMY AND CHOLECYSTECTOMY PROCEDURE, THE DEVICE GRIPPER BROKE DESPITE PROPER USE. ANOTHER HANDLE OF THE SAME TYPE AND LOT NUMBER WAS THEN USED, BUT IT ALSO BROKE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE RELOADS. IT WAS NOTED THAT THERE WAS A LOSS OF PROCEDURE TIME.
07-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A CEPHALIC DUODENOPANCREATECTOMY AND CHOLECYSTECTOMY PROCEDURE, THE DEVICE GRIPPER BROKE DESPITE PROPER USE. ANOTHER HANDLE OF THE SAME TYPE AND LOT NUMBER WAS THEN USED, BUT IT ALSO BROKE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE RELOADS. IT WAS NOTED THAT THERE WAS A LOSS OF PROCEDURE TIME.
07-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A CEPHALIC DUODENOPANCREATECTOMY AND CHOLECYSTECTOMY PROCEDURE, THE DEVICE GRIPPER BROKE DESPITE PROPER USE. ANOTHER HANDLE OF THE SAME TYPE AND LOT NUMBER WAS THEN USED, BUT IT ALSO BROKE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE RELOADS. IT WAS NOTED THAT THERE WAS A LOSS OF PROCEDURE TIME.
06-12-2023 | Injury | intuitive surgical | 04-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, TISSUE WAS PUSHED OUT AND NOT CUT WHEN A GREEN SUREFORM 60 RELOAD WAS FIRED WITH THE SUREFORM 60 STAPLER INSTRUMENT. THE EVENT OCCURRED WHILE THE SURGEON WAS CREATING THE GASTRIC POUCH AND STAPLING ACROSS THE STOMACH. THE STOMACH TISSUE WAS NOTICEABLY PUSHED OUT OF THE RELOAD AND THE STAPLES WERE DUMPED INTO THE ABDOMEN WHICH WERE RETRIEVED. A NEW SUREFORM 60 STAPLER INSTRUMENT AND RELOAD WERE OPENED. THE SURGEON THEN STAPLED MEDIAL TO THE PUSHED OUT TISSUE AND SUCCESSFULLY RESECTED THE ADDITIONAL TISSUE, WHICH REPAIRED THE DEFECT.
06-12-2023 | Injury | intuitive surgical | 04-01-2024 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, TISSUE WAS PUSHED OUT AND NOT CUT WHEN A GREEN SUREFORM 60 RELOAD WAS FIRED WITH THE SUREFORM 60 STAPLER INSTRUMENT. THE EVENT OCCURRED WHILE THE SURGEON WAS CREATING THE GASTRIC POUCH AND STAPLING ACROSS THE STOMACH. THE STOMACH TISSUE WAS NOTICEABLY PUSHED OUT OF THE RELOAD AND THE STAPLES WERE DUMPED INTO THE ABDOMEN WHICH WERE RETRIEVED. A NEW SUREFORM 60 STAPLER INSTRUMENT AND RELOAD WERE OPENED. THE SURGEON THEN STAPLED MEDIAL TO THE PUSHED OUT TISSUE AND SUCCESSFULLY RESECTED THE ADDITIONAL TISSUE, WHICH REPAIRED THE DEFECT.
02-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 04-01-2024 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING TRANSABDOMINAL PREPERITONEAL (TAPP) HERNIA REPAIR PROCEDURE, THE OPTIFIX STRAIGHT FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS ALSO REPORTED THAT NO FASTENERS HAVE FIXATED THE MESH, AND SOME DAMAGED FASTENERS WERE PRESENT WHICH HAVE BEEN REMOVED FROM THE PATIENT'S BODY. THERE WAS NO REPORTED PATIENT INJURY.
20-11-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOBECTOMY, ON RESECTION OF THE LUNG LOBE, THE FIRST HANDLE WAS USED WITH A PURPLE RELOAD AND FIRED HALFWAY ACROSS THE TUMOR BUT WAS NOT ABLE TO FIRE ACROSS FULLY IN THE TISSUE. A SECOND HANDLE WAS USED WITH BLACK RELOAD BUT THE SAME ISSUE OCCURRED. A THIRD RELOAD WAS USED AND STILL THE SAME ISSUE OCCURRED. THE STAPLE FORMATION WAS INCOMPLETE AND NO B SHAPE FORMATION RESULTED ON THE SPECIMEN. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED NORMALLY. ANOTHER MANUFACTURER'S PRODUCT WAS USED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD WAS PARTIALLY FIRED TO THE 2CM CUT LINE. THE CLAMPING MECHANISM WAS DEFORMED.
20-11-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOBECTOMY, ON RESECTION OF THE LUNG LOBE, THE FIRST HANDLE WAS USED WITH A PURPLE RELOAD AND FIRED HALFWAY ACROSS THE TUMOR BUT WAS NOT ABLE TO FIRE ACROSS FULLY IN THE TISSUE. A SECOND HANDLE WAS USED WITH BLACK RELOAD BUT THE SAME ISSUE OCCURRED. A THIRD RELOAD WAS USED AND STILL THE SAME ISSUE OCCURRED. THE STAPLE FORMATION WAS INCOMPLETE AND NO B SHAPE FORMATION RESULTED ON THE SPECIMEN. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED NORMALLY. ANOTHER MANUFACTURER'S PRODUCT WAS USED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD WAS PARTIALLY FIRED TO THE 2CM CUT LINE. THE CLAMPING MECHANISM WAS DEFORMED.
06-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE OPERATING-HOLE THORACOSCOPIC LUNG WEDGE RESECTION, IN DETACHING LUNG TISSUE, THE STAPLE BURST OUT WHEN THE FIRST RELOAD WAS FIRED. A SECOND RELOAD WAS THEN USED; THE GREEN BUTTON WAS PRESSED AND THE HANDLE COULD BE SQUEEZED, BUT THE DEVICE STILL FAILED TO FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED.
06-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 04-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE OPERATING-HOLE THORACOSCOPIC LUNG WEDGE RESECTION, IN DETACHING LUNG TISSUE, THE STAPLE BURST OUT WHEN THE FIRST RELOAD WAS FIRED. A SECOND RELOAD WAS THEN USED; THE GREEN BUTTON WAS PRESSED AND THE HANDLE COULD BE SQUEEZED, BUT THE DEVICE STILL FAILED TO FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , THE LOAD STOPPED IN THE MIDDLE OF THE FIRING. THE STAPLES THAT ALREADY FIRED HAD NO FORMATION PROBLEM. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , THE LOAD STOPPED IN THE MIDDLE OF THE FIRING. THE STAPLES THAT ALREADY FIRED HAD NO FORMATION PROBLEM. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , THE LOAD STOPPED IN THE MIDDLE OF THE FIRING. THE STAPLES THAT ALREADY FIRED HAD NO FORMATION PROBLEM. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIXON SURGERY TO DETACH THE INTESTINES, DUE TO THE EDEMATOUS AND THICK INTESTINAL WALL, THE PURPLE RELOAD DID NOT CLOSE WELL AND THE COMPRESSION WAS NOT SUFFICIENT. AFTER THE GREEN SAFETY WAS PRESSED, IT WAS UNABLE TO FIRE BY HOLDING AND PULLING THE HANDLE. THE HANDLE MADE AN ABNORMAL CRACKLING SOUND AFTER BEING FORCIBLY HELD AND PULLED. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD OF THE SAME TYPE WERE USED TO COMPLETE THE OPERATION. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIXON SURGERY TO DETACH THE INTESTINES, DUE TO THE EDEMATOUS AND THICK INTESTINAL WALL, THE PURPLE RELOAD DID NOT CLOSE WELL AND THE COMPRESSION WAS NOT SUFFICIENT. AFTER THE GREEN SAFETY WAS PRESSED, IT WAS UNABLE TO FIRE BY HOLDING AND PULLING THE HANDLE. THE HANDLE MADE AN ABNORMAL CRACKLING SOUND AFTER BEING FORCIBLY HELD AND PULLED. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD OF THE SAME TYPE WERE USED TO COMPLETE THE OPERATION. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIXON SURGERY TO DETACH THE INTESTINES, DUE TO THE EDEMATOUS AND THICK INTESTINAL WALL, THE PURPLE RELOAD DID NOT CLOSE WELL AND THE COMPRESSION WAS NOT SUFFICIENT. AFTER THE GREEN SAFETY WAS PRESSED, IT WAS UNABLE TO FIRE BY HOLDING AND PULLING THE HANDLE. THE HANDLE MADE AN ABNORMAL CRACKLING SOUND AFTER BEING FORCIBLY HELD AND PULLED. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD OF THE SAME TYPE WERE USED TO COMPLETE THE OPERATION. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIXON SURGERY TO DETACH THE INTESTINES, DUE TO THE EDEMATOUS AND THICK INTESTINAL WALL, THE PURPLE RELOAD DID NOT CLOSE WELL AND THE COMPRESSION WAS NOT SUFFICIENT. AFTER THE GREEN SAFETY WAS PRESSED, IT WAS UNABLE TO FIRE BY HOLDING AND PULLING THE HANDLE. THE HANDLE MADE AN ABNORMAL CRACKLING SOUND AFTER BEING FORCIBLY HELD AND PULLED. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD OF THE SAME TYPE WERE USED TO COMPLETE THE OPERATION. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC DIXON SURGERY TO DETACH THE INTESTINES, DUE TO THE EDEMATOUS AND THICK INTESTINAL WALL, THE PURPLE RELOAD DID NOT CLOSE WELL AND THE COMPRESSION WAS NOT SUFFICIENT. AFTER THE GREEN SAFETY WAS PRESSED, IT WAS UNABLE TO FIRE BY HOLDING AND PULLING THE HANDLE. THE HANDLE MADE AN ABNORMAL CRACKLING SOUND AFTER BEING FORCIBLY HELD AND PULLED. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD OF THE SAME TYPE WERE USED TO COMPLETE THE OPERATION. THERE WAS NO PATIENT INJURY.
19-09-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, AT THE BEGINNING OF THE PROCEDURE, THE TWO STAPLERS DID NOT WORK PROPERLY. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENTS WERE LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
19-09-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, AT THE BEGINNING OF THE PROCEDURE, THE TWO STAPLERS DID NOT WORK PROPERLY. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENTS WERE LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
12-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC COLORECTAL SURGERY, WHILE ON DISSECTION OF TRANSVERSE COLON TISSUE, THE STAPLER CAN BE FIRED BUT THERE WAS POOR STAPLE FORMATION. THE STAPLES DID NOT FORM B SHAPE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO SEW AROUND THE INCISION.
20-10-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. THE RELOAD WAS REPLACED AND SAME HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-10-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. THE RELOAD WAS REPLACED AND SAME HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-10-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE.
20-10-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE.
11-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT UPPER LOBE RESECTION, AT THE TIME OF DISSOCIATING THE INTERLOBULAR FISSURES, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG TUMOR REMOVAL, THE RELOAD COULD NOT PRESS THE NEEDLES AND THERE WAS AN ISSUE INSTALLING THE DEVICE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG TUMOR REMOVAL, THE RELOAD COULD NOT PRESS THE NEEDLES AND THERE WAS AN ISSUE INSTALLING THE DEVICE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG TUMOR REMOVAL, THE RELOAD COULD NOT PRESS THE NEEDLES AND THERE WAS AN ISSUE INSTALLING THE DEVICE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT UPPER LOBE RESECTION, AT THE TIME OF DISSOCIATING THE INTERLOBULAR FISSURES, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT UPPER LOBE RESECTION, AT THE TIME OF DISSOCIATING THE INTERLOBULAR FISSURES, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-10-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE HANDLE WAS REPLACED. THERE WAS NO PATIENT INJURY.
17-10-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE HANDLE WAS REPLACED. THERE WAS NO PATIENT INJURY.
17-10-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZED THE HANDLE, BUT THE DEVICE DID NOT FIRE.
17-10-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZED THE HANDLE, BUT THE DEVICE DID NOT FIRE.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG TUMOR REMOVAL, THE RELOAD COULD NOT PRESS THE NEEDLES AND THERE WAS AN ISSUE INSTALLING THE DEVICE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG TUMOR REMOVAL, THE RELOAD COULD NOT PRESS THE NEEDLES AND THERE WAS AN ISSUE INSTALLING THE DEVICE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, WHILE PULLING DOWN THE ORVIL THROUGH THE ESOPHAGUS TO THE ESOPHAGUS STUMP, THE ORVIL (ANVIL) COULD BE PULL DOWN THROUGH THE ESOPHAGUS, THE CATHETER WAS DETACHED FROM THE ORVIL AND THE ORVIL WAS FINALLY TAKEN OUT FROM THE MOUTH BY PULLING THE SUTURE. THE SURGEON USED ANOTHER NEW ORVIL AND RESET THE NECK POSITION OF THE PATIENT, TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT CLOSE PROPERLY. THERE WAS NO PATIENT INVOLVED.
07-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE DID NOT CLOSE PROPERLY. THERE WAS NO PATIENT INVOLVED.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, AT THE TIME OF GASTROINTESTINAL RECONSTRUCTION OF THE ESOPHAGUS AND RESIDUAL STOMACH, THE COMPONENT DISENGAGED. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | EEA ORVIL | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRIC CANCER SURGERY, AT THE TIME OF GASTROINTESTINAL RECONSTRUCTION OF THE ESOPHAGUS AND RESIDUAL STOMACH, THE COMPONENT DISENGAGED. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOW RADICAL RESECTION OF RECTAL CANCER, WHILE SEVERING THE RECTUM, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE FROM THE SAME COMPANY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-09-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS UNLOADING PROBLEM, THE STAPLER WAS STUCK. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
26-09-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS UNLOADING PROBLEM, THE STAPLER WAS STUCK. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
12-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOW RADICAL RESECTION OF RECTAL CANCER, WHILE SEVERING THE RECTUM, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE FROM THE SAME COMPANY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-11-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE RELOAD WAS DAMAGED BEFORE IT WAS FIRED. THERE WAS NO PATIENT INVOLVED.
30-10-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | UNKNOWN ENDO GIA SULU | Device Displays Incorrect Message | A090201 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A ROBOTIC LAPAROSCOPIC RESECTION OF THE APICAL SEGMENT OF A LUNG CANCER PATIENT WITH PULMONARY FIBROSIS, DURING THE STAPLING RESECTION OF THE CANCEROUS SEGMENT OF THE LUNG, THE FIRST RELOAD FIRED CORRECTLY. THE SECOND RELOAD FIRED BUT STOPPED. THE SURGICAL ASSISTANT RELAYED THAT THE STAPLER WAS SHOWING IN ZONE THREE AT THE CLOSURE OF THE RELOAD. FIRING THE DEVICE RESULTED IN THE KNIFE BEING UNABLE TO PROGRESS MID-FIRE, GENERATING THE EXCESSIVE FORCE THICK TISSUE MESSAGE. THE STAPLING STOPPED LEAVING A WAD OF STAPLES AND REINFORCEMENT MATERIAL. NO DEVICE COMPONENT FELL INTO THE PATIENT¿S CAVITY. THERE WERE ALSO MALFORMED STAPLES IN THE MIDDLE OF THE FIRED STAPLE LINE. A NONREINFORCED BLACK LOAD WAS THEN TRIED, BUT IT DID NOT FIRE. THE BLACK LOAD COULD NOT ENTER THE GREEN ZONES, AND IF IT WAS TRIED, IT DID NOT WORK. ANOTHER RELOAD, SIMILAR TO THE SECOND RELOAD, WAS USED IN A DIFFERENT FIRING ANGLE BY GOING HIGHER TOWARDS THE APEX OF THE SEGMENT TO ENCOUNTER TISSUE WITHIN THE FIRING RANGE. THE SURGICAL TIME WAS EXTENDED FOR APPROXIMATELY 45 MINUTES. THERE WAS NO AIR LEAK AND THE CASE WAS COMPLETED, BUT THE PATIENT EXPERIENCED POST-OPERATIVE BLEEDING THE NIGHT AFTER THE PROCEDURE, AND WAS BROUGHT BACK TO THE OPERATING ROOM. THE ORIGINAL STAPLE LINE WAS INCOMPLETE AND ULTIMATELY BLED AT THE CONFLUENCE OF THE TWO STAPLE LINES. THERE WAS BLOOD LOSS OF MORE THAN 500CC AND MULTIPLE BLOOD TRANSFUSION WAS DONE. THE PATIENT'S CHEST WAS OPENED. THE SURGEON PERFORMED A THORACOTOMY. ANOTHER STAPLING RELOAD WAS USED TO RESECT THE EXISTING STAPLE LINE WHERE THE BLEEDING WAS COMING FROM. THE PATIENT IS STILL IN THE INTENSIVE CARE UNIT (ICU), ON THE THRESHOLD OF MOVING TO A CARE FACILITY.
30-10-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A ROBOTIC LAPAROSCOPIC RESECTION OF THE APICAL SEGMENT OF A LUNG CANCER PATIENT WITH PULMONARY FIBROSIS, DURING THE STAPLING RESECTION OF THE CANCEROUS SEGMENT OF THE LUNG, THE FIRST RELOAD FIRED CORRECTLY. THE SECOND RELOAD FIRED BUT STOPPED. THE SURGICAL ASSISTANT RELAYED THAT THE STAPLER WAS SHOWING IN ZONE THREE AT THE CLOSURE OF THE RELOAD. FIRING THE DEVICE RESULTED IN THE KNIFE BEING UNABLE TO PROGRESS MID-FIRE, GENERATING THE EXCESSIVE FORCE THICK TISSUE MESSAGE. THE STAPLING STOPPED LEAVING A WAD OF STAPLES AND REINFORCEMENT MATERIAL. NO DEVICE COMPONENT FELL INTO THE PATIENT¿S CAVITY. THERE WERE ALSO MALFORMED STAPLES IN THE MIDDLE OF THE FIRED STAPLE LINE. A NONREINFORCED BLACK LOAD WAS THEN TRIED, BUT IT DID NOT FIRE. THE BLACK LOAD COULD NOT ENTER THE GREEN ZONES, AND IF IT WAS TRIED, IT DID NOT WORK. ANOTHER RELOAD, SIMILAR TO THE SECOND RELOAD, WAS USED IN A DIFFERENT FIRING ANGLE BY GOING HIGHER TOWARDS THE APEX OF THE SEGMENT TO ENCOUNTER TISSUE WITHIN THE FIRING RANGE. THE SURGICAL TIME WAS EXTENDED FOR APPROXIMATELY 45 MINUTES. THERE WAS NO AIR LEAK AND THE CASE WAS COMPLETED, BUT THE PATIENT EXPERIENCED POST-OPERATIVE BLEEDING THE NIGHT AFTER THE PROCEDURE, AND WAS BROUGHT BACK TO THE OPERATING ROOM. THE ORIGINAL STAPLE LINE WAS INCOMPLETE AND ULTIMATELY BLED AT THE CONFLUENCE OF THE TWO STAPLE LINES. THERE WAS BLOOD LOSS OF MORE THAN 500CC AND MULTIPLE BLOOD TRANSFUSION WAS DONE. THE PATIENT'S CHEST WAS OPENED. THE SURGEON PERFORMED A THORACOTOMY. ANOTHER STAPLING RELOAD WAS USED TO RESECT THE EXISTING STAPLE LINE WHERE THE BLEEDING WAS COMING FROM. THE PATIENT IS STILL IN THE INTENSIVE CARE UNIT (ICU), ON THE THRESHOLD OF MOVING TO A CARE FACILITY.
11-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | EEA | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL TUMOR RESECTION, WHEN RECONSTRUCTING THE DIGESTIVE TRACT, THE ANVIL WAS DIFFICULT TO ATTACH ON THE DEVICE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING VESSEL TRANSECTION TO HELP REMOVE THE LOBE, THE JAWS DID NOT OPEN AFTER FIRING AND THE STAPLE LINE WAS INCOMPLETE. THE SURGEON HAD TO USE A CLIP APPLIER TO CLIP THE VESSELS, THEN SCISSOR TO TRANSECT THE TISSUE. THE RELOAD COULD NOT BE REMOVED FROM THE HANDLE.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING VESSEL TRANSECTION TO HELP REMOVE THE LOBE, THE JAWS DID NOT OPEN AFTER FIRING AND THE STAPLE LINE WAS INCOMPLETE. THE SURGEON HAD TO USE A CLIP APPLIER TO CLIP THE VESSELS, THEN SCISSOR TO TRANSECT THE TISSUE. THE RELOAD COULD NOT BE REMOVED FROM THE HANDLE.
11-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING VESSEL TRANSECTION TO HELP REMOVE THE LOBE, THE JAWS DID NOT OPEN AFTER FIRING AND THE STAPLE LINE WAS INCOMPLETE. THE SURGEON HAD TO USE A CLIP APPLIER TO CLIP THE VESSELS, THEN SCISSOR TO TRANSECT THE TISSUE. THE RELOAD COULD NOT BE REMOVED FROM THE HANDLE.
26-09-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, A FILE ERROR MESSAGE WAS SHOWN AS SOON AS THE HANDLE WAS LIFTED FROM THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
26-09-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | SIGNIA | Display or Visual Feedback Problem | A0902 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, A FILE ERROR MESSAGE WAS SHOWN AS SOON AS THE HANDLE WAS LIFTED FROM THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
11-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING VESSEL TRANSECTION TO HELP REMOVE THE LOBE, THE JAWS DID NOT OPEN AFTER FIRING AND THE STAPLE LINE WAS INCOMPLETE. THE SURGEON HAD TO USE A CLIP APPLIER TO CLIP THE VESSELS, THEN SCISSOR TO TRANSECT THE TISSUE. THE RELOAD COULD NOT BE REMOVED FROM THE HANDLE.
11-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, DURING VESSEL TRANSECTION TO HELP REMOVE THE LOBE, THE JAWS DID NOT OPEN AFTER FIRING AND THE STAPLE LINE WAS INCOMPLETE. THE SURGEON HAD TO USE A CLIP APPLIER TO CLIP THE VESSELS, THEN SCISSOR TO TRANSECT THE TISSUE. THE RELOAD COULD NOT BE REMOVED FROM THE HANDLE.
23-10-2023 | Injury | us surgical puerto rico | 04-01-2024 | UNKNOWN GIA PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A RECENT STUDY DESCRIBED AN INSTITUTION¿S SURGICAL TECHNIQUE AND ASSOCIATED OUTCOMES FOR PATIENTS WHO UNDERWENT INDIANA POUCH CONTINENT CUTANEOUS URINARY DIVERSION FOLLOWING ROBOTIC-ASSISTED RADICAL CYSTECTOMY FOR BLADDER CANCER BETWEEN 2004 AND 2020. IT WAS NOTED THAT AN OPEN GASTRO-INTESTINAL STAPLER WAS USED TO PERFORM THE ILEO-COLONIC ANASTOMOSIS IN A STAPLED SIDE TO SIDE MANNER, AS DEMONSTRATED IN THE SUPPLEMENTAL VIDEO. THERE WERE 97 PATIENTS INCLUDED IN THE STUDY AND COMPLICATIONS INCLUDED ABDOMINAL INFECTION, BOWEL ANASTOMOTIC LEAK AND SEPSIS. INTERVENTIONS AND PATIENT OUTCOMES ARE UNKNOWN.
23-10-2023 | Injury | us surgical puerto rico | 04-01-2024 | UNKNOWN GIA PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: ACCORDING TO THE LITERATURE, A RECENT STUDY DESCRIBED AN INSTITUTION¿S SURGICAL TECHNIQUE AND ASSOCIATED OUTCOMES FOR PATIENTS WHO UNDERWENT INDIANA POUCH CONTINENT CUTANEOUS URINARY DIVERSION FOLLOWING ROBOTIC-ASSISTED RADICAL CYSTECTOMY FOR BLADDER CANCER BETWEEN 2004 AND 2020. IT WAS NOTED THAT AN OPEN GASTRO-INTESTINAL STAPLER WAS USED TO PERFORM THE ILEO-COLONIC ANASTOMOSIS IN A STAPLED SIDE TO SIDE MANNER, AS DEMONSTRATED IN THE SUPPLEMENTAL VIDEO. THERE WERE 97 PATIENTS INCLUDED IN THE STUDY AND COMPLICATIONS INCLUDED ABDOMINAL INFECTION, BOWEL ANASTOMOTIC LEAK AND SEPSIS. INTERVENTIONS AND PATIENT OUTCOMES ARE UNKNOWN.
23-10-2023 | Injury | us surgical puerto rico | 04-01-2024 | UNKNOWN GIA PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: ACCORDING TO THE LITERATURE, A RECENT STUDY DESCRIBED AN INSTITUTION¿S SURGICAL TECHNIQUE AND ASSOCIATED OUTCOMES FOR PATIENTS WHO UNDERWENT INDIANA POUCH CONTINENT CUTANEOUS URINARY DIVERSION FOLLOWING ROBOTIC-ASSISTED RADICAL CYSTECTOMY FOR BLADDER CANCER BETWEEN 2004 AND 2020. IT WAS NOTED THAT AN OPEN GASTRO-INTESTINAL STAPLER WAS USED TO PERFORM THE ILEO-COLONIC ANASTOMOSIS IN A STAPLED SIDE TO SIDE MANNER, AS DEMONSTRATED IN THE SUPPLEMENTAL VIDEO. THERE WERE 97 PATIENTS INCLUDED IN THE STUDY AND COMPLICATIONS INCLUDED ABDOMINAL INFECTION, BOWEL ANASTOMOTIC LEAK AND SEPSIS. INTERVENTIONS AND PATIENT OUTCOMES ARE UNKNOWN.
11-12-2023 | Injury | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS, DURING GASTRO ANASTOMOSIS, THE STAPLE LINE WAS INCOMPLETE. THE SURGEON CLOSED THE ENTIRE GASTRO ANASTOMOSIS BY USING A MANUAL SUTURE WITH A POSSIBILITY OF STENOSIS. SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES AND THE HOSPITALIZATION WAS EXTENDED FOR FIVE MORE DAYS WITH RECOVERY IN THE INTENSIVE CARE UNIT.
07-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE END OF THE STOMACH, THE STAPLER FIRED, BUT WHILE TRYING TO FIRE THE STAPLER IT WAS NOTED THAT THE TISSUE WAS TOO THICK AND THE STAPLER GAVE THE EXCESS TISSUE ERROR. A MANUAL HANDLE AND NEW RELOAD WERE USED COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | UNKNOWN SIGNIA EGIA ADAPTER | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE END OF THE STOMACH, WHILE TRYING TO FIRE THE STAPLER IT WAS NOTED THAT THE TISSUE WAS TOO THICK AND THE STAPLER GAVE THE EXCESS TISSUE ERROR. A MANUAL HANDLE AND NEW RELOAD WERE USED COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 04-01-2024 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHEN STAPLING THE END OF THE STOMACH, WHILE TRYING TO FIRE THE STAPLER IT WAS NOTED THAT THE TISSUE WAS TOO THICK AND THE STAPLER GAVE THE EXCESS TISSUE ERROR. A MANUAL HANDLE AND NEW RELOAD WERE USED COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-07-2024 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-07-2024 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY APPENDECTOMY, WHEN THE RELOAD WAS FIXATED IN THE TISSUE IN ORDER TO RESECT THE APPENDIX, THE SURGEON PRESSED THE GREEN BUTTON AND ATTEMPTED TO FIRE, BUT THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2023 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY APPENDECTOMY, WHEN THE RELOAD WAS FIXATED IN THE TISSUE IN ORDER TO RESECT THE APPENDIX, THE SURGEON PRESSED THE GREEN BUTTON AND ATTEMPTED TO FIRE, BUT THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER THE FIRE, NOTED THE STAPLES MALFORMED WITH BLEEDING. USED SUTURE SEWING TO CONTROL THE BLEEDING AND COMPLETE THE SURGERY. THERE IS NO OTHER INFORMATION PROVIDED.
08-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER THE FIRE, NOTED THE STAPLES MALFORMED WITH BLEEDING. USED SUTURE SEWING TO CONTROL THE BLEEDING AND COMPLETE THE SURGERY. THERE IS NO OTHER INFORMATION PROVIDED.
08-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Cut | A050702 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER THE FIRE, NOTED THE STAPLES MALFORMED WITH BLEEDING. USED SUTURE SEWING TO CONTROL THE BLEEDING AND COMPLETE THE SURGERY. THERE IS NO OTHER INFORMATION PROVIDED.
08-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER THE FIRE, NOTED THE STAPLES MALFORMED WITH BLEEDING. USED SUTURE SEWING TO CONTROL THE BLEEDING AND COMPLETE THE SURGERY. THERE IS NO OTHER INFORMATION PROVIDED.
23-07-2024 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-07-2024 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2023 | Injury | ethicon endo-surgery | 04-01-2024 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A BOWEL RESECTION THE DEVICE FIRED, STAPLE LINE NORMAL, BUT WITHIN 24 HOURS REOPERATION HAD TO BE DONE WHERE THERE WERE 2 STAPLES MALFORMED IN THE MIDDLE OF THE STAPLE LINE THAT CAUSED LEAKAGE. NO PATIENT HARM. PRODUCT NOT AVAILABLE FOR RETURN.
22-07-2024 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-07-2024 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Injury | ethicon endo-surgery | 04-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Cut | A050702 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE A CDH29P WAS USED, INCOMPLETE DOUGHNUTS, DID A LEAK TEST, THE LEAK TEST LOOKED FINE. IT WAS FOUND THAT THE PATIENT HAD A LEAK A WEEK LATER.
19-12-2023 | Injury | ethicon endo-surgery | 04-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE A CDH29P WAS USED, INCOMPLETE DOUGHNUTS, DID A LEAK TEST, THE LEAK TEST LOOKED FINE. IT WAS FOUND THAT THE PATIENT HAD A LEAK A WEEK LATER.
22-07-2024 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK AND COULD NOT BE FIRED. THE RELOADS WERE ALSO DIFFICULT TO UNLOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE HANDLES WERE CHANGED TWICE. NO PATIENT INJURY.
22-07-2024 | Malfunction | us surgical puerto rico | 04-01-2024 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK AND COULD NOT BE FIRED. THE RELOADS WERE ALSO DIFFICULT TO UNLOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE HANDLES WERE CHANGED TWICE. NO PATIENT INJURY.
05-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | ENDO ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2024 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS AN UNKNOWN DEVICE ISSUE. THE RELOAD WAS REPLACED AND SAME HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
04-01-2024 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS AN UNKNOWN DEVICE ISSUE. THE RELOAD WAS REPLACED AND SAME HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
04-01-2024 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: PRODUCTS RECEIVED WITHOUT ACCOMPANYING INFORMATION MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED AND APPLICATION OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
04-01-2024 | Malfunction | covidien lp llc north haven | 04-01-2024 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: PRODUCTS RECEIVED WITHOUT ACCOMPANYING INFORMATION MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED AND APPLICATION OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
06-12-2023 | Injury | intuitive surgical | 04-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT, WITH A GREEN RELOAD PAUSED WHILE FIRING AND THERE WAS BLEEDING FROM THE STAPLE LINE. THE BLEEDING WAS RESOLVED BY USING CLIPS. THE PROCEDURE WAS COMPLETED ROBOTICALLY.
08-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | GOLD RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | PWRD ECH FLEX 45MM 34CM SHAFT | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, NOTED THAT CARTRIDGE BASE LOOSE AND COULD NOT INSTALL. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, NOTED THAT CARTRIDGE BASE LOOSE AND COULD NOT INSTALL. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2023 | Malfunction | ethicon endo-surgery | 04-01-2024 | PWRD ECH FLEX 45MM 34CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, NOTED THAT CARTRIDGE BASE LOOSE AND COULD NOT INSTALL. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2023 | Malfunction | covidien | 04-01-2024 | COVIDIEN TRI-STAPLE 2.0 REINFORCED INTELLIGENT RELOAD |  |  | Device Embedded In Tissue or Plaque | Event Description: SIGNIA STAPLER MALFUNCTIONED. 45 PURPLE REINFORCED LOAD WAS FIRED THEN ATTEMPTED TO REFIRE ON ITS OWN. THE LOAD WAS IN FIRE MODE AND WAS STUCK IN FIRE MODE AND WOULD NOT RETRACT THE BLADE OR RELEASE THE LUNG TISSUE. AN EMERGENCY KEY WAS USED BUT WOULDN'T OPEN STAPLE LOAD JAWS. A SECOND SIGNIA STAPLER WAS OPENED BUT EMERGENCY KEY WAS NOT PRESENT IN THE SET. ADDITIONAL LUNG TISSUE STAPLED WAS REMOVED WITH STAPLER STILL ATTACHED. REFERENCE REPORT: MW5149819.
18-12-2023 | Malfunction | covidien | 04-01-2024 | COVIDIEN TRI-STAPLE 2.0 REINFORCED INTELLIGENT RELOAD |  |  | Unspecified Tissue Injury | Event Description: SIGNIA STAPLER MALFUNCTIONED. 45 PURPLE REINFORCED LOAD WAS FIRED THEN ATTEMPTED TO REFIRE ON ITS OWN. THE LOAD WAS IN FIRE MODE AND WAS STUCK IN FIRE MODE AND WOULD NOT RETRACT THE BLADE OR RELEASE THE LUNG TISSUE. AN EMERGENCY KEY WAS USED BUT WOULDN'T OPEN STAPLE LOAD JAWS. A SECOND SIGNIA STAPLER WAS OPENED BUT EMERGENCY KEY WAS NOT PRESENT IN THE SET. ADDITIONAL LUNG TISSUE STAPLED WAS REMOVED WITH STAPLER STILL ATTACHED. REFERENCE REPORT: MW5149819.
18-12-2023 | Malfunction | covidien | 04-01-2024 | COVIDIEN TRI-STAPLE 2.0 REINFORCED INTELLIGENT RELOAD |  |  | Device Embedded In Tissue or Plaque | Event Description: SIGNIA STAPLER MALFUNCTIONED. 45 PURPLE REINFORCED LOAD WAS FIRED THEN ATTEMPTED TO REFIRE ON ITS OWN. THE LOAD WAS IN FIRE MODE AND WAS STUCK IN FIRE MODE AND WOULD NOT RETRACT THE BLADE OR RELEASE THE LUNG TISSUE. AN EMERGENCY KEY WAS USED BUT WOULDN'T OPEN STAPLE LOAD JAWS. A SECOND SIGNIA STAPLER WAS OPENED BUT EMERGENCY KEY WAS NOT PRESENT IN THE SET. ADDITIONAL LUNG TISSUE STAPLED WAS REMOVED WITH STAPLER STILL ATTACHED. REFERENCE REPORT: MW5149818.
18-12-2023 | Malfunction | covidien | 04-01-2024 | COVIDIEN TRI-STAPLE 2.0 REINFORCED INTELLIGENT RELOAD |  |  | Unspecified Tissue Injury | Event Description: SIGNIA STAPLER MALFUNCTIONED. 45 PURPLE REINFORCED LOAD WAS FIRED THEN ATTEMPTED TO REFIRE ON ITS OWN. THE LOAD WAS IN FIRE MODE AND WAS STUCK IN FIRE MODE AND WOULD NOT RETRACT THE BLADE OR RELEASE THE LUNG TISSUE. AN EMERGENCY KEY WAS USED BUT WOULDN'T OPEN STAPLE LOAD JAWS. A SECOND SIGNIA STAPLER WAS OPENED BUT EMERGENCY KEY WAS NOT PRESENT IN THE SET. ADDITIONAL LUNG TISSUE STAPLED WAS REMOVED WITH STAPLER STILL ATTACHED. REFERENCE REPORT: MW5149818.
13-11-2023 | Injury | ethicon endo-surgery | 04-01-2024 | ECHELON ENDOPATH¿ 35MM RELOADS |  |  | Chest Pain | Event Description: PRODUCTS USED: ETHICON STAPLER (MFG# PVE35A, 1 EACH) 35MM WHITE VASCULAR RELOAD (MFG# VASECR35, 2 EACH) ON (B)(6) 2023 ESOPHAGOSCOPY TRANSORAL WITH DIVERTICULECTOMY WAS PERFORMED. AFTER THE PROCEDURE THE PATIENT WAS COMPLAINING OF SEVERE CHEST PAIN. THE SURGEON BELIEVES THAT THE 35MM RELOADS FAILED AND CAUSED THE ESOPHAGUS TO PERFORATE. REFERENCE REPORTS: MW5149835, MW5149836.
13-11-2023 | Injury | ethicon endo-surgery | 04-01-2024 | ECHELON ENDOPATH¿ 35MM RELOADS |  |  | Perforation of Esophagus | Event Description: PRODUCTS USED: ETHICON STAPLER (MFG# PVE35A, 1 EACH) 35MM WHITE VASCULAR RELOAD (MFG# VASECR35, 2 EACH) ON (B)(6) 2023 ESOPHAGOSCOPY TRANSORAL WITH DIVERTICULECTOMY WAS PERFORMED. AFTER THE PROCEDURE THE PATIENT WAS COMPLAINING OF SEVERE CHEST PAIN. THE SURGEON BELIEVES THAT THE 35MM RELOADS FAILED AND CAUSED THE ESOPHAGUS TO PERFORATE. REFERENCE REPORTS: MW5149835, MW5149836.
13-11-2023 | Injury | ethicon endo-surgery | 04-01-2024 | ECHELON ENDOPATH¿ 35MM RELOADS |  |  | Chest Pain | Event Description: PRODUCTS USED: ETHICON STAPLER (MFG# PVE35A, 1 EACH) 35MM WHITE VASCULAR RELOAD (MFG# VASECR35, 2 EACH) ON (B)(6) 2023 ESOPHAGOSCOPY TRANSORAL WITH DIVERTICULECTOMY WAS PERFORMED. AFTER THE PROCEDURE THE PATIENT WAS COMPLAINING OF SEVERE CHEST PAIN. THE SURGEON BELIEVES THAT THE 35MM RELOADS FAILED AND CAUSED THE ESOPHAGUS TO PERFORATE. REFERENCE REPORTS: MW5149834, MW5149836.
13-11-2023 | Injury | ethicon endo-surgery | 04-01-2024 | ECHELON ENDOPATH¿ 35MM RELOADS |  |  | Perforation of Esophagus | Event Description: PRODUCTS USED: ETHICON STAPLER (MFG# PVE35A, 1 EACH) 35MM WHITE VASCULAR RELOAD (MFG# VASECR35, 2 EACH) ON (B)(6) 2023 ESOPHAGOSCOPY TRANSORAL WITH DIVERTICULECTOMY WAS PERFORMED. AFTER THE PROCEDURE THE PATIENT WAS COMPLAINING OF SEVERE CHEST PAIN. THE SURGEON BELIEVES THAT THE 35MM RELOADS FAILED AND CAUSED THE ESOPHAGUS TO PERFORATE. REFERENCE REPORTS: MW5149834, MW5149836.
13-11-2023 | Injury | ethicon endo-surgery | 04-01-2024 | ECHELON ENDOPATH¿ 35MM RELOADS |  |  | Chest Pain | Event Description: PRODUCTS USED: ETHICON STAPLER (MFG# PVE35A, 1 EACH) 35MM WHITE VASCULAR RELOAD (MFG# VASECR35, 2 EACH) ON (B)(6) 2023 ESOPHAGOSCOPY TRANSORAL WITH DIVERTICULECTOMY WAS PERFORMED. AFTER THE PROCEDURE THE PATIENT WAS COMPLAINING OF SEVERE CHEST PAIN. THE SURGEON BELIEVES THAT THE 35MM RELOADS FAILED AND CAUSED THE ESOPHAGUS TO PERFORATE. REFERENCE REPORTS: MW5149834, MW5149835.
13-11-2023 | Injury | ethicon endo-surgery | 04-01-2024 | ECHELON ENDOPATH¿ 35MM RELOADS |  |  | Perforation of Esophagus | Event Description: PRODUCTS USED: ETHICON STAPLER (MFG# PVE35A, 1 EACH) 35MM WHITE VASCULAR RELOAD (MFG# VASECR35, 2 EACH) ON (B)(6) 2023 ESOPHAGOSCOPY TRANSORAL WITH DIVERTICULECTOMY WAS PERFORMED. AFTER THE PROCEDURE THE PATIENT WAS COMPLAINING OF SEVERE CHEST PAIN. THE SURGEON BELIEVES THAT THE 35MM RELOADS FAILED AND CAUSED THE ESOPHAGUS TO PERFORATE. REFERENCE REPORTS: MW5149834, MW5149835.
10-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY, AT THE BEGINNING, THE STAPLER GOT STUCK AND DID NOT FIRE. THE MATERIAL WAS CHANGED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY, AT THE BEGINNING, THE STAPLER GOT STUCK AND DID NOT FIRE. THE MATERIAL WAS CHANGED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-10-2023 | Injury | intuitive surgical | 05-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ON 12/08/2023, INTUITIVE SURGICAL, INC. (ISI) RECEIVED MW5148325, MW5148326, MW5148327 STATING: DURING ROBOTIC THORACIC SURGICAL PROCEDURE, ROBOTIC STAPLING DEVICE DID NOT FIRE COMPLETELY. OUTERMOST ROW OF STAPLES WAS STILL IN THE STAPLE LOAD AFTER STAPLER FIRED RESULTING IN SMALL AMOUNT OF BLEEDING AT STAPLE LINE. NO ERROR MESSAGE GENERATED BY DA VINCI SYSTEM WHEN STAPLER FIRED. NEW STAPLE LOAD OPENED AND INSERTED INTO SAME STAPLE GUN. WHEN SURGEON ATTEMPTED TO USE STAPLER, SYSTEM RESPONDED THAT THERE WAS TOO MUCH TISSUE BETWEEN THE JAWS OF THE STAPLER. SURGEON REPOSITIONED STAPLER. SAME ERROR MESSAGE RECEIVED. STAPLER REMOVED FROM SERVICE. NEW STAPLER WITH NEW LOAD OPENED AND USED WITHOUT INCIDENT. EVENT REPORTED TO INTUITIVE REP. ON 12/15/2023, INTUITIVE SURGICAL, INC. (ISI) RECEIVED MW5148317 STATING: DURING ROBOTIC THORACIC SURGICAL PROCEDURE, ROBOTIC STAPLING DEVICE DID NOT FIRE COMPLETELY. OUTERMOST ROW OF STAPLES WAS STILL IN THE STAPLE LOAD AFTER STAPLER FIRED RESULTING IN SMALL AMOUNT OF BLEEDING AT STAPLE LINE. NO ERROR MESSAGES GENERATED BY DA VINCI SYSTEM WHEN STAPLER FIRED. NEW STAPLE LOAD OPENED AND INSERTED INTO SAME STAPLE GUN. WHEN SURGEON ATTEMPTED TO USE STAPLER, SYSTEM RESPONDED THAT THERE WAS TOO MUCH TISSUE BETWEEN THE JAWS OF THE STAPLER. SURGEON REPOSITIONED STAPLER. SAME ERROR MESSAGE RECEIVED. STAPLER REMOVED FROM SERVICE. NEW STAPLER WITH NEW LOAD OPENED AND USED WITHOUT INCIDENT. EVENT REPORTED TO INTUITIVE REP.
26-10-2023 | Injury | intuitive surgical | 05-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ON 12/08/2023, INTUITIVE SURGICAL, INC. (ISI) RECEIVED MW5148325, MW5148326, MW5148327 STATING: DURING ROBOTIC THORACIC SURGICAL PROCEDURE, ROBOTIC STAPLING DEVICE DID NOT FIRE COMPLETELY. OUTERMOST ROW OF STAPLES WAS STILL IN THE STAPLE LOAD AFTER STAPLER FIRED RESULTING IN SMALL AMOUNT OF BLEEDING AT STAPLE LINE. NO ERROR MESSAGE GENERATED BY DA VINCI SYSTEM WHEN STAPLER FIRED. NEW STAPLE LOAD OPENED AND INSERTED INTO SAME STAPLE GUN. WHEN SURGEON ATTEMPTED TO USE STAPLER, SYSTEM RESPONDED THAT THERE WAS TOO MUCH TISSUE BETWEEN THE JAWS OF THE STAPLER. SURGEON REPOSITIONED STAPLER. SAME ERROR MESSAGE RECEIVED. STAPLER REMOVED FROM SERVICE. NEW STAPLER WITH NEW LOAD OPENED AND USED WITHOUT INCIDENT. EVENT REPORTED TO INTUITIVE REP. ON 12/15/2023, INTUITIVE SURGICAL, INC. (ISI) RECEIVED MW5148317 STATING: DURING ROBOTIC THORACIC SURGICAL PROCEDURE, ROBOTIC STAPLING DEVICE DID NOT FIRE COMPLETELY. OUTERMOST ROW OF STAPLES WAS STILL IN THE STAPLE LOAD AFTER STAPLER FIRED RESULTING IN SMALL AMOUNT OF BLEEDING AT STAPLE LINE. NO ERROR MESSAGES GENERATED BY DA VINCI SYSTEM WHEN STAPLER FIRED. NEW STAPLE LOAD OPENED AND INSERTED INTO SAME STAPLE GUN. WHEN SURGEON ATTEMPTED TO USE STAPLER, SYSTEM RESPONDED THAT THERE WAS TOO MUCH TISSUE BETWEEN THE JAWS OF THE STAPLER. SURGEON REPOSITIONED STAPLER. SAME ERROR MESSAGE RECEIVED. STAPLER REMOVED FROM SERVICE. NEW STAPLER WITH NEW LOAD OPENED AND USED WITHOUT INCIDENT. EVENT REPORTED TO INTUITIVE REP.
26-10-2023 | Injury | intuitive surgical | 05-01-2024 | SUREFORM | Insufficient Information | E2401 | Hemorrhage/Blood Loss/Bleeding | Event Description: ON 12/08/2023, INTUITIVE SURGICAL, INC. (ISI) RECEIVED MW5148325, MW5148326, MW5148327 STATING: DURING ROBOTIC THORACIC SURGICAL PROCEDURE, ROBOTIC STAPLING DEVICE DID NOT FIRE COMPLETELY. OUTERMOST ROW OF STAPLES WAS STILL IN THE STAPLE LOAD AFTER STAPLER FIRED RESULTING IN SMALL AMOUNT OF BLEEDING AT STAPLE LINE. NO ERROR MESSAGE GENERATED BY DA VINCI SYSTEM WHEN STAPLER FIRED. NEW STAPLE LOAD OPENED AND INSERTED INTO SAME STAPLE GUN. WHEN SURGEON ATTEMPTED TO USE STAPLER, SYSTEM RESPONDED THAT THERE WAS TOO MUCH TISSUE BETWEEN THE JAWS OF THE STAPLER. SURGEON REPOSITIONED STAPLER. SAME ERROR MESSAGE RECEIVED. STAPLER REMOVED FROM SERVICE. NEW STAPLER WITH NEW LOAD OPENED AND USED WITHOUT INCIDENT. EVENT REPORTED TO INTUITIVE REP. ON 12/15/2023, INTUITIVE SURGICAL, INC. (ISI) RECEIVED MW5148317 STATING: DURING ROBOTIC THORACIC SURGICAL PROCEDURE, ROBOTIC STAPLING DEVICE DID NOT FIRE COMPLETELY. OUTERMOST ROW OF STAPLES WAS STILL IN THE STAPLE LOAD AFTER STAPLER FIRED RESULTING IN SMALL AMOUNT OF BLEEDING AT STAPLE LINE. NO ERROR MESSAGES GENERATED BY DA VINCI SYSTEM WHEN STAPLER FIRED. NEW STAPLE LOAD OPENED AND INSERTED INTO SAME STAPLE GUN. WHEN SURGEON ATTEMPTED TO USE STAPLER, SYSTEM RESPONDED THAT THERE WAS TOO MUCH TISSUE BETWEEN THE JAWS OF THE STAPLER. SURGEON REPOSITIONED STAPLER. SAME ERROR MESSAGE RECEIVED. STAPLER REMOVED FROM SERVICE. NEW STAPLER WITH NEW LOAD OPENED AND USED WITHOUT INCIDENT. EVENT REPORTED TO INTUITIVE REP.
26-10-2023 | Injury | intuitive surgical | 05-01-2024 | SUREFORM | Insufficient Information | E2401 | Insufficient Information | Event Description: ON 12/08/2023, INTUITIVE SURGICAL, INC. (ISI) RECEIVED MW5148325, MW5148326, MW5148327 STATING: DURING ROBOTIC THORACIC SURGICAL PROCEDURE, ROBOTIC STAPLING DEVICE DID NOT FIRE COMPLETELY. OUTERMOST ROW OF STAPLES WAS STILL IN THE STAPLE LOAD AFTER STAPLER FIRED RESULTING IN SMALL AMOUNT OF BLEEDING AT STAPLE LINE. NO ERROR MESSAGE GENERATED BY DA VINCI SYSTEM WHEN STAPLER FIRED. NEW STAPLE LOAD OPENED AND INSERTED INTO SAME STAPLE GUN. WHEN SURGEON ATTEMPTED TO USE STAPLER, SYSTEM RESPONDED THAT THERE WAS TOO MUCH TISSUE BETWEEN THE JAWS OF THE STAPLER. SURGEON REPOSITIONED STAPLER. SAME ERROR MESSAGE RECEIVED. STAPLER REMOVED FROM SERVICE. NEW STAPLER WITH NEW LOAD OPENED AND USED WITHOUT INCIDENT. EVENT REPORTED TO INTUITIVE REP. ON 12/15/2023, INTUITIVE SURGICAL, INC. (ISI) RECEIVED MW5148317 STATING: DURING ROBOTIC THORACIC SURGICAL PROCEDURE, ROBOTIC STAPLING DEVICE DID NOT FIRE COMPLETELY. OUTERMOST ROW OF STAPLES WAS STILL IN THE STAPLE LOAD AFTER STAPLER FIRED RESULTING IN SMALL AMOUNT OF BLEEDING AT STAPLE LINE. NO ERROR MESSAGES GENERATED BY DA VINCI SYSTEM WHEN STAPLER FIRED. NEW STAPLE LOAD OPENED AND INSERTED INTO SAME STAPLE GUN. WHEN SURGEON ATTEMPTED TO USE STAPLER, SYSTEM RESPONDED THAT THERE WAS TOO MUCH TISSUE BETWEEN THE JAWS OF THE STAPLER. SURGEON REPOSITIONED STAPLER. SAME ERROR MESSAGE RECEIVED. STAPLER REMOVED FROM SERVICE. NEW STAPLER WITH NEW LOAD OPENED AND USED WITHOUT INCIDENT. EVENT REPORTED TO INTUITIVE REP.
11-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC BYPASS, THE DEVICE WAS INTRODUCED THROUGH THE UMBILICAL PORT TO PERFORM THE FIRST SHOT. WHEN IT WAS ACTIVATED, THE 20 SECONDS WERE WAITED AND WHEN THE TRIGGER WAS FIRST CLOSED, THE STAPLER SOUNDED AS IF IT BROKE. STAPLING COULD NOT BE PERFORMED AS A RESULT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Injury | intuitive surgical | 05-01-2024 | ENDOWRIST | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC BYPASS, THE DEVICE WAS INTRODUCED THROUGH THE UMBILICAL PORT TO PERFORM THE FIRST SHOT. WHEN IT WAS ACTIVATED, THE 20 SECONDS WERE WAITED AND WHEN THE TRIGGER WAS FIRST CLOSED, THE STAPLER SOUNDED AS IF IT BROKE. STAPLING COULD NOT BE PERFORMED AS A RESULT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC BYPASS, THE DEVICE WAS INTRODUCED THROUGH THE UMBILICAL PORT TO PERFORM THE FIRST SHOT. WHEN IT WAS ACTIVATED, THE 20 SECONDS WERE WAITED AND WHEN THE TRIGGER WAS FIRST CLOSED, THE STAPLER SOUNDED AS IF IT BROKE. STAPLING COULD NOT BE PERFORMED AS A RESULT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | CARTRIDGE FINAL ASSY 60MM WHITE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-09-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, AT THE BEGINNING OF THE PROCEDURE, THE TWO STAPLERS DID NOT WORK PROPERLY. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENTS WERE LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
19-09-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTECTOMY, AT THE BEGINNING OF THE PROCEDURE, THE TWO STAPLERS DID NOT WORK PROPERLY. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENTS WERE LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
10-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY, AT THE BEGINNING, THE STAPLER GOT STUCK AND DID NOT FIRE. THE HANDLE WAS CHANGED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-10-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE-PORT THORACOSCOPIC LEFT UPPER LOBECTOMY, WHEN DETACHING THE LUNG LOBE, THE STAPLER FIRED ALL THE WAY, BUT THE DISTAL STAPLE FORMATION WAS POOR AND THE DISTAL STAPLE LINE WAS INCOMPLETE. USE THREAD WAS USED FOR REINFORCE SUTURING TO FIX THE STAPLE LINE. ANOTHER MANUFACTURER'S PRODUCT WAS USED TO COMPLETE THE CASE. PLI 20- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
10-10-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE-PORT THORACOSCOPIC LEFT UPPER LOBECTOMY, WHEN DETACHING THE LUNG LOBE, THE STAPLER FIRED ALL THE WAY, BUT THE DISTAL STAPLE FORMATION WAS POOR AND THE DISTAL STAPLE LINE WAS INCOMPLETE. USE THREAD WAS USED FOR REINFORCE SUTURING TO FIX THE STAPLE LINE. ANOTHER MANUFACTURER'S PRODUCT WAS USED TO COMPLETE THE CASE. PLI 20- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
08-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTROINTESTINAL SUTURE SEPARATION, BEFORE THE GASTROINTESTINAL SUTURE RESECTION, EVEN AFTER PROPERLY SETTING THE DEVICE UP, THE HANDLE DID NOT MOVE. FIRING COULD NOT BE PERFORMED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | UNKNOWN EGIA TRI STAPLE | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTROINTESTINAL SUTURE SEPARATION, BEFORE THE GASTROINTESTINAL SUTURE RESECTION, EVEN AFTER PROPERLY SETTING THE DEVICE UP, THE HANDLE DID NOT MOVE. FIRING COULD NOT BE PERFORMED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | ECHELON 3000 45 STANDARD | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE ECHELON 3000 WOULDN'T ARTICULATE OR OPEN FOR THE FIRST FIRING. A SECOND DEVICE WAS PULLED AND IT WORKED FINE. THERE WAS NO OTHER ISSUES DURING THE CASE. NO PATIENT HARM.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE ECHELON 3000 WOULDN'T ARTICULATE OR OPEN FOR THE FIRST FIRING. A SECOND DEVICE WAS PULLED AND IT WORKED FINE. THERE WAS NO OTHER ISSUES DURING THE CASE. NO PATIENT HARM.
18-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST OP OF A ROBOTIC OPEN RIGHT HEMICOLECTOMY THAT THE SALES REP WAS PRESENT FOR THAT FOUR DAYS POST OP THE PATIENT WAS SHOWING SIGNS OF AN ANASTOMOTIC LEAK AND WAS RETURNED TO THE OR. SURGEON FOUND THAT THERE WAS A SUBSTANTIAL LEAK IN THE ANASTOMOSIS. LEAK WAS REPAIRED BY PERFORMING ANOTHER RIGHT HEMICOLECTOMY. PATIENT IS CURRENTLY STABLE BUT VERY UPSET. SURGEON IS A VERY EXPERIENCED ETHICON DEVICE OPERATOR AND DOES NOT TYPICALLY PLACE BLAME ON THE DEVICES WHEN SOMETHING GOES WRONG BUT FEELS THAT THIS PARTICULAR INSTANCE WAS THE RESULT OF A DEVICE MALFUNCTION.
18-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: IT WAS REPORTED THAT POST OP OF A ROBOTIC OPEN RIGHT HEMICOLECTOMY THAT THE SALES REP WAS PRESENT FOR THAT FOUR DAYS POST OP THE PATIENT WAS SHOWING SIGNS OF AN ANASTOMOTIC LEAK AND WAS RETURNED TO THE OR. SURGEON FOUND THAT THERE WAS A SUBSTANTIAL LEAK IN THE ANASTOMOSIS. LEAK WAS REPAIRED BY PERFORMING ANOTHER RIGHT HEMICOLECTOMY. PATIENT IS CURRENTLY STABLE BUT VERY UPSET. SURGEON IS A VERY EXPERIENCED ETHICON DEVICE OPERATOR AND DOES NOT TYPICALLY PLACE BLAME ON THE DEVICES WHEN SOMETHING GOES WRONG BUT FEELS THAT THIS PARTICULAR INSTANCE WAS THE RESULT OF A DEVICE MALFUNCTION.
18-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Ischemia | Event Description: IT WAS REPORTED THAT POST OP OF A ROBOTIC OPEN RIGHT HEMICOLECTOMY THAT THE SALES REP WAS PRESENT FOR THAT FOUR DAYS POST OP THE PATIENT WAS SHOWING SIGNS OF AN ANASTOMOTIC LEAK AND WAS RETURNED TO THE OR. SURGEON FOUND THAT THERE WAS A SUBSTANTIAL LEAK IN THE ANASTOMOSIS. LEAK WAS REPAIRED BY PERFORMING ANOTHER RIGHT HEMICOLECTOMY. PATIENT IS CURRENTLY STABLE BUT VERY UPSET. SURGEON IS A VERY EXPERIENCED ETHICON DEVICE OPERATOR AND DOES NOT TYPICALLY PLACE BLAME ON THE DEVICES WHEN SOMETHING GOES WRONG BUT FEELS THAT THIS PARTICULAR INSTANCE WAS THE RESULT OF A DEVICE MALFUNCTION.
01-01-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THERE WERE NO PINS IN THE CARTRIDGE WHEN SURGEON FIRED THE RELOAD.
01-01-2023 | Death | ethicon endo-surgery | 05-01-2024 | ECHELON FLEX 60 COMPACT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IN THE PROCEDURE, THE STAPLER DOES NOT CUT THE TISSUE AND LIKEWISE THE RELOADING DOES NOT STAPLE. WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? : UNKNOWN, WAS PROCEDURE SUCCESSFULLY COMPLETED? : UNKNOWN, WERE FRAGMENTS GENERATED? : UNKNOWN, IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? : UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES : NO, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED : UNKNOWN, IS THE PATIENT PART OF A CLINICAL STUDY : UNKNOWN, (B)(4). DEVICE PROPERTY OF : NONE, DEVICE IN POSSESSION OF : NONE, (B)(4). DEVICE PROPERTY OF : NONE, DEVICE IN POSSESSION OF : NONE. BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST. TRUE.
01-01-2023 | Death | ethicon endo-surgery | 05-01-2024 | ECHELON FLEX 60 COMPACT | Failure to Cut | A050702 | Hemorrhage/Blood Loss/Bleeding | Event Description: IN THE PROCEDURE, THE STAPLER DOES NOT CUT THE TISSUE AND LIKEWISE THE RELOADING DOES NOT STAPLE. WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? : UNKNOWN, WAS PROCEDURE SUCCESSFULLY COMPLETED? : UNKNOWN, WERE FRAGMENTS GENERATED? : UNKNOWN, IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? : UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES : NO, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED : UNKNOWN, IS THE PATIENT PART OF A CLINICAL STUDY : UNKNOWN, (B)(4). DEVICE PROPERTY OF : NONE, DEVICE IN POSSESSION OF : NONE, (B)(4). DEVICE PROPERTY OF : NONE, DEVICE IN POSSESSION OF : NONE. BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST. TRUE.
01-01-2023 | Death | ethicon endo-surgery | 05-01-2024 | ECHELON FLEX 60 COMPACT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IN THE PROCEDURE, THE STAPLER DOES NOT CUT THE TISSUE AND LIKEWISE THE RELOADING DOES NOT STAPLE. WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? : UNKNOWN, WAS PROCEDURE SUCCESSFULLY COMPLETED? : UNKNOWN, WERE FRAGMENTS GENERATED? : UNKNOWN, IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? : UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES : NO, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED : UNKNOWN, IS THE PATIENT PART OF A CLINICAL STUDY : UNKNOWN, (B)(4). DEVICE PROPERTY OF : NONE, DEVICE IN POSSESSION OF : NONE, (B)(4). DEVICE PROPERTY OF : NONE, DEVICE IN POSSESSION OF : NONE. BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST. TRUE.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AFTER THE STAPLER WAS FIRED A LEAK TEST WAS DONE AND A LARGE LEAK WAS FOUND. A SECOND DEVICE WAS USED TO CREATE A NEW ANASTOMOSIS BELOW THE FIRST. A SECOND LEAK TEST WAS DONE AND A SMALL LEAK WAS DETECTED. THE SURGEON OVER SEWED THE STAPLE LINE TO REPAIR THE LEAK, A THIRD LEAK TEST WAS DONE AND NO BUBBLES WERE DETECTED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AFTER THE STAPLER WAS FIRED A LEAK TEST WAS DONE AND A LARGE LEAK WAS FOUND. A SECOND DEVICE WAS USED TO CREATE A NEW ANASTOMOSIS BELOW THE FIRST. A SECOND LEAK TEST WAS DONE AND A SMALL LEAK WAS DETECTED. THE SURGEON OVER SEWED THE STAPLE LINE TO REPAIR THE LEAK, A THIRD LEAK TEST WAS DONE AND NO BUBBLES WERE DETECTED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AFTER THE STAPLER WAS FIRED A LEAK TEST WAS DONE AND A LARGE LEAK WAS FOUND. A SECOND DEVICE WAS USED TO CREATE A NEW ANASTOMOSIS BELOW THE FIRST. A SECOND LEAK TEST WAS DONE AND A SMALL LEAK WAS DETECTED. THE SURGEON OVER SEWED THE STAPLE LINE TO REPAIR THE LEAK, A THIRD LEAK TEST WAS DONE AND NO BUBBLES WERE DETECTED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AFTER THE STAPLER WAS FIRED A LEAK TEST WAS DONE AND A LARGE LEAK WAS FOUND. A SECOND DEVICE WAS USED TO CREATE A NEW ANASTOMOSIS BELOW THE FIRST. A SECOND LEAK TEST WAS DONE AND A SMALL LEAK WAS DETECTED. THE SURGEON OVER SEWED THE STAPLE LINE TO REPAIR THE LEAK, A THIRD LEAK TEST WAS DONE AND NO BUBBLES WERE DETECTED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-11-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 05-01-2024 | SORBAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING LAPAROSCOPIC HERNIA REPAIR PROCEDURE, THE SORBAFIX ENHANCED (DEVICE #1) DID NOT FIRE THE FASTENERS SMOOTHLY; HOWEVER, SOME OF THE FASTENERS DEPLOYED SUCCESSFULLY. IT WAS REPORTED THAT THE SAME PROBLEM OCCURRED IN A SECOND DEVICE (DEVICE #2) AND THE PROCEDURE WAS COMPLETED WITH A THIRD DEVICE (DEVICE #3). PHOTO PROVIDED SHOW MULTIPLE LOOSE FASTENERS THAT HAVE BEEN REMOVED FROM THE BODY AND A PROUD FASTENER CAN BE SEEN FIXATED TO THE MESH. THERE WAS NO REPORTED PATIENT INJURY.
15-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 05-01-2024 | CAPSURE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A ROBOTIC VENTRAL HERNIA REPAIR A CAPSURE FIXATION DEVICE WAS USED. IT WAS REPORTED THAT THERE WAS "INADEQUATE DEPTH WITH THE TACK" AND THE SURGEON USED A NEW FIXATION DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO REPORTED PATIENT INJURY.
14-11-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 05-01-2024 | SORBAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING LAPAROSCOPIC HERNIA REPAIR PROCEDURE, THE SORBAFIX ENHANCED (DEVICE #1) DID NOT FIRE THE FASTENERS SMOOTHLY; HOWEVER, SOME OF THE FASTENERS DEPLOYED SUCCESSFULLY. IT WAS REPORTED THAT THE SAME PROBLEM OCCURRED IN A SECOND DEVICE (DEVICE #2) AND THE PROCEDURE WAS COMPLETED WITH A THIRD DEVICE (DEVICE #3). PHOTO PROVIDED SHOW MULTIPLE LOOSE FASTENERS THAT HAVE BEEN REMOVED FROM THE BODY AND A PROUD FASTENER CAN BE SEEN FIXATED TO THE MESH. THERE WAS NO REPORTED PATIENT INJURY.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AFTER THE STAPLER WAS FIRED A LEAK TEST WAS DONE AND A LARGE LEAK WAS FOUND. A SECOND DEVICE WAS USED TO CREATE A NEW ANASTOMOSIS BELOW THE FIRST. A SECOND LEAK TEST WAS DONE AND A SMALL LEAK WAS DETECTED. THE SURGEON OVER SEWED THE STAPLE LINE TO REPAIR THE LEAK, A THIRD LEAK TEST WAS DONE AND NO BUBBLES WERE DETECTED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AFTER THE STAPLER WAS FIRED A LEAK TEST WAS DONE AND A LARGE LEAK WAS FOUND. A SECOND DEVICE WAS USED TO CREATE A NEW ANASTOMOSIS BELOW THE FIRST. A SECOND LEAK TEST WAS DONE AND A SMALL LEAK WAS DETECTED. THE SURGEON OVER SEWED THE STAPLE LINE TO REPAIR THE LEAK, A THIRD LEAK TEST WAS DONE AND NO BUBBLES WERE DETECTED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AFTER THE STAPLER WAS FIRED A LEAK TEST WAS DONE AND A LARGE LEAK WAS FOUND. A SECOND DEVICE WAS USED TO CREATE A NEW ANASTOMOSIS BELOW THE FIRST. A SECOND LEAK TEST WAS DONE AND A SMALL LEAK WAS DETECTED. THE SURGEON OVER SEWED THE STAPLE LINE TO REPAIR THE LEAK, A THIRD LEAK TEST WAS DONE AND NO BUBBLES WERE DETECTED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AFTER THE STAPLER WAS FIRED A LEAK TEST WAS DONE AND A LARGE LEAK WAS FOUND. A SECOND DEVICE WAS USED TO CREATE A NEW ANASTOMOSIS BELOW THE FIRST. A SECOND LEAK TEST WAS DONE AND A SMALL LEAK WAS DETECTED. THE SURGEON OVER SEWED THE STAPLE LINE TO REPAIR THE LEAK, A THIRD LEAK TEST WAS DONE AND NO BUBBLES WERE DETECTED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
19-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH STD 29MM | Failure to Cut | A050702 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AND UPON FIRING THE DEVICE THE BLADE FAILED TO CUT THE LUMEN. THE STAPLES WORKED BUT THE BLADE DIDN'T CUT. ONCE RELEASED FROM THE PATIENT THE SURGEONS DID TRY TO CUT THROUGH SOME PAPER TO SEE IF THE BLADE WAS SHARP AND IT DIDN¿T CUT AT ALL.
19-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH STD 29MM | Failure to Cut | A050702 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AND UPON FIRING THE DEVICE THE BLADE FAILED TO CUT THE LUMEN. THE STAPLES WORKED BUT THE BLADE DIDN'T CUT. ONCE RELEASED FROM THE PATIENT THE SURGEONS DID TRY TO CUT THROUGH SOME PAPER TO SEE IF THE BLADE WAS SHARP AND IT DIDN¿T CUT AT ALL.
19-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH STD 29MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AND UPON FIRING THE DEVICE THE BLADE FAILED TO CUT THE LUMEN. THE STAPLES WORKED BUT THE BLADE DIDN'T CUT. ONCE RELEASED FROM THE PATIENT THE SURGEONS DID TRY TO CUT THROUGH SOME PAPER TO SEE IF THE BLADE WAS SHARP AND IT DIDN¿T CUT AT ALL.
19-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH STD 29MM | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AND UPON FIRING THE DEVICE THE BLADE FAILED TO CUT THE LUMEN. THE STAPLES WORKED BUT THE BLADE DIDN'T CUT. ONCE RELEASED FROM THE PATIENT THE SURGEONS DID TRY TO CUT THROUGH SOME PAPER TO SEE IF THE BLADE WAS SHARP AND IT DIDN¿T CUT AT ALL.
19-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH STD 29MM | Failure to Eject | A050301 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AND UPON FIRING THE DEVICE THE BLADE FAILED TO CUT THE LUMEN. THE STAPLES WORKED BUT THE BLADE DIDN'T CUT. ONCE RELEASED FROM THE PATIENT THE SURGEONS DID TRY TO CUT THROUGH SOME PAPER TO SEE IF THE BLADE WAS SHARP AND IT DIDN¿T CUT AT ALL.
19-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH STD 29MM | Failure to Eject | A050301 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION AND UPON FIRING THE DEVICE THE BLADE FAILED TO CUT THE LUMEN. THE STAPLES WORKED BUT THE BLADE DIDN'T CUT. ONCE RELEASED FROM THE PATIENT THE SURGEONS DID TRY TO CUT THROUGH SOME PAPER TO SEE IF THE BLADE WAS SHARP AND IT DIDN¿T CUT AT ALL.
21-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTOCOLECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE WHEN CONNECTED WITH THE BATTERY IT ENDED UP NOT WORKING FROM DEPLETION.
18-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST OP OF A ROBOTIC OPEN RIGHT HEMICOLECTOMY THAT THE SALES REP WAS PRESENT FOR THAT FOUR DAYS POST OP THE PATIENT WAS SHOWING SIGNS OF AN ANASTOMOTIC LEAK AND WAS RETURNED TO THE OR. SURGEON FOUND THAT THERE WAS A SUBSTANTIAL LEAK IN THE ANASTOMOSIS. LEAK WAS REPAIRED BY PERFORMING ANOTHER RIGHT HEMICOLECTOMY. PATIENT IS CURRENTLY STABLE BUT VERY UPSET. SURGEON IS A VERY EXPERIENCED ETHICON DEVICE OPERATOR AND DOES NOT TYPICALLY PLACE BLAME ON THE DEVICES WHEN SOMETHING GOES WRONG BUT FEELS THAT THIS PARTICULAR INSTANCE WAS THE RESULT OF A DEVICE MALFUNCTION.
18-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: IT WAS REPORTED THAT POST OP OF A ROBOTIC OPEN RIGHT HEMICOLECTOMY THAT THE SALES REP WAS PRESENT FOR THAT FOUR DAYS POST OP THE PATIENT WAS SHOWING SIGNS OF AN ANASTOMOTIC LEAK AND WAS RETURNED TO THE OR. SURGEON FOUND THAT THERE WAS A SUBSTANTIAL LEAK IN THE ANASTOMOSIS. LEAK WAS REPAIRED BY PERFORMING ANOTHER RIGHT HEMICOLECTOMY. PATIENT IS CURRENTLY STABLE BUT VERY UPSET. SURGEON IS A VERY EXPERIENCED ETHICON DEVICE OPERATOR AND DOES NOT TYPICALLY PLACE BLAME ON THE DEVICES WHEN SOMETHING GOES WRONG BUT FEELS THAT THIS PARTICULAR INSTANCE WAS THE RESULT OF A DEVICE MALFUNCTION.
18-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | PROXIMATE*75MM LINEAR CUTTER | Adverse Event Without Identified Device or Use Problem | A24 | Ischemia | Event Description: IT WAS REPORTED THAT POST OP OF A ROBOTIC OPEN RIGHT HEMICOLECTOMY THAT THE SALES REP WAS PRESENT FOR THAT FOUR DAYS POST OP THE PATIENT WAS SHOWING SIGNS OF AN ANASTOMOTIC LEAK AND WAS RETURNED TO THE OR. SURGEON FOUND THAT THERE WAS A SUBSTANTIAL LEAK IN THE ANASTOMOSIS. LEAK WAS REPAIRED BY PERFORMING ANOTHER RIGHT HEMICOLECTOMY. PATIENT IS CURRENTLY STABLE BUT VERY UPSET. SURGEON IS A VERY EXPERIENCED ETHICON DEVICE OPERATOR AND DOES NOT TYPICALLY PLACE BLAME ON THE DEVICES WHEN SOMETHING GOES WRONG BUT FEELS THAT THIS PARTICULAR INSTANCE WAS THE RESULT OF A DEVICE MALFUNCTION.
13-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE FIRING ON STOMACH, THE STAPLE LINE WAS INCOMPLETE PROXIMALLY. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
01-01-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PROXIMATE*RL LINEAR STAPLER | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE ANY PROCEDURE THERE IS NO SEAL ON THE OUTER CARTON TO TELL ME IF IT WAS PREVIOUSLY OPENED EITHER. NO PATIENT INVOLVEMENT.
01-01-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PROXIMATE*RL LINEAR STAPLER | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE ANY PROCEDURE THERE IS NO SEAL ON THE OUTER CARTON TO TELL ME IF IT WAS PREVIOUSLY OPENED EITHER. NO PATIENT INVOLVEMENT.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE FIRING ON ANTRUM, THERE WERE 2 RELOADS AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DETACHING THE GASTRIC BODY, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DETACHING THE GASTRIC BODY, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DETACHING THE GASTRIC BODY, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DETACHING THE GASTRIC BODY, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DETACHING THE GASTRIC BODY, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DETACHING THE GASTRIC BODY, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH XL SEAL 33MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE END TO END ANASTOMOSIS BROKE DOWN, POSSIBLY DUE TO A MALFUNCTION OF THE CIRCULAR STAPLER. DR. THEN HAD TO PERFORM AN ILEOSTOMY.
08-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH XL SEAL 33MM | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE END TO END ANASTOMOSIS BROKE DOWN, POSSIBLY DUE TO A MALFUNCTION OF THE CIRCULAR STAPLER. DR. THEN HAD TO PERFORM AN ILEOSTOMY.
08-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH XL SEAL 33MM | Insufficient Information | E2401 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE END TO END ANASTOMOSIS BROKE DOWN, POSSIBLY DUE TO A MALFUNCTION OF THE CIRCULAR STAPLER. DR. THEN HAD TO PERFORM AN ILEOSTOMY.
08-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH XL SEAL 33MM | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE END TO END ANASTOMOSIS BROKE DOWN, POSSIBLY DUE TO A MALFUNCTION OF THE CIRCULAR STAPLER. DR. THEN HAD TO PERFORM AN ILEOSTOMY.
08-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH XL SEAL 33MM | Ejection Problem | A0503 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE END TO END ANASTOMOSIS BROKE DOWN, POSSIBLY DUE TO A MALFUNCTION OF THE CIRCULAR STAPLER. DR. THEN HAD TO PERFORM AN ILEOSTOMY.
08-12-2023 | Injury | ethicon endo-surgery | 05-01-2024 | CIRCULAR MECH XL SEAL 33MM | Ejection Problem | A0503 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE END TO END ANASTOMOSIS BROKE DOWN, POSSIBLY DUE TO A MALFUNCTION OF THE CIRCULAR STAPLER. DR. THEN HAD TO PERFORM AN ILEOSTOMY.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE FIRING ON ANTRUM, THERE WERE 2 RELOADS AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
21-12-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY + ILEORECTAL ANASTOMOSIS- REGISTRAR WAS MANAGING THE STAPLER FROM THE BOTTOM END. SURGEON AND REGISTRAR REPORTED A CLICK WAS HEARD WHEN THE ANVIL WAS CONNECTED TO THE SPIKE. THE STAPLER WAS FIRED, BUT THE REGISTRAR THEN HAD TROUBLE REMOVING THE STAPLER OUT. THE SURGEONS HAD TO MILK THE BOWEL AND DO A CONSIDERABLE AMOUNT OF TUGGING TO REMOVE THE STAPLER. THE ANVIL WAS NOT CONNECTED TO THE STAPLER WHEN THE STAPLER CAME OUT. THE ANVIL HAD TO BE RETRIEVED USING A SPONGE HOLDER. THERE WAS A SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRECTOMY PROCEDURE, WHEN THE JAW OF THE DEVICE WAS CLOSED, THE JAW OPENING WAS LARGER THAN USUAL. HENCE, FIRING WAS NOT PERFORMED. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE FIRING ON ANTRUM, THERE WAS POOR STAPLE FORMATION, THE STAPLES WERE OPEN, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE AND RESOLVE THE ISSUE.
07-12-2023 | Injury | intuitive surgical | 05-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED DUODENAL SWITCH SURGICAL PROCEDURE, THE SUREFORM STAPLER 60 INSTRUMENT FIRED AN UNKNOWN COLORED SUREFORM 60 RELOAD AND PUSHED THE TARGET TISSUE, CREATING MALFORMED STAPLES. DUE TO THIS EVENT, THE SURGEON CHANGED THE PROCEDURE TO A GASTRIC SLEEVE SURGICAL PROCEDURE. THE PROCEDURE WAS COMPLETED ROBOTICALLY.
07-12-2023 | Injury | intuitive surgical | 05-01-2024 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED DUODENAL SWITCH SURGICAL PROCEDURE, THE SUREFORM STAPLER 60 INSTRUMENT FIRED AN UNKNOWN COLORED SUREFORM 60 RELOAD AND PUSHED THE TARGET TISSUE, CREATING MALFORMED STAPLES. DUE TO THIS EVENT, THE SURGEON CHANGED THE PROCEDURE TO A GASTRIC SLEEVE SURGICAL PROCEDURE. THE PROCEDURE WAS COMPLETED ROBOTICALLY.
06-11-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHILE ON PULMONARY VEINS, THE STAPLER FIRED COMPLETELY, BUT THERE WAS BLEEDING ON ALL VASCULAR STUMPS AFTER STAPLING. THE BLOOD LOSS WAS ESTIMATED AT 250 ML. TO RESOLVE THE ISSUE, CHANGED CONSUMABLE AND RELOAD WAS CHANGED AFTER THE STAPLE LINE. THE OPERATIVE TIME WAS EXTENDED FOR 30 MINUTES OR MORE DUE TO THE PRODUCT ISSUE.
06-11-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHILE ON PULMONARY VEINS, THE STAPLER FIRED COMPLETELY, BUT THERE WAS BLEEDING ON ALL VASCULAR STUMPS AFTER STAPLING. THE BLOOD LOSS WAS ESTIMATED AT 250 ML. TO RESOLVE THE ISSUE, CHANGED CONSUMABLE AND RELOAD WAS CHANGED AFTER THE STAPLE LINE. THE OPERATIVE TIME WAS EXTENDED FOR 30 MINUTES OR MORE DUE TO THE PRODUCT ISSUE.
13-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ANASTOMOSIS WAS PERFORMED USING REINFORCE ON THE GASTRODUODENAL SIDE WITH A POWERED HANDLE. WHILE FIRING WAS PERFORMED, CONTINUED PRESSING THE BUTTON AND WHEN STOPPED, IT WAS THOUGHT THAT IT WAS COMPLETED AND THE JAWS WERE RELEASED. HOWEVER, THE FIRING HAD ONLY ADVANCED TILL HALFWAY, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE TISSUE WAS DAMAGED, DELTA ANASTOMOSIS WAS ABANDONED, THE OPERATION WAS COMPLETED WITH ADDITIONAL RESECTION OF THE DUODENUM AND PERFORMED RECOVERY WITH R-Y RECONSTRUCTION OR GASTRECTOMY R-Y ANASTOMOSIS AND INCOMPLETE TISSUE SPECIMENS WERE ALSO COLLECTED WITH RELOAD.
13-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ANASTOMOSIS WAS PERFORMED USING REINFORCE ON THE GASTRODUODENAL SIDE WITH A POWERED HANDLE. WHILE FIRING WAS PERFORMED, CONTINUED PRESSING THE BUTTON AND WHEN STOPPED, IT WAS THOUGHT THAT IT WAS COMPLETED AND THE JAWS WERE RELEASED. HOWEVER, THE FIRING HAD ONLY ADVANCED TILL HALFWAY, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE TISSUE WAS DAMAGED, DELTA ANASTOMOSIS WAS ABANDONED, THE OPERATION WAS COMPLETED WITH ADDITIONAL RESECTION OF THE DUODENUM AND PERFORMED RECOVERY WITH R-Y RECONSTRUCTION OR GASTRECTOMY R-Y ANASTOMOSIS AND INCOMPLETE TISSUE SPECIMENS WERE ALSO COLLECTED WITH RELOAD.
13-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ANASTOMOSIS WAS PERFORMED USING REINFORCE ON THE GASTRODUODENAL SIDE W ITH A POWERED HANDLE. WHILE FIRING WAS PERFORMED, CONTINUED PRESSING THE BUTTON AND WHEN STOPPED, IT WAS THOUGHT THAT IT WAS COMPLETED AND THE JAWS WERE RELEASED. HOWEVER, THE FIRING HAD ONLY ADVANCED TILL HALFWAY, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE TISSUE WAS DAMAGED, DELTA ANASTOMOSIS WAS ABANDONED, THE OPERATION WAS COMPLETED WITH ADDITIONAL RESECTION OF THE DUODENUM AND PERFORMED RECOVERY WITH R-Y RECONSTRUCTION OR GASTRECTOMY R-Y ANASTOMOSIS AND INCOMPLETE TISSUE SPECIMENS WERE ALSO COLLECTED WITH RELOAD.
13-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ANASTOMOSIS WAS PERFORMED USING REINFORCE ON THE GASTRODUODENAL SIDE W ITH A POWERED HANDLE. WHILE FIRING WAS PERFORMED, CONTINUED PRESSING THE BUTTON AND WHEN STOPPED, IT WAS THOUGHT THAT IT WAS COMPLETED AND THE JAWS WERE RELEASED. HOWEVER, THE FIRING HAD ONLY ADVANCED TILL HALFWAY, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE TISSUE WAS DAMAGED, DELTA ANASTOMOSIS WAS ABANDONED, THE OPERATION WAS COMPLETED WITH ADDITIONAL RESECTION OF THE DUODENUM AND PERFORMED RECOVERY WITH R-Y RECONSTRUCTION OR GASTRECTOMY R-Y ANASTOMOSIS AND INCOMPLETE TISSUE SPECIMENS WERE ALSO COLLECTED WITH RELOAD.
13-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ANASTOMOSIS WAS PERFORMED USING REINFORCE ON THE GASTRODUODENAL SIDE WITH A POWERED HANDLE. WHILE FIRING WAS PERFORMED, CONTINUED PRESSING THE BUTTON AND WHEN STOPPED, IT WAS THOUGHT THAT IT WAS COMPLETED AND THE JAWS WERE RELEASED. HOWEVER, THE FIRING HAD ONLY ADVANCED TILL HALFWAY, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE TISSUE WAS DAMAGED, DELTA ANASTOMOSIS WAS ABANDONED, THE OPERATION WAS COMPLETED WITH ADDITIONAL RESECTION OF THE DUODENUM AND PERFORMED RECOVERY WITH R-Y RECONSTRUCTION OR GASTRECTOMY R-Y ANASTOMOSIS AND INCOMPLETE TISSUE SPECIMENS WERE ALSO COLLECTED WITH RELOAD.
13-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, ANASTOMOSIS WAS PERFORMED USING REINFORCE ON THE GASTRODUODENAL SIDE WITH A POWERED HANDLE. WHILE FIRING WAS PERFORMED, CONTINUED PRESSING THE BUTTON AND WHEN STOPPED, IT WAS THOUGHT THAT IT WAS COMPLETED AND THE JAWS WERE RELEASED. HOWEVER, THE FIRING HAD ONLY ADVANCED TILL HALFWAY, AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE TISSUE WAS DAMAGED, DELTA ANASTOMOSIS WAS ABANDONED, THE OPERATION WAS COMPLETED WITH ADDITIONAL RESECTION OF THE DUODENUM AND PERFORMED RECOVERY WITH R-Y RECONSTRUCTION OR GASTRECTOMY R-Y ANASTOMOSIS AND INCOMPLETE TISSUE SPECIMENS WERE ALSO COLLECTED WITH RELOAD.
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 05-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC PROCEDURE, AFTER FIRING, THE POSITION OF THE PATIENT WAS CHANGED, AND THE SURGEON NOTICED AN AIR LEAKAGE FROM THE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE TWO RELOADS WERE FIRED. THE SURGEON CLOSED THE TISSUE BY USING A BLACK 60MM RELOAD. IT WAS REPORTED THAT THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES, THE INCISION WAS EXTENDED DUE TO THE PRODUCT PROBLEM AND THE PATIENT NEEDED TO UNDERGO OPERATION THE SECOND TIME ON (B)(6) 2023 DUE TO CONTINUING BLEEDING. APPROACH WAS THORACOTOMY USING OPEN STAPLERS FOR THE PROCEDURES AND THE EVENT LED TO EXTENDED HOSPITALIZATION.
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 05-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC PROCEDURE, AFTER FIRING, THE POSITION OF THE PATIENT WAS CHANGED, AND THE SURGEON NOTICED AN AIR LEAKAGE FROM THE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE TWO RELOADS WERE FIRED. THE SURGEON CLOSED THE TISSUE BY USING A BLACK 60MM RELOAD. IT WAS REPORTED THAT THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES, THE INCISION WAS EXTENDED DUE TO THE PRODUCT PROBLEM AND THE PATIENT NEEDED TO UNDERGO OPERATION THE SECOND TIME ON (B)(6) 2023 DUE TO CONTINUING BLEEDING. APPROACH WAS THORACOTOMY USING OPEN STAPLERS FOR THE PROCEDURES AND THE EVENT LED TO EXTENDED HOSPITALIZATION.
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 05-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC PROCEDURE, AFTER FIRING, THE POSITION OF THE PATIENT WAS CHANGED, AND THE SURGEON NOTICED AN AIR LEAKAGE FROM THE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE TWO RELOADS WERE FIRED. THE SURGEON CLOSED THE TISSUE BY USING A BLACK 60MM RELOAD. IT WAS REPORTED THAT THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES, THE INCISION WAS EXTENDED DUE TO THE PRODUCT PROBLEM AND THE PATIENT NEEDED TO UNDERGO OPERATION THE SECOND TIME ON (B)(6) 2023 DUE TO CONTINUING BLEEDING. APPROACH WAS THORACOTOMY USING OPEN STAPLERS FOR THE PROCEDURES AND THE EVENT LED TO EXTENDED HOSPITALIZATION.
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 05-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC PROCEDURE, AFTER FIRING, THE POSITION OF THE PATIENT WAS CHANGED, AND THE SURGEON NOTICED AN AIR LEAKAGE FROM THE STAPLE LINE AND POOR STAPLE FORMATION WHEREIN THE TWO RELOADS WERE FIRED. THE SURGEON CLOSED THE TISSUE BY USING A BLACK 60MM RELOAD. IT WAS REPORTED THAT THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES, THE INCISION WAS EXTENDED DUE TO THE PRODUCT PROBLEM AND THE PATIENT NEEDED TO UNDERGO OPERATION THE SECOND TIME ON (B)(6) 2023 DUE TO CONTINUING BLEEDING. APPROACH WAS THORACOTOMY USING OPEN STAPLERS FOR THE PROCEDURES AND THE EVENT LED TO EXTENDED HOSPITALIZATION.
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 05-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC PROCEDURE. AFTER FIRING, THE POSITION OF THE PATIENT WAS CHANGED, AND THE SURGEON NOTICED AN AIR LEAKAGE FROM THE STAPLE LINE. AND POOR STAPLE FORMATION WHEREIN THE TWO RELOADS WERE FIRED. THE SURGEON CLOSED THE TISSUE BY USING A BLACK 60MM RELOAD. IT WAS REPORTED, THAT THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES. THE INCISION WAS EXTENDED, DUE TO THE PRODUCT PROBLEM. AND THE PATIENT NEEDED TO UNDERGO OPERATION THE SECOND TIME ON (B)(6) 2023, DUE TO CONTINUING BLEEDING. APPROACH WAS THORACOTOMY USING OPEN STAPLERS FOR THE PROCEDURES. AND THE EVENT LED TO EXTENDED HOSPITALIZATION.
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 05-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC PROCEDURE. AFTER FIRING, THE POSITION OF THE PATIENT WAS CHANGED, AND THE SURGEON NOTICED AN AIR LEAKAGE FROM THE STAPLE LINE. AND POOR STAPLE FORMATION WHEREIN THE TWO RELOADS WERE FIRED. THE SURGEON CLOSED THE TISSUE BY USING A BLACK 60MM RELOAD. IT WAS REPORTED, THAT THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES. THE INCISION WAS EXTENDED, DUE TO THE PRODUCT PROBLEM. AND THE PATIENT NEEDED TO UNDERGO OPERATION THE SECOND TIME ON (B)(6) 2023, DUE TO CONTINUING BLEEDING. APPROACH WAS THORACOTOMY USING OPEN STAPLERS FOR THE PROCEDURES. AND THE EVENT LED TO EXTENDED HOSPITALIZATION.
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 05-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC PROCEDURE. AFTER FIRING, THE POSITION OF THE PATIENT WAS CHANGED, AND THE SURGEON NOTICED AN AIR LEAKAGE FROM THE STAPLE LINE. AND POOR STAPLE FORMATION WHEREIN THE TWO RELOADS WERE FIRED. THE SURGEON CLOSED THE TISSUE BY USING A BLACK 60MM RELOAD. IT WAS REPORTED, THAT THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES. THE INCISION WAS EXTENDED, DUE TO THE PRODUCT PROBLEM. AND THE PATIENT NEEDED TO UNDERGO OPERATION THE SECOND TIME ON (B)(6) 2023, DUE TO CONTINUING BLEEDING. APPROACH WAS THORACOTOMY USING OPEN STAPLERS FOR THE PROCEDURES. AND THE EVENT LED TO EXTENDED HOSPITALIZATION.
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 05-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC PROCEDURE. AFTER FIRING, THE POSITION OF THE PATIENT WAS CHANGED, AND THE SURGEON NOTICED AN AIR LEAKAGE FROM THE STAPLE LINE. AND POOR STAPLE FORMATION WHEREIN THE TWO RELOADS WERE FIRED. THE SURGEON CLOSED THE TISSUE BY USING A BLACK 60MM RELOAD. IT WAS REPORTED, THAT THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES. THE INCISION WAS EXTENDED, DUE TO THE PRODUCT PROBLEM. AND THE PATIENT NEEDED TO UNDERGO OPERATION THE SECOND TIME ON (B)(6) 2023, DUE TO CONTINUING BLEEDING. APPROACH WAS THORACOTOMY USING OPEN STAPLERS FOR THE PROCEDURES. AND THE EVENT LED TO EXTENDED HOSPITALIZATION.
12-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOW RECTAL CANCER SURGERY, WHEN DETACHING THE RECTUM, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE SAME RELOAD WITH A NEW HANDLE FROM THE SAME COMPANY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRECTOMY, DURING GASTRIC RESECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SUPERIOR POLAR OESOGASTRECTOMY FOR A LARGE PART OF THE DISSECTION WAS PERFORMED WITH CO ELIO, FOLLOWED BY A LAPAROTOMY TO BRING THE TWO DISSECTION PLANES TOGETHER, THE DEVICE MALFUNCTIONED. GASTRIC STAPLING WAS PLANNED IN ORDER TO MAKE A GASTRIC TUBE BUT THE RELOAD SEIZED BEFORE STAPLING COULD BE COMPLETED. THE 45 MM RELOAD HAS DETACHED FROM THE FORCEPS AND HAD TO BE REMOVED BY THE SURGEON FROM INSIDE THE PATIENT.
07-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SUPERIOR POLAR OESOGASTRECTOMY FOR A LARGE PART OF THE DISSECTION WAS PERFORMED WITH CO ELIO, FOLLOWED BY A LAPAROTOMY TO BRING THE TWO DISSECTION PLANES TOGETHER, THE DEVICE MALFUNCTIONED. GASTRIC STAPLING WAS PLANNED IN ORDER TO MAKE A GASTRIC TUBE BUT THE RELOAD SEIZED BEFORE STAPLING COULD BE COMPLETED. THE 45 MM RELOAD HAS DETACHED FROM THE FORCEPS AND HAD TO BE REMOVED BY THE SURGEON FROM INSIDE THE PATIENT.
07-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SUPERIOR POLAR OESOGASTRECTOMY FOR A LARGE PART OF THE DISSECTION WAS PERFORMED WITH CO ELIO, FOLLOWED BY A LAPAROTOMY TO BRING THE TWO DISSECTION PLANES TOGETHER, THE DEVICE MALFUNCTIONED. GASTRIC STAPLING WAS PLANNED IN ORDER TO MAKE A GASTRIC TUBE BUT THE RELOAD SEIZED BEFORE STAPLING COULD BE COMPLETED. THE 45 MM RELOAD HAS DETACHED FROM THE FORCEPS AND HAD TO BE REMOVED BY THE SURGEON FROM INSIDE THE PATIENT.
07-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SUPERIOR POLAR OESOGASTRECTOMY FOR A LARGE PART OF THE DISSECTION WAS PERFORMED WITH CO ELIO, FOLLOWED BY A LAPAROTOMY TO BRING THE TWO DISSECTION PLANES TOGETHER, THE DEVICE MALFUNCTIONED. GASTRIC STAPLING WAS PLANNED IN ORDER TO MAKE A GASTRIC TUBE BUT THE RELOAD SEIZED BEFORE STAPLING COULD BE COMPLETED. THE 45 MM RELOAD HAS DETACHED FROM THE FORCEPS AND HAD TO BE REMOVED BY THE SURGEON FROM INSIDE THE PATIENT.
30-11-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LIVER RESECTION, THE STAPLER WAS APPLIED TO THE TISSUE, AND IT CLOSED COMFORTABLY WITHOUT FORCE. THE GREEN BUTTON WAS PUSHED, AND THE TRIGGER EXTENDED FORWARD. ON THE FIRST SQUEEZE, THE STAPLER MAKES A ¿CRUNCHING¿ SOUND. THE STAPLER DID NOT FIRE ANY STAPLES OR CUT SINCE THE TRIGGER OF THE HANDLE WAS LOOSE AND FLOPPY. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LIVER RESECTION, THE STAPLER WAS APPLIED TO THE TISSUE, AND IT CLOSED COMFORTABLY WITHOUT FORCE. THE GREEN BUTTON WAS PUSHED, AND THE TRIGGER EXTENDED FORWARD. ON THE FIRST SQUEEZE, THE STAPLER MAKES A ¿CRUNCHING¿ SOUND. THE STAPLER DID NOT FIRE ANY STAPLES OR CUT SINCE THE TRIGGER OF THE HANDLE WAS LOOSE AND FLOPPY. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LIVER RESECTION, THE STAPLER WAS APPLIED TO THE TISSUE, AND IT CLOSED COMFORTABLY WITHOUT FORCE. THE GREEN BUTTON WAS PUSHED, AND THE TRIGGER EXTENDED FORWARD. ON THE FIRST SQUEEZE, THE STAPLER MAKES A ¿CRUNCHING¿ SOUND. THE STAPLER DID NOT FIRE ANY STAPLES OR CUT SINCE THE TRIGGER OF THE HANDLE WAS LOOSE AND FLOPPY. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LIVER RESECTION, THE STAPLER WAS APPLIED TO THE TISSUE, AND IT CLOSED COMFORTABLY WITHOUT FORCE. THE GREEN BUTTON WAS PUSHED, AND THE TRIGGER EXTENDED FORWARD. ON THE FIRST SQUEEZE, THE STAPLER MAKES A ¿CRUNCHING¿ SOUND. THE STAPLER DID NOT FIRE ANY STAPLES OR CUT SINCE THE TRIGGER OF THE HANDLE WAS LOOSE AND FLOPPY. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LIVER RESECTION, THE STAPLER WAS APPLIED TO THE TISSUE, AND IT CLOSED COMFORTABLY WITHOUT FORCE. THE GREEN BUTTON WAS PUSHED, AND THE TRIGGER EXTENDED FORWARD. ON THE FIRST SQUEEZE, THE STAPLER MAKES A ¿CRUNCHING¿ SOUND. THE STAPLER DID NOT FIRE ANY STAPLES OR CUT SINCE THE TRIGGER OF THE HANDLE WAS LOOSE AND FLOPPY. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LIVER RESECTION, THE STAPLER'S HANDLE BROKE AND THE RELOAD WAS LOCKED ON ONE OF THE VEINS. THE RELOAD HAD NOT FIRED, SO IT WAS REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LIVER RESECTION, THE STAPLER'S HANDLE BROKE AND THE RELOAD WAS LOCKED ON ONE OF THE VEINS. THE RELOAD HAD NOT FIRED, SO IT WAS REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LIVER RESECTION, THE STAPLER'S HANDLE BROKE AND THE RELOAD WAS LOCKED ON ONE OF THE VEINS. THE RELOAD HAD NOT FIRED, SO IT WAS REMOVED. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, HANDLE WAS USED TO RESECT AND ANASTOMOSED THE STOMACH, BUT WHEN THE HANDLE WAS GRIPPED AND TRIED TO FIRE THE STAPLER, THERE WAS A CRACKING SOUND AND THE STAPLER DID NOT FIRE DUE. THE SURGEON COULD PRESS THE GREEN BUTTON. WHEN THE SECOND HANDLE WAS USED, THE SAME PROBLEM OCCURRED, THE STAPLER DID NOT FIRE WITH CRACKING SOUND. THE OPERATION COULD BE COMPLETED EVENTUALLY USING THE THIRD HANDLE AND SAME RELOAD. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, HANDLE WAS USED TO RESECT AND ANASTOMOSED THE STOMACH, BUT WHEN THE HANDLE WAS GRIPPED AND TRIED TO FIRE THE STAPLER, THERE WAS A CRACKING SOUND AND THE STAPLER DID NOT FIRE DUE. THE SURGEON COULD PRESS THE GREEN BUTTON. WHEN THE SECOND HANDLE WAS USED, THE SAME PROBLEM OCCURRED, THE STAPLER DID NOT FIRE WITH CRACKING SOUND. THE OPERATION COULD BE COMPLETED EVENTUALLY USING THE THIRD HANDLE AND SAME RELOAD. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, HANDLE WAS USED TO RESECT AND ANASTOMOSED THE STOMACH, BUT WHEN THE HANDLE WAS GRIPPED AND TRIED TO FIRE THE STAPLER, THERE WAS A CRACKING SOUND AND THE STAPLER DID NOT FIRE DUE. THE SURGEON COULD PRESS THE GREEN BUTTON. WHEN THE SECOND HANDLE WAS USED, THE SAME PROBLEM OCCURRED, THE STAPLER DID NOT FIRE WITH CRACKING SOUND. THE OPERATION COULD BE COMPLETED EVENTUALLY USING THE THIRD HANDLE AND SAME RELOAD. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY, HANDLE WAS USED TO RESECT AND ANASTOMOSED THE STOMACH, BUT WHEN THE HANDLE WAS GRIPPED AND TRIED TO FIRE THE STAPLER, THERE WAS A CRACKING SOUND AND THE STAPLER DID NOT FIRE DUE. THE SURGEON COULD PRESS THE GREEN BUTTON. WHEN THE SECOND HANDLE WAS USED, THE SAME PROBLEM OCCURRED, THE STAPLER DID NOT FIRE WITH CRACKING SOUND. THE OPERATION COULD BE COMPLETED EVENTUALLY USING THE THIRD HANDLE AND SAME RELOAD. THERE WAS NO PATIENT INJURY.
11-07-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AFTER SEVERAL STAPLES WERE USED, THE DEVICE DID NOT FIRE. WHEN THE RELOAD WAS REINSTALLED, IT STILL COULD NOT BE USED. A NEW HANDLE WAS USED WITH THE SAME RELOAD TO CONTINUE THE SURGERY.
17-11-2023 | Malfunction | celestica de monterrey s.a de c.v | 05-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL CANCER SURGERY, THE JAWS OF THE RELOAD COULD NOT BE COMPLETELY CLOSED DURING CUTTING. A NEW RELOAD WAS USED WITH THE SAME MANUAL HANDLE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
21-09-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC APPENDECTOMY, THE STAPLER WAS NOT WORKING AND WOULD NOT FIRE. IT WAS ALSO MENTIONED THAT THE STAPLER GOT STUCK TOGETHER, SO IT COULD NOT BE RELOADED. THE USER, ANOTHER TECH, REGISTERED NURSE AND THE SURGEON TRIED TO TAKE IT APART TO RELOAD A NEW STAPLE ROW BUT NONE COULD GET IT DETACHED. THE BUTTON WAS STUCK AND NOT MOVING AT ALL. THERE WAS A SHORT DELAY AROUND 90 SECONDS IN CARE BECAUSE THE EMPLOYEE HAD TO GET A NEW STAPLER MID SURGERY.
21-09-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | UNKNOWN ENDO GIA SULU | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC APPENDECTOMY, THE STAPLER WAS NOT WORKING AND WOULD NOT FIRE. IT WAS ALSO MENTIONED THAT THE STAPLER GOT STUCK TOGETHER, SO IT COULD NOT BE RELOADED. THE USER, ANOTHER TECH, REGISTERED NURSE AND THE SURGEON TRIED TO TAKE IT APART TO RELOAD A NEW STAPLE ROW BUT NONE COULD GET IT DETACHED. THE BUTTON WAS STUCK AND NOT MOVING AT ALL. THERE WAS A SHORT DELAY AROUND 90 SECONDS IN CARE BECAUSE THE EMPLOYEE HAD TO GET A NEW STAPLER MID SURGERY.
12-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC, WHILE DETACHING THE JEJUNUM, AFTER PRESSING THE GREEN BUTTON, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
21-09-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Difficult or Delayed Activation | A150102 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC APPENDECTOMY, THE STAPLER WAS NOT WORKING AND WOULD NOT FIRE. IT WAS ALSO MENTIONED THAT THE STAPLER GOT STUCK TOGETHER, SO IT COULD NOT BE RELOADED. THE USER, ANOTHER TECH, REGISTERED NURSE AND THE SURGEON TRIED TO TAKE IT APART TO RELOAD A NEW STAPLE ROW BUT NONE COULD GET IT DETACHED. THE BUTTON WAS STUCK AND NOT MOVING AT ALL. THERE WAS A SHORT DELAY AROUND 90 SECONDS IN CARE BECAUSE THE EMPLOYEE HAD TO GET A NEW STAPLER MID SURGERY.
21-09-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC APPENDECTOMY, THE STAPLER WAS NOT WORKING AND WOULD NOT FIRE. IT WAS ALSO MENTIONED THAT THE STAPLER GOT STUCK TOGETHER, SO IT COULD NOT BE RELOADED. THE USER, ANOTHER TECH, REGISTERED NURSE AND THE SURGEON TRIED TO TAKE IT APART TO RELOAD A NEW STAPLE ROW BUT NONE COULD GET IT DETACHED. THE BUTTON WAS STUCK AND NOT MOVING AT ALL. THERE WAS A SHORT DELAY AROUND 90 SECONDS IN CARE BECAUSE THE EMPLOYEE HAD TO GET A NEW STAPLER MID SURGERY.
21-09-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC APPENDECTOMY, THE STAPLER WAS NOT WORKING AND WOULD NOT FIRE. IT WAS ALSO MENTIONED THAT THE STAPLER GOT STUCK TOGETHER, SO IT COULD NOT BE RELOADED. THE USER, ANOTHER TECH, REGISTERED NURSE AND THE SURGEON TRIED TO TAKE IT APART TO RELOAD A NEW STAPLE ROW BUT NONE COULD GET IT DETACHED. THE BUTTON WAS STUCK AND NOT MOVING AT ALL. THERE WAS A SHORT DELAY AROUND 90 SECONDS IN CARE BECAUSE THE EMPLOYEE HAD TO GET A NEW STAPLER MID SURGERY.
01-01-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: THE FAILED STAPLE FORMATION HAPPENED DURING A SLEEVE GASTRECTOMY, THE RELOAD WAS USED WITH STAPLE LINE REINFORCEMENT FROM SEAMGUARD. IN THIS PROCEDURE THE DOCTOR USED GREEN GST RELOAD, AFTER FIRING AT THE END OF THE STAPLE LINE THE B SHAPED WAS NOT FORMED INSTEAD IT WAS STRAIGHT FORM. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-11-2023 | Injury | ethicon endo-surgery | 05-01-2024 | GST WHITE RELOAD, 45MM, 6 ROW | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST OP OF A LAPAROSCOPIC COLON RESECTION PROCEDURE THAT THE PATIENT WAS RETURNED TO THE OR WITH A POST OP BLEED. ONCE IN THE OR IT WAS FOUND THAT THE VASCULAR STAPLE LINE WAS OOZING. THERE WERE NO ISSUES NOTICED DURING THE PROCEDURE.
01-01-2023 | Malfunction | ethicon endo-surgery | 05-01-2024 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE A BLUE RELOAD WAS USED, AND AFTER FIRING THE B SHAPED WAS NOT FORMED INSTEAD IT WAS A STRAIGHT LINE AT THE END OF THE STAPLE LINE, AND IN THE MID OF THE STAPLE LINE AS WELL. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC COLON TUMOR RESECTION, AT THE PART OF REMOVING THE PART OF THE COLON, THERE WAS POOR STAPLE FORMATION. THE FIRING WAS A LITTLE STUCK AND THE STAPLES WERE NOT FORMED WELL. THE STAPLER FORMATION WAS INCOMPLETE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Firing Problem | A0505 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC COLON TUMOR RESECTION, AT THE PART OF REMOVING THE PART OF THE COLON, THERE WAS POOR STAPLE FORMATION. THE FIRING WAS A LITTLE STUCK AND THE STAPLES WERE NOT FORMED WELL. THE STAPLER FORMATION WAS INCOMPLETE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT UPPER LOBECTOMY, WHEN DETACHING THE INTERLOBE FISSURE OF THE RIGHT LUNG, WHEN FIRING THE SECOND STAPLE, IT COULD NOT BE FIRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-10-2023 | Injury | ethicon endo-surgery | 05-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT PULMONARY BULLECTOMY, THE OPERATION WENT SMOOTHLY. NO OBVIOUS BLEEDING AT THE RESECTION MARGIN WAS FOUND DURING THE OPERATION. THE CHEST DRAINAGE WAS INCREASED ON THE 2ND DAY AFTER THE OPERATION. ON OCTOBER 6, RE-EXAMINATION OF CHEST X-RAY AND CHEST CT SHOWED INCREASED HEMOTHORAX ON THE RIGHT SIDE AND POOR PATENCY OF CHEST TUBE DRAINAGE. IT WAS CONSIDERED TO HAVE POSTOPERATIVE INTRATHORACIC HEMORRHAGE WITH BLOOD CLOTS BLOCKING THE CHEST TUBE. THE PATIENT WAS GIVEN INTRAPLEURAL INJECTION OF UROKINASE AND ECG MONITORING. 230ML BLOODY FLUID COULD BE EXTRACTED FROM THE CHEST TUBE. REEXAMINATION OF BLOOD ROUTINE SHOWED HEMOGLOBIN CONCENTRATION 74 (G/L) , WHICH WAS CONSIDERED TO BE CAUSED BY INTRATHORACIC HEMORRHAGE. ON OCTOBER 8, RED BLOOD CELLS 2U AND FRESH FROZEN PLASMA 200ML WERE INFUSED. LATER RE-EXAMINATION OF CHEST X-RAY SHOWED DECREASED PLEURAL EFFUSION ON THE RIGHT SIDE AND THE RIGHT CHEST TUBE WAS REMOVED ON (B)(6), THE PATIENT HAD HIGH FEVER, WITH THE HIGHEST BODY TEMPERATURE OF 39.0, OCCASIONAL CHEST TIGHTNESS. THE RE-EXAMINATION SHOWED INFECTION INDICATORS INCREASED. THE ANTI-INFECTIVE TREATMENT WAS CONTINUED. THE RE-EXAMINATION OF CHEST CT SUGGESTED RIGHT INTRATHORACIC EFFUSION. ON (B)(6), THE PATIENT WAS GIVEN ULTRASOUND-GUIDED CLOSED THORACIC DRAINAGE. AFTER THE OPERATION, REDDISH FLUID WAS EXTRACTED AND THE CHEST TIGHTNESS WAS RELIEVED. ON (B)(6) THE RIGHT THORACIC DRAINAGE WAS REMOVED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-10-2023 | Injury | ethicon endo-surgery | 05-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT PULMONARY BULLECTOMY, THE OPERATION WENT SMOOTHLY. NO OBVIOUS BLEEDING AT THE RESECTION MARGIN WAS FOUND DURING THE OPERATION. THE CHEST DRAINAGE WAS INCREASED ON THE 2ND DAY AFTER THE OPERATION. ON OCTOBER 6, RE-EXAMINATION OF CHEST X-RAY AND CHEST CT SHOWED INCREASED HEMOTHORAX ON THE RIGHT SIDE AND POOR PATENCY OF CHEST TUBE DRAINAGE. IT WAS CONSIDERED TO HAVE POSTOPERATIVE INTRATHORACIC HEMORRHAGE WITH BLOOD CLOTS BLOCKING THE CHEST TUBE. THE PATIENT WAS GIVEN INTRAPLEURAL INJECTION OF UROKINASE AND ECG MONITORING. 230ML BLOODY FLUID COULD BE EXTRACTED FROM THE CHEST TUBE. REEXAMINATION OF BLOOD ROUTINE SHOWED HEMOGLOBIN CONCENTRATION 74 (G/L) , WHICH WAS CONSIDERED TO BE CAUSED BY INTRATHORACIC HEMORRHAGE. ON OCTOBER 8, RED BLOOD CELLS 2U AND FRESH FROZEN PLASMA 200ML WERE INFUSED. LATER RE-EXAMINATION OF CHEST X-RAY SHOWED DECREASED PLEURAL EFFUSION ON THE RIGHT SIDE AND THE RIGHT CHEST TUBE WAS REMOVED ON (B)(6), THE PATIENT HAD HIGH FEVER, WITH THE HIGHEST BODY TEMPERATURE OF 39.0, OCCASIONAL CHEST TIGHTNESS. THE RE-EXAMINATION SHOWED INFECTION INDICATORS INCREASED. THE ANTI-INFECTIVE TREATMENT WAS CONTINUED. THE RE-EXAMINATION OF CHEST CT SUGGESTED RIGHT INTRATHORACIC EFFUSION. ON (B)(6), THE PATIENT WAS GIVEN ULTRASOUND-GUIDED CLOSED THORACIC DRAINAGE. AFTER THE OPERATION, REDDISH FLUID WAS EXTRACTED AND THE CHEST TIGHTNESS WAS RELIEVED. ON (B)(6) THE RIGHT THORACIC DRAINAGE WAS REMOVED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-10-2023 | Injury | ethicon endo-surgery | 05-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT PULMONARY BULLECTOMY, THE OPERATION WENT SMOOTHLY. NO OBVIOUS BLEEDING AT THE RESECTION MARGIN WAS FOUND DURING THE OPERATION. THE CHEST DRAINAGE WAS INCREASED ON THE 2ND DAY AFTER THE OPERATION. ON OCTOBER 6, RE-EXAMINATION OF CHEST X-RAY AND CHEST CT SHOWED INCREASED HEMOTHORAX ON THE RIGHT SIDE AND POOR PATENCY OF CHEST TUBE DRAINAGE. IT WAS CONSIDERED TO HAVE POSTOPERATIVE INTRATHORACIC HEMORRHAGE WITH BLOOD CLOTS BLOCKING THE CHEST TUBE. THE PATIENT WAS GIVEN INTRAPLEURAL INJECTION OF UROKINASE AND ECG MONITORING. 230ML BLOODY FLUID COULD BE EXTRACTED FROM THE CHEST TUBE. REEXAMINATION OF BLOOD ROUTINE SHOWED HEMOGLOBIN CONCENTRATION 74 (G/L) , WHICH WAS CONSIDERED TO BE CAUSED BY INTRATHORACIC HEMORRHAGE. ON OCTOBER 8, RED BLOOD CELLS 2U AND FRESH FROZEN PLASMA 200ML WERE INFUSED. LATER RE-EXAMINATION OF CHEST X-RAY SHOWED DECREASED PLEURAL EFFUSION ON THE RIGHT SIDE AND THE RIGHT CHEST TUBE WAS REMOVED ON (B)(6), THE PATIENT HAD HIGH FEVER, WITH THE HIGHEST BODY TEMPERATURE OF 39.0, OCCASIONAL CHEST TIGHTNESS. THE RE-EXAMINATION SHOWED INFECTION INDICATORS INCREASED. THE ANTI-INFECTIVE TREATMENT WAS CONTINUED. THE RE-EXAMINATION OF CHEST CT SUGGESTED RIGHT INTRATHORACIC EFFUSION. ON (B)(6), THE PATIENT WAS GIVEN ULTRASOUND-GUIDED CLOSED THORACIC DRAINAGE. AFTER THE OPERATION, REDDISH FLUID WAS EXTRACTED AND THE CHEST TIGHTNESS WAS RELIEVED. ON (B)(6) THE RIGHT THORACIC DRAINAGE WAS REMOVED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-10-2023 | Injury | ethicon endo-surgery | 05-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT PULMONARY BULLECTOMY, THE OPERATION WENT SMOOTHLY. NO OBVIOUS BLEEDING AT THE RESECTION MARGIN WAS FOUND DURING THE OPERATION. THE CHEST DRAINAGE WAS INCREASED ON THE 2ND DAY AFTER THE OPERATION. ON OCTOBER 6, RE-EXAMINATION OF CHEST X-RAY AND CHEST CT SHOWED INCREASED HEMOTHORAX ON THE RIGHT SIDE AND POOR PATENCY OF CHEST TUBE DRAINAGE. IT WAS CONSIDERED TO HAVE POSTOPERATIVE INTRATHORACIC HEMORRHAGE WITH BLOOD CLOTS BLOCKING THE CHEST TUBE. THE PATIENT WAS GIVEN INTRAPLEURAL INJECTION OF UROKINASE AND ECG MONITORING. 230ML BLOODY FLUID COULD BE EXTRACTED FROM THE CHEST TUBE. REEXAMINATION OF BLOOD ROUTINE SHOWED HEMOGLOBIN CONCENTRATION 74 (G/L) , WHICH WAS CONSIDERED TO BE CAUSED BY INTRATHORACIC HEMORRHAGE. ON OCTOBER 8, RED BLOOD CELLS 2U AND FRESH FROZEN PLASMA 200ML WERE INFUSED. LATER RE-EXAMINATION OF CHEST X-RAY SHOWED DECREASED PLEURAL EFFUSION ON THE RIGHT SIDE AND THE RIGHT CHEST TUBE WAS REMOVED ON (B)(6), THE PATIENT HAD HIGH FEVER, WITH THE HIGHEST BODY TEMPERATURE OF 39.0, OCCASIONAL CHEST TIGHTNESS. THE RE-EXAMINATION SHOWED INFECTION INDICATORS INCREASED. THE ANTI-INFECTIVE TREATMENT WAS CONTINUED. THE RE-EXAMINATION OF CHEST CT SUGGESTED RIGHT INTRATHORACIC EFFUSION. ON (B)(6), THE PATIENT WAS GIVEN ULTRASOUND-GUIDED CLOSED THORACIC DRAINAGE. AFTER THE OPERATION, REDDISH FLUID WAS EXTRACTED AND THE CHEST TIGHTNESS WAS RELIEVED. ON (B)(6) THE RIGHT THORACIC DRAINAGE WAS REMOVED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-10-2023 | Injury | ethicon endo-surgery | 05-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Patient Device Interaction Problem | A01 | Fever | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT PULMONARY BULLECTOMY, THE OPERATION WENT SMOOTHLY. NO OBVIOUS BLEEDING AT THE RESECTION MARGIN WAS FOUND DURING THE OPERATION. THE CHEST DRAINAGE WAS INCREASED ON THE 2ND DAY AFTER THE OPERATION. ON OCTOBER 6, RE-EXAMINATION OF CHEST X-RAY AND CHEST CT SHOWED INCREASED HEMOTHORAX ON THE RIGHT SIDE AND POOR PATENCY OF CHEST TUBE DRAINAGE. IT WAS CONSIDERED TO HAVE POSTOPERATIVE INTRATHORACIC HEMORRHAGE WITH BLOOD CLOTS BLOCKING THE CHEST TUBE. THE PATIENT WAS GIVEN INTRAPLEURAL INJECTION OF UROKINASE AND ECG MONITORING. 230ML BLOODY FLUID COULD BE EXTRACTED FROM THE CHEST TUBE. REEXAMINATION OF BLOOD ROUTINE SHOWED HEMOGLOBIN CONCENTRATION 74 (G/L) , WHICH WAS CONSIDERED TO BE CAUSED BY INTRATHORACIC HEMORRHAGE. ON OCTOBER 8, RED BLOOD CELLS 2U AND FRESH FROZEN PLASMA 200ML WERE INFUSED. LATER RE-EXAMINATION OF CHEST X-RAY SHOWED DECREASED PLEURAL EFFUSION ON THE RIGHT SIDE AND THE RIGHT CHEST TUBE WAS REMOVED ON (B)(6), THE PATIENT HAD HIGH FEVER, WITH THE HIGHEST BODY TEMPERATURE OF 39.0, OCCASIONAL CHEST TIGHTNESS. THE RE-EXAMINATION SHOWED INFECTION INDICATORS INCREASED. THE ANTI-INFECTIVE TREATMENT WAS CONTINUED. THE RE-EXAMINATION OF CHEST CT SUGGESTED RIGHT INTRATHORACIC EFFUSION. ON (B)(6), THE PATIENT WAS GIVEN ULTRASOUND-GUIDED CLOSED THORACIC DRAINAGE. AFTER THE OPERATION, REDDISH FLUID WAS EXTRACTED AND THE CHEST TIGHTNESS WAS RELIEVED. ON (B)(6) THE RIGHT THORACIC DRAINAGE WAS REMOVED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-10-2023 | Injury | ethicon endo-surgery | 05-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT PULMONARY BULLECTOMY, THE OPERATION WENT SMOOTHLY. NO OBVIOUS BLEEDING AT THE RESECTION MARGIN WAS FOUND DURING THE OPERATION. THE CHEST DRAINAGE WAS INCREASED ON THE 2ND DAY AFTER THE OPERATION. ON OCTOBER 6, RE-EXAMINATION OF CHEST X-RAY AND CHEST CT SHOWED INCREASED HEMOTHORAX ON THE RIGHT SIDE AND POOR PATENCY OF CHEST TUBE DRAINAGE. IT WAS CONSIDERED TO HAVE POSTOPERATIVE INTRATHORACIC HEMORRHAGE WITH BLOOD CLOTS BLOCKING THE CHEST TUBE. THE PATIENT WAS GIVEN INTRAPLEURAL INJECTION OF UROKINASE AND ECG MONITORING. 230ML BLOODY FLUID COULD BE EXTRACTED FROM THE CHEST TUBE. REEXAMINATION OF BLOOD ROUTINE SHOWED HEMOGLOBIN CONCENTRATION 74 (G/L) , WHICH WAS CONSIDERED TO BE CAUSED BY INTRATHORACIC HEMORRHAGE. ON OCTOBER 8, RED BLOOD CELLS 2U AND FRESH FROZEN PLASMA 200ML WERE INFUSED. LATER RE-EXAMINATION OF CHEST X-RAY SHOWED DECREASED PLEURAL EFFUSION ON THE RIGHT SIDE AND THE RIGHT CHEST TUBE WAS REMOVED ON (B)(6), THE PATIENT HAD HIGH FEVER, WITH THE HIGHEST BODY TEMPERATURE OF 39.0, OCCASIONAL CHEST TIGHTNESS. THE RE-EXAMINATION SHOWED INFECTION INDICATORS INCREASED. THE ANTI-INFECTIVE TREATMENT WAS CONTINUED. THE RE-EXAMINATION OF CHEST CT SUGGESTED RIGHT INTRATHORACIC EFFUSION. ON (B)(6), THE PATIENT WAS GIVEN ULTRASOUND-GUIDED CLOSED THORACIC DRAINAGE. AFTER THE OPERATION, REDDISH FLUID WAS EXTRACTED AND THE CHEST TIGHTNESS WAS RELIEVED. ON (B)(6) THE RIGHT THORACIC DRAINAGE WAS REMOVED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-10-2023 | Injury | ethicon endo-surgery | 05-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Patient Device Interaction Problem | A01 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT PULMONARY BULLECTOMY, THE OPERATION WENT SMOOTHLY. NO OBVIOUS BLEEDING AT THE RESECTION MARGIN WAS FOUND DURING THE OPERATION. THE CHEST DRAINAGE WAS INCREASED ON THE 2ND DAY AFTER THE OPERATION. ON OCTOBER 6, RE-EXAMINATION OF CHEST X-RAY AND CHEST CT SHOWED INCREASED HEMOTHORAX ON THE RIGHT SIDE AND POOR PATENCY OF CHEST TUBE DRAINAGE. IT WAS CONSIDERED TO HAVE POSTOPERATIVE INTRATHORACIC HEMORRHAGE WITH BLOOD CLOTS BLOCKING THE CHEST TUBE. THE PATIENT WAS GIVEN INTRAPLEURAL INJECTION OF UROKINASE AND ECG MONITORING. 230ML BLOODY FLUID COULD BE EXTRACTED FROM THE CHEST TUBE. REEXAMINATION OF BLOOD ROUTINE SHOWED HEMOGLOBIN CONCENTRATION 74 (G/L) , WHICH WAS CONSIDERED TO BE CAUSED BY INTRATHORACIC HEMORRHAGE. ON OCTOBER 8, RED BLOOD CELLS 2U AND FRESH FROZEN PLASMA 200ML WERE INFUSED. LATER RE-EXAMINATION OF CHEST X-RAY SHOWED DECREASED PLEURAL EFFUSION ON THE RIGHT SIDE AND THE RIGHT CHEST TUBE WAS REMOVED ON (B)(6), THE PATIENT HAD HIGH FEVER, WITH THE HIGHEST BODY TEMPERATURE OF 39.0, OCCASIONAL CHEST TIGHTNESS. THE RE-EXAMINATION SHOWED INFECTION INDICATORS INCREASED. THE ANTI-INFECTIVE TREATMENT WAS CONTINUED. THE RE-EXAMINATION OF CHEST CT SUGGESTED RIGHT INTRATHORACIC EFFUSION. ON (B)(6), THE PATIENT WAS GIVEN ULTRASOUND-GUIDED CLOSED THORACIC DRAINAGE. AFTER THE OPERATION, REDDISH FLUID WAS EXTRACTED AND THE CHEST TIGHTNESS WAS RELIEVED. ON (B)(6) THE RIGHT THORACIC DRAINAGE WAS REMOVED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-10-2023 | Injury | ethicon endo-surgery | 05-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Patient Device Interaction Problem | A01 | Pleural Effusion | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT PULMONARY BULLECTOMY, THE OPERATION WENT SMOOTHLY. NO OBVIOUS BLEEDING AT THE RESECTION MARGIN WAS FOUND DURING THE OPERATION. THE CHEST DRAINAGE WAS INCREASED ON THE 2ND DAY AFTER THE OPERATION. ON OCTOBER 6, RE-EXAMINATION OF CHEST X-RAY AND CHEST CT SHOWED INCREASED HEMOTHORAX ON THE RIGHT SIDE AND POOR PATENCY OF CHEST TUBE DRAINAGE. IT WAS CONSIDERED TO HAVE POSTOPERATIVE INTRATHORACIC HEMORRHAGE WITH BLOOD CLOTS BLOCKING THE CHEST TUBE. THE PATIENT WAS GIVEN INTRAPLEURAL INJECTION OF UROKINASE AND ECG MONITORING. 230ML BLOODY FLUID COULD BE EXTRACTED FROM THE CHEST TUBE. REEXAMINATION OF BLOOD ROUTINE SHOWED HEMOGLOBIN CONCENTRATION 74 (G/L) , WHICH WAS CONSIDERED TO BE CAUSED BY INTRATHORACIC HEMORRHAGE. ON OCTOBER 8, RED BLOOD CELLS 2U AND FRESH FROZEN PLASMA 200ML WERE INFUSED. LATER RE-EXAMINATION OF CHEST X-RAY SHOWED DECREASED PLEURAL EFFUSION ON THE RIGHT SIDE AND THE RIGHT CHEST TUBE WAS REMOVED ON (B)(6), THE PATIENT HAD HIGH FEVER, WITH THE HIGHEST BODY TEMPERATURE OF 39.0, OCCASIONAL CHEST TIGHTNESS. THE RE-EXAMINATION SHOWED INFECTION INDICATORS INCREASED. THE ANTI-INFECTIVE TREATMENT WAS CONTINUED. THE RE-EXAMINATION OF CHEST CT SUGGESTED RIGHT INTRATHORACIC EFFUSION. ON (B)(6), THE PATIENT WAS GIVEN ULTRASOUND-GUIDED CLOSED THORACIC DRAINAGE. AFTER THE OPERATION, REDDISH FLUID WAS EXTRACTED AND THE CHEST TIGHTNESS WAS RELIEVED. ON (B)(6) THE RIGHT THORACIC DRAINAGE WAS REMOVED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-10-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A COLOSTOMY TAKEDOWN PROCEDURE, WHEN THE SURGEON USED THE DEVICE, THE STAPLE LINE FAILED. THE SURGEON ATTEMPTED IT WITH TWO DIFFERENT STAPLERS AND TWO DIFFERENT LOADS, BUT THE STAPLER WOULD NOT CLOSE. THE SURGEON HAD TO RECONNECT THE BOWEL BY HAND.
16-10-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | GIA | Difficult to Open or Close | A051104 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A COLOSTOMY TAKEDOWN PROCEDURE, WHEN THE SURGEON USED THE DEVICE, THE STAPLE LINE FAILED. THE SURGEON ATTEMPTED IT WITH TWO DIFFERENT STAPLERS AND TWO DIFFERENT LOADS, BUT THE STAPLER WOULD NOT CLOSE. THE SURGEON HAD TO RECONNECT THE BOWEL BY HAND.
16-10-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A COLOSTOMY TAKEDOWN PROCEDURE, WHEN THE SURGEON USED THE DEVICE, THE STAPLE LINE FAILED. THE SURGEON ATTEMPTED IT WITH TWO DIFFERENT STAPLERS AND TWO DIFFERENT LOADS, BUT THE STAPLER WOULD NOT CLOSE. THE SURGEON HAD TO RECONNECT THE BOWEL BY HAND.
16-10-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | GIA | Difficult to Open or Close | A051104 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A COLOSTOMY TAKEDOWN PROCEDURE, WHEN THE SURGEON USED THE DEVICE, THE STAPLE LINE FAILED. THE SURGEON ATTEMPTED IT WITH TWO DIFFERENT STAPLERS AND TWO DIFFERENT LOADS, BUT THE STAPLER WOULD NOT CLOSE. THE SURGEON HAD TO RECONNECT THE BOWEL BY HAND.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOTIC LAPAROSCOPIC SLEEVE GASTRECTOMY, ON FIRING THE POWERED STAPLER, ON THE FIRST LOAD, THERE WAS AN ODD CRUNCHING SOUND, SO IT WAS PULLED OUT AND CHECKED AND WENT BACK AT IT AGAIN. AS WHEN THE USER STARTED TO FIRE, THERE WAS A CRUNCHED SOUND AGAIN AND THE RELOAD JERKS, AND ARTICULATED ALL THE WAY TO THE RIGHT ON ITS OWN AND THEN NEUTRAL. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY USING A MANUAL RETRACTION TOOL. A NEW STAPLER WAS THEN USED TO RESOLVE THE ISSUE AND FINISHED THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOTIC LAPAROSCOPIC SLEEVE GASTRECTOMY, ON FIRING THE POWERED STAPLER, ON THE FIRST LOAD, THERE WAS AN ODD CRUNCHING SOUND, SO IT WAS PULLED OUT AND CHECKED AND WENT BACK AT IT AGAIN. AS WHEN THE USER STARTED TO FIRE, THERE WAS A CRUNCHED SOUND AGAIN AND THE RELOAD JERKS, AND ARTICULATED ALL THE WAY TO THE RIGHT ON ITS OWN AND THEN NEUTRAL. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY USING A MANUAL RETRACTION TOOL. A NEW STAPLER WAS THEN USED TO RESOLVE THE ISSUE AND FINISHED THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOTIC LAPAROSCOPIC SLEEVE GASTRECTOMY, ON FIRING THE POWERED STAPLER, ON THE FIRST LOAD, THERE WAS AN ODD CRUNCHING SOUND, SO IT WAS PULLED OUT AND CHECKED AND WENT BACK AT IT AGAIN. AS WHEN THE USER STARTED TO FIRE, THERE WAS A CRUNCHED SOUND AGAIN AND THE RELOAD JERKS, AND ARTICULATED ALL THE WAY TO THE RIGHT ON ITS OWN AND THEN NEUTRAL. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY USING A MANUAL RETRACTION TOOL. A NEW STAPLER WAS THEN USED TO RESOLVE THE ISSUE AND FINISHED THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOTIC LAPAROSCOPIC SLEEVE GASTRECTOMY, ON FIRING THE POWERED STAPLER, ON THE FIRST LOAD, THERE WAS AN ODD CRUNCHING SOUND, SO IT WAS PULLED OUT AND CHECKED AND WENT BACK AT IT AGAIN. AS WHEN THE USER STARTED TO FIRE, THERE WAS A CRUNCHED SOUND AGAIN AND THE RELOAD JERKS, AND ARTICULATED ALL THE WAY TO THE RIGHT ON ITS OWN AND THEN NEUTRAL. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY USING A MANUAL RETRACTION TOOL. A NEW STAPLER WAS THEN USED TO RESOLVE THE ISSUE AND FINISHED THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOTIC LAPAROSCOPIC SLEEVE GASTRECTOMY, ON FIRING THE POWERED STAPLER, ON THE FIRST LOAD, THERE WAS AN ODD CRUNCHING SOUND, SO IT WAS PULLED OUT AND CHECKED AND WENT BACK AT IT AGAIN. AS WHEN THE USER STARTED TO FIRE, THERE WAS A CRUNCHED SOUND AGAIN AND THE RELOAD JERKS, AND ARTICULATED ALL THE WAY TO THE RIGHT ON ITS OWN AND THEN NEUTRAL. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY USING A MANUAL RETRACTION TOOL. A NEW STAPLER WAS THEN USED TO RESOLVE THE ISSUE AND FINISHED THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOTIC LAPAROSCOPIC SLEEVE GASTRECTOMY, ON FIRING THE POWERED STAPLER, ON THE FIRST LOAD, THERE WAS AN ODD CRUNCHING SOUND, SO IT WAS PULLED OUT AND CHECKED AND WENT BACK AT IT AGAIN. AS WHEN THE USER STARTED TO FIRE, THERE WAS A CRUNCHED SOUND AGAIN AND THE RELOAD JERKS, AND ARTICULATED ALL THE WAY TO THE RIGHT ON ITS OWN AND THEN NEUTRAL. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY USING A MANUAL RETRACTION TOOL. A NEW STAPLER WAS THEN USED TO RESOLVE THE ISSUE AND FINISHED THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOTIC LAPAROSCOPIC SLEEVE GASTRECTOMY, ON FIRING THE POWERED STAPLER, ON THE FIRST LOAD, THERE WAS AN ODD CRUNCHING SOUND, SO IT WAS PULLED OUT AND CHECKED AND WENT BACK AT IT AGAIN. AS WHEN THE USER STARTED TO FIRE, THERE WAS A CRUNCHED SOUND AGAIN AND THE RELOAD JERKS, AND ARTICULATED ALL THE WAY TO THE RIGHT ON ITS OWN AND THEN NEUTRAL. THE JAWS OF THE DEVICE LOCKED ON TISSUE AND WAS REMOVED BY USING A MANUAL RETRACTION TOOL. A NEW STAPLER WAS THEN USED TO RESOLVE THE ISSUE AND FINISHED THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROTOMY WITH LEFT ATRIAL APPENDAGE RESECTION, THE RELOAD WAS ATTACHED TO THE HANDLE AND THE TISSUE WAS CLAMPED BUT THE FIRING COULD NOT BE DONE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. BOTH THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROTOMY WITH LEFT ATRIAL APPENDAGE RESECTION, THE RELOAD WAS ATTACHED TO THE HANDLE AND THE TISSUE WAS CLAMPED BUT THE FIRING COULD NOT BE DONE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. BOTH THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-10-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY, BEFORE THE STAPLER WAS INSERTED TO THORACIC CAVITY, THE TIP JAWS WERE UNABLE TO BE OPEN/CLOSE. ANOTHER POWERED STAPLER WAS USED TO COMPLETE THE CASE. AFTER THE SURGERY, THERE WAS A POWER HANDLE ERROR WITH RED CIRCLE, SLASH AND EXCLAMATION MARK ON THE REAR MONITOR. THE SURGERY WAS PERFORMED USING ANOTHER HANDLE AND SHELL. THE ADAPTER WAS STILL USED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AN ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PR OGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
26-10-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY, BEFORE THE STAPLER WAS INSERTED TO THORACIC CAVITY, THE TIP JAWS WERE UNABLE TO BE OPEN/CLOSE. ANOTHER POWERED STAPLER WAS USED TO COMPLETE THE CASE. AFTER THE SURGERY, THERE WAS A POWER HANDLE ERROR WITH RED CIRCLE, SLASH AND EXCLAMATION MARK ON THE REAR MONITOR. THE SURGERY WAS PERFORMED USING ANOTHER HANDLE AND SHELL. THE ADAPTER WAS STILL USED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AN ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PR OGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
26-10-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY, BEFORE THE STAPLER WAS INSERTED TO THORACIC CAVITY, THE TIP JAWS WERE UNABLE TO BE OPEN/CLOSE. ANOTHER POWERED STAPLER WAS USED TO COMPLETE THE CASE. AFTER THE SURGERY, THERE WAS A POWER HANDLE ERROR WITH RED CIRCLE, SLASH AND EXCLAMATION MARK ON THE REAR MONITOR. THE SURGERY WAS PERFORMED USING ANOTHER HANDLE AND SHELL. THE ADAPTER WAS STILL USED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AN ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PR OGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
29-11-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBE THORACOTOMY, WHILE RESECTIONING THE LUNG PARENCHYMAL, A LOUD CRACKING SOUND WAS HEARD COMING FROM THE HANDLE. THE STAPLER DID NOT FIRE, IT WAS NOT POSSIBLE TO TRIGGER HANDLE AND RELOAD COULD NOT BE RELEASED. A NEW HANDLE WAS USED TO COMPLETE THE CASE. THE SURGERY WAS DELAYED BY A FEW MINUTES. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED THE CLAMPING MECHANISM WAS DEFORMED AND THICK TISSUE.
29-11-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBE THORACOTOMY, WHILE RESECTIONING THE LUNG PARENCHYMAL, A LOUD CRACKING SOUND WAS HEARD COMING FROM THE HANDLE. THE STAPLER DID NOT FIRE, IT WAS NOT POSSIBLE TO TRIGGER HANDLE AND RELOAD COULD NOT BE RELEASED. A NEW HANDLE WAS USED TO COMPLETE THE CASE. THE SURGERY WAS DELAYED BY A FEW MINUTES. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED THE CLAMPING MECHANISM WAS DEFORMED AND THICK TISSUE.
06-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRECTOMY, DURING GASTRIC RESECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRECTOMY, DURING GASTRIC RESECTION, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-10-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN CHECKING BEFORE STARTING THE OPERATION, THE HANDLE WAS NOT WORKING. WHEN IT WAS DIS CONNECTED FROM THE CHARGER, THE POWER COULD NOT BE TURNED ON AND THERE WAS NO SOUND. IT WAS NOTED THAT THERE WAS NO STRANGE INDICATION AT THE CHARGER. IT WAS NOT USED ON THE OPERATION AND A MANUAL HANDLE WAS USED INSTEAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AN ANALYSIS OF THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
10-10-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN CHECKING BEFORE STARTING THE OPERATION, THE HANDLE WAS NOT WORKING. WHEN IT WAS DIS CONNECTED FROM THE CHARGER, THE POWER COULD NOT BE TURNED ON AND THERE WAS NO SOUND. IT WAS NOTED THAT THERE WAS NO STRANGE INDICATION AT THE CHARGER. IT WAS NOT USED ON THE OPERATION AND A MANUAL HANDLE WAS USED INSTEAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AN ANALYSIS OF THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
02-11-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE PACKAGING WAS TORN IN THE BOX. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
02-11-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Packaging Problem | A0205 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE PACKAGING WAS TORN IN THE BOX. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, WHILE CLAMPING PYLORUS, DURING THE FIRST FIRING, THE STAPLER SHOWED EXCESS FORCE WHILE FIRING WAS PERFORMED, STOPPING IN THE MIDDLE OF THE STAPLE AND NOT FINISHING. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, WHILE CLAMPING PYLORUS, DURING THE FIRST FIRING, THE STAPLER SHOWED EXCESS FORCE WHILE FIRING WAS PERFORMED, STOPPING IN THE MIDDLE OF THE STAPLE AND NOT FINISHING. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-10-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE RELOAD COULD NOT BE LOADED ONTO THE HANDLE. HENCE, THE DEVICE COULD NOT BE FIRED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE AND FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
08-10-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE RELOAD COULD NOT BE LOADED ONTO THE HANDLE. HENCE, THE DEVICE COULD NOT BE FIRED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE AND FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
11-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | UNKNOWN STAPLING DEVICE | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, WHILE CLAMPING PYLORUS, DURING THE FIRST FIRING, THE STAPLER SHOWED EXCESS FORCE WHILE FIRING WAS PERFORMED, STOPPING IN THE MIDDLE OF THE STAPLE AND NOT FINISHING. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. ANOTHER CARTRIDGE WAS USED TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | UNKNOWN EGIA ADAPTER | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, WHILE CLAMPING PYLORUS, DURING THE FIRST FIRING, THE STAPLER SHOWED EXCESS FORCE WHILE FIRING WAS PERFORMED, STOPPING IN THE MIDDLE OF THE STAPLE AND NOT FINISHING. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THERE WAS NO PROBLEM WITH THE CARTRIDGE CONNECTION AND OPENING/CLOSING CHECK, BUT WHEN THE TISSUE WAS CLAMPED, CARTRIDGE ERROR WITH YELLOW CARTRIDGE STATUS INDICATOR AND EXCLAMATION MARK WAS DISPLAYED AND THE RELEASE WAS NOT POSSIBLE. WHEN THE TISSUE WAS LOCKED, A POWERED APPROACH WAS IMPLEMENTED, AFTER RELEASE, MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THERE WAS NO PROBLEM WITH THE CARTRIDGE CONNECTION AND OPENING/CLOSING CHECK, BUT WHEN THE TISSUE WAS CLAMPED, CARTRIDGE ERROR WITH YELLOW CARTRIDGE STATUS INDICATOR AND EXCLAMATION MARK WAS DISPLAYED AND THE RELEASE WAS NOT POSSIBLE. WHEN THE TISSUE WAS LOCKED, A POWERED APPROACH WAS IMPLEMENTED, AFTER RELEASE, MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THERE WAS NO PROBLEM WITH THE CARTRIDGE CONNECTION AND OPENING/CLOSING CHECK, BUT WHEN THE TISSUE WAS CLAMPED, CARTRIDGE ERROR WITH YELLOW CARTRIDGE STATUS INDICATOR AND EXCLAMATION MARK WAS DISPLAYED AND THE RELEASE WAS NOT POSSIBLE. WHEN THE TISSUE WAS LOCKED, A POWERED APPROACH WAS IMPLEMENTED, AFTER RELEASE, MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THERE WAS NO PROBLEM WITH THE CARTRIDGE CONNECTION AND OPENING/CLOSING CHECK, BUT WHEN THE TISSUE WAS CLAMPED, CARTRIDGE ERROR WITH YELLOW CARTRIDGE STATUS INDICATOR AND EXCLAMATION MARK WAS DISPLAYED AND THE RELEASE WAS NOT POSSIBLE. WHEN THE TISSUE WAS LOCKED, A POWERED APPROACH WAS IMPLEMENTED, AFTER RELEASE, MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THERE WAS NO PROBLEM WITH THE CARTRIDGE CONNECTION AND OPENING/CLOSING CHECK, BUT WHEN THE TISSUE WAS CLAMPED, CARTRIDGE ERROR WITH YELLOW CARTRIDGE STATUS INDICATOR AND EXCLAMATION MARK WAS DISPLAYED AND THE RELEASE WAS NOT POSSIBLE. WHEN THE TISSUE WAS LOCKED, A POWERED APPROACH WAS IMPLEMENTED, AFTER RELEASE, MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THERE WAS NO PROBLEM WITH THE CARTRIDGE CONNECTION AND OPENING/CLOSING CHECK, BUT WHEN THE TISSUE WAS CLAMPED, CARTRIDGE ERROR WITH YELLOW CARTRIDGE STATUS INDICATOR AND EXCLAMATION MARK WAS DISPLAYED AND THE RELEASE WAS NOT POSSIBLE. WHEN THE TISSUE WAS LOCKED, A POWERED APPROACH WAS IMPLEMENTED, AFTER RELEASE, MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMICOLECTOMY, THERE WAS NO PROBLEM WITH THE CARTRIDGE CONNECTION AND OPENING/CLOSING CHECK, BUT WHEN THE TISSUE WAS CLAMPED, CARTRIDGE ERROR WITH YELLOW CARTRIDGE STATUS INDICATOR AND EXCLAMATION MARK WAS DISPLAYED AND THE RELEASE WAS NOT POSSIBLE. WHEN THE TISSUE WAS LOCKED, A POWERED APPROACH WAS IMPLEMENTED, AFTER RELEASE, MANUAL HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-10-2023 | Malfunction | intuitive surgical | 05-01-2024 | SUREFORM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LIVER RESECTION SURGICAL PROCEDURE, SUREFORM 60 STAPLER ONLY FIRED 9 TIMES OUT OF IT'S INTENDED 12. UPON LOADING IT AFTER THE 9TH FIRE THE SUREFORM WOULD NOT FIRE. THE CUSTOMER REMOVED THE STAPLER AND CONTINUED WITH ANOTHER OF THE SAME. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY OR DELAY.
08-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LAPAROTOMY TOTAL GASTRECTOMY, WHILE USING THE MANUAL HANDLE WITH A 45MM RELOAD, THE TISSUE WAS APPROACHED AND THE GREEN BUTTON WAS PRESSED, BUT THE HANDLE WAS STIFF AND COULD NOT BE SQUEEZED, SO THE HANDLE WAS REPLACED, AND THE SAME CARTRIDGE WAS USED WITHOUT ANY PROBLEM. NO PATIENT INJURY.
30-11-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE OF THE DEVICE WHEN CHECKING THE OPENING AND CLOSING OF THE RELOAD, THE JAW COULD NOT BE COMPLETELY TIGHTENED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE, SHELL AND ADAPTER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
08-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTROINTESTINAL SUTURE SEPARATION, BEFORE THE GASTROINTESTINAL SUTURE RESECTION, EVEN AFTER PROPERLY SETTING THE DEVICE UP, THE HANDLE DID NOT MOVE. FIRING COULD NOT BE PERFORMED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY, WHILE ON SUTURING, RESECTION, FIRING, THE DEVICE WAS CLAMPED AND WAS EVEN ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. IT WAS MENTIONED THAT THE JAW OF THE DEVICE WAS FIXATED TO TISSUE AND WAS ABLE TO BE REMOVED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY, WHILE ON SUTURING, RESECTION, FIRING, THE DEVICE WAS CLAMPED AND WAS EVEN ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. IT WAS MENTIONED THAT THE JAW OF THE DEVICE WAS FIXATED TO TISSUE AND WAS ABLE TO BE REMOVED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHEN THE MESH WAS FIXATED, AFTER FIRING THE TACK, IT DID NOT DETACH FROM THE SHAFT, AND THE TACK WAS PULLED AND REMOVED. THE FIRING WAS ALSO STIFF. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHEN THE MESH WAS FIXATED, AFTER FIRING THE TACK, IT DID NOT DETACH FROM THE SHAFT, AND THE TACK WAS PULLED AND REMOVED. THE FIRING WAS ALSO STIFF. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY FOR LUNG MALIGNANT TUMOR RESECTION, WHILE ON SUTURING, RESECTION, FIRING, THE DEVICE WAS CLAMPED AND WAS EVEN ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. IT WAS MENTIONED THAT THE JAW OF THE DEVICE WAS FIXATED TO TISSUE AND WAS ABLE TO BE REMOVED. A NEW RELOAD WAS USED WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY FOR LUNG MALIGNANT TUMOR RESECTION, WHILE ON SUTURING, RESECTION, FIRING, THE DEVICE WAS CLAMPED AND WAS EVEN ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. IT WAS MENTIONED THAT THE JAW OF THE DEVICE WAS FIXATED TO TISSUE AND WAS ABLE TO BE REMOVED. A NEW RELOAD WAS USED WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY FOR LUNG MALIGNANT TUMOR RESECTION, WHILE ON SUTURING, RESECTION, FIRING, THE DEVICE WAS CLAMPED AND WAS EVEN ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. IT WAS MENTIONED THAT THE JAW OF THE DEVICE WAS FIXATED TO TISSUE AND WAS ABLE TO BE REMOVED. A NEW RELOAD WAS USED WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY FOR LUNG MALIGNANT TUMOR RESECTION, WHILE ON SUTURING, RESECTION, FIRING, THE DEVICE WAS CLAMPED AND WAS EVEN ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. IT WAS MENTIONED THAT THE JAW OF THE DEVICE WAS FIXATED TO TISSUE AND WAS ABLE TO BE REMOVED. A NEW RELOAD WAS USED WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY FOR LUNG MALIGNANT TUMOR RESECTION, WHILE ON SUTURING, RESECTION, FIRING, THE DEVICE WAS CLAMPED AND WAS EVEN ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. IT WAS MENTIONED THAT THE JAW OF THE DEVICE WAS FIXATED TO TISSUE AND WAS ABLE TO BE REMOVED. A NEW RELOAD WAS USED WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 05-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY FOR LUNG MALIGNANT TUMOR RESECTION, WHILE ON SUTURING, RESECTION, FIRING, THE DEVICE WAS CLAMPED AND WAS EVEN ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED. IT WAS MENTIONED THAT THE JAW OF THE DEVICE WAS FIXATED TO TISSUE AND WAS ABLE TO BE REMOVED. A NEW RELOAD WAS USED WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SUPERIOR POLAR OESOGASTRECTOMY FOR A LARGE PART OF THE DISSECTION WAS PERFORMED WITH CO ELIO, FOLLOWED BY A LAPAROTOMY TO BRING THE TWO DISSECTION PLANES TOGETHER, THE DEVICE MALFUNCTIONED. GASTRIC STAPLING WAS PLANNED IN ORDER TO MAKE A GASTRIC TUBE BUT THE RELOAD SEIZED BEFORE STAPLING COULD BE COMPLETED. THE 45 MM RELOAD DETACHED FROM THE FORCEPS AND HAD TO BE REMOVED BY THE SURGEON FROM INSIDE THE PATIENT. ANOTHER HANDLE FROM THE SAME REFERENCE AND A NEW RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SUPERIOR POLAR OESOGASTRECTOMY FOR A LARGE PART OF THE DISSECTION WAS PERFORMED WITH CO ELIO, FOLLOWED BY A LAPAROTOMY TO BRING THE TWO DISSECTION PLANES TOGETHER, THE DEVICE MALFUNCTIONED. GASTRIC STAPLING WAS PLANNED IN ORDER TO MAKE A GASTRIC TUBE BUT THE RELOAD SEIZED BEFORE STAPLING COULD BE COMPLETED. THE 45 MM RELOAD DETACHED FROM THE FORCEPS AND HAD TO BE REMOVED BY THE SURGEON FROM INSIDE THE PATIENT. ANOTHER HANDLE FROM THE SAME REFERENCE AND A NEW RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SUPERIOR POLAR OESOGASTRECTOMY FOR A LARGE PART OF THE DISSECTION WAS PERFORMED WITH CO ELIO, FOLLOWED BY A LAPAROTOMY TO BRING THE TWO DISSECTION PLANES TOGETHER, THE DEVICE MALFUNCTIONED. GASTRIC STAPLING WAS PLANNED IN ORDER TO MAKE A GASTRIC TUBE BUT THE RELOAD SEIZED BEFORE STAPLING COULD BE COMPLETED. THE 45 MM RELOAD DETACHED FROM THE FORCEPS AND HAD TO BE REMOVED BY THE SURGEON FROM INSIDE THE PATIENT. ANOTHER HANDLE FROM THE SAME REFERENCE AND A NEW RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 05-01-2024 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SUPERIOR POLAR OESOGASTRECTOMY FOR A LARGE PART OF THE DISSECTION WAS PERFORMED WITH CO ELIO, FOLLOWED BY A LAPAROTOMY TO BRING THE TWO DISSECTION PLANES TOGETHER, THE DEVICE MALFUNCTIONED. GASTRIC STAPLING WAS PLANNED IN ORDER TO MAKE A GASTRIC TUBE BUT THE RELOAD SEIZED BEFORE STAPLING COULD BE COMPLETED. THE 45 MM RELOAD DETACHED FROM THE FORCEPS AND HAD TO BE REMOVED BY THE SURGEON FROM INSIDE THE PATIENT. ANOTHER HANDLE FROM THE SAME REFERENCE AND A NEW RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
16-11-2023 | Malfunction | us surgical puerto rico | 06-01-2024 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, AT THE TIME OF DETACHING THE BODY OF THE STOMACH, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. THE DEVICE WAS RETURNED WITHOUT ACCOMPANYING INFORMATION. PLI:20 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AND FIRING OVER TISSUE THAT IS BEYONDTHE RECOMMENDED THICKNESS RANGE.
16-11-2023 | Malfunction | us surgical puerto rico | 06-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, AT THE TIME OF DETACHING THE BODY OF THE STOMACH, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. THE DEVICE WAS RETURNED WITHOUT ACCOMPANYING INFORMATION. PLI:20 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AND FIRING OVER TISSUE THAT IS BEYONDTHE RECOMMENDED THICKNESS RANGE.
11-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AND UNKNOWN SURGERY, THE DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AFTER THE PULMONARY PARENCHYMA WAS CLAMPED, THE DEVICE FIRED, BUT IT STOPPED MOVING MIDWAY. HENCE, NO STAPLES ARE DEPLOYED ON THE DISTAL PART. THE JAWS ALSO LOCKED ON TISSUE. THE BLADE WORKED BY REBOOTING THE DEVICE. TO COMPLETE THE CASE, ANOTHER RELOAD WAS USED. IT WAS NOTED THAT THE HANDLE AND ADAPTER WORKED WITH OTHER DEVICES. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AFTER THE PULMONARY PARENCHYMA WAS CLAMPED, THE DEVICE FIRED, BUT IT STOPPED MOVING MIDWAY. HENCE, NO STAPLES ARE DEPLOYED ON THE DISTAL PART. THE JAWS ALSO LOCKED ON TISSUE. THE BLADE WORKED BY REBOOTING THE DEVICE. TO COMPLETE THE CASE, ANOTHER RELOAD WAS USED. IT WAS NOTED THAT THE HANDLE AND ADAPTER WORKED WITH OTHER DEVICES. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE A BLUE RELOAD WAS USED, AND AFTER FIRING THE B SHAPED WAS NOT FORMED INSTEAD IT WAS A STRAIGHT LINE AT THE END OF THE STAPLE LINE, AND IN THE MID OF THE STAPLE LINE AS WELL. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGUS SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | BLUE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-11-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 08-01-2024 | CAPSURE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR THE CAPSURE FASTENERS WERE NOT FIXATING THE MESH ADEQUATELY AND HAD TO BE REMOVED. ANOTHER CAPSURE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | ILS 21MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY SURGERY, COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-11-2023 | Injury | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A RESPIRATORY SURGERY, LEAKAGE OCCURRED FROM THE LUNG PARENCHYMA RESECTION. PNEUMOTHORAX OCCURRED. RE-OPERATION WAS PERFORMED. THE BLEEDING OCCURRED FROM THE STAPLE CUT END. IT WAS CLAMPED BEFORE FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. NO FURTHER INFORMATION IS AVAILABLE.
29-11-2023 | Injury | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER A RESPIRATORY SURGERY, LEAKAGE OCCURRED FROM THE LUNG PARENCHYMA RESECTION. PNEUMOTHORAX OCCURRED. RE-OPERATION WAS PERFORMED. THE BLEEDING OCCURRED FROM THE STAPLE CUT END. IT WAS CLAMPED BEFORE FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. NO FURTHER INFORMATION IS AVAILABLE.
29-11-2023 | Injury | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Pneumothorax | Event Description: IT WAS REPORTED THAT AFTER A RESPIRATORY SURGERY, LEAKAGE OCCURRED FROM THE LUNG PARENCHYMA RESECTION. PNEUMOTHORAX OCCURRED. RE-OPERATION WAS PERFORMED. THE BLEEDING OCCURRED FROM THE STAPLE CUT END. IT WAS CLAMPED BEFORE FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. NO FURTHER INFORMATION IS AVAILABLE.
22-11-2023 | Injury | ethicon endo-surgery | 08-01-2024 | FLEX 60 ARTICULATING | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN RESPIRATORY SURGERY, A BLACK CARTRIDGE WAS USED. A POST-OPERATIVE LEAKAGE OCCURRED FROM THE REAL INCISION AREA. POSTOPERATIVE HEMOTHORAX OCCURRED, RESULTED IN REOPERATION, BUT THERE WAS BLEEDING FROM THE STAPLE DISSECTION (NOT THE STAPLE LINE, BUT THE TRANSECTED EDGE). THERE WAS PRIOR COMPRESSION. THE SURGEON SAID, ¿THE CARTRIDGE BLACK MAY NOT HAVE BEEN SUFFICIENTLY COMPRESSED.¿ ADDITIONAL SUTURES WERE PERFORMED TO COMPLETE THE CASE. NO FURTHER INFORMATION IS AVAILABLE.
22-11-2023 | Injury | ethicon endo-surgery | 08-01-2024 | FLEX 60 ARTICULATING | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN RESPIRATORY SURGERY, A BLACK CARTRIDGE WAS USED. A POST-OPERATIVE LEAKAGE OCCURRED FROM THE REAL INCISION AREA. POSTOPERATIVE HEMOTHORAX OCCURRED, RESULTED IN REOPERATION, BUT THERE WAS BLEEDING FROM THE STAPLE DISSECTION (NOT THE STAPLE LINE, BUT THE TRANSECTED EDGE). THERE WAS PRIOR COMPRESSION. THE SURGEON SAID, ¿THE CARTRIDGE BLACK MAY NOT HAVE BEEN SUFFICIENTLY COMPRESSED.¿ ADDITIONAL SUTURES WERE PERFORMED TO COMPLETE THE CASE. NO FURTHER INFORMATION IS AVAILABLE.
27-10-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, WHEN EXAMINING THE POWER HANDLE, A HANDLE ERROR (A GRAPHIC WITH A HANDLE AND A RED CIRCLE WITH A LINE GOING THROUGH IT) APPEARED IMMEDIATELY AFTER REMOVING THE HANDLE FROM THE BATTERY CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL. THE CAUSE OF THE OBSERVED ERROR WAS DUE TO A SOFTWARE RESTRICTIONS ON THE CAPACITY OF THE QUEUE.
27-10-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING SERVICING, WHEN EXAMINING THE POWER HANDLE, A HANDLE ERROR (A GRAPHIC WITH A HANDLE AND A RED CIRCLE WITH A LINE GOING THROUGH IT) APPEARED IMMEDIATELY AFTER REMOVING THE HANDLE FROM THE BATTERY CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL. THE CAUSE OF THE OBSERVED ERROR WAS DUE TO A SOFTWARE RESTRICTIONS ON THE CAPACITY OF THE QUEUE.
27-10-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE WOULD NOT FIRE AND MAKING AN UNUSUAL CLICK/GRINDING NOISE. THEY CHANGED STAPLE LOAD AND IT DID THE SAME THING. GOT A NEW DEVICE TO FINISH CASE. NO PATIENT HARM WAS REPORTED.
27-10-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE WOULD NOT FIRE AND MAKING AN UNUSUAL CLICK/GRINDING NOISE. THEY CHANGED STAPLE LOAD AND IT DID THE SAME THING. GOT A NEW DEVICE TO FINISH CASE. NO PATIENT HARM WAS REPORTED.
06-09-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, WHILE CHARGING, THE DEVICE FAILED TO CHARGE. A MANUAL HANDLE WAS USED INSTEAD. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE VENTED.
06-09-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, WHILE CHARGING, THE DEVICE FAILED TO CHARGE. A MANUAL HANDLE WAS USED INSTEAD. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE VENTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, SURGEON WENT TO FIRE BLUE RELOAD WITH OUR SLR BUTTRESS. IT WAS THE THIRD RELOAD/FIRING. AS HE FIRED WE NOTICED A SLIGHT MOVEMENT OF THE STALE LINE WHILE THE DEVICE WAS FIRING. IT APPEARED AS IF THE KNIFE BLADE WAS DRAGGING THE SLR/STAPLE LINE. AS HE OPENED THE JAWS WE NOTICED MALFORMED STAPLES WITH OOZING. HE PROCEEDED WITH THE CASE USING THE SAME STAPLER AND MORE BLUE BUT WITHOUT SLR AND HAD NO ISSUES. THE TECH SWISHED AND DRIED THE DEVICE BETWEEN FIRINGS. NO PATIENT HARM WAS REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | GST BLUE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, SURGEON WENT TO FIRE BLUE RELOAD WITH OUR SLR BUTTRESS. IT WAS THE THIRD RELOAD/FIRING. AS HE FIRED WE NOTICED A SLIGHT MOVEMENT OF THE STALE LINE WHILE THE DEVICE WAS FIRING. IT APPEARED AS IF THE KNIFE BLADE WAS DRAGGING THE SLR/STAPLE LINE. AS HE OPENED THE JAWS WE NOTICED MALFORMED STAPLES WITH OOZING. HE PROCEEDED WITH THE CASE USING THE SAME STAPLER AND MORE BLUE BUT WITHOUT SLR AND HAD NO ISSUES. THE TECH SWISHED AND DRIED THE DEVICE BETWEEN FIRINGS. NO PATIENT HARM WAS REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | GST BLUE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, SURGEON WENT TO FIRE BLUE RELOAD WITH OUR SLR BUTTRESS. IT WAS THE THIRD RELOAD/FIRING. AS HE FIRED WE NOTICED A SLIGHT MOVEMENT OF THE STALE LINE WHILE THE DEVICE WAS FIRING. IT APPEARED AS IF THE KNIFE BLADE WAS DRAGGING THE SLR/STAPLE LINE. AS HE OPENED THE JAWS WE NOTICED MALFORMED STAPLES WITH OOZING. HE PROCEEDED WITH THE CASE USING THE SAME STAPLER AND MORE BLUE BUT WITHOUT SLR AND HAD NO ISSUES. THE TECH SWISHED AND DRIED THE DEVICE BETWEEN FIRINGS. NO PATIENT HARM WAS REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | GST BLUE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY, SURGEON WENT TO FIRE BLUE RELOAD WITH OUR SLR BUTTRESS. IT WAS THE THIRD RELOAD/FIRING. AS HE FIRED WE NOTICED A SLIGHT MOVEMENT OF THE STALE LINE WHILE THE DEVICE WAS FIRING. IT APPEARED AS IF THE KNIFE BLADE WAS DRAGGING THE SLR/STAPLE LINE. AS HE OPENED THE JAWS WE NOTICED MALFORMED STAPLES WITH OOZING. HE PROCEEDED WITH THE CASE USING THE SAME STAPLER AND MORE BLUE BUT WITHOUT SLR AND HAD NO ISSUES. THE TECH SWISHED AND DRIED THE DEVICE BETWEEN FIRINGS. NO PATIENT HARM WAS REPORTED.
01-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC PROCEDURE CONVERTED TO OPEN PANCREATECTOMY AND SPLENECTOMY, THE STAPLES THAT HAD PROPER FORMATION MISFIRED AND DID NOT FIRE ALL THE WAY. THE OTHER STAPLES HAD POOR FORMATION. A DIFFERENT STAPLER FROM ANOTHER MANUFACTURER WAS USED TO REPAIR THE PREVIOUSLY MISFIRED STAPLE LINE.
25-10-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | ENDOPATH |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: SECOND DEVICE USED ON THE SAME CASE DID NOT WORK. FIRST ONE WAS REPORTED ON (B)(4).
07-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE COULD NOT BE FIRED DURING USE. BOTH THE KNIFE REVERSE SWITCH AND THE MANUAL OVERRIDE LEVER WERE USED BUT THE JAWS DID NOT OPEN. AFTER THAT, AS THE SURGEON MOVED HANDLE BACK AND FORTH, THE JAWS OPENED. THE DEVICE WAS USED ON JEJUNUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE COULD NOT BE FIRED DURING USE. BOTH THE KNIFE REVERSE SWITCH AND THE MANUAL OVERRIDE LEVER WERE USED BUT THE JAWS DID NOT OPEN. AFTER THAT, AS THE SURGEON MOVED HANDLE BACK AND FORTH, THE JAWS OPENED. THE DEVICE WAS USED ON JEJUNUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE COULD NOT BE FIRED DURING USE. BOTH THE KNIFE REVERSE SWITCH AND THE MANUAL OVERRIDE LEVER WERE USED BUT THE JAWS DID NOT OPEN. AFTER THAT, AS THE SURGEON MOVED HANDLE BACK AND FORTH, THE JAWS OPENED. THE DEVICE WAS USED ON JEJUNUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
07-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLECTOMY, THE DEVICE COULD NOT BE FIRED DURING USE. BOTH THE KNIFE REVERSE SWITCH AND THE MANUAL OVERRIDE LEVER WERE USED BUT THE JAWS DID NOT OPEN. AFTER THAT, AS THE SURGEON MOVED HANDLE BACK AND FORTH, THE JAWS OPENED. THE DEVICE WAS USED ON JEJUNUM. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
25-10-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | ENDOPATH |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: DEVICE IS DEFECTIVE DID NOT WORK.
06-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK SURGERY, THE DEVICE, ONCE ACTIVATED, DID NOT REOPEN. NO FURTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCES REPORTED.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY COLON RESECTION (COLECTOMY), DURING INTESTINAL RESECTION, THE GREEN LIGHT DID NOT ILLUMINATE AND THE DEVICE WAS UNABLE TO ENTER FIRING MODE WITH THE THREE RELOADS. A NEW HANDLE, SHELL AND ADAPTER WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT CLAMPING MECHANISMS WERE DEFORMED AND THE DEVICE WAS FIRED IN TISSUE THICKER THAN INDICATED.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY COLON RESECTION (COLECTOMY), DURING INTESTINAL RESECTION, THE GREEN LIGHT DID NOT ILLUMINATE AND THE DEVICE WAS UNABLE TO ENTER FIRING MODE WITH THE THREE RELOADS. A NEW HANDLE, SHELL AND ADAPTER WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT CLAMPING MECHANISMS WERE DEFORMED AND THE DEVICE WAS FIRED IN TISSUE THICKER THAN INDICATED.
15-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TRANSVERSE COLECTOMY, IN FUNCTIONAL END-TO-END ANASTOMOSIS OF THE INTESTINAL TRACT, THE DEVICE FIRED, BUT THE MUCOUS MEMBRANE GOT STUCK AT THE TIP. WHEN THE USER PULLED THE STAPLER TO REMOVE THE DEVICE, THE STUCK MEMBRANE WAS TORN, AND THE INTESTINAL TRACT WAS DAMAGED. THE MUCOSAL REGION WAS RESECTED WITH A SCISSOR, AND A SUTURE WAS USED TO FIX THE DAMAGE. IT WAS NOTED THAT THE HANDLE WORKED WITH OTHER RELOADS OF THE SAME TYPE.
15-09-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE UNIT WOULD NOT TAKE CHARGE AS THERE WAS A BATTERY ERROR SHOWING. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. THE BATTERY VENTS WERE NOTED TO BE OPEN.
15-09-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE UNIT WOULD NOT TAKE CHARGE AS THERE WAS A BATTERY ERROR SHOWING. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. THE BATTERY VENTS WERE NOTED TO BE OPEN.
01-01-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: WHEN THE SURGEON FIRED THE EQUIPMENT, THE KNIFE WAS STOCK IN THE MIDDLE AND COULD NOT REVERSE BY REVERSE BUTTON. AT THE END, THE SURGEON REMOVES THE EQUIPMENT BY USING THE MANUAL OVERRIDE BUTTON. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: WHEN THE SURGEON FIRED THE EQUIPMENT, THE KNIFE WAS STOCK IN THE MIDDLE AND COULD NOT REVERSE BY REVERSE BUTTON. AT THE END, THE SURGEON REMOVES THE EQUIPMENT BY USING THE MANUAL OVERRIDE BUTTON. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: WHEN THE SURGEON FIRED THE EQUIPMENT, THE KNIFE WAS STOCK IN THE MIDDLE AND COULD NOT REVERSE BY REVERSE BUTTON. AT THE END, THE SURGEON REMOVES THE EQUIPMENT BY USING THE MANUAL OVERRIDE BUTTON. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN TOTAL GASTRECTOMY, THERE WAS AN ISSUE WHEN LOADING THE DEVICE. THE LOCK RING WAS DAMAGED, AND THE RELOAD COULD NOT BE CONNECTED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN TOTAL GASTRECTOMY, THERE WAS AN ISSUE WHEN LOADING THE DEVICE. THE LOCK RING WAS DAMAGED, AND THE RELOAD COULD NOT BE CONNECTED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Component or Accessory Incompatibility | A1701 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN TOTAL GASTRECTOMY, THERE WAS AN ISSUE WHEN LOADING THE DEVICE. THE LOCK RING WAS DAMAGED, AND THE RELOAD COULD NOT BE CONNECTED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN TOTAL GASTRECTOMY, THERE WAS AN ISSUE WHEN LOADING THE DEVICE. THE LOCK RING WAS DAMAGED, AND THE RELOAD COULD NOT BE CONNECTED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
01-01-2024 | Malfunction | ethicon endo-surgery | 08-01-2024 | GST BLUE RELOAD, 60MM, 6 ROW | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RECTAL RESECTION, THERE WERE NO PROBLEMS WITH THE MOVEMENT OF THE MAIN DEVICE, STAPLE FIRING, AND STAPLE FORMATION. HOWEVER, AFTER FIRING, SHAVINGS-LIKE MATERIAL FROM THE GST60B FELL INTO THE BODY AND WAS RETRIEVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
10-09-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | BLUE RELOAD FOR ECHELON 60 | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, OPEN THE PACKING AND NOTED THAT THE STAPLE DRIVERS FELL OFF . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Failure to Anastomose | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Abscess | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Adhesion(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Anemia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Bacterial Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Dehydration | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Purulent Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Emotional Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Erythema | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Foreign Body Reaction | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Hematoma | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Hemorrhage/Blood Loss/Bleeding | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Unspecified Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Muscle Weakness/Atrophy | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Necrosis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Pleural Effusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Rash | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Shock | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Electrolyte Imbalance | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Hernia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Ulcer | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Anxiety | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Distress | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Discomfort | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Depression | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Impaired Healing | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Ambulation or Postural Difficulties | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Ascites | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Blister | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Unspecified Tissue Injury | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Mechanics Altered | A0507 | Pneumoperitoneum | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Adhesion(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Anemia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Bacterial Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Dehydration | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Purulent Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Emotional Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Erythema | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Foreign Body Reaction | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Muscle Weakness/Atrophy | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Pleural Effusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Rash | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Shock | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Electrolyte Imbalance | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Hernia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Ulcer | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Anxiety | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Distress | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Discomfort | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Depression | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Impaired Healing | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Ambulation or Postural Difficulties | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Ascites | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Blister | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Adverse Event Without Identified Device or Use Problem | A24 | Pneumoperitoneum | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Failure to Anastomose | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Abscess | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Adhesion(s) | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Anemia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Bacterial Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Dehydration | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Purulent Discharge | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Emotional Changes | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Erythema | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Foreign Body Reaction | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Hematoma | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Hemorrhage/Blood Loss/Bleeding | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Unspecified Infection | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Inflammation | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Muscle Weakness/Atrophy | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Necrosis | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Pain | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Pleural Effusion | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Rash | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Shock | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Electrolyte Imbalance | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Hernia | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Ulcer | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Anxiety | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Distress | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Discomfort | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Depression | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Impaired Healing | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Ambulation or Postural Difficulties | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Ascites | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Abdominal Distention | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Blister | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Appropriate Term / Code Not Available | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
12-09-2021 | Death | us surgical puerto rico | 08-01-2024 | TA | Insufficient Information | E2401 | Pneumoperitoneum | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, AND DIVERTING LOOP ILEOSTOMY. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.
14-11-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | GST BLUE RELOAD, 45MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-09-2023 | Injury | us surgical puerto rico | 08-01-2024 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, PATIENT WAS RETURNED BACK FOR SURGERY DUE TO SYMPTOMS ANASTOMOTIC LEAK. PATIENT HOSPITALIZATION WAS NOTED.
08-09-2023 | Injury | us surgical puerto rico | 08-01-2024 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, PATIENT WAS RETURNED BACK FOR SURGERY DUE TO SYMPTOMS ANASTOMOTIC LEAK. PATIENT HOSPITALIZATION WAS NOTED.
08-09-2023 | Injury | us surgical puerto rico | 08-01-2024 | UNKNOWN GIA DST PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, PATIENT WAS RETURNED BACK FOR SURGERY DUE TO SYMPTOMS ANASTOMOTIC LEAK. PATIENT HOSPITALIZATION WAS NOTED.
08-09-2023 | Injury | us surgical puerto rico | 08-01-2024 | UNKNOWN GIA DST PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, PATIENT WAS RETURNED BACK FOR SURGERY DUE TO SYMPTOMS ANASTOMOTIC LEAK. PATIENT HOSPITALIZATION WAS NOTED.
01-12-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | CARTRIDGE FINAL ASSY 60MM GREEN | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 08-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, ON POST OPERATIVE LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE PATIENT HAD TO BE CONTROLLED AND RE-OPERATED DUE TO A LARGE AMOUNT OF BLOOD IN THE DRAIN (ABOUT 50 ML) AND A CONTINUOUS AIR LEAK. THERE WAS BLEEDING FROM THE STAPLE LINE. ON THE REOPERATION, THE SURGEON FOUND DEHISCENCE IN STAPLE LINE. THE SURGEON NOTICED THAT STAPLES WERE NOT FORMED IN B-SHAPE, STAPLES WERE CRUSHED AND TWISTED. THE SURGEON RESECTED ADDITIONAL TISSUE WITH DAMAGED STAPLE LINE TO CLOSE. THE SURGEON USED A NEW BLACK RELOAD TO CLOSE AND CUT OFF THE TISSUE THAT HAD SEPARATED. THERE WAS A TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE EVENT HAS LEAD TO OR EXTEND PATIENT HOSPITALIZATION
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 08-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, ON POST OPERATIVE LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE PATIENT HAD TO BE CONTROLLED AND RE-OPERATED DUE TO A LARGE AMOUNT OF BLOOD IN THE DRAIN (ABOUT 50 ML) AND A CONTINUOUS AIR LEAK. THERE WAS BLEEDING FROM THE STAPLE LINE. ON THE REOPERATION, THE SURGEON FOUND DEHISCENCE IN STAPLE LINE. THE SURGEON NOTICED THAT STAPLES WERE NOT FORMED IN B-SHAPE, STAPLES WERE CRUSHED AND TWISTED. THE SURGEON RESECTED ADDITIONAL TISSUE WITH DAMAGED STAPLE LINE TO CLOSE. THE SURGEON USED A NEW BLACK RELOAD TO CLOSE AND CUT OFF THE TISSUE THAT HAD SEPARATED. THERE WAS A TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE EVENT HAS LEAD TO OR EXTEND PATIENT HOSPITALIZATION
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 08-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, ON POST OPERATIVE LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE PATIENT HAD TO BE CONTROLLED AND RE-OPERATED DUE TO A LARGE AMOUNT OF BLOOD IN THE DRAIN (ABOUT 50 ML) AND A CONTINUOUS AIR LEAK. THERE WAS BLEEDING FROM THE STAPLE LINE. ON THE REOPERATION, THE SURGEON FOUND DEHISCENCE IN STAPLE LINE. THE SURGEON NOTICED THAT STAPLES WERE NOT FORMED IN B-SHAPE, STAPLES WERE CRUSHED AND TWISTED. THE SURGEON RESECTED ADDITIONAL TISSUE WITH DAMAGED STAPLE LINE TO CLOSE. THE SURGEON USED A NEW BLACK RELOAD TO CLOSE AND CUT OFF THE TISSUE THAT HAD SEPARATED. THERE WAS A TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE EVENT HAS LEAD TO OR EXTEND PATIENT HOSPITALIZATION
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 08-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, ON POST OPERATIVE LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE PATIENT HAD TO BE CONTROLLED AND RE-OPERATED DUE TO A LARGE AMOUNT OF BLOOD IN THE DRAIN (ABOUT 50 ML) AND A CONTINUOUS AIR LEAK. THERE WAS BLEEDING FROM THE STAPLE LINE. ON THE REOPERATION, THE SURGEON FOUND DEHISCENCE IN STAPLE LINE. THE SURGEON NOTICED THAT STAPLES WERE NOT FORMED IN B-SHAPE, STAPLES WERE CRUSHED AND TWISTED. THE SURGEON RESECTED ADDITIONAL TISSUE WITH DAMAGED STAPLE LINE TO CLOSE. THE SURGEON USED A NEW BLACK RELOAD TO CLOSE AND CUT OFF THE TISSUE THAT HAD SEPARATED. THERE WAS A TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE EVENT HAS LEAD TO OR EXTEND PATIENT HOSPITALIZATION
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 08-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, ON POST OPERATIVE LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE PATIENT HAD TO BE CONTROLLED AND RE-OPERATED DUE TO A LARGE AMOUNT OF BLOOD IN THE DRAIN (ABOUT 50 ML) AND A CONTINUOUS AIR LEAK. THERE WAS BLEEDING FROM THE STAPLE LINE. ON THE REOPERATION, THE SURGEON FOUND DEHISCENCE IN STAPLE LINE. THE SURGEON NOTICED THAT STAPLES WERE NOT FORMED IN B-SHAPE, STAPLES WERE CRUSHED AND TWISTED. THE SURGEON RESECTED ADDITIONAL TISSUE WITH DAMAGED STAPLE LINE TO CLOSE. THE SURGEON USED A NEW BLACK RELOAD TO CLOSE AND CUT OFF THE TISSUE THAT HAD SEPARATED. THERE WAS A TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE EVENT HAS LEAD TO OR EXTEND PATIENT HOSPITALIZATION
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 08-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, ON POST OPERATIVE LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE PATIENT HAD TO BE CONTROLLED AND RE-OPERATED DUE TO A LARGE AMOUNT OF BLOOD IN THE DRAIN (ABOUT 50 ML) AND A CONTINUOUS AIR LEAK. THERE WAS BLEEDING FROM THE STAPLE LINE. ON THE REOPERATION, THE SURGEON FOUND DEHISCENCE IN STAPLE LINE. THE SURGEON NOTICED THAT STAPLES WERE NOT FORMED IN B-SHAPE, STAPLES WERE CRUSHED AND TWISTED. THE SURGEON RESECTED ADDITIONAL TISSUE WITH DAMAGED STAPLE LINE TO CLOSE. THE SURGEON USED A NEW BLACK RELOAD TO CLOSE AND CUT OFF THE TISSUE THAT HAD SEPARATED. THERE WAS A TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE EVENT HAS LEAD TO OR EXTEND PATIENT HOSPITALIZATION
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 08-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, ON POST OPERATIVE LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE PATIENT HAD TO BE CONTROLLED AND RE-OPERATED DUE TO A LARGE AMOUNT OF BLOOD IN THE DRAIN (ABOUT 50 ML) AND A CONTINUOUS AIR LEAK. THERE WAS BLEEDING FROM THE STAPLE LINE. ON THE REOPERATION, THE SURGEON FOUND DEHISCENCE IN STAPLE LINE. THE SURGEON NOTICED THAT STAPLES WERE NOT FORMED IN B-SHAPE, STAPLES WERE CRUSHED AND TWISTED. THE SURGEON RESECTED ADDITIONAL TISSUE WITH DAMAGED STAPLE LINE TO CLOSE. THE SURGEON USED A NEW BLACK RELOAD TO CLOSE AND CUT OFF THE TISSUE THAT HAD SEPARATED. THERE WAS A TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE EVENT HAS LEAD TO OR EXTEND PATIENT HOSPITALIZATION
12-12-2023 | Injury | celestica de monterrey s.a de c.v | 08-01-2024 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, ON POST OPERATIVE LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY, THE PATIENT HAD TO BE CONTROLLED AND RE-OPERATED DUE TO A LARGE AMOUNT OF BLOOD IN THE DRAIN (ABOUT 50 ML) AND A CONTINUOUS AIR LEAK. THERE WAS BLEEDING FROM THE STAPLE LINE. ON THE REOPERATION, THE SURGEON FOUND DEHISCENCE IN STAPLE LINE. THE SURGEON NOTICED THAT STAPLES WERE NOT FORMED IN B-SHAPE, STAPLES WERE CRUSHED AND TWISTED. THE SURGEON RESECTED ADDITIONAL TISSUE WITH DAMAGED STAPLE LINE TO CLOSE. THE SURGEON USED A NEW BLACK RELOAD TO CLOSE AND CUT OFF THE TISSUE THAT HAD SEPARATED. THERE WAS A TISSUE DAMAGE AND UNANTICIPATED TISSUE LOSS AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE EVENT HAS LEAD TO OR EXTEND PATIENT HOSPITALIZATION
12-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLORECTAL SURGERY, THE SURGICAL NURSE DISCOVERED THAT THE WHITE TIP OF THE OPEN STAPLER WAS MISSING. THE SURGICAL STAFF STARTED LOOKING FOR IT AND FOUND THE WHITE TIP ON THE FLOOR. THERE WAS NO ACTION NEEDED; THE LAST FIRING OF THE STAPLER WORKED FINE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | intuitive surgical | 08-01-2024 | ENDOWRIST | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
14-11-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL GASTRECTOMY WITH GASTROJEJUNAL ANASTOMOSIS, WHEN THE STAPLER WAS USED FOR CLOSURE, IT WAS FOUND THAT THE HANDLE COULD NOT BE SQUEEZE AND DID NOT FIRE AFTER PRESSING THE GREEN BUTTON, AND STILL FAILED TO FIRE AFTER REPLACING THE RELOAD. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-11-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL GASTRECTOMY WITH GASTROJEJUNAL ANASTOMOSIS, WHEN THE STAPLER WAS USED FOR CLOSURE, IT WAS FOUND THAT THE HANDLE COULD NOT BE SQUEEZE AND DID NOT FIRE AFTER PRESSING THE GREEN BUTTON, AND STILL FAILED TO FIRE AFTER REPLACING THE RELOAD. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-11-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL GASTRECTOMY WITH GASTROJEJUNAL ANASTOMOSIS, WHEN THE STAPLER WAS USED FOR CLOSURE, IT WAS FOUND THAT THE HANDLE COULD NOT BE SQUEEZE AND DID NOT FIRE AFTER PRESSING THE GREEN BUTTON, AND STILL FAILED TO FIRE AFTER REPLACING THE RELOAD. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-11-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL GASTRECTOMY WITH GASTROJEJUNAL ANASTOMOSIS, WHEN THE STAPLER WAS USED FOR CLOSURE, IT WAS FOUND THAT THE HANDLE COULD NOT BE SQUEEZE AND DID NOT FIRE AFTER PRESSING THE GREEN BUTTON, AND STILL FAILED TO FIRE AFTER REPLACING THE RELOAD. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
01-01-2023 | Injury | ethicon endo-surgery | 08-01-2024 | ECHELON 3000 60MM STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC CYST PROSTATECTOMY, THERE WAS A LEAK FROM STAPLER USAGE AFTER AN UNKNOWN UROLOGY CASE. PATIENT WAS ADMITTED DUE TO BLADDER CANCER AND BROUGHT IN FOR A RADICAL CYSTECTOMY. STAPLE LEAK NEAR THE BOWEL WAS DISCOVERED ON PATIENT AFTER A CT SCAN. IT RESULTED IN AN ELEVATED WHITE BLOOD CELL COUNT AND THE 4PO INTAKE GI REFUNCTION TO RETURN. PATIENT REQUIRED SURGERY CONSULTATION AND HAD TO HAVE A ILEOSTOMY. SURGEON MENTIONED AFTER THAT FOLLOWING HIS 5 YEARS OF USING THIS PRODUCT, HE FEELS THAT IT ISN'T AS USER FRIENDLY LIKE IT USED TO BE.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, AT THE TIME OF FIRST FIRING, THE STAPLER DID NOT ALLOW FI RING, THE HANDLE COULD NOT BE SQUEEZED. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE SAME RELOAD. THIS RESULTED IN PROCEDURE TIME EXTENSION OF MORE THAN 30 MINUTES.
11-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN AN OPEN LAPAROTOMY GASTROSTOMY, DURING DUODENECTOMY, FROM THE MIDDLE OF THE FIRING, THE HANDLE COULD NOT BE SQUEEZED AND UNABLE TO FIRE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-10-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE ANVIL DID NOT CLOSE. THERE WAS NO PATIENT INVOLVED.
12-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT LOWER LOBECTOMY PROCEDURE, IN DETACHING THE INTERLOBAR FISSURE IN THE LOWER LEFT LUNG, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON; HOWEVER, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
25-10-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE ANVIL DID NOT CLOSE. THERE WAS NO PATIENT INVOLVED.
14-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC SEGMENTECTOMY, WHEN DETACHING THE LUNG SEGMENT TISSUE, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 08-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE WAS CLOSED AND READY TO FIRE, BUT THE DEVICE WOULD NOT FIRE. SURGEON TRIED TO OPEN THE DEVICE JAWS AND IT WOULD NOT OPEN. SURGEON EVENTUALLY WAS ABLE TO OPEN THE JAWS OF THE DEVICE TO REMOVE FROM THE PATIENT. NO FURTHER INFORMATION WAS PROVIDED. THERE WERE NO ADVERSE CONSEQUENCE REPORTED.
01-01-2023 | Injury | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, INITIALLY THE STAPLER DID NOT TRIGGER/FIRE, BUT AFTER SEVERAL ATTEMPTS WITH A GREEN MAGAZINE THERE WAS A CRACKING NOISE IN THE INSTRUMENT. THE STAPLER TRIGGERED, BUT WITHOUT CUTTING - IT WAS ONLY STAPLED. THE POSTERIOR WALL OF THE STOMACH WAS TORN. THE PATIENT BE TREATED SURGICALLY.
01-01-2023 | Injury | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, INITIALLY THE STAPLER DID NOT TRIGGER/FIRE, BUT AFTER SEVERAL ATTEMPTS WITH A GREEN MAGAZINE THERE WAS A CRACKING NOISE IN THE INSTRUMENT. THE STAPLER TRIGGERED, BUT WITHOUT CUTTING - IT WAS ONLY STAPLED. THE POSTERIOR WALL OF THE STOMACH WAS TORN. THE PATIENT BE TREATED SURGICALLY.
01-01-2023 | Injury | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, INITIALLY THE STAPLER DID NOT TRIGGER/FIRE, BUT AFTER SEVERAL ATTEMPTS WITH A GREEN MAGAZINE THERE WAS A CRACKING NOISE IN THE INSTRUMENT. THE STAPLER TRIGGERED, BUT WITHOUT CUTTING - IT WAS ONLY STAPLED. THE POSTERIOR WALL OF THE STOMACH WAS TORN. THE PATIENT BE TREATED SURGICALLY.
01-01-2023 | Injury | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Noise, Audible | A0508 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, INITIALLY THE STAPLER DID NOT TRIGGER/FIRE, BUT AFTER SEVERAL ATTEMPTS WITH A GREEN MAGAZINE THERE WAS A CRACKING NOISE IN THE INSTRUMENT. THE STAPLER TRIGGERED, BUT WITHOUT CUTTING - IT WAS ONLY STAPLED. THE POSTERIOR WALL OF THE STOMACH WAS TORN. THE PATIENT BE TREATED SURGICALLY.
01-01-2023 | Injury | ethicon endo-surgery | 08-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, INITIALLY THE STAPLER DID NOT TRIGGER/FIRE, BUT AFTER SEVERAL ATTEMPTS WITH A GREEN MAGAZINE THERE WAS A CRACKING NOISE IN THE INSTRUMENT. THE STAPLER TRIGGERED, BUT WITHOUT CUTTING - IT WAS ONLY STAPLED. THE POSTERIOR WALL OF THE STOMACH WAS TORN. THE PATIENT BE TREATED SURGICALLY.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR GASTRIC CANCER, WHILE DETACHING THE GASTRIC BODY, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-11-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE OPERATOR COULD NOT LOAD THE RELOAD. ADDITIONAL RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS RELOAD THICK TISSUE. THE RELOAD WAS PRE-FIRED. THE CLAMPING MECHANISM WAS DEFORMED.
01-11-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE OPERATOR COULD NOT LOAD THE RELOAD. ADDITIONAL RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS RELOAD THICK TISSUE. THE RELOAD WAS PRE-FIRED. THE CLAMPING MECHANISM WAS DEFORMED.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLORECTAL CANCER RESECTION, IN DETACHING THE COLON TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER PURPLE RELOAD WAS USED WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRECTOMY PROCEDURE, WHEN THE JAW OF THE DEVICE WAS CLOSED, THE JAW OPENING WAS LARGER THAN USUAL. HENCE, FIRING WAS NOT PERFORMED. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
08-01-2024 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS UNABLE TO SHOT THE RELOAD. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS RELOAD THICK TISSUE. THE CLAMPING MECHANISM WAS DEFORMED.
08-01-2024 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS UNABLE TO SHOT THE RELOAD. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS RELOAD THICK TISSUE. THE CLAMPING MECHANISM WAS DEFORMED.
08-01-2024 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS UNABLE TO SHOT THE RELOAD. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS RELOAD THICK TISSUE. THE CLAMPING MECHANISM WAS DEFORMED.
08-01-2024 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS UNABLE TO SHOT THE RELOAD. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS RELOAD THICK TISSUE. THE CLAMPING MECHANISM WAS DEFORMED.
13-10-2023 | Malfunction | intuitive surgical | 08-01-2024 | SUREFORM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LIVER RESECTION SURGICAL PROCEDURE, SUREFORM 60 STAPLER BLACK RELOAD ONLY FIRED 9 TIMES OUT OF IT'S INTENDED 12. UPON LOADING IT AFTER THE 9TH FIRE THE SUREFORM BLACK RELOAD WOULD NOT FIRE. THE CUSTOMER REMOVED THE STAPLER INSTRUMENT AND CONTINUED WITH ANOTHER OF THE SAME. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY OR DELAY.
28-11-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY WITH ROBOT SUPPORT, WHILE USING THE FIRST RELOAD, THE DUODENUM WAS APPROACHED, BUT FIRING COULD NOT BE PERFORMED AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. WHEN THE PROXIMAL SIDE OF THE CARTRIDGE WAS CHECKED, TWO STAPLES WERE PROTRUDING. A NEW RELOAD WITH THE SAME HANDLE WAS USED TO RESOLVE THE FIRST ISSUE. WHEN THE SECOND RELOAD WAS USED FOR ESOPHAGUS-JEJUNUM RECONSTRUCTION, THE JAWS WERE COMPLETELY CLOSED, BUT COULD NOT BE INSERTED THROUGH THE 12MM PORT DUE TO POOR JAWS CLOSURE. ALTHOUGH IT WAS POSSIBLE TO INSERT THE DEVICE THROUGH THE 12MM PORT BY MANUALLY CLOSING THE JAWS, THE USED OF THE DEVICE WAS STOPPED. A NEW 60MM RELOAD WITH THE SAME HANDLE WAS TAKEN OUT AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-11-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY WITH ROBOT SUPPORT, WHILE USING THE FIRST RELOAD, THE DUODENUM WAS APPROACHED, BUT FIRING COULD NOT BE PERFORMED AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. WHEN THE PROXIMAL SIDE OF THE CARTRIDGE WAS CHECKED, TWO STAPLES WERE PROTRUDING. A NEW RELOAD WITH THE SAME HANDLE WAS USED TO RESOLVE THE FIRST ISSUE. WHEN THE SECOND RELOAD WAS USED FOR ESOPHAGUS-JEJUNUM RECONSTRUCTION, THE JAWS WERE COMPLETELY CLOSED, BUT COULD NOT BE INSERTED THROUGH THE 12MM PORT DUE TO POOR JAWS CLOSURE. ALTHOUGH IT WAS POSSIBLE TO INSERT THE DEVICE THROUGH THE 12MM PORT BY MANUALLY CLOSING THE JAWS, THE USED OF THE DEVICE WAS STOPPED. A NEW 60MM RELOAD WITH THE SAME HANDLE WAS TAKEN OUT AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) RIGHT LOWER LOBECTOMY, WHEN THE RELOAD WAS CONNECTED TO THE STAPLER, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE, ADAPTER, AND SHELL TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) RIGHT LOWER LOBECTOMY, WHEN THE RELOAD WAS CONNECTED TO THE STAPLER, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE, ADAPTER, AND SHELL TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A SUPERIOR LOBECTOMY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS STUCK. ANOTHER HANDLE WAS OPENED AND USED WITH THE SAME RELOAD BUT THE ISSUE STILL OCCURRED. ANOTHER RELOAD WAS THEN USED WITH THE FIRST HANDLE TO COMPLETE THE CASE. THE SECOND HANDLE ALSO WORKED WITH OTHER RELOADS. THERE WAS NO PATIENT INJURY.
07-03-2023 | Injury | us surgical puerto rico | 08-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE POSTERIOR PORTION OF THE ANASTOMOSIS HAD A BLEEDING, WHICH WAS RESOLVED VIA MANUAL SUTURING AND A GAS INSTALLATION TEST WAS PERFORMED THROUGH THE ANUS WHICH CONFIRMED THAT THERE WAS NO LEAK. ON THE 3RD POST-OPERATIVE DAY, THERE WAS WOUND DEHISCENCE IN THE OPENING OF THE ANASTOMOSIS AT THE INTERSECTION OF THE STAPLE LINE, AN EXPLORATORY LAPAROTOMY, AND A COLOSTOMY WAS MADE IN THE DESCENDING COLON LOOP. THEN 4 DAYS AFTER THE INITIAL LAPAROTOMY, THE PATIENT HAD FEVER, TACHYCARDIA, AND SIGNIFICANT ABDOMINAL PAIN. ANOTHER EXPLORATORY LAPAROTOMY, WAS THEN PERFORMED, WHICH SHOWED EVIDENCE OF NECROSIS OF THE EFFERENT SEGMENT OF THE COLON OF THE LOOP COLOSTOMY, RESECTION OF THE NECROTIC SEGMENT WAS THEN PERFORMED AND THE COLORECTAL ANASTOMOSIS WAS UNDONE AND TERMINAL COLOSTOMY WAS PERFORMED AND THE RECTAL STUMP CLOSURE WAS ALSO PERFORMED. THE PATIENT'S CONDITION PROGRESSED TO DAILY FEVER AND LEUKOCYTOSIS. A CONTROL TOTAL ABDOMEN WAS PERFORMED, TOMOGRAPHY WAS PERFORMED WITH MULTIPLE INTRA-ABDOMINAL COLLECTIONS AND A NEW EXPLORATORY LAPAROTOMY WAS INDICATED TO HEAL THE CAVITY. THE PATIENT'S CONDITION PROGRESSED TO CLINICAL STABILITY AND IMPROVEMENT IN THE CLINICAL CONDITION, HOWEVER ON THE 5TH PART OF THE LAST APPROACH, THE PATIENT WAS PRESENTED WITH A NEW FEVER AND RECTAL CLISTER WAS PERFORMED WITH A IODINATED CONTRAST AND A SMALL RECTAL STUMP FISTULA WITH PERIRECTAL COLLECTION WAS EVIDENT, WHICH WAS DRAINED BY RECTAL TUBE. THE PATIENT THEN PROGRESSED TO A CLINICALLY STABLE CONDITION, AND WAS AFEBRILE POST-OPERATIVELY. THE TREATMENT WAS COMPLETED WITH ANTI-BIOTIC THERAPY, IMPROVED LABORATORY TEST AND PRESENTING CONDITIONS OF MEDICAL DISCHARGE FOR CONTINUITY OF OUTPATIENT TREATMENT.
07-03-2023 | Injury | us surgical puerto rico | 08-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abdominal Pain | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE POSTERIOR PORTION OF THE ANASTOMOSIS HAD A BLEEDING, WHICH WAS RESOLVED VIA MANUAL SUTURING AND A GAS INSTALLATION TEST WAS PERFORMED THROUGH THE ANUS WHICH CONFIRMED THAT THERE WAS NO LEAK. ON THE 3RD POST-OPERATIVE DAY, THERE WAS WOUND DEHISCENCE IN THE OPENING OF THE ANASTOMOSIS AT THE INTERSECTION OF THE STAPLE LINE, AN EXPLORATORY LAPAROTOMY, AND A COLOSTOMY WAS MADE IN THE DESCENDING COLON LOOP. THEN 4 DAYS AFTER THE INITIAL LAPAROTOMY, THE PATIENT HAD FEVER, TACHYCARDIA, AND SIGNIFICANT ABDOMINAL PAIN. ANOTHER EXPLORATORY LAPAROTOMY, WAS THEN PERFORMED, WHICH SHOWED EVIDENCE OF NECROSIS OF THE EFFERENT SEGMENT OF THE COLON OF THE LOOP COLOSTOMY, RESECTION OF THE NECROTIC SEGMENT WAS THEN PERFORMED AND THE COLORECTAL ANASTOMOSIS WAS UNDONE AND TERMINAL COLOSTOMY WAS PERFORMED AND THE RECTAL STUMP CLOSURE WAS ALSO PERFORMED. THE PATIENT'S CONDITION PROGRESSED TO DAILY FEVER AND LEUKOCYTOSIS. A CONTROL TOTAL ABDOMEN WAS PERFORMED, TOMOGRAPHY WAS PERFORMED WITH MULTIPLE INTRA-ABDOMINAL COLLECTIONS AND A NEW EXPLORATORY LAPAROTOMY WAS INDICATED TO HEAL THE CAVITY. THE PATIENT'S CONDITION PROGRESSED TO CLINICAL STABILITY AND IMPROVEMENT IN THE CLINICAL CONDITION, HOWEVER ON THE 5TH PART OF THE LAST APPROACH, THE PATIENT WAS PRESENTED WITH A NEW FEVER AND RECTAL CLISTER WAS PERFORMED WITH A IODINATED CONTRAST AND A SMALL RECTAL STUMP FISTULA WITH PERIRECTAL COLLECTION WAS EVIDENT, WHICH WAS DRAINED BY RECTAL TUBE. THE PATIENT THEN PROGRESSED TO A CLINICALLY STABLE CONDITION, AND WAS AFEBRILE POST-OPERATIVELY. THE TREATMENT WAS COMPLETED WITH ANTI-BIOTIC THERAPY, IMPROVED LABORATORY TEST AND PRESENTING CONDITIONS OF MEDICAL DISCHARGE FOR CONTINUITY OF OUTPATIENT TREATMENT.
07-03-2023 | Injury | us surgical puerto rico | 08-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE POSTERIOR PORTION OF THE ANASTOMOSIS HAD A BLEEDING, WHICH WAS RESOLVED VIA MANUAL SUTURING AND A GAS INSTALLATION TEST WAS PERFORMED THROUGH THE ANUS WHICH CONFIRMED THAT THERE WAS NO LEAK. ON THE 3RD POST-OPERATIVE DAY, THERE WAS WOUND DEHISCENCE IN THE OPENING OF THE ANASTOMOSIS AT THE INTERSECTION OF THE STAPLE LINE, AN EXPLORATORY LAPAROTOMY, AND A COLOSTOMY WAS MADE IN THE DESCENDING COLON LOOP. THEN 4 DAYS AFTER THE INITIAL LAPAROTOMY, THE PATIENT HAD FEVER, TACHYCARDIA, AND SIGNIFICANT ABDOMINAL PAIN. ANOTHER EXPLORATORY LAPAROTOMY, WAS THEN PERFORMED, WHICH SHOWED EVIDENCE OF NECROSIS OF THE EFFERENT SEGMENT OF THE COLON OF THE LOOP COLOSTOMY, RESECTION OF THE NECROTIC SEGMENT WAS THEN PERFORMED AND THE COLORECTAL ANASTOMOSIS WAS UNDONE AND TERMINAL COLOSTOMY WAS PERFORMED AND THE RECTAL STUMP CLOSURE WAS ALSO PERFORMED. THE PATIENT'S CONDITION PROGRESSED TO DAILY FEVER AND LEUKOCYTOSIS. A CONTROL TOTAL ABDOMEN WAS PERFORMED, TOMOGRAPHY WAS PERFORMED WITH MULTIPLE INTRA-ABDOMINAL COLLECTIONS AND A NEW EXPLORATORY LAPAROTOMY WAS INDICATED TO HEAL THE CAVITY. THE PATIENT'S CONDITION PROGRESSED TO CLINICAL STABILITY AND IMPROVEMENT IN THE CLINICAL CONDITION, HOWEVER ON THE 5TH PART OF THE LAST APPROACH, THE PATIENT WAS PRESENTED WITH A NEW FEVER AND RECTAL CLISTER WAS PERFORMED WITH A IODINATED CONTRAST AND A SMALL RECTAL STUMP FISTULA WITH PERIRECTAL COLLECTION WAS EVIDENT, WHICH WAS DRAINED BY RECTAL TUBE. THE PATIENT THEN PROGRESSED TO A CLINICALLY STABLE CONDITION, AND WAS AFEBRILE POST-OPERATIVELY. THE TREATMENT WAS COMPLETED WITH ANTI-BIOTIC THERAPY, IMPROVED LABORATORY TEST AND PRESENTING CONDITIONS OF MEDICAL DISCHARGE FOR CONTINUITY OF OUTPATIENT TREATMENT.
07-03-2023 | Injury | us surgical puerto rico | 08-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE POSTERIOR PORTION OF THE ANASTOMOSIS HAD A BLEEDING, WHICH WAS RESOLVED VIA MANUAL SUTURING AND A GAS INSTALLATION TEST WAS PERFORMED THROUGH THE ANUS WHICH CONFIRMED THAT THERE WAS NO LEAK. ON THE 3RD POST-OPERATIVE DAY, THERE WAS WOUND DEHISCENCE IN THE OPENING OF THE ANASTOMOSIS AT THE INTERSECTION OF THE STAPLE LINE, AN EXPLORATORY LAPAROTOMY, AND A COLOSTOMY WAS MADE IN THE DESCENDING COLON LOOP. THEN 4 DAYS AFTER THE INITIAL LAPAROTOMY, THE PATIENT HAD FEVER, TACHYCARDIA, AND SIGNIFICANT ABDOMINAL PAIN. ANOTHER EXPLORATORY LAPAROTOMY, WAS THEN PERFORMED, WHICH SHOWED EVIDENCE OF NECROSIS OF THE EFFERENT SEGMENT OF THE COLON OF THE LOOP COLOSTOMY, RESECTION OF THE NECROTIC SEGMENT WAS THEN PERFORMED AND THE COLORECTAL ANASTOMOSIS WAS UNDONE AND TERMINAL COLOSTOMY WAS PERFORMED AND THE RECTAL STUMP CLOSURE WAS ALSO PERFORMED. THE PATIENT'S CONDITION PROGRESSED TO DAILY FEVER AND LEUKOCYTOSIS. A CONTROL TOTAL ABDOMEN WAS PERFORMED, TOMOGRAPHY WAS PERFORMED WITH MULTIPLE INTRA-ABDOMINAL COLLECTIONS AND A NEW EXPLORATORY LAPAROTOMY WAS INDICATED TO HEAL THE CAVITY. THE PATIENT'S CONDITION PROGRESSED TO CLINICAL STABILITY AND IMPROVEMENT IN THE CLINICAL CONDITION, HOWEVER ON THE 5TH PART OF THE LAST APPROACH, THE PATIENT WAS PRESENTED WITH A NEW FEVER AND RECTAL CLISTER WAS PERFORMED WITH A IODINATED CONTRAST AND A SMALL RECTAL STUMP FISTULA WITH PERIRECTAL COLLECTION WAS EVIDENT, WHICH WAS DRAINED BY RECTAL TUBE. THE PATIENT THEN PROGRESSED TO A CLINICALLY STABLE CONDITION, AND WAS AFEBRILE POST-OPERATIVELY. THE TREATMENT WAS COMPLETED WITH ANTI-BIOTIC THERAPY, IMPROVED LABORATORY TEST AND PRESENTING CONDITIONS OF MEDICAL DISCHARGE FOR CONTINUITY OF OUTPATIENT TREATMENT.
07-03-2023 | Injury | us surgical puerto rico | 08-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE POSTERIOR PORTION OF THE ANASTOMOSIS HAD A BLEEDING, WHICH WAS RESOLVED VIA MANUAL SUTURING AND A GAS INSTALLATION TEST WAS PERFORMED THROUGH THE ANUS WHICH CONFIRMED THAT THERE WAS NO LEAK. ON THE 3RD POST-OPERATIVE DAY, THERE WAS WOUND DEHISCENCE IN THE OPENING OF THE ANASTOMOSIS AT THE INTERSECTION OF THE STAPLE LINE, AN EXPLORATORY LAPAROTOMY, AND A COLOSTOMY WAS MADE IN THE DESCENDING COLON LOOP. THEN 4 DAYS AFTER THE INITIAL LAPAROTOMY, THE PATIENT HAD FEVER, TACHYCARDIA, AND SIGNIFICANT ABDOMINAL PAIN. ANOTHER EXPLORATORY LAPAROTOMY, WAS THEN PERFORMED, WHICH SHOWED EVIDENCE OF NECROSIS OF THE EFFERENT SEGMENT OF THE COLON OF THE LOOP COLOSTOMY, RESECTION OF THE NECROTIC SEGMENT WAS THEN PERFORMED AND THE COLORECTAL ANASTOMOSIS WAS UNDONE AND TERMINAL COLOSTOMY WAS PERFORMED AND THE RECTAL STUMP CLOSURE WAS ALSO PERFORMED. THE PATIENT'S CONDITION PROGRESSED TO DAILY FEVER AND LEUKOCYTOSIS. A CONTROL TOTAL ABDOMEN WAS PERFORMED, TOMOGRAPHY WAS PERFORMED WITH MULTIPLE INTRA-ABDOMINAL COLLECTIONS AND A NEW EXPLORATORY LAPAROTOMY WAS INDICATED TO HEAL THE CAVITY. THE PATIENT'S CONDITION PROGRESSED TO CLINICAL STABILITY AND IMPROVEMENT IN THE CLINICAL CONDITION, HOWEVER ON THE 5TH PART OF THE LAST APPROACH, THE PATIENT WAS PRESENTED WITH A NEW FEVER AND RECTAL CLISTER WAS PERFORMED WITH A IODINATED CONTRAST AND A SMALL RECTAL STUMP FISTULA WITH PERIRECTAL COLLECTION WAS EVIDENT, WHICH WAS DRAINED BY RECTAL TUBE. THE PATIENT THEN PROGRESSED TO A CLINICALLY STABLE CONDITION, AND WAS AFEBRILE POST-OPERATIVELY. THE TREATMENT WAS COMPLETED WITH ANTI-BIOTIC THERAPY, IMPROVED LABORATORY TEST AND PRESENTING CONDITIONS OF MEDICAL DISCHARGE FOR CONTINUITY OF OUTPATIENT TREATMENT.
07-03-2023 | Injury | us surgical puerto rico | 08-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Tachycardia | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE POSTERIOR PORTION OF THE ANASTOMOSIS HAD A BLEEDING, WHICH WAS RESOLVED VIA MANUAL SUTURING AND A GAS INSTALLATION TEST WAS PERFORMED THROUGH THE ANUS WHICH CONFIRMED THAT THERE WAS NO LEAK. ON THE 3RD POST-OPERATIVE DAY, THERE WAS WOUND DEHISCENCE IN THE OPENING OF THE ANASTOMOSIS AT THE INTERSECTION OF THE STAPLE LINE, AN EXPLORATORY LAPAROTOMY, AND A COLOSTOMY WAS MADE IN THE DESCENDING COLON LOOP. THEN 4 DAYS AFTER THE INITIAL LAPAROTOMY, THE PATIENT HAD FEVER, TACHYCARDIA, AND SIGNIFICANT ABDOMINAL PAIN. ANOTHER EXPLORATORY LAPAROTOMY, WAS THEN PERFORMED, WHICH SHOWED EVIDENCE OF NECROSIS OF THE EFFERENT SEGMENT OF THE COLON OF THE LOOP COLOSTOMY, RESECTION OF THE NECROTIC SEGMENT WAS THEN PERFORMED AND THE COLORECTAL ANASTOMOSIS WAS UNDONE AND TERMINAL COLOSTOMY WAS PERFORMED AND THE RECTAL STUMP CLOSURE WAS ALSO PERFORMED. THE PATIENT'S CONDITION PROGRESSED TO DAILY FEVER AND LEUKOCYTOSIS. A CONTROL TOTAL ABDOMEN WAS PERFORMED, TOMOGRAPHY WAS PERFORMED WITH MULTIPLE INTRA-ABDOMINAL COLLECTIONS AND A NEW EXPLORATORY LAPAROTOMY WAS INDICATED TO HEAL THE CAVITY. THE PATIENT'S CONDITION PROGRESSED TO CLINICAL STABILITY AND IMPROVEMENT IN THE CLINICAL CONDITION, HOWEVER ON THE 5TH PART OF THE LAST APPROACH, THE PATIENT WAS PRESENTED WITH A NEW FEVER AND RECTAL CLISTER WAS PERFORMED WITH A IODINATED CONTRAST AND A SMALL RECTAL STUMP FISTULA WITH PERIRECTAL COLLECTION WAS EVIDENT, WHICH WAS DRAINED BY RECTAL TUBE. THE PATIENT THEN PROGRESSED TO A CLINICALLY STABLE CONDITION, AND WAS AFEBRILE POST-OPERATIVELY. THE TREATMENT WAS COMPLETED WITH ANTI-BIOTIC THERAPY, IMPROVED LABORATORY TEST AND PRESENTING CONDITIONS OF MEDICAL DISCHARGE FOR CONTINUITY OF OUTPATIENT TREATMENT.
07-03-2023 | Injury | us surgical puerto rico | 08-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE POSTERIOR PORTION OF THE ANASTOMOSIS HAD A BLEEDING, WHICH WAS RESOLVED VIA MANUAL SUTURING AND A GAS INSTALLATION TEST WAS PERFORMED THROUGH THE ANUS WHICH CONFIRMED THAT THERE WAS NO LEAK. ON THE 3RD POST-OPERATIVE DAY, THERE WAS WOUND DEHISCENCE IN THE OPENING OF THE ANASTOMOSIS AT THE INTERSECTION OF THE STAPLE LINE, AN EXPLORATORY LAPAROTOMY, AND A COLOSTOMY WAS MADE IN THE DESCENDING COLON LOOP. THEN 4 DAYS AFTER THE INITIAL LAPAROTOMY, THE PATIENT HAD FEVER, TACHYCARDIA, AND SIGNIFICANT ABDOMINAL PAIN. ANOTHER EXPLORATORY LAPAROTOMY, WAS THEN PERFORMED, WHICH SHOWED EVIDENCE OF NECROSIS OF THE EFFERENT SEGMENT OF THE COLON OF THE LOOP COLOSTOMY, RESECTION OF THE NECROTIC SEGMENT WAS THEN PERFORMED AND THE COLORECTAL ANASTOMOSIS WAS UNDONE AND TERMINAL COLOSTOMY WAS PERFORMED AND THE RECTAL STUMP CLOSURE WAS ALSO PERFORMED. THE PATIENT'S CONDITION PROGRESSED TO DAILY FEVER AND LEUKOCYTOSIS. A CONTROL TOTAL ABDOMEN WAS PERFORMED, TOMOGRAPHY WAS PERFORMED WITH MULTIPLE INTRA-ABDOMINAL COLLECTIONS AND A NEW EXPLORATORY LAPAROTOMY WAS INDICATED TO HEAL THE CAVITY. THE PATIENT'S CONDITION PROGRESSED TO CLINICAL STABILITY AND IMPROVEMENT IN THE CLINICAL CONDITION, HOWEVER ON THE 5TH PART OF THE LAST APPROACH, THE PATIENT WAS PRESENTED WITH A NEW FEVER AND RECTAL CLISTER WAS PERFORMED WITH A IODINATED CONTRAST AND A SMALL RECTAL STUMP FISTULA WITH PERIRECTAL COLLECTION WAS EVIDENT, WHICH WAS DRAINED BY RECTAL TUBE. THE PATIENT THEN PROGRESSED TO A CLINICALLY STABLE CONDITION, AND WAS AFEBRILE POST-OPERATIVELY. THE TREATMENT WAS COMPLETED WITH ANTI-BIOTIC THERAPY, IMPROVED LABORATORY TEST AND PRESENTING CONDITIONS OF MEDICAL DISCHARGE FOR CONTINUITY OF OUTPATIENT TREATMENT.
07-03-2023 | Injury | us surgical puerto rico | 08-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE POSTERIOR PORTION OF THE ANASTOMOSIS HAD A BLEEDING, WHICH WAS RESOLVED VIA MANUAL SUTURING AND A GAS INSTALLATION TEST WAS PERFORMED THROUGH THE ANUS WHICH CONFIRMED THAT THERE WAS NO LEAK. ON THE 3RD POST-OPERATIVE DAY, THERE WAS WOUND DEHISCENCE IN THE OPENING OF THE ANASTOMOSIS AT THE INTERSECTION OF THE STAPLE LINE, AN EXPLORATORY LAPAROTOMY, AND A COLOSTOMY WAS MADE IN THE DESCENDING COLON LOOP. THEN 4 DAYS AFTER THE INITIAL LAPAROTOMY, THE PATIENT HAD FEVER, TACHYCARDIA, AND SIGNIFICANT ABDOMINAL PAIN. ANOTHER EXPLORATORY LAPAROTOMY, WAS THEN PERFORMED, WHICH SHOWED EVIDENCE OF NECROSIS OF THE EFFERENT SEGMENT OF THE COLON OF THE LOOP COLOSTOMY, RESECTION OF THE NECROTIC SEGMENT WAS THEN PERFORMED AND THE COLORECTAL ANASTOMOSIS WAS UNDONE AND TERMINAL COLOSTOMY WAS PERFORMED AND THE RECTAL STUMP CLOSURE WAS ALSO PERFORMED. THE PATIENT'S CONDITION PROGRESSED TO DAILY FEVER AND LEUKOCYTOSIS. A CONTROL TOTAL ABDOMEN WAS PERFORMED, TOMOGRAPHY WAS PERFORMED WITH MULTIPLE INTRA-ABDOMINAL COLLECTIONS AND A NEW EXPLORATORY LAPAROTOMY WAS INDICATED TO HEAL THE CAVITY. THE PATIENT'S CONDITION PROGRESSED TO CLINICAL STABILITY AND IMPROVEMENT IN THE CLINICAL CONDITION, HOWEVER ON THE 5TH PART OF THE LAST APPROACH, THE PATIENT WAS PRESENTED WITH A NEW FEVER AND RECTAL CLISTER WAS PERFORMED WITH A IODINATED CONTRAST AND A SMALL RECTAL STUMP FISTULA WITH PERIRECTAL COLLECTION WAS EVIDENT, WHICH WAS DRAINED BY RECTAL TUBE. THE PATIENT THEN PROGRESSED TO A CLINICALLY STABLE CONDITION, AND WAS AFEBRILE POST-OPERATIVELY. THE TREATMENT WAS COMPLETED WITH ANTI-BIOTIC THERAPY, IMPROVED LABORATORY TEST AND PRESENTING CONDITIONS OF MEDICAL DISCHARGE FOR CONTINUITY OF OUTPATIENT TREATMENT.
07-03-2023 | Injury | us surgical puerto rico | 08-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | High White Blood Cell Count | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO LAPAROSCOPIC RECTOSIGMOIDECTOMY, THE POSTERIOR PORTION OF THE ANASTOMOSIS HAD A BLEEDING, WHICH WAS RESOLVED VIA MANUAL SUTURING AND A GAS INSTALLATION TEST WAS PERFORMED THROUGH THE ANUS WHICH CONFIRMED THAT THERE WAS NO LEAK. ON THE 3RD POST-OPERATIVE DAY, THERE WAS WOUND DEHISCENCE IN THE OPENING OF THE ANASTOMOSIS AT THE INTERSECTION OF THE STAPLE LINE, AN EXPLORATORY LAPAROTOMY, AND A COLOSTOMY WAS MADE IN THE DESCENDING COLON LOOP. THEN 4 DAYS AFTER THE INITIAL LAPAROTOMY, THE PATIENT HAD FEVER, TACHYCARDIA, AND SIGNIFICANT ABDOMINAL PAIN. ANOTHER EXPLORATORY LAPAROTOMY, WAS THEN PERFORMED, WHICH SHOWED EVIDENCE OF NECROSIS OF THE EFFERENT SEGMENT OF THE COLON OF THE LOOP COLOSTOMY, RESECTION OF THE NECROTIC SEGMENT WAS THEN PERFORMED AND THE COLORECTAL ANASTOMOSIS WAS UNDONE AND TERMINAL COLOSTOMY WAS PERFORMED AND THE RECTAL STUMP CLOSURE WAS ALSO PERFORMED. THE PATIENT'S CONDITION PROGRESSED TO DAILY FEVER AND LEUKOCYTOSIS. A CONTROL TOTAL ABDOMEN WAS PERFORMED, TOMOGRAPHY WAS PERFORMED WITH MULTIPLE INTRA-ABDOMINAL COLLECTIONS AND A NEW EXPLORATORY LAPAROTOMY WAS INDICATED TO HEAL THE CAVITY. THE PATIENT'S CONDITION PROGRESSED TO CLINICAL STABILITY AND IMPROVEMENT IN THE CLINICAL CONDITION, HOWEVER ON THE 5TH PART OF THE LAST APPROACH, THE PATIENT WAS PRESENTED WITH A NEW FEVER AND RECTAL CLISTER WAS PERFORMED WITH A IODINATED CONTRAST AND A SMALL RECTAL STUMP FISTULA WITH PERIRECTAL COLLECTION WAS EVIDENT, WHICH WAS DRAINED BY RECTAL TUBE. THE PATIENT THEN PROGRESSED TO A CLINICALLY STABLE CONDITION, AND WAS AFEBRILE POST-OPERATIVELY. THE TREATMENT WAS COMPLETED WITH ANTI-BIOTIC THERAPY, IMPROVED LABORATORY TEST AND PRESENTING CONDITIONS OF MEDICAL DISCHARGE FOR CONTINUITY OF OUTPATIENT TREATMENT.
10-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | UNKNOWN GIA DST PRODUCT | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | UNKNOWN GIA DST PRODUCT | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | UNKNOWN GIA DST PRODUCT | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | UNKNOWN GIA DST PRODUCT | Firing Problem | A0505 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRECTOMY, WHEN THE GASTRIC TISSUE WAS REMOVED, THE TWO HANDLES DID NOT FIRE AT ALL. TO RESOLVE THE ISSUE, A DEVICE FROM ANOTHER MANUFACTURER WAS USED. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, AT THE TIME OF FIRST FIRING, THE STAPLER DID NOT ALLOW FI RING, THE HANDLE COULD NOT BE SQUEEZED. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE SAME RELOAD. THIS RESULTED IN PROCEDURE TIME EXTENSION OF MORE THAN 30 MINUTES.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, AT THE TIME OF FIRST FIRING, THE STAPLER DID NOT ALLOW FI RING, THE HANDLE COULD NOT BE SQUEEZED. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE SAME RELOAD. THIS RESULTED IN PROCEDURE TIME EXTENSION OF MORE THAN 30 MINUTES.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, AT THE TIME OF FIRST FIRING, THE STAPLER DID NOT ALLOW FI RING, THE HANDLE COULD NOT BE SQUEEZED. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE SAME RELOAD. THIS RESULTED IN PROCEDURE TIME EXTENSION OF MORE THAN 30 MINUTES.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, AT THE TIME OF FIRST FIRING, THE STAPLER DID NOT ALLOW FI RING, THE HANDLE COULD NOT BE SQUEEZED. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE SAME RELOAD. THIS RESULTED IN PROCEDURE TIME EXTENSION OF MORE THAN 30 MINUTES.
17-09-2023 | Injury | us surgical puerto rico | 08-01-2024 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, POSTOPERATIVELY, SYMPTOMS OF ANASTOMOTIC LEAK LED TO TAKING THE PATIENT BACK TO SURGERY.
14-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION PROCEDURE, AFTER THE FIRST SQUEEZE WHILE FIRING THE FIRST RELOAD ON LUNG TISSUE, THE HANDLE'S KNIFE RETRACTED ON ITS OWN TO THE BEGINNING WITHOUT ANY ACTIVITIES IN THE SURGICAL SITE. PRE-FIRED STAPLES WERE OBSERVED IN THE RELOAD. TO RESOLVE THE ISSUE, A NEW HANDLE WITH A DIFFERENT LOT NUMBER WAS USED TO FIRE THE OTHER RELOADS. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRECTOMY, WHEN THE GASTRIC TISSUE WAS REMOVED, THE TWO HANDLES DID NOT FIRE AT ALL. TO RESOLVE THE ISSUE, A DEVICE FROM ANOTHER MANUFACTURER WAS USED. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRECTOMY, WHEN THE GASTRIC TISSUE WAS REMOVED, THE TWO HANDLES DID NOT FIRE AT ALL. TO RESOLVE THE ISSUE, A DEVICE FROM ANOTHER MANUFACTURER WAS USED. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC RADICAL SURGERY OF RECTAL CANCER, DURING THE REMOVAL OF RECTAL TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE ISSUE ALSO OCCURRED IN ANOTHER HANDLE AND RELOAD. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC RADICAL SURGERY OF RECTAL CANCER, DURING THE REMOVAL OF RECTAL TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE ISSUE ALSO OCCURRED IN ANOTHER HANDLE AND RELOAD. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | UNKNOWN DST TA STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN COLON RESECTION, THE PATIENT HAD A HEMATOMA.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC RADICAL SURGERY OF RECTAL CANCER, DURING THE REMOVAL OF RECTAL TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE ISSUE ALSO OCCURRED IN ANOTHER HANDLE AND RELOAD. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC RADICAL SURGERY OF RECTAL CANCER, DURING THE REMOVAL OF RECTAL TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE ISSUE ALSO OCCURRED IN ANOTHER HANDLE AND RELOAD. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) RIGHT LOWER LOBECTOMY, WHEN THE RELOAD WAS CONNECTED TO THE STAPLER, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE, ADAPTER, AND SHELL TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A SUPERIOR LOBECTOMY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE DEVICE WAS STUCK. ANOTHER HANDLE WAS OPENED AND USED WITH THE SAME RELOAD BUT THE ISSUE STILL OCCURRED. ANOTHER RELOAD WAS THEN USED WITH THE FIRST HANDLE TO COMPLETE THE CASE. THE SECOND HANDLE ALSO WORKED WITH OTHER RELOADS. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC UMBILICAL HERNIA REPAIR PROCEDURE WHEN FIXATING THE MESH, NONE OF THE TACKS WERE ABLE TO BE FIRE NORMALLY FROM THE DEVICE. THE TACKS ALSO DISENGAGE INTO THE CAVITY OF THE PATIENT BUT WAS ABLE TO BE RETRIEVED. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM AS THERE WAS A DELAY TO GET ANOTHER DEVICE.
13-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC UMBILICAL HERNIA REPAIR PROCEDURE WHEN FIXATING THE MESH, NONE OF THE TACKS WERE ABLE TO BE FIRE NORMALLY FROM THE DEVICE. THE TACKS ALSO DISENGAGE INTO THE CAVITY OF THE PATIENT BUT WAS ABLE TO BE RETRIEVED. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM AS THERE WAS A DELAY TO GET ANOTHER DEVICE.
13-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY, WHILE DETACHING THE LUNG TISSUE, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZE AND THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
04-10-2022 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Application Program Version or Upgrade Problem | A110206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING/MAINTENANCE, THE HANDLE WOULD PRODUCE "VER" CODE WHEN UPDATING IN VALLEYLAB EXPLORER (VLEX). THE UPDATE WOULD STOP AT 35 PERCENT AND IT WOULD TAKE QUITE A BIT OF TIME TO DISPLAY. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE).
04-10-2022 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING/MAINTENANCE, THE HANDLE WOULD PRODUCE "VER" CODE WHEN UPDATING IN VALLEYLAB EXPLORER (VLEX). THE UPDATE WOULD STOP AT 35 PERCENT AND IT WOULD TAKE QUITE A BIT OF TIME TO DISPLAY. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE).
04-10-2022 | Malfunction | covidien lp llc north haven | 08-01-2024 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING/MAINTENANCE, THE HANDLE WOULD PRODUCE "VER" CODE WHEN UPDATING IN VALLEYLAB EXPLORER (VLEX). THE UPDATE WOULD STOP AT 35 PERCENT AND IT WOULD TAKE QUITE A BIT OF TIME TO DISPLAY. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE).
14-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LOW RECTAL RESECTION, WHILE DETACHING THE RECTUM, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LOW RECTAL RESECTION, WHILE DETACHING THE RECTUM, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBECTOMY, THE STAPLER WAS NOT ABLE FIRE AND A CLICKING SOUND WAS HEARD, AND IT COULD NOT BE USED. A NEW HANDLE WAS USED TO CONTINUE THE SURGICAL TREATMENT. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC LOBECTOMY, THE STAPLER WAS NOT ABLE FIRE AND A CLICKING SOUND WAS HEARD, AND IT COULD NOT BE USED. A NEW HANDLE WAS USED TO CONTINUE THE SURGICAL TREATMENT. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | covidien medical products | 08-01-2024 | PREMIUM PLUS CEEA | Failure to Cut | A050702 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE ANVIL OF THE DEVICE DISENGAGED INTO THE PATIENT'S CAVITY. THE ANVIL WAS REMOVED USING A SNARE. IT WAS ALSO REPORTED THAT THE ANVIL COULD NOT BE TILTED AS THE ANVIL WAS HELD UP BY TISSUE THAT WAS NOT CUT THROUGH. THE DEVICE DID NOT CUT FULLY THROUGH THE TISSUE, BUT THE STAPLE LINES OUTLINING THE DONUTS WERE INTACT. A HOT SNARE WAS USED TO LASSO AND FULGURATE THE AREA.
13-12-2023 | Malfunction | covidien medical products | 08-01-2024 | PREMIUM PLUS CEEA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE ANVIL OF THE DEVICE DISENGAGED INTO THE PATIENT'S CAVITY. THE ANVIL WAS REMOVED USING A SNARE. IT WAS ALSO REPORTED THAT THE ANVIL COULD NOT BE TILTED AS THE ANVIL WAS HELD UP BY TISSUE THAT WAS NOT CUT THROUGH. THE DEVICE DID NOT CUT FULLY THROUGH THE TISSUE, BUT THE STAPLE LINES OUTLINING THE DONUTS WERE INTACT. A HOT SNARE WAS USED TO LASSO AND FULGURATE THE AREA.
13-12-2023 | Malfunction | covidien medical products | 08-01-2024 | PREMIUM PLUS CEEA | Detachment of Device or Device Component | A0501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE ANVIL OF THE DEVICE DISENGAGED INTO THE PATIENT'S CAVITY. THE ANVIL WAS REMOVED USING A SNARE. IT WAS ALSO REPORTED THAT THE ANVIL COULD NOT BE TILTED AS THE ANVIL WAS HELD UP BY TISSUE THAT WAS NOT CUT THROUGH. THE DEVICE DID NOT CUT FULLY THROUGH THE TISSUE, BUT THE STAPLE LINES OUTLINING THE DONUTS WERE INTACT. A HOT SNARE WAS USED TO LASSO AND FULGURATE THE AREA.
13-12-2023 | Malfunction | covidien medical products | 08-01-2024 | PREMIUM PLUS CEEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE ANVIL OF THE DEVICE DISENGAGED INTO THE PATIENT'S CAVITY. THE ANVIL WAS REMOVED USING A SNARE. IT WAS ALSO REPORTED THAT THE ANVIL COULD NOT BE TILTED AS THE ANVIL WAS HELD UP BY TISSUE THAT WAS NOT CUT THROUGH. THE DEVICE DID NOT CUT FULLY THROUGH THE TISSUE, BUT THE STAPLE LINES OUTLINING THE DONUTS WERE INTACT. A HOT SNARE WAS USED TO LASSO AND FULGURATE THE AREA.
14-12-2023 | Malfunction | us surgical puerto rico | 08-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, AFTER VASCULAR TREATMENT, ON THE FOURTH FIRING, DURING FORMATION OF INTERLOBAR SPACE, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT FIRING COULD NOT BE PERFORMED. ANOTHER HANDLE WAS BROUGHT IN, AND ALTHOUGH IT WAS DONE USING THE SAME RELOAD, IT COULD NOT BE FIRED. IN THE END, A NEW REINFORCED 60 MILLIMETER BLACK WAS BROUGHT IN, AND THE PROBLEM WAS SOLVED. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, AFTER VASCULAR TREATMENT, ON THE FOURTH FIRING, DURING FORMATION OF INTERLOBAR SPACE, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT FIRING COULD NOT BE PERFORMED. ANOTHER HANDLE WAS BROUGHT IN, AND ALTHOUGH IT WAS DONE USING THE SAME RELOAD, IT COULD NOT BE FIRED. IN THE END, A NEW REINFORCED 60 MILLIMETER BLACK WAS BROUGHT IN, AND THE PROBLEM WAS SOLVED. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | covidien lp llc north haven | 08-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) LOBECTOMY PROCEDURE, AT THE TIME OF BRONCHIAL RESECTION, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOTIC LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT CUT THE TISSUE COMPLETELY, AND FAILED TO FIRE TO THE END. THERE WAS NO PROBLEM AT THE STAPLE LINE TO THE PART WHERE THE TISSUE WAS CUT. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE.
16-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Cut | A050702 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE ROBOTIC LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT CUT THE TISSUE COMPLETELY, AND FAILED TO FIRE TO THE END. THERE WAS NO PROBLEM AT THE STAPLE LINE TO THE PART WHERE THE TISSUE WAS CUT. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE.
28-11-2023 | Injury | covidien lp llc north haven | 09-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LYMPHADENECTOMY, WITH PLEURODESIS, AND THORACOSTOMY PROCEDURE, AFTER FIRING THE FOURTH RELOAD, THE HANDLE'S TRIGGER BROKE. THEREFORE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED WITH THE SAME HANDLE, BUT THE TRIGGER STILL DID NOT WORK. ANOTHER HANDLE AND RELOAD WERE THEN USED AND FIRING WAS ATTEMPTED MULTIPLE TIMES, BUT THE DEVICE DID NOT FIRE. THE SURGEON THEN EXTENDED THE INCISION AND CONVERTED THE PROCEDURE INTO AN OPEN PROCEDURE AND EXTENDED THE SURGICAL TIME BY 30 MINUTES TO RESOLVE THE ISSUE. THE PATIENT'S HOSPITALIZATION WAS ALSO EXTENDED.
28-11-2023 | Injury | covidien lp llc north haven | 09-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LYMPHADENECTOMY, WITH PLEURODESIS, AND THORACOSTOMY PROCEDURE, AFTER FIRING THE FOURTH RELOAD, THE HANDLE'S TRIGGER BROKE. THEREFORE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED WITH THE SAME HANDLE, BUT THE TRIGGER STILL DID NOT WORK. ANOTHER HANDLE AND RELOAD WERE THEN USED AND FIRING WAS ATTEMPTED MULTIPLE TIMES, BUT THE DEVICE DID NOT FIRE. THE SURGEON THEN EXTENDED THE INCISION AND CONVERTED THE PROCEDURE INTO AN OPEN PROCEDURE AND EXTENDED THE SURGICAL TIME BY 30 MINUTES TO RESOLVE THE ISSUE. THE PATIENT'S HOSPITALIZATION WAS ALSO EXTENDED.
28-11-2023 | Injury | covidien lp llc north haven | 09-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LYMPHADENECTOMY, WITH PLEURODESIS, AND THORACOSTOMY PROCEDURE, AFTER FIRING THE FOURTH RELOAD, THE HANDLE'S TRIGGER BROKE. THEREFORE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED WITH THE SAME HANDLE, BUT THE TRIGGER STILL DID NOT WORK. ANOTHER HANDLE AND RELOAD WERE THEN USED AND FIRING WAS ATTEMPTED MULTIPLE TIMES, BUT THE DEVICE DID NOT FIRE. THE SURGEON THEN EXTENDED THE INCISION AND CONVERTED THE PROCEDURE INTO AN OPEN PROCEDURE AND EXTENDED THE SURGICAL TIME BY 30 MINUTES TO RESOLVE THE ISSUE. THE PATIENT'S HOSPITALIZATION WAS ALSO EXTENDED.
28-11-2023 | Injury | covidien lp llc north haven | 09-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LYMPHADENECTOMY, WITH PLEURODESIS, AND THORACOSTOMY PROCEDURE, AFTER FIRING THE FOURTH RELOAD, THE HANDLE'S TRIGGER BROKE. THEREFORE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED WITH THE SAME HANDLE, BUT THE TRIGGER STILL DID NOT WORK. ANOTHER HANDLE AND RELOAD WERE THEN USED AND FIRING WAS ATTEMPTED MULTIPLE TIMES, BUT THE DEVICE DID NOT FIRE. THE SURGEON THEN EXTENDED THE INCISION AND CONVERTED THE PROCEDURE INTO AN OPEN PROCEDURE AND EXTENDED THE SURGICAL TIME BY 30 MINUTES TO RESOLVE THE ISSUE. THE PATIENT'S HOSPITALIZATION WAS ALSO EXTENDED.
10-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | GIA | Firing Problem | A0505 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | GIA | Firing Problem | A0505 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN GIA DST PRODUCT | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN GIA DST PRODUCT | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN GIA DST PRODUCT | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
10-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN GIA DST PRODUCT | Firing Problem | A0505 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, SURGEON SAID THE TISSUE WAS NOT TOO THICK, THAT THE FIRST FIRE IS A LITTLE HARDER TO FIRE, BUT WITH EACH RELOAD IT WAS GETTING HARDER AND HARDER. THE SURGEON SAID BY THE TIME THE USER GOT THE THIRD LOAD IT COULDN'T EVEN GET IT TO FIRE COMPLETELY TO THE END AND WAS NOT ABLE TO PULL IT BACK TO NEUTRAL POSITION TO REMOVE IT, SO THE SURGEON HAD TO OPEN ANOTHER STAPLER AND RESECT MORE BOWEL AND THE STAPLER THAT WAS STILL ATTACHED TO IT.
12-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AFTER THE DEVICE WAS CONNECTED TO THE HANDLE AND AN ATTEMPT TO FIRE WAS MADE, THE SURGEON COULD PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
28-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AT THE VERY EARLY STAGE, THE STAPLER AND RELOAD WOULD NOT CLOSE FULLY ON THE INTENDED TISSUE. TO RESOLVE THE ISSUE, A SECOND HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
28-07-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ABSORBATACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC REPAIR OF ANTERIOR ABDOMINAL HERNIA PROCEDURE, THE DEVICE BROKE, FIRED A FEW TIMES, AND WOULD NOT CONTINUE TO FUNCTION. NO DEVICE COMPONENT FELL INTO CAVITY OF PATIENT. A NEW DEVICE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-07-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC REPAIR OF ANTERIOR ABDOMINAL HERNIA PROCEDURE, THE DEVICE BROKE, FIRED A FEW TIMES, AND WOULD NOT CONTINUE TO FUNCTION. NO DEVICE COMPONENT FELL INTO CAVITY OF PATIENT. A NEW DEVICE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THAT ON THE FIRST FIRING WITH A BLACK RELOAD THE ANVIL BENT. A SECOND LIKE DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. PROCEDURE WAS COMPLETED WITHOUT FURTHER ISSUE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
09-01-2024 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY, THE HANDLE WAS NO LONGER ABLE TO BE CHARGED. IT WAS NOTED THAT THE CHARGER WORKED WITH ANOTHER HANDLE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
09-01-2024 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, POST-OPERATIVELY, THE HANDLE WAS NO LONGER ABLE TO BE CHARGED. IT WAS NOTED THAT THE CHARGER WORKED WITH ANOTHER HANDLE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
25-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ANTERIOR RESECTION, THE DEVICE COULD NOT BE FIRED AT 1ST FIRING. THE BATTERY WAS REINSERTED, BUT THE ISSUE DID NOT IMPROVE. THE BATTERY WAS REPLACED, BUT THE ISSUE DID NOT IMPROVE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
19-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THAT AFTER THE FIRST FIRING WITH A GREEN RELOAD THE ANVIL WAS BENT AND THE HOME BUTTON WAS NOT FUNCTIONING PROPERLY AND HAD TO HIT IT TWICE TO RETURN TO NEUTRAL. A SECOND DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITH NO FURTHER ISSUES. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
09-11-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE, THE CARTRIDGE COULD NOT BE REMOVED AFTER 1ST FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
13-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THAT AFTER THE SECOND FIRING THAT AFTER REMOVING THE SPENT CARTRIDGE THE UPPER JAW WOULD NOT OPEN ENOUGH TO INSERT A NEW CARTRIDGE. WHILE THE REP WAS GETTING A SECOND DEVICE THE SURGEON CLOSED THE JAWS AND DEPRESSED THE FIRING TRIGGER. THE DEVICE WENT THROUGH THE FULL FIRING SEQUENCE WITH NO CARTRIDGE INSIDE. SURGEON OPENED THE JAWS AND THE DEVICE FULLY OPENED. SURGEON THEN CLOSED THE JAWS AGAIN AND ATTEMPTED TO FIRE (NO CARTRIDGE INSIDE) AND THE DEVICE WOULD NOT FIRE. THE DEVICE DID ITS NORMAL SAFETY PROTOCOL AND FUNCTIONED AS IT SHOULD WITH NO CARTRIDGE INSIDE. A SECOND DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
13-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THAT AFTER THE SECOND FIRING THAT AFTER REMOVING THE SPENT CARTRIDGE THE UPPER JAW WOULD NOT OPEN ENOUGH TO INSERT A NEW CARTRIDGE. WHILE THE REP WAS GETTING A SECOND DEVICE THE SURGEON CLOSED THE JAWS AND DEPRESSED THE FIRING TRIGGER. THE DEVICE WENT THROUGH THE FULL FIRING SEQUENCE WITH NO CARTRIDGE INSIDE. SURGEON OPENED THE JAWS AND THE DEVICE FULLY OPENED. SURGEON THEN CLOSED THE JAWS AGAIN AND ATTEMPTED TO FIRE (NO CARTRIDGE INSIDE) AND THE DEVICE WOULD NOT FIRE. THE DEVICE DID ITS NORMAL SAFETY PROTOCOL AND FUNCTIONED AS IT SHOULD WITH NO CARTRIDGE INSIDE. A SECOND DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
13-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Difficult or Delayed Activation | A150102 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE THAT THE STAPLER SEEMED TO BE STRUGGLING TO FINISH STAPLING SEQUENCE BUT DID FINISHED. ON ANOTHER ATTEMPT THE DEVICE DID NOT FINISH THE STAPLING SEQUENCES AND A "REBOOT" WAS NEEDED. BATTERY WAS REMOVED AND REINSERTED TO "REBOOT" THE DEVICE. THIS WAS NEEDED A FEW TIMES THROUGH OUT THE PROCEDURE. PROCEDURE COMPLETED WITH THE SAME DEVICE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
13-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE THAT THE STAPLER SEEMED TO BE STRUGGLING TO FINISH STAPLING SEQUENCE BUT DID FINISHED. ON ANOTHER ATTEMPT THE DEVICE DID NOT FINISH THE STAPLING SEQUENCES AND A "REBOOT" WAS NEEDED. BATTERY WAS REMOVED AND REINSERTED TO "REBOOT" THE DEVICE. THIS WAS NEEDED A FEW TIMES THROUGH OUT THE PROCEDURE. PROCEDURE COMPLETED WITH THE SAME DEVICE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
25-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTRECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | GST BLUE RELOAD, 45MM, 6 ROW | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THE CARTRIDGE PAN SEPARATION AFTER OPENED THE PACKING. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
26-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTRECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ENDO ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS STUCK WHEN TRYING TO REOPEN AND THE STAPLES DID NOT FORM BUT IT DID NOT CUT PROPERLY DURING A VANKERS PROCEDURE. THEY USED A NEW DEVICE TO COMPLETE CASE WITH NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS STUCK WHEN TRYING TO REOPEN AND THE STAPLES DID NOT FORM BUT IT DID NOT CUT PROPERLY DURING A VANKERS PROCEDURE. THEY USED A NEW DEVICE TO COMPLETE CASE WITH NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS STUCK WHEN TRYING TO REOPEN AND THE STAPLES DID NOT FORM BUT IT DID NOT CUT PROPERLY DURING A VANKERS PROCEDURE. THEY USED A NEW DEVICE TO COMPLETE CASE WITH NO PATIENT HARM.
27-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ARTHROSCOPIC DONOR LEFT NEPHRECTOMY, THE KNIFE MOVED FORWARD BUT STOPPED MOVING AT THE 2ND FIRING ON RENAL VEIN. THE PATIENT RECOVERED WITHOUT SERIOUS INJURY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Injury | ethicon endo-surgery | 09-01-2024 | ILS 21MM CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE FACILITY EXPERIENCED AN UNKNOWN ISSUE WITH A CIRCULAR STAPLER.
01-01-2023 | Injury | ethicon endo-surgery | 09-01-2024 | ILS 21MM CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE FACILITY EXPERIENCED AN UNKNOWN ISSUE WITH A CIRCULAR STAPLER.
01-01-2023 | Injury | ethicon endo-surgery | 09-01-2024 | ILS 21MM CURVED | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE FACILITY EXPERIENCED AN UNKNOWN ISSUE WITH A CIRCULAR STAPLER.
01-01-2023 | Injury | ethicon endo-surgery | 09-01-2024 | ILS 21MM CURVED | Patient Device Interaction Problem | A01 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE FACILITY EXPERIENCED AN UNKNOWN ISSUE WITH A CIRCULAR STAPLER.
01-01-2023 | Injury | ethicon endo-surgery | 09-01-2024 | ILS 21MM CURVED | Ejection Problem | A0503 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE FACILITY EXPERIENCED AN UNKNOWN ISSUE WITH A CIRCULAR STAPLER.
01-01-2023 | Injury | ethicon endo-surgery | 09-01-2024 | ILS 21MM CURVED | Ejection Problem | A0503 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE FACILITY EXPERIENCED AN UNKNOWN ISSUE WITH A CIRCULAR STAPLER.
14-10-2023 | Injury | ethicon endo-surgery | 09-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Adverse Event Without Identified Device or Use Problem | A24 | Chest Pain | Event Description: IT WAS REPORTED THAT AFTER A THORACOSCOPIC LEFT UPPER PULMONARY BULLECTOMY, THE PATIENT WAS ADMITTED DUE TO "CHEST DISTRESS FOR MORE THAN 1 MONTH", AND UNDERWENT THORACOSCOPIC LEFT UPPER BULLAE RESECTION UNDER GENERAL ANESTHESIA ON (B)(6) 2023 . IT WAS REPORTED THAT THE OPERATION WENT SMOOTHLY. NO OBVIOUS AIR LEAKAGE WAS FOUND DURING THE OPERATION AND NO OBVIOUS GAS ESCAPE WAS OBSERVED FROM THE CHEST BOTTLE ON THE 1ST DAY AFTER THE OPERATION. A LARGE AMOUNT OF GAS ESCAPE WAS OBSERVED FROM THE CHEST BOTTLE ON THE 4TH DAY AFTER THE OPERATION. THORACENTESIS WAS PERFORMED AGAIN TO DRAIN THE GAS. THE PATIENT WAS IMPROVED AND DISCHARGED.
14-10-2023 | Injury | ethicon endo-surgery | 09-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Adverse Event Without Identified Device or Use Problem | A24 | Pneumothorax | Event Description: IT WAS REPORTED THAT AFTER A THORACOSCOPIC LEFT UPPER PULMONARY BULLECTOMY, THE PATIENT WAS ADMITTED DUE TO "CHEST DISTRESS FOR MORE THAN 1 MONTH", AND UNDERWENT THORACOSCOPIC LEFT UPPER BULLAE RESECTION UNDER GENERAL ANESTHESIA ON (B)(6) 2023 . IT WAS REPORTED THAT THE OPERATION WENT SMOOTHLY. NO OBVIOUS AIR LEAKAGE WAS FOUND DURING THE OPERATION AND NO OBVIOUS GAS ESCAPE WAS OBSERVED FROM THE CHEST BOTTLE ON THE 1ST DAY AFTER THE OPERATION. A LARGE AMOUNT OF GAS ESCAPE WAS OBSERVED FROM THE CHEST BOTTLE ON THE 4TH DAY AFTER THE OPERATION. THORACENTESIS WAS PERFORMED AGAIN TO DRAIN THE GAS. THE PATIENT WAS IMPROVED AND DISCHARGED.
06-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT TRIGGER, AND IT DID NOT WORK. IT DID NOT COMPLY WITH ITS FUNCTION. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHILE DETACHING THE TISSUE, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN REPAIR OF PARALYZED DIAPHRAGM, THE DEVICE DID NOT FIRE. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN REPAIR OF PARALYZED DIAPHRAGM, THE DEVICE DID NOT FIRE. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-11-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, DURING THE PREPARATION OF THE DEVICE, THE ADAPTER DID NOT SEEM TO BE COMPLETELY CONNECTED TO THE HANDLE. THE HANDLE THEN WOULD NOT READ ANY OF THE RELOADS. THE SHELL WAS TRIED TO BE CHANGE BUT THE ISSUE PERSISTS. ANOTHER ADAPTER WAS USED WITH THE SAME HANDLE, SHELL AND RELOAD WHICH RESOLVED THE ISSUE. THERE WAS NO PATIENT INVOLVED.
29-11-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, DURING THE PREPARATION OF THE DEVICE, THE ADAPTER DID NOT SEEM TO BE COMPLETELY CONNECTED TO THE HANDLE. THE HANDLE THEN WOULD NOT READ ANY OF THE RELOADS. THE SHELL WAS TRIED TO BE CHANGE BUT THE ISSUE PERSISTS. ANOTHER ADAPTER WAS USED WITH THE SAME HANDLE, SHELL AND RELOAD WHICH RESOLVED THE ISSUE. THERE WAS NO PATIENT INVOLVED.
03-11-2023 | Malfunction | covidien lp | 09-01-2024 | COVIDIEN LP |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: SCRUB PERSON WAS UNLOADING A FIRED STAPLER ON FIELD, AND A PIECE OF IT BROKE ONTO THE FIELD AND FELL ON THE FLOOR. SURGEON AWARE.
14-09-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE HANDLE WAS PULLED OUT FROM THE CHARGER DURING THE PREPARATION STAGE, THERE WAS HANDLE ERROR MESSAGE SHOWING EXCLAMATION MARK WHICH AND HANDLE INSIDE THE NO SYMBOL/RED CIRCLE WITH LINE. THE HANDLE WAS FULLY CHARGED WITH THE CHARGER AND THE CHARGER WAS ABLE TO FULLY CHARGE OTHER HANDLE. ANOTHER HANDLE WAS TAKEN OUT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT EVALUATION REVEALED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
14-09-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE HANDLE WAS PULLED OUT FROM THE CHARGER DURING THE PREPARATION STAGE, THERE WAS HANDLE ERROR MESSAGE SHOWING EXCLAMATION MARK WHICH AND HANDLE INSIDE THE NO SYMBOL/RED CIRCLE WITH LINE. THE HANDLE WAS FULLY CHARGED WITH THE CHARGER AND THE CHARGER WAS ABLE TO FULLY CHARGE OTHER HANDLE. ANOTHER HANDLE WAS TAKEN OUT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. PLI10: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT EVALUATION REVEALED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
12-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. AS THE TISSUE WAS CUT OFF, REMOVED THE DEVICE FROM THE PATIENT THROUGH TISSUE GAP. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
15-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN UROLOGY, THERE WAS A PROBLEM LOADING THE DEVICE; THAT STAPLER WAS OPENED AND LOADED BY AN AGENCY NURSE, BUT IT WAS DONE INCORRECTLY. IT WAS NOTED THAT THE STAPLER WAS ABLE TO BE CYCLED, IT WOULD OPEN AND CLOSED. THE SURGEON TRIED TO FIRE THE DEVICE, DEVICE WAS CLAMPED ONTO TISSUE AT ONE POINT BUT WHEN IT DID NOT FIRE, IT WAS THEN REMOVED FROM THE TISSUE AND IT WAS NOT LOCKED ONTO THE TISSUE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT USED IN THE PATIENT. ANOTHER STAPLER FROM ANOTHER COMPANY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN UROLOGY, THERE WAS A PROBLEM LOADING THE DEVICE; THAT STAPLER WAS OPENED AND LOADED BY AN AGENCY NURSE, BUT IT WAS DONE INCORRECTLY. IT WAS NOTED THAT THE STAPLER WAS ABLE TO BE CYCLED, IT WOULD OPEN AND CLOSED. THE SURGEON TRIED TO FIRE THE DEVICE, DEVICE WAS CLAMPED ONTO TISSUE AT ONE POINT BUT WHEN IT DID NOT FIRE, IT WAS THEN REMOVED FROM THE TISSUE AND IT WAS NOT LOCKED ONTO THE TISSUE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT USED IN THE PATIENT. ANOTHER STAPLER FROM ANOTHER COMPANY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREATIC HEAD AND DUODENAL RESECTION (PANCREATICODUODENECTOMY) PROCEDURE, AT THE TIME OF THE FIRST FIRING OF THE GASTRECTOMY (GASTRIC RESECTION), THE FIRING WAS STARTED IN SLOW MODE. HOWEVER, IT STOPPED HALFWAY THROUGH, SO THE KNIFE WAS THEN PULLED BACK. TO RESOLVE THE ISSUE, THE SURGEON USED A NEW RELOAD WITH THE SAME POWERED DEVICES AND RESECTED FROM WHERE THE FIRING STOPPED. ACCORDING TO THE SURGEON, THERE WAS SEVERE INFLAMMATION, AND THE GASTRIC WALL HAD CONSIDERABLY THICKENED; AND WHEN IT WAS CLAMP THE SURGEON WAS UNABLE TO CHECK THE HANDLE'S ZONE DISPLAY AND IT WAS ALSO UNCLEAR WHETHER AN ALERT HAD SOUNDED. THE SEVERE INFLAMMATION AND THICKENED GASTRIC WALL WAS NOT DUE TO THE DEVICE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE DEVICE WOULD NOT FIRE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
01-01-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE KNIFE DID NOT RETURN TO HOME POSITION. KNIFE REVERSE SWITCH WAS USED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Failure to Cut | A050702 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC CONVERSION OF A SLEEVE TO A LAPAROSCOPIC ROUX-N-Y, THE STAPLE WAS FIRED, THE STAPLES DEPLOYED, BUT IT LOOKED LIKE THE KNIFE EITHER GOT CAUGHT ON SOMETHING OR SOMETHING PREVENTED THE BLADE FROM DISSECTING IN BETWEEN THE STAPLE LINES AS IT NORMALLY DOES. THERE WAS NOT AN ANASTOMOSIS FORMED WHEN IT SHOULD HAVE BEEN. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN UROLOGY, THERE WAS A PROBLEM LOADING THE DEVICE; THAT STAPLER WAS OPENED AND LOADED BY AN AGENCY NURSE, BUT IT WAS DONE INCORRECTLY. IT WAS NOTED THAT THE STAPLER WAS ABLE TO BE CYCLED, IT WOULD OPEN AND CLOSED. THE SURGEON TRIED TO FIRE THE DEVICE, DEVICE WAS CLAMPED ONTO TISSUE AT ONE POINT BUT WHEN IT DID NOT FIRE, IT WAS THEN REMOVED FROM THE TISSUE AND IT WAS NOT LOCKED ONTO THE TISSUE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT USED IN THE PATIENT. ANOTHER STAPLER FROM ANOTHER COMPANY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | TRI-STAPLE 2.0 | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN UROLOGY, THERE WAS A PROBLEM LOADING THE DEVICE; THAT STAPLER WAS OPENED AND LOADED BY AN AGENCY NURSE, BUT IT WAS DONE INCORRECTLY. IT WAS NOTED THAT THE STAPLER WAS ABLE TO BE CYCLED, IT WOULD OPEN AND CLOSED. THE SURGEON TRIED TO FIRE THE DEVICE, DEVICE WAS CLAMPED ONTO TISSUE AT ONE POINT BUT WHEN IT DID NOT FIRE, IT WAS THEN REMOVED FROM THE TISSUE AND IT WAS NOT LOCKED ONTO THE TISSUE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS NOT USED IN THE PATIENT. ANOTHER STAPLER FROM ANOTHER COMPANY WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Failure to Cut | A050702 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC CONVERSION OF A SLEEVE TO A RYN, THE STAPLE WAS FIRED, THE STAPLES DEPLOYED, BUT IT LOOKED LIKE THE KNIFE EITHER GOT CAUGHT ON SOMETHING OR SOMETHING PREVENTED THE BLADE FROM DISSECTING IN BETWEEN THE STAPLE LINES AS IT NORMALLY DOES. THERE WAS NOT AN ANASTOMOSIS FORMED WHEN IT SHOULD HAVE BEEN. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
12-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | EEA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT COLECTOMY PROCEDURE, AFTER THE STAPLING PHASE, THE SURGEON TRIED TO REMOVE THE STAPLER. THE STAPLER WAS HARD TO REMOVE AFTER 2 FULL TURNS. THE SURGEON WAS ABLE TO REMOVE THE STAPLER NORMALLY, BUT THE SURGEON FELT LIKE SOMETHING WAS RETAINING THE ANVIL. IT AS NOTED THAT THE ANVIL WAS BENT. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | EEA | Material Twisted/Bent | A040609 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT COLECTOMY PROCEDURE, AFTER THE STAPLING PHASE, THE SURGEON TRIED TO REMOVE THE STAPLER. THE STAPLER WAS HARD TO REMOVE AFTER 2 FULL TURNS. THE SURGEON WAS ABLE TO REMOVE THE STAPLER NORMALLY, BUT THE SURGEON FELT LIKE SOMETHING WAS RETAINING THE ANVIL. IT AS NOTED THAT THE ANVIL WAS BENT. THERE WAS NO PATIENT INJURY.
02-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, NOTED THE CARTRIDGE DAMAGED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | GREEN RELOAD FOR ECHELON 60 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, NOTED THE CARTRIDGE DAMAGED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, NOTED THE CARTRIDGE DAMAGED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | GREEN RELOAD FOR ECHELON 60 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, NOTED THE CARTRIDGE DAMAGED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Excessive Heating | A1002 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN SIGMOIDECTOMY PROCEDURE, THE DEVICE WORKED WELL WHEN IT WAS USED DURING THE OPERATION. THE DEVICE WAS USED ON COLON. HOWEVER, WHEN A NURSE TRIED TO REMOVE THE BATTERY TO DISPOSE OF THE DEVICE AFTER SURGERY, THE NURSE NOTICED THAT THE BATTERY WAS HOT. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
08-11-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Difficult or Delayed Activation | A150102 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THAT ON THE FIRST FIRING WITHOUT SEAM GUARD THE DEVICE WAS PUT DOWN THE TROCAR, ARTICULATED, THEN WHEN ATTEMPTING TO FIRE THERE WAS NO POWER. DEVICE WAS UNCLAMPED FROM TISSUE, PULLED OUT OF THE TROCAR, THE BATTERY WAS FLIPPED THEN PUT BACK AND ON THE SECOND ATTEMPT THE DEVICE FIRED. STAPLE LINE WAS INTACT WITH NO ISSUES. SAME DEVICE WAS USED FOR THE REMAINDER OF THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
08-11-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THAT ON THE FIRST FIRING WITHOUT SEAM GUARD THE DEVICE WAS PUT DOWN THE TROCAR, ARTICULATED, THEN WHEN ATTEMPTING TO FIRE THERE WAS NO POWER. DEVICE WAS UNCLAMPED FROM TISSUE, PULLED OUT OF THE TROCAR, THE BATTERY WAS FLIPPED THEN PUT BACK AND ON THE SECOND ATTEMPT THE DEVICE FIRED. STAPLE LINE WAS INTACT WITH NO ISSUES. SAME DEVICE WAS USED FOR THE REMAINDER OF THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
14-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPY PULMONARY RESECTION PROCEDURE, AFTER CONNECTING HANDLE AND RELOAD AT STAND BY STAGE PRIOR TO USE IN PATIENT, WHEN THE YELLOW TAB WAS REMOVED, A PIECE OF PLASTIC FELL OUTSIDE THE SURGICAL FIELD. NOTHING WAS DONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-10-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE DEVICE COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. USED RBRB AND MANUAL OVERRIDE TO OPEN THE JAW AND REMOVED THE DEVICE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-10-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE DEVICE COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. USED RBRB AND MANUAL OVERRIDE TO OPEN THE JAW AND REMOVED THE DEVICE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SUBTOTAL GASTRECTOMY PROCEDURE, ON THE STOMACH, AFTER THE RELOAD WAS LOADED TO THE HANDLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE DEVICE DID NOT FIRE AND MADE A STRANGE NOISE, AS IF SOMETHING HAD BROKEN. TO RESOLVE THE ISSUE, A POWERED HANDLE WAS USED WITH A NEW RELOAD. THERE WAS NO PATIENT INJURY.
28-01-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM STANDARD | Difficult or Delayed Activation | A150102 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK PROCEDURE, THE STAPLER WAS WORKING THEN IT LOST POWER. ANOTHER DEVICE WAS USED TO COMPLETE PROCEDURE. NO PATIENT CONSEQUENCES WERE REPORTED.
07-11-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE THAT ON THE FIRST FIRING AT THE ANTRUM, DEVICE WAS CLAMPED DOWN ON TISSUE BUT THERE WAS NO POWER. BATTERY WAS REMOVED AND REINSERTED STILL NO POWER. BATTERY WAS REMOVED AND FLIPPED OVER BUT STILL NO POWER. A SECOND DEVICE WAS OPENED AND THE NEW BATTER WAS INSERTED TO THE ORIGINAL DEVICE STILL CLAMPED ON TISSUE, AGAIN THERE WAS NO POWER TO THE DEVICE. THE DEVICE WAS REMOVED FROM TISSUE AND THE NEW DEVICE AND BATTERY WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | EEA ORVIL | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, WHEN INSERTING THE ORAL ANVIL AND TRYING TO DELIVER THE TUBE, THE SURGEON KEPT ENCOUNTERING RESISTANCE OR A BLOCKAGE, AND THE ANVIL WOULD POP OFF THE TUBING. ANOTHER ORAL ANVIL WAS USED FROM DEMO STOCK AND THE SAME THING HAPPENED. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, WHEN INSERTING THE ORAL ANVIL AND TRYING TO DELIVER THE TUBE, THE SURGEON KEPT ENCOUNTERING RESISTANCE OR A BLOCKAGE, AND THE ANVIL WOULD POP OFF THE TUBING. ANOTHER ORAL ANVIL WAS USED FROM DEMO STOCK AND THE SAME THING HAPPENED. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 60MM STANDARD | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, UPON THE FIRST FIRING THERE WAS NO ISSUE BUT AS THE SCRUB TECH WAS LOADING THE DEVICE IT COULD NOT BE LOADED. NO FURTHER INFORMATION WAS PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | EEA ORVIL | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, WHEN INSERTING THE ORAL ANVIL AND TRYING TO DELIVER THE TUBE, THE SURGEON KEPT ENCOUNTERING RESISTANCE OR A BLOCKAGE, AND THE ANVIL WOULD POP OFF THE TUBING. ANOTHER ORAL ANVIL WAS USED FROM DEMO STOCK AND THE SAME THING HAPPENED. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | EEA ORVIL | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC ESOPHAGECTOMY PROCEDURE, WHEN INSERTING THE ORAL ANVIL AND TRYING TO DELIVER THE TUBE, THE SURGEON KEPT ENCOUNTERING RESISTANCE OR A BLOCKAGE, AND THE ANVIL WOULD POP OFF THE TUBING. ANOTHER ORAL ANVIL WAS USED FROM DEMO STOCK AND THE SAME THING HAPPENED. A COMPETITOR'S DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTAL PULMONARY RESECTION PROCEDURE, FIRING COULD NOT BE DONE WHEN RESECTING THE PULMONARY PARENCHYMA. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AND DURING 2ND PASS, THERE WAS CRACKING NOISE THEN STAPLING IMPOSSIBLE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AND DURING 2ND PASS, THERE WAS CRACKING NOISE THEN STAPLING IMPOSSIBLE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AND DURING 2ND PASS, THERE WAS CRACKING NOISE THEN STAPLING IMPOSSIBLE. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE SECOND FIRING OF THE 60 MM. RELOAD, THE INTERNAL COG WAS BROKEN. THE SURGEON HEARD THE COG SNAP BUT WAS ABLE TO CONTINUE TO FIRE. COG SNAPPED AGAIN AND WAS STILL ABLE TO CONTINUE TO FIRE. THE PHYSICIAN WAS ABLE TO COMPLETE THE PROCEDURE WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE SECOND FIRING OF THE 60 MM. RELOAD, THE INTERNAL COG WAS BROKEN. THE SURGEON HEARD THE COG SNAP BUT WAS ABLE TO CONTINUE TO FIRE. COG SNAPPED AGAIN AND WAS STILL ABLE TO CONTINUE TO FIRE. THE PHYSICIAN WAS ABLE TO COMPLETE THE PROCEDURE WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTAL PULMONARY RESECTION PROCEDURE, FIRING COULD NOT BE DONE WHEN RESECTING THE PULMONARY PARENCHYMA. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A LAPAROSCOPIC PROCEDURE, WHILE PREPARING THE DEVICE, FRAGMENTS FELL OFF THE TIP OF THE HANDLE WHEN THE NURSE TOOK THE DEVICE OFF THE PACKAGING. A NEW DEVICE OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
15-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHILE DETACHING LUNG LOBE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHILE DETACHING LUNG LOBE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW RADICAL RESECTION OF RECTAL CANCER, IN DETACHING THE TISSUE, THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREATIC HEAD AND DUODENAL RESECTION (PANCREATICODUODENECTOMY) PROCEDURE, AT THE TIME OF THE FIRST FIRING OF THE GASTRECTOMY (GASTRIC RESECTION), THE FIRING WAS STARTED IN SLOW MODE. HOWEVER, IT STOPPED HALFWAY THROUGH, SO THE KNIFE WAS THEN PULLED BACK. TO RESOLVE THE ISSUE, THE SURGEON USED A NEW RELOAD WITH THE SAME POWERED DEVICES AND RESECTED FROM WHERE THE FIRING STOPPED. ACCORDING TO THE SURGEON, THERE WAS SEVERE INFLAMMATION, AND THE GASTRIC WALL HAD CONSIDERABLY THICKENED; AND WHEN IT WAS CLAMP THE SURGEON WAS UNABLE TO CHECK THE HANDLE'S ZONE DISPLAY AND IT WAS ALSO UNCLEAR WHETHER AN ALERT HAD SOUNDED. THE SEVERE INFLAMMATION AND THICKENED GASTRIC WALL WAS NOT DUE TO THE DEVICE ISSUE. THERE WAS NO PATIENT INJURY.
09-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS DEFECTIVE; THE TIP WAS OPEN. THE ISSUE OCCURRED IN TWO RELOADS. THE SAME HANDLE WAS USED WITH A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
09-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS DEFECTIVE; THE TIP WAS OPEN. THE ISSUE OCCURRED IN TWO RELOADS. THE SAME HANDLE WAS USED WITH A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
11-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN AN OPEN LAPAROTOMY GASTROSTOMY, DURING DUODENECTOMY, FROM THE MIDDLE OF THE FIRING, THE HANDLE COULD NOT BE SQUEEZED AND UNABLE TO FIRE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
09-11-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS DEFECTIVE; THE TIP WAS OPEN. THE ISSUE OCCURRED IN TWO RELOADS. THERE WAS NO PATIENT INVOLVED.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREATIC HEAD AND DUODENAL RESECTION (PANCREATICODUODENECTOMY) PROCEDURE, AT THE TIME OF THE FIRST FIRING OF THE GASTRECTOMY (GASTRIC RESECTION), THE FIRING WAS STARTED IN SLOW MODE. HOWEVER, IT STOPPED HALFWAY THROUGH, SO THE KNIFE WAS THEN PULLED BACK. TO RESOLVE THE ISSUE, THE SURGEON USED A NEW RELOAD WITH THE SAME POWERED DEVICES AND RESECTED FROM WHERE THE FIRING STOPPED. ACCORDING TO THE SURGEON, THERE WAS SEVERE INFLAMMATION, AND THE GASTRIC WALL HAD CONSIDERABLY THICKENED; AND WHEN IT WAS CLAMP THE SURGEON WAS UNABLE TO CHECK THE HANDLE'S ZONE DISPLAY AND IT WAS ALSO UNCLEAR WHETHER AN ALERT HAD SOUNDED. THE SEVERE INFLAMMATION AND THICKENED GASTRIC WALL WAS NOT DUE TO THE DEVICE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 09-01-2024 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY WHEN THE DEVICE WAS FIRED WITH A GOLD RELOAD, THE STAPLER DIDN¿T ACTUALLY FIRE. INSTEAD THE MOTOR OF THE DEVICE MADE A SOUND AS IF IT WERE FIRING BUT THE DEVICE DID NOT ACTUALLY FIRE. THE DEVICE WAS OPENED ON TISSUE AND REMOVED FROM THE PATIENT. UPON EXAMINATION OF THE RELOAD, THE SLED WAS OBSERVED TO STILL BE IN THE HOME POSITION. ANOTHER DEVICE WAS OPENED AND THE CASE WAS COMPLETED WITH NO ISSUES. NO PATIENT HARM WAS REPORTED.
26-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY TO RELEASE INTESTINAL OBSTRUCTION, AT THE TIME OF THE SECOND FIRING DURING INTESTINAL RESECTION, THE HANDLE COULD NOT BE SQUEEZED AND THE FIRING COULD NOT BE DONE. THERE WAS EVIDENCE OF PRE-FIRING PRESENT. ANOTHER RELOAD WAS LOADED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY TO RELEASE INTESTINAL OBSTRUCTION, AT THE TIME OF THE SECOND FIRING DURING INTESTINAL RESECTION, THE HANDLE COULD NOT BE SQUEEZED AND THE FIRING COULD NOT BE DONE. THERE WAS EVIDENCE OF PRE-FIRING PRESENT. ANOTHER RELOAD WAS LOADED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Break | A0401 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, AFTER PERFORMING THE SECOND FIRING OF THE PULMONARY PARENCHYMA RESECTION, A PIECE OF YELLOW TAB FELL INTO PATIENT CAVITY, BUT THE SURGEON WAS ABLE TO BE COLLECT THE FRAGMENT BY USING A FORCEPS. THERE WERE STAPLES THAT HAD POOR FORMATION AS WELL. ADDITIONAL RESECTION WAS PERFORMED, AND THE PROCEDURE WAS COMPLETED WITHOUT ANY PROBLEM.
13-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, AFTER PERFORMING THE SECOND FIRING OF THE PULMONARY PARENCHYMA RESECTION, A PIECE OF YELLOW TAB FELL INTO PATIENT CAVITY, BUT THE SURGEON WAS ABLE TO BE COLLECT THE FRAGMENT BY USING A FORCEPS. THERE WERE STAPLES THAT HAD POOR FORMATION AS WELL. ADDITIONAL RESECTION WAS PERFORMED, AND THE PROCEDURE WAS COMPLETED WITHOUT ANY PROBLEM.
13-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Detachment of Device or Device Component | A0501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PARTIAL PULMONARY RESECTION, AFTER PERFORMING THE SECOND FIRING OF THE PULMONARY PARENCHYMA RESECTION, A PIECE OF YELLOW TAB FELL INTO PATIENT CAVITY, BUT THE SURGEON WAS ABLE TO BE COLLECT THE FRAGMENT BY USING A FORCEPS. THERE WERE STAPLES THAT HAD POOR FORMATION AS WELL. ADDITIONAL RESECTION WAS PERFORMED, AND THE PROCEDURE WAS COMPLETED WITHOUT ANY PROBLEM.
19-03-2024 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, THE FIRING WAS PERFORMED BUT THE USER FELT THAT IT STOPPED MIDWAY. THE JAW WAS OPENED AND WAS RELEASED FROM THE TISSUE. THERE SEEMED TO BE NO PROBLEMS WITH THE STAPLE LINE, BUT THE HANDLE AND THE ADAPTER WERE REPLACED. THE FOLLOW FIRING WAS PERFORMED FROM WHERE THE LINE STOPPED. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, AT THE TIME OF COLON RESECTION, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, AT THE TIME OF COLON RESECTION, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), AT THE TIME OF LUNG RESECTION DURING 1ST AND 2ND FIRING, FIRING COULD NOT BE PERFORMED USING TWO 60MM PURPLE RELOADS. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT THE INSTALLATION WAS ALSO PERFORMED AS USUAL, AS WELL AS THE STAPLES WERE FOUND TO BE EXPOSED FROM THE ROOT PART. THE HANDLE AND RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT INJURY.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), AT THE TIME OF LUNG RESECTION DURING 1ST AND 2ND FIRING, FIRING COULD NOT BE PERFORMED USING TWO 60MM PURPLE RELOADS. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT THE INSTALLATION WAS ALSO PERFORMED AS USUAL, AS WELL AS THE STAPLES WERE FOUND TO BE EXPOSED FROM THE ROOT PART. THE HANDLE AND RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT INJURY.
01-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | UNKNOWN SIGNIA EGIA ADAPTER | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC PROCEDURE CONVERTED TO OPEN PANCREATECTOMY AND SPLENECTOMY, THE STAPLES THAT HAD PROPER FORMATION MISFIRED AND DID NOT FIRE ALL THE WAY. THE OTHER STAPLES HAD POOR FORMATION. A DIFFERENT STAPLER FROM ANOTHER MANUFACTURER WAS USED TO REPAIR THE PREVIOUSLY MISFIRED STAPLE LINE.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), AT THE TIME OF LUNG RESECTION DURING 1ST AND 2ND FIRING, FIRING COULD NOT BE PERFORMED USING TWO 60MM PURPLE RELOADS. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT THE INSTALLATION WAS ALSO PERFORMED AS USUAL. THE HANDLE AND RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT INJURY.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), AT THE TIME OF LUNG RESECTION DURING 1ST AND 2ND FIRING, FIRING COULD NOT BE PERFORMED USING TWO 60MM PURPLE RELOADS. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT THE INSTALLATION WAS ALSO PERFORMED AS USUAL. THE HANDLE AND RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT INJURY.
15-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS), AT THE TIME OF LUNG RESECTION DURING 1ST AND 2ND FIRING, FIRING COULD NOT BE PERFORMED USING TWO 60MM PURPLE RELOADS. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. IT WAS NOTED THAT THE INSTALLATION WAS ALSO PERFORMED AS USUAL, AS WELL AS THE STAPLES WERE FOUND TO BE EXPOSED FROM THE ROOT PART. THE HANDLE AND RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ULTRA-LOW ANTERIOR PROCTECTOMY AND PARTIAL INTERNAL SPHINCTEROTOMY, AFTER PRESSING THE GREEN BUTTON, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND COULD NOT BE FIRED. AFTER SQUEEZING HARD, TEH HANDLE MADE A CLICKING SOUND AND COULD NOT BE USED NORMALLY. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ULTRA-LOW ANTERIOR PROCTECTOMY AND PARTIAL INTERNAL SPHINCTEROTOMY, AFTER PRESSING THE GREEN BUTTON, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND COULD NOT BE FIRED. AFTER SQUEEZING HARD, TEH HANDLE MADE A CLICKING SOUND AND COULD NOT BE USED NORMALLY. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
16-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN EGIA ULTRA STAPLER | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER LOADING THE FIRST RELOAD, THE DEVICE WAS LOCKED ON THE TISSUE. THE SURGEON PULLED THE TISSUE, BUT IT WAS STILL LOCKED AND COULD NOT BE FIRED. THE SECOND RELOAD FIRED BUT LOCKED HALFWAY. TO RESOLVE THE ISSUE, MANUAL SUTURE WAS USED.
16-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN EGIA ULTRA STAPLER | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER LOADING THE FIRST RELOAD, THE DEVICE WAS LOCKED ON THE TISSUE. THE SURGEON PULLED THE TISSUE, BUT IT WAS STILL LOCKED AND COULD NOT BE FIRED. THE SECOND RELOAD FIRED BUT LOCKED HALFWAY. TO RESOLVE THE ISSUE, MANUAL SUTURE WAS USED.
16-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN EGIA ULTRA STAPLER | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER LOADING THE FIRST RELOAD, THE DEVICE WAS LOCKED ON THE TISSUE. THE SURGEON PULLED THE TISSUE, BUT IT WAS STILL LOCKED AND COULD NOT BE FIRED. THE SECOND RELOAD FIRED BUT LOCKED HALFWAY. TO RESOLVE THE ISSUE, MANUAL SUTURE WAS USED.
16-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER LOADING THE FIRST RELOAD, THE DEVICE WAS LOCKED ON THE TISSUE. THE SURGEON PULLED THE TISSUE, BUT IT WAS STILL LOCKED AND COULD NOT BE FIRED. THE SECOND RELOAD FIRED BUT LOCKED HALFWAY. TO RESOLVE THE ISSUE, MANUAL SUTURE WAS USED.
16-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS PROCEDURE, AFTER LOADING THE FIRST RELOAD, THE DEVICE WAS LOCKED ON THE TISSUE. THE SURGEON PULLED THE TISSUE, BUT IT WAS STILL LOCKED AND COULD NOT BE FIRED. THE DEVICE WAS FORCIBLY REMOVED. THE SECOND RELOAD FIRED BUT LOCKED HALFWAY. TO RESOLVE THE ISSUE, MANUAL SUTURE WAS USED.
16-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC BYPASS PROCEDURE, AFTER LOADING THE FIRST RELOAD, THE DEVICE WAS LOCKED ON THE TISSUE. THE SURGEON PULLED THE TISSUE, BUT IT WAS STILL LOCKED AND COULD NOT BE FIRED. THE DEVICE WAS FORCIBLY REMOVED. THE SECOND RELOAD FIRED BUT LOCKED HALFWAY. TO RESOLVE THE ISSUE, MANUAL SUTURE WAS USED.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PARTIAL RESECTION OF THE LUNG, DURING LUNG PARENCHYMA RESECTION, THE HANDLE BECAME HARDER TO SQUEEZE AND THE DEVICE SLIGHTLY FIRED THEN THEN THE FIRING STOPPED. A NEW 45MM RELOAD WITH THE SAME HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
15-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE TIP OF THE JAWS WERE NOT CLOSED PROPERLY. THERE WAS NO PATIENT INVOLVED.
09-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, SOME OF THE STAPLES DID NOT FORM THE B-SHAPE AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE SURGEON HAND SEWN THE TISSUE AND A NEW RELOAD WAS USED TO COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED. FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
09-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, SOME OF THE STAPLES DID NOT FORM THE B-SHAPE AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THE SURGEON HAND SEWN THE TISSUE AND A NEW RELOAD WAS USED TO COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED. FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
15-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE TIP OF THE JAWS WERE NOT CLOSED PROPERLY. THERE WAS NO PATIENT INVOLVED.
11-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC PROCEDURE, WHILE AT THE ESOPHAGUS, THERE WAS AN UNKNOWN ISSUE. THE PRODUCT WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
11-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC PROCEDURE, WHILE AT THE ESOPHAGUS, THERE WAS AN UNKNOWN ISSUE. THE PRODUCT WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
09-01-2024 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS AN UNKNOWN DEVICE ISSUE. THE RELOAD WAS REPLACED AND USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
09-01-2024 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THERE WAS AN UNKNOWN DEVICE ISSUE. THE RELOAD WAS REPLACED AND USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A VISUAL INSPECTION OF INTERNAL COMPONENTS REVEALED THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
14-11-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRECTOMY FOR GASTRIC CANCER, WHEN FIRING WAS BEING PERFORMED ON THE SMALL INTESTINE, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED, AND THE PHENOMENON OF FAILURE TO FIRE OCCURRED. A 45-MILLIMETER RELOAD WAS USED, BUT THE DEVICE DID NOT FIRE AGAIN. TO RESOLVE THE ISSUE, A NEW 60-MILLIMETER RELOAD OF THE SAME MODEL AND HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
16-12-2023 | Malfunction | us surgical puerto rico | 09-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A NEPHRECTOMY PROCEDURE, THE USER WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. THE ISSUE OCCURRED IN ONE HANDLE AND THREE RELOADS. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A NEPHRECTOMY PROCEDURE, THE USER WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. THE ISSUE OCCURRED IN ONE HANDLE AND THREE RELOADS. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE USER WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. THE ISSUE OCCURRED IN ONE HANDLE AND THREE RELOADS. THERE WAS NO PATIENT INVOLVEMENT.
18-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, ON THE THIRD FIRING OF COLONIC ANASTOMOSIS (FUNCTIONAL END-TO-END (FEE) RECONSTRUCTION 4 FIRING METHOD), AFTER LESION REMOVAL, THE DEVICE FIRED ALL THE WAY, BUT THE STAPLE FORMATION OF THE TIP PART WAS POORLY FORMED. SINCE THIS COULD BE CONFIRMED VISUALLY, THE POORLY FORMED PORTION WAS REINFORCED WITH ADDITIONAL SUTURES.
18-12-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, ON THE THIRD FIRING OF COLONIC ANASTOMOSIS (FUNCTIONAL END-TO-END (FEE) RECONSTRUCTION 4 FIRING METHOD), AFTER LESION REMOVAL, THE DEVICE FIRED ALL THE WAY, BUT THE STAPLE FORMATION OF THE TIP PART WAS POORLY FORMED. SINCE THIS COULD BE CONFIRMED VISUALLY, THE POORLY FORMED PORTION WAS REINFORCED WITH ADDITIONAL SUTURES.
16-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 09-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A NEPHRECTOMY PROCEDURE, THE USER WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. THE ISSUE OCCURRED IN ONE HANDLE AND THREE RELOADS. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2023 | Injury | us surgical puerto rico | 09-01-2024 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING THE INITIAL ROBOTIC LOW ANTERIOR COLON RESECTION PROCEDURE, WHEN THE DEVICE WAS USED FOR END-TO-END ANASTOMOSIS, THERE WAS NO ISSUE REPORTED, AND THE SURGEON CLAIMED THE DEVICE FUNCTIONED NORMALLY, AND INTRAOPERATIVE COLONOSCOPY SHOWED NO EVIDENCE OF MALFORMITY ON THE STAPLE LINE OR ANY CONCERNS REGARDING TISSUE DONUTS. HENCE, 12 DAYS POSTOPERATIVELY, AN ANASTOMOTIC LEAK WAS IDENTIFIED. THE PATIENT HAD TO UNDERGO A RE-OPERATION, WHICH WAS AN ADDITIONAL 3-HOUR PROCEDURE, AND A POOR STAPLE FORMATION WAS IDENTIFIED. THE STAPLE LINE WAS FIXED BY OVER SEWING. THE PATIENT LEFT WITH COLOSTOMY AFTER THE SECOND SURGERY AND EXTENDED THE HOSPITALIZATION.
18-12-2023 | Injury | us surgical puerto rico | 09-01-2024 | EEA | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE INITIAL ROBOTIC LOW ANTERIOR COLON RESECTION PROCEDURE, WHEN THE DEVICE WAS USED FOR END-TO-END ANASTOMOSIS, THERE WAS NO ISSUE REPORTED, AND THE SURGEON CLAIMED THE DEVICE FUNCTIONED NORMALLY, AND INTRAOPERATIVE COLONOSCOPY SHOWED NO EVIDENCE OF MALFORMITY ON THE STAPLE LINE OR ANY CONCERNS REGARDING TISSUE DONUTS. HENCE, 12 DAYS POSTOPERATIVELY, AN ANASTOMOTIC LEAK WAS IDENTIFIED. THE PATIENT HAD TO UNDERGO A RE-OPERATION, WHICH WAS AN ADDITIONAL 3-HOUR PROCEDURE, AND A POOR STAPLE FORMATION WAS IDENTIFIED. THE STAPLE LINE WAS FIXED BY OVER SEWING. THE PATIENT LEFT WITH COLOSTOMY AFTER THE SECOND SURGERY AND EXTENDED THE HOSPITALIZATION.
18-12-2023 | Injury | us surgical puerto rico | 09-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING THE INITIAL ROBOTIC LOW ANTERIOR COLON RESECTION PROCEDURE, WHEN THE DEVICE WAS USED FOR END-TO-END ANASTOMOSIS, THERE WAS NO ISSUE REPORTED, AND THE SURGEON CLAIMED THE DEVICE FUNCTIONED NORMALLY, AND INTRAOPERATIVE COLONOSCOPY SHOWED NO EVIDENCE OF MALFORMITY ON THE STAPLE LINE OR ANY CONCERNS REGARDING TISSUE DONUTS. HENCE, 12 DAYS POSTOPERATIVELY, AN ANASTOMOTIC LEAK WAS IDENTIFIED. THE PATIENT HAD TO UNDERGO A RE-OPERATION, WHICH WAS AN ADDITIONAL 3-HOUR PROCEDURE, AND A POOR STAPLE FORMATION WAS IDENTIFIED. THE STAPLE LINE WAS FIXED BY OVER SEWING. THE PATIENT LEFT WITH COLOSTOMY AFTER THE SECOND SURGERY AND EXTENDED THE HOSPITALIZATION.
18-12-2023 | Injury | us surgical puerto rico | 09-01-2024 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING THE INITIAL ROBOTIC LOW ANTERIOR COLON RESECTION PROCEDURE, WHEN THE DEVICE WAS USED FOR END-TO-END ANASTOMOSIS, THERE WAS NO ISSUE REPORTED, AND THE SURGEON CLAIMED THE DEVICE FUNCTIONED NORMALLY, AND INTRAOPERATIVE COLONOSCOPY SHOWED NO EVIDENCE OF MALFORMITY ON THE STAPLE LINE OR ANY CONCERNS REGARDING TISSUE DONUTS. HENCE, 12 DAYS POSTOPERATIVELY, AN ANASTOMOTIC LEAK WAS IDENTIFIED. THE PATIENT HAD TO UNDERGO A RE-OPERATION, WHICH WAS AN ADDITIONAL 3-HOUR PROCEDURE, AND A POOR STAPLE FORMATION WAS IDENTIFIED. THE STAPLE LINE WAS FIXED BY OVER SEWING. THE PATIENT LEFT WITH COLOSTOMY AFTER THE SECOND SURGERY AND EXTENDED THE HOSPITALIZATION.
11-12-2023 | Injury | davis & geck caribe | 09-01-2024 | UNKNOWN VLOC ENDO STITCH RELOAD ABS | Detachment of Device or Device Component | A0501 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HYSTERECTOMY PROCEDURE, WHEN CLOSING THE VAGINAL CUFF, THE SUTURE NEEDLE BROKE OFF INTO THE CAVITY OF THE PATIENT. AN X-RAY WAS PERFORMED TO LOCATE THE NEEDLE. THE SURGEON WAS ABLE TO RETRIEVE THE NEEDLE, AND THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES. MOREOVER, THE EVENT LEAD TO EXTEND PATIENT'S HOSPITALIZATION.
17-12-2023 | Injury | covidien lp llc north haven | 09-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: ACCORDING TO THE REPORTER, TWO DAYS AFTER THE DEVICE WAS USED IN A PARTIAL GASTRECTOMY DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT HAD TO BE RE-OPERATED AS THERE WAS BLOOD IN THE DRAIN. THE ISSUE INVOLVED TWO RELOADS. DURING REOPERATION, THE PATIENT BLED AT THE STAPLE LINE, AND THERE WAS A HEMATOMA ALONG THE STAPLE LINE. TO RESOLVE THE ISSUE, A CLIP APPLIER WAS USED ALONG THE LINES. IT WAS NOTED THAT THE PATIENT HAD LOSS ABOUT 1000-1200MILLILTERS (ML) OF BLOOD BUT DID NOT HAVE ANY TRANSFUSIONS. HOSPITALIZATION WAS EXTENDED AS WELL.
17-12-2023 | Injury | covidien lp llc north haven | 09-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, TWO DAYS AFTER THE DEVICE WAS USED IN A PARTIAL GASTRECTOMY DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT HAD TO BE RE-OPERATED AS THERE WAS BLOOD IN THE DRAIN. THE ISSUE INVOLVED TWO RELOADS. DURING REOPERATION, THE PATIENT BLED AT THE STAPLE LINE, AND THERE WAS A HEMATOMA ALONG THE STAPLE LINE. TO RESOLVE THE ISSUE, A CLIP APPLIER WAS USED ALONG THE LINES. IT WAS NOTED THAT THE PATIENT HAD LOSS ABOUT 1000-1200MILLILTERS (ML) OF BLOOD BUT DID NOT HAVE ANY TRANSFUSIONS. HOSPITALIZATION WAS EXTENDED AS WELL.
14-11-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Display or Visual Feedback Problem | A0902 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, HANDLE WAS UNABLE TO RECOGNIZE THE ADAPTER AND SHOWING A FILE ERROR. THE ADAPTER WAS USED WITHOUT ISSUE ON A DIFFERENT HANDLE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
14-11-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Electrical /Electronic Property Problem | A07 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, HANDLE WAS UNABLE TO RECOGNIZE THE ADAPTER AND SHOWING A FILE ERROR. THE ADAPTER WAS USED WITHOUT ISSUE ON A DIFFERENT HANDLE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
14-11-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, HANDLE WAS UNABLE TO RECOGNIZE THE ADAPTER AND SHOWING A FILE ERROR. THE ADAPTER WAS USED WITHOUT ISSUE ON A DIFFERENT HANDLE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
09-11-2023 | Malfunction | covidien lp llc north haven | 09-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE DEVICE WAS DEFECTIVE; THE TIP WAS OPEN. THE ISSUE OCCURRED IN TWO RELOADS. THERE WAS NO PATIENT INVOLVEMENT.
03-01-2024 | Malfunction | covidien | 09-01-2024 | COVIDIEN GIA STAPLER WITH TRI-STAPLER TECHNOLOGY |  |  | Failure to Anastomose | Event Description: 2 MESENTERIC WINDOWS WERE MADE IMMEDIATELY PROXIMAL AND DISTAL TO THIS AND THE BOWEL DIVIDED HERE WITH A GIA STAPLER PURPLE LOAD. MESENTERY WAS TAKEN WITH LIGASURE DEVICE. I THEN FASHIONED A SIDE-TO-SIDE FUNCTIONAL END-TO-END ANASTOMOSIS WITH ANOTHER FIRING OF THE GIA PURPLE STAPLER. COMMON ENTEROTOMY WAS CLOSED WITH A TA 60. ANASTOMOSIS WAS PALPATED AND PATENT. I THEN RAN THE BOWEL AGAIN AND NO FURTHER EVIDENCE OF INJURY WAS NOTED AGAIN. HOWEVER WHEN I REACHED MY SMALL-BOWEL ANASTOMOSIS THERE WAS SMALL BUBBLES OF AIR COMING FROM IT. IT WAS THEREFORE RESECTED. IT WAS THEN REPERFORMED WITH ANOTHER GIA 80 AND A SIDE-TO-SIDE FUNCTIONAL END-TO-END FASHION. THIS TIME THE COMMON ENTEROTOMY WAS CLOSED WITH ANOTHER FIRING OF THE 80 GIA STAPLER. THIS WAS VERY CLOSELY INSPECTED AND THIS TIME WAS LEAKING FROM 2 SEPARATE AREAS OF THE STAPLE LINE FROM THE ENTEROTOMY CLOSURE. THIS IS AN OCCURRENCE OF UNKNOWN STAPLE FAILURE. THE PROXIMAL BOWEL WAS FAIRLY DISTENDED BUT I WOULD STILL EXPECT THE STAPLE LINE TO HOLD. I CONSIDERED REDOING THE ANASTOMOSIS AGAIN AND HAND SEWING IT BUT DECIDED AGAINST IT. GIVEN ALL OF THIS AND THE FACT THAT THE PATIENT HAD BEEN UNDER ANESTHESIA FOR ALMOST 9 HOURS AT THIS POINT, I ELECTED TO LEAVE HER IN DISCONTINUITY AND PLACE AN ABTHERA WITH PLANS TO TAKE HER TO THE (INTENSIVE CARE UNIT), RESUSCITATE HER POTENTIALLY DIURESE HER, DECOMPRESS THE BOWEL AND RETURNED TO COMPLETE THE ANASTOMOSIS ANOTHER DAY.
06-12-2023 | Injury | unknown | 09-01-2024 | BLUE LOAD GIA STAPLER |  |  | Bacterial Infection | Event Description: SURGICAL NOTES: WE THEN IDENTIFIED THE LOOP OF SIGMOID COLON FOR THE COLOSTOMY. WE ENSURED THAT THERE WAS NO TWISTING. THE BOWEL WAS ISOLATED WITH ELECTROCAUTERY AND THEN TRANSECTED WITH A SINGLE FIRING OF A BLUE LOAD GIA (GASTROINTESTINAL ANASTOMOSIS) STAPLER. THE PROXIMAL END WAS BROUGHT THROUGH THE TREPHINE. THE DISTAL END WAS BROUGHT UP JUST AS A SMALL CORNER THROUGH THE TREPHINE. THE MIDLINE FASCIA WAS CLOSED WITH SINGLE-STRANDED 0 PDS (POLYDIOXANONE). THE SKIN AND SUBCUTANEOUS FAT WAS IRRIGATED AND CLOSED WITH 4-0 MONOCRYL IN A SUBCUTICULAR FASHION. DERMABOND WAS PLACED AS A DRESSING. THE COLOSTOMY WAS MATURED BY EXCISING THE PROXIMAL STAPLE LINE. IT WAS MATURED IN A BROOKE FASHION WITH INTERRUPTED 3-0 VICRYL SUTURES WITH THE EXCEPTION OF THE INFERIOR PORTION. AT THE INFERIOR PORTION OF THE TREPHINE THE MUCOUS FISTULA WAS CREATED BY EXCISING A CORNER OF THE STAPLE LINE OF THE DISTAL SIGMOID COLON. THIS WAS MATURED TO THE SKIN WITH INTERRUPTED 3-0 VICRYL SUTURES. THE INFERIOR PORTION OF THE PROXIMAL SIGMOID COLON WAS THEN MATURED DOWN TO THE MUCOUS FISTULA AS WELL AS TO THE SKIN WITH INTERRUPTED 3-0 VICRYL SUTURES. A STOMA APPLIANCE WAS APPLIED. HE WAS AWAKENED EXTUBATED AND TRANSFERRED TO THE PACU (POST-ANESTHESIA CARE UNIT) IN STABLE CONDITION. COUNTS WERE CORRECT AT THE END OF THE CASE. EXPLORATORY LAPAROTOMY: ABDOMINAL WALL WASHOUT WAS NEEDED DUE TO STAPLES FAILING INDINGS: FECULENT FLUID THROUGHOUT THE ABDOMEN, FOUND TO HAVE MUCUS FISTULA SEPARATED FROM ABDOMINAL WALL WITH PERFORATION JUST DISTAL TO THE STAPLE LINE. ABDOMEN IRRIGATED WITH 10+L OF FLUID, COLOSTOMY IRRIGATED WITH 3L SALINE TO PROMOTE DRAINAGE OF STOOL. AFTER THE ABDOMINAL WASHOUT AN INFECTION WAS DEVELPOED. PATIENT IS STILL FIGHTING INFECTION IN HOSPITAL.
06-12-2023 | Injury | unknown | 09-01-2024 | BLUE LOAD GIA STAPLER |  |  | Fistula | Event Description: SURGICAL NOTES: WE THEN IDENTIFIED THE LOOP OF SIGMOID COLON FOR THE COLOSTOMY. WE ENSURED THAT THERE WAS NO TWISTING. THE BOWEL WAS ISOLATED WITH ELECTROCAUTERY AND THEN TRANSECTED WITH A SINGLE FIRING OF A BLUE LOAD GIA (GASTROINTESTINAL ANASTOMOSIS) STAPLER. THE PROXIMAL END WAS BROUGHT THROUGH THE TREPHINE. THE DISTAL END WAS BROUGHT UP JUST AS A SMALL CORNER THROUGH THE TREPHINE. THE MIDLINE FASCIA WAS CLOSED WITH SINGLE-STRANDED 0 PDS (POLYDIOXANONE). THE SKIN AND SUBCUTANEOUS FAT WAS IRRIGATED AND CLOSED WITH 4-0 MONOCRYL IN A SUBCUTICULAR FASHION. DERMABOND WAS PLACED AS A DRESSING. THE COLOSTOMY WAS MATURED BY EXCISING THE PROXIMAL STAPLE LINE. IT WAS MATURED IN A BROOKE FASHION WITH INTERRUPTED 3-0 VICRYL SUTURES WITH THE EXCEPTION OF THE INFERIOR PORTION. AT THE INFERIOR PORTION OF THE TREPHINE THE MUCOUS FISTULA WAS CREATED BY EXCISING A CORNER OF THE STAPLE LINE OF THE DISTAL SIGMOID COLON. THIS WAS MATURED TO THE SKIN WITH INTERRUPTED 3-0 VICRYL SUTURES. THE INFERIOR PORTION OF THE PROXIMAL SIGMOID COLON WAS THEN MATURED DOWN TO THE MUCOUS FISTULA AS WELL AS TO THE SKIN WITH INTERRUPTED 3-0 VICRYL SUTURES. A STOMA APPLIANCE WAS APPLIED. HE WAS AWAKENED EXTUBATED AND TRANSFERRED TO THE PACU (POST-ANESTHESIA CARE UNIT) IN STABLE CONDITION. COUNTS WERE CORRECT AT THE END OF THE CASE. EXPLORATORY LAPAROTOMY: ABDOMINAL WALL WASHOUT WAS NEEDED DUE TO STAPLES FAILING INDINGS: FECULENT FLUID THROUGHOUT THE ABDOMEN, FOUND TO HAVE MUCUS FISTULA SEPARATED FROM ABDOMINAL WALL WITH PERFORATION JUST DISTAL TO THE STAPLE LINE. ABDOMEN IRRIGATED WITH 10+L OF FLUID, COLOSTOMY IRRIGATED WITH 3L SALINE TO PROMOTE DRAINAGE OF STOOL. AFTER THE ABDOMINAL WASHOUT AN INFECTION WAS DEVELPOED. PATIENT IS STILL FIGHTING INFECTION IN HOSPITAL.
06-12-2023 | Injury | unknown | 09-01-2024 | BLUE LOAD GIA STAPLER |  |  | Bowel Perforation | Event Description: SURGICAL NOTES: WE THEN IDENTIFIED THE LOOP OF SIGMOID COLON FOR THE COLOSTOMY. WE ENSURED THAT THERE WAS NO TWISTING. THE BOWEL WAS ISOLATED WITH ELECTROCAUTERY AND THEN TRANSECTED WITH A SINGLE FIRING OF A BLUE LOAD GIA (GASTROINTESTINAL ANASTOMOSIS) STAPLER. THE PROXIMAL END WAS BROUGHT THROUGH THE TREPHINE. THE DISTAL END WAS BROUGHT UP JUST AS A SMALL CORNER THROUGH THE TREPHINE. THE MIDLINE FASCIA WAS CLOSED WITH SINGLE-STRANDED 0 PDS (POLYDIOXANONE). THE SKIN AND SUBCUTANEOUS FAT WAS IRRIGATED AND CLOSED WITH 4-0 MONOCRYL IN A SUBCUTICULAR FASHION. DERMABOND WAS PLACED AS A DRESSING. THE COLOSTOMY WAS MATURED BY EXCISING THE PROXIMAL STAPLE LINE. IT WAS MATURED IN A BROOKE FASHION WITH INTERRUPTED 3-0 VICRYL SUTURES WITH THE EXCEPTION OF THE INFERIOR PORTION. AT THE INFERIOR PORTION OF THE TREPHINE THE MUCOUS FISTULA WAS CREATED BY EXCISING A CORNER OF THE STAPLE LINE OF THE DISTAL SIGMOID COLON. THIS WAS MATURED TO THE SKIN WITH INTERRUPTED 3-0 VICRYL SUTURES. THE INFERIOR PORTION OF THE PROXIMAL SIGMOID COLON WAS THEN MATURED DOWN TO THE MUCOUS FISTULA AS WELL AS TO THE SKIN WITH INTERRUPTED 3-0 VICRYL SUTURES. A STOMA APPLIANCE WAS APPLIED. HE WAS AWAKENED EXTUBATED AND TRANSFERRED TO THE PACU (POST-ANESTHESIA CARE UNIT) IN STABLE CONDITION. COUNTS WERE CORRECT AT THE END OF THE CASE. EXPLORATORY LAPAROTOMY: ABDOMINAL WALL WASHOUT WAS NEEDED DUE TO STAPLES FAILING INDINGS: FECULENT FLUID THROUGHOUT THE ABDOMEN, FOUND TO HAVE MUCUS FISTULA SEPARATED FROM ABDOMINAL WALL WITH PERFORATION JUST DISTAL TO THE STAPLE LINE. ABDOMEN IRRIGATED WITH 10+L OF FLUID, COLOSTOMY IRRIGATED WITH 3L SALINE TO PROMOTE DRAINAGE OF STOOL. AFTER THE ABDOMINAL WASHOUT AN INFECTION WAS DEVELPOED. PATIENT IS STILL FIGHTING INFECTION IN HOSPITAL.
19-12-2023 | Malfunction | intuitive surgical | 10-01-2024 | SUREFORM | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
14-12-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | ENDO ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTRECTOMY THE DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | 75MM SELECTABLE NEW TLC | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE INTERVENTION OF DUODENUM-CEFALOPANCREASECTOMIA, THE 3TH ROW OF CLIPS, MADE WITH LINEAR CUTTER, GREEN CLIPS, WAS NOT DONE WITH CORRECT CLOSURE OF THE SAME THAT IS TO SAY IS NOT COMPLETED AS EXPECTED. CLIPS NOT CLOSED. NECESSARY RE-RESECTION OF INTESTINAL LOOP IN CUL DE SAC. SURGERY COMPLETED WITH OTHER DEVICE.
06-12-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTRECTOMY THE DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2023 | Injury | ethicon endo-surgery | 10-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST OP TO A LOW ANTERIOR RESECTION THE PATIENT HAD A POST-OP LEAK. PATIENT CAME BACK ON (B)(6) 2023 FOR THE LEAK TO BE TAKEN CARE OF WITH NO ISSUES. NO ISSUES WERE DETECTED WITH THE STAPLER ITSELF, ACCORDING TO THE SURGEON.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | CIRCULAR MECH STD 25MM | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TOTAL GASTRECTOMY PROCEDURE, THE CIRCULAR STAPLER HEAD DOES NOT ENGAGE WITH THE MACHINE BODY. NO ISSUES FOR THE PATIENT.
01-12-2023 | Injury | ethicon endo-surgery | 10-01-2024 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE PRODUCT DOES NOT STAPLE WELL, THE MACHINE HAS SOME CARTRIDGES OF RELOADS STAPLES THAT ARE DEFECTIVE, THE LEG OF THE STAPLES DO NOT CLOSE COMPLETELY, IT SEEMS LIKE IT IS NOT LONG ENOUGH. THE PROCEDURE WAS DELAYED BY 30 MINUTES. THE LAP SURGERY HAD TO BE CONVERTED TO OPEN SURGERY AND WAS COMPLETED SUCCESSFULLY.
01-12-2023 | Injury | ethicon endo-surgery | 10-01-2024 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE PRODUCT DOES NOT STAPLE WELL, THE MACHINE HAS SOME CARTRIDGES OF RELOADS STAPLES THAT ARE DEFECTIVE, THE LEG OF THE STAPLES DO NOT CLOSE COMPLETELY, IT SEEMS LIKE IT IS NOT LONG ENOUGH. THE PROCEDURE WAS DELAYED BY 30 MINUTES. THE LAP SURGERY HAD TO BE CONVERTED TO OPEN SURGERY AND WAS COMPLETED SUCCESSFULLY.
19-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | UNKNOWN ENDO GIA SULU | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHEN POSITIONING THE DEVICE INSIDE THE PATIENT PRIOR TO FIRING, THE DEVICE HAD UNEXPECTED ARTICULATION/ROTATION. THEY HAVE TRIED TO REPLACE THE RELOADS BUT WITH THE SAME RESULTS. THE HANDLE WAS CONTINUED TO BE USED DESPITE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHEN POSITIONING THE DEVICE INSIDE THE PATIENT PRIOR TO FIRING, THE DEVICE HAD UNEXPECTED ARTICULATION/ROTATION. THEY HAVE TRIED TO REPLACE THE RELOADS BUT WITH THE SAME RESULTS. THE HANDLE WAS CONTINUED TO BE USED DESPITE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-05-2023 | Malfunction | intuitive surgical | 10-01-2024 | SUREFORM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS SURGICAL PROCEDURE, THE STAPLER EXPERIENCED A PARTIAL FIRE. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY.
19-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN AN ENDOSCOPIC THORACIC LOBECTOMY, WHILE DETACHING TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | EEA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMORRHOIDECTOMY, DURING RESECTION OR INCISION, THE DEVICE WAS DIFFICULT TO LOAD. IT WAS DIFFICULT TO CLOSE THE DEVICE. THERE WAS AN UNUSUAL EFFORT REQUIRED TO TURN THE TWIST KNOB TO VISUALIZE THE GREEN BAR. THE DEVICE DID NOT FIRE. THE SURGEON ATTEMPTED TO FIRE THE DEVICE TWICE WITHOUT SUCCESS. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN OPEN HEMORRHOIDECTOMY, DURING RESECTION OR INCISION, THE DEVICE WAS DIFFICULT TO LOAD. IT WAS DIFFICULT TO CLOSE THE DEVICE. THERE WAS AN UNUSUAL EFFORT REQUIRED TO TURN THE TWIST KNOB TO VISUALIZE THE GREEN BAR. THE DEVICE DID NOT FIRE. THE SURGEON ATTEMPTED TO FIRE THE DEVICE TWICE WITHOUT SUCCESS. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
28-11-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRIC BYPASS, DURING THE FIRST SHOT IN THE GASTRIC POUCH, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, BUT THE STAPLER DID NOT FIRE. THE DEVICE JAMMED. THERE WAS A DELAY IN THE SURGERY FOR 30 MINUTES AS THE SURGEON REVIEWED THE DEVICE ISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THE RELOAD WAS USED IN ANOTHER STAPLER.
14-11-2023 | Malfunction | davis & geck caribe | 10-01-2024 | UNKNOWN ENDO STITCH INSTRUMENT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, THE NEEDLE BROKE IN HALF. THERE WAS NO PATIENT INJURY.
30-11-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SECOND BLUE MAGAZINE, THE STAPLER CUT PERFECTLY, BUT DID NOT CLAMPED AT ALL. NO PATIENT HARM NOR SIGNIFICANT DELAY.
07-12-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | CARTRIDGE FINAL ASSY 60MM BLUE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2023 | Injury | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: ACCORDING TO THE REPORTER, TWO DAYS AFTER THE DEVICE WAS USED IN A PARTIAL GASTRECTOMY DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT HAD TO BE RE-OPERATED AS THERE WAS BLOOD IN THE DRAIN. THE ISSUE INVOLVED TWO RELOADS. DURING REOPERATION, THE PATIENT BLED AT THE STAPLE LINE, AND THERE WAS A HEMATOMA ALONG THE STAPLE LINE. TO RESOLVE THE ISSUE, A CLIP APPLIER WAS USED ALONG THE LINES. IT WAS NOTED THAT THE PATIENT HAD LOSS ABOUT 1000-1200MILLILTERS (ML) OF BLOOD BUT DID NOT HAVE ANY TRANSFUSIONS. HOSPITALIZATION WAS EXTENDED AS WELL.
17-12-2023 | Injury | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, TWO DAYS AFTER THE DEVICE WAS USED IN A PARTIAL GASTRECTOMY DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE PATIENT HAD TO BE RE-OPERATED AS THERE WAS BLOOD IN THE DRAIN. THE ISSUE INVOLVED TWO RELOADS. DURING REOPERATION, THE PATIENT BLED AT THE STAPLE LINE, AND THERE WAS A HEMATOMA ALONG THE STAPLE LINE. TO RESOLVE THE ISSUE, A CLIP APPLIER WAS USED ALONG THE LINES. IT WAS NOTED THAT THE PATIENT HAD LOSS ABOUT 1000-1200MILLILTERS (ML) OF BLOOD BUT DID NOT HAVE ANY TRANSFUSIONS. HOSPITALIZATION WAS EXTENDED AS WELL.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A MINIMALLY INVASIVE ESOPHAGECTOMY, THE FIRST STAPLER FIRED BUT WOULDN'T OPEN. THEY HAD TO MANUALLY CRANK TO OPEN JAWS AFTER OVERRIDE/HOME BUTTON DIDN'T WORK. THE SECOND STAPLER WOULDN'T CLOSE WITH RELOAD IN IT. DEVICE WAS NOT USED AND WAS SET TO THE SIDE. FOR THE ENSEAL DEVICE THEY STATED A SMALL PIECE OF METAL CAME OFF THE TIP. THERE WAS A DELAY DURING PROCEDURE. PROCEDURE WAS COMPLETED WITH A NEW DEVICE WITH NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A MINIMALLY INVASIVE ESOPHAGECTOMY, THE FIRST STAPLER FIRED BUT WOULDN'T OPEN. THEY HAD TO MANUALLY CRANK TO OPEN JAWS AFTER OVERRIDE/HOME BUTTON DIDN'T WORK. THE SECOND STAPLER WOULDN'T CLOSE WITH RELOAD IN IT. DEVICE WAS NOT USED AND WAS SET TO THE SIDE. FOR THE ENSEAL DEVICE THEY STATED A SMALL PIECE OF METAL CAME OFF THE TIP. THERE WAS A DELAY DURING PROCEDURE. PROCEDURE WAS COMPLETED WITH A NEW DEVICE WITH NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A MINIMALLY INVASIVE ESOPHAGECTOMY, THE FIRST STAPLER FIRED BUT WOULDN'T OPEN. THEY HAD TO MANUALLY CRANK TO OPEN JAWS AFTER OVERRIDE/HOME BUTTON DIDN'T WORK. THE SECOND STAPLER WOULDN'T CLOSE WITH RELOAD IN IT. DEVICE WAS NOT USED AND WAS SET TO THE SIDE. FOR THE ENSEAL DEVICE THEY STATED A SMALL PIECE OF METAL CAME OFF THE TIP. THERE WAS A DELAY DURING PROCEDURE. PROCEDURE WAS COMPLETED WITH A NEW DEVICE WITH NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS STUCK WHEN TRYING TO REOPEN AND THE STAPLES DID NOT FORM BUT IT DID NOT CUT PROPERLY DURING A VANKERS PROCEDURE. THEY USED A NEW DEVICE TO COMPLETE CASE WITH NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS STUCK WHEN TRYING TO REOPEN AND THE STAPLES DID NOT FORM BUT IT DID NOT CUT PROPERLY DURING A VANKERS PROCEDURE. THEY USED A NEW DEVICE TO COMPLETE CASE WITH NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE WAS STUCK WHEN TRYING TO REOPEN AND THE STAPLES DID NOT FORM BUT IT DID NOT CUT PROPERLY DURING A VANKERS PROCEDURE. THEY USED A NEW DEVICE TO COMPLETE CASE WITH NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | ECHELON 3000 60MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY FIRED 3 TIMES ON WEDGE RESECTION PROCEDURE BUT ON 4TH FIRE JAWS WOULDN'T FIRE ON THE BUTTRESS. THEY CLOSED HANDLE, HIT HOME BUTTON. HANDLE RELEASED BUT JAWS STILL DID NOT OPEN. THEY GOT ANOTHER DEVICE TO COMPLETE THE CASE WITH NO PATIENT HARM.
01-01-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | ECHELON 3000 60MM COMPACT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THEY FIRED 3 TIMES ON WEDGE RESECTION PROCEDURE BUT ON 4TH FIRE JAWS WOULDN'T FIRE ON THE BUTTRESS. THEY CLOSED HANDLE, HIT HOME BUTTON. HANDLE RELEASED BUT JAWS STILL DID NOT OPEN. THEY GOT ANOTHER DEVICE TO COMPLETE THE CASE WITH NO PATIENT HARM.
06-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOTOMY PROCEDURE, THE DEVICE WAS USED WITH NO ISSUE WITH SIX RELOADS. HOWEVER, ON THE SEVENTH RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE; BUT, THE DEVICE WAS MAKING A "CLICKING SOUND" AND WOULD NOT FIRE THE STAPLE. SEVERAL ATTEMPTS WERE MADE TO CORRECT THE ISSUE WITHOUT RESOLUTION. THE SURGEON REQUESTED ANOTHER 45-MILLIMETER RELOAD AND HANDLE BOTH OF THE SAME TYPE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOTOMY PROCEDURE, THE DEVICE WAS USED WITH NO ISSUE WITH SIX RELOADS. HOWEVER, ON THE SEVENTH RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE; BUT, THE DEVICE WAS MAKING A "CLICKING SOUND" AND WOULD NOT FIRE THE STAPLE. SEVERAL ATTEMPTS WERE MADE TO CORRECT THE ISSUE WITHOUT RESOLUTION. THE SURGEON REQUESTED ANOTHER 45-MILLIMETER RELOAD AND HANDLE BOTH OF THE SAME TYPE. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION PROCEDURE, AFTER THE FIRST SQUEEZE WHILE FIRING THE FIRST RELOAD ON LUNG TISSUE, THE HANDLE'S KNIFE RETRACTED ON ITS OWN TO THE BEGINNING WITHOUT ANY ACTIVITIES IN THE SURGICAL SITE. PRE-FIRED STAPLES WERE OBSERVED IN THE RELOAD. TO RESOLVE THE ISSUE, A NEW HANDLE WITH A DIFFERENT LOT NUMBER WAS USED TO FIRE THE OTHER RELOADS. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WEDGE RESECTION PROCEDURE, AFTER THE FIRST SQUEEZE WHILE FIRING THE FIRST RELOAD ON LUNG TISSUE, THE HANDLE'S KNIFE RETRACTED ON ITS OWN TO THE BEGINNING WITHOUT ANY ACTIVITIES IN THE SURGICAL SITE. PRE-FIRED STAPLES WERE OBSERVED IN THE RELOAD. TO RESOLVE THE ISSUE, A NEW HANDLE WITH A DIFFERENT LOT NUMBER WAS USED TO FIRE THE OTHER RELOADS. THERE WAS NO PATIENT INJURY.
10-01-2024 | Malfunction | covidien lp llc north haven | 10-01-2024 | UNKNOWN STAPLING DEVICE | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE JAWS WERE NOT FULLY CLOSED.
10-01-2024 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE JAWS WERE NOT FULLY CLOSED.
18-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, THE LOADING UNIT WAS USED TO REMOVE THE VESSELS. WHEN THE VESSELS WERE SET DOWN, THE LOADING UNIT COULD NOT BE TRIGGERED. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. A NEW LOADING UNIT HAS BEEN INSTALLED AND SAME HANDLE WAS USED TO COMPLETE THE OPERATION. THERE WAS NO PATIENT INJURY
07-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING FIRING, THE DEVICE WAS ABLE TO FIRE BUT STOPPED AFTER APPROXIMATELY 15MM. THE STAPLE LINE WAS INCOMPLETE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, ON GASTRIC TISSUE, AFTER INSERTING THE RELOAD, THE STAPLER COULD NOT BE FIRED. THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE HANDLE WAS CHANGED WITH THE SAME RELOAD WHICH CAN BE FIRE. THERE WAS NO PATIENT INJURY.
16-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC TAIL RESECTION, THE OPENING AND CLOSING OF THE JAWS WERE CHECKED, BUT WHEN THE DEVICE WAS INSERTED INTO THE ABDOMINAL CAVITY AND AN ATTEMPT WAS MADE TO CLOSE AFTER CLAMPING THE TISSUE, THE JAWS COULD NOT BE CLOSED EVEN THOUGH THE HANDLE WAS SQUEEZED. THE DEVICE WAS REMOVED FROM THE TROCAR AND WHEN IT WAS CHECKED EXTERNALLY, THE JAWS CLOSED. THE USAGE WAS STOPPED, AND A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
15-12-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO CONTINUE THE SURGERY. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO CONTINUE THE SURGERY. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC TAIL RESECTION, THE OPENING AND CLOSING OF THE JAWS WERE CHECKED, BUT WHEN THE DEVICE WAS INSERTED INTO THE ABDOMINAL CAVITY AND AN ATTEMPT WAS MADE TO CLOSE AFTER CLAMPING THE TISSUE, THE JAWS COULD NOT BE CLOSED EVEN THOUGH THE HANDLE WAS SQUEEZED. THE DEVICE WAS REMOVED FROM THE TROCAR AND WHEN IT WAS CHECKED EXTERNALLY, THE JAWS CLOSED. THE USAGE WAS STOPPED, AND A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-12-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | GST GREEN RELOAD, 60MM, 6 ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING BARIATRIC PROCEDURE, BEFORE PUTTING IT IN THE GUN A SMALL BROWN PIECE FELL OF THE END. THEY GOT ANOTHER ONE TO COMPLETE THE CASE WITH NO PATIENT HARM WAS REPORTED.
16-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | UNKNOWN ENDO GIA SULU | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, A PART OF THE RELOAD BECAME STUCK INSIDE. THERE WAS A PIN TO RELEASE THE RELOAD, BUT IT DID NOT RELEASE. IT WAS JAMMED IN PRETTY WELL AND COULD NOT BE REMOVED WITH GENTLE FORCE. ANOTHER ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | UNKNOWN ENDO GIA SULU | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, A PART OF THE RELOAD BECAME STUCK INSIDE. THERE WAS A PIN TO RELEASE THE RELOAD, BUT IT DID NOT RELEASE. IT WAS JAMMED IN PRETTY WELL AND COULD NOT BE REMOVED WITH GENTLE FORCE. ANOTHER ADAPTER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | ethicon endo-surgery | 10-01-2024 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, AFTER STAPLING, THE DOCTOR FOUND THAT IN THE THIRD LINE OF DISTAL STAPLES, 2 STAPLES REMAINED WITHOUT CLOSURE. IT WAS NOT STAPLED AGAIN, BUT THERE WAS A NEED TO REINFORCE THE SUTURE IN THIS STAPLE LINE. NO PATIENT CONSEQUENCES WERE REPORTED.
19-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 10-01-2024 | ENDO GIA | Failure to Cycle | A050703 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, WHEN THE GASTRIC TISSUE WAS STAPLED, IT WAS FOUND OUT THAT THE STAPLES ON THE STOMACH TISSUE WAS PILED UP, PUSHING THE TISSUE FORWARD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
11-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOT-ASSISTED TOTAL GASTRECTOMY, WHEN THE RELOAD WAS LOADED INTO THE ADAPTER DURING SMALL INTESTINAL RESECTION, THE OPENING/CLOSING CHECK WAS PERFORMED WITHOUT AN ISSUE, THE GREEN BUTTON WAS PRESSED, AND THE CONTROL BUTTON WAS PRESSED DOWN. HOWEVER, THE DEVICE COULD NOT BE FIRED. A NEW 45MM RELOAD WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
14-11-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL GASTRECTOMY FOR GASTRIC CANCER, WHEN FIRING WAS BEING PERFORMED ON THE SMALL INTESTINE, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED, AND THE PHENOMENON OF FAILURE TO FIRE OCCURRED. A 45-MILLIMETER RELOAD WAS USED, BUT THE DEVICE DID NOT FIRE AGAIN. TO RESOLVE THE ISSUE, A NEW 60-MILLIMETER RELOAD OF THE SAME MODEL AND HANDLE WAS USED. THERE WAS NO PATIENT INJURY.
14-12-2023 | Injury | intuitive surgical | 10-01-2024 | SUREFORM | Failure to Cut | A050702 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, TISSUE WAS DRAGGED WHEN A GREEN SUREFORM 60 RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. ON THE FIRST FIRE, THE SURGEON SUCCESSFULLY FIRED A BLACK STAPLER RELOAD ON STOMACH TISSUE WITH NO ISSUES. ON THE SECOND FIRE, UPON FIRING THE STAPLER, THE BLADE STARTED DRAGGING THE TISSUE INSTEAD OF CUTTING THROUGH IT. THE SURGEON HIT THE EMERGENCY STOP BUTTON AND REMOVED THE STAPLER. THE GREEN STAPLER RELOAD HAD FIRED AROUND 10MM OR LESS THAN A QUARTER OF THE WAY; THE STAPLES DEPLOYED WERE FULLY FORMED, BUT NOT DEPLOYED INTO THE TISSUE. THERE WAS A SMALL AMOUNT OF BLEEDING WHICH THE SURGEON STATED WAS NOT CLINICALLY IMPORTANT; AND A BIPOLAR FORCEPS INSTRUMENT WAS USED FOR COAGULATION. A BLUE STAPLER RELOAD WAS THEN USED TO FINISH THE STAPLE LINE, AND A SMALL AMOUNT OF ADDITIONAL TISSUE NEEDED TO BE REMOVED DUE TO THE FIRING FAILURE; BUT THE SURGEON STATED IT WAS NOT CLINICALLY IMPORTANT. THE STAPLER WAS USED FOR THE REMAINDER OF THE PROCEDURE WITH NO FURTHER ISSUES. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE SURGEON STATED THERE WAS NO HARM TO THE PATIENT AND THERE WERE NO POST OPERATIVE COMPLICATIONS.
14-12-2023 | Injury | intuitive surgical | 10-01-2024 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, TISSUE WAS DRAGGED WHEN A GREEN SUREFORM 60 RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. ON THE FIRST FIRE, THE SURGEON SUCCESSFULLY FIRED A BLACK STAPLER RELOAD ON STOMACH TISSUE WITH NO ISSUES. ON THE SECOND FIRE, UPON FIRING THE STAPLER, THE BLADE STARTED DRAGGING THE TISSUE INSTEAD OF CUTTING THROUGH IT. THE SURGEON HIT THE EMERGENCY STOP BUTTON AND REMOVED THE STAPLER. THE GREEN STAPLER RELOAD HAD FIRED AROUND 10MM OR LESS THAN A QUARTER OF THE WAY; THE STAPLES DEPLOYED WERE FULLY FORMED, BUT NOT DEPLOYED INTO THE TISSUE. THERE WAS A SMALL AMOUNT OF BLEEDING WHICH THE SURGEON STATED WAS NOT CLINICALLY IMPORTANT; AND A BIPOLAR FORCEPS INSTRUMENT WAS USED FOR COAGULATION. A BLUE STAPLER RELOAD WAS THEN USED TO FINISH THE STAPLE LINE, AND A SMALL AMOUNT OF ADDITIONAL TISSUE NEEDED TO BE REMOVED DUE TO THE FIRING FAILURE; BUT THE SURGEON STATED IT WAS NOT CLINICALLY IMPORTANT. THE STAPLER WAS USED FOR THE REMAINDER OF THE PROCEDURE WITH NO FURTHER ISSUES. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE SURGEON STATED THERE WAS NO HARM TO THE PATIENT AND THERE WERE NO POST OPERATIVE COMPLICATIONS.
14-12-2023 | Injury | intuitive surgical | 10-01-2024 | SUREFORM | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, TISSUE WAS DRAGGED WHEN A GREEN SUREFORM 60 RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. ON THE FIRST FIRE, THE SURGEON SUCCESSFULLY FIRED A BLACK STAPLER RELOAD ON STOMACH TISSUE WITH NO ISSUES. ON THE SECOND FIRE, UPON FIRING THE STAPLER, THE BLADE STARTED DRAGGING THE TISSUE INSTEAD OF CUTTING THROUGH IT. THE SURGEON HIT THE EMERGENCY STOP BUTTON AND REMOVED THE STAPLER. THE GREEN STAPLER RELOAD HAD FIRED AROUND 10MM OR LESS THAN A QUARTER OF THE WAY; THE STAPLES DEPLOYED WERE FULLY FORMED, BUT NOT DEPLOYED INTO THE TISSUE. THERE WAS A SMALL AMOUNT OF BLEEDING WHICH THE SURGEON STATED WAS NOT CLINICALLY IMPORTANT; AND A BIPOLAR FORCEPS INSTRUMENT WAS USED FOR COAGULATION. A BLUE STAPLER RELOAD WAS THEN USED TO FINISH THE STAPLE LINE, AND A SMALL AMOUNT OF ADDITIONAL TISSUE NEEDED TO BE REMOVED DUE TO THE FIRING FAILURE; BUT THE SURGEON STATED IT WAS NOT CLINICALLY IMPORTANT. THE STAPLER WAS USED FOR THE REMAINDER OF THE PROCEDURE WITH NO FURTHER ISSUES. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE SURGEON STATED THERE WAS NO HARM TO THE PATIENT AND THERE WERE NO POST OPERATIVE COMPLICATIONS.
14-12-2023 | Injury | intuitive surgical | 10-01-2024 | SUREFORM | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, TISSUE WAS DRAGGED WHEN A GREEN SUREFORM 60 RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT. ON THE FIRST FIRE, THE SURGEON SUCCESSFULLY FIRED A BLACK STAPLER RELOAD ON STOMACH TISSUE WITH NO ISSUES. ON THE SECOND FIRE, UPON FIRING THE STAPLER, THE BLADE STARTED DRAGGING THE TISSUE INSTEAD OF CUTTING THROUGH IT. THE SURGEON HIT THE EMERGENCY STOP BUTTON AND REMOVED THE STAPLER. THE GREEN STAPLER RELOAD HAD FIRED AROUND 10MM OR LESS THAN A QUARTER OF THE WAY; THE STAPLES DEPLOYED WERE FULLY FORMED, BUT NOT DEPLOYED INTO THE TISSUE. THERE WAS A SMALL AMOUNT OF BLEEDING WHICH THE SURGEON STATED WAS NOT CLINICALLY IMPORTANT; AND A BIPOLAR FORCEPS INSTRUMENT WAS USED FOR COAGULATION. A BLUE STAPLER RELOAD WAS THEN USED TO FINISH THE STAPLE LINE, AND A SMALL AMOUNT OF ADDITIONAL TISSUE NEEDED TO BE REMOVED DUE TO THE FIRING FAILURE; BUT THE SURGEON STATED IT WAS NOT CLINICALLY IMPORTANT. THE STAPLER WAS USED FOR THE REMAINDER OF THE PROCEDURE WITH NO FURTHER ISSUES. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE SURGEON STATED THERE WAS NO HARM TO THE PATIENT AND THERE WERE NO POST OPERATIVE COMPLICATIONS.
21-10-2023 | Injury | covidien lp llc north haven | 10-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF THE STUDY PERFORMED BETWEEN JANUARY 2009 AND JUNE 2022, A RETROSPECTIVE STUDY AIMED TO ESTABLISH THE ¿BEST POSSIBLE¿ OUTCOMES FOR ROBOTIC BARIATRIC SURGERY AND COMPARE THESE OUTCOMES WITH ESTABLISHED LAPAROSCOPIC BENCHMARKS. IT WAS NOTED THAT THE ANASTOMOTIC TECHNIQUE WAS PERFORMED WITH A MANUAL STAPLER. THERE WERE 9097 VARIOUS ROBOTIC-ASSISTED BARIATRIC PROCEDURES PERFORMED, AND COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, POSTOPERATIVE BLEEDING, AND MARGINAL ULCER. INTERVENTION AND PATIENT OUTCOMES ARE UNKNOWN. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE WERE MOBILITY DISORDER, SMALL BOWEL OBSTRUCTION/INTERNAL HERNIA, DYSPHAGIA/GASTRO-OESOPHAGEAL REFLUX DISEASE/STENOSIS, ABDOMINAL OR OSTEO-ARTICULAR PAIN, AND DEEP-VEIN THROMBOSIS/PULMONARY EMBOLISM.
21-10-2023 | Injury | covidien lp llc north haven | 10-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE OF THE STUDY PERFORMED BETWEEN JANUARY 2009 AND JUNE 2022, A RETROSPECTIVE STUDY AIMED TO ESTABLISH THE ¿BEST POSSIBLE¿ OUTCOMES FOR ROBOTIC BARIATRIC SURGERY AND COMPARE THESE OUTCOMES WITH ESTABLISHED LAPAROSCOPIC BENCHMARKS. IT WAS NOTED THAT THE ANASTOMOTIC TECHNIQUE WAS PERFORMED WITH A MANUAL STAPLER. THERE WERE 9097 VARIOUS ROBOTIC-ASSISTED BARIATRIC PROCEDURES PERFORMED, AND COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, POSTOPERATIVE BLEEDING, AND MARGINAL ULCER. INTERVENTION AND PATIENT OUTCOMES ARE UNKNOWN. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE WERE MOBILITY DISORDER, SMALL BOWEL OBSTRUCTION/INTERNAL HERNIA, DYSPHAGIA/GASTRO-OESOPHAGEAL REFLUX DISEASE/STENOSIS, ABDOMINAL OR OSTEO-ARTICULAR PAIN, AND DEEP-VEIN THROMBOSIS/PULMONARY EMBOLISM.
21-10-2023 | Injury | covidien lp llc north haven | 10-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Ulcer | Event Description: ACCORDING TO THE LITERATURE SOURCE OF THE STUDY PERFORMED BETWEEN JANUARY 2009 AND JUNE 2022, A RETROSPECTIVE STUDY AIMED TO ESTABLISH THE ¿BEST POSSIBLE¿ OUTCOMES FOR ROBOTIC BARIATRIC SURGERY AND COMPARE THESE OUTCOMES WITH ESTABLISHED LAPAROSCOPIC BENCHMARKS. IT WAS NOTED THAT THE ANASTOMOTIC TECHNIQUE WAS PERFORMED WITH A MANUAL STAPLER. THERE WERE 9097 VARIOUS ROBOTIC-ASSISTED BARIATRIC PROCEDURES PERFORMED, AND COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, POSTOPERATIVE BLEEDING, AND MARGINAL ULCER. INTERVENTION AND PATIENT OUTCOMES ARE UNKNOWN. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE WERE MOBILITY DISORDER, SMALL BOWEL OBSTRUCTION/INTERNAL HERNIA, DYSPHAGIA/GASTRO-OESOPHAGEAL REFLUX DISEASE/STENOSIS, ABDOMINAL OR OSTEO-ARTICULAR PAIN, AND DEEP-VEIN THROMBOSIS/PULMONARY EMBOLISM.
19-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VATS (VIDEO-ASSISTED THORACIC SURGERY) SEGMENTECTOMY PROCEDURE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VATS (VIDEO-ASSISTED THORACIC SURGERY) SEGMENTECTOMY PROCEDURE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VATS (VIDEO-ASSISTED THORACIC SURGERY) SEGMENTECTOMY PROCEDURE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VATS (VIDEO-ASSISTED THORACIC SURGERY) SEGMENTECTOMY PROCEDURE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PANCREATECTOMY, WHEN CLOSING THE HANDLE, THE STAPLER COULD NOT BE FIRED. THE ISSUE OCCURRED ON TWO MORE HANDLES. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AFTER FIRING TWO BLACK REINFORCED RELOADS, THE SURGEON CHANGED IT TO THE PURPLE RELOAD. WHEN FIRING, THE BLADE STOPPED AND DISPLAYED AN ERROR MESSAGE WHICH SHOWED THE JAWS WITH THE SIGN OF OPENING IT. THE SURGEON TRIED IT WITH A NEW PURPLE RELOAD AGAIN AND HAD THE SAME ISSUE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE AND NO OTHER STEPS WERE TAKEN TO REMOVED IT. THE KNIFE WAS ABLE TO BE RETRACTED WITHOUT OPENING IT MANUALLY. THE SURGEON END UP USING A THIRD BLACK NON REINFORCED RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AFTER FIRING TWO BLACK REINFORCED RELOADS, THE SURGEON CHANGED IT TO THE PURPLE RELOAD. WHEN FIRING, THE BLADE STOPPED AND DISPLAYED AN ERROR MESSAGE WHICH SHOWED THE JAWS WITH THE SIGN OF OPENING IT. THE SURGEON TRIED IT WITH A NEW PURPLE RELOAD AGAIN AND HAD THE SAME ISSUE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE AND NO OTHER STEPS WERE TAKEN TO REMOVED IT. THE KNIFE WAS ABLE TO BE RETRACTED WITHOUT OPENING IT MANUALLY. THE SURGEON END UP USING A THIRD BLACK NON REINFORCED RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-11-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING/MAINTENANCE, THERE WAS POWER HANDLE ERROR. THE REBOOT FAILED AND HANDLE WOULD NOT TURN ON. CHARGER BAY, PORTS, CABLES ALL WORK WELL AND NO ISSUES CHARGING THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
03-11-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING/MAINTENANCE, THERE WAS POWER HANDLE ERROR. THE REBOOT FAILED AND HANDLE WOULD NOT TURN ON. CHARGER BAY, PORTS, CABLES ALL WORK WELL AND NO ISSUES CHARGING THE HANDLE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT ANALYSIS OF DATA STORED ON THE HANDLE SHOWS EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
21-11-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY, THE JAWS OF THE DEVICE LOCK ON TISSUE. ANOTHER STAPLER WAS THEN USED TO CUT OUT THE DEVICE. ADDITIONALLY, THE USER HAD DIFFICULTY UNLOADING THE DEVICE AS THE HANDLE WOULD NOT RELEASE.
21-11-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | UNKNOWN ENDO GIA SULU | Human-Device Interface Problem | A22 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY, THE JAWS OF THE DEVICE LOCK ON TISSUE. ANOTHER STAPLER WAS THEN USED TO CUT OUT THE DEVICE. ADDITIONALLY, THE USER HAD DIFFICULTY UNLOADING THE DEVICE AS THE HANDLE WOULD NOT RELEASE.
12-12-2023 | Injury | intuitive surgical | 10-01-2024 | SUREFORM | Retraction Problem | A0510 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PROCEDURE, THE BLADE OF THE BLUE SUREFORM 60 STAPLER RELOAD FAILED TO RETURN TO ITS PROPER POSITION, AND A NURSE SUSTAINED AN UNSPECIFIED INJURY AS A RESULT. INTUITIVE SURGICAL, INC. (ISI) CONTACTED THE SITE FOR ADDITIONAL INFORMATION; HOWEVER, AT THE TIME OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN PROVIDED.
12-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AFTER FIRING TWO BLACK REINFORCED RELOADS, THE SURGEON CHANGED IT TO THE PURPLE RELOAD. WHEN FIRING, THE BLADE STOPPED AND DISPLAYED AN ERROR MESSAGE WHICH SHOWED THE JAWS WITH THE SIGN OF OPENING IT. THE SURGEON TRIED IT WITH A NEW PURPLE RELOAD AGAIN AND HAD THE SAME ISSUE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE AND NO OTHER STEPS WERE TAKEN TO REMOVED IT. THE KNIFE WAS ABLE TO BE RETRACTED WITHOUT OPENING IT MANUALLY. THE SURGEON END UP USING A THIRD BLACK NON REINFORCED RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AFTER FIRING TWO BLACK REINFORCED RELOADS, THE SURGEON CHANGED IT TO THE PURPLE RELOAD. WHEN FIRING, THE BLADE STOPPED AND DISPLAYED AN ERROR MESSAGE WHICH SHOWED THE JAWS WITH THE SIGN OF OPENING IT. THE SURGEON TRIED IT WITH A NEW PURPLE RELOAD AGAIN AND HAD THE SAME ISSUE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE AND NO OTHER STEPS WERE TAKEN TO REMOVED IT. THE KNIFE WAS ABLE TO BE RETRACTED WITHOUT OPENING IT MANUALLY. THE SURGEON END UP USING A THIRD BLACK NON REINFORCED RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDICECTOMY PROCEDURE, IN THE REMOVAL OF THE APPENDIX, THE DEVICE FAILED TO FIRE AND THE JAWS LOCKED ON THE TISSUE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDICECTOMY PROCEDURE, IN THE REMOVAL OF THE APPENDIX, THE DEVICE FAILED TO FIRE AND THE JAWS LOCKED ON THE TISSUE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AFTER FIRING TWO BLACK REINFORCED RELOADS, THE SURGEON CHANGED IT TO THE PURPLE RELOAD. WHEN FIRING, THE BLADE STOPPED AND DISPLAYED AN ERROR MESSAGE WHICH SHOWED THE JAWS WITH THE SIGN OF OPENING IT. THE SURGEON TRIED IT WITH A NEW PURPLE RELOAD AGAIN AND HAD THE SAME ISSUE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE AND NO OTHER STEPS WERE TAKEN TO REMOVED IT. THE KNIFE WAS ABLE TO BE RETRACTED WITHOUT OPENING IT MANUALLY. THE SURGEON END UP USING A THIRD BLACK NON REINFORCED RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AFTER FIRING TWO BLACK REINFORCED RELOADS, THE SURGEON CHANGED IT TO THE PURPLE RELOAD. WHEN FIRING, THE BLADE STOPPED AND DISPLAYED AN ERROR MESSAGE WHICH SHOWED THE JAWS WITH THE SIGN OF OPENING IT. THE SURGEON TRIED IT WITH A NEW PURPLE RELOAD AGAIN AND HAD THE SAME ISSUE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE AND NO OTHER STEPS WERE TAKEN TO REMOVED IT. THE KNIFE WAS ABLE TO BE RETRACTED WITHOUT OPENING IT MANUALLY. THE SURGEON END UP USING A THIRD BLACK NON REINFORCED RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDICECTOMY PROCEDURE, IN THE REMOVAL OF THE APPENDIX, THE DEVICE FAILED TO FIRE AND THE JAWS LOCKED ON THE TISSUE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDICECTOMY PROCEDURE, IN THE REMOVAL OF THE APPENDIX, THE DEVICE FAILED TO FIRE AND THE JAWS LOCKED ON THE TISSUE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
21-11-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY, THE JAWS OF THE DEVICE LOCK ON TISSUE. ANOTHER STAPLER WAS THEN USED TO CUT OUT THE DEVICE. ADDITIONALLY, THE USER HAD DIFFICULTY UNLOADING THE DEVICE AS THE HANDLE WOULD NOT RELEASE.
21-11-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC DISTAL PANCREATECTOMY, THE JAWS OF THE DEVICE LOCK ON TISSUE. ANOTHER STAPLER WAS THEN USED TO CUT OUT THE DEVICE. ADDITIONALLY, THE USER HAD DIFFICULTY UNLOADING THE DEVICE AS THE HANDLE WOULD NOT RELEASE.
06-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ULTRA-LOW ANTERIOR PROCTECTOMY AND PARTIAL INTERNAL SPHINCTEROTOMY, AFTER PRESSING THE GREEN BUTTON, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AND COULD NOT BE FIRED. AFTER SQUEEZING HARD, THE HANDLE MADE A CLICKING SOUND AND COULD NOT BE USED NORMALLY. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 10-01-2024 | ENDO GIA | Failure to Cycle | A050703 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, WHEN APPLIED ON GASTRIC TISSUE, THE RELOAD WAS FIRED BUT THE STAPLES ON THE STOMACH TISSUE PILED UP, PUSHING THE TISSUE FORWARD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
19-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 10-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VATS (VIDEO-ASSISTED THORACIC SURGERY) LOBECTOMY PROCEDURE, THE USER WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 10-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VATS (VIDEO-ASSISTED THORACIC SURGERY) LOBECTOMY PROCEDURE, THE USER WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PANCREATECTOMY, WHEN CLOSING THE HANDLE, THE STAPLER COULD NOT BE FIRED. THE ISSUE OCCURRED ON TWO MORE HANDLES. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PANCREATECTOMY, WHEN CLOSING THE HANDLE, THE STAPLER COULD NOT BE FIRED. THE ISSUE OCCURRED ON TWO MORE HANDLES. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) SEGMENTECTOMY, IN STAPLING THE UPPER RIGHT LOBE TISSUE, THE HANDLE SUCCESSFULLY FIRED THE TWO RELOADS. HOWEVER, ON THE THREE PURPLE RELOADS, THE SURGEON WAS NOT ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE DEVICES DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE SAME THREE RELOADS. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHEN POSITIONING THE DEVICE INSIDE THE PATIENT PRIOR TO FIRING THE DEVICE HAD UNEXPECTED ARTICULATION/ROTATION. THEY HAVE TRIED TO REPLACE THE RELOADS BUT WITH THE SAME RESULTS. THE HANDLE WAS CONTINUED TO BE USE DESPITE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | us surgical puerto rico | 10-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC PARTIAL PULMONARY RESECTION, DURING PULMONARY PARENCHYMA RESECTION, THE TISSUE WAS SET ONTO THE JAWS, AND AN ATTEMPT WAS MADE TO PERFORM FIRING. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED, AND FIRING COULD NOT BE PERFORMED. THE ISSUE OCCURRED WITH TWO. IN ONE OF THE TWO RELOADS, THE HANDLE WAS ABLE TO BE SQUEEZED, BUT THE DEVICE DID NOT FIRE. THE SURGERY WAS COMPLETED USING A NEW PRODUCT. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | covidien lp llc north haven | 10-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC PARTIAL PULMONARY RESECTION, DURING PULMONARY PARENCHYMA RESECTION, THE TISSUE WAS SET ONTO THE JAWS, AND AN ATTEMPT WAS MADE TO PERFORM FIRING. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED, AND FIRING COULD NOT BE PERFORMED. THE ISSUE OCCURRED WITH TWO. IN ONE OF THE TWO RELOADS, THE HANDLE WAS ABLE TO BE SQUEEZED, BUT THE DEVICE DID NOT FIRE. THE SURGERY WAS COMPLETED USING A NEW PRODUCT. THERE WAS NO PATIENT INJURY.
30-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ESOPHAGUS PROCEDURE THE SURGERY, COULD NOT FIRE . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE THE DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE THE DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE THE DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-11-2023 | Injury | ethicon endo-surgery | 11-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST OP TO A LOW ANTERIOR RESECTION THE PATIENT PRESENTS WITH A LOW GRADE FEVER WHICH ADVANCES TO A HIGH GRADE FEVER. PATIENT IS THEN BROUGHT BACK TO THE OR WHERE A LEAK WAS FOUND AND THE PATIENT RECEIVED A COLOSTOMY. SURGEON WILL GO BACK IN ABOUT SIX WEEKS TO ASSESS AND REVERSE THE COLOSTOMY. DURING THE INITIAL PROCEDURE THE STAPLING SEQUENCES WENT SMOOTHLY, THE DONUTS WERE INTACT AND NOTHING APPEARED TO BE WRONG AT THE TIME.
01-11-2023 | Injury | ethicon endo-surgery | 11-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: IT WAS REPORTED THAT POST OP TO A LOW ANTERIOR RESECTION THE PATIENT PRESENTS WITH A LOW GRADE FEVER WHICH ADVANCES TO A HIGH GRADE FEVER. PATIENT IS THEN BROUGHT BACK TO THE OR WHERE A LEAK WAS FOUND AND THE PATIENT RECEIVED A COLOSTOMY. SURGEON WILL GO BACK IN ABOUT SIX WEEKS TO ASSESS AND REVERSE THE COLOSTOMY. DURING THE INITIAL PROCEDURE THE STAPLING SEQUENCES WENT SMOOTHLY, THE DONUTS WERE INTACT AND NOTHING APPEARED TO BE WRONG AT THE TIME.
01-01-2024 | Malfunction | ethicon endo-surgery | 11-01-2024 | CIRCULAR MECH STD 29MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE TISSUE WAS NOT CUT COMPLETE, ANOTHER ONE WAS SUBSTITUTED TO COMPLETE THE SURGERY.THERE WERE NO PATIENT CONSEQUENCES.
27-10-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ENDO ILS 29MM, CURVED | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION THE ANVIL HEAD AND BASE OF THE STAPLER DID NOT FIT WELL, IT WAS VERY LOOSE AND COULD NOT BE JAMMED, BUT IT COULD FIRE, SO THE STAPLER WAS NOT REPLACED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2023 | Injury | covidien lp llc north haven | 11-01-2024 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WITH ROBOT SUPPORT LEFT LIVER LOBECTOMY, WHEN THE GLISSON WAS RESECTED, IT WAS AS IF A HOLE WAS OPENED IN THE STAPLE LINE, THE STAPLE POPPED, AND A BLEEDING OCCURRED AFTER RELEASE. BLEEDING WAS MORE THAN 500 CC AND IT WAS NEAR TO THE TUMOR SITE. A TRANSITION TO LAPAROTOMY WAS PERFORMED. IT WAS UNKNOWN WHETHER THE CAUSE WAS CO2 BLOOD VESSEL CONTAMINATION BY THE PNEUMOPERITONEUM DEVICE. HOWEVER, THE PATIENT'S SPO2 DID NOT RISE, AND THE SURGERY WAS DIFFICULT. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES AND THERE WAS TISSUE LOSS OF MORE THAN ONE INCH. EVEN AFTER THE TUMOR SITE WAS REMOVED, AND THE PATIENT WAS MANAGED IN THE INTENSIVE CARE UNIT (ICU) FOR FIVE DAYS, THE HOSPITALIZATION WAS EXTENDED, THE THERE WAS A BLOOD LOSS OF MORE THAN 500CC THAT NEEDED A TRANSFUSION DUE TO THE PRODUCT ISSUE. THE PATIENT DIED BUT THE SURGEON BELIEVED THAT DEATH WAS NOT RELATED TO THE DEVICE ISSUE.
14-12-2023 | Injury | covidien lp llc north haven | 11-01-2024 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WITH ROBOT SUPPORT LEFT LIVER LOBECTOMY, WHEN THE GLISSON WAS RESECTED, IT WAS AS IF A HOLE WAS OPENED IN THE STAPLE LINE, THE STAPLE POPPED, AND A BLEEDING OCCURRED AFTER RELEASE. BLEEDING WAS MORE THAN 500 CC AND IT WAS NEAR TO THE TUMOR SITE. A TRANSITION TO LAPAROTOMY WAS PERFORMED. IT WAS UNKNOWN WHETHER THE CAUSE WAS CO2 BLOOD VESSEL CONTAMINATION BY THE PNEUMOPERITONEUM DEVICE. HOWEVER, THE PATIENT'S SPO2 DID NOT RISE, AND THE SURGERY WAS DIFFICULT. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES AND THERE WAS TISSUE LOSS OF MORE THAN ONE INCH. EVEN AFTER THE TUMOR SITE WAS REMOVED, AND THE PATIENT WAS MANAGED IN THE INTENSIVE CARE UNIT (ICU) FOR FIVE DAYS, THE HOSPITALIZATION WAS EXTENDED, THE THERE WAS A BLOOD LOSS OF MORE THAN 500CC THAT NEEDED A TRANSFUSION DUE TO THE PRODUCT ISSUE. THE PATIENT DIED BUT THE SURGEON BELIEVED THAT DEATH WAS NOT RELATED TO THE DEVICE ISSUE.
14-12-2023 | Injury | covidien lp llc north haven | 11-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WITH ROBOT SUPPORT LEFT LIVER LOBECTOMY, WHEN THE GLISSON WAS RESECTED, IT WAS AS IF A HOLE WAS OPENED IN THE STAPLE LINE, THE STAPLE POPPED, AND A BLEEDING OCCURRED AFTER RELEASE. BLEEDING WAS MORE THAN 500 CC AND IT WAS NEAR TO THE TUMOR SITE. A TRANSITION TO LAPAROTOMY WAS PERFORMED. IT WAS UNKNOWN WHETHER THE CAUSE WAS CO2 BLOOD VESSEL CONTAMINATION BY THE PNEUMOPERITONEUM DEVICE. HOWEVER, THE PATIENT'S SPO2 DID NOT RISE, AND THE SURGERY WAS DIFFICULT. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES AND THERE WAS TISSUE LOSS OF MORE THAN ONE INCH. EVEN AFTER THE TUMOR SITE WAS REMOVED, AND THE PATIENT WAS MANAGED IN THE INTENSIVE CARE UNIT (ICU) FOR FIVE DAYS, THE HOSPITALIZATION WAS EXTENDED, THE THERE WAS A BLOOD LOSS OF MORE THAN 500CC THAT NEEDED A TRANSFUSION DUE TO THE PRODUCT ISSUE. THE PATIENT DIED BUT THE SURGEON BELIEVED THAT DEATH WAS NOT RELATED TO THE DEVICE ISSUE.
14-12-2023 | Injury | covidien lp llc north haven | 11-01-2024 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC WITH ROBOT SUPPORT LEFT LIVER LOBECTOMY, WHEN THE GLISSON WAS RESECTED, IT WAS AS IF A HOLE WAS OPENED IN THE STAPLE LINE, THE STAPLE POPPED, AND A BLEEDING OCCURRED AFTER RELEASE. BLEEDING WAS MORE THAN 500 CC AND IT WAS NEAR TO THE TUMOR SITE. A TRANSITION TO LAPAROTOMY WAS PERFORMED. IT WAS UNKNOWN WHETHER THE CAUSE WAS CO2 BLOOD VESSEL CONTAMINATION BY THE PNEUMOPERITONEUM DEVICE. HOWEVER, THE PATIENT'S SPO2 DID NOT RISE, AND THE SURGERY WAS DIFFICULT. SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES AND THERE WAS TISSUE LOSS OF MORE THAN ONE INCH. EVEN AFTER THE TUMOR SITE WAS REMOVED, AND THE PATIENT WAS MANAGED IN THE INTENSIVE CARE UNIT (ICU) FOR FIVE DAYS, THE HOSPITALIZATION WAS EXTENDED, THE THERE WAS A BLOOD LOSS OF MORE THAN 500CC THAT NEEDED A TRANSFUSION DUE TO THE PRODUCT ISSUE. THE PATIENT DIED BUT THE SURGEON BELIEVED THAT DEATH WAS NOT RELATED TO THE DEVICE ISSUE.
02-01-2024 | Injury | ethicon endo-surgery | 11-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT ABOUT 4-5 DAYS POST OP FROM A LOW ANTERIOR RESECTION THE PATIENT PRESENTS WITH A LOW GRADE FEVER WHICH ADVANCES TO A HIGH GRADE FEVER. PATIENT IS THEN BROUGHT BACK TO THE OR WHERE NO LEAK COULD BE SEEN BUT THE SURGEON DECIDED TO GIVE THE PATIENT A COLOSTOMY. SURGEON WILL GO BACK IN ABOUT SIX WEEKS TO ASSESS AND REVERSE THE COLOSTOMY. DURING THE INITIAL PROCEDURE THE STAPLING SEQUENCES WENT SMOOTHLY, THE DONUTS WERE INTACT AND NOTHING APPEARED TO BE WRONG AT THE TIME.
02-01-2024 | Injury | ethicon endo-surgery | 11-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: IT WAS REPORTED THAT ABOUT 4-5 DAYS POST OP FROM A LOW ANTERIOR RESECTION THE PATIENT PRESENTS WITH A LOW GRADE FEVER WHICH ADVANCES TO A HIGH GRADE FEVER. PATIENT IS THEN BROUGHT BACK TO THE OR WHERE NO LEAK COULD BE SEEN BUT THE SURGEON DECIDED TO GIVE THE PATIENT A COLOSTOMY. SURGEON WILL GO BACK IN ABOUT SIX WEEKS TO ASSESS AND REVERSE THE COLOSTOMY. DURING THE INITIAL PROCEDURE THE STAPLING SEQUENCES WENT SMOOTHLY, THE DONUTS WERE INTACT AND NOTHING APPEARED TO BE WRONG AT THE TIME.
01-08-2023 | Injury | ethicon endo-surgery | 11-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED POST OP TO A LOW ANTERIOR RESECTION THE PATIENT PRESENTS WITH A LOW GRADE FEVER WHICH ADVANCES TO A HIGH GRADE FEVER. PATIENT IS THEN BROUGHT BACK TO THE OR WHERE A LEAK WAS FOUND AND THE PATIENT RECEIVED A COLOSTOMY. SURGEON WILL GO BACK IN ABOUT SIX WEEKS TO ASSESS AND REVERSE THE COLOSTOMY. DURING THE INITIAL PROCEDURE THE STAPLING SEQUENCES WENT SMOOTHLY, THE DONUTS WERE INTACT AND NOTHING APPEARED TO BE WRONG AT THE TIME.
01-08-2023 | Injury | ethicon endo-surgery | 11-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Fever | Event Description: IT WAS REPORTED POST OP TO A LOW ANTERIOR RESECTION THE PATIENT PRESENTS WITH A LOW GRADE FEVER WHICH ADVANCES TO A HIGH GRADE FEVER. PATIENT IS THEN BROUGHT BACK TO THE OR WHERE A LEAK WAS FOUND AND THE PATIENT RECEIVED A COLOSTOMY. SURGEON WILL GO BACK IN ABOUT SIX WEEKS TO ASSESS AND REVERSE THE COLOSTOMY. DURING THE INITIAL PROCEDURE THE STAPLING SEQUENCES WENT SMOOTHLY, THE DONUTS WERE INTACT AND NOTHING APPEARED TO BE WRONG AT THE TIME.
01-01-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE STAPLER FIRED AND THEN WOULDN'T OPEN. WAS REPLACED WITH ANOTHER STAPLER PLEE60A TO COMPLETE THE CASE. IT IS UNKNOWN IF THERE WERE ANY PATIENT CONSEQUENCES.
19-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 11-01-2024 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UMBILICAL HERNIA REPAIR PROCEDURE WHEN USING AN OPTIFIX FIXATION DEVICE, THERE WAS DELAY IN FIRING AND THEN THE FASTENERS WOULD BREAK AND BURST INTO PIECES. IT WAS REPORTED THAT THE BROKEN FASTENERS WERE REMOVED AND ANOTHER OPTIFIX DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO CONTINUE THE SURGERY. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO CONTINUE THE SURGERY. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHILE THE SURGEON WAS FIRING THE PVE35A AND THE STAPLE LINE DID NOT COMPLETE THE FIRING AND THAT IT WAS TEARING TISSUE. THE SURGEON SUTURED THE DEFECT AND DISCARDED MALFUNCTIONING DEVICE. SURGEON REQUESTED A 2ND DEVICE TO FINISH THE CASE. THE CASE WAS COMPLETED SUCCESSFULLY AND PATIENT IMPACT WAS NONE.
14-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHILE THE SURGEON WAS FIRING THE PVE35A AND THE STAPLE LINE DID NOT COMPLETE THE FIRING AND THAT IT WAS TEARING TISSUE. THE SURGEON SUTURED THE DEFECT AND DISCARDED MALFUNCTIONING DEVICE. SURGEON REQUESTED A 2ND DEVICE TO FINISH THE CASE. THE CASE WAS COMPLETED SUCCESSFULLY AND PATIENT IMPACT WAS NONE.
11-01-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | POWERED PLUS STANDARD 45 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED FROM THE CUSTOMER THAT A SCREW FELL OUT OF THE DEVICE AFTER IT WAS FIRED. NO PATIENT HARM.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, AT THE TIME OF ENTERAL WHOLE ANASTOMOSIS, METHYLENE BLUE WAS FILTERED INTO THE ENTERAL ANASTOMOSIS IN RELATION TO THE MECHANICAL SUTURE AND THEN THE LAPAROSCOPIC POINT PLACEMENT WAS PERFORMED. HOWEVER, THE DEVICE DID NOT FIRE. THE SURGEON USED SUTURE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-11-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, AT THE TIME OF ENTERAL WHOLE ANASTOMOSIS, METHYLENE BLUE WAS FILTERED INTO THE ENTERAL ANASTOMOSIS IN RELATION TO THE MECHANICAL SUTURE AND THEN THE LAPAROSCOPIC POINT PLACEMENT WAS PERFORMED. HOWEVER, THE DEVICE DID NOT FIRE. THE SURGEON USED SUTURE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC PARTIAL PULMONARY RESECTION, DURING PULMONARY PARENCHYMA RESECTION, THE TISSUE WAS SET ONTO THE JAWS, AND AN ATTEMPT WAS MADE TO PERFORM FIRING. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED, AND FIRING COULD NOT BE PERFORMED. THE ISSUE OCCURRED WITH TWO. IN ONE OF THE TWO RELOADS, THE HANDLE WAS ABLE TO BE SQUEEZED, BUT THE DEVICE DID NOT FIRE. THE SURGERY WAS COMPLETED USING A NEW PRODUCT. THERE WAS NO PATIENT INJURY.
02-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE.
02-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE.
02-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE.
02-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE.
01-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | GST WHITE RELOAD, 60MM, 6 ROW | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE THE SURGERY, NOTED THE SEAL OPENED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE WITH A POSITIVE METHYLENE BLUE TEST IN THE ANASTOMOSIS. NO PATIENT INJURY.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE WITH A POSITIVE METHYLENE BLUE TEST IN THE ANASTOMOSIS. NO PATIENT INJURY.
16-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | POWERED PLUS STANDARD 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE 3RD FIRING. ANOTHER DEVICE WAS USED TO CONTINUE THE SURGERY. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE. AND THEN DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE 3RD FIRING. ANOTHER DEVICE WAS USED TO CONTINUE THE SURGERY. OUR CUSTOMER SUSPECT IT WAS BATTERY ISSUE. AND THEN DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE WITH A POSITIVE METHYLENE BLUE TEST IN THE ANASTOMOSIS. NO PATIENT INJURY.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE WITH A POSITIVE METHYLENE BLUE TEST IN THE ANASTOMOSIS. NO PATIENT INJURY.
05-10-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: DUPLICATE PE OF (B)(4) ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE DEVICE WAS SUCCESSFULLY FIRED ACROSS THE RECTUM, IT WAS ATTEMPTED TO RETRACT THE ANVIL, HOWEVER, THE SURGEON MET SIGNIFICANT RESISTANCE ON THE HANDLE OF THE STAPLER. THE ANVIL COULD ONLY PULL BACK HALFWAY WHICH MADE IT IMPOSSIBLE TO PULL OUT. THE ANVIL WAS MOBILE FORWARD, BUT NOT BACKWARDS. WHEN THE SALES REPRESENTATIVE CAME TO THE OPERATING ROOM, THE DEVICE WAS SUCCESSFULLY REMOVED FROM THE PATIENT AFTER TROUBLESHOOTING THE ISSUE. NO PATIENT INJURY.
05-10-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: DUPLICATE PE OF (B)(4) ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE DEVICE WAS SUCCESSFULLY FIRED ACROSS THE RECTUM, IT WAS ATTEMPTED TO RETRACT THE ANVIL, HOWEVER, THE SURGEON MET SIGNIFICANT RESISTANCE ON THE HANDLE OF THE STAPLER. THE ANVIL COULD ONLY PULL BACK HALFWAY WHICH MADE IT IMPOSSIBLE TO PULL OUT. THE ANVIL WAS MOBILE FORWARD, BUT NOT BACKWARDS. WHEN THE SALES REPRESENTATIVE CAME TO THE OPERATING ROOM, THE DEVICE WAS SUCCESSFULLY REMOVED FROM THE PATIENT AFTER TROUBLESHOOTING THE ISSUE. NO PATIENT INJURY.
05-10-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: DUPLICATE PE OF (B)(4) ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE DEVICE WAS SUCCESSFULLY FIRED ACROSS THE RECTUM, IT WAS ATTEMPTED TO RETRACT THE ANVIL, HOWEVER, THE SURGEON MET SIGNIFICANT RESISTANCE ON THE HANDLE OF THE STAPLER. THE ANVIL COULD ONLY PULL BACK HALFWAY WHICH MADE IT IMPOSSIBLE TO PULL OUT. THE ANVIL WAS MOBILE FORWARD, BUT NOT BACKWARDS. WHEN THE SALES REPRESENTATIVE CAME TO THE OPERATING ROOM, THE DEVICE WAS SUCCESSFULLY REMOVED FROM THE PATIENT AFTER TROUBLESHOOTING THE ISSUE. NO PATIENT INJURY.
05-10-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE DEVICE WAS SUCCESSFULLY FIRED ACROSS THE RECTUM, IT WAS ATTEMPTED TO RETRACT THE ANVIL, HOWEVER, THE SURGEON MET SIGNIFICANT RESISTANCE ON THE HANDLE OF THE STAPLER. THE ANVIL COULD ONLY PULL BACK HALFWAY WHICH MADE IT IMPOSSIBLE TO PULL OUT. THE ANVIL WAS MOBILE FORWARD, BUT NOT BACKWARDS. WHEN THE SALES REPRESENTATIVE CAME TO THE OPERATING ROOM, THE DEVICE WAS SUCCESSFULLY REMOVED FROM THE PATIENT AFTER TROUBLESHOOTING THE ISSUE. THERE WAS NO PATIENT INJURY.
05-10-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE DEVICE WAS SUCCESSFULLY FIRED ACROSS THE RECTUM, IT WAS ATTEMPTED TO RETRACT THE ANVIL, HOWEVER, THE SURGEON MET SIGNIFICANT RESISTANCE ON THE HANDLE OF THE STAPLER. THE ANVIL COULD ONLY PULL BACK HALFWAY WHICH MADE IT IMPOSSIBLE TO PULL OUT. THE ANVIL WAS MOBILE FORWARD, BUT NOT BACKWARDS. WHEN THE SALES REPRESENTATIVE CAME TO THE OPERATING ROOM, THE DEVICE WAS SUCCESSFULLY REMOVED FROM THE PATIENT AFTER TROUBLESHOOTING THE ISSUE. THERE WAS NO PATIENT INJURY.
05-10-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE DEVICE WAS SUCCESSFULLY FIRED ACROSS THE RECTUM, IT WAS ATTEMPTED TO RETRACT THE ANVIL, HOWEVER, THE SURGEON MET SIGNIFICANT RESISTANCE ON THE HANDLE OF THE STAPLER. THE ANVIL COULD ONLY PULL BACK HALFWAY WHICH MADE IT IMPOSSIBLE TO PULL OUT. THE ANVIL WAS MOBILE FORWARD, BUT NOT BACKWARDS. WHEN THE SALES REPRESENTATIVE CAME TO THE OPERATING ROOM, THE DEVICE WAS SUCCESSFULLY REMOVED FROM THE PATIENT AFTER TROUBLESHOOTING THE ISSUE. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ECHELON 3000 45 STANDARD | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON WOULD NOT USE THE STAPLER BECAUSE IT ¿MISFIRED¿. HE STATED THE JAWS WOULDN¿T OPEN WHEN HE TRIED TO REPOSITION THE STAPLER AT FIRST. THEN WHEN HE GOT IT OFF, THE STAPLER FIRED ON ITS OWN. HE DID NOT PULL THE FIRING TRIGGER. NURSE STATED THAT IT WAS SITTING IN HER OFFICE AND SHE COULD HEAR IT TRYING TO FIRE IN THE BOX AFTER THE EVENT ALL ON ITS OWN. PROCEDURE WAS COMPLETED WITH NO PATIENT HARM OR ANY OTHER ISSUES.
15-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ECHELON 3000 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON WOULD NOT USE THE STAPLER BECAUSE IT ¿MISFIRED¿. HE STATED THE JAWS WOULDN¿T OPEN WHEN HE TRIED TO REPOSITION THE STAPLER AT FIRST. THEN WHEN HE GOT IT OFF, THE STAPLER FIRED ON ITS OWN. HE DID NOT PULL THE FIRING TRIGGER. NURSE STATED THAT IT WAS SITTING IN HER OFFICE AND SHE COULD HEAR IT TRYING TO FIRE IN THE BOX AFTER THE EVENT ALL ON ITS OWN. PROCEDURE WAS COMPLETED WITH NO PATIENT HARM OR ANY OTHER ISSUES.
15-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON WOULD NOT USE THE STAPLER BECAUSE IT ¿MISFIRED¿. HE STATED THE JAWS WOULDN¿T OPEN WHEN HE TRIED TO REPOSITION THE STAPLER AT FIRST. THEN WHEN HE GOT IT OFF, THE STAPLER FIRED ON ITS OWN. HE DID NOT PULL THE FIRING TRIGGER. NURSE STATED THAT IT WAS SITTING IN HER OFFICE AND SHE COULD HEAR IT TRYING TO FIRE IN THE BOX AFTER THE EVENT ALL ON ITS OWN. PROCEDURE WAS COMPLETED WITH NO PATIENT HARM OR ANY OTHER ISSUES.
13-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY COLON RESECTION (COLECTOMY), DURING INTESTINAL RESECTION, THE GREEN LIGHT DID NOT ILL UMINATE AND THE DEVICE WAS UNABLE TO ENTER FIRING MODE WITH THE THREE RELOADS. A NEW HANDLE, SHELL AND ADAPTER WERE USED WITH THE SAME RELOADS TO RESOLVE THE ISSUE. NO PATIENT INJURY.
13-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | TRI-STAPLE 2.0 | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY COLON RESECTION (COLECTOMY), DURING INTESTINAL RESECTION, THE GREEN LIGHT DID NOT ILL UMINATE AND THE DEVICE WAS UNABLE TO ENTER FIRING MODE WITH THE THREE RELOADS. A NEW HANDLE, SHELL AND ADAPTER WERE USED WITH THE SAME RELOADS TO RESOLVE THE ISSUE. NO PATIENT INJURY.
28-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY SEGMENTECTOMY, THE 60MM RELOAD WAS CONNECTED INTO THE HANDLE, POSITIONED THE DEVICE IN THE LUNG AND THE DEVICE MADE A SNAP AND DID NOT STAPLE. THE USER WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW RELOAD WAS CONNECTED BUT STILL DID NOT WORK. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY SEGMENTECTOMY, THE 60MM RELOAD WAS CONNECTED INTO THE HANDLE, POSITIONED THE DEVICE IN THE LUNG AND THE DEVICE MADE A SNAP AND DID NOT STAPLE. THE USER WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW RELOAD WAS CONNECTED BUT STILL DID NOT WORK. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE IS NO PATIENT INJURY.
28-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY SEGMENTECTOMY, THE 60MM RELOAD WAS CONNECTED INTO THE HANDLE, POSITIONED THE DEVICE IN THE LUNG AND THE DEVICE MADE A SNAP AND DID NOT STAPLE. THE USER WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. A NEW RELOAD WAS CONNECTED BUT STILL DID NOT WORK. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE IS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon | 11-01-2024 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP FIXATION DEVICE | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2023 AND ABSORBABLE STRAPS WERE USED. THE DEVICE JAMMED DURING THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION WAS PROVIDED.
02-11-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ECHELON 3000 60MM STANDARD | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER WAS FIRED WITH THE FIRST RELOAD, BUT THE SECOND RELOAD WOULD NOT FIT INTO THE JAWS OF THE STAPLER FOR THE SECOND FIRING. THERE WAS A DELAY AND A SECOND DEVICE WAS USED TO COMPLETE CASE. NO PATIENT HARM.
05-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC COLON CLIPPING PROCEDURE CONVERTED TO OPEN UNRELATED TO DEVICE PROBLEM, THERE WAS A VERY THICK TISSUE. THE DEVICE DID NOT FIRE. AT THE MOMENT OF STAPLING, THE STAPLER WAS FORCED AND THE HANDLE BROKE OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC COLON CLIPPING PROCEDURE CONVERTED TO OPEN UNRELATED TO DEVICE PROBLEM, THERE WAS A VERY THICK TISSUE. THE DEVICE DID NOT FIRE. AT THE MOMENT OF STAPLING, THE STAPLER WAS FORCED AND THE HANDLE BROKE OFF. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
08-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | EEA | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN DISCOID RESECTION PROCEDURE, WHEN REMOVING THE CIRCULAR STAPLER AFTER STAPLING AND PERFORMING THE ANASTOMOSIS, THE ANVIL TILTED AND GOT STUCK ABOVE THE ANASTOMOSIS. THERE WAS DIFFICULTY REMOVING THE CIRCULAR STAPLER FROM THE CAVITY, AND IT WAS NOTED THAT THE ANVIL WAS BENT. TO RESOLVE THE ISSUE AND COMPLETE THE CASE, STITCHES WERE APPLIED AT THE ANASTOMOSIS.
08-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | EEA | Material Twisted/Bent | A040609 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN DISCOID RESECTION PROCEDURE, WHEN REMOVING THE CIRCULAR STAPLER AFTER STAPLING AND PERFORMING THE ANASTOMOSIS, THE ANVIL TILTED AND GOT STUCK ABOVE THE ANASTOMOSIS. THERE WAS DIFFICULTY REMOVING THE CIRCULAR STAPLER FROM THE CAVITY, AND IT WAS NOTED THAT THE ANVIL WAS BENT. TO RESOLVE THE ISSUE AND COMPLETE THE CASE, STITCHES WERE APPLIED AT THE ANASTOMOSIS.
12-11-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Arthritis | Event Description: IT WAS REPORTED THAT POST OP A GASTRIC BYPASS, THE PATIENT EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY. THE PATIENT HAD REPORTED BACK PAIN, OA KNEES, EPISODE OF URINE RETENTION (8 MONTHS AGO), REFLUX OCCASIONAL (HIATUS HERNIA FROM PREVIOUS GASTRIC SURGERY), AND CHRONIC CONSTIPATION.
12-11-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Heartburn/Indigestion | Event Description: IT WAS REPORTED THAT POST OP A GASTRIC BYPASS, THE PATIENT EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY. THE PATIENT HAD REPORTED BACK PAIN, OA KNEES, EPISODE OF URINE RETENTION (8 MONTHS AGO), REFLUX OCCASIONAL (HIATUS HERNIA FROM PREVIOUS GASTRIC SURGERY), AND CHRONIC CONSTIPATION.
12-11-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST OP A GASTRIC BYPASS, THE PATIENT EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY. THE PATIENT HAD REPORTED BACK PAIN, OA KNEES, EPISODE OF URINE RETENTION (8 MONTHS AGO), REFLUX OCCASIONAL (HIATUS HERNIA FROM PREVIOUS GASTRIC SURGERY), AND CHRONIC CONSTIPATION.
12-11-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: IT WAS REPORTED THAT POST OP A GASTRIC BYPASS, THE PATIENT EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY. THE PATIENT HAD REPORTED BACK PAIN, OA KNEES, EPISODE OF URINE RETENTION (8 MONTHS AGO), REFLUX OCCASIONAL (HIATUS HERNIA FROM PREVIOUS GASTRIC SURGERY), AND CHRONIC CONSTIPATION.
12-11-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Urinary Retention | Event Description: IT WAS REPORTED THAT POST OP A GASTRIC BYPASS, THE PATIENT EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY. THE PATIENT HAD REPORTED BACK PAIN, OA KNEES, EPISODE OF URINE RETENTION (8 MONTHS AGO), REFLUX OCCASIONAL (HIATUS HERNIA FROM PREVIOUS GASTRIC SURGERY), AND CHRONIC CONSTIPATION.
12-11-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hernia | Event Description: IT WAS REPORTED THAT POST OP A GASTRIC BYPASS, THE PATIENT EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY. THE PATIENT HAD REPORTED BACK PAIN, OA KNEES, EPISODE OF URINE RETENTION (8 MONTHS AGO), REFLUX OCCASIONAL (HIATUS HERNIA FROM PREVIOUS GASTRIC SURGERY), AND CHRONIC CONSTIPATION.
12-11-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Obstruction/Occlusion | Event Description: IT WAS REPORTED THAT POST OP A GASTRIC BYPASS, THE PATIENT EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY. THE PATIENT HAD REPORTED BACK PAIN, OA KNEES, EPISODE OF URINE RETENTION (8 MONTHS AGO), REFLUX OCCASIONAL (HIATUS HERNIA FROM PREVIOUS GASTRIC SURGERY), AND CHRONIC CONSTIPATION.
01-11-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | 75MM SELECTABLE NEW TLC | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE ANVIL HALF ATTACHMENTS ARE DISPLACED. THERE WERE NO PATIENT CONSEQUENCES.
19-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN OPENING THE FRONT HANDLE PART AND BACK HANDLE PART OF THE POWER SHELL FOR HANDLE IN SERTION, THE FRONT/BACK PART WERE FOUND TO BE SEPARATED. IT WAS NOTICED THAT THE CONNECTION REGION OF THE SHELL WAS DISCONNECTED WHEN IT WAS REMOVED. A NEW SHELL WAS TAKEN OUT TO FIX THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
20-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG RESECTION, THE RELOAD DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE ASSOCIATED HANDLE WAS ABLE TO FIRE. THERE WAS NO PATIENT INJURY.
01-11-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | 55MM SELECTABLE NEW TLC | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE ANVIL HALF ATTACHMENTS ARE DISPLACED. THERE WERE NO PATIENT CONSEQUENCES.
01-01-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE PROCEDURE, THERE WAS LEAK FOLLOWING THE USAGE OF THE DEVICE. NO FURTHER DETAILS WERE PROVIDED. NO PATIENT CONSEQUENCES REPORTED.
20-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC THORASCOPIC PULMONARY WEDGE RESECTION, DURING THE REMOVAL OF TARGET LUNG TISSUE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. BUT THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LOBECTOMY, IN THE LUNG TISSUE, THE LOAD WOULD NOT CUT THE TISSUE AND FIRE STAPLES. THE SURGEON TRIED TO APPLY EXTRA STRENGTH. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE HANDLE MADE A STRANGE NOISE AS IT WAS SQUEEZED WHICH MADE THE SURGEON SUSPECT THAT IT WAS BROKEN. THE SURGEON CHANGED BOTH THE RELOAD AND THE HANDLE THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AFTER THE PULMONARY PARENCHYMA WAS CLAMPED, THE DEVICE FIRED, BUT IT STOPPED MOVING MIDWAY. HENCE, NO STAPLES ARE DEPLOYED ON THE DISTAL PART. THE JAWS ALSO LOCKED ON TISSUE. THE BLADE WORKED BY REBOOTING THE DEVICE. TO COMPLETE THE CASE, ANOTHER RELOAD WAS USED. IT WAS NOTED THAT THE HANDLE AND ADAPTER WORKED WITH OTHER DEVICES. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AFTER THE PULMONARY PARENCHYMA WAS CLAMPED, THE DEVICE FIRED, BUT IT STOPPED MOVING MIDWAY. HENCE, NO STAPLES ARE DEPLOYED ON THE DISTAL PART. THE JAWS ALSO LOCKED ON TISSUE. THE BLADE WORKED BY REBOOTING THE DEVICE. TO COMPLETE THE CASE, ANOTHER RELOAD WAS USED. IT WAS NOTED THAT THE HANDLE AND ADAPTER WORKED WITH OTHER DEVICES. THERE WAS NO PATIENT INJURY.
28-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AT THE VERY EARLY STAGE, THE STAPLER AND RELOAD WOULD NOT CLOSE FULLY ON THE INTENDED TISSUE. TO RESOLVE THE ISSUE, A SECOND HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AFTER FIRING TWO BLACK REINFORCED RELOADS, THE SURGEON CHANGED IT TO THE PURPLE RELOAD. WHEN FIRING, THE BLADE STOPPED AND DISPLAYED AN ERROR MESSAGE WHICH SHOWED THE JAWS WITH THE SIGN OF OPENING IT. THE SURGEON TRIED IT WITH A NEW PURPLE RELOAD AGAIN AND HAD THE SAME ISSUE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE AND NO OTHER STEPS WERE TAKEN TO REMOVED IT. THE KNIFE WAS ABLE TO BE RETRACTED WITHOUT OPENING IT MANUALLY. THE SURGEON END UP USING A THIRD BLACK NON REINFORCED RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE PROCEDURE, AFTER FIRING TWO BLACK REINFORCED RELOADS, THE SURGEON CHANGED IT TO THE PURPLE RELOAD. WHEN FIRING, THE BLADE STOPPED AND DISPLAYED AN ERROR MESSAGE WHICH SHOWED THE JAWS WITH THE SIGN OF OPENING IT. THE SURGEON TRIED IT WITH A NEW PURPLE RELOAD AGAIN AND HAD THE SAME ISSUE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE AND NO OTHER STEPS WERE TAKEN TO REMOVED IT. THE KNIFE WAS ABLE TO BE RETRACTED WITHOUT OPENING IT MANUALLY. THE SURGEON END UP USING A THIRD BLACK NON REINFORCED RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Malfunction | intuitive surgical | 11-01-2024 | SUREFORM | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE SUREFORM 60 STAPLER HAD STAPLES HUNG UP IN TISSUE IN THE 4TH STAPLE BLUE RELOAD. THE PROCEDURE WAS COMPLETED AS PLANNED WITH NO REPORTED INJURY.
07-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SMALL ATYPICAL RESECTION TO STAPLE THE LUNG PARENCHYMA, HE 1ST STAPLES HAVE STARTED TO COME OUT, BUT THEN THE DEVICE JAMMED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-12-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AFTER A GASTRIC BYPASS, EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY, POST OP. REPORTED PATIENT CONSEQUENCES WERE BACK PAIN AND DEPRESSION.
06-12-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: IT WAS REPORTED THAT AFTER A GASTRIC BYPASS, EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY, POST OP. REPORTED PATIENT CONSEQUENCES WERE BACK PAIN AND DEPRESSION.
06-12-2023 | Injury | ethicon endo-surgery | 11-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Depression | Event Description: IT WAS REPORTED THAT AFTER A GASTRIC BYPASS, EXPERIENCED STAPLE LINE BLEEDING AT THE GASTROJEJUNOSTOMY, POST OP. REPORTED PATIENT CONSEQUENCES WERE BACK PAIN AND DEPRESSION.
27-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC SURGERY, WHEN CUTTING PULMONARY LOBE TISSUE, THERE WAS POOR STAPLE FORMATION. THE STAPLES BURST OUT AND THE TISSUE AT THE CUTTING SITE SPLIT. WHEN USING ANOTHER RELOAD TO CUT THE LUNG TISSUE, THE STAPLES IN THE RELOAD DID NOT COME OUT. ANOTHER BRAND OF DEVICE WAS USED TO RESOLVE THE ISSUE.
13-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRECTOMY PROCEDURE, WHEN THE JAW OF THE DEVICE WAS CLOSED, THE JAW OPENING WAS LARGER THAN USUAL. HENCE, FIRING WAS NOT PERFORMED. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC SURGERY, WHEN CUTTING PULMONARY LOBE TISSUE, THERE WAS POOR STAPLE FORMATION. THE STAPLES BURST OUT AND THE TISSUE AT THE CUTTING SITE SPLIT. WHEN USING ANOTHER RELOAD TO CUT THE LUNG TISSUE, THE STAPLES IN THE RELOAD DID NOT COME OUT. ANOTHER BRAND OF DEVICE WAS USED TO RESOLVE THE ISSUE.
27-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC SURGERY, WHEN CUTTING PULMONARY LOBE TISSUE, THERE WAS POOR STAPLE FORMATION. THE STAPLES BURST OUT AND THE TISSUE AT THE CUTTING SITE SPLIT. WHEN USING ANOTHER RELOAD TO CUT THE LUNG TISSUE, THE STAPLES IN THE RELOAD DID NOT COME OUT. ANOTHER BRAND OF DEVICE WAS USED TO RESOLVE THE ISSUE.
01-01-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, WHEN REMOVING THE DEVICE FROM THE KIT BE USED IN SURGERY, IT DIDN¿T BRING THE MANUAL OVERRIDE STOPPER, AND WHEN SURGEONS TRIED TO GIVE THE FIRST FIRE AFTER ASSEMBLING WITH A RELOAD, IT DIDN¿T WORK. THEY HAD TO OPEN ANOTHER DEVICE, AND USING THE SAME RELOAD IT WORKED NORMALLY. THE SURGERY WAS COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | POWERED PLUS STANDARD 45 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, WHEN REMOVING THE DEVICE FROM THE KIT BE USED IN SURGERY, IT DIDN¿T BRING THE MANUAL OVERRIDE STOPPER, AND WHEN SURGEONS TRIED TO GIVE THE FIRST FIRE AFTER ASSEMBLING WITH A RELOAD, IT DIDN¿T WORK. THEY HAD TO OPEN ANOTHER DEVICE, AND USING THE SAME RELOAD IT WORKED NORMALLY. THE SURGERY WAS COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES REPORTED.
07-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SMALL ATYPICAL RESECTION TO STAPLE THE LUNG PARENCHYMA, THE 1ST STAPLES HAVE STARTED TO COME OUT, BUT THEN THE DEVICE JAMMED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACIC SURGERY, WHILE ON LUNG PARENCHYMA RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACIC SURGERY, WHILE ON LUNG PARENCHYMA RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACIC SURGERY, WHILE ON LUNG PARENCHYMA RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACIC SURGERY, WHILE ON LUNG PARENCHYMA RESECTION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE DEVICE LOCKED ON TISSUE. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ENDO ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, THE DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AFTER THE PULMONARY PARENCHYMA WAS CLAMPED, THE DEVICE FIRED, BUT IT STOPPED MOVING MIDWAY. HENCE, NO STAPLES ARE DEPLOYED ON THE DISTAL PART. THE JAWS ALSO LOCKED ON TISSUE. THE BLADE WORKED BY REBOOTING THE DEVICE. TO COMPLETE THE CASE, ANOTHER RELOAD WAS USED. IT WAS NOTED THAT THE HANDLE AND ADAPTER WORKED WITH OTHER DEVICES. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, AFTER THE PULMONARY PARENCHYMA WAS CLAMPED, THE DEVICE FIRED, BUT IT STOPPED MOVING MIDWAY. HENCE, NO STAPLES ARE DEPLOYED ON THE DISTAL PART. THE JAWS ALSO LOCKED ON TISSUE. THE BLADE WORKED BY REBOOTING THE DEVICE. TO COMPLETE THE CASE, ANOTHER RELOAD WAS USED. IT WAS NOTED THAT THE HANDLE AND ADAPTER WORKED WITH OTHER DEVICES. THERE WAS NO PATIENT INJURY.
18-10-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-11-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | 75MM SELECTABLE NEW TLC | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON FOUND THE CARTRIDGE WING OF SR75 EASILY TO BE TORN, AND THEY FELL OFF FROM THE CARTRIDGE. THE SURGEON FOUND HE WAS NOT ABLE TO REMOVE THE REMAINING SR75 FROM NTLC75. HE OPENED ANOTHER NEW NTLC75 AND DECIDED TO OPEN ANOTHER SR75. TWO FRAGMENTS WERE GENERATED (BOTH CARTRIDGE WINGS). NO PATIENT CONSEQUENCE REPORTED.
15-12-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | PROXIMATE*ILS CURVED CIR STAPL | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-08-2023 | Injury | us surgical puerto rico | 11-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
04-08-2023 | Injury | us surgical puerto rico | 11-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
04-08-2023 | Injury | us surgical puerto rico | 11-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
04-08-2023 | Injury | us surgical puerto rico | 11-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
15-12-2023 | Malfunction | intuitive surgical | 11-01-2024 | ENDOWRIST | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SURGICAL PROCEDURE, THE CUSTOMER ENCOUNTERED A PARTIAL FIRE WITH THE CURVED-TIP STAPLER 30 INSTRUMENT. A GREEN RELOAD WAS INSTALLED INSIDE THE INSTRUMENT. ERROR MESSAGE DISPLAYED WAS ¿BLADE MAY BE EXPOSED¿. THE PROCEDURE WAS COMPLETED WITH NO REPORTED INJURY. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION: THE INSTRUMENT WAS INSPECTED PRIOR TO USE AND NO DAMAGE WAS NOTED. THE ISSUE OCCURRED ON THE SECOND STAPLER FIRE. THE SURGICAL TASK BEING PERFORMED WAS STAPLING. THE TARGET TISSUE WAS AN ARTERY. THE TISSUE WAS NOT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY. THE TISSUE WAS NOT CALCIFIED. THERE WAS NO TISSUE TENSION AND NO TISSUE BUNCHING. THE CUSTOMER CONFIRMED THE ISSUE WAS FAILURE TO COMPLETE THE FIRING CYCLE AND THE INSTRUMENT PARTIALLY FIRED. THE BLADE WAS EXPOSED. THE RELOAD WAS NOT STUCK TO TISSUE. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. THERE WAS NO BUTTRESS MATERIAL USED. THERE WERE NO CLAMPING ISSUES PRIOR TO FIRING. THERE WAS NO OBSERVED BLEEDING. THERE WAS NO UNPLANNED TISSUE REMOVAL. THE BACKUP STAPLER HAD THE SAME ISSUES. THEY WERE UNABLE TO TELL IF STAPLES WERE MISSING FROM THE STAPLE LINE THAT FIRED DUE TO THE MISFIRE. THE PROCEDURE WAS COMPLETED ROBOTICALLY USING A THIRD STAPLER.
04-08-2023 | Injury | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
04-08-2023 | Injury | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
04-08-2023 | Injury | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
04-08-2023 | Injury | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
04-08-2023 | Injury | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
20-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC WEDGE RESECTION, DURING FIRING OF THE LUNG PARENCHYMA, THE RELOAD WAS LOADED, THE GREEN BUTTON WAS ACTIVATED BUT THE DEVICE DID NOT FIRE AFTER THE FIRST SQUEEZED AND UNABLE TO FIRE ALL FOUR SQUEEZED. TWO ADDITIONAL RELOADS WERE USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC COLON CLIPPING PROCEDURE CONVERTED TO OPEN UNRELATED TO DEVICE PROBLEM, THERE WAS A VERY THICK TISSUE. THE DEVICE DID NOT FIRE. AT THE MOMENT OF STAPLING, THE STAPLER WAS FORCED AND THE HANDLE BROKE OFF. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC COLON CLIPPING PROCEDURE CONVERTED TO OPEN UNRELATED TO DEVICE PROBLEM, THERE WAS A VERY THICK TISSUE. THE DEVICE DID NOT FIRE. AT THE MOMENT OF STAPLING, THE STAPLER WAS FORCED AND THE HANDLE BROKE OFF. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
15-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANASTOMOSIS ENTER ENTERO PROCEDURE, THE STAPLER DID NOT FIRE AND NO SUTURE LINE WAS FORMED, BUT THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
15-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANASTOMOSIS ENTER ENTERO PROCEDURE, THE STAPLER DID NOT FIRE AND NO SUTURE LINE WAS FORMED, BUT THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE. NO PATIENT INJURY.
15-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANASTOMOSIS ENTER ENTERO PROCEDURE, THE STAPLER DID NOT FIRE, BUT THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE.
15-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANASTOMOSIS ENTER ENTERO PROCEDURE, THE STAPLER DID NOT FIRE, BUT THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE.
30-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANASTOMOSIS ENTER ENTERO BY PASS PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE.
30-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ANASTOMOSIS ENTER ENTERO BY PASS PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE.
20-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRIC SLEEVE, WHILE DETACHING GREATER CURVATURE OF STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRIC SLEEVE, WHILE DETACHING GREATER CURVATURE OF STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-11-2023 | Injury | covidien lp llc north haven | 11-01-2024 | ENDO GIA | Break | A0401 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC COLECTOMY, DURING INTESTINAL RESECTION, AFTER WOUND OR SURGICAL INCISION CLOSURE, A YELLOW PIECE OF DEBRIS WAS FOUND ON THE SURGICAL DRAPE. WHEN THE YELLOW TABS OF THE FOUR CARTRIDGES USED IN THIS CASE WERE CHECKED, THERE WERE TWO RELOADS WITH PROTRUDING PARTS THAT WERE MISSING, AND ONE RELOAD WHERE IT SEEMED THAT IT MIGHT HAVE BEEN SCRAPED. THERE WERE 4 RELOADS WITH YELLOW TABS DAMAGED; HOWEVER, NO SCRAPED NOR CHIPPED FRAGMENTS WERE FOUND. TO SEARCH FOR THE YELLOW TAB'S FRAGMENT, AFTER WOUND OR SURGICAL INCISION CLOSURE, LAPAROSCOPIC LAPAROTOMY WAS PERFORMED AGAIN. THE PROCEDURE WAS EXTENDED BY ONE HOUR. A FAIRLY SMALL FRAGMENT OF ABOUT 1MM WAS FOUND ON TOP OF THE SURGICAL DRAPE, BUT NO FURTHER FRAGMENT WAS FOUND INSIDE THE CAVITY.
27-11-2023 | Injury | covidien lp llc north haven | 11-01-2024 | ENDO GIA | Break | A0401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC COLECTOMY, DURING INTESTINAL RESECTION, AFTER WOUND OR SURGICAL INCISION CLOSURE, A YELLOW PIECE OF DEBRIS WAS FOUND ON THE SURGICAL DRAPE. WHEN THE YELLOW TABS OF THE FOUR CARTRIDGES USED IN THIS CASE WERE CHECKED, THERE WERE TWO RELOADS WITH PROTRUDING PARTS THAT WERE MISSING, AND ONE RELOAD WHERE IT SEEMED THAT IT MIGHT HAVE BEEN SCRAPED. THERE WERE 4 RELOADS WITH YELLOW TABS DAMAGED; HOWEVER, NO SCRAPED NOR CHIPPED FRAGMENTS WERE FOUND. TO SEARCH FOR THE YELLOW TAB'S FRAGMENT, AFTER WOUND OR SURGICAL INCISION CLOSURE, LAPAROSCOPIC LAPAROTOMY WAS PERFORMED AGAIN. THE PROCEDURE WAS EXTENDED BY ONE HOUR. A FAIRLY SMALL FRAGMENT OF ABOUT 1MM WAS FOUND ON TOP OF THE SURGICAL DRAPE, BUT NO FURTHER FRAGMENT WAS FOUND INSIDE THE CAVITY.
20-12-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC GASTRIC BYPASS, DURING THE INITIAL STAPLING TO CREATE THE GASTRIC POUCH ON THE STOMACH, TISSUE BECAME STUCK TO THE CARTRIDGE WHEN IT OPENED AFTER FIRING. THE SURGEON HAD TO USE A GRASPER TO REMOVE THE TISSUE FROM THE CARTRIDGE AND THE STAPLER. THERE WAS NO PATIENT INJURY.
11-01-2024 | Malfunction | covidien lp llc north haven | 11-01-2024 | UNKNOWN ENDO GIA SULU | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PNEUMONECTOMY, THERE WERE NO STAPLES PRESENT AFTER FIRING.
01-01-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE SURGEON ATTEMPTED TO FIRE THE DEVICE ACROSS APPENDIX BUT DID NOT FIRE. DUMPED UNFORMED AND PARTIALLY FORMED STAPLES INTO FIELD. NEW DEVICE OPENED TO COMPLETE PROCEDURE. PROCEDURE WAS COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 11-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE SURGEON ATTEMPTED TO FIRE THE DEVICE ACROSS APPENDIX BUT DID NOT FIRE. DUMPED UNFORMED AND PARTIALLY FORMED STAPLES INTO FIELD. NEW DEVICE OPENED TO COMPLETE PROCEDURE. PROCEDURE WAS COMPLETED SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
11-01-2024 | Malfunction | covidien lp llc north haven | 11-01-2024 | GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SEGMENTAL RESECTION OF THE SMALL INTESTINE, WHEN CUTTING THE SMALL INTESTINE WITH THE STAPLER, AFTER THE STAPLER WAS FIRED, ONLY ONE SIDE WAS CLOSED WITH THE STAPLES. THE DEVICE WAS FIRED DOWN COMPLETELY, BUT SOME STAPLES WERE MISSING FROM THE STAPLE LINE. IT WAS NOTED THAT IT ALSO HAPPENED ON THE RELOAD UNITS, AND THERE WAS NO CORRECT CLOSURE OF BOTH SIDES WITH STAPLES OR DID NOT FORM PROPER SHAPE. TO RESOLVE THE ISSUE, THE SMALL INTESTINE WAS CLOSED WITH A CLAMP, AND THE DESIRED MANUAL ANASTOMOSIS WAS PERFORMED. THERE WAS NO PATIENT INJURY.
11-01-2024 | Malfunction | covidien lp llc north haven | 11-01-2024 | GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SEGMENTAL RESECTION OF THE SMALL INTESTINE, WHEN CUTTING THE SMALL INTESTINE WITH THE STAPLER, AFTER THE STAPLER WAS FIRED, ONLY ONE SIDE WAS CLOSED WITH THE STAPLES. THE DEVICE WAS FIRED DOWN COMPLETELY, BUT SOME STAPLES WERE MISSING FROM THE STAPLE LINE. IT WAS NOTED THAT IT ALSO HAPPENED ON THE RELOAD UNITS, AND THERE WAS NO CORRECT CLOSURE OF BOTH SIDES WITH STAPLES OR DID NOT FORM PROPER SHAPE. TO RESOLVE THE ISSUE, THE SMALL INTESTINE WAS CLOSED WITH A CLAMP, AND THE DESIRED MANUAL ANASTOMOSIS WAS PERFORMED. THERE WAS NO PATIENT INJURY.
18-08-2023 | Injury | us surgical puerto rico | 11-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO LITERATURE SOURCE, RETROSPECTIVE STUDY EVALUATED THE OUTCOMES OF PATIENTS WHO UNDERWENT TOTAL GASTRECTOMY WITH CURATIVE INTENT FOR GASTRIC CANCER BETWEEN 2017 AND 2021. CIRCULAR STAPLER OR ANOTHER COMPETITOR CIRCULAR STAPLER WAS USED TO PERFORM AN END TO SIDE ESOPHAGOJEJUNOSTOMY AND A COMPETITOR LINEAR STAPLER WERE USED TO PERFORM LINEAR ANASTOMOSIS. THE INTRAOPERATIVE ENDOSCOPY AND AIR LEAK TESTS WERE USED TO CONFIRM ANASTOMOTIC INTEGRITY. THERE WERE 99 PATIENTS IN THE STUDY AND POSTOPERATIVE COMPLICATIONS INCLUDED: 6 ANASTOMOTIC LEAKS AND 8 INFECTIONS. REINTERVENTIONS INCLUDED 10 ESOPHAGOGASTRODUODENOSCOPY AND 14 COMPUTED TOMOGRAPHY GUIDED DRAINAGE OF INTRAABDOMINAL FLUID COLLECTION. THERE WERE 14 UNPLANNED ICU STAY, 15 EXTENDED HOSPITALIZATIONS AND 7 READMISSIONS WITHIN 30-DAYS THAT WERE ALSO REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED: PULMONARY, CARDIAC, RENAL, HEPATIC, THROMBOEMBOLIC COMPLICATIONS AND PNEUMONIA.
18-08-2023 | Injury | us surgical puerto rico | 11-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO LITERATURE SOURCE, RETROSPECTIVE STUDY EVALUATED THE OUTCOMES OF PATIENTS WHO UNDERWENT TOTAL GASTRECTOMY WITH CURATIVE INTENT FOR GASTRIC CANCER BETWEEN 2017 AND 2021. CIRCULAR STAPLER OR ANOTHER COMPETITOR CIRCULAR STAPLER WAS USED TO PERFORM AN END TO SIDE ESOPHAGOJEJUNOSTOMY AND A COMPETITOR LINEAR STAPLER WERE USED TO PERFORM LINEAR ANASTOMOSIS. THE INTRAOPERATIVE ENDOSCOPY AND AIR LEAK TESTS WERE USED TO CONFIRM ANASTOMOTIC INTEGRITY. THERE WERE 99 PATIENTS IN THE STUDY AND POSTOPERATIVE COMPLICATIONS INCLUDED: 6 ANASTOMOTIC LEAKS AND 8 INFECTIONS. REINTERVENTIONS INCLUDED 10 ESOPHAGOGASTRODUODENOSCOPY AND 14 COMPUTED TOMOGRAPHY GUIDED DRAINAGE OF INTRAABDOMINAL FLUID COLLECTION. THERE WERE 14 UNPLANNED ICU STAY, 15 EXTENDED HOSPITALIZATIONS AND 7 READMISSIONS WITHIN 30-DAYS THAT WERE ALSO REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED: PULMONARY, CARDIAC, RENAL, HEPATIC, THROMBOEMBOLIC COMPLICATIONS AND PNEUMONIA.
18-08-2023 | Injury | us surgical puerto rico | 11-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: ACCORDING TO LITERATURE SOURCE, RETROSPECTIVE STUDY EVALUATED THE OUTCOMES OF PATIENTS WHO UNDERWENT TOTAL GASTRECTOMY WITH CURATIVE INTENT FOR GASTRIC CANCER BETWEEN 2017 AND 2021. CIRCULAR STAPLER OR ANOTHER COMPETITOR CIRCULAR STAPLER WAS USED TO PERFORM AN END TO SIDE ESOPHAGOJEJUNOSTOMY AND A COMPETITOR LINEAR STAPLER WERE USED TO PERFORM LINEAR ANASTOMOSIS. THE INTRAOPERATIVE ENDOSCOPY AND AIR LEAK TESTS WERE USED TO CONFIRM ANASTOMOTIC INTEGRITY. THERE WERE 99 PATIENTS IN THE STUDY AND POSTOPERATIVE COMPLICATIONS INCLUDED: 6 ANASTOMOTIC LEAKS AND 8 INFECTIONS. REINTERVENTIONS INCLUDED 10 ESOPHAGOGASTRODUODENOSCOPY AND 14 COMPUTED TOMOGRAPHY GUIDED DRAINAGE OF INTRAABDOMINAL FLUID COLLECTION. THERE WERE 14 UNPLANNED ICU STAY, 15 EXTENDED HOSPITALIZATIONS AND 7 READMISSIONS WITHIN 30-DAYS THAT WERE ALSO REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED: PULMONARY, CARDIAC, RENAL, HEPATIC, THROMBOEMBOLIC COMPLICATIONS AND PNEUMONIA.
18-08-2023 | Injury | us surgical puerto rico | 11-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO LITERATURE SOURCE, RETROSPECTIVE STUDY EVALUATED THE OUTCOMES OF PATIENTS WHO UNDERWENT TOTAL GASTRECTOMY WITH CURATIVE INTENT FOR GASTRIC CANCER BETWEEN 2017 AND 2021. CIRCULAR STAPLER OR ANOTHER COMPETITOR CIRCULAR STAPLER WAS USED TO PERFORM AN END TO SIDE ESOPHAGOJEJUNOSTOMY AND A COMPETITOR LINEAR STAPLER WERE USED TO PERFORM LINEAR ANASTOMOSIS. THE INTRAOPERATIVE ENDOSCOPY AND AIR LEAK TESTS WERE USED TO CONFIRM ANASTOMOTIC INTEGRITY. THERE WERE 99 PATIENTS IN THE STUDY AND POSTOPERATIVE COMPLICATIONS INCLUDED: 6 ANASTOMOTIC LEAKS AND 8 INFECTIONS. REINTERVENTIONS INCLUDED 10 ESOPHAGOGASTRODUODENOSCOPY AND 14 COMPUTED TOMOGRAPHY GUIDED DRAINAGE OF INTRAABDOMINAL FLUID COLLECTION. THERE WERE 14 UNPLANNED ICU STAY, 15 EXTENDED HOSPITALIZATIONS AND 7 READMISSIONS WITHIN 30-DAYS THAT WERE ALSO REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED: PULMONARY, CARDIAC, RENAL, HEPATIC, THROMBOEMBOLIC COMPLICATIONS AND PNEUMONIA.
04-08-2023 | Injury | davis & geck caribe | 11-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
04-08-2023 | Injury | davis & geck caribe | 11-01-2024 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2016 AND 2021, A RETROSPECTIVE STUDY EVALUATED THE COMPLICATION RATE OF MINIMAL INVASIVE ESOPHAGECTOMY IN PATIENTS WITH ESOPHAGUS OR GASTROESOPHAGEAL CANCERS. A TOTAL OF 120 PATIENTS WERE INCLUDED IN THE STUDY. THE MANUAL STAPLER WAS USED TO CONSTRUCT THE GASTRIC CONDUIT. THE ANVIL OF THE CIRCULAR STAPLER WAS PLACED ORALLY AND SECURED BY PURSE-STRING SUTURES USING ENDOSCOPIC SUTURING DEVICE. THE INTRATHORACIC ANASTOMOSIS WAS PERFORMED END-TO-SIDE USING THE CIRCULAR STAPLER INTRODUCED THROUGH THE GASTRIC CONDUIT TUBE. THE COMPLICATIONS INCLUDED STAPLE LINE OR ANASTOMOTIC LEAK AND CONDUIT NECROSIS. CONVERSION WAS REQUIRED IN THREE PATIENTS FOR ADDITIONAL SUTURING ON THE ANASTOMOSIS OR STAPLE LINE. ANASTOMOTIC LEAKAGE OCCURRED IN NINE PATIENTS AND WAS DIAGNOSED BY COMPUTED TOMOGRAPHY AND TREATED WITH ENDOSCOPIC STENT OR SURGICAL TREATMENT. PULMONARY, CARDIAC, THROMBOEMBOLIC, NEUROLOGIC AND INFECTION RELATED COMPLICATIONS WERE NOT RELATED TO THE DEVICE.
24-10-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, AT THE TIME OF ENTERAL WHOLE ANASTOMOSIS, METHYLENE BLUE WAS FILTERED INTO THE ENTERAL ANASTOMOSIS IN RELATION TO THE MECHANICAL SUTURE AND THEN THE LAPAROSCOPIC POINT PLACEMENT WAS PERFORMED. HOWEVER, THE DEVICE DID NOT FIRE. THE SURGEON USED MANUAL SUTURE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-10-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, AT THE TIME OF ENTERAL WHOLE ANASTOMOSIS, METHYLENE BLUE WAS FILTERED INTO THE ENTERAL ANASTOMOSIS IN RELATION TO THE MECHANICAL SUTURE AND THEN THE LAPAROSCOPIC POINT PLACEMENT WAS PERFORMED. HOWEVER, THE DEVICE DID NOT FIRE. THE SURGEON USED MANUAL SUTURE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-10-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS SURGERY, AT THE TIME OF ENTERAL WHOLE ANASTOMOSIS, METHYLENE BLUE WAS FILTERED INTO THE ENTERAL ANASTOMOSIS IN RELATION TO THE MECHANICAL SUTURE AND THEN THE LAPAROSCOPIC POINT PLACEMENT WAS PERFORMED. HOWEVER, THE DEVICE DID NOT FIRE. THIS LED TO BLEEDING. THE ISSUE OCCURRED IN TWO RELOADS. THE SURGEON USED A MANUAL SUTURE TO RESOLVE THE ISSUE.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS NOT ABLE TO FIRE. IT WAS MENTIONED THAT THERE WAS MODERATE BLEEDING. MANUAL SUTURING WAS PERFORMED TO RESOLVE THE ISSUE.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS NOT ABLE TO FIRE. IT WAS MENTIONED THAT THERE WAS MODERATE BLEEDING. MANUAL SUTURING WAS PERFORMED TO RESOLVE THE ISSUE.
18-10-2023 | Malfunction | covidien lp llc north haven | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS SURGERY, AT THE TIME OF ENTERAL WHOLE ANASTOMOSIS, METHYLENE BLUE WAS FILTERED INTO THE ENTERAL ANASTOMOSIS IN RELATION TO THE MECHANICAL SUTURE AND THEN THE LAPAROSCOPIC POINT PLACEMENT WAS PERFORMED. HOWEVER, THE DEVICE DID NOT FIRE AND THIS LED TO BLEEDING. THE ISSUE OCCURRED IN TWO RELOADS. THE SURGEON USED A MANUAL SUTURE TO RESOLVE THE ISSUE.
18-10-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS SURGERY, AT THE TIME OF ENTERAL WHOLE ANASTOMOSIS, METHYLENE BLUE WAS FILTERED INTO THE ENTERAL ANASTOMOSIS IN RELATION TO THE MECHANICAL SUTURE AND THEN THE LAPAROSCOPIC POINT PLACEMENT WAS PERFORMED. HOWEVER, THE DEVICE DID NOT FIRE AND THIS LED TO BLEEDING. THE ISSUE OCCURRED IN TWO RELOADS. THE SURGEON USED A MANUAL SUTURE TO RESOLVE THE ISSUE.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS NOT ABLE TO FIRE. IT WAS MENTIONED THAT THERE WAS MODERATE BLEEDING. MANUAL SUTURING WAS PERFORMED TO RESOLVE THE ISSUE.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | ENDO GIA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS NOT ABLE TO FIRE. IT WAS MENTIONED THAT THERE WAS MODERATE BLEEDING. MANUAL SUTURING WAS PERFORMED TO RESOLVE THE ISSUE.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS NOT ABLE TO FIRE. IT WAS MENTIONED THAT THERE WAS MODERATE BLEEDING. MANUAL SUTURING WAS PERFORMED TO RESOLVE THE ISSUE.
07-11-2023 | Malfunction | us surgical puerto rico | 11-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS NOT ABLE TO FIRE. IT WAS MENTIONED THAT THERE WAS MODERATE BLEEDING. MANUAL SUTURING WAS PERFORMED TO RESOLVE THE ISSUE.
18-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | GST GOLD RELOAD, 60MM, 6 ROW | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE KNIFE BLADE CAME OUT BEFORE FIRING FROM CARTRIDGE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC CHOLECYSTECTOMY - LAPAROSCOPIC STAPLER NOT FIRING (RELEASING SUTURES). NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
14-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 12-01-2024 | SORBAFIX ENHANCED | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNSPECIFIED PROCEDURE, THE SORBAFIX ENHANCED FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE EXPOSURE OF ALL THE FASTENERS MADE THE DEVICE UNUSABLE. THE PROVIDED VIDEO SHOWS THAT THE MANDREL FELL OUT OF THE CANNULA. THERE WAS NO PATIENT INJURY.
14-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 12-01-2024 | SORBAFIX ENHANCED | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNSPECIFIED PROCEDURE, THE SORBAFIX ENHANCED FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE EXPOSURE OF ALL THE FASTENERS MADE THE DEVICE UNUSABLE. THE PROVIDED VIDEO SHOWS THAT THE MANDREL FELL OUT OF THE CANNULA. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 12-01-2024 | SORBAFIX ENHANCED | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNSPECIFIED PROCEDURE, THE SORBAFIX ENHANCED FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE EXPOSURE OF ALL THE FASTENERS MADE THE DEVICE UNUSABLE. THE PROVIDED VIDEO SHOWS THAT THE MANDREL FELL OUT OF THE CANNULA. THERE WAS NO PATIENT INJURY.
12-12-2023 | Injury | intuitive surgical | 12-01-2024 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
18-08-2023 | Death | us surgical puerto rico | 12-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO LITERATURE SOURCE, RETROSPECTIVE STUDY EVALUATED THE OUTCOMES OF PATIENTS WHO UNDERWENT TOTAL GASTRECTOMY WITH CURATIVE INTENT FOR GASTRIC CANCER BETWEEN 2017 AND 2021. CIRCULAR STAPLER OR ANOTHER COMPETITOR CIRCULAR STAPLER WAS USED TO PERFORM AN END TO SIDE ESOPHAGOJEJUNOSTOMY AND A COMPETITOR LINEAR STAPLER WAS USED TO PERFORM LINEAR ANASTOMOSIS. THE INTRAOPERATIVE ENDOSCOPY AND AIR LEAK TESTS WERE USED TO CONFIRM ANASTOMOTIC INTEGRITY. THERE WERE 99 PATIENTS IN THE STUDY. THERE WERE TWO DEATHS DURING THE FIRST 30-DAYS AFTER SURGERY, BUT CAUSE OF DEATH WAS NOT SPECIFIED.
15-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER SIX TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY. NO PATIENT CONSEQUENCES REPORTED.
18-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ABDOMINAL WALL SCAR HERNIA REPAIR, WHEN ATTACHING THE RELOAD THAT HAS 10 FIRINGS TO THE HANDLE AND WHILE PRESSING THE RED FIRE BUTTON, THE BUTTON WAS HARDER TO PRESS THAN USUAL. THE USER MANAGED TO FORCEFULLY PUSH IT IN AND, FORTUNATELY, WAS ABLE TO USE IT NORMALLY. AFTER THAT, THE RELOAD WAS REPLACED, BUT IT WAS DIFFICULT TO RELOAD. IT WAS NOTED THAT THERE WAS DIFFICULTY PUSHING THE "PUSH TO RELOAD" BUTTON, AND THE LOCK AND UNLOCK BUTTON WAS DIFFICULT TO OPERATE AS WELL. FIRING WAS CONTINUED, AND AS EXPECTED, THE RED BUTTON WAS STILL HARD, AND THE DOCTOR HAD TO APPLY A LOT OF FORCE TO PUSH IT IN. THE OPERATION ITSELF WAS COMPLETED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | RELIATACK | Difficult or Delayed Activation | A150102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ABDOMINAL WALL SCAR HERNIA REPAIR, WHEN ATTACHING THE RELOAD THAT HAS 10 FIRINGS TO THE HANDLE AND WHILE PRESSING THE RED FIRE BUTTON, THE BUTTON WAS HARDER TO PRESS THAN USUAL. THE USER MANAGED TO FORCEFULLY PUSH IT IN AND, FORTUNATELY, WAS ABLE TO USE IT NORMALLY. AFTER THAT, THE RELOAD WAS REPLACED, BUT IT WAS DIFFICULT TO RELOAD. IT WAS NOTED THAT THERE WAS DIFFICULTY PUSHING THE "PUSH TO RELOAD" BUTTON, AND THE LOCK AND UNLOCK BUTTON WAS DIFFICULT TO OPERATE AS WELL. FIRING WAS CONTINUED, AND AS EXPECTED, THE RED BUTTON WAS STILL HARD, AND THE DOCTOR HAD TO APPLY A LOT OF FORCE TO PUSH IT IN. THE OPERATION ITSELF WAS COMPLETED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | RELIATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ABDOMINAL WALL SCAR HERNIA REPAIR, WHEN ATTACHING THE RELOAD THAT HAS 10 FIRINGS TO THE HANDLE AND WHILE PRESSING THE RED FIRE BUTTON, THE BUTTON WAS HARDER TO PRESS THAN USUAL. THE USER MANAGED TO FORCEFULLY PUSH IT IN AND, FORTUNATELY, WAS ABLE TO USE IT NORMALLY. AFTER THAT, THE RELOAD WAS REPLACED, BUT IT WAS DIFFICULT TO RELOAD. IT WAS NOTED THAT THERE WAS DIFFICULTY PUSHING THE "PUSH TO RELOAD" BUTTON, AND THE LOCK AND UNLOCK BUTTON WAS DIFFICULT TO OPERATE AS WELL. FIRING WAS CONTINUED, AND AS EXPECTED, THE RED BUTTON WAS STILL HARD, AND THE DOCTOR HAD TO APPLY A LOT OF FORCE TO PUSH IT IN. THE OPERATION ITSELF WAS COMPLETED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
18-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE PLASTIC BAG ON THE TIP OF THE DEVICE WAS TORN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOMETRIOSIS WITH RECTOSIGMOIDECTOMY, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE TISSUE WAS NOT STAPLED. THE STAPLER BECAME RESISTANT AND MADE A CLICK SOUND. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 12-01-2024 | OPTIFIX STRAIGHT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR PROCEDURE THE OPTIFIX DEVICE WAS MISFIRING AND DEPLOYING BROKEN FASTENERS. IT WAS REPORTED THAT THE PIECES OF FASTENER WERE REMOVED FROM PATIENT AND THE CASE WAS COMPLETED USING ANOTHER DEVICE. THERE WAS NO PATIENT HARM OR INJURY.
18-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. IT IS UNKNOWN THAT HOW TO OPEN THE JAW. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | FLEX 60 ARTICULATING | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE DEVICE COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. THEN CUT SOME TISSUE AND REMOVED THE DEVICE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE DEVICE COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. THEN CUT SOME TISSUE AND REMOVED THE DEVICE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE DEVICE COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. THEN CUT SOME TISSUE AND REMOVED THE DEVICE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE DEVICE COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. THEN CUT SOME TISSUE AND REMOVED THE DEVICE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, DURING PANCREATIC DUODENECTOMY, THE STAPLER WITH 2 RELOADS DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-09-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE HAS NOT TRIGGED. SURGERY WAS COMPLETED WITH A NEW INSTRUMENT. NO PATIENT HARM.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, DURING PANCREATIC DUODENECTOMY, THE STAPLER WITH 2 RELOADS DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-07-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-10-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CHOLECYSTECTOMY, DURING PANCREATIC DUODENECTOMY, THE STAPLER WITH 2 RELOADS DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, WHILE DISSOCIATING THE LIVER PEDICLE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. THE PRODUCT WAS CHANGED WITH THE SAME MODEL. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, WHILE DISSOCIATING THE LIVER PEDICLE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. THE PRODUCT WAS CHANGED WITH THE SAME MODEL. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, WHILE DISSOCIATING THE LIVER PEDICLE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. THE PRODUCT WAS CHANGED WITH THE SAME MODEL. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, WHILE DISSOCIATING THE LIVER PEDICLE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. THE PRODUCT WAS CHANGED WITH THE SAME MODEL. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, WHILE DISSOCIATING THE LIVER PEDICLE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. THE PRODUCT WAS CHANGED WITH THE SAME MODEL. THERE WAS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | GST GOLD RELOAD, 60MM, 6 ROW | Structural Problem | A0511 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, STAPLE FELL OFF INSIDE THE PATIENT. THE STAPLE WAS RETRIEVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | GST GOLD RELOAD, 45MM, 6 ROW | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: THE ACCESSORIES OF THE NEW RELOAD FELL OFF SO THE RELOADS FAILED TO INSTALL ONTO THE ECHELON. THE HCP CHANGED A NEW RELOAD AND SAFELY COMPLETED THE SURGERY. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | GST GOLD RELOAD, 45MM, 6 ROW | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: THE ACCESSORIES OF THE NEW RELOAD FELL OFF SO THE RELOADS FAILED TO INSTALL ONTO THE ECHELON. THE HCP CHANGED A NEW RELOAD AND SAFELY COMPLETED THE SURGERY. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
25-10-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | FLEX 60 ARTICULATING | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THAT THE TISSUE WAS NOT CUT THEN CHANGED ANOTHER DEVICE TO CONTINUE THE PROCEDURE. AND THEN DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THE CARTRIDGE PAN SEPARATION AFTER OPENED THE PACKING. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-10-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | FLEX 60 ARTICULATING | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THAT THE TISSUE WAS NOT CUT THEN CHANGED ANOTHER DEVICE TO CONTINUE THE PROCEDURE. AND THEN DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THE CARTRIDGE PAN SEPARATION AFTER OPENED THE PACKING. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC RIGHT UPPER LOBE RESECTION, ON BRONCHOTOMY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC RIGHT UPPER LOBE RESECTION, ON BRONCHOTOMY, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-09-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC TOTAL MESORECTAL RESECTION, WHILE ON SUPERIOR RECTAL ARTERY CLOSURE AND DISCONNECTION OF BLOOD VESSELS, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT THE STAPLER DID NOT FIRE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TISSUE RESECTION PROCEDURE, IN CUTTING AND STAPLING THE TISSUE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
22-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, DURING COLON RESECTION, AFTER CONNECTING THE RELOAD INTO THE HANDLE, THE GREEN BUTTON WAS PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED, AND THE DEVICE DID NOT FIRE. A NEW HANDLE AND 60MM RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
20-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOMETRIOSIS WITH RECTOSIGMOIDECTOMY, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE TISSUE WAS NOT STAPLED. THE STAPLER BECAME RESISTANT AND MADE A CLICK SOUND. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENDOMETRIOSIS WITH RECTOSIGMOIDECTOMY, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE TISSUE WAS NOT STAPLED. THE STAPLER BECAME RESISTANT AND MADE A CLICK SOUND. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, DURING COLON RESECTION, AFTER CONNECTING THE RELOAD INTO THE HANDLE, THE GREEN BUTTON WAS PRESSED BUT THE HANDLE COULD NOT BE SQUEEZED, AND THE DEVICE DID NOT FIRE. A NEW HANDLE AND 60MM RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
20-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. IT WAS STATED THAT THE REINFORCING TAB ON THE UPPER PART OF THE MAGAZINE WAS POORLY GLUED AND PREVENTED IT FROM STAPLING PROPERLY. ANOTHER RELOAD FROM THE SAME REFERENCE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN AND SQUEEZE THE HANDLE, BUT THE DEVICE DID NOT FIRE. IT WAS STATED THAT THE REINFORCING TAB ON THE UPPER PART OF THE MAGAZINE WAS POORLY GLUED AND PREVENTED IT FROM STAPLING PROPERLY. ANOTHER RELOAD FROM THE SAME REFERENCE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE ATTEMPTING TO STAPLE ACROSS THE ANTRUM OF THE STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE ATTEMPTING TO STAPLE ACROSS THE ANTRUM OF THE STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | UNKNOWN EGIA TRI STAPLE | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE ATTEMPTING TO STAPLE ACROSS THE ANTRUM OF THE STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | UNKNOWN EGIA TRI STAPLE | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE ATTEMPTING TO STAPLE ACROSS THE ANTRUM OF THE STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN AN OPEN EXPLORATORY BOWEL RESECTION, THE DEVICE WAS CLOSED AND THE BLACK PUSHER THUMB PIECE WAS MOVED, BUT THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SLEEVE GASTRECTOMY, WHILE ATTEMPTING TO STAPLE ACROSS THE ANTRUM OF THE STOMACH, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-11-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | ILS 29MM, CURVED | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY OPENED THE PACKING AND NOTED THAT SOME STAPLES FELL OFF. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-10-2023 | Injury | us surgical puerto rico | 12-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS CONDUCTED USING ELECTRONIC HEALTHCARE DATA TO ASSESS THE RISK OF ANASTOMOTIC LEAK IN PATIENTS WHO UNDERWENT LEFT-SIDED COLORECTAL SURGERY BETWEEN (B)(6) 2019 AND (B)(6) 2022. PATIENTS WHO HAD A BILLING CHARGE FOR A STUDY DEVICE WERE INCLUDED. THE DEVICES USED WERE EEA DST, TRISTAPLE EEA, AND A COMPETITOR STAPLER. THERE WERE 7928 PATIENTS IN THE MEDTRONIC EEA DST COHORT, 1306 PATIENTS IN THE TRISTAPLE EEA COHORT AND 8337 PATIENTS IN THE COMPETITOR COHORT. WITHIN 30 DAYS, THERE WERE 598 ANASTOMOTIC LEAKS REPORTED IN THE EEA DST GROUP AND 107 ANASTOMOTIC LEAKS REPORTED IN THE TRISTAPLE EEA GROUP. BECAUSE THERE IS NOT A SPECIFIC CODE TO DIAGNOSE ANASTOMOTIC LEAK, DIAGNOSIS WAS MADE BASED ON DIAGNOSIS CODES FOR PERITONEAL ABSCESS, PERITONITIS, FISTULA OF INTESTINE, OR STOMA FORMATION POST COLORECTAL ANASTOMOSIS PROCEDURE AS THESE USUALLY OCCUR WITH AN ANASTOMOTIC LEAK. NO INTERVENTIONS WERE REPORTED.
18-10-2023 | Injury | us surgical puerto rico | 12-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS CONDUCTED USING ELECTRONIC HEALTHCARE DATA TO ASSESS THE RISK OF ANASTOMOTIC LEAK IN PATIENTS WHO UNDERWENT LEFT-SIDED COLORECTAL SURGERY BETWEEN (B)(6) 2019 AND (B)(6) 2022. PATIENTS WHO HAD A BILLING CHARGE FOR A STUDY DEVICE WERE INCLUDED. THE DEVICES USED WERE EEA DST, TRISTAPLE EEA, AND A COMPETITOR STAPLER. THERE WERE 7928 PATIENTS IN THE MEDTRONIC EEA DST COHORT, 1306 PATIENTS IN THE TRISTAPLE EEA COHORT AND 8337 PATIENTS IN THE COMPETITOR COHORT. WITHIN 30 DAYS, THERE WERE 598 ANASTOMOTIC LEAKS REPORTED IN THE EEA DST GROUP AND 107 ANASTOMOTIC LEAKS REPORTED IN THE TRISTAPLE EEA GROUP. BECAUSE THERE IS NOT A SPECIFIC CODE TO DIAGNOSE ANASTOMOTIC LEAK, DIAGNOSIS WAS MADE BASED ON DIAGNOSIS CODES FOR PERITONEAL ABSCESS, PERITONITIS, FISTULA OF INTESTINE, OR STOMA FORMATION POST COLORECTAL ANASTOMOSIS PROCEDURE AS THESE USUALLY OCCUR WITH AN ANASTOMOTIC LEAK. NO INTERVENTIONS WERE REPORTED.
18-10-2023 | Injury | us surgical puerto rico | 12-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS CONDUCTED USING ELECTRONIC HEALTHCARE DATA TO ASSESS THE RISK OF ANASTOMOTIC LEAK IN PATIENTS WHO UNDERWENT LEFT-SIDED COLORECTAL SURGERY BETWEEN (B)(6) 2019 AND (B)(6) 2022. PATIENTS WHO HAD A BILLING CHARGE FOR A STUDY DEVICE WERE INCLUDED. THE DEVICES USED WERE EEA DST, TRISTAPLE EEA, AND A COMPETITOR STAPLER. THERE WERE 7928 PATIENTS IN THE MEDTRONIC EEA DST COHORT, 1306 PATIENTS IN THE TRISTAPLE EEA COHORT AND 8337 PATIENTS IN THE COMPETITOR COHORT. WITHIN 30 DAYS, THERE WERE 598 ANASTOMOTIC LEAKS REPORTED IN THE EEA DST GROUP AND 107 ANASTOMOTIC LEAKS REPORTED IN THE TRISTAPLE EEA GROUP. BECAUSE THERE IS NOT A SPECIFIC CODE TO DIAGNOSE ANASTOMOTIC LEAK, DIAGNOSIS WAS MADE BASED ON DIAGNOSIS CODES FOR PERITONEAL ABSCESS, PERITONITIS, FISTULA OF INTESTINE, OR STOMA FORMATION POST COLORECTAL ANASTOMOSIS PROCEDURE AS THESE USUALLY OCCUR WITH AN ANASTOMOTIC LEAK. NO INTERVENTIONS WERE REPORTED.
18-10-2023 | Injury | us surgical puerto rico | 12-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS CONDUCTED USING ELECTRONIC HEALTHCARE DATA TO ASSESS THE RISK OF ANASTOMOTIC LEAK IN PATIENTS WHO UNDERWENT LEFT-SIDED COLORECTAL SURGERY BETWEEN (B)(6) 2019 AND (B)(6) 2022. PATIENTS WHO HAD A BILLING CHARGE FOR A STUDY DEVICE WERE INCLUDED. THE DEVICES USED WERE EEA DST, TRISTAPLE EEA, AND A COMPETITOR STAPLER. THERE WERE 7928 PATIENTS IN THE MEDTRONIC EEA DST COHORT, 1306 PATIENTS IN THE TRISTAPLE EEA COHORT AND 8337 PATIENTS IN THE COMPETITOR COHORT. WITHIN 30 DAYS, THERE WERE 598 ANASTOMOTIC LEAKS REPORTED IN THE EEA DST GROUP AND 107 ANASTOMOTIC LEAKS REPORTED IN THE TRISTAPLE EEA GROUP. BECAUSE THERE IS NOT A SPECIFIC CODE TO DIAGNOSE ANASTOMOTIC LEAK, DIAGNOSIS WAS MADE BASED ON DIAGNOSIS CODES FOR PERITONEAL ABSCESS, PERITONITIS, FISTULA OF INTESTINE, OR STOMA FORMATION POST COLORECTAL ANASTOMOSIS PROCEDURE AS THESE USUALLY OCCUR WITH AN ANASTOMOTIC LEAK. NO INTERVENTIONS WERE REPORTED.
18-10-2023 | Injury | us surgical puerto rico | 12-01-2024 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Peritonitis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS CONDUCTED USING ELECTRONIC HEALTHCARE DATA TO ASSESS THE RISK OF ANASTOMOTIC LEAK IN PATIENTS WHO UNDERWENT LEFT-SIDED COLORECTAL SURGERY BETWEEN (B)(6) 2019 AND (B)(6) 2022. PATIENTS WHO HAD A BILLING CHARGE FOR A STUDY DEVICE WERE INCLUDED. THE DEVICES USED WERE EEA DST, TRISTAPLE EEA, AND A COMPETITOR STAPLER. THERE WERE 7928 PATIENTS IN THE MEDTRONIC EEA DST COHORT, 1306 PATIENTS IN THE TRISTAPLE EEA COHORT AND 8337 PATIENTS IN THE COMPETITOR COHORT. WITHIN 30 DAYS, THERE WERE 598 ANASTOMOTIC LEAKS REPORTED IN THE EEA DST GROUP AND 107 ANASTOMOTIC LEAKS REPORTED IN THE TRISTAPLE EEA GROUP. BECAUSE THERE IS NOT A SPECIFIC CODE TO DIAGNOSE ANASTOMOTIC LEAK, DIAGNOSIS WAS MADE BASED ON DIAGNOSIS CODES FOR PERITONEAL ABSCESS, PERITONITIS, FISTULA OF INTESTINE, OR STOMA FORMATION POST COLORECTAL ANASTOMOSIS PROCEDURE AS THESE USUALLY OCCUR WITH AN ANASTOMOTIC LEAK. NO INTERVENTIONS WERE REPORTED.
21-12-2023 | Injury | covidien lp llc north haven | 12-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL SWITCH PROCEDURE, WHEN THE FIRE BUTTON WAS PUSHED AND BEGAN TO STAPLE, THE SURGEON HEARD AN AUDIBLE "CLICK" OR "CLANK." THE RELOAD FIRED, PLACING A COMPLETE STAPLE LINE BUT WOULD NOT RETRACT THE KNIFE BLADE. THE RELOAD BECAME LOCKED ON THE TISSUE. THE RETRACTION DID NOT OPEN THE JAWS, AND THE RELOAD HAD TO BE CUT OUT. THE SURGEON CONVERTED TO A LAPAROSCOPIC GASTRIC BYPASS AND COMPLETED THE PROCEDURE USING A MANUAL STAPLER. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
21-12-2023 | Injury | covidien lp llc north haven | 12-01-2024 | TRI-STAPLE 2.0 | Noise, Audible | A0508 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL SWITCH PROCEDURE, WHEN THE FIRE BUTTON WAS PUSHED AND BEGAN TO STAPLE, THE SURGEON HEARD AN AUDIBLE "CLICK" OR "CLANK." THE RELOAD FIRED, PLACING A COMPLETE STAPLE LINE BUT WOULD NOT RETRACT THE KNIFE BLADE. THE RELOAD BECAME LOCKED ON THE TISSUE. THE RETRACTION DID NOT OPEN THE JAWS, AND THE RELOAD HAD TO BE CUT OUT. THE SURGEON CONVERTED TO A LAPAROSCOPIC GASTRIC BYPASS AND COMPLETED THE PROCEDURE USING A MANUAL STAPLER. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
01-01-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE SURGEON TRIED CLOSE THE HANDLES ON THE DEVICE SEVERAL TIMES WITHOUT SUCCESS. AS THE OR-TEAM COMMENTED, IT WAS KIND OF BLOCKED. NO OBVIOUS REASON MET THE EYE. NO HARM OCCURRED TO THE PATIENT.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL SWITCH PROCEDURE, WHEN THE FIRE BUTTON WAS PUSHED AND BEGAN TO STAPLE, THE SURGEON HEARD AN AUDIBLE "CLICK" OR "CLANK." THE RELOAD FIRED, PLACING A COMPLETE STAPLE LINE BUT WOULD NOT RETRACT THE KNIFE BLADE. THE RELOAD BECAME LOCKED ON THE TISSUE. THE RETRACTION DID NOT OPEN THE JAWS, AND THE RELOAD HAD TO BE CUT OUT. THE SURGEON CONVERTED TO A LAPAROSCOPIC GASTRIC BYPASS AND COMPLETED THE PROCEDURE USING A MANUAL STAPLER. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Noise, Audible | A0508 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL SWITCH PROCEDURE, WHEN THE FIRE BUTTON WAS PUSHED AND BEGAN TO STAPLE, THE SURGEON HEARD AN AUDIBLE "CLICK" OR "CLANK." THE RELOAD FIRED, PLACING A COMPLETE STAPLE LINE BUT WOULD NOT RETRACT THE KNIFE BLADE. THE RELOAD BECAME LOCKED ON THE TISSUE. THE RETRACTION DID NOT OPEN THE JAWS, AND THE RELOAD HAD TO BE CUT OUT. THE SURGEON CONVERTED TO A LAPAROSCOPIC GASTRIC BYPASS AND COMPLETED THE PROCEDURE USING A MANUAL STAPLER. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
15-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RESPIRATORY SURGERY, CARTRIDGE FELL OFF INSIDE THE PATIENT AT 2ND FIRING. THE CARTRIDGE WAS RETRIEVED AND WAS NOT USED ON THE PATIENT. THE CARTRIDGE WAS USED WITH PCEE60A. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. RECEIVED FOUR CARTRIDGES, AND 1 IS THE CARTRIDGE IN THE QUESTION, 2 ARE THE CARTRIDGES WHICH WAS ALREADY USED, 1 IS THE CARTRIDGE WHICH IS NOT USED AND WAS ON THE MAYO-TABLE DURING THE OPERATION. NO FURTHER INFORMATION IS AVAILABLE.
21-11-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | ILS 29MM, CURVED | Device Misassembled During Manufacturing /Shipping | A0206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY OPENED THE PACKING AND NOTED THAT SOME STAPLES FELL OFF. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL SWITCH PROCEDURE, WHEN THE FIRE BUTTON WAS PUSHED AND BEGAN TO STAPLE, THE SURGEON HEARD AN AUDIBLE "CLICK" OR "CLANK." THE RELOAD FIRED, PLACING A COMPLETE STAPLE LINE BUT WOULD NOT RETRACT THE KNIFE BLADE. THE RELOAD BECAME LOCKED ON THE TISSUE. THE RETRACTION DID NOT OPEN THE JAWS, AND THE RELOAD HAD TO BE CUT OUT. THE SURGEON CONVERTED TO A LAPAROSCOPIC GASTRIC BYPASS AND COMPLETED THE PROCEDURE USING A MANUAL STAPLER. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Noise, Audible | A0508 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL SWITCH PROCEDURE, WHEN THE FIRE BUTTON WAS PUSHED AND BEGAN TO STAPLE, THE SURGEON HEARD AN AUDIBLE "CLICK" OR "CLANK." THE RELOAD FIRED, PLACING A COMPLETE STAPLE LINE BUT WOULD NOT RETRACT THE KNIFE BLADE. THE RELOAD BECAME LOCKED ON THE TISSUE. THE RETRACTION DID NOT OPEN THE JAWS, AND THE RELOAD HAD TO BE CUT OUT. THE SURGEON CONVERTED TO A LAPAROSCOPIC GASTRIC BYPASS AND COMPLETED THE PROCEDURE USING A MANUAL STAPLER. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES.
06-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT TRIGGER, AND IT DID NOT WORK. IT DID NOT COMPLY WITH ITS FUNCTION. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC DUODENECTOMY PROCEDURE, WHILE RESECTING THE GASTRODUODENAL ARTERY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE HANDLE WAS REPLACED. THERE WAS NO PATIENT INJURY.
22-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, IN STOMACH RESECTION, THE OPENING WAS BIG EVEN WHEN THE JAWS WERE FULLY CLOSED. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER. THERE WAS NO PATIENT INJURY.
13-11-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A EXPLORATORY LAPAROTOMY PROCEDURE THAT WHILE USING THE FIRST DEVICE WHEN CLOSING THEY HEARD A LOUD CRACK AND IT FELT LIKE IT WASN'T CLOSING ALL THE WAY. SURGEON ATTEMPTED TO FIRE BUT COULDN'T. DEVICE WAS REMOVED AND ATTEMPTED TO FIRE AGAIN ON A TOWEL ON THE BACK TABLE AND THEY WERE NOT ABLE TO GET IT TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
13-11-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A EXPLORATORY LAPAROTOMY PROCEDURE THAT WHILE USING THE FIRST DEVICE WHEN CLOSING THEY HEARD A LOUD CRACK AND IT FELT LIKE IT WASN'T CLOSING ALL THE WAY. SURGEON ATTEMPTED TO FIRE BUT COULDN'T. DEVICE WAS REMOVED AND ATTEMPTED TO FIRE AGAIN ON A TOWEL ON THE BACK TABLE AND THEY WERE NOT ABLE TO GET IT TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
13-11-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | ECHELON CONTOUR W/ GREEN RELOAD | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A EXPLORATORY LAPAROTOMY PROCEDURE THAT WHILE USING THE FIRST DEVICE WHEN CLOSING THEY HEARD A LOUD CRACK AND IT FELT LIKE IT WASN'T CLOSING ALL THE WAY. SURGEON ATTEMPTED TO FIRE BUT COULDN'T. DEVICE WAS REMOVED AND ATTEMPTED TO FIRE AGAIN ON A TOWEL ON THE BACK TABLE AND THEY WERE NOT ABLE TO GET IT TO FIRE. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
22-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, IN STOMACH RESECTION, THE OPENING WAS BIG EVEN WHEN THE JAWS WERE FULLY CLOSED. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER. THERE WAS NO PATIENT INJURY.
20-11-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | 55MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SPLENECTOMY AN ISSUE AROSE WITH THE NTLC75 DEVICE; IT BECAME JAMMED, AND NONE OF THE HEALTHCARE PROFESSIONALS PRESENT IN THE OPERATING ROOM COULD DISCHARGE (FIRE) IT. WHEN USING THE LINEAR STAPLER, TISSUE STITCHING DID NOT OCCUR, AND SUBSEQUENT INCISION WAS NOT POSSIBLE. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH ANOTHER DEVICE WITHOUT PATIENT CONSEQUENCES.
22-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, IN STOMACH RESECTION, THE OPENING WAS BIG EVEN WHEN THE JAWS WERE FULLY CLOSED. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER. THERE WAS NO PATIENT INJURY.
22-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY (BARIATRIC SURGERY), AT THE TIME OF FIRING THE SECOND 45MM RELOAD AT THE BOTTOM OF THE STOMACH, THERE WAS A FEELING WITH THE HANDLE WHICH HAD MORE LOUD NOISE, AND WHEN RE-SQUEEZES THE TRIGGER, THE HANDLE SOUNDED LOUDER, AT WHICH THE SURGEON DECIDED NOT TO CONTINUE TO TRY TO STAPLE. WHEN THE DOCTOR RETRACTS THE BLACK KNOB ON THE HANDLE, IT WAS FOUND THAT THE STAPLES HAD NOT BEEN FORMED, REFERRED TO BY THE SURGEON AS ¿AN EXPLOSION" WITHIN THE PATIENT. IT WAS MENTIONED THAT THE STAPLE LINE WAS INCOMPLETE DISTALLY AND SOME OF THE STAPLES FELL INTO THE PATIENT CAVITY. THE TISSUE HAD TO BE STAPLED HIGHER THAN EXPECTED, LEAVING THE STOMACH SMALLER. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE SURGEON HAD TO USED ANOTHER HANDLE AND RELOAD TO RESOLVE THE ISSUE.
22-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY (BARIATRIC SURGERY), AT THE TIME OF FIRING THE SECOND 45MM RELOAD AT THE BOTTOM OF THE STOMACH, THERE WAS A FEELING WITH THE HANDLE WHICH HAD MORE LOUD NOISE, AND WHEN RE-SQUEEZES THE TRIGGER, THE HANDLE SOUNDED LOUDER, AT WHICH THE SURGEON DECIDED NOT TO CONTINUE TO TRY TO STAPLE. WHEN THE DOCTOR RETRACTS THE BLACK KNOB ON THE HANDLE, IT WAS FOUND THAT THE STAPLES HAD NOT BEEN FORMED, REFERRED TO BY THE SURGEON AS ¿AN EXPLOSION" WITHIN THE PATIENT. IT WAS MENTIONED THAT THE STAPLE LINE WAS INCOMPLETE DISTALLY AND SOME OF THE STAPLES FELL INTO THE PATIENT CAVITY. THE TISSUE HAD TO BE STAPLED HIGHER THAN EXPECTED, LEAVING THE STOMACH SMALLER. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE SURGEON HAD TO USED ANOTHER HANDLE AND RELOAD TO RESOLVE THE ISSUE.
22-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY (BARIATRIC SURGERY), AT THE TIME OF FIRING THE SECOND 45MM RELOAD AT THE BOTTOM OF THE STOMACH, THERE WAS A FEELING WITH THE HANDLE WHICH HAD MORE LOUD NOISE, AND WHEN RE-SQUEEZES THE TRIGGER, THE HANDLE SOUNDED LOUDER, AT WHICH THE SURGEON DECIDED NOT TO CONTINUE TO TRY TO STAPLE. WHEN THE DOCTOR RETRACTS THE BLACK KNOB ON THE HANDLE, IT WAS FOUND THAT THE STAPLES HAD NOT BEEN FORMED, REFERRED TO BY THE SURGEON AS ¿AN EXPLOSION" WITHIN THE PATIENT. IT WAS MENTIONED THAT THE STAPLE LINE WAS INCOMPLETE DISTALLY AND SOME OF THE STAPLES FELL INTO THE PATIENT CAVITY. THE TISSUE HAD TO BE STAPLED HIGHER THAN EXPECTED, LEAVING THE STOMACH SMALLER. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE SURGEON HAD TO USED ANOTHER HANDLE AND RELOAD TO RESOLVE THE ISSUE.
06-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE RESECTION, AT THE 3RD AND 4TH FIRING OF STOMACH RESECTION, EXCESSIVE FORCE INDICATION WAS OBSERVED IN THE HANDLE, AND THE FIRING STOPPED HALFWAY THROUGH. THIS ISSUE OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, ANOTHER RELOAD OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE RESECTION, AT THE 3RD AND 4TH FIRING OF STOMACH RESECTION, EXCESSIVE FORCE INDICATION WAS OBSERVED IN THE HANDLE, AND THE FIRING STOPPED HALFWAY THROUGH. THIS ISSUE OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, ANOTHER RELOAD OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
22-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY, THE OTHER STAPLE WAS USED NORMALLY AND WHILE FINISHING THE PROCEDURE, THE LAST STAPLE FOR CONFECTION ENTERO-ENTERO, THE RELOAD FIRED HALFWAY ONLY AND RETREATED THE BLADE. IT WAS MENTIONED THAT THE STAPLE LINE UNTIL HALFWAY THROUGH THE STAPLE WERE UNIFORM AND WITHOUT ANY ABNORMALITIES. THE SURGEON USED SUTURE TO RESOLVE THE ISSUE AND FINISHED THE CASE.
22-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY, THE OTHER STAPLE WAS USED NORMALLY AND WHILE FINISHING THE PROCEDURE, THE LAST STAPLE FOR CONFECTION ENTERO-ENTERO, THE RELOAD FIRED HALFWAY ONLY AND RETREATED THE BLADE. IT WAS MENTIONED THAT THE STAPLE LINE UNTIL HALFWAY THROUGH THE STAPLE WERE UNIFORM AND WITHOUT ANY ABNORMALITIES. THE SURGEON USED SUTURE TO RESOLVE THE ISSUE AND FINISHED THE CASE.
22-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY, THE OTHER STAPLE WAS USED NORMALLY AND WHILE FINISHING THE PROCEDURE, THE LAST STAPLE FOR CONFECTION ENTERO-ENTERO, THE RELOAD FIRED HALFWAY ONLY AND RETREATED THE BLADE. IT WAS MENTIONED THAT THE STAPLE LINE UNTIL HALFWAY THROUGH THE STAPLE WERE UNIFORM AND WITHOUT ANY ABNORMALITIES. THE SURGEON USED SUTURE TO RESOLVE THE ISSUE AND FINISHED THE CASE.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | EEA | Failure to Cut | A050702 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A HEMORROIDECTOMY PROCEDURE, AFTER DOING THE PURSE STRING AND FIRING THE DEVICE, THE STAPLES SEEMED TO DEPLOY PROPERLY BUT IF DID NOT CUT PROPERLY. THE ANVIL WAS TAKEN OUT.
01-01-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | 75MM SELECTABLE NEW TLC | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT STAFF FIRED THREE DIFFERENT TIMES AND THERE WAS OOZING AT THE STAPLE LINE FOR ALL THREE FIRINGS. THEY HAD IT ON THE BLUE RELOAD SETTING. STAFF ENDED UP PULLING A COMPETITOR'S VERSION, THE GIA AND THERE WAS ALSO OOZING AT THE STAPLE LINE WHEN USING ALTERNATIVE PRODUCT. HOW THE OOZING WAS CONTROLLED WAS NOT PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | 75MM SELECTABLE NEW TLC | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT STAFF FIRED THREE DIFFERENT TIMES AND THERE WAS OOZING AT THE STAPLE LINE FOR ALL THREE FIRINGS. THEY HAD IT ON THE BLUE RELOAD SETTING. STAFF ENDED UP PULLING A COMPETITOR'S VERSION, THE GIA AND THERE WAS ALSO OOZING AT THE STAPLE LINE WHEN USING ALTERNATIVE PRODUCT. HOW THE OOZING WAS CONTROLLED WAS NOT PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
13-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN THORACOTOMY PROCEDURE, AT THE TIME OF FIRING THE FIFTH RELOAD, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. THEREFORE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS LOADED AND FIRED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT A SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
06-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE RESECTION, AT THE 3RD AND 4TH FIRING OF STOMACH RESECTION, EXCESSIVE FORCE INDICATION WAS OBSERVED IN THE HANDLE, AND THE FIRING STOPPED HALFWAY THROUGH. THIS ISSUE OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, ANOTHER RELOAD OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE RESECTION, AT THE 3RD AND 4TH FIRING OF STOMACH RESECTION, EXCESSIVE FORCE INDICATION WAS OBSERVED IN THE HANDLE, AND THE FIRING STOPPED HALFWAY THROUGH. THIS ISSUE OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, ANOTHER RELOAD OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE RESECTION, AT THE 3RD AND 4TH FIRING OF STOMACH RESECTION, EXCESSIVE FORCE INDICATION WAS OBSERVED IN THE HANDLE, AND THE FIRING STOPPED HALFWAY THROUGH. THIS ISSUE OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, ANOTHER PRODUCT OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE RESECTION, AT THE 3RD AND 4TH FIRING OF STOMACH RESECTION, EXCESSIVE FORCE INDICATION WAS OBSERVED IN THE HANDLE, AND THE FIRING STOPPED HALFWAY THROUGH. THIS ISSUE OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, ANOTHER PRODUCT OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
02-10-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS PROCEDURE, AT THE TIME OF WHOLE ANASTOMOSIS, THE DEVICE DID NOT FIRE. THE SURGEON USED SUTURE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-10-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS PROCEDURE, AT THE TIME OF WHOLE ANASTOMOSIS, THE DEVICE DID NOT FIRE. THE SURGEON USED SUTURE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, AT THE TIME OF DUODENECTOMY, THE DEVICE WAS NOT POSSIBLE TO FIRE EVEN WHEN THE HANDLE WAS SQUEEZED. ANOTHER DEVICE WITH THE SAME PRODUCT NUMBER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AFTER FIRING STAPLER, THERE WAS A WOVEN METAL SCRAPS INSIDE THE PATIENT AND IN/AROUND THE STAPLE LINE. THEY WEREN'T ABLE TO CONFIRM IF IT WAS FROM THE DEVICE OR A FOREIGN BODY. NO PATIENT HARM.
12-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION PROCEDURE, WHEN RESECTING THE COLON, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, WHEN AN ATTEMPT WAS MADE TO PERFORM A RESECTION OF THE MOUTH SIDE INTESTINAL TRACT, THE JAWS OF THE RELOAD COULD NOT BE CLOSED COMPLETELY. THE INSERTION OF THE RELOAD INTO THE ABDOMINAL CAVITY WAS DIFFICULT. AT THAT TIME, THE HANDLE DISPLAY ON THE SCREEN WAS AT ZONE ONE. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
20-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, WHEN AN ATTEMPT WAS MADE TO PERFORM A RESECTION OF THE MOUTH SIDE INTESTINAL TRACT, THE JAWS OF THE RELOAD COULD NOT BE CLOSED COMPLETELY. THE INSERTION OF THE RELOAD INTO THE ABDOMINAL CAVITY WAS DIFFICULT. AT THAT TIME, THE HANDLE DISPLAY ON THE SCREEN WAS AT ZONE ONE. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ENTERECTOMY AND PANCREATECTOMY PROCEDURE, THE DEVICE DID FIRE, EVEN WITH CORRECT AS ASSEMBLY. A NEW STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, IT WAS IMPOSSIBLE TO ACTIVE THE MECHANISM OF THE DEVICE. NO CLINICAL CONSEQUENCE FOR THE PATIENT. USE OF ANOTHER FUNCTIONAL DEVICE FROM SAME PRODUCT CODE.
11-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, AT THE TIME OF DUODENECTOMY, THE DEVICE WAS NOT POSSIBLE TO FIRE EVEN WHEN THE HANDLE WAS SQUEEZED. ANOTHER DEVICE WITH THE SAME PRODUCT NUMBER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, AT THE TIME OF DUODENECTOMY, THE DEVICE WAS NOT POSSIBLE TO FIRE EVEN WHEN THE HANDLE WAS SQUEEZED. ANOTHER DEVICE WITH THE SAME PRODUCT NUMBER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTROPLASTY, WHILE ON AN ATTEMPT TO TRIGGER THE LOAD, THE STAPLER DID NOT FIRE. THE DISPLAY SHOWED NO RECOGNITION OF THE LOAD OR EMPTY LOAD. THE RELOAD WAS REPLACED AND WORKED NORMALLY. THERE WAS NO PATIENT INJURY.
17-11-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE STAPLER DID NOT FIRE AND HAD NO POWER. THERE WAS A 5-MINUTE DELAY. NO FRAGMENTS MADE. THE PROCEDURE WAS COMPLETED WITH A NEW DEVICE WITH NO PATIENT HARM.
20-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY, WHEN AN ATTEMPT WAS MADE TO PERFORM A RESECTION OF THE MOUTH SIDE INTESTINAL TRACT, THE JAWS OF THE RELOAD COULD NOT BE CLOSED COMPLETELY. THE INSERTION OF THE RELOAD INTO THE ABDOMINAL CAVITY WAS DIFFICULT. AT THAT TIME, THE HANDLE DISPLAY ON THE SCREEN WAS AT ZONE ONE. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE RESECTION, AT THE 3RD AND 4TH FIRING OF STOMACH RESECTION, EXCESSIVE FORCE INDICATION WAS OBSERVED IN THE HANDLE, AND THE FIRING STOPPED HALFWAY THROUGH. THIS ISSUE OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, ANOTHER RELOAD OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
06-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE RESECTION, AT THE 3RD AND 4TH FIRING OF STOMACH RESECTION, EXCESSIVE FORCE INDICATION WAS OBSERVED IN THE HANDLE, AND THE FIRING STOPPED HALFWAY THROUGH. THIS ISSUE OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, ANOTHER RELOAD OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, WHEN STAPLING AT THE BEGINNING, THE PRODUCT CRASHED AND THE DEVICE DID NOT FIRE. THE DEVICE WAS CHANGED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-10-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE WAS NOT FIRED. NO PATIENT CONSEQUENCE REPORTED.
27-12-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-11-2023 | Malfunction | ethicon endo-surgery | 12-01-2024 | CONTOUR CURVED CUTTER STAPLER | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE PIN DOES NOT WORK. NO PATIENT CONSEQUENCE REPORTED.
21-12-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE GREEN BUTTON WAS PRESSED, AND THE HANDLE WAS SQUEEZED, HOWEVER, THE SUTURE RESECTION COULD NOT BE PERFORMED. BOTH HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
11-10-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE AND COULD NOT BE UNLOADED. NO PATIENT INJURY.
11-10-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE AND COULD NOT BE UNLOADED. NO PATIENT INJURY.
11-10-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE AND COULD NOT BE UNLOADED. NO PATIENT INJURY.
11-10-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE AND COULD NOT BE UNLOADED. NO PATIENT INJURY.
28-11-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PARTIAL ROBOTIC HEPATECTOMY WITH HYBRID STAPLING, THE FIRST HANDLE WAS HARD TO CLOSE, BUT MANAGED AND PRESSED GREEN BUTTON. SURGEON NOT ABLE TO SQUEEZE THE HANDLE AND UNABLE TO FIRE. INSTRUMENT STUCK IN TISSUE AND RELEASED AFTER TURNING AND TWITCHING. SECOND HANDLE WITH NEW RELOAD WERE USED, HANDLE WAS CLOSED AND GREEN BUTTON ACTIVATED, THEN THE BLACK HANDLE SHOULD GO INTO FIRING MODE, BUT DID NOT, IT WAS NOT STABLED AND HANGING LOOSE, NOT BEING ABLE TO FIRE AND THE SURGEON COULD NOT SQUEEZE THE HANDLE. THIRD HANDLE WAS TAKEN OUT, SAME RELOAD WAS USED AND FIRING WAS FINE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP AND FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
28-11-2023 | Malfunction | us surgical puerto rico | 12-01-2024 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PARTIAL ROBOTIC HEPATECTOMY WITH HYBRID STAPLING, THE FIRST HANDLE WAS HARD TO CLOSE, BUT MANAGED AND PRESSED GREEN BUTTON. SURGEON NOT ABLE TO SQUEEZE THE HANDLE AND UNABLE TO FIRE. INSTRUMENT STUCK IN TISSUE AND RELEASED AFTER TURNING AND TWITCHING. SECOND HANDLE WITH NEW RELOAD WERE USED, HANDLE WAS CLOSED AND GREEN BUTTON ACTIVATED, THEN THE BLACK HANDLE SHOULD GO INTO FIRING MODE, BUT DID NOT, IT WAS NOT STABLED AND HANGING LOOSE, NOT BEING ABLE TO FIRE AND THE SURGEON COULD NOT SQUEEZE THE HANDLE. THIRD HANDLE WAS TAKEN OUT, SAME RELOAD WAS USED AND FIRING WAS FINE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP AND FIRING OVER TISSUE THAT IS BEYOND THE RECOMMENDED THICKNESS RANGE.
18-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | ENDO GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) WITH LEFT UPPER LOBE LOBECTOMY, WHEN THE DOCTOR APPLIED RELOAD FOR TRANSECTION AT BRONCHUS, THE DEVICE FIRED, BUT FIRING WAS INCOMPLETE PROXIMALLY AND DISTALLY. THERE WAS MALFORMATION OF STAPLE AND STAPLE LINE WAS INCOMPLETE PROXIMALLY. THE STAPLE LINE WAS FIXED WHEN THE SURGEON OVERSEW ALL STAPLE LINE AT BRONCHUS WITH SUTURE. AFTER THAT SURGEON USED SAME POWERED STAPLER SYSTEM WITH ANOTHER RELOAD TO TRANSECT VESSEL WITH GOOD STAPLE FORMATION.
18-12-2023 | Malfunction | covidien lp llc north haven | 12-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) WITH LEFT UPPER LOBE LOBECTOMY, WHEN THE DOCTOR APPLIED RELOAD FOR TRANSECTION AT BRONCHUS, THE DEVICE FIRED, BUT FIRING WAS INCOMPLETE PROXIMALLY AND DISTALLY. THERE WAS MALFORMATION OF STAPLE AND STAPLE LINE WAS INCOMPLETE PROXIMALLY. THE STAPLE LINE WAS FIXED WHEN THE SURGEON OVERSEW ALL STAPLE LINE AT BRONCHUS WITH SUTURE. AFTER THAT SURGEON USED SAME POWERED STAPLER SYSTEM WITH ANOTHER RELOAD TO TRANSECT VESSEL WITH GOOD STAPLE FORMATION.
05-12-2023 | Injury | standard bariatrics | 12-01-2024 | TITAN SGS | Adverse Event Without Identified Device or Use Problem | A24 | Intra-Abdominal Hemorrhage | Event Description: A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE WAS COMPLETED WITHOUT ANY REPORTED INCIDENTS. AFTER AN UNDISCLOSED AMOUNT OF TIME, THE PATIENT WAS BROUGHT BACK TO THE OR FOR A SECOND SURGERY DUE TO A SUSPECTED BLEED DUE TO THE STAPLER. ADDITIONAL INFORMATION WAS NOT PROVIDED.
14-12-2023 | Malfunction | us surgical puerto rico | 14-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY, DURING STOMACH STAPLING, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 14-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE.
21-12-2023 | Malfunction | us surgical puerto rico | 14-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE.
21-12-2023 | Malfunction | us surgical puerto rico | 14-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE.
21-12-2023 | Malfunction | us surgical puerto rico | 14-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLER DID NOT FIRE. THE SURGEON COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE.
18-12-2023 | Malfunction | us surgical puerto rico | 14-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY WITH ROBOT SUPPORT TOTAL GASTRECTOMY, AT THE TIME OF GASTRIC INCISION, THE STAPLER DID NOT FIRE. ONLY THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-09-2023 | Malfunction | covidien lp llc north haven | 15-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, DURING CHARGING, ONCE THE START UP WAS CHECKED, AND UPON PUTTING THE HANDLE BACK TO THE CHARGER, ONLY THE RED FRAME APPEARED IN THE BLACK DISPLAY WITH NO SYMBOL. AND A "WHIRRING" AUDIBLE ERROR TONE WAS HEARD. WHEN THE HANDLE WAS REMOVED FROM THE CHARGER AND PUT BACK IN, IT STARTED CHARGING WITH NO PROBLEM. THE INVOLVED CHARGER STILL ABLE TO CHARGE OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE CAUSE OF THE OBSERVED ERROR WAS DUE TO A SOFTWARE RESTRICTIONS ON THE CAPACITY OF THE QUEUE. THE HANDLE QUEUE FILLED UP DUE TO MULTIPLE BUTTON PRESSES IN SHORT SUCCESSION.
08-09-2023 | Malfunction | covidien lp llc north haven | 15-01-2024 | SIGNIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, DURING CHARGING, ONCE THE START UP WAS CHECKED, AND UPON PUTTING THE HANDLE BACK TO THE CHARGER, ONLY THE RED FRAME APPEARED IN THE BLACK DISPLAY WITH NO SYMBOL. AND A "WHIRRING" AUDIBLE ERROR TONE WAS HEARD. WHEN THE HANDLE WAS REMOVED FROM THE CHARGER AND PUT BACK IN, IT STARTED CHARGING WITH NO PROBLEM. THE INVOLVED CHARGER STILL ABLE TO CHARGE OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE CAUSE OF THE OBSERVED ERROR WAS DUE TO A SOFTWARE RESTRICTIONS ON THE CAPACITY OF THE QUEUE. THE HANDLE QUEUE FILLED UP DUE TO MULTIPLE BUTTON PRESSES IN SHORT SUCCESSION.
08-09-2023 | Malfunction | covidien lp llc north haven | 15-01-2024 | SIGNIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, DURING CHARGING, ONCE THE START UP WAS CHECKED, AND UPON PUTTING THE HANDLE BACK TO THE CHARGER, ONLY THE RED FRAME APPEARED IN THE BLACK DISPLAY WITH NO SYMBOL. AND A "WHIRRING" AUDIBLE ERROR TONE WAS HEARD. WHEN THE HANDLE WAS REMOVED FROM THE CHARGER AND PUT BACK IN, IT STARTED CHARGING WITH NO PROBLEM. THE INVOLVED CHARGER STILL ABLE TO CHARGE OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE CAUSE OF THE OBSERVED ERROR WAS DUE TO A SOFTWARE RESTRICTIONS ON THE CAPACITY OF THE QUEUE. THE HANDLE QUEUE FILLED UP DUE TO MULTIPLE BUTTON PRESSES IN SHORT SUCCESSION.
26-12-2023 | Malfunction | us surgical puerto rico | 15-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION WHEN RESECTING THE ILEUM, THE FIRING WAS STARTED WITH ALL GREEN DISPLAY, BUT AN ERROR APPEARED, AND FIRING COULD NOT BE PERFORMED. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE, ADAPTER AND SHELL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 16-01-2024 | OPTIFIX AT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN INGUINAL HERNIA PROCEDURE, THE OPTIFIX AT FIXATION DEVICE DID NOT EXHIBIT FORCE TO RELEASE FASTENERS PROPERLY, CAUSING THE FASTENERS TO FALL OFF THE MESH/TISSUE WHEN ATTEMPTING TO FIXATE. IT WAS REPORTED THAT THE LOOSE FASTENERS WERE REMOVED FROM THE PATIENT. IT WAS ALSO REPORTED THAT WHEN ATTEMPTING TO DRY FIRE OUTSIDE OF THE PATIENT, THE DEVICE DEPLOYED BROKEN FASTENERS. THERE WAS NO PATIENT INJURY.
20-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 16-01-2024 | OPTIFIX AT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN INGUINAL HERNIA PROCEDURE, THE OPTIFIX AT FIXATION DEVICE DID NOT EXHIBIT FORCE TO RELEASE FASTENERS PROPERLY, CAUSING THE FASTENERS TO FALL OFF THE MESH/TISSUE WHEN ATTEMPTING TO FIXATE. IT WAS REPORTED THAT THE LOOSE FASTENERS WERE REMOVED FROM THE PATIENT. IT WAS ALSO REPORTED THAT WHEN ATTEMPTING TO DRY FIRE OUTSIDE OF THE PATIENT, THE DEVICE DEPLOYED BROKEN FASTENERS. THERE WAS NO PATIENT INJURY.
05-10-2023 | Malfunction | johnson & johnson medical (suzhou | 16-01-2024 | CARTRIDGE FINAL ASSY 60MM WHITE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE WHEN SEVERING DUODENUM TISSUE ON THE FIRING. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED WITH IRREGULAR SHAPE. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 16-01-2024 | OPTIFIX STRAIGHT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNSPECIFIED PROCEDURE, THE OPTIFIX STRAIGHT FAILED TO DEPLOY THE FASTENER FROM THE FIRST SHOT. IT WAS REPORTED THAT THE "BOLT CAME OUT AT THE TIP OF THE DEVICE. THERE WAS NO PATIENT INJURY. ADDENDUM PER ADDITIONAL INFORMATION: PHOTOS RECEIVED SHOW SIGNIFICANT DAMAGE TO THE DISTAL TIP OF THE DEVICE. THERE ARE NO DETACHED COMPONENTS. IT WAS REPORTED THAT THE SURGEON BECAME ANNOYED AS THERE WAS DELAY IN PROCEDURE DUE TO THE DEVICE ISSUE AND NO ADDITIONAL MEDICATION WAS NEEDED.
14-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 16-01-2024 | OPTIFIX STRAIGHT | Material Twisted/Bent | A040609 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNSPECIFIED PROCEDURE, THE OPTIFIX STRAIGHT FAILED TO DEPLOY THE FASTENER FROM THE FIRST SHOT. IT WAS REPORTED THAT THE "BOLT CAME OUT AT THE TIP OF THE DEVICE. THERE WAS NO PATIENT INJURY. ADDENDUM PER ADDITIONAL INFORMATION: PHOTOS RECEIVED SHOW SIGNIFICANT DAMAGE TO THE DISTAL TIP OF THE DEVICE. THERE ARE NO DETACHED COMPONENTS. IT WAS REPORTED THAT THE SURGEON BECAME ANNOYED AS THERE WAS DELAY IN PROCEDURE DUE TO THE DEVICE ISSUE AND NO ADDITIONAL MEDICATION WAS NEEDED.
20-10-2023 | Malfunction | johnson & johnson medical (suzhou | 16-01-2024 | CARTRIDGE FINAL ASSY 45MM BLUE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK PROCEDURE, DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-10-2023 | Malfunction | johnson & johnson medical (suzhou | 16-01-2024 | CARTRIDGE FINAL ASSY 60MM GOLD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, DEVICE COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER WOULD NOT OPEN AFTER USE. THIS HAPPENED INTRA-OP COMPLETED WITH LIKE DEVICE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
14-11-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY, THE JAWS DID NOT OPEN AFTER 1ST FIRING. MANUAL OVERRIDE LEVER WAS USED TO OPEN THE JAWS. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
29-11-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THE CARTRIDGE PAN SEPARATION AFTER OPENED THE PACKING. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AND THEN DURING THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. UNKNOWN HOW TO OPEN THE JAW FINALLY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
29-11-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Manufacturing, Packaging or Shipping Problem | A02 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THE CARTRIDGE PAN SEPARATION AFTER OPENED THE PACKING. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AND THEN DURING THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. UNKNOWN HOW TO OPEN THE JAW FINALLY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
29-11-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, BEFORE USED ON THE PATIENT, NOTED THE CARTRIDGE PAN SEPARATION AFTER OPENED THE PACKING. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AND THEN DURING THE SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. UNKNOWN HOW TO OPEN THE JAW FINALLY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, THERE WAS NO PROBLEM ON THE FIRST FIRING OF THE RECTAL RESECTION. HOWEVER, WHEN THE DEVICE WAS PULLED BACK, IT STOPPED, AND THE JAWS COULD NOT BE OPENED. AFTER A WHILE, WHEN AN ATTEMPT WAS MADE TO PULL IT BACK AGAIN, THE PULL-BACK OR WITHDRAWAL RESUMED, AND THE JAWS COULD BE SUCCESSFULLY OPENED. THE SECOND AND SUBSEQUENT FIRINGS WITH BOTH THE HANDLE AND RELOAD WORKED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, THERE WAS NO PROBLEM ON THE FIRST FIRING OF THE RECTAL RESECTION. HOWEVER, WHEN THE DEVICE WAS PULLED BACK, IT STOPPED, AND THE JAWS COULD NOT BE OPENED. AFTER A WHILE, WHEN AN ATTEMPT WAS MADE TO PULL IT BACK AGAIN, THE PULL-BACK OR WITHDRAWAL RESUMED, AND THE JAWS COULD BE SUCCESSFULLY OPENED. THE SECOND AND SUBSEQUENT FIRINGS WITH BOTH THE HANDLE AND RELOAD WORKED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) PULMONARY SEGMENTAL RESECTION PROCEDURE, THE SURGEON COULD PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) PULMONARY SEGMENTAL RESECTION PROCEDURE, THE SURGEON COULD PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON COULD PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE.
25-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SURGEON COULD PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE.
25-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | EEA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE BOUNCE CAP DROPPED AND BECAME CONTAMINATED. THERE WAS NO PATIENT INVOLVEMENT.
22-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR) PROCEDURE, AFTER FIRING, THE KNIFE BLADE COULD NOT BE PULLED BACK. TO RESOLVE THE ISSUE, ANOTHER HANDLE OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, THERE WAS NO PROBLEM ON THE FIRST FIRING OF THE RECTAL RESECTION. HOWEVER, WHEN THE DEVICE WAS PULLED BACK, IT STOPPED, AND THE JAWS COULD NOT BE OPENED. AFTER A WHILE, WHEN AN ATTEMPT WAS MADE TO PULL IT BACK AGAIN, THE PULL-BACK OR WITHDRAWAL RESUMED, AND THE JAWS COULD BE SUCCESSFULLY OPENED. THE SECOND AND SUBSEQUENT FIRINGS WITH BOTH THE HANDLE AND RELOAD WORKED WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
25-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) LOBECTOMY PROCEDURE, WHILE ASSEMBLING THE DEVICE FOR PULMONARY PARENCHYMA RESECTION, THE DEVICE WAS DIFFICULT TO LOAD. UNNECESSARY TENSION WAS APPLIED AND THE PROXIMAL SIDE (LOCK RING'S PERIPHERAL PART) OF THE RELOAD WAS DAMAGED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE; AND THE ADAPTER COULD BE USED CONTINUOUSLY. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) LOBECTOMY PROCEDURE, WHILE ASSEMBLING THE DEVICE FOR PULMONARY PARENCHYMA RESECTION, THE DEVICE WAS DIFFICULT TO LOAD. UNNECESSARY TENSION WAS APPLIED AND THE PROXIMAL SIDE (LOCK RING'S PERIPHERAL PART) OF THE RELOAD WAS DAMAGED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE; AND THE ADAPTER COULD BE USED CONTINUOUSLY. THERE WAS NO PATIENT INJURY.
01-12-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | BLUE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE HANDLE WAS BROKEN. THERE WAS NO PATIENT INVOLVED.
15-04-2022 | Malfunction | us surgical puerto rico | 16-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER FIRED BUT THERE WAS AN UNLOADING PROBLEM WHEN THE DOCTOR WANTED TO STAPLE. IN THE SECOND RELOAD, THE STAPLER STUCK (JAMMED) AND STAPLING WAS NO LONGER POSSIBLE. THE RELOADS WERE REPLACED, BUT THE STAPLER STILL DID NOT WORK. THERE WAS NO PATIENT INJURY.
15-04-2022 | Malfunction | us surgical puerto rico | 16-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER FIRED BUT THERE WAS AN UNLOADING PROBLEM WHEN THE DOCTOR WANTED TO STAPLE. IN THE SECOND RELOAD, THE STAPLER STUCK (JAMMED) AND STAPLING WAS NO LONGER POSSIBLE. THE RELOADS WERE REPLACED, BUT THE STAPLER STILL DID NOT WORK. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, IN DETACHING THE INTERLOBAR FISSURE, THE SURGEON COULD PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE OF THE SAME TYPE WAS USED WITH THE SAME RELOAD. THERE WAS NO PATIENT INJURY.
15-04-2022 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE STAPLER FIRED BUT THERE WAS AN UNLOADING PROBLEM WHEN THE DOCTOR WANTED TO STAPLE. IN THE SECOND RELOAD, THE STAPLER STUCK (JAMMED) AND STAPLING WAS NO LONGER POSSIBLE. THE RELOADS WERE REPLACED, BUT THE STAPLER STILL DID NOT WORK. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
01-01-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | ETS*45 ENDO LIN CUT RELOAD WH | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: THE DEVICE WAS RECEIVED WITH NO COMPLAINT INFORMATION. NO FURTHER INFORMATION IS AVAILABLE.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN NEPHRECTOMY, IN THE FIRST ATTEMPT AT FIRING ON THE VESSEL, THE DEVICE PRODUCED UNUSUAL NOISES. THE STAPLES DID NOT FIRE. THE JAWS OF THE DEVICE LOCKED ON TISSUE. THIS RESULTED IN A DAMAGED VEIN AND A TORN-OFF RENAL ARTERY. THERE WAS MAJOR BLEEDING; 500 MILLIMETERS OF BLOOD WERE LOST. A SECOND RELOAD WAS USED, BUT THE DEVICE COULD NO LONGER BE CLOSED. TO RESOLVE THE ISSUE, A VASCULAR SURGEON WAS CALLED IN TO SEW THE VESSELS BACK TOGETHER.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, AFTER INSTALLING THE RELOAD, THE DEVICE WAS INSERTED INTO THE ABDOMINAL CAVITY, OPENED THE JAWS AND APPROACHED, AND WHEN ATTEMPTED TO FIRE, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED, BUT THE DEVICE DID NOT FIRE AGAIN. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, THE THICK PULMONARY PARENCHYMA WAS RESECTED, ON MIDWAY, THOUGH HEAVY, THE HANDLE MANAGED TO FIRE UNTIL THE END. IT WAS ATTEMPTED TO RETURN THE BLACK RETURN KNOB, BUT IT WAS EXTREMELY HARD. SEVERAL ATTEMPTS WERE MADE AND THE BLACK RETURN KNOB MANAGED TO RETURN. IT WAS DISCOVERED THAT THE JAWS COULD NOT BE OPENED. THE RELOAD WAS THEN REMOVED FROM THE TISSUE WHEN PULLED WHILE MOVING THE HAND LITTLE BY LITTLE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | TRI-STAPLE 2.0 | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, THE THICK PULMONARY PARENCHYMA WAS RESECTED, ON MIDWAY, THOUGH HEAVY, THE HANDLE MANAGED TO FIRE UNTIL THE END. IT WAS ATTEMPTED TO RETURN THE BLACK RETURN KNOB, BUT IT WAS EXTREMELY HARD. SEVERAL ATTEMPTS WERE MADE AND THE BLACK RETURN KNOB MANAGED TO RETURN. IT WAS DISCOVERED THAT THE JAWS COULD NOT BE OPENED. THE RELOAD WAS THEN REMOVED FROM THE TISSUE WHEN PULLED WHILE MOVING THE HAND LITTLE BY LITTLE. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF PERFORMING LOBECTOMY, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE ISSUE OCCURRED IN THREE RELOADS. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Injury | ethicon endo-surgery | 16-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AT THE END OF A NEPHRECTOMY PROCEDURE, THE PATIENT GOES OUT TO RECOVERY, IN RECOVERY THE PATIENT BEGINS TO BLEED, DECOMPENSATES AND REINTERVENES THE PATIENT AT 2:00 PM, WHEN OPENING THE PATIENT HAD A LOT OF BLOOD, WHAT DR. DID WAS TAKE THE AORTA TO STOP THE BLEEDING, THEY CHECK THE PATIENT AND REALIZE THAT THE ENTIRE STAPLE LINE WAS OPENED. THE PATIENT WAS UNEMPLOYED AND WAS RESUSCITATED FOR 10 MINUTES, HAD ACTIVE BLEEDING FROM THE ARTERY AND RENAL VEIN. HE HAD A HEMOPERITONEUM OF 4000 C/C AND HIGH CLOTTING TIMES, ONCE STABILIZED THE PATIENT HE WAS RAISED TO THE ICU AND BEGAN TO CONVULSE. THE PATIENT IS CURRENTLY IN THE ICU WITH A RESERVED PROGNOSIS.
04-12-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE THE ECHELON GOT STUCK WHILE FIRING, IMPOSSIBLE FOR THE SURGEON TO RE-OPEN DEVICE. OPENED ANOTHER ECHELON TO RESECT TISSUE, AND TO FREE THE STUCK DEVICE. PROCEDURE DELAYED BY 15 MINUTES. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
04-12-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE THE ECHELON GOT STUCK WHILE FIRING, IMPOSSIBLE FOR THE SURGEON TO RE-OPEN DEVICE. OPENED ANOTHER ECHELON TO RESECT TISSUE, AND TO FREE THE STUCK DEVICE. PROCEDURE DELAYED BY 15 MINUTES. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
27-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) SEGMENTECTOMY, WHILE DETACHING THE LUNG PARENCHYMA, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) SEGMENTECTOMY, WHILE DETACHING THE LUNG PARENCHYMA, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, THE THICK PULMONARY PARENCHYMA WAS RESECTED, ON MIDWAY, THOUGH HEAVY, THE HANDLE MANAGED TO FIRE UNTIL THE END. IT WAS ATTEMPTED TO RETURN THE BLACK RETURN KNOB, BUT IT WAS EXTREMELY HARD. SEVERAL ATTEMPTS WERE MADE AND THE BLACK RETURN KNOB MANAGED TO RETURN. IT WAS DISCOVERED THAT THE JAWS COULD NOT BE OPENED. THE RELOAD WAS THEN REMOVED FROM THE TISSUE WHEN PULLED WHILE MOVING THE HAND LITTLE BY LITTLE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, THE THICK PULMONARY PARENCHYMA WAS RESECTED, ON MIDWAY, THOUGH HEAVY, THE HANDLE MANAGED TO FIRE UNTIL THE END. IT WAS ATTEMPTED TO RETURN THE BLACK RETURN KNOB, BUT IT WAS EXTREMELY HARD. SEVERAL ATTEMPTS WERE MADE AND THE BLACK RETURN KNOB MANAGED TO RETURN. IT WAS DISCOVERED THAT THE JAWS COULD NOT BE OPENED. THE RELOAD WAS THEN REMOVED FROM THE TISSUE WHEN PULLED WHILE MOVING THE HAND LITTLE BY LITTLE. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF PERFORMING LOBECTOMY, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE ISSUE OCCURRED IN THREE RELOADS. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE GREEN BUTTON WAS PRESSED, AND THE HANDLE WAS SQUEEZED, HOWEVER, THE SUTURE RESECTION COULD NOT BE PERFORMED. BOTH HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF PERFORMING LOBECTOMY, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE ISSUE OCCURRED IN THREE RELOADS. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2023 | Injury | ethicon endo-surgery | 16-01-2024 | POWERED PLUS COMPACT 45 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY ¿ WHEN STAPLING. A LOT OF OOZING OCCURRED ¿ BOTH WITH WHITE CARTRIDGE AS WELL AS GREEN CARTRIDGE. HAD TO STOP BLEEDING WITH CLIPS. POSTOP ON DAY 5 PT IS READMITTED DUE TO AIR LEAK ¿ DRAIN IS BEING PLACED AND PT. IS DISCHARGED FROM THE HOSPITAL THE DAY AFTER WITHOUT FURTHER INTERVENTIONS.
13-12-2023 | Injury | ethicon endo-surgery | 16-01-2024 | POWERED PLUS COMPACT 45 | Adverse Event Without Identified Device or Use Problem | A24 | Pneumothorax | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY ¿ WHEN STAPLING. A LOT OF OOZING OCCURRED ¿ BOTH WITH WHITE CARTRIDGE AS WELL AS GREEN CARTRIDGE. HAD TO STOP BLEEDING WITH CLIPS. POSTOP ON DAY 5 PT IS READMITTED DUE TO AIR LEAK ¿ DRAIN IS BEING PLACED AND PT. IS DISCHARGED FROM THE HOSPITAL THE DAY AFTER WITHOUT FURTHER INTERVENTIONS.
13-12-2023 | Injury | ethicon endo-surgery | 16-01-2024 | POWERED PLUS COMPACT 45 | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY ¿ WHEN STAPLING. A LOT OF OOZING OCCURRED ¿ BOTH WITH WHITE CARTRIDGE AS WELL AS GREEN CARTRIDGE. HAD TO STOP BLEEDING WITH CLIPS. POSTOP ON DAY 5 PT IS READMITTED DUE TO AIR LEAK ¿ DRAIN IS BEING PLACED AND PT. IS DISCHARGED FROM THE HOSPITAL THE DAY AFTER WITHOUT FURTHER INTERVENTIONS.
13-12-2023 | Injury | ethicon endo-surgery | 16-01-2024 | POWERED PLUS COMPACT 45 | No Apparent Adverse Event | A25 | Pneumothorax | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY ¿ WHEN STAPLING. A LOT OF OOZING OCCURRED ¿ BOTH WITH WHITE CARTRIDGE AS WELL AS GREEN CARTRIDGE. HAD TO STOP BLEEDING WITH CLIPS. POSTOP ON DAY 5 PT IS READMITTED DUE TO AIR LEAK ¿ DRAIN IS BEING PLACED AND PT. IS DISCHARGED FROM THE HOSPITAL THE DAY AFTER WITHOUT FURTHER INTERVENTIONS.
26-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 16-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LAR (LOW ANTERIOR RESECTION) PROCEDURE, WHILE TRANSECTIONING THE TISSUE, THE STAPLES OF THE TWO RELOADS WERE FIRED INCOMPLETELY. ANOTHER RELOAD WAS FIRED TO RESOLVE THE ISSUE.
26-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC HEPATECTOMY WITH POSSIBLE COLECTOMY, AFTER THE STAPLER FIRED, THE BLACK LEVERS USED TO PULL BACK THE BLADE AND OPEN THE JAWS WOULD NOT OPEN UP TO DISENGAGE THE TISSUE. ANOTHER HANDLE AND RELOAD WERE OPENED AND USED TO RESECT THE TISSUE THAT WAS BEING HELD FROM THE LOCKED STAPLER.
26-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 16-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LAR (LOW ANTERIOR RESECTION) PROCEDURE, WHILE TRANSITIONING THE TISSUE, THE STAPLES OF THE TWO RELOADS WERE FIRED INCOMPLETELY. ANOTHER RELOAD WAS FIRED TO RESOLVE THE ISSUE.
26-11-2023 | Malfunction | ethicon endo-surgery | 16-01-2024 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A APPENDICECTOMY PROCEDURE, THE WRONG APPLICATION OF THE STAPLES OCCURRED DUE TO A MALFUNCTION OF THE DEVICE. THIS EVENT ALLOWED STAPLES TO BE APPLIED ONLY TO THE THREE ROWS OF ONE SIDE OF THE BLADE. THE CUT WAS CARRIED OUT AS USUAL, BUT THE CONTRALATERAL EDGE OF THE ORGAN REMAINED OPEN. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
26-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LOWER ANTERIOR RESECTION, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Injury | ethicon endo-surgery | 16-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT AT THE END OF A NEPHRECTOMY PROCEDURE, THE PATIENT GOES OUT TO RECOVERY, IN RECOVERY THE PATIENT BEGINS TO BLEED, DECOMPENSATES AND REINTERVENES THE PATIENT AT 2:00 PM, WHEN OPENING THE PATIENT HAD A LOT OF BLOOD, WHAT DR. DID WAS TAKE THE AORTA TO STOP THE BLEEDING, THEY CHECK THE PATIENT AND REALIZE THAT THE ENTIRE STAPLE LINE WAS OPENED. THE PATIENT WAS UNEMPLOYED AND WAS RESUSCITATED FOR 10 MINUTES, HAD ACTIVE BLEEDING FROM THE ARTERY AND RENAL VEIN. HE HAD A HEMOPERITONEUM OF 4000 C/C AND HIGH CLOTTING TIMES, ONCE STABILIZED THE PATIENT HE WAS RAISED TO THE ICU AND BEGAN TO CONVULSE. THE PATIENT IS CURRENTLY IN THE ICU WITH A RESERVED PROGNOSIS.
26-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC LAPAROSCOPIC HEPATECTOMY WITH POSSIBLE COLECTOMY, AFTER THE STAPLER FIRED, THE BLACK LEVERS USED TO PULL BACK THE BLADE AND OPEN THE JAWS WOULD NOT OPEN UP TO DISENGAGE THE TISSUE. ANOTHER HANDLE AND RELOAD WERE OPENED AND USED TO RESECT THE TISSUE THAT WAS BEING HELD FROM THE LOCKED STAPLER.
27-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, ON THE FIRST RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE; HOWEVER, THE DEVICE DID NOT FIRE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE AGAIN.
27-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, ON THE FIRST RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE; HOWEVER, THE DEVICE DID NOT FIRE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE AGAIN.
27-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, ON THE FIRST RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE; HOWEVER, THE DEVICE DID NOT FIRE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE AGAIN.
28-11-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AT THE VERY EARLY STAGE, THE STAPLER AND RELOAD WOULD NOT CLOSE FULLY ON THE INTENDED TISSUE. TO RESOLVE THE ISSUE, A SECOND HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED DUE TO THICK TISSUE.
28-11-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, AT THE VERY EARLY STAGE, THE STAPLER AND RELOAD WOULD NOT CLOSE FULLY ON THE INTENDED TISSUE. TO RESOLVE THE ISSUE, A SECOND HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED DUE TO THICK TISSUE.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, AFTER INSTALLING THE RELOAD, THE DEVICE WAS INSERTED INTO THE ABDOMINAL CAVITY, OPENED THE JAWS AND APPROACHED, AND WHEN ATTEMPTED TO FIRE, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED, BUT THE DEVICE DID NOT FIRE AGAIN. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) SEGMENTECTOMY SURGERY, WHEN THE BLOOD VESSEL WAS APPROACHED AND THE NURSE WAS REMOVING THE CARTRIDGES, THE CONNECTION BETWEEN THE CARTRIDGE AND THE ADAPTER WAS DAMAGED. HENCE, AN ARTICULATION WAS NECESSARY WHEN THE 45MM WAS USED; THE CARTRIDGE AND ADAPTER CONNECTION PART WAS TOUCHING AND PRESSED AGAINST THE RIBS. WHEN IT WAS FIRED, THE CONNECTION PART MADE A CRACKING SOUND. WHEN THE CARTRIDGE WAS REMOVED, THE BLACK PLASTIC HAD CRACKED AND COME OFF. AFTER THAT, IT WAS CHANGED TO A 60MM RELOAD, AND ARTICULATION WAS STILL NECESSARY. SIMILARLY, WHEN IT WAS FIRED WHILE PRESSING AGAINST THE RIBS, THE PLASTIC OF THE CONNECTION PART GOT DAMAGED. THE STAPLES WERE FIRED ACCURATELY AND IT WAS ALSO CONFIRMED THAT NO PARTS FELL FROM THE CONNECTION PART INTO THE THORACIC CAVITY. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) SEGMENTECTOMY SURGERY, WHEN THE BLOOD VESSEL WAS APPROACHED AND THE NURSE WAS REMOVING THE CARTRIDGES, THE CONNECTION BETWEEN THE CARTRIDGE AND THE ADAPTER WAS DAMAGED. HENCE, AN ARTICULATION WAS NECESSARY WHEN THE 45MM WAS USED; THE CARTRIDGE AND ADAPTER CONNECTION PART WAS TOUCHING AND PRESSED AGAINST THE RIBS. WHEN IT WAS FIRED, THE CONNECTION PART MADE A CRACKING SOUND. WHEN THE CARTRIDGE WAS REMOVED, THE BLACK PLASTIC HAD CRACKED AND COME OFF. AFTER THAT, IT WAS CHANGED TO A 60MM RELOAD, AND ARTICULATION WAS STILL NECESSARY. SIMILARLY, WHEN IT WAS FIRED WHILE PRESSING AGAINST THE RIBS, THE PLASTIC OF THE CONNECTION PART GOT DAMAGED. THE STAPLES WERE FIRED ACCURATELY AND IT WAS ALSO CONFIRMED THAT NO PARTS FELL FROM THE CONNECTION PART INTO THE THORACIC CAVITY. THERE WAS NO PATIENT INJURY.
19-12-2023 | Injury | intuitive surgical | 16-01-2024 | SUREFORM | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
19-12-2023 | Injury | intuitive surgical | 16-01-2024 | SUREFORM | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: REFER TO H10/H11 FOR FOLLOW-UP INFORMATION.
19-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RECTAL RESECTION, AFTER INSTALLING THE RELOAD, THE DEVICE WAS INSERTED INTO THE ABDOMINAL CAVITY, OPENED THE JAWS AND APPROACHED, AND WHEN ATTEMPTED TO FIRE, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED, BUT THE DEVICE DID NOT FIRE AGAIN. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
30-01-2024 | Malfunction | covidien lp llc north haven | 16-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, THE FIRING WAS PERFORMED BUT THE USER FELT THAT IT STOPPED MIDWAY. THE JAW WAS OPENED AND WAS RELEASED FROM THE TISSUE. THERE SEEMED TO BE NO PROBLEMS WITH THE STAPLE LINE, BUT THE HANDLE AND THE ADAPTER WERE REPLACED. THE FOLLOW FIRING WAS PERFORMED FROM WHERE THE LINE STOPPED. THERE WAS NO PATIENT INJURY.
30-01-2024 | Malfunction | covidien lp llc north haven | 16-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, THE FIRING WAS PERFORMED BUT THE USER FELT THAT IT STOPPED MIDWAY. THE JAW WAS OPENED AND WAS RELEASED FROM THE TISSUE. THERE SEEMED TO BE NO PROBLEMS WITH THE STAPLE LINE, BUT THE HANDLE AND THE ADAPTER WERE REPLACED. THE FOLLOW FIRING WAS PERFORMED FROM WHERE THE LINE STOPPED. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) SEGMENTECTOMY SURGERY, WHEN THE BLOOD V ESSEL WAS APPROACHED AND THE NURSE WAS REMOVING THE CARTRIDGES, THE CONNECTION BETWEEN THE CARTRIDGE AND THE ADAPTER WAS DAMAGED. HENCE, AN ARTICULATION WAS NECESSARY WHEN THE 45MM WAS USED; THE CARTRIDGE AND ADAPTER CONNECTION PART WAS TOUCHING AND PRESSED AGAINST THE RIBS. WHEN IT WAS FIRED, THE CONNECTION PART MADE A CRACKING SOUND. WHEN THE CARTRIDGE WAS REMOVED, THE BLACK PLASTIC HAD CRACKED AND COME OFF. AFTER THAT, IT WAS CHANGED TO A 60MM RELOAD, AND ARTICULATION WAS STILL NECESSARY. SIMILARLY, WHEN IT WAS FIRED WHILE PRESSING AGAINST THE RIBS, THE PLASTIC OF THE CONNECTION PART GOT DAMAGED. THE STAPLES WERE FIRED ACCURATELY AND IT WAS ALSO CONFIRMED THAT NO PARTS FELL FROM THE CONNECTION PART INTO THE THORACIC CAVITY. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) SEGMENTECTOMY SURGERY, WHEN THE BLOOD V ESSEL WAS APPROACHED AND THE NURSE WAS REMOVING THE CARTRIDGES, THE CONNECTION BETWEEN THE CARTRIDGE AND THE ADAPTER WAS DAMAGED. HENCE, AN ARTICULATION WAS NECESSARY WHEN THE 45MM WAS USED; THE CARTRIDGE AND ADAPTER CONNECTION PART WAS TOUCHING AND PRESSED AGAINST THE RIBS. WHEN IT WAS FIRED, THE CONNECTION PART MADE A CRACKING SOUND. WHEN THE CARTRIDGE WAS REMOVED, THE BLACK PLASTIC HAD CRACKED AND COME OFF. AFTER THAT, IT WAS CHANGED TO A 60MM RELOAD, AND ARTICULATION WAS STILL NECESSARY. SIMILARLY, WHEN IT WAS FIRED WHILE PRESSING AGAINST THE RIBS, THE PLASTIC OF THE CONNECTION PART GOT DAMAGED. THE STAPLES WERE FIRED ACCURATELY AND IT WAS ALSO CONFIRMED THAT NO PARTS FELL FROM THE CONNECTION PART INTO THE THORACIC CAVITY. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) SEGMENTECTOMY, WHILE DETACHING THE LUNG PAR ENCHYMA, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, AT THE TIME OF STOPPING THE MESH AT THE UPPER EDGE OF THE PUBIC BONE, WHEN THE DEVICE WAS TRIED TO FIRE, A SNAPPING SOUND WAS HEARD, AND ONLY THE TACK FELL OUT INTO THE ABDOMINAL CAVITY. IT WAS CONTINUED THREE TIMES. IT WAS NOTED THAT THE TACKS WERE NOT ABLE TO PENETRATE THE TISSUE AND WERE NOT ABLE TO HOLD CORRECTLY ONTO THE TISSUE. THE TACKS WERE COLLECTED COMPLETELY BY VISUAL INSPECTION, AND THERE WERE NO LEFTOVERS. A NEW TACKER WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, AT THE TIME OF STOPPING THE MESH AT THE UPPER EDGE OF THE PUBIC BONE, WHEN THE DEVICE WAS TRIED TO FIRE, A SNAPPING SOUND WAS HEARD, AND ONLY THE TACK FELL OUT INTO THE ABDOMINAL CAVITY. IT WAS CONTINUED THREE TIMES. IT WAS NOTED THAT THE TACKS WERE NOT ABLE TO PENETRATE THE TISSUE AND WERE NOT ABLE TO HOLD CORRECTLY ONTO THE TISSUE. THE TACKS WERE COLLECTED COMPLETELY BY VISUAL INSPECTION, AND THERE WERE NO LEFTOVERS. A NEW TACKER WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | ABSORBATACK | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, AT THE TIME OF STOPPING THE MESH AT THE UPPER EDGE OF THE PUBIC BONE, WHEN THE DEVICE WAS TRIED TO FIRE, A SNAPPING SOUND WAS HEARD, AND ONLY THE TACK FELL OUT INTO THE ABDOMINAL CAVITY. IT WAS CONTINUED THREE TIMES. IT WAS NOTED THAT THE TACKS WERE NOT ABLE TO PENETRATE THE TISSUE AND WERE NOT ABLE TO HOLD CORRECTLY ONTO THE TISSUE. THE TACKS WERE COLLECTED COMPLETELY BY VISUAL INSPECTION, AND THERE WERE NO LEFTOVERS. A NEW TACKER WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LOBECTOMY, IN THE LUNG TISSUE, THE LOAD WOULD NOT CUT THE TISSUE AND FIRE STAPLES. THE SURGEON TRIED TO APPLY EXTRA STRENGTH. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE HANDLE MADE A STRANGE NOISE AS IT WAS SQUEEZED WHICH MADE THE SURGEON SUSPECT THAT IT WAS BROKEN. THE SURGEON CHANGED BOTH THE RELOAD AND THE HANDLE. THERE WAS NO PATIENT INJURY.
05-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LOBECTOMY, IN THE LUNG TISSUE, THE LOAD WOULD NOT CUT THE TISSUE AND FIRE STAPLES. THE SURGEON TRIED TO APPLY EXTRA STRENGTH. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE HANDLE MADE A STRANGE NOISE AS IT WAS SQUEEZED WHICH MADE THE SURGEON SUSPECT THAT IT WAS BROKEN. THE SURGEON CHANGED BOTH THE RELOAD AND THE HANDLE. THERE WAS NO PATIENT INJURY.
01-11-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. NO PATIENT WAS INVOLVED.
01-11-2023 | Malfunction | covidien lp llc north haven | 16-01-2024 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. NO PATIENT WAS INVOLVED.
01-11-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. NO PATIENT WAS INVOLVED.
21-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC PARTIAL PULMONARY RESECTION, DURING THE FOURTH FIRING IN THE PULMONARY PARENCHYMA, THE RELOAD WAS CONNECTED TO THE ADAPTER. THE RELOAD HAS NOT BEEN FIRED BUT ALREADY SHOWED AS ¿0¿ OR FIRED IN THE DISPLAY. WHEN AN ATTEMPT WAS MADE TO REPLACE THE RELOAD, THE RELOAD COULD NOT BE REMOVED FROM THE ADAPTER. ANOTHER ADAPTER AND RELOAD WERE USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WERE DEFORMED ANTI-FRICTION NYLON PADS ON I-BEAM DUE TO THICK TISSUE.
21-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC PARTIAL PULMONARY RESECTION, DURING THE FOURTH FIRING IN THE PULMONARY PARENCHYMA, THE RELOAD WAS CONNECTED TO THE ADAPTER. THE RELOAD HAS NOT BEEN FIRED BUT ALREADY SHOWED AS ¿0¿ OR FIRED IN THE DISPLAY. WHEN AN ATTEMPT WAS MADE TO REPLACE THE RELOAD, THE RELOAD COULD NOT BE REMOVED FROM THE ADAPTER. ANOTHER ADAPTER AND RELOAD WERE USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WERE DEFORMED ANTI-FRICTION NYLON PADS ON I-BEAM DUE TO THICK TISSUE.
21-12-2023 | Malfunction | us surgical puerto rico | 16-01-2024 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC PARTIAL PULMONARY RESECTION, DURING THE FOURTH FIRING IN THE PULMONARY PARENCHYMA, THE RELOAD WAS CONNECTED TO THE ADAPTER. THE RELOAD HAS NOT BEEN FIRED BUT ALREADY SHOWED AS ¿0¿ OR FIRED IN THE DISPLAY. WHEN AN ATTEMPT WAS MADE TO REPLACE THE RELOAD, THE RELOAD COULD NOT BE REMOVED FROM THE ADAPTER. ANOTHER ADAPTER AND RELOAD WERE USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WERE DEFORMED ANTI-FRICTION NYLON PADS ON I-BEAM DUE TO THICK TISSUE.
21-12-2023 | Injury | covidien | 16-01-2024 | COVIDIEN IDRIVE ULTRA POWERED STAPLING DEVICE |  |  | Device Embedded In Tissue or Plaque | Event Description: PATIENT UNDERGOING LAPAROSCOPIC ILEAL PANCREATIC DIVERSION WITH DUODENAL SWITCH FOR TREATMENT OF MORBID OBESITY. DURING PROCEDURE, THE COVIDIEN IDRIVE ULTRA POWERED STAPLING DEVICE WOULD NOT UNLOCK WITH DRIVER OR WITH MANUAL SCREWDRIVER. STAPLE LOAD HAD TO BE CUT OUT OF PATIENT'S STOMACH AND THE PROCEDURE CONVERTED TO ROUX-EN-Y GASTRIC BYPASS.
21-12-2023 | Injury | covidien | 16-01-2024 | COVIDIEN IDRIVE ULTRA POWERED STAPLING DEVICE |  |  | Unspecified Tissue Injury | Event Description: PATIENT UNDERGOING LAPAROSCOPIC ILEAL PANCREATIC DIVERSION WITH DUODENAL SWITCH FOR TREATMENT OF MORBID OBESITY. DURING PROCEDURE, THE COVIDIEN IDRIVE ULTRA POWERED STAPLING DEVICE WOULD NOT UNLOCK WITH DRIVER OR WITH MANUAL SCREWDRIVER. STAPLE LOAD HAD TO BE CUT OUT OF PATIENT'S STOMACH AND THE PROCEDURE CONVERTED TO ROUX-EN-Y GASTRIC BYPASS.
18-12-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | BLUE RELOAD FOR ECHELON 60 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT CLOSE THE JAW BEFORE INSERT INTO TROCAR. CHANGED THE NEW ONE TO COMPLETE SURGERY. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, AFTER FIRED, NOTED THE METAL PART OF CARTRIDGE WAS DEFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | ECHELON 3000 60MM STANDARD | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SCRUB TECH COULD NOT GET THE DEVICE LOADED. NO FURTHER INFORMATION WAS PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
20-11-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | GST WHITE RELOAD, 45MM, 6 ROW | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT CLOSE THE JAW. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: THE STAPLER MISFIRED ON THE LAST FIRING OF A SLEEVE GASTRECTOMY. THE STAPLER FIRED AND IT SOUNDED LIKE THE BATTERY WAS GOING TO DIE. AFTER THE FIRING SEQUENCE COMPLETED, IT DID NOT CAPTURE ALL OF THE TISSUE AND THERE WAS A MASSIVE HOLE LEFT IN THE STOMACH. SURGEON HAD SO OVERSEW THE STAPLE LINE. WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? YES. IF YES, NUMBER OF MINUTES: 20. ACTION TAKEN WHEN EVENT OCCURRED? OVERSEW THE STAPLE LINE AND OPEN A NEW STAPLER, WAS PROCEDURE SUCCESSFULLY COMPLETED? YES. WERE FRAGMENTS GENERATED? UNKNOWN. IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? YES. PATIENT STATUS/ OUTCOME / CONSEQUENCES: UNKNOWN. WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN. IS THE PATIENT PART OF A CLINICAL STUDY: UNKNOWN. (B)(4). DEVICE PROPERTY OF: NONE. DEVICE IN POSSESSION OF: NONE. BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST: TRUE.
01-01-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: THE STAPLER MISFIRED ON THE LAST FIRING OF A SLEEVE GASTRECTOMY. THE STAPLER FIRED AND IT SOUNDED LIKE THE BATTERY WAS GOING TO DIE. AFTER THE FIRING SEQUENCE COMPLETED, IT DID NOT CAPTURE ALL OF THE TISSUE AND THERE WAS A MASSIVE HOLE LEFT IN THE STOMACH. SURGEON HAD SO OVERSEW THE STAPLE LINE. WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? YES. IF YES, NUMBER OF MINUTES: 20. ACTION TAKEN WHEN EVENT OCCURRED? OVERSEW THE STAPLE LINE AND OPEN A NEW STAPLER, WAS PROCEDURE SUCCESSFULLY COMPLETED? YES. WERE FRAGMENTS GENERATED? UNKNOWN. IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? YES. PATIENT STATUS/ OUTCOME / CONSEQUENCES: UNKNOWN. WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN. IS THE PATIENT PART OF A CLINICAL STUDY: UNKNOWN. (B)(4). DEVICE PROPERTY OF: NONE. DEVICE IN POSSESSION OF: NONE. BY CHECKING THIS BOX I CERTIFY THAT ALL INFORMATION THAT ARE KNOWN/AVAILABLE HAS BEEN DISCLOSED. IF ANY NEW INFORMATION WILL BE MADE AVAILABLE, THE ADDITIONAL INFORMATION WILL BE SUBMITTED THROUGH CST: TRUE.
16-11-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | 55MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OSTOMY CLOSURE SURGERY, THE DEVICE COULD NOT FIRE. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE COULDN'T MAKE FIRING ACTIONS WHILE OPERATION. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
19-10-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | ECHELON CONTOUR W/ BLUE RELOAD | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE SURGEON AND THE RESIDENT WAS UNSURE IF THE DEVICE FIRED CORRECTLY. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
20-10-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | CONTOUR CURVED CUTTER STAPLER | Appropriate Term/Code Not Available |  | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER TRIPPED WITHOUT LOWERING THE SECURITY LATCH , THE STAPLES DID NOT CLOSE, BUT DID NOT CUT, CAUSING NO DAMAGE . ONLY A FEW CLIPS FELL INTO THE CAVITY. THE SURGERY WAS DELAYED. THERE WAS NO PATIENT CONSEQUENCE.
17-10-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | PWRD ECH FLEX 60MM 28CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2023 | Injury | intuitive surgical | 17-01-2024 | SUREFORM | Misfire | A050502 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY SURGICAL PROCEDURE, WHILE USING THE SUREFORM 60 STAPLER INSTRUMENT WITH A BLUE SUREFORM 60 RELOAD, THERE WAS BLEEDING ON THE RESECTED END OF THE SMALL BOWEL ANASTOMOSIS. THIS WAS THE FIRST OCCURRENCE OF BLEEDING, BUT WAS RESOLVED WITH THE PLACEMENT OF CLIPS. THE SECOND OCCURRENCE WAS AFTER THE SIDE-TO-SIDE ANASTOMOSIS; THERE WAS A SMALL AMOUNT OF BLEEDING WHICH WAS RESOLVED USING A COUPLE OF CLIPS. HOWEVER, WHEN THE SURGEON VISUALIZED BOTH ANASTOMOSES, THERE WAS STILL SOME BLEEDING. THE SURGEON MADE THE DECISION TO CONVERT TO OPEN SURGERY TO BETTER VISUALIZE THE ANASTOMOSES. THE SURGEON DID NOT WANT TO PROCEED WITH THE ROBOT OR A THIRD PARTY LAPAROSCOPIC STAPLER. THE OPEN SURGERY WAS COMPLETED. THE PATIENT DID NOT REQUIRE ANY BLOOD TRANSFUSIONS AND HAS HAD NO POST-OPERATIVE COMPLICATIONS SINCE. HISTORICALLY, THE SURGEON DID NOT PREVIOUSLY USE CLIPS AFTER STAPLING. HOWEVER, THE SURGEON HAS SINCE CHANGED PRACTICES AND USES CLIPS AFTER STAPLING FOR ALL THE RIGHT HEMICOLECTOMY CASES, WHETHER THE PATIENT IS BLEEDING FROM THE STAPLE LINE OR NOT. REFER TO MEDWATCH REPORT (MDR) WITH MFR. REPORT #2955842-2024-10122 (PATIENT IDENTIFIER # (B)(6) FOR ADDITIONAL INFORMATION REGARDING THE SECOND OCCURRENCE OF BLEEDING AFTER USING THE SUREFORM 60 STAPLER INSTRUMENT WITH A BLUE SUREFORM 60 RELOAD.
19-12-2023 | Injury | intuitive surgical | 17-01-2024 | SUREFORM | Misfire | A050502 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY SURGICAL PROCEDURE, WHILE USING THE SUREFORM 60 STAPLER INSTRUMENT WITH A BLUE SUREFORM 60 RELOAD, THERE WAS BLEEDING ON THE RESECTED END OF THE SMALL BOWEL ANASTOMOSIS. THIS WAS THE FIRST OCCURRENCE OF BLEEDING, BUT WAS RESOLVED WITH THE PLACEMENT OF CLIPS. THE SECOND OCCURRENCE WAS AFTER THE SIDE-TO-SIDE ANASTOMOSIS; THERE WAS A SMALL AMOUNT OF BLEEDING WHICH WAS RESOLVED USING A COUPLE OF CLIPS. HOWEVER, WHEN THE SURGEON VISUALIZED BOTH ANASTOMOSES, THERE WAS STILL SOME BLEEDING. THE SURGEON MADE THE DECISION TO CONVERT TO OPEN SURGERY TO BETTER VISUALIZE THE ANASTOMOSES. THE SURGEON DID NOT WANT TO PROCEED WITH THE ROBOT OR A THIRD PARTY LAPAROSCOPIC STAPLER. THE OPEN SURGERY WAS COMPLETED. THE PATIENT DID NOT REQUIRE ANY BLOOD TRANSFUSIONS AND HAS HAD NO POST-OPERATIVE COMPLICATIONS SINCE. HISTORICALLY, THE SURGEON DID NOT PREVIOUSLY USE CLIPS AFTER STAPLING. HOWEVER, THE SURGEON HAS SINCE CHANGED PRACTICES AND USES CLIPS AFTER STAPLING FOR ALL THE RIGHT HEMICOLECTOMY CASES, WHETHER THE PATIENT IS BLEEDING FROM THE STAPLE LINE OR NOT. REFER TO MEDWATCH REPORT (MDR) WITH MFR. REPORT #2955842-2024-10122 (PATIENT IDENTIFIER # (B)(6) FOR ADDITIONAL INFORMATION REGARDING THE SECOND OCCURRENCE OF BLEEDING AFTER USING THE SUREFORM 60 STAPLER INSTRUMENT WITH A BLUE SUREFORM 60 RELOAD.
03-11-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | FLEX 60 ARTICULATING | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE MISFIRED AND THE STAPLER JAMMED. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2024 | Malfunction | covidien lp llc north haven | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY PROCEDURE, WHILE DETACHING THE BLOOD VESSELS, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEREFORE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THE ISSUE OCCURRED IN THREE RELOADS. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY, BUT AT THAT STAGE SURGEON WAS NOT FIRING THE STAPLER BEFORE USING IT TO THE PATIENT, THE ARTICULATING KNOB OF THE TWO HANDLES BROKE OFF AND THIS HAPPENED IN JUST BETWEEN 5-8 FIRING. ANOTHER STAPLER WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY, BUT AT THAT STAGE SURGEON WAS NOT FIRING THE STAPLER BEFORE USING IT TO THE PATIENT, THE ARTICULATING KNOB OF THE TWO HANDLES BROKE OFF AND THIS HAPPENED IN JUST BETWEEN 5-8 FIRING. ANOTHER STAPLER WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY, BUT AT THAT STAGE SURGEON WAS NOT FIRING THE STAPLER BEFORE USING IT TO THE PATIENT, THE ARTICULATING KNOB OF THE TWO HANDLES BROKE OFF AND THIS HAPPENED IN JUST BETWEEN 5-8 FIRING. ANOTHER STAPLER WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HEMICOLECTOMY, BUT AT THAT STAGE SURGEON WAS NOT FIRING THE STAPLER BEFORE USING IT TO THE PATIENT, THE ARTICULATING KNOB OF THE TWO HANDLES BROKE OFF AND THIS HAPPENED IN JUST BETWEEN 5-8 FIRING. ANOTHER STAPLER WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2024 | Malfunction | covidien lp llc north haven | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY PROCEDURE, WHILE DETACHING THE BLOOD VESSELS, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEREFORE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THE ISSUE OCCURRED IN THREE RELOADS. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC GASTRECTOMY, WHEN ON GASTRIC TISSUE, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-10-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE THAT THE DEVICE WAS UNABLE TO FIRE. NO FURTHER INFORMATION WAS PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT.
21-03-2024 | Malfunction | covidien lp llc north haven | 17-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORASCOPIC SEGMENTECTOMY, DURING SEGMENTAL RESECTION, THE RELOAD COULD NOT BE CLOSED AFTER ITS INSTALLATION ON A MANUAL STAPLER. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC APPENDECTOMY, THE RELOAD WAS ATTACHED, AND OPENING/CLOSING WAS CHECKED. THE TISSUE WAS APPROACHED, THE DEVICE WAS CLOSED, THE GREEN BUTTON WAS PRESSED, ATTEMPTED TO FIRE, BUT IT DID NOT ADVANCE, AND FIRING COULD NOT BE PERFORMED. WHEN THE LCD WAS CHECKED, THE RELOAD WAS SHOWN TO BE 0. A NEW RELOAD WAS USED WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC APPENDECTOMY, THE RELOAD WAS ATTACHED, AND OPENING/CLOSING WAS CHECKED. THE TISSUE WAS APPROACHED, THE DEVICE WAS CLOSED, THE GREEN BUTTON WAS PRESSED, ATTEMPTED TO FIRE, BUT IT DID NOT ADVANCE, AND FIRING COULD NOT BE PERFORMED. WHEN THE LCD WAS CHECKED, THE RELOAD WAS SHOWN TO BE 0. A NEW RELOAD WAS USED WITHOUT ANY PROBLEM. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | ENDO GIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY, THE RELOAD WAS FIXED TO THE HANDLE BUT THERE WAS A GAP OF SEVERAL MILLIMETERS SO THE PHYSICIAN WAS NOT ABLE TO COMPLETELY CLOSE THE JAWS AND UNABLE TO INSERT THE DEVICE INTO THE TROCAR. THE RELOAD AND HANDLE WERE CHANGED TO FIX THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING APPENDECTOMY, THE RELOAD WAS FIXED TO THE HANDLE BUT THERE WAS A GAP OF SEVERAL MILLIMETERS SO THE PHYSICIAN WAS NOT ABLE TO COMPLETELY CLOSE THE JAWS AND UNABLE TO INSERT THE DEVICE INTO THE TROCAR. THE RELOAD AND HANDLE WERE CHANGED TO FIX THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE ANVIL WAS NOT FIX AS IT USED TO BE AND THE NURSE WAS NOT ABLE TO CLOSE AND OPEN PROPERLY THE RELOAD BEFORE HANDING IT OVER TO THE SURGEON. THE RELOAD WAS UNSTABLE AND WAS IMPOSSIBLE TO PUT THE TROUGH THE TROCAR. THE RELOAD AND HANDLE WERE CHANGED TO FIX THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE ANVIL WAS NOT FIX AS IT USED TO BE AND THE NURSE WAS NOT ABLE TO CLOSE AND OPEN PROPERLY THE RELOAD BEFORE HANDING IT OVER TO THE SURGEON. THE RELOAD WAS UNSTABLE AND WAS IMPOSSIBLE TO PUT THE TROUGH THE TROCAR. THE RELOAD AND HANDLE WERE CHANGED TO FIX THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE ANVIL WAS NOT FIX AS IT USED TO BE AND THE NURSE WAS NOT ABLE TO CLOSE AND OPEN PROPERLY THE RELOAD BEFORE HANDING IT OVER TO THE SURGEON. THE RELOAD WAS UNSTABLE AND WAS IMPOSSIBLE TO PUT THE TROUGH THE TROCAR. THE RELOAD AND HANDLE WERE CHANGED TO FIX THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
19-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE ANVIL WAS NOT FIX AS IT USED TO BE AND THE NURSE WAS NOT ABLE TO CLOSE AND OPEN PROPERLY THE RELOAD BEFORE HANDING IT OVER TO THE SURGEON. THE RELOAD WAS UNSTABLE AND WAS IMPOSSIBLE TO PUT THE TROUGH THE TROCAR. THE RELOAD AND HANDLE WERE CHANGED TO FIX THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
18-10-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE INITIAL FIRING WAS NORMAL, THE DEVICE WAS UNABLE TO FIRE FOR THE SECOND FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
28-12-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL SWITCH (SADIS), WHEN THE SECOND RELOAD WAS FULLY FIRED ON THE DISTAL STOMACH, IT WAS LOCKED DOWN ON THE DISTAL STOMACH. TO RESOLVE THE ISSUE, A RETRACTION TOOL HAD TO BE USED TO OPEN THE JAWS OF THE RELOAD. IT WAS NOTED THAT THE STAPLE LINE LOOKED GOOD. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-10-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE INITIAL FIRING WAS NORMAL, THE DEVICE WAS UNABLE TO FIRE FOR THE SECOND FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
28-12-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | UNKNOWN SIGNIA EGIA ADAPTER | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL SWITCH (SADIS), WHEN THE SECOND RELOAD WAS FULLY FIRED ON THE DISTAL STOMACH, IT WAS LOCKED DOWN ON THE DISTAL STOMACH. TO RESOLVE THE ISSUE, A RETRACTION TOOL HAD TO BE USED TO OPEN THE JAWS OF THE RELOAD. IT WAS NOTED THAT THE STAPLE LINE LOOKED GOOD. A MANUAL HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SINGLE ANASTOMOSIS DUODENAL SWITCH (SADIS), WHEN THE SECOND RELOAD WAS FULLY FIRED ON THE DISTAL STOMACH, IT WAS LOCKED DOWN ON THE DISTAL STOMACH. TO RESOLVE THE ISSUE, A RETRACTION TOOL HAD TO BE USED TO OPEN THE JAWS OF THE RELOAD. IT WAS NOTED THAT THE STAPLE LINE LOOKED GOOD. A MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-10-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Difficult to Remove | A150207 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Difficult to Remove | A150207 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Firing Problem | A0505 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Firing Problem | A0505 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
27-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC RADICAL GASTRECTOMY, ON PREPARING TO CUT GASTRIC BODY LESIONS, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WAS NO PATIENT INJURY.
21-08-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WAS CLOSED, SECURED, AND THEN IT WOULD NOT RELEASE. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED.
04-09-2023 | Malfunction | ethicon endo-surgery | 17-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE INSTRUMENT COULD NOT BE TRIGGERED. SURGERY WAS COMPLETED WITH NEW INSTRUMENT. NO PATIENT HARM. NO SURGERY DELAY.
25-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC ROBOT-ASSISTED RADICAL CYSTECTOMY, WHILE PROCESSING TISSUE, WHEN AN ATTEMPT WAS MADE TO PERFORM FIRING ON THE TISSUE, FIRING COULD NOT BE PERFORMED. THE SURGEON COULD NOT SQUEEZE THE HANDLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2024 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY PROCEDURE, WHILE DETACHING THE BLOOD VESSELS, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEREFORE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THE ISSUE OCCURRED IN THREE RELOADS. THERE WAS NO PATIENT INJURY.
01-01-2024 | Malfunction | covidien lp llc north haven | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY PROCEDURE, WHILE DETACHING THE BLOOD VESSELS, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THEREFORE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THE ISSUE OCCURRED IN THREE RELOADS. THERE WAS NO PATIENT INJURY.
20-11-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PYLORUS PRESERVING PANCREATICODUODENECTOMY PROCEDURE, THE DEVICE DID NOT FIRE. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE.
21-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE, AND THE LOAD STUCK WHEN UNLOADING. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE, AND THE LOAD STUCK WHEN UNLOADING. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2022 | Malfunction | covidien lp llc north haven | 17-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, THE RELOAD DID NOT OPEN. THE SURGEON STARTED THE STAPLING PROCEDURE AWAY FROM THE PATIENT IN ORDER TO OPEN THE STAPLER. TO COMPLETE THE CASE, ANOTHER RELOAD WAS USED WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
27-10-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DEVICE FOLLOW-UP, THE DEVICE HAD NO SIGNAL AND COULD NOT BE CHARGED. NO PATIENT WAS INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. THE BATTERY VENTS WERE NOTED TO BE OPEN.
27-10-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DEVICE FOLLOW-UP, THE DEVICE HAD NO SIGNAL AND COULD NOT BE CHARGED. NO PATIENT WAS INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. THE BATTERY VENTS WERE NOTED TO BE OPEN.
21-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE, AND THE LOAD STUCK WHEN UNLOADING. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER DID NOT FIRE, AND THE LOAD STUCK WHEN UNLOADING. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | SIGNIA | Unraveled Material | A040508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, WHILE PREPARING FOR THE OPERATION, THE USER TRIED SETTING UP THE POWER HANDLE AND POWER SHELL, BUT IT WAS DIFFICULT TO CLOSE THE SHELL. THE CLAWS OF THE POWER SHELL WERE STUCK INSIDE AND WOULD NOT CLOSE. TO RESOLVE THE ISSUE, ANOTHER SHELL WAS USED. THERE WAS NO PATIENT INVOLVEMENT.
20-09-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, AN ERROR OCCURRED WHEN ATTEMPTING TO USE THE HANDLE AFTER RECOGNIZING THE CLAMSHELL, RESULTING IN AN ERROR MESSAGE INDICATING A CIRCLE INTERSECTING A PROHIBITION LINE. INITIALLY, THE ISSUE WAS ADDRESSED BY ATTEMPTING TO CHANGE THE CLAMSHELL, BUT THIS DID NOT RESOLVE THE PROBLEM. AFTERWARDS, A DIFFERENT HANDLE WAS OBTAINED, AND WHEN PAIRED WITH A NEWCLAMSHELL, IT SUCCESSFULLY RESOLVED THE ISSUE, ENABLING THE SURGERY TO PROCEED. THERE WAS NO PATIENT INVOLVEMENT. PLI 10- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THE CAUSE OF THE REPORTED CONDITION CAN BE TRACED TO FPGA RESET/REPROGRAM FAILURES.
20-09-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, AN ERROR OCCURRED WHEN ATTEMPTING TO USE THE HANDLE AFTER RECOGNIZING THE CLAMSHELL, RESULTING IN AN ERROR MESSAGE INDICATING A CIRCLE INTERSECTING A PROHIBITION LINE. INITIALLY, THE ISSUE WAS ADDRESSED BY ATTEMPTING TO CHANGE THE CLAMSHELL, BUT THIS DID NOT RESOLVE THE PROBLEM. AFTERWARDS, A DIFFERENT HANDLE WAS OBTAINED, AND WHEN PAIRED WITH A NEWCLAMSHELL, IT SUCCESSFULLY RESOLVED THE ISSUE, ENABLING THE SURGERY TO PROCEED. THERE WAS NO PATIENT INVOLVEMENT. PLI 10- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THE CAUSE OF THE REPORTED CONDITION CAN BE TRACED TO FPGA RESET/REPROGRAM FAILURES.
18-12-2023 | Injury | covidien lp llc north haven | 17-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG VOLUME REDUCTION SURGERY, UPON LUNG WEDGE RESECTION, THE JAWS WOULD NOT OPEN AND WERE LOCKED ON THE TISSUE. IT WAS REPORTED THAT THE LAPAROSCOPIC PROCEDURE WAS CONVERTED INTO AN OPEN SURGERY DUE TO THE DEVICE PROBLEM, AND THE INCISION WAS EXTENDED BY MORE THAN ONE INCH, WHICH LED TO AN EXTENDED SURGICAL TIME OF MORE THAN 30 MINUTES. AN EXTRA PIECE OF LUNG TISSUE HAD TO BE STAPLED OFF IN ORDER TO REMOVE THE STAPLER, AND ANOTHER STAPLER WAS USED TO STAPLE THE STUCK RELOAD AND ADAPTER OUT.
18-12-2023 | Injury | covidien lp llc north haven | 17-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG VOLUME REDUCTION SURGERY, UPON LUNG WEDGE RESECTION, THE JAWS WOULD NOT OPEN AND WERE LOCKED ON THE TISSUE. IT WAS REPORTED THAT THE LAPAROSCOPIC PROCEDURE WAS CONVERTED INTO AN OPEN SURGERY DUE TO THE DEVICE PROBLEM, AND THE INCISION WAS EXTENDED BY MORE THAN ONE INCH, WHICH LED TO AN EXTENDED SURGICAL TIME OF MORE THAN 30 MINUTES. AN EXTRA PIECE OF LUNG TISSUE HAD TO BE STAPLED OFF IN ORDER TO REMOVE THE STAPLER, AND ANOTHER STAPLER WAS USED TO STAPLE THE STUCK RELOAD AND ADAPTER OUT.
18-12-2023 | Injury | covidien lp llc north haven | 17-01-2024 | SIGNIA | Entrapment of Device | A150208 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG VOLUME REDUCTION SURGERY, UPON LUNG WEDGE RESECTION, THE JAWS WOULD NOT OPEN AND WERE LOCKED ON THE TISSUE. IT WAS REPORTED THAT THE LAPAROSCOPIC PROCEDURE WAS CONVERTED INTO AN OPEN SURGERY DUE TO THE DEVICE PROBLEM, AND THE INCISION WAS EXTENDED BY MORE THAN ONE INCH, WHICH LED TO AN EXTENDED SURGICAL TIME OF MORE THAN 30 MINUTES. AN EXTRA PIECE OF LUNG TISSUE HAD TO BE STAPLED OFF IN ORDER TO REMOVE THE STAPLER, AND ANOTHER STAPLER WAS USED TO STAPLE THE STUCK RELOAD AND ADAPTER OUT.
18-12-2023 | Injury | covidien lp llc north haven | 17-01-2024 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG VOLUME REDUCTION SURGERY, UPON LUNG WEDGE RESECTION, THE JAWS WOULD NOT OPEN AND WERE LOCKED ON THE TISSUE. IT WAS REPORTED THAT THE LAPAROSCOPIC PROCEDURE WAS CONVERTED INTO AN OPEN SURGERY DUE TO THE DEVICE PROBLEM, AND THE INCISION WAS EXTENDED BY MORE THAN ONE INCH, WHICH LED TO AN EXTENDED SURGICAL TIME OF MORE THAN 30 MINUTES. AN EXTRA PIECE OF LUNG TISSUE HAD TO BE STAPLED OFF IN ORDER TO REMOVE THE STAPLER, AND ANOTHER STAPLER WAS USED TO STAPLE THE STUCK RELOAD AND ADAPTER OUT.
18-12-2023 | Injury | covidien lp llc north haven | 17-01-2024 | SIGNIA | Entrapment of Device | A150208 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG VOLUME REDUCTION SURGERY, UPON LUNG WEDGE RESECTION, THE JAWS WOULD NOT OPEN AND WERE LOCKED ON THE TISSUE. IT WAS REPORTED THAT THE LAPAROSCOPIC PROCEDURE WAS CONVERTED INTO AN OPEN SURGERY DUE TO THE DEVICE PROBLEM, AND THE INCISION WAS EXTENDED BY MORE THAN ONE INCH, WHICH LED TO AN EXTENDED SURGICAL TIME OF MORE THAN 30 MINUTES. AN EXTRA PIECE OF LUNG TISSUE HAD TO BE STAPLED OFF IN ORDER TO REMOVE THE STAPLER, AND ANOTHER STAPLER WAS USED TO STAPLE THE STUCK RELOAD AND ADAPTER OUT.
18-12-2023 | Injury | covidien lp llc north haven | 17-01-2024 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG VOLUME REDUCTION SURGERY, UPON LUNG WEDGE RESECTION, THE JAWS WOULD NOT OPEN AND WERE LOCKED ON THE TISSUE. IT WAS REPORTED THAT THE LAPAROSCOPIC PROCEDURE WAS CONVERTED INTO AN OPEN SURGERY DUE TO THE DEVICE PROBLEM, AND THE INCISION WAS EXTENDED BY MORE THAN ONE INCH, WHICH LED TO AN EXTENDED SURGICAL TIME OF MORE THAN 30 MINUTES. AN EXTRA PIECE OF LUNG TISSUE HAD TO BE STAPLED OFF IN ORDER TO REMOVE THE STAPLER, AND ANOTHER STAPLER WAS USED TO STAPLE THE STUCK RELOAD AND ADAPTER OUT.
18-12-2023 | Injury | covidien lp llc north haven | 17-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG VOLUME REDUCTION SURGERY, UPON LUNG WEDGE RESECTION, THE JAWS WOULD NOT OPEN AND WERE LOCKED ON THE TISSUE. IT WAS REPORTED THAT THE LAPAROSCOPIC PROCEDURE WAS CONVERTED INTO AN OPEN SURGERY DUE TO THE DEVICE PROBLEM, AND THE INCISION WAS EXTENDED BY MORE THAN ONE INCH, WHICH LED TO AN EXTENDED SURGICAL TIME OF MORE THAN 30 MINUTES. AN EXTRA PIECE OF LUNG TISSUE HAD TO BE STAPLED OFF IN ORDER TO REMOVE THE STAPLER, AND ANOTHER STAPLER WAS USED TO STAPLE THE STUCK RELOAD AND ADAPTER OUT.
18-12-2023 | Injury | covidien lp llc north haven | 17-01-2024 | SIGNIA | Difficult to Open or Close | A051104 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG VOLUME REDUCTION SURGERY, UPON LUNG WEDGE RESECTION, THE JAWS WOULD NOT OPEN AND WERE LOCKED ON THE TISSUE. IT WAS REPORTED THAT THE LAPAROSCOPIC PROCEDURE WAS CONVERTED INTO AN OPEN SURGERY DUE TO THE DEVICE PROBLEM, AND THE INCISION WAS EXTENDED BY MORE THAN ONE INCH, WHICH LED TO AN EXTENDED SURGICAL TIME OF MORE THAN 30 MINUTES. AN EXTRA PIECE OF LUNG TISSUE HAD TO BE STAPLED OFF IN ORDER TO REMOVE THE STAPLER, AND ANOTHER STAPLER WAS USED TO STAPLE THE STUCK RELOAD AND ADAPTER OUT.
28-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLORECTAL SURGERY, IN DETACHING THE COLON TISSUE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER 60-MILLIMETER RELOAD WAS USED WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
28-12-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, IN STAPLING THE DUODENAL STUM, THE DEVICE FIRED BUT THE STAPLES WAS NOT FULLY CLOSED. TO RESOLVE THE ISSUE, ONLY THE RELOAD WAS REPLACED.
24-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW HANDLE WAS USED WITH THE SAME RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
30-11-2023 | Malfunction | covidien lp llc north haven | 17-01-2024 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC BYPASS GASTROPLASTY, DURING THE SECOND STAPLE OF THE STOMACH, AFTER STAPLING WHEN REMOVING THE LOAD, SURGEON EXPERIENCED DIFFICULTY UNSTICKING FROM THE TISSUE. THE SURGEON MANEUVERED AND WAS ABLE TO REMOVE THE JAWS THAT WERE LOCKED ON TISSUE. THERE WAS NO PATIENT INJURY.
29-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PULMONARY LOBECTOMY, IN DETACHING THE LUNG LOBE, THE USER COULD PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE SAME RELOAD. THERE WAS NO PATIENT INJURY.
28-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, IN DETACHING THE RIGHT HEPATIC PEDICLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. THE ISSUE OCCURRED ON TWO DEVICES. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
28-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, IN DETACHING THE RIGHT HEPATIC PEDICLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. THE ISSUE OCCURRED ON TWO DEVICES. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
28-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, IN DETACHING THE RIGHT HEPATIC PEDICLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. THE ISSUE OCCURRED ON TWO DEVICES. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
28-12-2023 | Malfunction | us surgical puerto rico | 17-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMIHEPATECTOMY, IN DETACHING THE RIGHT HEPATIC PEDICLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. THE ISSUE OCCURRED ON TWO DEVICES. TO RESOLVE THE ISSUE, ANOTHER HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
15-12-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. STAPLES MALFORMED. CHANGED ANOTHER THREE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. STAPLES MALFORMED. CHANGED ANOTHER THREE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-08-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WAS CLOSED, SECURED, AND THEN IT WOULD NOT RELEASE. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED.
28-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 18-01-2024 | SORBAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING LAPAROSCOPIC INGUINAL HERNIA REPAIR PROCEDURE, SORBAFIX ENHANCED FIXATION DEVICE WAS USED TO FIXATE THE MESH AT BILATERAL GROIN. IT WAS REPORTED THAT AFTER 3 TO 4 FASTENERS DEPLOYED SUCCESSFULLY INTO THE MESH, THE FASTENER DID NOT DEPLOY IN NEXT ATTEMPT AND ANOTHER ATTEMPT WAS MADE IN WHICH THE FASTENER DEPLOYED BUT DID NOT FIXATE THE MESH. IT WAS REPORTED THAT THE PROCEDURE WAS COMPLETED WITH ANOTHER DEVICE. THERE WAS NO PATIENT HARM OR INJURY.
20-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, BLEEDING FROM THE STAPLE LINE WOULD NOT STOP.
21-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE RELOAD WAS CONNECTED TO THE STAPLER. OPENING AND CLOSING OF THE JAWS WERE CHECKED BUT THE JAWS DID NOT CLOSE UNTIL THE END. WHEN INSERTING THE STAPLER INTO THE ABDOMINAL CAVITY, COULD NOT BE INSERTED THROUGH THE TROCAR. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE RELOAD WAS CONNECTED TO THE STAPLER. OPENING AND CLOSING OF THE JAWS WERE CHECKED BUT THE JAWS DID NOT CLOSE UNTIL THE END. WHEN INSERTING THE STAPLER INTO THE ABDOMINAL CAVITY, COULD NOT BE INSERTED THROUGH THE TROCAR. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC PROXIMAL GASTRECTOMY WITH ROBOT SUPPORT, THE RELOAD DID NOT FULLY CLOSE DURING THE OPENING AND CLOSING OF THE DEVICE. THE DEVICE COULD NOT BE INSERTED INTO THE ABDOMINAL CAVITY THROUGH THE TROCAR. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC PROXIMAL GASTRECTOMY WITH ROBOT SUPPORT, THE RELOAD DID NOT FULLY CLOSE DURING THE OPENING AND CLOSING OF THE DEVICE. THE DEVICE COULD NOT BE INSERTED INTO THE ABDOMINAL CAVITY THROUGH THE TROCAR. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-08-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | PROXIMATE |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: VASCULAR GIA STAPLING DEVICE WAS THEN USED TO COME ACROSS THE FISTULA FROM THE LATERAL SIDE, BUT IT CLEARLY FULLY MISFIRED--WHEN HAD COULD CUT ACROSS ONLY APPROXIMATELY HALF OF THE TRACT WITH THE STAPLER BECAME LOCKED UP. AS SUCH, THE STAPLER WAS DISENGAGED. WE THEN USED A COMPLETELY NEW VASCULAR GIA STAPLER DEVICE WITH A NEW LOAD OF STAPLES, AND THIS TIME TRIED TO USE THAT STAPLER COMING FROM THE MEDIAL SIDE. WE WAITED AT LEAST 30 SECONDS PRIOR TO FIRING THE STAPLER AS WE HAD WITH THE FIRST STAPLER, AND AGAIN WHEN TRYING TO COMPLETE THE CUT WITH THE STAPLER, THE MECHANISMS SIMPLY WOULD NOT ADVANCE AFTER THE STAPLER HAD CUT ACROSS APPROXIMATELY HALF THE WIDTH OF THE FISTULA TRACT TISSUES ON THE MEDIAL SIDE. AS SUCH, THAT STAPLER WAS REMOVED.
19-12-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE DEVICE WOULD NOT OPEN TO RELEASE THE RELOAD IN PLACE. NO FURTHER INFORMATION WAS PROVIDED TO THE SALES REP. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
31-10-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2023 | Injury | intuitive surgical | 18-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY PROCEDURE, THERE WAS A STAPLING ISSUE WHILE USING A SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT LOADED WITH A WHITE SUREFORM 45 CURVED-TIP STAPLER RELOAD. THE ISSUE OCCURRED ON THE THIRD FIRE, WHILE STAPLING BLOOD VESSELS. THE FIRE STOPPED MIDWAY THROUGH, AFTER "TISSUE TOO THICK" AND "BLADE MAY BE EXPOSED" MESSAGES WERE PRESENTED. IT WAS UNKNOWN WHICH BLOOD VESSEL WAS THE TARGET TISSUE. THE VESSEL WAS NOT CALCIFIED NOR WAS IT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY PRIOR TO THE PROCEDURE. THERE WAS NO TISSUE TENSION, AND IT WAS UNKNOWN IF THERE WAS ANY TISSUE BUNCHING. THERE WAS NO UNPLANNED TISSUE REMOVAL DUE TO THE FIRING ISSUE. THE EVENT DID NOT INVOLVE A FAILURE TO FIRE, TISSUE BEING PUSHED FORWARD/DRAGGED DURING THE FIRING SEQUENCE, NOR THE STAPLER JAWS OPENING/RELEASING DURING THE FIRING SEQUENCE. MALFORMED STAPLES WERE PRESENT AS A RESULT OF THIS ISSUE, WHICH APPEARED AS OPEN STAPLES WITH PRONG TIPS THAT WERE BENT SLIGHTLY INWARD. THE RELOAD DID NOT HAVE AN EXPOSED BLADE NOR WAS IT STUCK TO THE TISSUE. STAPLES WERE MISSING FROM THE STAPLE LINE; HOWEVER, NO HOLES/GAPS WERE OBSERVED WITHIN THE STAPLE LINE. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS, AND THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. BUTTRESS MATERIAL WAS NOT USED. THERE WERE NO CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD. THE INSTRUMENT AND ACCESSORY WERE INSPECTED PRIOR TO USE, AND NOTHING ABNORMAL WAS OBSERVED. THE SURGEON CALLED A TECHNICAL SUPPORT ENGINEER (TSE) WHO ADVISED THEM TO UNCLAMP THE STAPLER BY TAPPING THE BLUE PEDAL AND TO CONSIDER CHANGING THE RELOAD. HOWEVER, THE SURGEON STATED THAT THE STAPLER WAS CLAMPED ONTO AN IMPORTANT VESSEL, AND THEY WERE CONCERNED ABOUT THE POTENTIAL FOR BLEEDING IF THE STAPLER WERE TO BE UNCLAMPED. THE TSE RECOMMENDED THAT THE SURGEON CLAMP THE VESSEL IN ANOTHER AREA BEFORE UNCLAMPING THE STAPLER, AND THE SURGEON DID SO, TO PREVENT BLEEDING. THERE WAS NO UNEXPECTED BLEEDING OBSERVED DUE TO THE FIRING ISSUE. TO RESOLVE THE ISSUE, THE SURGEON PERFORMED VASCULAR LIGATION. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE PATIENT HAS BEEN DISCHARGED.
20-12-2023 | Injury | intuitive surgical | 18-01-2024 | SUREFORM | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY PROCEDURE, THERE WAS A STAPLING ISSUE WHILE USING A SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT LOADED WITH A WHITE SUREFORM 45 CURVED-TIP STAPLER RELOAD. THE ISSUE OCCURRED ON THE THIRD FIRE, WHILE STAPLING BLOOD VESSELS. THE FIRE STOPPED MIDWAY THROUGH, AFTER "TISSUE TOO THICK" AND "BLADE MAY BE EXPOSED" MESSAGES WERE PRESENTED. IT WAS UNKNOWN WHICH BLOOD VESSEL WAS THE TARGET TISSUE. THE VESSEL WAS NOT CALCIFIED NOR WAS IT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY PRIOR TO THE PROCEDURE. THERE WAS NO TISSUE TENSION, AND IT WAS UNKNOWN IF THERE WAS ANY TISSUE BUNCHING. THERE WAS NO UNPLANNED TISSUE REMOVAL DUE TO THE FIRING ISSUE. THE EVENT DID NOT INVOLVE A FAILURE TO FIRE, TISSUE BEING PUSHED FORWARD/DRAGGED DURING THE FIRING SEQUENCE, NOR THE STAPLER JAWS OPENING/RELEASING DURING THE FIRING SEQUENCE. MALFORMED STAPLES WERE PRESENT AS A RESULT OF THIS ISSUE, WHICH APPEARED AS OPEN STAPLES WITH PRONG TIPS THAT WERE BENT SLIGHTLY INWARD. THE RELOAD DID NOT HAVE AN EXPOSED BLADE NOR WAS IT STUCK TO THE TISSUE. STAPLES WERE MISSING FROM THE STAPLE LINE; HOWEVER, NO HOLES/GAPS WERE OBSERVED WITHIN THE STAPLE LINE. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS, AND THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. BUTTRESS MATERIAL WAS NOT USED. THERE WERE NO CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD. THE INSTRUMENT AND ACCESSORY WERE INSPECTED PRIOR TO USE, AND NOTHING ABNORMAL WAS OBSERVED. THE SURGEON CALLED A TECHNICAL SUPPORT ENGINEER (TSE) WHO ADVISED THEM TO UNCLAMP THE STAPLER BY TAPPING THE BLUE PEDAL AND TO CONSIDER CHANGING THE RELOAD. HOWEVER, THE SURGEON STATED THAT THE STAPLER WAS CLAMPED ONTO AN IMPORTANT VESSEL, AND THEY WERE CONCERNED ABOUT THE POTENTIAL FOR BLEEDING IF THE STAPLER WERE TO BE UNCLAMPED. THE TSE RECOMMENDED THAT THE SURGEON CLAMP THE VESSEL IN ANOTHER AREA BEFORE UNCLAMPING THE STAPLER, AND THE SURGEON DID SO, TO PREVENT BLEEDING. THERE WAS NO UNEXPECTED BLEEDING OBSERVED DUE TO THE FIRING ISSUE. TO RESOLVE THE ISSUE, THE SURGEON PERFORMED VASCULAR LIGATION. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE PATIENT HAS BEEN DISCHARGED.
01-01-2024 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE APPENDECTOMY, THE STAPLER DID NOT FIRE. THE GREEN BUTTON WAS PRESSED, BUT THE HANDLE JAMMED. THE RELOAD WAS REPLACED, AND THE SAME HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
29-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | EEA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY FOR MALIGNANT TUMOR OF THE COLON, WHEN THE DEVICE WAS INSERTED THROUGH THE PATIENT'S ANUS AT THE TIME OF ANASTOMOSIS, THE WHITE PART AT THE PROXIMAL BASE OF THE RED LEVER THAT PREVENTS MISFIRING BROKE. THE BROKEN COMPONENT DID NOT FALL INTO THE PATIENT'S CAVITY. ANOTHER PRODUCT WAS USED WITH THE INITIAL ANVIL TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE PROCEDURE, WHILE ON STOMACH, THE SURGEON LOADED STAPLER AND CLOSED ON TISSUE BUT THE STAPLER WOULD NOT FIRE. THE GREEN FIRING BUTTON WAS PUSHED. WHEN SURGEON DEPRESSED FIRING BUTTON, NOTHING HAPPENED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-01-2024 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A VIDEO-ASSISTED THORACIC SURGERY (VATS) LEFT LUNG SEGMENTECTOMY, THE DEVICE WAS FIRED TO STAPLE THE LUNG LOBE. WHEN REMOVING THE STAPLE BOARD, A SMALL PIECE OF TISSUE WAS INADVERTENTLY PULLED DOWN ALONG WITH THE STAPLE BOARD. THERE WAS NO PATIENT INJURY.
30-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC UMBILICAL HERNIA PROCEDURE, WHILE USING THE TACKER FOR FIRING THE MESH, IT DID NOT FIX WELL. A NEW TACKER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-02-2024 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK AND COULD NOT BE FIRED. THE RELOADS WERE ALSO DIFFICULT TO UNLOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE HANDLES WERE CHANGED TWICE. NO PATIENT INJURY.
01-02-2024 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK AND COULD NOT BE FIRED. THE RELOADS WERE ALSO DIFFICULT TO UNLOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE HANDLES WERE CHANGED TWICE. NO PATIENT INJURY.
01-02-2024 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK AND COULD NOT BE FIRED. THE RELOADS WERE ALSO DIFFICULT TO UNLOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE HANDLES WERE CHANGED TWICE. NO PATIENT INJURY.
01-02-2024 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SLEEVE, AT THE BEGINNING, WHILE USING RELOAD WAS PLACED INTO THE TISSUE, THE DEVICE DID NOT FIRE AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE. THE SAME ISSUE OCCURRED WITH THE OTHER THREE RELOADS. ALL RELOADS WERE STUCK AND COULD NOT BE FIRED. THE RELOADS WERE ALSO DIFFICULT TO UNLOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THE HANDLES WERE CHANGED TWICE. NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Expulsion | A010401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
22-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY (BARIATRIC SURGERY), AT THE TIME OF FIRING THE SECOND 45MM. RELOAD AT THE BOTTOM OF THE STOMACH, THERE WAS A FEELING WITH THE HANDLE WHICH HAD MORE LOUD NOISE, AND WHEN RE-SQUEEZES THE TRIGGER, THE HANDLE SOUNDS LOUDER, AT WHICH THE SURGEON DECIDES NOT TO CONTINUE TO TRY TO STAPLE. WHEN THE DOCTOR RETRACTS THE BLACK KNOB ON THE HANDLE, IT WAS FOUND THAT THE STAPLES HAD NOT BEEN FORMED, REFERRED TO BY THE SURGEON AS ¿AN EXPLOSION" WITHIN THE PATIENT. IT WAS MENTIONED THAT THE STAPLE LINE WAS INCOMPLETE DISTALLY AND SOME OF THE STAPLES FELL INTO THE PATIENT CAVITY. THE TISSUE HAD TO BE STAPLED HIGHER THAN EXPECTED, LEAVING THE STOMACH SMALLER. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE SURGEON HAD TO USED ANOTHER HANDLE AND RELOAD TO RESOLVE THE ISSUE.
22-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY (BARIATRIC SURGERY), AT THE TIME OF FIRING THE SECOND 45MM. RELOAD AT THE BOTTOM OF THE STOMACH, THERE WAS A FEELING WITH THE HANDLE WHICH HAD MORE LOUD NOISE, AND WHEN RE-SQUEEZES THE TRIGGER, THE HANDLE SOUNDS LOUDER, AT WHICH THE SURGEON DECIDES NOT TO CONTINUE TO TRY TO STAPLE. WHEN THE DOCTOR RETRACTS THE BLACK KNOB ON THE HANDLE, IT WAS FOUND THAT THE STAPLES HAD NOT BEEN FORMED, REFERRED TO BY THE SURGEON AS ¿AN EXPLOSION" WITHIN THE PATIENT. IT WAS MENTIONED THAT THE STAPLE LINE WAS INCOMPLETE DISTALLY AND SOME OF THE STAPLES FELL INTO THE PATIENT CAVITY. THE TISSUE HAD TO BE STAPLED HIGHER THAN EXPECTED, LEAVING THE STOMACH SMALLER. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE SURGEON HAD TO USED ANOTHER HANDLE AND RELOAD TO RESOLVE THE ISSUE.
22-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Detachment of Device or Device Component | A0501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY (BARIATRIC SURGERY), AT THE TIME OF FIRING THE SECOND 45MM. RELOAD AT THE BOTTOM OF THE STOMACH, THERE WAS A FEELING WITH THE HANDLE WHICH HAD MORE LOUD NOISE, AND WHEN RE-SQUEEZES THE TRIGGER, THE HANDLE SOUNDS LOUDER, AT WHICH THE SURGEON DECIDES NOT TO CONTINUE TO TRY TO STAPLE. WHEN THE DOCTOR RETRACTS THE BLACK KNOB ON THE HANDLE, IT WAS FOUND THAT THE STAPLES HAD NOT BEEN FORMED, REFERRED TO BY THE SURGEON AS ¿AN EXPLOSION" WITHIN THE PATIENT. IT WAS MENTIONED THAT THE STAPLE LINE WAS INCOMPLETE DISTALLY AND SOME OF THE STAPLES FELL INTO THE PATIENT CAVITY. THE TISSUE HAD TO BE STAPLED HIGHER THAN EXPECTED, LEAVING THE STOMACH SMALLER. THERE WAS AN UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. THE SURGEON HAD TO USED ANOTHER HANDLE AND RELOAD TO RESOLVE THE ISSUE.
27-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE ILEOCECAL RESECTION, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ILEOCECAL RESECTION, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE ILEOCECAL RESECTION, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ILEOCECAL RESECTION, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Expulsion | A010401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE ILEOCECAL RESECTION, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED. THERE WAS NO PATIENT INJURY.
27-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE ILEOCECAL RESECTION, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE KNIFE DID NOT ADVANCE DURING THE TISSUE RESECTION. WHEN THE CARTRIDGE PROXIMAL BASE WAS CHECKED, IT WAS FOUND THAT STAPLE HAS COME OUT, SO THERE WAS A POSSIBILITY OF PRE-FIRE. IT WAS ALSO MENTIONED THAT THE JAWS OF THE RELOAD HAS BEEN FIXATED TO THE TISSUE. THE DEVICE WAS THEN REMOVED WHEN THE JAWS WERE OPENED AND ANCHORING WAS RELEASED. THERE WAS NO PATIENT INJURY.
19-10-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT COME OUT WHEN FIRED.
19-10-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY, THE STAPLES DID NOT COME OUT WHEN FIRED.
02-01-2024 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A LAPAROSCOPIC ROBOTIC PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE DEVICE DID NOT FIRE. FIRING WAS ATTEMPTED BY ANOTHER SURGEON AND AS WELL AT THE BACK TABLE, BUT THE DEVICE STILL WOULD NOT FIRE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
04-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT LUNG RESECTION, PLEURODESIS, AND LEFT CHEST CLOSED DRAINAGE UNDER GENERAL ANESTHESIA, WHEN THE DEVICE WAS USED TO CUT THE DAMAGED TISSUE, THE DEVICE WAS DIFFICULT TO FIRE AND AFTER FORCED FIRING, THE STAPLES WERE POORLY FORMED. AFTER CHECKING THAT THE APPEARANCE AND USAGE OF THE STAPLER WERE CORRECT, THE RELOAD WAS REPLACED, BUT THE SAME ISSUE OCCURRED. ANOTHER BRAND OF STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO BLEEDING AND THE INCISION AT THE ANASTOMOSIS WAS COMPLETE.
04-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Firing Problem | A0505 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT LUNG RESECTION, PLEURODESIS, AND LEFT CHEST CLOSED DRAINAGE UNDER GENERAL ANESTHESIA, WHEN THE DEVICE WAS USED TO CUT THE DAMAGED TISSUE, THE DEVICE WAS DIFFICULT TO FIRE AND AFTER FORCED FIRING, THE STAPLES WERE POORLY FORMED. AFTER CHECKING THAT THE APPEARANCE AND USAGE OF THE STAPLER WERE CORRECT, THE RELOAD WAS REPLACED, BUT THE SAME ISSUE OCCURRED. ANOTHER BRAND OF STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO BLEEDING AND THE INCISION AT THE ANASTOMOSIS WAS COMPLETE.
01-01-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | ECHELON 3000 60MM LONG | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT WHEN THEY TRIED TO RELOAD AND IT WOULDN'T LOAD. THEY CHECKED THE CARTRIDGE/ANVIL AND NOTHING SEEMED TO HAVE SHOWN AS TO WHY IT SHOULDN'T HAVE RELOADED PROPERLY. USED ANOTHER DEVICE TO COMPLETE. NO FRAGMENTS MADE. NO PATIENT HARM.
04-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT LUNG RESECTION, PLEURODESIS, AND LEFT CHEST CLOSED DRAINAGE UNDER GENERAL ANESTHESIA, WHEN THE DEVICE WAS USED TO CUT THE DAMAGED TISSUE, THE DEVICE WAS DIFFICULT TO FIRE AND AFTER FORCED FIRING, THE STAPLES WERE POORLY FORMED. AFTER CHECKING THAT THE APPEARANCE AND USAGE OF THE STAPLER WERE CORRECT, THE RELOAD WAS REPLACED, BUT THE SAME ISSUE OCCURRED. ANOTHER BRAND OF STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO BLEEDING AND THE INCISION AT THE ANASTOMOSIS WAS COMPLETE.
04-12-2023 | Malfunction | covidien lp llc north haven | 18-01-2024 | ENDO GIA | Firing Problem | A0505 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LEFT LUNG RESECTION, PLEURODESIS, AND LEFT CHEST CLOSED DRAINAGE UNDER GENERAL ANESTHESIA, WHEN THE DEVICE WAS USED TO CUT THE DAMAGED TISSUE, THE DEVICE WAS DIFFICULT TO FIRE AND AFTER FORCED FIRING, THE STAPLES WERE POORLY FORMED. AFTER CHECKING THAT THE APPEARANCE AND USAGE OF THE STAPLER WERE CORRECT, THE RELOAD WAS REPLACED, BUT THE SAME ISSUE OCCURRED. ANOTHER BRAND OF STAPLER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO BLEEDING AND THE INCISION AT THE ANASTOMOSIS WAS COMPLETE.
29-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC PULMONARY SEGMENTECTOMY PROCEDURE, AT THE TIME OF DETACHING THE PULMONARY SEGMENTAL BRONCHUS, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW HANDLE WAS USED WITH THE SAME RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-08-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING SURGERY, THE COLON PROFESSOR TRIED TO CONNECT THE ANVIL AND BODY APPROXIMATION. THE LIGHT CAME ON IN THE BODY AT FIRST, BUT IT DID NOT LIGHT UP WHEN CONNECTED. SO, THEY THINK IT'S A BATTERY PROBLEM AND PUT IN A NEW BATTERY. HOWEVER, IT DIDN'T WORK. SO THEY TOOK OUT A NEW BODY AND USED IT, AND IT WORKED WELL. THERE WAS NO PATIENT CONSEQUENCE REPORTED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY.
01-01-2024 | Injury | ethicon endo-surgery | 18-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED DURING A LAR ON ROBOT THAT NOT MUCH SKELETONIZATION WAS DONE PRIOR TO CLOSING DOWN AND FIRING ON TISSUE BUT WALKED THROUGH STEPS, AND SHE FIRED DEVICE CORRECTLY. NOTED DONUTS WERE PERFECTLY FORMED, THE BOWEL WAS PREPPED PRE-OP ALONG WITH A BETADYNE WASH OUT. ANASTOMOSIS WAS TESTED VIA COLONOSCOPE TO CHECK FOR BUBBLES. NO LEAK NOTED. THE PATIENT WAS BROUGHT BACK WITH AN ANASTOMOTIC LEAK.
16-11-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, ON CUTTING AND CLOSING THE GASTRIC BODY, THE STAPLES WERE POORLY FORMED. THE TISSUE WAS NOT COMPLETELY CLOSED, AND BLEEDING OCCURRED. ADDITIONAL MANUAL SUTURING USING THREADS WAS PERFORMED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
16-11-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY, ON CUTTING AND CLOSING THE GASTRIC BODY, THE STAPLES WERE POORLY FORMED. THE TISSUE WAS NOT COMPLETELY CLOSED, AND BLEEDING OCCURRED. ADDITIONAL MANUAL SUTURING USING THREADS WAS PERFORMED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
14-08-2023 | Injury | ethicon endo-surgery | 18-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED FROM A HEALTH AUTHORITY THAT DURING AFTER A THORACOSCOPIC LEFT LOWER BULLAE AND LEFT UPPER BULLAE RESECTION, NO OBVIOUS AIR LEAKAGE WAS FOUND IN THE EXPANDED LUNG DURING THE OPERATION. THE PATIENT'S CHEST BOTTLE DID NOT SHOW OBVIOUS GAS ESCAPE ON THE 1ST DAY AFTER THE OPERATION. THE PATIENT'S CHEST BOTTLE SHOWED A LARGE AMOUNT OF GAS ESCAPE ON THE 3RD DAY AFTER THE OPERATION. THE THORACENTESIS WAS PERFORMED AGAIN TO DRAIN THE GAS. AFTER SUPPORTIVE TREATMENT INCLUDING CLOSED THORACIC DRAINAGE AND VACUUM ASPIRATION, THE PATIENT WAS IMPROVED AND DISCHARGED.
14-08-2023 | Injury | ethicon endo-surgery | 18-01-2024 | GST GOLD RELOAD, 45MM, 6-ROW | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: IT WAS REPORTED FROM A HEALTH AUTHORITY THAT DURING AFTER A THORACOSCOPIC LEFT LOWER BULLAE AND LEFT UPPER BULLAE RESECTION, NO OBVIOUS AIR LEAKAGE WAS FOUND IN THE EXPANDED LUNG DURING THE OPERATION. THE PATIENT'S CHEST BOTTLE DID NOT SHOW OBVIOUS GAS ESCAPE ON THE 1ST DAY AFTER THE OPERATION. THE PATIENT'S CHEST BOTTLE SHOWED A LARGE AMOUNT OF GAS ESCAPE ON THE 3RD DAY AFTER THE OPERATION. THE THORACENTESIS WAS PERFORMED AGAIN TO DRAIN THE GAS. AFTER SUPPORTIVE TREATMENT INCLUDING CLOSED THORACIC DRAINAGE AND VACUUM ASPIRATION, THE PATIENT WAS IMPROVED AND DISCHARGED.
27-12-2023 | Malfunction | us surgical puerto rico | 18-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ESOPHAGEAL CANCER RESECTION, IN DETACHING ESOPHAGEAL TISSUE, THE SURGEON COULD PRESS THE GREEN BUTTON. HOWEVER, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER HANDLE WAS USED WITH THE SAME RELOAD. THERE WAS NO PATIENT INJURY.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Difficult to Remove | A150207 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Difficult to Remove | A150207 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Failure to Align | A050701 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Failure to Align | A050701 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Misfire | A050502 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Failure to Cut | A050702 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Failure to Cut | A050702 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Material Protrusion/Extrusion | A0411 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Material Protrusion/Extrusion | A0411 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Firing Problem | A0505 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
17-10-2023 | Malfunction | celestica de monterrey s.a de c.v | 18-01-2024 | ENDO GIA | Firing Problem | A0505 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE CLAMPING FORCE AND CUTTING FORCE OF THE STAPLER CARTRIDGE WERE INSUFFICIENT. THE THICK TISSUE COULD NOT BE CLAMPED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
21-12-2023 | Injury | intuitive surgical | 18-01-2024 | SUREFORM | Material Fragmentation | A040103 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS PROCEDURE, A WHITE RELOAD WAS FIRED AND THE TARGET TISSUE WAS PUSHED. THE SURGEON CONTINUED TO USE THE SAME STAPLER INSTRUMENT FOR THE REMAINDER OF THE PROCEDURE; NO FURTHER COMPLICATIONS WERE EXPERIENCED. THE PROCEDURE WAS COMPLETED ROBOTICALLY.
21-12-2023 | Injury | intuitive surgical | 18-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS PROCEDURE, A WHITE RELOAD WAS FIRED AND THE TARGET TISSUE WAS PUSHED. THE SURGEON CONTINUED TO USE THE SAME STAPLER INSTRUMENT FOR THE REMAINDER OF THE PROCEDURE; NO FURTHER COMPLICATIONS WERE EXPERIENCED. THE PROCEDURE WAS COMPLETED ROBOTICALLY.
21-12-2023 | Injury | intuitive surgical | 18-01-2024 | SUREFORM | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS PROCEDURE, A WHITE RELOAD WAS FIRED AND THE TARGET TISSUE WAS PUSHED. THE SURGEON CONTINUED TO USE THE SAME STAPLER INSTRUMENT FOR THE REMAINDER OF THE PROCEDURE; NO FURTHER COMPLICATIONS WERE EXPERIENCED. THE PROCEDURE WAS COMPLETED ROBOTICALLY.
18-12-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER WOULD NOT OPEN AFTER USE. THIS HAPPENED INTRA-OP COMPLETED WITH LIKE DEVICE. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE JAWS OF THE GUN WERE STUCK AND CANNOT BE LOOSENED WHEN LOADING NAILS. NO PATIENT CONSEQUENCES WERE REPORTED. NO FURTHER DETAILS WERE PROVIDED.
06-11-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SYGMOID PROCEDURE, THE FIRST 4 STAPLES WENT WELL, BUT ON THE 5TH STAPLE THE STAPLER STOPPED IN MID-STAPLE. THEY COULDN'T GO BACK NORMALLY AND HAD TO OPEN THE STAPLER MANUALLY USING THE TOP FLAP. THE STAPLER BECAME UNUSABLE. THE DEVICE WAS MANUALLY OPENED AND A NEW STAPLER WAS USED TO COMPLETE THE SURGERY SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
06-11-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SYGMOID PROCEDURE, THE FIRST 4 STAPLES WENT WELL, BUT ON THE 5TH STAPLE THE STAPLER STOPPED IN MID-STAPLE. THEY COULDN'T GO BACK NORMALLY AND HAD TO OPEN THE STAPLER MANUALLY USING THE TOP FLAP. THE STAPLER BECAME UNUSABLE. THE DEVICE WAS MANUALLY OPENED AND A NEW STAPLER WAS USED TO COMPLETE THE SURGERY SUCCESSFULLY. NO PATIENT CONSEQUENCES WERE REPORTED.
18-12-2023 | Injury | ethicon endo-surgery | 18-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE STAPLING DEVICE WAS USED 3 TIMES DURING THE SURGERY. WHEN STAPLING THE RIGHT HEPATIC VEIN, DEVICE BLOCKED AFTER STAPLING. UNABLE TO OPEN THE DEVICE DESPITE THE USE OF THE EMERGENCY BUTTON. AT THE SAME TIME, HEMORRHAGE ON A WOUND OF THE INFERIOR VENA CAVA. HYPOTHESIS OF THE DEVICE WHICH "SLIPPED" DURING THE OPENING MANEUVERS. AFTER CONTROL OF THE BLEEDING, OPENING OF THE DEVICE BY OPENING THE "COVER". PARTIAL STAPLING. PATIENT WAS REPORTED TO HAVE HEMORRHAGIC SHOCK. PATIENT STAY IN INTENSIVE CARE.
18-12-2023 | Injury | ethicon endo-surgery | 18-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE STAPLING DEVICE WAS USED 3 TIMES DURING THE SURGERY. WHEN STAPLING THE RIGHT HEPATIC VEIN, DEVICE BLOCKED AFTER STAPLING. UNABLE TO OPEN THE DEVICE DESPITE THE USE OF THE EMERGENCY BUTTON. AT THE SAME TIME, HEMORRHAGE ON A WOUND OF THE INFERIOR VENA CAVA. HYPOTHESIS OF THE DEVICE WHICH "SLIPPED" DURING THE OPENING MANEUVERS. AFTER CONTROL OF THE BLEEDING, OPENING OF THE DEVICE BY OPENING THE "COVER". PARTIAL STAPLING. PATIENT WAS REPORTED TO HAVE HEMORRHAGIC SHOCK. PATIENT STAY IN INTENSIVE CARE.
18-12-2023 | Injury | ethicon endo-surgery | 18-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE STAPLING DEVICE WAS USED 3 TIMES DURING THE SURGERY. WHEN STAPLING THE RIGHT HEPATIC VEIN, DEVICE BLOCKED AFTER STAPLING. UNABLE TO OPEN THE DEVICE DESPITE THE USE OF THE EMERGENCY BUTTON. AT THE SAME TIME, HEMORRHAGE ON A WOUND OF THE INFERIOR VENA CAVA. HYPOTHESIS OF THE DEVICE WHICH "SLIPPED" DURING THE OPENING MANEUVERS. AFTER CONTROL OF THE BLEEDING, OPENING OF THE DEVICE BY OPENING THE "COVER". PARTIAL STAPLING. PATIENT WAS REPORTED TO HAVE HEMORRHAGIC SHOCK. PATIENT STAY IN INTENSIVE CARE.
18-12-2023 | Injury | ethicon endo-surgery | 18-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Activation Failure | A150101 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE STAPLING DEVICE WAS USED 3 TIMES DURING THE SURGERY. WHEN STAPLING THE RIGHT HEPATIC VEIN, DEVICE BLOCKED AFTER STAPLING. UNABLE TO OPEN THE DEVICE DESPITE THE USE OF THE EMERGENCY BUTTON. AT THE SAME TIME, HEMORRHAGE ON A WOUND OF THE INFERIOR VENA CAVA. HYPOTHESIS OF THE DEVICE WHICH "SLIPPED" DURING THE OPENING MANEUVERS. AFTER CONTROL OF THE BLEEDING, OPENING OF THE DEVICE BY OPENING THE "COVER". PARTIAL STAPLING. PATIENT WAS REPORTED TO HAVE HEMORRHAGIC SHOCK. PATIENT STAY IN INTENSIVE CARE.
25-10-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED PLUS STANDARD 45 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE WAS DEFECTIVE. THE DEVICE BLOCKED DURING SURGERY. NO PATIENT CONSEQUENCES REPORTED.
19-12-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BILATERAL LUNG TRANSPLANT DR. WERE USING A PVE35A AND WHEN THEY FIRED THE STAPLER IT MADE A STRANGE NOISE AND DIDN'T FULLY FIRE AND REQUIRED THEM TO USE THE MANUAL OVERRIDE TO BACK OFF THE KNIFE BLADE, OPEN THE JAWS AND REMOVE THE DEVICE. NO PATIENT CONSEQUENCES.
19-12-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BILATERAL LUNG TRANSPLANT DR. WERE USING A PVE35A AND WHEN THEY FIRED THE STAPLER IT MADE A STRANGE NOISE AND DIDN'T FULLY FIRE AND REQUIRED THEM TO USE THE MANUAL OVERRIDE TO BACK OFF THE KNIFE BLADE, OPEN THE JAWS AND REMOVE THE DEVICE. NO PATIENT CONSEQUENCES.
19-12-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A BILATERAL LUNG TRANSPLANT DR. WERE USING A PVE35A AND WHEN THEY FIRED THE STAPLER IT MADE A STRANGE NOISE AND DIDN'T FULLY FIRE AND REQUIRED THEM TO USE THE MANUAL OVERRIDE TO BACK OFF THE KNIFE BLADE, OPEN THE JAWS AND REMOVE THE DEVICE. NO PATIENT CONSEQUENCES.
26-10-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOSTOMY PROCEDURE WITH DRAINAGE, THE STAPLER GOT STUCK. THE REVERSE BUTTON WAS PRESSED, BUT WITHOUT SUCCESS. SURGEON HAD TO OPEN THE LID AND DO THE MANUAL MANEUVER. OPEN ANOTHER STAPLER TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES REPORTED.
26-10-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED PLUS STANDARD 45 | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE THORACOSTOMY PROCEDURE WITH DRAINAGE, THE STAPLER GOT STUCK. THE REVERSE BUTTON WAS PRESSED, BUT WITHOUT SUCCESS. SURGEON HAD TO OPEN THE LID AND DO THE MANUAL MANEUVER. OPEN ANOTHER STAPLER TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS AND CHOLECYSTECTOMY, SURGEON REPORTED THAT AFTER COMPLETE FIRING, THE STICKY OF TISSUE ON THE RELOAD OF STAPLER WHEN OPENING THE JAW OF THE DEVICE. IT WAS ALLEGED THAT THE DEVICE DID NOT PERFORM A CLEAN CUT TO THE TISSUE. STAPLERS ARE WELL FORMED, IT JUST THE STICKY TISSUE ON THE JAW OF THE DEVICE AFTER FIRING OF STAPLER IN LAPAROSCOPIC GASTRIC BYPASS CASE. SURGERY WAS COMPLETED SUCCESSFULLY AND WAS NOT DELAYED. NO PATIENT CONSEQUENCES REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS AND CHOLECYSTECTOMY, SURGEON REPORTED THAT AFTER COMPLETE FIRING, THE STICKY OF TISSUE ON THE RELOAD OF STAPLER WHEN OPENING THE JAW OF THE DEVICE. IT WAS ALLEGED THAT THE DEVICE DID NOT PERFORM A CLEAN CUT TO THE TISSUE. STAPLERS ARE WELL FORMED, IT JUST THE STICKY TISSUE ON THE JAW OF THE DEVICE AFTER FIRING OF STAPLER IN LAPAROSCOPIC GASTRIC BYPASS CASE. SURGERY WAS COMPLETED SUCCESSFULLY AND WAS NOT DELAYED. NO PATIENT CONSEQUENCES REPORTED.
06-11-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK PROCEDURE, DEVICE COULD NOT BE TRIGGERED. NO PATIENT CONSEQUENCES REPORTED. THERE WAS NO SIGNIFICANT DELAY REPORTED FOR THE SURGERY. A NEW INSTRUMENT AND A NEW MAGAZINE WERE TAKEN AND THE PROCEDURE FINISHED SUCCESSFULLY.
22-12-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, THE SURGEON REPORTED THAT THE STAPLER MALFUNCTIONED. THEY FIRED THE STAPLER, THE BLADE WENT ACROSS AND THEN WHEN THE BLADE WAS COMING BACK ACROSS, IT MADE A LOUD CRUNCHY TYPE SOUND AND THE BATTERY CAME HALFWAY OUT, SO THEY HAD TO DO THE MANUAL RELEASE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT HARM REPORTED.
22-12-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | ECHELON 3000 45 STANDARD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, THE SURGEON REPORTED THAT THE STAPLER MALFUNCTIONED. THEY FIRED THE STAPLER, THE BLADE WENT ACROSS AND THEN WHEN THE BLADE WAS COMING BACK ACROSS, IT MADE A LOUD CRUNCHY TYPE SOUND AND THE BATTERY CAME HALFWAY OUT, SO THEY HAD TO DO THE MANUAL RELEASE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT HARM REPORTED.
01-01-2023 | Malfunction | ethicon endo-surgery | 18-01-2024 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, WITH THE SECOND MAGAZINE, THE STAPLER RELEASED BRIEFLY AND THEN STOPPED. EXEMPLARY HANDLING INTRAOPERATIVELY. NO PATIENT CONSEQUENCES WERE REPORTED.
22-12-2023 | Injury | ethicon endo-surgery | 18-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ZANKERS DIVERTICULUM PROCEDURE, THAT AFTER THE PROCEDURE THE PATIENT IS NOW NOT DOING WELL. THE PATIENT HAS BEEN IN AND OUT OF THE HOSPITAL WITH A LEAK. ORIGINAL PROCEDURE DID NOT PRESENT ANY ISSUES. NO FURTHER INFORMATION WAS PROVIDED.
22-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP APPENDECTOMY CASE THE DEVICE WOULD NOT FIRE. SURGEON USED ANOTHER DEVICE TO COMPLETE CASE. NO PATIENT HARM WAS REPORTED.
19-01-2024 | Injury | covidien lp llc north haven | 19-01-2024 | UNKNOWN SURGICAL INNOVATIONS NH PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE CLINICAL STUDY, POSTOPERATIVELY, THERE WERE THREE COMPLICATIONS USING THE DEVICES. A BLOOD TRANSFUSION, BLEEDING AND ANASTOMOTIC LEAK.
19-01-2024 | Injury | covidien lp llc north haven | 19-01-2024 | UNKNOWN SURGICAL INNOVATIONS NH PRODUCT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE CLINICAL STUDY, POSTOPERATIVELY, THERE WERE THREE COMPLICATIONS USING THE DEVICES. A BLOOD TRANSFUSION, BLEEDING AND ANASTOMOTIC LEAK.
27-08-2023 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTROPLASTY BYPASS PROCEDURE, THE STAPLE LINE WAS NOT COMPLETE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
15-12-2023 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SUBTOTAL GASTRECTOMY PROCEDURE, ON THE STOMACH, AFTER THE RELOAD WAS LOADED TO THE HANDLE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. HOWEVER, THE DEVICE DID NOT FIRE AND MADE A STRANGE NOISE, AS IF SOMETHING HAD BROKEN. TO RESOLVE THE ISSUE, A POWERED HANDLE WAS USED WITH A NEW RELOAD. THERE WAS NO PATIENT INJURY.
22-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL REVERSE BUTTON TO RETURN KNIFE. IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL REVERSE BUTTON TO RETURN KNIFE. IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL REVERSE BUTTON TO RETURN KNIFE. IT WAS REPORTED THAT DURING THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLER COULDN'T BE FIRED. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
21-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGERY, THE KNIFE DID NOT MOVE FORWARD FROM MIDDLE PART OF STAPLER. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
02-11-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | CARTRIDGE FINAL ASSY 60MM BLUE | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLE CARTRIDGE WAS USED TO CUT THE HEPATIC PEDICLE (WITH THE THICKNESS OF 1-1.5 CM), BUT IT COULD NOT BE CUT AND THE STAPLE CARTRIDGE COULD NOT BE OPENED. FINALLY, THE DEVICE WAS HARD PULLED OUT, CAUSING THAT 1/3 OF THE STAPLE CARTRIDGE FELL OFF AND THE WOUND SURFACE WAS ABOUT 1.5 CM. SUTURE WAS USED TO OVERSEW. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-11-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | CARTRIDGE FINAL ASSY 60MM BLUE | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLE CARTRIDGE WAS USED TO CUT THE HEPATIC PEDICLE (WITH THE THICKNESS OF 1-1.5 CM), BUT IT COULD NOT BE CUT AND THE STAPLE CARTRIDGE COULD NOT BE OPENED. FINALLY, THE DEVICE WAS HARD PULLED OUT, CAUSING THAT 1/3 OF THE STAPLE CARTRIDGE FELL OFF AND THE WOUND SURFACE WAS ABOUT 1.5 CM. SUTURE WAS USED TO OVERSEW. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-11-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | CARTRIDGE FINAL ASSY 60MM BLUE | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLE CARTRIDGE WAS USED TO CUT THE HEPATIC PEDICLE (WITH THE THICKNESS OF 1-1.5 CM), BUT IT COULD NOT BE CUT AND THE STAPLE CARTRIDGE COULD NOT BE OPENED. FINALLY, THE DEVICE WAS HARD PULLED OUT, CAUSING THAT 1/3 OF THE STAPLE CARTRIDGE FELL OFF AND THE WOUND SURFACE WAS ABOUT 1.5 CM. SUTURE WAS USED TO OVERSEW. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-11-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | CARTRIDGE FINAL ASSY 60MM BLUE | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE STAPLE CARTRIDGE WAS USED TO CUT THE HEPATIC PEDICLE (WITH THE THICKNESS OF 1-1.5 CM), BUT IT COULD NOT BE CUT AND THE STAPLE CARTRIDGE COULD NOT BE OPENED. FINALLY, THE DEVICE WAS HARD PULLED OUT, CAUSING THAT 1/3 OF THE STAPLE CARTRIDGE FELL OFF AND THE WOUND SURFACE WAS ABOUT 1.5 CM. SUTURE WAS USED TO OVERSEW. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE WHITE PART AND THE PLASTIC PART OF THE TYVEK WAS ATTACHED DIAGONALLY, THE ATTACHED AREA WAS SO SMALL. THE DOCTOR CONCERNED THAT IF THE STERILITY OF THE PACKAGE WAS MAINTAINED OR NOT. THE DEVICE WAS NOT USED FOR THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PANCREATECTOMY, THE KNIFE DID NOT RETURN PROPERLY AT 5TH FIRING OF COMMON HOLE CLOSURE AT STOMACH-JEJUNUM ANASTOMOSIS. GST60D WAS USED. KNIFE REVERSE SWITCH WAS USED TO OPEN THE JAWS. THE SAME ISSUE OCCURRED AT THE COMMON HOLE CLOSURE AT BRAUN ANASTOMOSIS. GST60W WAS USED. KNIFE REVERSE SWITCH WAS USED, BUT THE ISSUE DID NOT IMPROVE. MANUAL OVERRIDE LEVER WAS STIFF AND COULD NOT USE FOR ABOUT 5 MINUTES. OPEN BUTTON WAS PRESSED, AND THE JAWS OPENED, AND THE CASE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PANCREATECTOMY, THE KNIFE DID NOT RETURN PROPERLY AT 5TH FIRING OF COMMON HOLE CLOSURE AT STOMACH-JEJUNUM ANASTOMOSIS. GST60D WAS USED. KNIFE REVERSE SWITCH WAS USED TO OPEN THE JAWS. THE SAME ISSUE OCCURRED AT THE COMMON HOLE CLOSURE AT BRAUN ANASTOMOSIS. GST60W WAS USED. KNIFE REVERSE SWITCH WAS USED, BUT THE ISSUE DID NOT IMPROVE. MANUAL OVERRIDE LEVER WAS STIFF AND COULD NOT USE FOR ABOUT 5 MINUTES. OPEN BUTTON WAS PRESSED, AND THE JAWS OPENED, AND THE CASE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PANCREATECTOMY, THE KNIFE DID NOT RETURN PROPERLY AT 5TH FIRING OF COMMON HOLE CLOSURE AT STOMACH-JEJUNUM ANASTOMOSIS. GST60D WAS USED. KNIFE REVERSE SWITCH WAS USED TO OPEN THE JAWS. THE SAME ISSUE OCCURRED AT THE COMMON HOLE CLOSURE AT BRAUN ANASTOMOSIS. GST60W WAS USED. KNIFE REVERSE SWITCH WAS USED, BUT THE ISSUE DID NOT IMPROVE. MANUAL OVERRIDE LEVER WAS STIFF AND COULD NOT USE FOR ABOUT 5 MINUTES. OPEN BUTTON WAS PRESSED, AND THE JAWS OPENED, AND THE CASE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | POWERED 60 ECHELON +, 280MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PANCREATECTOMY, THE KNIFE DID NOT RETURN PROPERLY AT 5TH FIRING OF COMMON HOLE CLOSURE AT STOMACH-JEJUNUM ANASTOMOSIS. GST60D WAS USED. KNIFE REVERSE SWITCH WAS USED TO OPEN THE JAWS. THE SAME ISSUE OCCURRED AT THE COMMON HOLE CLOSURE AT BRAUN ANASTOMOSIS. GST60W WAS USED. KNIFE REVERSE SWITCH WAS USED, BUT THE ISSUE DID NOT IMPROVE. MANUAL OVERRIDE LEVER WAS STIFF AND COULD NOT USE FOR ABOUT 5 MINUTES. OPEN BUTTON WAS PRESSED, AND THE JAWS OPENED, AND THE CASE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-12-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 19-01-2024 | CAPSURE STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING TAPP (TRANSABDOMINAL PREPERITONEAL) HERNIA REPAIR PROCEDURE, THE CAPSURE STRAIGHT FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE DEPLOYED FASTENER FAILED TO FIXATE THE MESH AND WAS REMOVED FROM THE PATIENT'S BODY. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | ethicon endo-surgery | 19-01-2024 | ECHELON ENDOPATH WHITE, 60MM |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: 60 WHITE ETHICON STAPLE LOAD WOULD NOT FIRE.
21-11-2023 | Malfunction | davol inc sub c.r bard inc -1213643 | 19-01-2024 | SORBAFIX ENHANCED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING AN UNKNOWN PROCEDURE SORBAFIX ENHANCED FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE FIXATION DEVICE FIRED 7 TIMES AND THE FASTENERS WERE NOT HOLDING (FIXATING) THE MESH. IT WAS ALSO REPORTED THAT THE PROCEDURE WAS COMPLETED WITH NON-BARD FIXATION DEVICE. THERE WAS NO PATIENT HARM OR INJURY.
31-08-2023 | Malfunction | davis & geck caribe | 19-01-2024 | GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT ALLOW PROPER STAPLER CLOSURE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. A HAND SEWN WAS PERFORMED TO RESOLVE THE ISSUE.
31-08-2023 | Malfunction | davis & geck caribe | 19-01-2024 | GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT ALLOW PROPER STAPLER CLOSURE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. A HAND SEWN WAS PERFORMED TO RESOLVE THE ISSUE.
31-08-2023 | Malfunction | davis & geck caribe | 19-01-2024 | GIA | Difficult to Open or Close | A051104 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT ALLOW PROPER STAPLER CLOSURE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. A HAND SEWN WAS PERFORMED TO RESOLVE THE ISSUE.
31-08-2023 | Malfunction | davis & geck caribe | 19-01-2024 | GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT ALLOW PROPER STAPLER CLOSURE. THE SURGICAL TIME WAS EXTENDED 30 MINUTES OR MORE DUE TO THE PRODUCT PROBLEM. A HAND SEWN WAS PERFORMED TO RESOLVE THE ISSUE.
11-12-2023 | Injury | davis & geck caribe | 19-01-2024 | UNKNOWN ENDO STITCH INSTRUMENT | Break | A0401 | Unspecified Tissue Injury | Event Description: ACCORDING TO THE REPORTER, DURING A HYSTERECTOMY PROCEDURE, WHEN SUTURING THE VAGINAL CUFF USING THE SUTURING DEVICE, THE NEEDLE BROKE IN HALF. AN X-RAY WAS PERFORMED AND SHOWED THAT THE NEEDLE FRAGMENT WAS STILL IN THE BODY. THE SURGEON REOPENED THE WOUND CLOSURE. A FLUOROSCOPY WAS PERFORMED, AND A TWO HOURS OF LAPAROSCOPIC WORK TO LOCATE THE RETAINED NEEDLE FRAGMENT. THE NEEDLE FRAGMENT WAS LOCATED IN THE BOWEL MESENTERY AND WAS ABLE TO BE REMOVED.
11-12-2023 | Injury | davis & geck caribe | 19-01-2024 | UNKNOWN ENDO STITCH INSTRUMENT | Break | A0401 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER, DURING A HYSTERECTOMY PROCEDURE, WHEN SUTURING THE VAGINAL CUFF USING THE SUTURING DEVICE, THE NEEDLE BROKE IN HALF. AN X-RAY WAS PERFORMED AND SHOWED THAT THE NEEDLE FRAGMENT WAS STILL IN THE BODY. THE SURGEON REOPENED THE WOUND CLOSURE. A FLUOROSCOPY WAS PERFORMED, AND A TWO HOURS OF LAPAROSCOPIC WORK TO LOCATE THE RETAINED NEEDLE FRAGMENT. THE NEEDLE FRAGMENT WAS LOCATED IN THE BOWEL MESENTERY AND WAS ABLE TO BE REMOVED.
03-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. A NEW 60MM RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-12-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, WHILE WEANING LUNG PARENCHYMA, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE 2 HANDLES. THE HANDLES COULD NOT BE TRIGGERED BECAUSE THERE WAS NO RESISTANCE. THE 2 HANDLES AND 2 RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WAS SHEARED TEETH WERE VISIBLE ON THE FIRING RACK DUE TO THICK TISSUE.
08-12-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, WHILE WEANING LUNG PARENCHYMA, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE 2 HANDLES. THE HANDLES COULD NOT BE TRIGGERED BECAUSE THERE WAS NO RESISTANCE. THE 2 HANDLES AND 2 RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WAS SHEARED TEETH WERE VISIBLE ON THE FIRING RACK DUE TO THICK TISSUE.
25-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | GST BLUE RELOAD, 45MM, 6-ROW | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING PROCEDURE, NOTED THAT CARTRIDGE BASE LOOSE AND COULD NOT INSTALL. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, WHILE WEANING LUNG PARENCHYMA, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE 2 HANDLES. THE HANDLES COULD NOT BE TRIGGERED BECAUSE THERE WAS NO RESISTANCE. THE 2 HANDLES AND 2 RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WAS SHEARED TEETH WERE VISIBLE ON THE FIRING RACK DUE TO THICK TISSUE.
08-12-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, WHILE WEANING LUNG PARENCHYMA, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE 2 HANDLES. THE HANDLES COULD NOT BE TRIGGERED BECAUSE THERE WAS NO RESISTANCE. THE 2 HANDLES AND 2 RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WAS SHEARED TEETH WERE VISIBLE ON THE FIRING RACK DUE TO THICK TISSUE.
20-10-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, WHEN THE SURGEON APPLIED PRESSURE TO CLOSE THE STAPLES, A LOUD CRACK NOISE WAS HEARD. THE STAPLE LINE LOOKED LIKE IT WAS SUCCESSFULLY DEPLOYED AND FORMED. WHEN ANOTHER RELOAD WAS BEING LOADED, THE NURSE WAS UNABLE TO LOAD THE DEVICE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. PLI 10- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
20-10-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, WHEN THE SURGEON APPLIED PRESSURE TO CLOSE THE STAPLES, A LOUD CRACK NOISE WAS HEARD. THE STAPLE LINE LOOKED LIKE IT WAS SUCCESSFULLY DEPLOYED AND FORMED. WHEN ANOTHER RELOAD WAS BEING LOADED, THE NURSE WAS UNABLE TO LOAD THE DEVICE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. PLI 10- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
20-10-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE STAPLING THE STOMACH, WHEN THE SURGEON APPLIED PRESSURE TO CLOSE THE STAPLES, A LOUD CRACK NOISE WAS HEARD. THE STAPLE LINE LOOKED LIKE IT WAS SUCCESSFULLY DEPLOYED AND FORMED. WHEN ANOTHER RELOAD WAS BEING LOADED, THE NURSE WAS UNABLE TO LOAD THE DEVICE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. PLI 10- MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
08-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 19-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, WHILE WEANING LUNG PARENCHYMA, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE 2 HANDLES. THE HANDLES COULD NOT BE TRIGGERED BECAUSE THERE WAS NO RESISTANCE. THE 2 HANDLES AND 2 RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED.
08-12-2023 | Malfunction | celestica de monterrey s.a de c.v | 19-01-2024 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO-ASSISTED THORACOSCOPIC SURGERY LOBECTOMY, WHILE WEANING LUNG PARENCHYMA, THE SURGEON COULD PRESS THE GREEN BUTTON AND SQUEEZE THE 2 HANDLES. THE HANDLES COULD NOT BE TRIGGERED BECAUSE THERE WAS NO RESISTANCE. THE 2 HANDLES AND 2 RELOADS WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED.
04-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHEN DETACHING THE PULMONARY ARTERY IN THE UPPER LOBE OF THE RIGHT LUNG, AFTER INSTALLING THE RELOAD, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZE AND THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
07-12-2023 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE KNIFE DID NOT GO THROUGH AND STOPPED THE RELOAD. THERE WAS NO PATIENT INVOLVED.
22-12-2023 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
04-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC WEDGE RESECTION, WHEN PREPARING TO ASSEMBLE THE ANVIL, IT WAS FOUND THAT THE TOP OF THE HANDLE ROD WAS DAMAGED. THE OPERATOR WAS WORRIED THAT THE ANVIL COULD NOT BE INSTALLED PROPERLY, SO IT WAS REPLACED WITH THE SAME TYPE OF PRODUCT. THERE WAS NO PATIENT INJURY.
25-12-2023 | Malfunction | ethicon endo-surgery | 19-01-2024 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IN THE STOMACH TAKEDOWN, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOWER ANTERIOR RESECTION, IN RESECTION OF THE RECTUM, THE DEVICE DID NOT FIRE PROPERLY AS THE STAPLE LINE WAS DISTALLY INCOMPLETE. ADDITIONALLY, THE STAPLE FORMATION WAS ALSO INCOMPLETE. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED.
31-12-2023 | Injury | ethicon endo-surgery | 19-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLOSTOMY CLOSURE PROCEDURE THE CDH29P USED TO CREATE ANASTOMOSES USING DOUBLE STAPLE TECHNIQUE. NOTED THE ANVIL PENETRATED JUST ABOVE THE STAPLE LINE NOT THROUGH IT AND TISSUE WAS SKELETONIZED PRIOR TO CREATING THE ANASTOMOSIS. ONCE FIRED, THE STAPLE LINE LOOKED CLEAN WITH GOOD SUPPLY, NO TENSION, THE ANASTOMOSIS WAS TESTED VIA AIR LEAK TEST AND PASSED WITH CLEAN DONUTS. ON (B)(6) THE PATIENT WAS BROUGHT BACK INTO HOSPITAL WITH 600CC OF FLUID REMOVED FROM THE ABDOMEN WITH POOR URINE OUTPUT. THEY KEPT THE PATIENT IN THE HOSPITAL FOR OBSERVATION AND BROUGHT HIM BACK INTO SURGERY ON (B)(6) TO UNCOVER THAT STAPLES WERE PRESENT, THE TISSUE WAS PRESENT, BUT THE AREA WAS SEPARATED ALMOST AS IF ¿IT WAS FLOPPING IN THE BREEZE¿ PER SURGEON'S DESCRIPTION. THE PATIENT HAD A DIVERTED COLOSTOMY BAG.
31-12-2023 | Injury | ethicon endo-surgery | 19-01-2024 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Patient Device Interaction Problem | A01 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A COLOSTOMY CLOSURE PROCEDURE THE CDH29P USED TO CREATE ANASTOMOSES USING DOUBLE STAPLE TECHNIQUE. NOTED THE ANVIL PENETRATED JUST ABOVE THE STAPLE LINE NOT THROUGH IT AND TISSUE WAS SKELETONIZED PRIOR TO CREATING THE ANASTOMOSIS. ONCE FIRED, THE STAPLE LINE LOOKED CLEAN WITH GOOD SUPPLY, NO TENSION, THE ANASTOMOSIS WAS TESTED VIA AIR LEAK TEST AND PASSED WITH CLEAN DONUTS. ON (B)(6) THE PATIENT WAS BROUGHT BACK INTO HOSPITAL WITH 600CC OF FLUID REMOVED FROM THE ABDOMEN WITH POOR URINE OUTPUT. THEY KEPT THE PATIENT IN THE HOSPITAL FOR OBSERVATION AND BROUGHT HIM BACK INTO SURGERY ON (B)(6) TO UNCOVER THAT STAPLES WERE PRESENT, THE TISSUE WAS PRESENT, BUT THE AREA WAS SEPARATED ALMOST AS IF ¿IT WAS FLOPPING IN THE BREEZE¿ PER SURGEON'S DESCRIPTION. THE PATIENT HAD A DIVERTED COLOSTOMY BAG.
26-10-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | UNKNOWN DST TA STAPLER | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-10-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | UNKNOWN DST TA STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-10-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | UNKNOWN DST TA STAPLER | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-10-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | UNKNOWN DST TA STAPLER | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC ENDOSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF DISSOCIATING THE TISSUE, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-01-2024 | Injury | ethicon endo-surgery | 19-01-2024 | PROXIMATE*RL LINEAR STAPLER | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY THE TX 60B DID NOT FIRE OUT STAPLES. PATIENT REQUIRED A COLOSTOMY BAG. THERE WAS A 20 MINUTE DELAY FROM THE RESULT.
02-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL RECONSTRUCTION PROCEDURE, WHEN THE CIRCULAR STAPLER WAS USED FOR THE COLON AND RECTUM ANASTOMOSIS, THERE WAS A POOR STAPLED FORMATION AT A SMALL PART OF THE CIRCULAR STAPLED LINE, RESULTING IN A SMALL PORTION OF TISSUE BEING LOST. TO RESOLVE THE ISSUE, THE SURGEON PERFORMED A MANUAL SUTURE ABOVE THE STAPLED LINE FORMATION TO PREVENT LEAKS FROM OCCURRING.
02-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | EEA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTESTINAL RECONSTRUCTION PROCEDURE, WHEN THE CIRCULAR STAPLER WAS USED FOR THE COLON AND RECTUM ANASTOMOSIS, THERE WAS A POOR STAPLED FORMATION AT A SMALL PART OF THE CIRCULAR STAPLED LINE, RESULTING IN A SMALL PORTION OF TISSUE BEING LOST. TO RESOLVE THE ISSUE, THE SURGEON PERFORMED A MANUAL SUTURE ABOVE THE STAPLED LINE FORMATION TO PREVENT LEAKS FROM OCCURRING.
01-01-2024 | Malfunction | ethicon endo-surgery | 19-01-2024 | PROXIMATE*PPH PROCEDURE SET | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY THE PPH STAPLER CUT THROUGH TISSUE BUT DID NOT STAPLE. IT WAS PLACED ASIDE, AND THE PATIENT HAD UNKNOWN COMPLICATIONS.
03-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF DETACHING BRONCHUS (BRONCHIAL TISSUE), THE JAWS DID NOT CLOSE COMPLETELY. IT WAS ALSO OBSERVED THAT THERE WAS A GAP ON THE CONNECTION BETWEEN THE RELOAD AND THE HANDLE. THE SURGEON STILL TRIED FIRING THE DEVICE, BUT IT DID NOT OPERATE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF DETACHING BRONCHUS (BRONCHIAL TISSUE), THE JAWS DID NOT CLOSE COMPLETELY. IT WAS ALSO OBSERVED THAT THERE WAS A GAP ON THE CONNECTION BETWEEN THE RELOAD AND THE HANDLE. THE SURGEON STILL TRIED FIRING THE DEVICE, BUT IT DID NOT OPERATE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF DETACHING BRONCHUS (BRONCHIAL TISSUE), THE JAWS DID NOT CLOSE COMPLETELY. IT WAS ALSO OBSERVED THAT THERE WAS A GAP ON THE CONNECTION BETWEEN THE RELOAD AND THE HANDLE. THE SURGEON STILL TRIED FIRING THE DEVICE, BUT IT DID NOT OPERATE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF DETACHING BRONCHUS (BRONCHIAL TISSUE), THE JAWS DID NOT CLOSE COMPLETELY. IT WAS ALSO OBSERVED THAT THERE WAS A GAP ON THE CONNECTION BETWEEN THE RELOAD AND THE HANDLE. THE SURGEON STILL TRIED FIRING THE DEVICE, BUT IT DID NOT OPERATE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF DETACHING BRONCHUS (BRONCHIAL TISSUE), THE JAWS DID NOT CLOSE COMPLETELY. IT WAS ALSO OBSERVED THAT THERE WAS A GAP ON THE CONNECTION BETWEEN THE RELOAD AND THE HANDLE. THE SURGEON STILL TRIED FIRING THE DEVICE, BUT IT DID NOT OPERATE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, AT THE TIME OF DETACHING BRONCHUS (BRONCHIAL TISSUE), THE JAWS DID NOT CLOSE COMPLETELY. IT WAS ALSO OBSERVED THAT THERE WAS A GAP ON THE CONNECTION BETWEEN THE RELOAD AND THE HANDLE. THE SURGEON STILL TRIED FIRING THE DEVICE, BUT IT DID NOT OPERATE. A NEW RELOAD AND HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC EXPLORATORY LAPAROTOMY, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
02-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY AND WEDGE RESECTION, DURING LUNG RESECTION ACROSS THE PARENCHYMA, WHILE USING THE FIRST HANDLE WITH A BLACK 60MM RELOAD, THE SURGEON TRIED TO STAPLE BUT FOUND OUT THAT IT WAS EXTREMELY DIFFICULT TO STAPLE THE TISSUE. THE GREEN BUTTON WAS UNABLE TO BE PRESSED. THE SECOND RELOAD DID NOT LOCK AND THE GUN WAS UNUSABLE. A SECOND HANDLE WAS USED BUT THE SAME ISSUE OCCURRED. A NEW STAPLER FROM A DIFFERENT MANUFACTURER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | UNKNOWN ENDO GIA SULU | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY AND WEDGE RESECTION, DURING LUNG RESECTION ACROSS THE PARENCHYMA, WHILE USING THE FIRST HANDLE WITH A BLACK 60MM RELOAD, THE SURGEON TRIED TO STAPLE BUT FOUND OUT THAT IT WAS EXTREMELY DIFFICULT TO STAPLE THE TISSUE. THE GREEN BUTTON WAS UNABLE TO BE PRESSED. THE SECOND RELOAD DID NOT LOCK AND THE GUN WAS UNUSABLE. A SECOND HANDLE WAS USED BUT THE SAME ISSUE OCCURRED. A NEW STAPLER FROM A DIFFERENT MANUFACTURER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VATS (VIDEO-ASSISTED THORACOSCOPIC SURGERY) RIGHT UPPER LOBE WEDGE OF THE LUNG, WHEN THE DEVICE WAS REMOVED AFTER STAPLING A WEDGE-SHAPED SECTION OF THE LUNG AROUND THE TUMOR, IT WAS OBSERVED THAT SOME TISSUE WAS INADVERTENTLY CAUGHT AT THE DISTAL JAW OF THE RELOAD. NO INTERVENTION WAS PERFORMED. THERE WAS NO PATIENT INJURY.
19-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, IN A LOBECTOMY, WHEN THE RELOAD WAS RESECTING THE LUNG TISSUE, THE STAPLES IN THE CARTRIDGE BURST OUT. THE STAPLES WERE POORLY FORMED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
04-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC LUNG CANCER RADICAL SURGERY, WHILE CLOSING THE LOBE, THE RELOAD COULD NOT BE FIRED AND THE STAPLES DID NOT COME OUT. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE COMPLETELY. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC LUNG CANCER SURGERY, WHEN CLOSING THE LOBE, THE STAPLER DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY AND WEDGE RESECTION, DURING LUNG RESECTION ACROSS THE PARENCHYMA, WHILE USING THE FIRST HANDLE WITH A BLACK 60MM RELOAD, THE SURGEON TRIED TO STAPLE BUT FOUND OUT THAT IT WAS EXTREMELY DIFFICULT TO STAPLE THE TISSUE. THE GREEN BUTTON WAS UNABLE TO BE PRESSED. THE SECOND RELOAD DID NOT LOCK AND THE GUN WAS UNUSABLE. A SECOND HANDLE WAS USED BUT THE SAME ISSUE OCCURRED. A NEW STAPLER FROM A DIFFERENT MANUFACTURER WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
02-01-2024 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT UPPER LOBECTOMY AND WEDGE RESECTION, DURING LUNG RESECTION ACROSS THE PARENCHYMA, WHILE USING THE FIRST HANDLE WITH A BLACK 60MM RELOAD, THE SURGEON TRIED TO STAPLE BUT FOUND OUT THAT IT WAS EXTREMELY DIFFICULT TO STAPLE THE TISSUE. THE GREEN BUTTON WAS UNABLE TO BE PRESSED. THE SECOND RELOAD DID NOT LOCK AND THE GUN WAS UNUSABLE. A SECOND HANDLE WAS USED BUT THE SAME ISSUE OCCURRED. A NEW STAPLER FROM A DIFFERENT MANUFACTURER WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
31-12-2023 | Malfunction | covidien lp llc north haven | 19-01-2024 | RELIATACK | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC HERNIA REPAIR, UPON FIXATION, THE TACKS WOULD NOT LOAD CORRECTLY AND WAS UNABLE TO LOAD. ANOTHER DEVICE WAS OPENED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2024 | Malfunction | covidien lp llc north haven | 19-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC EXPLORATORY LAPAROTOMY, AFTER PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, THE DEVICE DID NOT FIRE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
22-12-2023 | Malfunction | covidien lp llc north haven | 19-01-2024 | TRI-STAPLE 2.0 | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC, WHEN TRYING TO CONNECT THE PIECES, THERE WAS A RESISTANCE, AND A SMALL PIECE OF THE DEVICE HAS BROKEN OFF. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
22-12-2023 | Malfunction | covidien lp llc north haven | 19-01-2024 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC, WHEN TRYING TO CONNECT THE PIECES, THERE WAS A RESISTANCE, AND A SMALL PIECE OF THE DEVICE HAS BROKEN OFF. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
22-12-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC, WHEN TRYING TO CONNECT THE PIECES, THERE WAS A RESISTANCE, AND A SMALL PIECE OF THE DEVICE HAS BROKEN OFF. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
22-12-2023 | Malfunction | us surgical puerto rico | 19-01-2024 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACIC, WHEN TRYING TO CONNECT THE PIECES, THERE WAS A RESISTANCE, AND A SMALL PIECE OF THE DEVICE HAS BROKEN OFF. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE FIRST RELOAD HAD AN UNKNOWN ISSUE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED. THERE WAS NO PATIENT INJURY.
22-01-2024 | Malfunction | ethicon endo-surgery | 22-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS DID NOT OPEN AFTER FIRING. THE DEVICE WAS REMOVED FROM THE TROCAR AND CLEANED, BUT THE JAWS COULD NOT BE OPENED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE FIRST RELOAD FAILED TO CUT THE TISSUE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE FIRST RELOAD FAILED TO CUT THE TISSUE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Mechanics Altered | A0507 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE FIRST RELOAD FAILED TO CUT THE TISSUE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE FIRST RELOAD FAILED TO CUT THE TISSUE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE FIRST RELOAD FAILED TO CUT THE TISSUE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE FIRST RELOAD FAILED TO CUT THE TISSUE. ON THE SECOND RELOAD, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ONLY THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOBECTOMY PROCEDURE, WHILE STAPLING THE PULMONARY ARTERY (PA), THE DEVICE FIRED. HOWEVER, WHILE RETRACTING THE I-BEAM/KNIFE BLADE THE TABS STOPPED SHORT OF RELEASING AND OPENING THE RELOAD; IT LOCKED ON TISSUE. THE SURGEON CLAMPED THE PA AND USED A CLAMP/GRASPER TO PRY OPEN THE RELOAD ON THE DISTAL END. THE SURGEON USED ENOUGH FORCE TO SLIGHTLY OPEN THE RELOAD AT WHICH POINT THE STAPLED PA FELL OUT OF THE JAWS OF THE RELOAD. IT WAS OBSERVED THAT THE STAPLES FORMED CORRECTLY. A NEW HANDLE AND RELOADS WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOBECTOMY PROCEDURE, WHILE STAPLING THE PULMONARY ARTERY (PA), THE DEVICE FIRED. HOWEVER, WHILE RETRACTING THE I-BEAM/KNIFE BLADE THE TABS STOPPED SHORT OF RELEASING AND OPENING THE RELOAD; IT LOCKED ON TISSUE. THE SURGEON CLAMPED THE PA AND USED A CLAMP/GRASPER TO PRY OPEN THE RELOAD ON THE DISTAL END. THE SURGEON USED ENOUGH FORCE TO SLIGHTLY OPEN THE RELOAD AT WHICH POINT THE STAPLED PA FELL OUT OF THE JAWS OF THE RELOAD. IT WAS OBSERVED THAT THE STAPLES FORMED CORRECTLY. A NEW HANDLE AND RELOADS WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOBECTOMY PROCEDURE, WHILE STAPLING THE PULMONARY ARTERY (PA), THE DEVICE FIRED. HOWEVER, WHILE RETRACTING THE I-BEAM/KNIFE BLADE THE TABS STOPPED SHORT OF RELEASING AND OPENING THE RELOAD; IT LOCKED ON TISSUE. THE SURGEON CLAMPED THE PA AND USED A CLAMP/GRASPER TO PRY OPEN THE RELOAD ON THE DISTAL END. THE SURGEON USED ENOUGH FORCE TO SLIGHTLY OPEN THE RELOAD AT WHICH POINT THE STAPLED PA FELL OUT OF THE JAWS OF THE RELOAD. IT WAS OBSERVED THAT THE STAPLES FORMED CORRECTLY. A NEW HANDLE AND RELOADS WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN LOBECTOMY PROCEDURE, WHILE STAPLING THE PULMONARY ARTERY (PA), THE DEVICE FIRED. HOWEVER, WHILE RETRACTING THE I-BEAM/KNIFE BLADE THE TABS STOPPED SHORT OF RELEASING AND OPENING THE RELOAD; IT LOCKED ON TISSUE. THE SURGEON CLAMPED THE PA AND USED A CLAMP/GRASPER TO PRY OPEN THE RELOAD ON THE DISTAL END. THE SURGEON USED ENOUGH FORCE TO SLIGHTLY OPEN THE RELOAD AT WHICH POINT THE STAPLED PA FELL OUT OF THE JAWS OF THE RELOAD. IT WAS OBSERVED THAT THE STAPLES FORMED CORRECTLY. A NEW HANDLE AND RELOADS WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, ANOTHER DEVICE WAS USED.
13-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | TA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN STAPLING, NOT ALL OF THE TISSUE WAS CAUGHT AND SUTURING WAS NOT POSSIBLE.
06-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, WHEN SEVERING LUNG TISSUE , AFTER FIRED, NOTED SOME STAPLES MALFORMED . CHANGED ANOTHER ONE TO USE , THEY STILL HAD THE SAME PROBLEM. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-11-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , ON FIRST FIRING AT THE TRANSECTION OF STOMACH, THE SUTURE AT END OF RELOAD NEVER GOT CUT BY THE KNIFE BLADE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO CUT THE SUTURE AT DISTAL END OF RELOAD WITH SHEARS. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT CONFIRMED THAT THE SYSTEM TERMINATED FIRING AT THE DETERMINED END STOP ZONE AS INTENDED BY THE SOFTWARE BUT PRIOR TO THE DISTAL SUTURE BEING RELEASED. THE CAUSE OF THE SYSTEM STOPPING THE FIRING PRIOR TO RELEASING THE DISTAL SUTURE COULD NOT BE DETERMINED.
28-11-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , ON FIRST FIRING AT THE TRANSECTION OF STOMACH, THE SUTURE AT END OF RELOAD NEVER GOT CUT BY THE KNIFE BLADE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO CUT THE SUTURE AT DISTAL END OF RELOAD WITH SHEARS. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT CONFIRMED THAT THE SYSTEM TERMINATED FIRING AT THE DETERMINED END STOP ZONE AS INTENDED BY THE SOFTWARE BUT PRIOR TO THE DISTAL SUTURE BEING RELEASED. THE CAUSE OF THE SYSTEM STOPPING THE FIRING PRIOR TO RELEASING THE DISTAL SUTURE COULD NOT BE DETERMINED.
22-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL COULD NOT FIRE AND WITHOUT MOTOR SOUND. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL COULD NOT FIRE AND WITHOUT MOTOR SOUND. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-01-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-11-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , ON FIRST FIRING AT THE TRANSECTION OF STOMACH, THE SUTURE AT END OF RELOAD NEVER GOT CUT BY THE KNIFE BLADE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO CUT THE SUTURE AT DISTAL END OF RELOAD WITH SHEARS. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT CONFIRMED THAT THE SYSTEM TERMINATED FIRING AT THE DETERMINED END STOP ZONE AS INTENDED BY THE SOFTWARE BUT PRIOR TO THE DISTAL SUTURE BEING RELEASED. THE CAUSE OF THE SYSTEM STOPPING THE FIRING PRIOR TO RELEASING THE DISTAL SUTURE COULD NOT BE DETERMINED.
28-11-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | TRI-STAPLE 2.0 | Separation Failure | A150301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SLEEVE GASTRECTOMY , ON FIRST FIRING AT THE TRANSECTION OF STOMACH, THE SUTURE AT END OF RELOAD NEVER GOT CUT BY THE KNIFE BLADE. TO RESOLVE THE ISSUE, THE SURGEON HAD TO CUT THE SUTURE AT DISTAL END OF RELOAD WITH SHEARS. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT CONFIRMED THAT THE SYSTEM TERMINATED FIRING AT THE DETERMINED END STOP ZONE AS INTENDED BY THE SOFTWARE BUT PRIOR TO THE DISTAL SUTURE BEING RELEASED. THE CAUSE OF THE SYSTEM STOPPING THE FIRING PRIOR TO RELEASING THE DISTAL SUTURE COULD NOT BE DETERMINED.
17-01-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | GST GREEN RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | GOLD RELOAD FOR ECHELON 60 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, NOTED THE CARTRIDGE DAMAGED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-11-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, AFTER THE TISSUE WAS DISSOCIATED AND BEFORE DETACHING THE PULMONARY VEIN, THERE WAS DIFFICULTY LOADING THE RELOAD TO THE HANDLE. TO RESOLVE THE ISSUE, ANOTHER HANDLE OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE.
14-11-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, AFTER THE TISSUE WAS DISSOCIATED AND BEFORE DETACHING THE PULMONARY VEIN, THERE WAS DIFFICULTY LOADING THE RELOAD TO THE HANDLE. TO RESOLVE THE ISSUE, ANOTHER HANDLE OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S EVALUATION FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | UNK_ECHELON CIRCULAR POWERED STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | UNK_ECHELON CIRCULAR POWERED STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Diarrhea | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | UNK_ECHELON CIRCULAR POWERED STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | UNK_ECHELON CIRCULAR POWERED STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | UNK_ECHELON CIRCULAR POWERED STAPLER | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
18-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY (THORACOSCOPIC LUNG SEGMENTECTOMY), AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. CHANGED ANOTHER ONE TO USE, THEY STILL HAD THE SAME PROBLEM. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY (THORACOSCOPIC LUNG SEGMENTECTOMY), AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. CHANGED ANOTHER ONE TO USE, THEY STILL HAD THE SAME PROBLEM. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY (THORACOSCOPIC LUNG SEGMENTECTOMY), AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. CHANGED ANOTHER ONE TO USE, THEY STILL HAD THE SAME PROBLEM. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORASCOPIC PULMONARY RESECTION FOR UPPER LOBE LUNG CANCER, FRAGMENTS OF THE YELLOW TAB REMAINED IN THE CHEST CAVITY, AND DAMAGE TO THE YELLOW TAB OF THE CARTRIDGE WAS CONFIRMED. THE FRAGMENTS WERE REMOVED WITH FORCEPS. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC THORASCOPIC PULMONARY RESECTION FOR UPPER LOBE LUNG CANCER, FRAGMENTS OF THE YELLOW TAB REMAINED IN THE CHEST CAVITY, AND DAMAGE TO THE YELLOW TAB OF THE CARTRIDGE WAS CONFIRMED. THE FRAGMENTS WERE REMOVED WITH FORCEPS. THERE WAS NO PATIENT INJURY.
13-10-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. USED CLIP AND ELECTROCOAGULATION ROD TO STOP OOZING. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-10-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. USED CLIP AND ELECTROCOAGULATION ROD TO STOP OOZING. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-10-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. USED CLIP AND ELECTROCOAGULATION ROD TO STOP OOZING. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-10-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. USED CLIP AND ELECTROCOAGULATION ROD TO STOP OOZING. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE BLADE COULDN'T RETURN DESPITE OF PRESSING THE REVERSE BUTTON SO SURGEON USED MANUAL HANDLE FOR OPENING THE JAW. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE BLADE COULDN'T RETURN DESPITE OF PRESSING THE REVERSE BUTTON SO SURGEON USED MANUAL HANDLE FOR OPENING THE JAW. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE BLADE COULDN'T RETURN DESPITE OF PRESSING THE REVERSE BUTTON SO SURGEON USED MANUAL HANDLE FOR OPENING THE JAW. NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE.
11-01-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER TWO TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Diarrhea | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
06-01-2024 | Injury | ethicon endo-surgery | 22-01-2024 | PROXIMATE*RL LINEAR STAPLER (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, USED ETHICON POWERED CIRCULAR TO REMOVE A SECTION OF THE COLON DEEP IN THE PELVIS ALONG WITH THE RECTUM AND A PORTION OF THE ANUS. PROCEDURE WAS BEING PERFORMED TO BOTH ADDRESS DIVERTICULITIS AND A SUSPECTED LEAK SITE. ALSO USED A LINEAR STAPLER DURING THE SAME PROCEDURE TO MAKE A SECOND, UNPLANNED ANASTOMOSIS WHERE THE BOWEL HAD GROWN INTO SURGICAL MESH FROM A PRIOR BLADDER-RELATED PROCEDURE. SURGEON COMMENTED ALL STAPLERS WORKED GREAT DURING THE PROCEDURE. PATIENT STARTED HAVING EXTREME PAIN (10/10 AS STATED BY THE PATIENT) WITHIN 48 HOURS POST-SURGERY. CT SCAN REVEALED A LEAK IN THE REGION OF THE CIRCULAR STAPLER LINE. PATIENT HAD LIQUID, BLOODY STOOL. SURGEON WAS UNABLE TO FIND THE EXACT SOURCE OF THE LEAK DUE TO HOW SMALL IT WAS, THOUGH THE LEAK WAS EITHER FROM OR IMMEDIATELY ADJACENT TO THE CIRCULAR STAPLE LINE. SURGEON STATED THE SECOND UNPLANNED ANASTOMOSIS LOOKED GREAT AND WAS CAUSING NO ISSUES.
11-01-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER TWO TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | UNKNOWN TRI STAPLE PRODUCT | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, IN AN OPEN LOWER ANTERIOR RESECTION, DURING RESECTION, THE DEVICE DID NOT FIRE. IMPROPER B SHAPED STAPLE FORMATION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
19-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | ILS 29MM, CURVED | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTOCOLECTOMY PROCEDURE, BEFORE USED ON THE PATIENT, NOTED THE PACKING WAS DAMAGE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
11-01-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER TWO TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2024 | Malfunction | ethicon endo-surgery | 22-01-2024 | CIRCULAR MECH STD 29MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE BLADE FAILED TO CUT THE TISSUE THOROUGHLY. ABOUT 1/3 PART OF TISSUE DID NOT CUT. THE DEVICE WAS REPLACED BY A NEW ONE, AND THE SURGERY WAS COMPLETED SUCCESSFULLY.
06-12-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE VENTRAL INCISIONAL HERNIA REPAIR WITH MESH, THE STAPLER FAILED TO RELEASE THE LOAD WHEN CLOSING ON A PORTION OF SMALL BOWEL ANASTOMOSIS.
06-12-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE VENTRAL INCISIONAL HERNIA REPAIR WITH MESH, THE STAPLER FAILED TO RELEASE THE LOAD WHEN CLOSING ON A PORTION OF SMALL BOWEL ANASTOMOSIS.
23-12-2023 | Malfunction | ethicon endo-surgery | 22-01-2024 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-09-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Display or Visual Feedback Problem | A0902 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, AFTER THE POWER HANDLE WAS CHARGED, IT DISPLAYED 'VER' ON THE SCREEN AND DID NOT WORK WHEN TAKEN OFF THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE THREAD LOCKER APPLIED TO THE SCREWS WAS NOT ACTIVATED PROPERLY.
06-09-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, AFTER THE POWER HANDLE WAS CHARGED, IT DISPLAYED 'VER' ON THE SCREEN AND DID NOT WORK WHEN TAKEN OFF THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE THREAD LOCKER APPLIED TO THE SCREWS WAS NOT ACTIVATED PROPERLY.
06-09-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | SIGNIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, AFTER THE POWER HANDLE WAS CHARGED, IT DISPLAYED 'VER' ON THE SCREEN AND DID NOT WORK WHEN TAKEN OFF THE CHARGER. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE THREAD LOCKER APPLIED TO THE SCREWS WAS NOT ACTIVATED PROPERLY.
06-12-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE VENTRAL INCISIONAL HERNIA REPAIR WITH MESH, THE STAPLER FAILED TO RELEASE THE LOAD WHEN CLOSING ON A PORTION OF SMALL BOWEL ANASTOMOSIS.
06-12-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE VENTRAL INCISIONAL HERNIA REPAIR WITH MESH, THE STAPLER FAILED TO RELEASE THE LOAD WHEN CLOSING ON A PORTION OF SMALL BOWEL ANASTOMOSIS.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC HERNIA, DURING MESH FIXATION, THERE WERE THREE TACKERS USED DURING THE PROCEDURE. ONLY HALF OF THE FIRED TACKS IN THREE TACKERS PENETRATED THE ABDOMINAL WALL CORRECTLY. ONE OF THE THREE TACKERS ALSO BROKE IN THE PATIENT DURING THE PROCEDURE. THE SHAFT CRACKED FELL INTO THE CAVITY. AFTER IDENTIFYING THE CRACK DURING THE LAPAROSCOPY, THE DEVICE WAS REMOVED FROM THE ABDOMINAL CAVITY AS QUICKLY AS POSSIBLE, AND TRACTION IN THE TROCAR USING STANDARD FENESTRATED FORCEPS WHICH HAS FINISHED CUTTING, TO RECOVER THE AXIS OF THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC HERNIA, DURING MESH FIXATION, THERE WERE THREE TACKERS USED DURING THE PROCEDURE. ONLY HALF OF THE FIRED TACKS IN THREE TACKERS PENETRATED THE ABDOMINAL WALL CORRECTLY. ONE OF THE THREE TACKERS ALSO BROKE IN THE PATIENT DURING THE PROCEDURE. THE SHAFT CRACKED FELL INTO THE CAVITY. AFTER IDENTIFYING THE CRACK DURING THE LAPAROSCOPY, THE DEVICE WAS REMOVED FROM THE ABDOMINAL CAVITY AS QUICKLY AS POSSIBLE, AND TRACTION IN THE TROCAR USING STANDARD FENESTRATED FORCEPS WHICH HAS FINISHED CUTTING, TO RECOVER THE AXIS OF THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ABSORBATACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC HERNIA, DURING MESH FIXATION, THERE WERE THREE TACKERS USED DURING THE PROCEDURE. ONLY HALF OF THE FIRED TACKS IN THREE TACKERS PENETRATED THE ABDOMINAL WALL CORRECTLY. ONE OF THE THREE TACKERS ALSO BROKE IN THE PATIENT DURING THE PROCEDURE. THE SHAFT CRACKED FELL INTO THE CAVITY. AFTER IDENTIFYING THE CRACK DURING THE LAPAROSCOPY, THE DEVICE WAS REMOVED FROM THE ABDOMINAL CAVITY AS QUICKLY AS POSSIBLE, AND TRACTION IN THE TROCAR USING STANDARD FENESTRATED FORCEPS WHICH HAS FINISHED CUTTING, TO RECOVER THE AXIS OF THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC HERNIA, DURING MESH FIXATION, THERE WERE THREE TACKERS USED DURING THE PROCEDURE. ONLY HALF OF THE FIRED TACKS IN THREE TACKERS PENETRATED THE ABDOMINAL WALL CORRECTLY. ONE OF THE THREE TACKERS ALSO BROKE IN THE PATIENT DURING THE PROCEDURE. THE SHAFT CRACKED FELL INTO THE CAVITY. AFTER IDENTIFYING THE CRACK DURING THE LAPAROSCOPY, THE DEVICE WAS REMOVED FROM THE ABDOMINAL CAVITY AS QUICKLY AS POSSIBLE, AND TRACTION IN THE TROCAR USING STANDARD FENESTRATED FORCEPS WHICH HAS FINISHED CUTTING, TO RECOVER THE AXIS OF THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ABSORBATACK | Loss of or Failure to Bond | A040102 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, LAPAROSCOPIC HERNIA, DURING MESH FIXATION, THERE WERE THREE TACKERS USED DURING THE PROCEDURE. ONLY HALF OF THE FIRED TACKS IN THREE TACKERS PENETRATED THE ABDOMINAL WALL CORRECTLY. ONE OF THE THREE TACKERS ALSO BROKE IN THE PATIENT DURING THE PROCEDURE. THE SHAFT CRACKED FELL INTO THE CAVITY. AFTER IDENTIFYING THE CRACK DURING THE LAPAROSCOPY, THE DEVICE WAS REMOVED FROM THE ABDOMINAL CAVITY AS QUICKLY AS POSSIBLE, AND TRACTION IN THE TROCAR USING STANDARD FENESTRATED FORCEPS WHICH HAS FINISHED CUTTING, TO RECOVER THE AXIS OF THE DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
22-11-2023 | Injury | intuitive surgical | 22-01-2024 | SUREFORM | Insufficient Information | E2401 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED ROUX-EN-Y PROCEDURE, THERE WAS UNEXPECTED BLEEDING FROM THE STAPLE LINE. THE PROCEDURE WAS COMPLETED AS PLANNED, BUT THE PATIENT REQUIRED A BLOOD TRANSFUSION.
22-12-2023 | Injury | ethicon endo-surgery | 22-01-2024 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE SURGEON FELT IT NOT SMOOTH ON THE 6TH FIRE. ON THE 7TH FIRE, NOTED THE KNIFE COULD NOT GO FARTHER AFTER FIRED 1/2. THE SURGEON TRIED THE REVERSE BUTTON, BUT THE KNIFE COULD NOT REVERSE. PULLED THE DEVICE FROM TISSUE FINALLY AND CHANGED TO OPEN OPERATION. USED SUTURE SEWING TO COMPLETE THE SURGERY.
22-12-2023 | Injury | ethicon endo-surgery | 22-01-2024 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE SURGEON FELT IT NOT SMOOTH ON THE 6TH FIRE. ON THE 7TH FIRE, NOTED THE KNIFE COULD NOT GO FARTHER AFTER FIRED 1/2. THE SURGEON TRIED THE REVERSE BUTTON, BUT THE KNIFE COULD NOT REVERSE. PULLED THE DEVICE FROM TISSUE FINALLY AND CHANGED TO OPEN OPERATION. USED SUTURE SEWING TO COMPLETE THE SURGERY.
22-12-2023 | Injury | ethicon endo-surgery | 22-01-2024 | FLEX 60 ARTICULATING | Mechanical Jam | A0506 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, THE SURGEON FELT IT NOT SMOOTH ON THE 6TH FIRE. ON THE 7TH FIRE, NOTED THE KNIFE COULD NOT GO FARTHER AFTER FIRED 1/2. THE SURGEON TRIED THE REVERSE BUTTON, BUT THE KNIFE COULD NOT REVERSE. PULLED THE DEVICE FROM TISSUE FINALLY AND CHANGED TO OPEN OPERATION. USED SUTURE SEWING TO COMPLETE THE SURGERY.
27-12-2023 | Injury | intuitive surgical | 22-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED IN A SOCIAL MEDIA POST, THAT AFTER A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THERE WAS BLEEDING FROM THE STAPLE LINE ON THE STOMACH. THE AMOUNT OF BLEEDING IN UNKNOWN BUT WAS DESCRIBED AS A SMALL AMOUNT. THE BLEEDING OCCURRED ON THE FIRST FIRE USING A SUREFORM 60 STAPLER INSTRUMENT, WITH A BLUE RELOAD ON STOMACH TISSUE. THE SURGEON STATED SPENDING ADDITIONAL SURGERY TIME SUTURE LIGATING TO RESOLVE THE BLEEDING. THE PROCEDURE WAS COMPLETED ROBOTICALLY AND THERE WERE NO POST-OPERATIVE COMPLICATIONS. HISTORICALLY, THE SURGEON USED A SUREFORM 60 STAPLER INSTRUMENT WITH A GREEN RELOAD THAT CONTAINED BUTTRESS MATERIAL. ANOTHER BARIATRIC SURGEON ADVISED TO NOT USE BUTTRESS MATERIAL; THE SURGEON HAS SINCE CHANGED TO USING A BLUE RELOAD. THE SURGEON CHOSE TO SWITCH TO A BLUE RELOAD BECAUSE IT WAS THE MOST COMPARABLE OPTION WHEN USING A 3RD PARTY LAPAROSCOPIC STAPLER AND A PURPLE STAPLE LOAD. THE CLINICAL SALES REPRESENTATIVE HAS DISCUSSED WITH THE SURGEON, TRANSITIONING FROM A BLUE RELOAD TO A WHITE RELOAD AS THE TISSUE THAT THE RELOAD IS BEING USED ON IS THIN.
27-12-2023 | Injury | intuitive surgical | 22-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED IN A SOCIAL MEDIA POST, THAT AFTER A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THERE WAS BLEEDING FROM THE STAPLE LINE ON THE STOMACH. THE AMOUNT OF BLEEDING IN UNKNOWN BUT WAS DESCRIBED AS A SMALL AMOUNT. THE BLEEDING OCCURRED ON THE FIRST FIRE USING A SUREFORM 60 STAPLER INSTRUMENT, WITH A BLUE RELOAD ON STOMACH TISSUE. THE SURGEON STATED SPENDING ADDITIONAL SURGERY TIME SUTURE LIGATING TO RESOLVE THE BLEEDING. THE PROCEDURE WAS COMPLETED ROBOTICALLY AND THERE WERE NO POST-OPERATIVE COMPLICATIONS. HISTORICALLY, THE SURGEON USED A SUREFORM 60 STAPLER INSTRUMENT WITH A GREEN RELOAD THAT CONTAINED BUTTRESS MATERIAL. ANOTHER BARIATRIC SURGEON ADVISED TO NOT USE BUTTRESS MATERIAL; THE SURGEON HAS SINCE CHANGED TO USING A BLUE RELOAD. THE SURGEON CHOSE TO SWITCH TO A BLUE RELOAD BECAUSE IT WAS THE MOST COMPARABLE OPTION WHEN USING A 3RD PARTY LAPAROSCOPIC STAPLER AND A PURPLE STAPLE LOAD. THE CLINICAL SALES REPRESENTATIVE HAS DISCUSSED WITH THE SURGEON, TRANSITIONING FROM A BLUE RELOAD TO A WHITE RELOAD AS THE TISSUE THAT THE RELOAD IS BEING USED ON IS THIN.
14-12-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHEN THE STAPLER WAS BEING FIRED ACROSS TISSUE, THE HANDLE WOULD NOT ACCEPT THE RELOAD DESPITE HAVING FIRED MULTIPLE RELOADS WITHOUT INCIDENT. THE STAPLES FIRED, BUT DID NOT ALL STAPLE STRAIGHT, SO SOME JAMMED AND THE KNIFE COULD NOT CUT THROUGH THE STAPLE LINE. ANOTHER HANDLE AND RELOAD WERE OPENED AND USED, AND THE CASE WAS COMPLETED.
14-12-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | UNKNOWN ENDO GIA SULU | Failure to Cut | A050702 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHEN THE STAPLER WAS BEING FIRED ACROSS TISSUE, THE HANDLE WOULD NOT ACCEPT THE RELOAD DESPITE HAVING FIRED MULTIPLE RELOADS WITHOUT INCIDENT. THE STAPLES FIRED, BUT DID NOT ALL STAPLE STRAIGHT, SO SOME JAMMED AND THE KNIFE COULD NOT CUT THROUGH THE STAPLE LINE. ANOTHER HANDLE AND RELOAD WERE OPENED AND USED, AND THE CASE WAS COMPLETED.
14-12-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | UNKNOWN ENDO GIA SULU | Component or Accessory Incompatibility | A1701 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHEN THE STAPLER WAS BEING FIRED ACROSS TISSUE, THE HANDLE WOULD NOT ACCEPT THE RELOAD DESPITE HAVING FIRED MULTIPLE RELOADS WITHOUT INCIDENT. THE STAPLES FIRED, BUT DID NOT ALL STAPLE STRAIGHT, SO SOME JAMMED AND THE KNIFE COULD NOT CUT THROUGH THE STAPLE LINE. ANOTHER HANDLE AND RELOAD WERE OPENED AND USED, AND THE CASE WAS COMPLETED.
05-01-2024 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TOTAL GASTRECTOMY PROCEDURE, WHILE DETACHING THE STOMACH, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE AFTER PRESSING THE GREEN BUTTON. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT UPPER LOBECTOMY, WHEN DEALING WITH THE INTERLOBAR FISSURE IN THE RIGHT UPPER LOBE OF THE LUNG, A PURPLE STAPLE HANDLE WAS USED, BUT IT COULD NOT BE FIRED. THE SURGEON COULD PRESS THE GREEN BUTTON, BUT COULD NOT SQUEEZE THE HANDLE. AFTER REPLACING WITH A NEW HANDLE, THE FIRING WAS COMPLETED. THERE WAS NO PATIENT INJURY.
03-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING COLECTOMY, THE RELOAD DID NOT KNOW HOW TO CLOSE AND STAPLER DID NOT FIRE. THE RELOAD WAS REPLACED AND SAME HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, AFTER THE THIRD FIRING ACROSS A FISSURE ON THE LUNG, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON PULLED BACK ON THE BLACK RETURN KNOBS A FEW TIMES AND THE RELOAD THEN OPENED. THERE WAS NO PATIENT INJURY.
24-12-2023 | Malfunction | intuitive surgical | 22-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED IN A SOCIAL MEDIA POST, THAT AFTER A DA VINCI-ASSISTED PROCEDURE, THERE WERE MALFORMED STAPLES AFTER USING A SUREFORM STAPLER WITH A WHITE RELOAD. AN ERROR MESSAGE WAS RECEIVED, TISSUE TOO THICK. ONCE THE STAPLER WAS REMOVED, THE MALFORMED STAPLES WERE SEEN. HOW THE PROCEDURE WAS COMPLETED, PATIENT¿S STATUS, OR ANY INTERVENTIONS TAKEN; ARE ALL UNKNOWN.
04-01-2024 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, AFTER THE THIRD FIRING ACROSS A FISSURE ON THE LUNG, THE JAWS OF THE DEVICE LOCKED ON TISSUE. THE SURGEON PULLED BACK ON THE BLACK RETURN KNOBS A FEW TIMES AND THE RELOAD THEN OPENED. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, IN AN OPEN LOWER ANTERIOR RESECTION, DURING RESECTION, THE DEVICE DID NOT FIRE. IMPROPER B SHAPED STAPLE FORMATION. THE SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES.
03-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-12-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE JAWS DID NOT CLOSE. A NEW RELOAD WAS USED ON THE SAME HANDLE TO RESOLVE THE ISSUE.
06-12-2023 | Malfunction | covidien lp llc north haven | 22-01-2024 | GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A CYSTOSCOPY WITH BILATERAL URETERAL STENT PLACEMENT, THE DEVICE WOULD NOT FULLY STAPLE.
04-01-2024 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT COLECTOMY, THE DEVICE WAS COMPLETELY BLOCKED AND CANNOT BE USED. THERE WAS NO PATIENT INJURY.
04-01-2024 | Malfunction | covidien lp llc north haven | 22-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LEFT COLECTOMY, THE DEVICE WAS COMPLETELY BLOCKED AND CANNOT BE USED. THERE WAS NO PATIENT INJURY.
03-01-2024 | Injury | us surgical puerto rico | 22-01-2024 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC BYPASS, DURING THE ANASTOMOSIS OF THE ENTERO-ENTERO, THE STAPLE LINE WAS INCOMPLETE AND THERE WAS A METHYLENE BLUE LEAKAGE FROM THE ENTIRE ANASTOMOSIS. THE ANASTOMOSIS WAS MANUALLY SUTURED TO RESOLVE THE ISSUE. THE HOSPITALIZATION WAS EXTENDED FOR ONE MORE NIGHT AND THE INCISION WAS EXTENDED.
27-12-2023 | Injury | intuitive surgical | 22-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED IN A SOCIAL MEDIA POST, THAT AFTER A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THERE WAS BLEEDING FROM THE STAPLE LINE ON THE STOMACH. THE AMOUNT OF BLEEDING IN UNKNOWN BUT WAS DESCRIBED AS A SMALL AMOUNT. THE BLEEDING OCCURRED ON THE FIRST FIRE USING A SUREFORM 60 STAPLER INSTRUMENT, WITH A BLUE RELOAD ON STOMACH TISSUE. THE SURGEON STATED SPENDING ADDITIONAL SURGERY TIME SUTURE LIGATING TO RESOLVE THE BLEEDING. THE PROCEDURE WAS COMPLETED ROBOTICALLY AND THERE WERE NO POST OPERATIVE COMPLICATIONS. HISTORICALLY, THE SURGEON USED A SUREFORM 60 STAPLER INSTRUMENT WITH A GREEN RELOAD THAT CONTAINED BUTTRESS MATERIAL. ANOTHER BARIATRIC SURGEON ADVISED TO NOT USE BUTTRESS MATERIAL; THE SURGEON HAS SINCE CHANGED TO USING A BLUE RELOAD. THE SURGEON CHOSE TO SWITCH TO A BLUE RELOAD BECAUSE IT WAS THE MOST COMPARABLE OPTION WHEN USING A 3RD PARTY LAPAROSCOPIC STAPLER AND A PURPLE STAPLE LOAD. THE CLINICAL SALES REPRESENTATIVE HAS DISCUSSED WITH THE SURGEON, TRANSITIONING FROM A BLUE RELOAD TO A WHITE RELOAD AS THE TISSUE THAT THE RELOAD IS BEING USED ON IS THIN.
27-12-2023 | Injury | intuitive surgical | 22-01-2024 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED IN A SOCIAL MEDIA POST, THAT AFTER A DA VINCI-ASSISTED SLEEVE GASTRECTOMY PROCEDURE, THERE WAS BLEEDING FROM THE STAPLE LINE ON THE STOMACH. THE AMOUNT OF BLEEDING IN UNKNOWN BUT WAS DESCRIBED AS A SMALL AMOUNT. THE BLEEDING OCCURRED ON THE FIRST FIRE USING A SUREFORM 60 STAPLER INSTRUMENT, WITH A BLUE RELOAD ON STOMACH TISSUE. THE SURGEON STATED SPENDING ADDITIONAL SURGERY TIME SUTURE LIGATING TO RESOLVE THE BLEEDING. THE PROCEDURE WAS COMPLETED ROBOTICALLY AND THERE WERE NO POST OPERATIVE COMPLICATIONS. HISTORICALLY, THE SURGEON USED A SUREFORM 60 STAPLER INSTRUMENT WITH A GREEN RELOAD THAT CONTAINED BUTTRESS MATERIAL. ANOTHER BARIATRIC SURGEON ADVISED TO NOT USE BUTTRESS MATERIAL; THE SURGEON HAS SINCE CHANGED TO USING A BLUE RELOAD. THE SURGEON CHOSE TO SWITCH TO A BLUE RELOAD BECAUSE IT WAS THE MOST COMPARABLE OPTION WHEN USING A 3RD PARTY LAPAROSCOPIC STAPLER AND A PURPLE STAPLE LOAD. THE CLINICAL SALES REPRESENTATIVE HAS DISCUSSED WITH THE SURGEON, TRANSITIONING FROM A BLUE RELOAD TO A WHITE RELOAD AS THE TISSUE THAT THE RELOAD IS BEING USED ON IS THIN.
22-11-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED, AND THE DEVICE DID NOT FIRE. IT WAS CHECKED LATER AND THE RELOAD WAS FIRED, AND THE STAPLE CAME OUTSIDE OF THE RELOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
22-11-2023 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SURGERY, AFTER PRESSING THE GREEN BUTTON, THE HANDLE COULD NOT BE SQUEEZED, AND THE DEVICE DID NOT FIRE. IT WAS CHECKED LATER AND THE RELOAD WAS FIRED, AND THE STAPLE CAME OUTSIDE OF THE RELOAD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2024 | Malfunction | us surgical puerto rico | 22-01-2024 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PROCEDURE, THE JAWS DID NOT CLOSE UNTIL THE END. TO RESOLVE THE ISSUE, ANOTHER RELOAD WAS USED WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
28-12-2023 | Injury | medtronic / covidien | 22-01-2024 | COVIDIEN TRI-STAPLE 2.0 REINFORCED RELOAD 60 MM MED/THICK |  |  | Device Embedded In Tissue or Plaque | Event Description: PATIENT UNDERGOING LAPAROSCOPIC ILEAL PANCREATIC DIVERSION WITH DUODENAL SWITCH FOR TREATMENT OF MORBID OBESITY. DURING PROCEDURE, THE MEDTRONIC COVIDIEN TRI-STAPLE 2.0 RELOAD WOULD NOT UNLOCK. MANUAL EXTRACTION TOOL REQUIRED TO UNLOCK RELOAD FROM TISSUE. THIS WAS THE SECOND SIMILAR EVENT INVOLVING THIS PRODUCT AT OUR FACILITY WITHIN 7 DAYS. SEE PREVIOUSLY SUBMITTED MEDWATCH FORM (SUBMITTTED ON 01/11/2024) REGARDING PATIENT (B)(6).
14-09-2023 | Malfunction | covidien lp llc north haven | 23-01-2024 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, DURING PREPARATION PRIOR TO PATIENT CONTACT, THE HANDLE DID NOT ACTIVATE. A MANUAL HANDLE WAS TAKEN INSTEAD. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE MCP (MASTER CONTROL PANEL) LIVE SCRIPT SHOWED EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
14-09-2023 | Malfunction | covidien lp llc north haven | 23-01-2024 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, DURING PREPARATION PRIOR TO PATIENT CONTACT, THE HANDLE DID NOT ACTIVATE. A MANUAL HANDLE WAS TAKEN INSTEAD. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC¿S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE MCP (MASTER CONTROL PANEL) LIVE SCRIPT SHOWED EVIDENCE OF FIELD PROGRAMMABLE GATE ARRAY (FPGA) RESET/REPROGRAM FAILURES.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Injury | covidien lp llc north haven | 02-01-2025 | UNKNOWN EGIA TRI STAPLE | Failure to Form Staple | A050704 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IN DETACHING THE TRACHEA USING A PURPLE 60 MILLIMETERS (MM) RELOAD, THE STAPLES HAD POOR STAPLE FORMATION (BURST OUT). THE HANDLE ALSO MADE A CRACK SOUND OF BREAKING. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED WITH A BLACK 60MM RELOAD, BUT THE SAME ISSUE OCCURRED. TO COMPLETE THE CASE, THE SURGERY WAS CONVERTED TO AN OPEN SURGERY, WHICH RESULTED IN MORE THAN ONE INCH OF INCISION EXTENSION.
09-12-2024 | Injury | covidien lp llc north haven | 02-01-2025 | UNKNOWN EGIA TRI STAPLE | Failure to Form Staple | A050704 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IN DETACHING THE TRACHEA USING A PURPLE 60 MILLIMETERS (MM) RELOAD, THE STAPLES HAD POOR STAPLE FORMATION (BURST OUT). THE HANDLE ALSO MADE A CRACK SOUND OF BREAKING. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED WITH A BLACK 60MM RELOAD, BUT THE SAME ISSUE OCCURRED. TO COMPLETE THE CASE, THE SURGERY WAS CONVERTED TO AN OPEN SURGERY, WHICH RESULTED IN MORE THAN ONE INCH OF INCISION EXTENSION.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PANCREATICODUODENECTOMY, THE RELOAD WAS UNABLE TO FIRE WHILE BEING APPLIED TO THE GASTRODUODENAL ARTERY. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER WITH A NEW DEVICE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PANCREATICODUODENECTOMY, THE RELOAD WAS UNABLE TO FIRE WHILE BEING APPLIED TO THE GASTRODUODENAL ARTERY. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER WITH A NEW DEVICE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PANCREATICODUODENECTOMY, THE RELOAD WAS UNABLE TO FIRE WHILE BEING APPLIED TO THE GASTRODUODENAL ARTERY. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER WITH A NEW DEVICE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PANCREATICODUODENECTOMY, THE RELOAD WAS UNABLE TO FIRE WHILE BEING APPLIED TO THE GASTRODUODENAL ARTERY. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER WITH A NEW DEVICE. THERE WAS NO PATIENT INJURY.
19-12-2024 | Injury | ethicon endo-surgery | 02-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN END-TO-END ANASTOMOSIS PROGRESSION IN SURGERY USING ECP29 THERE DIDN'T SEEM TO BE ANY PROBLEM WITH THE STAPLE LINE IN INTRA-OP, BUT FOUND POST-OP ABOUT 3 DAYS AFTER SURGERY. FOUND A LEAK IN THE POSTERIOR WALL OF THE ANASTOMOSIS SITE.
09-12-2024 | Malfunction | covidien | 02-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PANCREATICODUODENECTOMY, THE RELOAD WAS UNABLE TO FIRE WHILE BEING APPLIED TO THE GASTRODUODENAL ARTERY. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER WITH A NEW DEVICE. THERE WAS NO PATIENT INJURY.
01-01-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | ECHELON LINEAR CUTTER 80MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN CASE COLORECTAL ANASTOMOSIS PROCEDURE THE STAPLER WAS USED IN AND THE EXTRA STAPLES WERE NOT FORMED. THERE WAS NO HARM TO THE PATIENT. THE PROCEDURE WAS SUCCESSFUL.
09-07-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE RELOAD BURST WHEN THE LAPAROSCOPIC LINEAR CUTTING STAPLER WAS USED TO CUT THE LUNG TISSUE.
01-01-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | PROXIMATE*75MM LINEAR CUTTER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURES THE STAPLES FAILED TO FIRE EFFECTIVELY (DID NOT CREATE B-FORM) TO CREATE SEAL AT DISTAL END OF DEVICE. DEVICE WAS REPLACE AND THE FIRING WAS REPEATED PARALLEL TO ORIGINAL FIRING WITH SUCCESS. REPLACEMENT DEVICE WAS FROM SAME LOT #. THERE WAS NO SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE USE OF THE EGIA60AMT STAPLER RESULTED IN STAPLE LINE BLEEDING WHILE FORMING THE GASTRIC POUCH AND IN THE ANASTOMOSES. THIS ISSUE EXTENDED THE SURGICAL TIME BY 30 MINUTES. TO ADDRESS THE BLEEDING, SUTURES AND CLIPS WERE APPLIED ALONG THE SUTURE LINE AS REINFORCEMENT.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE, THE USE OF THE EGIA60AMT STAPLER RESULTED IN STAPLE LINE BLEEDING WHILE FORMING THE GASTRIC POUCH AND IN THE ANASTOMOSES. THIS ISSUE EXTENDED THE SURGICAL TIME BY 30 MINUTES. TO ADDRESS THE BLEEDING, SUTURES AND CLIPS WERE APPLIED ALONG THE SUTURE LINE AS REINFORCEMENT.
09-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, NOTED THAT CARTRIDGE BASE LOOSE AND COULD NOT INSTALL. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, NOTED THAT CARTRIDGE BASE LOOSE AND COULD NOT INSTALL. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2024 | Injury | ethicon endo-surgery | 02-01-2025 | 75MM SELECTABLE NEW TLC | Patient Device Interaction Problem | A01 | Fistula | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE THE LINEAR STAPLER STAPLING FAILED. THE PATIENT PRESENTED A FISTULA IN THE FIRST POSTOPERATIVE PERIOD REQUIRING RE-APPROACH. DURING THE SURGERY, A 0.5 CM FISTULOUS ORIFICE WAS OBSERVED IN THE STAPLE LINE.
29-10-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, WHEN USING A MANUAL HANDLE AND THREE RELOADS, THERE WAS DIFFICULTY IN FIRING. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA HANDLE THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
29-10-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY PROCEDURE, WHEN USING A MANUAL HANDLE AND THREE RELOADS, THERE WAS DIFFICULTY IN FIRING. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA HANDLE THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
18-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC ASSISTED ANTERIOR RESECTION AN ANVIL SECURED PROXIMALLY USING PURSE STRING TECHNIQUE (PROLENE). DEVICE INTRODUCED INTO RECTUM; TROCAR FULLY OPENED USING ADJUSTING KNOB. WHEN CONNECTING THE ANVIL TO TROCAR, THE TROCAR RETRACTED INTO DEVICE. SURGEON INSTRUCTED FIRST ASSIST TO RE-OPEN THE DEVICE TO PUSH OUT THE TROCAR. ANVIL AND TROCAR CONNECTED SUCCESSFULLY (ORANGE BAND NO LONGER VISIBLE). FIRING COMPLETED, GREEN TICK VISIBLE. ONCE DEVICE REMOVED, DONUT INSPECTION UNCOVERED INCOMPLETE PROXIMAL DONUT. PROXIMAL DONUT WAS A ¿U¿ SHAPE/HORSESHOE SHAPE. NO STAPLES SEEN AT THE TOP OF THE ¿U¿ ON EITHER PROXIMAL OR DISTAL COLON. UPON INSPECTION, A 1-1 ½ CM HOLE OBSERVED IN THE ANASTOMOSIS. SURGEON USED 5 INTERRUPTED SUTURE A ROBOTICALLY SECURED HOLE. ADDED ADDITION 1+ HOURS ONTO CASE. SURGEON OBSERVED THE PROLENE PURSE-STRING HAD NOT CUT AS IT NORMALLY WOULD WITH THE CIRCULAR KNIFE. PATIENTS CARE CONTINUED AS NORMAL, POTENTIAL DRAIN MUST REMAIN FOR LONGER. PATIENT WAS GIVEN A DE-FUNCTIONING ILEOSTOMY (PRE-PLANNED).
18-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC ASSISTED ANTERIOR RESECTION AN ANVIL SECURED PROXIMALLY USING PURSE STRING TECHNIQUE (PROLENE). DEVICE INTRODUCED INTO RECTUM; TROCAR FULLY OPENED USING ADJUSTING KNOB. WHEN CONNECTING THE ANVIL TO TROCAR, THE TROCAR RETRACTED INTO DEVICE. SURGEON INSTRUCTED FIRST ASSIST TO RE-OPEN THE DEVICE TO PUSH OUT THE TROCAR. ANVIL AND TROCAR CONNECTED SUCCESSFULLY (ORANGE BAND NO LONGER VISIBLE). FIRING COMPLETED, GREEN TICK VISIBLE. ONCE DEVICE REMOVED, DONUT INSPECTION UNCOVERED INCOMPLETE PROXIMAL DONUT. PROXIMAL DONUT WAS A ¿U¿ SHAPE/HORSESHOE SHAPE. NO STAPLES SEEN AT THE TOP OF THE ¿U¿ ON EITHER PROXIMAL OR DISTAL COLON. UPON INSPECTION, A 1-1 ½ CM HOLE OBSERVED IN THE ANASTOMOSIS. SURGEON USED 5 INTERRUPTED SUTURE A ROBOTICALLY SECURED HOLE. ADDED ADDITION 1+ HOURS ONTO CASE. SURGEON OBSERVED THE PROLENE PURSE-STRING HAD NOT CUT AS IT NORMALLY WOULD WITH THE CIRCULAR KNIFE. PATIENTS CARE CONTINUED AS NORMAL, POTENTIAL DRAIN MUST REMAIN FOR LONGER. PATIENT WAS GIVEN A DE-FUNCTIONING ILEOSTOMY (PRE-PLANNED).
18-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC ASSISTED ANTERIOR RESECTION AN ANVIL SECURED PROXIMALLY USING PURSE STRING TECHNIQUE (PROLENE). DEVICE INTRODUCED INTO RECTUM; TROCAR FULLY OPENED USING ADJUSTING KNOB. WHEN CONNECTING THE ANVIL TO TROCAR, THE TROCAR RETRACTED INTO DEVICE. SURGEON INSTRUCTED FIRST ASSIST TO RE-OPEN THE DEVICE TO PUSH OUT THE TROCAR. ANVIL AND TROCAR CONNECTED SUCCESSFULLY (ORANGE BAND NO LONGER VISIBLE). FIRING COMPLETED, GREEN TICK VISIBLE. ONCE DEVICE REMOVED, DONUT INSPECTION UNCOVERED INCOMPLETE PROXIMAL DONUT. PROXIMAL DONUT WAS A ¿U¿ SHAPE/HORSESHOE SHAPE. NO STAPLES SEEN AT THE TOP OF THE ¿U¿ ON EITHER PROXIMAL OR DISTAL COLON. UPON INSPECTION, A 1-1 ½ CM HOLE OBSERVED IN THE ANASTOMOSIS. SURGEON USED 5 INTERRUPTED SUTURE A ROBOTICALLY SECURED HOLE. ADDED ADDITION 1+ HOURS ONTO CASE. SURGEON OBSERVED THE PROLENE PURSE-STRING HAD NOT CUT AS IT NORMALLY WOULD WITH THE CIRCULAR KNIFE. PATIENTS CARE CONTINUED AS NORMAL, POTENTIAL DRAIN MUST REMAIN FOR LONGER. PATIENT WAS GIVEN A DE-FUNCTIONING ILEOSTOMY (PRE-PLANNED).
09-12-2024 | Malfunction | covidien | 02-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS UNABLE TO FIRE. IT WAS ALSO NOTED THAT THE SAFETY LOCK BUCKLE OF THE CIRCULAR STAPLER WAS BROKEN. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS UNABLE TO FIRE. THE HANDLE GOT STUCK AFTER BEING SQUEEZED, AND THE RELOADS COULD NOT BE FIRED. IT WAS ALSO NOTED THAT THE SAFETY LOCK BUCKLE OF THE CIRCULAR STAPLER WAS BROKEN. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE LAPAROSCOPIC SUBTOTAL GASTRECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS UNABLE TO FIRE. IT WAS ALSO NOTED THAT THE SAFETY LOCK BUCKLE OF THE CIRCULAR STAPLER WAS BROKEN. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | PROXIMATE |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: MISFIRED ON RELOAD IN PATIENT BODY.
01-01-2024 | Injury | ethicon endo-surgery | 02-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING AN OPEN SIGMOID RESECTION WITH A REPAIR FOR A COLO VESICLE FISTULA. THE STAPLER WAS USED FOR AN END-TO-END ANASTOMOSIS OF THE RECTUM. WHEN ATTEMPTS WERE MADE TO PULL THE STAPLER OUT OF THE RECTUM, THEY MET RESISTANCE. TRIED DIFFERENT MANEUVERS TO REMOVE STAPLER. AS THEY WERE DISSECTING AND FREEING UP THE STAPLER A LARGE SEPARATION OF THE STAPLE LINE OF THE ANASTOMOSIS WAS NOTED. AS A RESULT, THE TEAM PLACED AN UNPLANNED STOMA AND COLOSTOMY. CASE WAS COMPLETED. NO ADDITIONAL PATIENT CONSEQUENCE.
01-01-2024 | Injury | ethicon endo-surgery | 02-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Eject | A050301 | Appropriate Term / Code Not Available | Event Description: IT WAS REPORTED THAT DURING AN OPEN SIGMOID RESECTION WITH A REPAIR FOR A COLO VESICLE FISTULA. THE STAPLER WAS USED FOR AN END-TO-END ANASTOMOSIS OF THE RECTUM. WHEN ATTEMPTS WERE MADE TO PULL THE STAPLER OUT OF THE RECTUM, THEY MET RESISTANCE. TRIED DIFFERENT MANEUVERS TO REMOVE STAPLER. AS THEY WERE DISSECTING AND FREEING UP THE STAPLER A LARGE SEPARATION OF THE STAPLE LINE OF THE ANASTOMOSIS WAS NOTED. AS A RESULT, THE TEAM PLACED AN UNPLANNED STOMA AND COLOSTOMY. CASE WAS COMPLETED. NO ADDITIONAL PATIENT CONSEQUENCE.
22-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE FIRES AUTOMATICALLY AND NOTED THE STAPLES MALFORMED. CHANGE ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
22-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE FIRES AUTOMATICALLY AND NOTED THE STAPLES MALFORMED. CHANGE ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED.
23-01-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | PROXIMATE*RELOAD THICK TISSUE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE DID NOT LOOSE ANY STAPLES DURING USE. THERE WAS NO SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHEN PERFORMING AN OPEN/CLOSE CHECK, THE GAP BETWEEN THE ANVIL AND CARTRIDGE WAS EXCESSIVELY LARGE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHEN PERFORMING AN OPEN/CLOSE CHECK, THE GAP BETWEEN THE ANVIL AND CARTRIDGE WAS EXCESSIVELY LARGE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE FIRING STOPPED IN THE MIDDLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE FIRING STOPPED IN THE MIDDLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN PERFORMING A OPEN/CLOSE CHECK, THE GAP BETWEEN THE ANVIL AND CARTRIDGE WAS EXCESSIVELY LARGE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE.
09-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN PERFORMING A OPEN/CLOSE CHECK, THE GAP BETWEEN THE ANVIL AND CARTRIDGE WAS EXCESSIVELY LARGE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE FIRING STOPPED IN THE MIDDLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE FIRING STOPPED IN THE MIDDLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
10-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE DEVICE, THERE WAS A PRODUCT ISSUE WITH STAPLE FORMATION. THE STAPLE LINE WAS POORLY FORMED, AND THE THREE LINES ON EACH SIDE WERE NOT VISIBLE. VISUALLY, ONE STAPLE WAS DISENGAGED INTO THE PATIENT'S CAVITY. AS A RESULT, SUTURING WAS PERFORMED AS A STAPLE LINE REINFORCEMENT.
10-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Detachment of Device or Device Component | A0501 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE DEVICE, THERE WAS A PRODUCT ISSUE WITH STAPLE FORMATION. THE STAPLE LINE WAS POORLY FORMED, AND THE THREE LINES ON EACH SIDE WERE NOT VISIBLE. VISUALLY, ONE STAPLE WAS DISENGAGED INTO THE PATIENT'S CAVITY. AS A RESULT, SUTURING WAS PERFORMED AS A STAPLE LINE REINFORCEMENT.
03-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | CIRCULAR MECH STD 25MM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE FIRING FORCE WAS QUITE HIGHER THAN USUAL AT THE 1ST FIRING DURING ANASTOMOSIS. THERE WAS NO EFFECT ON THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
03-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-10-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | 75MM SELECTABLE NEW TLC | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A JOINT GYN AND COLORECTAL CASE THE DEVICE WAS NOT CLICKING AND ONE OF THE STAPLES PIERCED THE SCRUB NURSES THUMB BUT LUCKILY, IT WAS BEFORE IT WAS USED.
18-10-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | 75MM SELECTABLE NEW TLC | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A JOINT GYN AND COLORECTAL CASE THE DEVICE WAS NOT CLICKING AND ONE OF THE STAPLES PIERCED THE SCRUB NURSES THUMB BUT LUCKILY, IT WAS BEFORE IT WAS USED.
05-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 02-01-2025 | CAPSURE | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A TRANSABDOMINAL PREPERITONEAL (TAPP) PROCEDURE, THE CAPSURE DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT ON THE FIRST TRIGGER THE DEVICE DID NOT RELEASE A FASTENER AND ON THE SECOND TRIGGER TWO FASTENERS CAME OUT INTERTWINED WHICH WERE REMOVED FROM THE PATIENT'S BODY. THE USER REMOVED THE DEVICE FROM THE PATIENT TO CLEAR IT AND FOUND THE DEVICE TO THEN WORK PROPERLY DEPLOYING 3 FASTENERS SUCCESSFULLY. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE SMALL INTESTINAL RESECTION, THERE WAS PARTIAL FIRING. THE RELOAD WAS CHANGED WITHOUT CHANGING THE HANDLE, SHELL, OR ADAPTER. THERE WAS NO PATIENT INJURY.
16-12-2024 | Injury | ethicon endo-surgery | 02-01-2025 | ILS 33MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE THAT THE ANVIL WAS PLACED IN THE PROXIMAL COLON AND PUSHED THROUGH POSTERIOR TO THE STAPLE LINE BUT AFTER SEVERAL ATTEMPTS THE SPIKE WOULD NOT DISENGAGE FROM THE ANVIL, MULTIPLE ATTEMPTS WITH VARYING TECHNIQUES AND NO SUCCESS ULTIMATELY THE COLON WAS OPENED TO REMOVE THE ANVIL. A NEW STAPLER WAS OBTAINED. INTACT DISTAL DONUT, PROXIMAL DONUT NOT INTACT, LEAK TEST POSITIVE DIVERSION WAS THEN PLANNED.
16-12-2024 | Injury | ethicon endo-surgery | 02-01-2025 | ILS 33MM, CURVED | Ejection Problem | A0503 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY PROCEDURE THAT THE ANVIL WAS PLACED IN THE PROXIMAL COLON AND PUSHED THROUGH POSTERIOR TO THE STAPLE LINE BUT AFTER SEVERAL ATTEMPTS THE SPIKE WOULD NOT DISENGAGE FROM THE ANVIL, MULTIPLE ATTEMPTS WITH VARYING TECHNIQUES AND NO SUCCESS ULTIMATELY THE COLON WAS OPENED TO REMOVE THE ANVIL. A NEW STAPLER WAS OBTAINED. INTACT DISTAL DONUT, PROXIMAL DONUT NOT INTACT, LEAK TEST POSITIVE DIVERSION WAS THEN PLANNED.
16-12-2024 | Injury | ethicon endo-surgery | 02-01-2025 | ILS 33MM, CURVED | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY THE STAPLER CLOSED, FIRED, AND THEN THE ANASTOMOTIC LEAK WAS DISCOVERED SO THEY DIVERTED THE COLON. THIS ANVIL WAS PLACED WITHOUT THE SPIKE THE EEA WAS PLACED DISTAL SPIKE DEPLOYED, CONNECTED, STAPLER CLOSED, FIRED ETC. INTACT DISTAL DONUT, PROXIMAL DONUT NOT INTACT, LEAK TEST POSITIVE DIVERSION WAS THEN PLANNED.
16-12-2024 | Injury | ethicon endo-surgery | 02-01-2025 | ILS 33MM, CURVED | Ejection Problem | A0503 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY THE STAPLER CLOSED, FIRED, AND THEN THE ANASTOMOTIC LEAK WAS DISCOVERED SO THEY DIVERTED THE COLON. THIS ANVIL WAS PLACED WITHOUT THE SPIKE THE EEA WAS PLACED DISTAL SPIKE DEPLOYED, CONNECTED, STAPLER CLOSED, FIRED ETC. INTACT DISTAL DONUT, PROXIMAL DONUT NOT INTACT, LEAK TEST POSITIVE DIVERSION WAS THEN PLANNED.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE FIRING STOPPED IN THE MIDDLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE FIRING STOPPED IN THE MIDDLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
04-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 02-01-2025 | SORBAFIX ENHANCED | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A PROCEDURE THE SURGEON TRIED TO FIXATE THE MESH USING SORBAFIX ENHANCED FIXATION DEVICE. IT WAS REPORTED THAT THE DEVICE FAILED TO ACTIVATE THE STAPLE AND DEPLOYED BROKEN FASTENERS. IT WAS FURTHER REPORTED THAT THE PROCEDURE WAS COMPLETED WITH SUTURES. THERE WAS NO PATIENT INJURY.
26-11-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP HEMICOLECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-11-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, COULD NOT FIRE. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE FIRING KNOB WAS BROKEN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
25-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | 75MM SELECTABLE NEW TLC | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE FIRING KNOB WAS BROKEN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
25-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DISTAL GASTRECTOMY PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE ANVIL ON THE STAPLER LOOKED TO BE DAMAGED, IT WOULDN'T LOCK ON THE SITE. THE NEW DEVICE WAS OPENED. NO PATIENT CONSEQUENCES TO COMPLETE THE PROCEDURE.
27-12-2024 | Malfunction | ethicon endo-surgery | 02-01-2025 | ECHELON CONTOUR GST RELOAD GREEN | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE DEVICE WAS CLOSED ON TISSUE, FIRE THE STAPLER PROPERLY AND ALSO GOT AUDIBLE FEEDBACK. AFTER HE OPENED THE STAPLER HE SAW THAT COLON IS TRANSECTED BUT THERE WERE NO STAPLERS, SO STAPLER DID ONLY TRANSECTION WITHOUT CLAMPS IN IT. WE PUT A RELOAD ASIDE AND IT WILL BE SEND FOR INSPECTION. THE SURGERY WAS DELAYED 20 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
04-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE DEVICE PRIOR TO FIRING, SPECIFICALLY RELATED TO THE RELOAD COMPONENT. THE PROBLEMS INCLUDED JAW ISSUES, DIFFICULTY CLOSING THE JAWS, AND THE JAWS NOT CLOSING COMPLETELY. THE DEVICE WAS USED WITH A POWERED STAPLER AND A STANDARD ADAPTER. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER DEVICE FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
04-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE DEVICE PRIOR TO FIRING, SPECIFICALLY RELATED TO THE RELOAD COMPONENT. THE PROBLEMS INCLUDED JAW ISSUES, DIFFICULTY CLOSING THE JAWS, AND THE JAWS NOT CLOSING COMPLETELY. THE DEVICE WAS USED WITH A POWERED STAPLER AND A STANDARD ADAPTER. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER DEVICE FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
04-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE DEVICE PRIOR TO FIRING, SPECIFICALLY RELATED TO THE RELOAD COMPONENT. THE PROBLEMS INCLUDED JAW ISSUES, DIFFICULTY CLOSING THE JAWS, AND THE JAWS NOT CLOSING COMPLETELY. THE DEVICE WAS USED WITH A POWERED STAPLER AND A STANDARD ADAPTER. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER DEVICE FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
04-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE DEVICE PRIOR TO FIRING, SPECIFICALLY RELATED TO THE RELOAD COMPONENT. THE PROBLEMS INCLUDED JAW ISSUES, DIFFICULTY CLOSING THE JAWS, AND THE JAWS NOT CLOSING COMPLETELY. THE DEVICE WAS USED WITH A POWERED STAPLER AND A STANDARD ADAPTER. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER DEVICE FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
04-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE DEVICE PRIOR TO FIRING, SPECIFICALLY RELATED TO THE RELOAD COMPONENT. THE PROBLEMS INCLUDED JAW ISSUES, DIFFICULTY CLOSING THE JAWS, AND THE JAWS NOT CLOSING COMPLETELY. THE DEVICE WAS USED WITH A POWERED STAPLER AND A STANDARD ADAPTER. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER DEVICE FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
04-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE DEVICE PRIOR TO FIRING, SPECIFICALLY RELATED TO THE RELOAD COMPONENT. THE PROBLEMS INCLUDED JAW ISSUES, DIFFICULTY CLOSING THE JAWS, AND THE JAWS NOT CLOSING COMPLETELY. THE DEVICE WAS USED WITH A POWERED STAPLER AND A STANDARD ADAPTER. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER DEVICE FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
10-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY, WHEN CLAMPING THE LUNG LOBE, IT POPPED WHEN BEING FIRED WHEN THE HANDLE WAS USED. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE STAPLER DID NOT FIRE, AND IT OCCURRED ON TWO RELOADS. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY, WHEN CLAMPING THE LUNG LOBE, IT POPPED WHEN BEING FIRED WHEN THE HANDLE WAS USED. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE STAPLER DID NOT FIRE, AND IT OCCURRED ON TWO RELOADS. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY, WHEN CLAMPING THE LUNG LOBE, IT POPPED WHEN BEING FIRED WHEN THE HANDLE WAS USED. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE STAPLER DID NOT FIRE, AND IT OCCURRED ON TWO RELOADS. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PULMONARY LOBECTOMY, WHEN CLAMPING THE LUNG LOBE, IT POPPED WHEN BEING FIRED WHEN THE HANDLE WAS USED. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE STAPLER DID NOT FIRE, AND IT OCCURRED ON TWO RELOADS. A NEW DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2024 | Malfunction | covidien | 02-01-2025 | ENDO GIA UNIVERSAL STAPLER (XL) WITH A 60 PURPLE ENDO GIA RELOAD |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: THE STAPLE LOAD WAS ON THE STAPLER BUT WOULD NOT CLOSE COMPLETELY. BOTH THE STAPLER AND THE LOAD WERE SEQUESTERED IN THE COORDINATOR OFFICE.
26-12-2024 | Malfunction | covidien | 02-01-2025 | 60 PURPLE ENDO GIA RELOAD |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: (THIS REPORT IS FOR THE STAPLE LOAD) THE STAPLE LOAD WAS ON THE STAPLER BUT WOULD NOT CLOSE COMPLETELY. BOTH THE STAPLER AND THE LOAD WERE SEQUESTERED IN THE COORDINATOR OFFICE.
10-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW RECTAL RESECTION, ONE RELOAD WAS FIRED WHILE CUTTING THE RECTUM. AFTER THE SECOND RELOAD WAS CONNECTED TO THE SAME HANDLE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW HANDLE WAS USED WITH THE SAME RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE USING THE STAPLERS, POOR STAPLE FORMATION OCCURRED TWICE, AND AN INCOMPLETE STAPLE LINE WAS NOTED TWICE, ONCE CENTRALLY AND ONCE DISTALLY. THE PROCEDURE WAS LAPAROSCOPIC AND INVOLVED THE TRACHEA AND BLOOD VESSELS. THE ISSUE WAS RESOLVED BY REPLACING THE DEFECTIVE STAPLES WITH NEW ONES. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | us surgical puerto rico | 02-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE USING THE STAPLERS, POOR STAPLE FORMATION OCCURRED TWICE, AND AN INCOMPLETE STAPLE LINE WAS NOTED TWICE, ONCE CENTRALLY AND ONCE DISTALLY. THE PROCEDURE WAS LAPAROSCOPIC AND INVOLVED THE TRACHEA AND BLOOD VESSELS. THE ISSUE WAS RESOLVED BY REPLACING THE DEFECTIVE STAPLES WITH NEW ONES. THERE WAS NO PATIENT INJURY.
02-12-2024 | Malfunction | ethicon | 02-01-2025 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HERNIA REPAIR ON (B)(6)2024 AND ABSORBABLE STRAPS WERE USED. THE STAPLER STOPPED WORKING PROPERLY, THE STAPLES WERE ESCAPING. DEVICE WAS REPLACED AND SURGERY COMPLETED. THERE WAS NO DAMAGE TO SURGERY OR TO THE PATIENT. ADDITIONAL INFORMATION WAS REQUESTED.
10-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE USING THE EGIA60AXT AND EGIA45AMT STAPLERS, POOR STAPLE FORMATION OCCURRED TWICE, AND AN INCOMPLETE STAPLE LINE WAS NOTED TWICE, ONCE CENTRALLY AND ONCE DISTALLY. THE PROCEDURE WAS LAPAROSCOPIC AND INVOLVED THE TRACHEA AND BLOOD VESSELS. THE ISSUE WAS RESOLVED BY REPLACING THE DEFECTIVE STAPLES WITH NEW ONES. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE USING THE EGIA60AXT AND EGIA45AMT STAPLERS, POOR STAPLE FORMATION OCCURRED TWICE, AND AN INCOMPLETE STAPLE LINE WAS NOTED TWICE, ONCE CENTRALLY AND ONCE DISTALLY. THE PROCEDURE WAS LAPAROSCOPIC AND INVOLVED THE TRACHEA AND BLOOD VESSELS. THE ISSUE WAS RESOLVED BY REPLACING THE DEFECTIVE STAPLES WITH NEW ONES. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | ethicon | 02-01-2025 | ETHICON SECURESTRAP | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNSPECIFIED PROCEDURE ON (B)(6) 2024 AND ABSORBABLE STRAPS WERE USED. WHEN FIXING THE MESH WITH THE STRAPS THE DEVICE JAMS AND DID NOT MAKE ANY SHOT, THE STRAPS FELL AND DO NOT FIX. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION WAS PROVIDED.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE SMALL INTESTINAL RESECTION, THERE WAS PARTIAL FIRING. THE RELOAD WAS CHANGED WITHOUT CHANGING THE HANDLE, SHELL, OR ADAPTER. THERE WAS NO PATIENT INJURY.
02-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE DEVICE DURING SET-UP AND PRIOR TO FIRING, THE JAWS DID NOT CLOSE COMPLETELY. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
02-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE DEVICE DURING SET-UP AND PRIOR TO FIRING, THE JAWS DID NOT CLOSE COMPLETELY. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE SMALL INTESTINAL RESECTION, THERE WAS PARTIAL FIRING. THE RELOAD WAS CHANGED WITHOUT CHANGING THE HANDLE, SHELL, OR ADAPTER. THERE WAS NO PATIENT INJURY.
19-08-2024 | Injury | covidien lp llc north haven | 02-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A PROSPECTIVE AND RETROSPECTIVE STUDY PERFORMED BETWEEN 2013 AND 2023, INVOLVING 116 PATIENTS WHO UNDERWENT MINIMALLY INVASIVE DISTAL PANCREATECTOMY (MIDP) WITH OR WITHOUT SPLENECTOMY BETWEEN 2012 AND 2023. PANCREATIC TRANSECTION WAS PERFORMED USING A 60MM STAPLER. SPLENIC VESSELS WERE SECTIONED USING EITHER AN UNKNOWN VASCULAR STAPLER OR COMPETITOR CLIPS. DEVICE RELATED COMPLICATIONS INCLUDED POSTOPERATIVE PANCREATIC FISTULA (POPF), POSTOPERATIVE BLEEDING, AND ABDOMINAL FLUID COLLECTION. THE POPFS REQUIRED REOPERATION, ENDOSCOPY, OR PERCUTANEOUS DRAINAGE. THE POSTOPERATIVE BLEEDING REQUIRED REOPERATION. READMISSIONS WERE ALSO REPORTED.
19-08-2024 | Injury | covidien lp llc north haven | 02-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A PROSPECTIVE AND RETROSPECTIVE STUDY PERFORMED BETWEEN 2013 AND 2023, INVOLVING 116 PATIENTS WHO UNDERWENT MINIMALLY INVASIVE DISTAL PANCREATECTOMY (MIDP) WITH OR WITHOUT SPLENECTOMY BETWEEN 2012 AND 2023. PANCREATIC TRANSECTION WAS PERFORMED USING A 60MM STAPLER. SPLENIC VESSELS WERE SECTIONED USING EITHER AN UNKNOWN VASCULAR STAPLER OR COMPETITOR CLIPS. DEVICE RELATED COMPLICATIONS INCLUDED POSTOPERATIVE PANCREATIC FISTULA (POPF), POSTOPERATIVE BLEEDING, AND ABDOMINAL FLUID COLLECTION. THE POPFS REQUIRED REOPERATION, ENDOSCOPY, OR PERCUTANEOUS DRAINAGE. THE POSTOPERATIVE BLEEDING REQUIRED REOPERATION. READMISSIONS WERE ALSO REPORTED.
19-08-2024 | Injury | covidien lp llc north haven | 02-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A PROSPECTIVE AND RETROSPECTIVE STUDY PERFORMED BETWEEN 2013 AND 2023, INVOLVING 116 PATIENTS WHO UNDERWENT MINIMALLY INVASIVE DISTAL PANCREATECTOMY (MIDP) WITH OR WITHOUT SPLENECTOMY BETWEEN 2012 AND 2023. PANCREATIC TRANSECTION WAS PERFORMED USING A 60MM STAPLER. SPLENIC VESSELS WERE SECTIONED USING EITHER AN UNKNOWN VASCULAR STAPLER OR COMPETITOR CLIPS. DEVICE RELATED COMPLICATIONS INCLUDED POSTOPERATIVE PANCREATIC FISTULA (POPF), POSTOPERATIVE BLEEDING, AND ABDOMINAL FLUID COLLECTION. THE POPFS REQUIRED REOPERATION, ENDOSCOPY, OR PERCUTANEOUS DRAINAGE. THE POSTOPERATIVE BLEEDING REQUIRED REOPERATION. READMISSIONS WERE ALSO REPORTED.
02-12-2024 | Malfunction | covidien lp llc north haven | 02-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE DEVICE DURING SET-UP AND PRIOR TO FIRING, THE JAWS DID NOT CLOSE COMPLETELY. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
18-02-2025 | Injury | covidien lp llc north haven | 02-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PROCEDURE, THE STAPLER HAD MISFIRED WHICH LED TO EXCESSIVE BLEEDING ON THE STAPLE LINE.
18-12-2024 | Injury | covidien | 02-01-2025 | END TO END STAPLING DEVICE |  |  | Bowel Perforation | Event Description: ANASTOMOSIS STAPLER DID NOT FIRE PROPERLY, RESULTING IN CONTINUATION AND PROLONGED SURGERY, AS WELL AS REMOVING MORE BOWEL THAN WHAT WAS PLANNED AND MANUAL SUTURE REPAIR.
02-12-2024 | Death | intuitive surgical | 03-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED BARIATRIC GASTRIC REVISION AND LAP BAND REMOVAL PROCEDURE, A STAPLE LINE LEAK OCCURRED REQUIRING RE-OPERATION. THE PATIENT EXPERIENCED AN UNSPECIFIED COMPLICATION AND ULTIMATELY EXPIRED ON POST-OPERATIVE DAY 24. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS INFORMED BY THE SURGEON'S PHYSICIAN ASSISTANT THAT DURING THE ROBOTIC PROCEDURE, TWO SUREFORM 60 BLUE STAPLER RELOADS WERE FIRED. ON THE SECOND STAPLER FIRING, A TISSUE PUSH EVENT OCCURRED TOWARDS THE END OF THE STAPLE LINE. THE SURGEON OVERSEWED THE STAPLE LINE AT THE TIME. THERE WERE NO OTHER ISSUES REPORTED DURING THE DA VINCI PROCEDURE. ON AN UNSPECIFIED LATER DATE, A POST-OPERATIVE LEAK WAS DISCOVERED FROM THIS STAPLE LINE. THE PATIENT UNDERWENT A RE-OPERATION TO REPAIR THE LEAK. ON A LATER DATE, THE PATIENT DEVELOPED A COMPLICATION SECONDARY TO THE STAPLE LINE LEAK AND WAS REINTUBATED. THE PATIENT'S FAMILY ULTIMATELY ELECTED TO TAKE THE PATIENT OFF OF LIFE SUPPORT AND THE PATIENT PASSED. NO FURTHER SPECIFIC DETAILS WERE PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
02-12-2024 | Death | intuitive surgical | 03-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED BARIATRIC GASTRIC REVISION AND LAP BAND REMOVAL PROCEDURE, A STAPLE LINE LEAK OCCURRED REQUIRING RE-OPERATION. THE PATIENT EXPERIENCED AN UNSPECIFIED COMPLICATION AND ULTIMATELY EXPIRED ON POST-OPERATIVE DAY 24. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS INFORMED BY THE SURGEON'S PHYSICIAN ASSISTANT THAT DURING THE ROBOTIC PROCEDURE, TWO SUREFORM 60 BLUE STAPLER RELOADS WERE FIRED. ON THE SECOND STAPLER FIRING, A TISSUE PUSH EVENT OCCURRED TOWARDS THE END OF THE STAPLE LINE. THE SURGEON OVERSEWED THE STAPLE LINE AT THE TIME. THERE WERE NO OTHER ISSUES REPORTED DURING THE DA VINCI PROCEDURE. ON AN UNSPECIFIED LATER DATE, A POST-OPERATIVE LEAK WAS DISCOVERED FROM THIS STAPLE LINE. THE PATIENT UNDERWENT A RE-OPERATION TO REPAIR THE LEAK. ON A LATER DATE, THE PATIENT DEVELOPED A COMPLICATION SECONDARY TO THE STAPLE LINE LEAK AND WAS REINTUBATED. THE PATIENT'S FAMILY ULTIMATELY ELECTED TO TAKE THE PATIENT OFF OF LIFE SUPPORT AND THE PATIENT PASSED. NO FURTHER SPECIFIC DETAILS WERE PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
02-12-2024 | Death | intuitive surgical | 03-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED BARIATRIC GASTRIC REVISION AND LAP BAND REMOVAL PROCEDURE, A STAPLE LINE LEAK OCCURRED REQUIRING RE-OPERATION. THE PATIENT EXPERIENCED AN UNSPECIFIED COMPLICATION AND ULTIMATELY EXPIRED ON POST-OPERATIVE DAY 24. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS INFORMED BY THE SURGEON'S PHYSICIAN ASSISTANT THAT DURING THE ROBOTIC PROCEDURE, TWO SUREFORM 60 BLUE STAPLER RELOADS WERE FIRED. ON THE SECOND STAPLER FIRING, A TISSUE PUSH EVENT OCCURRED TOWARDS THE END OF THE STAPLE LINE. THE SURGEON OVERSEWED THE STAPLE LINE AT THE TIME. THERE WERE NO OTHER ISSUES REPORTED DURING THE DA VINCI PROCEDURE. ON AN UNSPECIFIED LATER DATE, A POST-OPERATIVE LEAK WAS DISCOVERED FROM THIS STAPLE LINE. THE PATIENT UNDERWENT A RE-OPERATION TO REPAIR THE LEAK. ON A LATER DATE, THE PATIENT DEVELOPED A COMPLICATION SECONDARY TO THE STAPLE LINE LEAK AND WAS REINTUBATED. THE PATIENT'S FAMILY ULTIMATELY ELECTED TO TAKE THE PATIENT OFF OF LIFE SUPPORT AND THE PATIENT PASSED. NO FURTHER SPECIFIC DETAILS WERE PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
02-12-2024 | Death | intuitive surgical | 03-01-2025 | SUREFORM | Failure to Cut | A050702 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED BARIATRIC GASTRIC REVISION AND LAP BAND REMOVAL PROCEDURE, A STAPLE LINE LEAK OCCURRED REQUIRING RE-OPERATION. THE PATIENT EXPERIENCED AN UNSPECIFIED COMPLICATION AND ULTIMATELY EXPIRED ON POST-OPERATIVE DAY 24. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS INFORMED BY THE SURGEON'S PHYSICIAN ASSISTANT THAT DURING THE ROBOTIC PROCEDURE, TWO SUREFORM 60 BLUE STAPLER RELOADS WERE FIRED. ON THE SECOND STAPLER FIRING, A TISSUE PUSH EVENT OCCURRED TOWARDS THE END OF THE STAPLE LINE. THE SURGEON OVERSEWED THE STAPLE LINE AT THE TIME. THERE WERE NO OTHER ISSUES REPORTED DURING THE DA VINCI PROCEDURE. ON AN UNSPECIFIED LATER DATE, A POST-OPERATIVE LEAK WAS DISCOVERED FROM THIS STAPLE LINE. THE PATIENT UNDERWENT A RE-OPERATION TO REPAIR THE LEAK. ON A LATER DATE, THE PATIENT DEVELOPED A COMPLICATION SECONDARY TO THE STAPLE LINE LEAK AND WAS REINTUBATED. THE PATIENT'S FAMILY ULTIMATELY ELECTED TO TAKE THE PATIENT OFF OF LIFE SUPPORT AND THE PATIENT PASSED. NO FURTHER SPECIFIC DETAILS WERE PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
02-12-2024 | Death | intuitive surgical | 03-01-2025 | SUREFORM | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED BARIATRIC GASTRIC REVISION AND LAP BAND REMOVAL PROCEDURE, A STAPLE LINE LEAK OCCURRED REQUIRING RE-OPERATION. THE PATIENT EXPERIENCED AN UNSPECIFIED COMPLICATION AND ULTIMATELY EXPIRED ON POST-OPERATIVE DAY 24. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS INFORMED BY THE SURGEON'S PHYSICIAN ASSISTANT THAT DURING THE ROBOTIC PROCEDURE, TWO SUREFORM 60 BLUE STAPLER RELOADS WERE FIRED. ON THE SECOND STAPLER FIRING, A TISSUE PUSH EVENT OCCURRED TOWARDS THE END OF THE STAPLE LINE. THE SURGEON OVERSEWED THE STAPLE LINE AT THE TIME. THERE WERE NO OTHER ISSUES REPORTED DURING THE DA VINCI PROCEDURE. ON AN UNSPECIFIED LATER DATE, A POST-OPERATIVE LEAK WAS DISCOVERED FROM THIS STAPLE LINE. THE PATIENT UNDERWENT A RE-OPERATION TO REPAIR THE LEAK. ON A LATER DATE, THE PATIENT DEVELOPED A COMPLICATION SECONDARY TO THE STAPLE LINE LEAK AND WAS REINTUBATED. THE PATIENT'S FAMILY ULTIMATELY ELECTED TO TAKE THE PATIENT OFF OF LIFE SUPPORT AND THE PATIENT PASSED. NO FURTHER SPECIFIC DETAILS WERE PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
02-12-2024 | Death | intuitive surgical | 03-01-2025 | SUREFORM | Failure to Cut | A050702 | Insufficient Information | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED BARIATRIC GASTRIC REVISION AND LAP BAND REMOVAL PROCEDURE, A STAPLE LINE LEAK OCCURRED REQUIRING RE-OPERATION. THE PATIENT EXPERIENCED AN UNSPECIFIED COMPLICATION AND ULTIMATELY EXPIRED ON POST-OPERATIVE DAY 24. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS INFORMED BY THE SURGEON'S PHYSICIAN ASSISTANT THAT DURING THE ROBOTIC PROCEDURE, TWO SUREFORM 60 BLUE STAPLER RELOADS WERE FIRED. ON THE SECOND STAPLER FIRING, A TISSUE PUSH EVENT OCCURRED TOWARDS THE END OF THE STAPLE LINE. THE SURGEON OVERSEWED THE STAPLE LINE AT THE TIME. THERE WERE NO OTHER ISSUES REPORTED DURING THE DA VINCI PROCEDURE. ON AN UNSPECIFIED LATER DATE, A POST-OPERATIVE LEAK WAS DISCOVERED FROM THIS STAPLE LINE. THE PATIENT UNDERWENT A RE-OPERATION TO REPAIR THE LEAK. ON A LATER DATE, THE PATIENT DEVELOPED A COMPLICATION SECONDARY TO THE STAPLE LINE LEAK AND WAS REINTUBATED. THE PATIENT'S FAMILY ULTIMATELY ELECTED TO TAKE THE PATIENT OFF OF LIFE SUPPORT AND THE PATIENT PASSED. NO FURTHER SPECIFIC DETAILS WERE PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | EEA ORVIL | Incomplete or Inadequate Connection | A1207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, THE ORAL ANVIL COULD NOT BE ATTACHED TO THE STAPLER. ANOTHER ANVIL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-07-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE COULD NOT BE PUSHED AND DELAYING THE TREATMENT. THERE WAS NO PATIENT INJURY.
02-10-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, DURING THE CALIBRATION OF THE STAPLER, THE STAPLER DID NOT PERFORM ITS INTENDED FUNCTION AFTER INSERTING THE LOAD. THE HANDLE WAS REPLACED FROM ANOTHER BATCH. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA HANDLE THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
02-10-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, DURING THE CALIBRATION OF THE STAPLER, THE STAPLER DID NOT PERFORM ITS INTENDED FUNCTION AFTER INSERTING THE LOAD. THE HANDLE WAS REPLACED FROM ANOTHER BATCH. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA HANDLE THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSOTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSOTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Perforation | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSOTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Ulcer | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSOTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Perforation | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Ulcer | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSOTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSOTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Perforation | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSOTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
15-07-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Ulcer | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY BETWEEN 2008 AND 2022 INCLUDED 158 CASES OF LAPAROSCOPIC SLEEVE GASTRECTOMY (LSG) AND 161 CASES OF LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS (LRYGB). IN THE LSG PROCEDURE, THE STOMACH WAS STAPLED USING 60 MM PURPLE AND TAN ENDOSCOPIC STAPLER RELOADS. A COMPETITOR SUTURE WAS USED TO THE LOWER STAPLE LINE. DURING THE LRYGB PROCEDURE, THE GASTRIC POUCH WAS CREATED USING 30 MM AND A TAN 60 MM ENDOSCOPIC STAPLER RELOADS. A 45 MM RELOAD WAS USED TO FORM THE GASTROJEJUNOSTOMY AND JEJUNOJEJUNOSOTOMY. THE OMEGA LIMB WAS DIVIDED USING A 60 MM ENDOSCOPIC STAPLER. THE JEJUNAL-JEJUNAL AND PETERSENS DEFECTS WERE CLOSED WITH AN UNSPECIFIED ENDO-HERNIA STAPLER. ACROSS THE COHORT POSTOPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK, STAPLER LINE LEAKS, POSTOPERATIVE BLEEDS, ANASTOMOTIC ULCER PERFORATION. INTERVENTIONS INCLUDED REOPERATION. NO INFORMATION WAS GIVEN INTO THE SPECIFICS OF EACH COMPLICATION. COMPLICATIONS OF INTERNAL HERNIAS WERE NOT DEVICE RELATED. LONG-TERM WEIGHT LOSS AND COMORBIDITY RESOLUTION OF LAPAROSCOPIC SLEEVE GASTRECTOMY AND LAPAROSCOPIC ROUXEN-Y GASTRIC BYPASS AND THE IMPACT OF PREOPERATIVE WEIGHT LOSS ON OVERALL OUTCOME DOI: 10.1097/SLE.0000000000001313.
05-03-2025 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A COLORECTAL PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-03-2025 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A COLORECTAL PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RECTAL CANCER SURGERY, WHEN CLOSING THE LOW RECTAL INCISION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON HOWEVER THE HANDLE COULD NOT BE SQUEEZED AND THEREFORE THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RECTAL CANCER SURGERY, WHEN CLOSING THE LOW RECTAL INCISION, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON HOWEVER THE HANDLE COULD NOT BE SQUEEZED AND THEREFORE THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | TRI-STAPLE 2.0 | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE RELOADS, THERE WERE FIRING ISSUES WITH THE HANDLE. SPECIFICALLY, THERE WAS UNCONTROLLED ARTICULATION DURING THE SECOND AND THIRD FIRINGS, AND THE ARTICULATION KNOB AUTO-RESUMED TO NEUTRAL AFTER THE LAST SQUEEZE. THE DEVICE WAS REPLACED WITH ANOTHER HANDLE FOR THE REMAINDER OF THE PROCEDURE. THE SUBSEQUENT FIRINGS (FOURTH TO SIXTH RELOADS) WERE COMPLETED WITHOUT ARTICULATION ISSUES. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | TRI-STAPLE 2.0 | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE RELOADS, THERE WERE FIRING ISSUES WITH THE HANDLE. SPECIFICALLY, THERE WAS UNCONTROLLED ARTICULATION DURING THE SECOND AND THIRD FIRINGS, AND THE ARTICULATION KNOB AUTO-RESUMED TO NEUTRAL AFTER THE LAST SQUEEZE. THE DEVICE WAS REPLACED WITH ANOTHER HANDLE FOR THE REMAINDER OF THE PROCEDURE. THE SUBSEQUENT FIRINGS (FOURTH TO SIXTH RELOADS) WERE COMPLETED WITHOUT ARTICULATION ISSUES. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE RELOADS, THERE WERE FIRING ISSUES WITH THE HANDLE. SPECIFICALLY, THERE WAS UNCONTROLLED ARTICULATION DURING THE SECOND AND THIRD FIRINGS, AND THE ARTICULATION KNOB AUTO-RESUMED TO NEUTRAL AFTER THE LAST SQUEEZE. THE DEVICE WAS REPLACED WITH ANOTHER HANDLE FOR THE REMAINDER OF THE PROCEDURE. THE SUBSEQUENT FIRINGS (FOURTH TO SIXTH RELOADS) WERE COMPLETED WITHOUT ARTICULATION ISSUES. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE RELOADS, THERE WERE FIRING ISSUES WITH THE HANDLE. SPECIFICALLY, THERE WAS UNCONTROLLED ARTICULATION DURING THE SECOND AND THIRD FIRINGS, AND THE ARTICULATION KNOB AUTO-RESUMED TO NEUTRAL AFTER THE LAST SQUEEZE. THE DEVICE WAS REPLACED WITH ANOTHER HANDLE FOR THE REMAINDER OF THE PROCEDURE. THE SUBSEQUENT FIRINGS (FOURTH TO SIXTH RELOADS) WERE COMPLETED WITHOUT ARTICULATION ISSUES. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | TRI-STAPLE 2.0 | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE RELOADS, THERE WERE FIRING ISSUES WITH THE HANDLE. SPECIFICALLY, THERE WAS UNCONTROLLED ARTICULATION DURING THE SECOND AND THIRD FIRINGS, AND THE ARTICULATION KNOB AUTO-RESUMED TO NEUTRAL AFTER THE LAST SQUEEZE. THE DEVICE WAS REPLACED WITH ANOTHER HANDLE FOR THE REMAINDER OF THE PROCEDURE. THE SUBSEQUENT FIRINGS (FOURTH TO SIXTH RELOADS) WERE COMPLETED WITHOUT ARTICULATION ISSUES. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | TRI-STAPLE 2.0 | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE RELOADS, THERE WERE FIRING ISSUES WITH THE HANDLE. SPECIFICALLY, THERE WAS UNCONTROLLED ARTICULATION DURING THE SECOND AND THIRD FIRINGS, AND THE ARTICULATION KNOB AUTO-RESUMED TO NEUTRAL AFTER THE LAST SQUEEZE. THE DEVICE WAS REPLACED WITH ANOTHER HANDLE FOR THE REMAINDER OF THE PROCEDURE. THE SUBSEQUENT FIRINGS (FOURTH TO SIXTH RELOADS) WERE COMPLETED WITHOUT ARTICULATION ISSUES. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 03-01-2025 | OPTIFIX STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE OPTIFIX DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE FIRST ATTEMPT WAS FIXED AS USUAL BUT THE SECOND TIME THE FASTENER FELL OUT AND AFTER THE THIRD ATTEMPT NO FASTENERS DEPLOYED. IT WAS ALSO REPORTED THAT A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
02-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | CIRCULAR MECH XL SEAL 25MM |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLER WAS BEING USED FOR SURGICAL PROCEDURE AND WAS NOT FUNCTIONING PROPERLY. STAPLER WOULD NOT FIRE AND DEPLOY A STAPLE. THE DEVICE WAS NOT ABLE TO COMPLETE ITS INTENDED TASK. STAPLER WAS REMOVED FROM THE FIELD AND A NEW STAPLER WAS USED SUCCESSFULLY. THE MALFUNCTIONING STAPLER WAS PULLED AS A PEEL PACK (REF #ECS25B, LOT #C9EN6F).
10-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHEN THE SURGEON CLAMPED DOWN ON APPENDIX AND ATTEMPTED TO STAPLE. THE GREEN BUTTON WAS PUSHED AND THE STAPLE MECHANISM WAS NOT DEPLOYING THE STAPLES. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE.
19-11-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ENDOPATH |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: DURING AN UNSPECIFIED SURGICAL PROCEDURE, THE CARTRIDGE FOR THE ENDOPATH ETS-FLEX 45 (ATS45) FAILED TO FIRE. THE DEVICE WAS CHANGED, AND A NEW CARTRIDGE WAS INSERTED. THE CARTRIDGE FAILED TO FIRE AGAIN. IT WAS NECESSARY TO SWITCH TO A DIFFERENT INSTRUMENT IN ORDER TO PROCEED WITH THE PROCEDURE.
09-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | UNKNOWN ENDO GIA INSTRUMENT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | UNKNOWN ENDO GIA INSTRUMENT | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | UNKNOWN ENDO GIA INSTRUMENT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
19-08-2024 | Death | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A PROSPECTIVE AND RETROSPECTIVE STUDY PERFORMED BETWEEN 2013 AND 2023, INVOLVING 116 PATIENTS WHO UNDERWENT MINIMALLY INVASIVE DISTAL PANCREATECTOMY (MIDP) WITH OR WITHOUT SPLENECTOMY BETWEEN 2012 AND 2023. PANCREATIC TRANSECTION WAS PERFORMED USING A 60MM STAPLER. SPLENIC VESSELS WERE SECTIONED USING EITHER AN UNKNOWN VASCULAR STAPLER OR COMPETITOR CLIPS. DEVICE RELATED COMPLICATIONS INCLUDED POSTOPERATIVE PANCREATIC FISTULA (POPF), POSTOPERATIVE BLEEDING, AND ABDOMINAL FLUID COLLECTION. ONE DEATH WAS REPORTED DUE TO LATE POSTOPERATIVE UNCONTROLLED BLEEDING.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THE JAWS OF THE 6 RELOADS WOULD NOT FULLY CLOSE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED BY OBTAINING A DIFFERENT BOX FROM THE SHELF. THERE WAS NO PATIENT INVOLVEMENT.
08-06-2016 | Injury | davol inc sub c.r bard | 03-01-2025 | SORBAFIX ENHANCED | Defective Device | A0203 | Neuralgia | Event Description: PER INFORMATION PROVIDED BY PATIENT'S ATTORNEY: ON (B)(6) 2016 - PATIENT WAS DIAGNOSED WITH LEFT INGUINAL HERNIA THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE IMPLANT OF 3DMAX LIGHT MESH (DEVICE #1) WHICH WAS SECURED IN PLACE USING SORBAFIX ENHANCED FASTENER (DEVICE #2). PER OPERATIVE NOTES, ¿THE PATIENT WAS NOTED TO HAVE A INDIRECT INGUINAL HERNIA. THE HERNIA SAC WAS REDUCED AND ABDOMINAL WALL WERE SEPARATED FROM THE ROUND LIGAMENT. A 3DMAX LIGHT MESH (DEVICE #1) WAS PLACED INTO THE PREPERITONEAL SPACE AND SECURED TO THE COOPER'S LIGAMENT AND THE MUSCLE OF THE ANTERIOR ABDOMINAL WALL WITH SORBAFIX ENHANCED ABSORBABLE FASTENERS (DEVICE #2).¿ ON (B)(6) -2016 - PATIENT WAS DIAGNOSED WITH MESHOMA AND NERVE PAIN THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE REMOVAL OF TACKS (DEVICE #2) AND PARTIAL REMOVAL OF MESH (DEVICE #1). PER OPERATIVE NOTES, ¿THERE ARE TWO TACKS STICKING AT THE SUPERIOR ASPECT AND THE LATERAL ASPECT OF MESH. ON THE MEDIAL SIDE IN THE AREA OF COOPER'S LIGAMENT AND IN THE INTERNAL RING, THE MESH LOOKS LIKE FOLDED. THE TACKS (DEVICE #2) IRRITATING SOME SENSITIVE NERVES WERE REMOVED. THE MESHOMA INTEGRATED WITH LOWER ASPECT OF MESH WAS REMOVED. THE MESH IS FOLDED AND CRUMBLED ON THE MIDLINE. THE LOWER PART OF MESHOMA AND MESH IS LEFT BEHIND.¿ ON (B)(6) 2016 - PATIENT WAS DIAGNOSED WITH LEFT INGUINODYNIA, CHRONIC NEUROPATHIC PAIN THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE REMOVAL OF MESH (DEVICE #1). PER OPERATIVE NOTES, ¿IIIOHYPOGASTRIC NERVE WAS TRAPPED IN SCAR. THE ILIOINGUINAL NERVE WAS NEURECTOMIZED. A BULK OF MESH CENTERED AROUND THE INTERNAL RING WAS ADHERENT AND CLOSE TO ILIAC VESSELS. THIS WAS DISSECTED ALONG THE MESH FROM THE FLOOR OF INGUINAL CANAL AND TACK WAS FOUND AT COOPER¿S LIGAMENT. THE MESH (DEVICE #1) WAS DISSECTED. THE ANTERIOR WALL OF THE MESH WAS INCORPORATED WITH THE VESSEL, THEN THE MESH WAS USED TO BUTTRESS THE REPAIR. THE MESH WAS SHAVING OFF, CUTTING AWAY THE BULK OF THE INGUINAL PORTION OF THE MESH OVER THE DIRECT SPACE AND INTERNAL RING WAS TRANSECTED OFF. THERE WAS A SMALL EDGE OF MESH WAS STILL ON THE ILIAC VESSELS WHICH LEFT IN PLACE.¿ ATTORNEY ALLEGES THAT THE PATIENT HAD ADHESIONS, MESH MIGRATION, PAIN & SUFFERING, MESH SHRINKAGE AND HERNIA RECURRENCE. IT IS ALSO ALLEGED DIFFICULT AND UNABLE TO REMOVAL ALL OF MESH, EMOTIONAL/PSYCHOLOGICAL INJURIES AND BLURRED VISION.
08-06-2016 | Injury | davol inc sub c.r bard | 03-01-2025 | SORBAFIX ENHANCED | Patient Device Interaction Problem | A01 | Neuralgia | Event Description: PER INFORMATION PROVIDED BY PATIENT'S ATTORNEY: ON (B)(6) 2016 - PATIENT WAS DIAGNOSED WITH LEFT INGUINAL HERNIA THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE IMPLANT OF 3DMAX LIGHT MESH (DEVICE #1) WHICH WAS SECURED IN PLACE USING SORBAFIX ENHANCED FASTENER (DEVICE #2). PER OPERATIVE NOTES, ¿THE PATIENT WAS NOTED TO HAVE A INDIRECT INGUINAL HERNIA. THE HERNIA SAC WAS REDUCED AND ABDOMINAL WALL WERE SEPARATED FROM THE ROUND LIGAMENT. A 3DMAX LIGHT MESH (DEVICE #1) WAS PLACED INTO THE PREPERITONEAL SPACE AND SECURED TO THE COOPER'S LIGAMENT AND THE MUSCLE OF THE ANTERIOR ABDOMINAL WALL WITH SORBAFIX ENHANCED ABSORBABLE FASTENERS (DEVICE #2).¿ ON (B)(6) -2016 - PATIENT WAS DIAGNOSED WITH MESHOMA AND NERVE PAIN THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE REMOVAL OF TACKS (DEVICE #2) AND PARTIAL REMOVAL OF MESH (DEVICE #1). PER OPERATIVE NOTES, ¿THERE ARE TWO TACKS STICKING AT THE SUPERIOR ASPECT AND THE LATERAL ASPECT OF MESH. ON THE MEDIAL SIDE IN THE AREA OF COOPER'S LIGAMENT AND IN THE INTERNAL RING, THE MESH LOOKS LIKE FOLDED. THE TACKS (DEVICE #2) IRRITATING SOME SENSITIVE NERVES WERE REMOVED. THE MESHOMA INTEGRATED WITH LOWER ASPECT OF MESH WAS REMOVED. THE MESH IS FOLDED AND CRUMBLED ON THE MIDLINE. THE LOWER PART OF MESHOMA AND MESH IS LEFT BEHIND.¿ ON (B)(6) 2016 - PATIENT WAS DIAGNOSED WITH LEFT INGUINODYNIA, CHRONIC NEUROPATHIC PAIN THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE REMOVAL OF MESH (DEVICE #1). PER OPERATIVE NOTES, ¿IIIOHYPOGASTRIC NERVE WAS TRAPPED IN SCAR. THE ILIOINGUINAL NERVE WAS NEURECTOMIZED. A BULK OF MESH CENTERED AROUND THE INTERNAL RING WAS ADHERENT AND CLOSE TO ILIAC VESSELS. THIS WAS DISSECTED ALONG THE MESH FROM THE FLOOR OF INGUINAL CANAL AND TACK WAS FOUND AT COOPER¿S LIGAMENT. THE MESH (DEVICE #1) WAS DISSECTED. THE ANTERIOR WALL OF THE MESH WAS INCORPORATED WITH THE VESSEL, THEN THE MESH WAS USED TO BUTTRESS THE REPAIR. THE MESH WAS SHAVING OFF, CUTTING AWAY THE BULK OF THE INGUINAL PORTION OF THE MESH OVER THE DIRECT SPACE AND INTERNAL RING WAS TRANSECTED OFF. THERE WAS A SMALL EDGE OF MESH WAS STILL ON THE ILIAC VESSELS WHICH LEFT IN PLACE.¿ ATTORNEY ALLEGES THAT THE PATIENT HAD ADHESIONS, MESH MIGRATION, PAIN & SUFFERING, MESH SHRINKAGE AND HERNIA RECURRENCE. IT IS ALSO ALLEGED DIFFICULT AND UNABLE TO REMOVAL ALL OF MESH, EMOTIONAL/PSYCHOLOGICAL INJURIES AND BLURRED VISION.
05-09-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Circuit Failure | A0721 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, IT WAS RESTARTED USING A POWER APPROACH (LONG PRESSED THE LEFT AND RIGHT GREEN BUTTONS AND PLUGGED IT BACK INTO THE CHARGER), BUT IT DID NOT RECOVER, THERE WAS A POWER HANDLE ERROR WITH RED CIRCLE SLASH AND YELLOW EXCLAMATION MARK WAS DISPLAYED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WERE MULTIPLE LOGS WHERE THE HANDLE FAILED MOTOR TEST ON INITIALIZATION.
05-09-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, IT WAS RESTARTED USING A POWER APPROACH (LONG PRESSED THE LEFT AND RIGHT GREEN BUTTONS AND PLUGGED IT BACK INTO THE CHARGER), BUT IT DID NOT RECOVER, THERE WAS A POWER HANDLE ERROR WITH RED CIRCLE SLASH AND YELLOW EXCLAMATION MARK WAS DISPLAYED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THERE WERE MULTIPLE LOGS WHERE THE HANDLE FAILED MOTOR TEST ON INITIALIZATION.
04-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A DUODENECTOMY, DURING JEJUNAL RESECTION, AFTER FIRING, THE STAPLE LINE GOT CAUGHT ON THE RIGHT SIDE OF THE CARTRIDGE ANVIL, CAUSING DIFFICULTY IN RELEASING SMOOTHLY. SINCE IT WAS ON THE SPECIMEN SIDE, IT WAS FORCIBLY DETACHED WITH SOME EFFORT. IT WAS NOTED THAT THERE WAS NO OPENING OF THE STAPLE LINE. THE SURGERY WAS COMPLETED WITHOUT ANY PROBLEMS. TO RESOLVE THE ISSUE, THE SAME HANDLE, ADAPTER, AND SHELL WERE USED WITH THE NEW RELOAD.
09-12-2024 | Injury | intuitive surgical | 03-01-2025 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, THE FIRST FIRE OF A SUREFORM 30 GREY RELOAD HAD AN INCOMPLETE STAPLE LINE WHILE STAPLING ACROSS A BRANCH OF THE PULMONARY ARTERY. THERE WERE NO COMPLICATIONS, PAUSES, OR ERROR MESSAGES DURING THE FIRING SEQUENCE. THE STAPLE LINE WAS INCOMPLETE AND MALFORMED STAPLES WERE PRODUCED. THE SURGEON RESOLVED THE DEFECT BY PLACING A CLIP AROUND THE ARTERY. THE SUREFORM 30 STAPLER INSTRUMENT WAS NOT USED AFTER THE EVENT. A SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT WAS UTILIZED FOR THE REMAINDER OF THE PROCEDURE, WHICH WAS COMPLETED ROBOTICALLY WITH NO FURTHER ISSUES. THE PATIENT IS RECOVERING WELL.
11-12-2024 | Malfunction | ethicon | 03-01-2025 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNSPECIFIED PROCEDURE ON (B)(6) 2024 AND ABSORBABLE STRAPS WERE USED. THE DEVICE DOES NOT SHOOT PROPERLY, SOMETIMES IT DOES NOT SHOOT AND SOMETIMES IT SHOOTS 2 STRAPS AND THEY DO NOT GRAB THE TISSUE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SPLENECTOMY, THE JAWS DID NOT CLOSE AT ALL. TO RESOLVE THE ISSUE, THE HANDLE WAS NOT CHANGED, BUT THE CARTRIDGE WAS CHANGED TO A NEW ONE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLER DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLER DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
07-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE STAPLER, THERE WERE FIRING ISSUES WHERE THE STAPLER DID NOT STAPLE OR CUT. THE STAPLER WAS REPLACED WITH ANOTHER STAPLER WHICH FUNCTIONED CORRECTLY. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT.
06-11-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC MEDIASTINAL TUMOR RESECTION, BEFORE CUTTING THE MEDIASTINUM TISSUE, THE NURSE CONNECTED THE RELOAD AND HELD THE GRIP TIGHTLY, BUT THE JAWS DID NOT CLOSE. TO RESOLVE THE ISSUE, A NEW RELOAD AND HANDLE WERE USED. THERE WAS NO PATIENT INJURY.
20-03-2024 | Injury | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY CONSISTED OF 92 PATIENTS WHO UNDERWENT ROBOT-ASSISTED THORACOSCOPIC SEGMENTECTOMY (RATS) AND 82 PATIENTS WHO UNDERWENT COMPLETE VIDEO-ASSISTED THORACOSCOPIC SEGMENTECTOMIES (CVATS) BETWEEN 2020 TO 2023, IN THE CVATS PROCEDURE, THE LUNG WAS STAPLED USING EITHER THE ENDO GIA STAPLER OR A COMPETITOR STAPLER. POSTOPERATIVE COMPLICATION INCLUDED PROLONGED AIR LEAK LASTING MORE THAN SEVEN DAYS OR REQUIRING PLEURODESIS IN NINE PATIENTS.
12-11-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC SEGMENTECTOMY, WHEN REMOVAL OF LUNG TISSUE, THE STAPLER WAS ABLE TO FIRE, THE STAPLES WERE NOT COMPLETELY STAPLED ON THE LUNG, AND ALL OF THEM FELL OFF, AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY. ANOTHER RELOAD WAS USED, BUT SAME ISSUE OCCURRED. A REPLACEMENT DEVICE FROM ANOTHER MANUFACTURER WAS DONE TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE SHEARED TEETH WERE VISIBLE ON EACH INSTRUMENTS FIRING RACK.
12-11-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC SEGMENTECTOMY, WHEN REMOVAL OF LUNG TISSUE, THE STAPLER WAS ABLE TO FIRE, THE STAPLES WERE NOT COMPLETELY STAPLED ON THE LUNG, AND ALL OF THEM FELL OFF, AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY. ANOTHER RELOAD WAS USED, BUT SAME ISSUE OCCURRED. A REPLACEMENT DEVICE FROM ANOTHER MANUFACTURER WAS DONE TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE SHEARED TEETH WERE VISIBLE ON EACH INSTRUMENTS FIRING RACK.
10-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE PROXIMAL STAPLES POKED OUT OF THE CARTRIDGE UNFORMED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Fail-Safe Problem | A1602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE PROXIMAL STAPLES POKED OUT OF THE CARTRIDGE UNFORMED. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHEN THE SURGEON CLAMPED DOWN ON APPENDIX AND ATTEMPTED TO STAPLE. THE GREEN BUTTON WAS PUSHED AND THE STAPLE MECHANISM WAS NOT DEPLOYING THE STAPLES. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE.
12-11-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC SEGMENTECTOMY, WHEN REMOVAL OF LUNG TISSUE, THE STAPLER WAS ABLE TO FIRE, THE STAPLES WERE NOT COMPLETELY STAPLED ON THE LUNG, AND ALL OF THEM FELL OFF, AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY. ANOTHER RELOAD WAS USED, BUT SAME ISSUE OCCURRED. A REPLACEMENT DEVICE FROM ANOTHER MANUFACTURER WAS DONE TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE SHEARED TEETH WERE VISIBLE ON EACH INSTRUMENTS FIRING RACK.
12-11-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC SEGMENTECTOMY, WHEN REMOVAL OF LUNG TISSUE, THE STAPLER WAS ABLE TO FIRE, THE STAPLES WERE NOT COMPLETELY STAPLED ON THE LUNG, AND ALL OF THEM FELL OFF, AND THE STAPLE LINE WAS INCOMPLETE CENTRALLY. ANOTHER RELOAD WAS USED, BUT SAME ISSUE OCCURRED. A REPLACEMENT DEVICE FROM ANOTHER MANUFACTURER WAS DONE TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT THE SHEARED TEETH WERE VISIBLE ON EACH INSTRUMENTS FIRING RACK.
02-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE EGIA60AMT EGIA 60 ARTIC MED THICK SULU DURING SET-UP AND PRIOR TO FIRING. THE SPECIFIC PROBLEMS INCLUDED JAW ISSUES, DIFFICULTY CLOSING THE JAWS, AND THE JAWS NOT CLOSING COMPLETELY. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH ANOTHER FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
02-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE EGIA60AMT EGIA 60 ARTIC MED THICK SULU DURING SET-UP AND PRIOR TO FIRING. THE SPECIFIC PROBLEMS INCLUDED JAW ISSUES, DIFFICULTY CLOSING THE JAWS, AND THE JAWS NOT CLOSING COMPLETELY. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH ANOTHER FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
20-12-2024 | Injury | ethicon endo-surgery | 03-01-2025 | ECHELON LINEAR CUTTER 60MM | Failure to Fire | A050501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN ABDOMINAL PERITONEUM UPON FIRING THE DEVICE FOR AN ENTEROTOMY CLOSURE DURING A SIDE TO SIDE ANASTOMOSES, HE NOTICED THERE WAS A HOLE ON THE OPPOSITE SIDE OF THE ENTEROTOMY. THIS NEEDED TO BE MANUALLY CLOSED WITH SUTURE. PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A DELAY OF TWENTY MINUTES. PATIENT CONSEQUENCES UNKNOWN.
20-12-2024 | Injury | ethicon endo-surgery | 03-01-2025 | ECHELON LINEAR CUTTER 60MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN ABDOMINAL PERITONEUM UPON FIRING THE DEVICE FOR AN ENTEROTOMY CLOSURE DURING A SIDE TO SIDE ANASTOMOSES, HE NOTICED THERE WAS A HOLE ON THE OPPOSITE SIDE OF THE ENTEROTOMY. THIS NEEDED TO BE MANUALLY CLOSED WITH SUTURE. PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A DELAY OF TWENTY MINUTES. PATIENT CONSEQUENCES UNKNOWN.
02-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE EGIA60AMT EGIA 60 ARTIC MED THICK SULU DURING SET-UP AND PRIOR TO FIRING. THE SPECIFIC PROBLEMS INCLUDED JAW ISSUES, DIFFICULTY CLOSING THE JAWS, AND THE JAWS NOT CLOSING COMPLETELY. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH ANOTHER FROM A DIFFERENT LOT NUMBER. THERE WAS NO PATIENT INVOLVEMENT.
08-10-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY ON 148 PATIENTS BETWEEN JULY 2023 AND NOVEMBER 2023 WERE DIAGNOSED WITH OBESITY WHO UNDERWENT LAPAROSCOPIC GASTRIC BYPASS WITH EITHER OMENTOPEXY OR CLIP PLACEMENT ALONG THE STAPLE LINE. THE MANUAL STAPLER WAS USED FOR GASTRIC RESECTION, THE OMENTOPEXY WAS COMPLETED USING A BARBED SUTURE, AND THE CLIPPING WAS PERFORMED WITH THE CLIP APPLIER. POSTOPERATIVE BLEEDING OCCURRED IN TWO OMENTOPEXY PATIENTS AND 11 CLIP PATIENTS. BLOOD TRANSFUSIONS AND COMPUTED TOMOGRAPHY SCANS WERE REQUIRED FOR TWO PATIENTS IN THE CLIP GROUP DUE TO A HEMATOMA ADJACENT TO THE STAPLE LINE. ADDITIONAL COMPLICATIONS UNRELATED TO THE DEVICE INCLUDED WOUND HEMATOMA, WOUND INFECTION AND READMISSION DUE TO NAUSEA AND VOMITING.
08-10-2024 | Injury | covidien lp llc north haven | 03-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY ON 148 PATIENTS BETWEEN JULY 2023 AND NOVEMBER 2023 WERE DIAGNOSED WITH OBESITY WHO UNDERWENT LAPAROSCOPIC GASTRIC BYPASS WITH EITHER OMENTOPEXY OR CLIP PLACEMENT ALONG THE STAPLE LINE. THE MANUAL STAPLER WAS USED FOR GASTRIC RESECTION, THE OMENTOPEXY WAS COMPLETED USING A BARBED SUTURE, AND THE CLIPPING WAS PERFORMED WITH THE CLIP APPLIER. POSTOPERATIVE BLEEDING OCCURRED IN TWO OMENTOPEXY PATIENTS AND 11 CLIP PATIENTS. BLOOD TRANSFUSIONS AND COMPUTED TOMOGRAPHY SCANS WERE REQUIRED FOR TWO PATIENTS IN THE CLIP GROUP DUE TO A HEMATOMA ADJACENT TO THE STAPLE LINE. ADDITIONAL COMPLICATIONS UNRELATED TO THE DEVICE INCLUDED WOUND HEMATOMA, WOUND INFECTION AND READMISSION DUE TO NAUSEA AND VOMITING.
01-01-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DELIVERED STAPLES BUT ONLY ON ONE SIDE OF THE CUT LINE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE CUT AFTER FIRING BUT THE INTEGRITY OF THE CUTLINE WAS QUESTIONABLE. THE CUTLINE AND STAPLE FORMATION APPEARED TO NOT DELIVER THE EXPECTED RESULT AS IN PREVIOUS SURGERIES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE CUT AFTER FIRING BUT THE INTEGRITY OF THE CUTLINE WAS QUESTIONABLE. THE CUTLINE AND STAPLE FORMATION APPEARED TO NOT DELIVER THE EXPECTED RESULT AS IN PREVIOUS SURGERIES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE CUT AFTER FIRING BUT THE INTEGRITY OF THE CUTLINE WAS QUESTIONABLE. THE CUTLINE AND STAPLE FORMATION APPEARED TO NOT DELIVER THE EXPECTED RESULT AS IN PREVIOUS SURGERIES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE CUT AFTER FIRING BUT THE INTEGRITY OF THE CUTLINE WAS QUESTIONABLE. THE CUTLINE AND STAPLE FORMATION APPEARED TO NOT DELIVER THE EXPECTED RESULT AS IN PREVIOUS SURGERIES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
18-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER THREE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
18-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER THREE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
18-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER THREE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
18-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | WHITE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER THREE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
01-01-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE FIRED FOR 1CM AND THEN LOCKED OUT WHEN TRYING TO FIRE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
06-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE PURPLE RELOAD PRESENTED A DEFECT, WHICH WAS NOT BEING ABLE TO CLOSE THE STAPLER. EVEN CHANGING THE STAPLER, THE PROBLEM CONTINUED AND DETECTED THAT THE PROBLEM WAS IN THE RELOAD. THERE WAS NO PATIENT INJURY.
11-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, IT WAS OBSERVED THAT THE STAPLER STOPPED FIRING HALFWAY THROUGH WHILE FIRING THROUGH THE BOWEL. SURGEON ATTEMPTED TO FINISH THE FIRING CYCLE BUT IT WOULD NOT. HOME BUTTON WAS ATTEMPTED BUT WOULD NOT RETURN THE KNIFE HOME AND THE DEVICE COULD NOT BE OPENED. MANUAL OVERRIDE WAS USED TO OPEN THE DEVICE. ANOTHER LIKE DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
11-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON 3000 60MM LONG | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, IT WAS OBSERVED THAT THE STAPLER STOPPED FIRING HALFWAY THROUGH WHILE FIRING THROUGH THE BOWEL. SURGEON ATTEMPTED TO FINISH THE FIRING CYCLE BUT IT WOULD NOT. HOME BUTTON WAS ATTEMPTED BUT WOULD NOT RETURN THE KNIFE HOME AND THE DEVICE COULD NOT BE OPENED. MANUAL OVERRIDE WAS USED TO OPEN THE DEVICE. ANOTHER LIKE DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
11-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON 3000 60MM LONG | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, IT WAS OBSERVED THAT THE STAPLER STOPPED FIRING HALFWAY THROUGH WHILE FIRING THROUGH THE BOWEL. SURGEON ATTEMPTED TO FINISH THE FIRING CYCLE BUT IT WOULD NOT. HOME BUTTON WAS ATTEMPTED BUT WOULD NOT RETURN THE KNIFE HOME AND THE DEVICE COULD NOT BE OPENED. MANUAL OVERRIDE WAS USED TO OPEN THE DEVICE. ANOTHER LIKE DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
21-01-2025 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN RESPIRATORY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE MOVED FORWARD SLIGHTLY AND STOPPED ON THE WAY AT THE 2ND FIRING. THE DEVICE WAS LOCKED OUT. THERE WAS NO DAMAGE TO THE TISSUE, AND THE PRODUCT WAS REPLACED WITH A NEW ONE, AND THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
06-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE PURPLE RELOAD PRESENTED A DEFECT, WHICH WAS NOT BEING ABLE TO CLOSE THE STAPLER. EVEN CHANGING THE STAPLER, THE PROBLEM CONTINUED AND DETECTED THAT THE PROBLEM WAS IN THE RELOAD. THERE WAS NO PATIENT INJURY.
11-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLING WAS NOT COMPLETED WHEN THE DEVICE WAS FIRED ON LUNG VEIN. STAPLER COULD NOT BE OPENED; AND THEREFORE, HAD TO BE OPEN MANUALLY WITH OVERRIDE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH AN UNSPECIFIED DELAY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. ADDITIONAL INFORMATION REQUESTED.
11-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VATS LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLING WAS NOT COMPLETED WHEN THE DEVICE WAS FIRED ON LUNG VEIN. STAPLER COULD NOT BE OPENED; AND THEREFORE, HAD TO BE OPEN MANUALLY WITH OVERRIDE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH AN UNSPECIFIED DELAY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. ADDITIONAL INFORMATION REQUESTED.
17-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A SEMI-OPEN PNEUMONECTOMY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE MOVED FORWARD SLIGHTLY AND STOPPED AT THE 3RD FIRING. THE DEVICE WAS LOCKED OUT. THE DEVICE WAS USED ON THE BLOOD VESSEL. THE CARTRIDGE COLOR WAS WHITE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
19-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | ECHELON 3000 60MM STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE JAW ON THE DEVICE DID NOT OPEN AFTER SECOND FIRING. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
18-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, A GOLD RELOAD WAS FIRED. DURING THE FIRING, THE SURGEON STATED THAT A GRINDING NOISE BEGAN AND THE KNIFE SLOWED DOWN CONSIDERABLY. SOME OF THE STAPLES WERE FORMED, SOME WERE COMPLETELY OPEN, AND SOME WERE MANGLED. A HOLE WAS LEFT THAT HAD TO BE OVERSEWN ADDING TIME. DR STATES THE TISSUE WAS NOT THICK AND NOTHING WAS IN THE STOMACH OTHER THAN THE BOUGIE AND NOTHING WAS CLAMPED OR FIRED UPON OTHER THAN TISSUE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY COMPLETED WITH A DELAY OF FIVE MINUTES. ADDITIONAL INFORMATION REQUESTED.
18-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, A GOLD RELOAD WAS FIRED. DURING THE FIRING, THE SURGEON STATED THAT A GRINDING NOISE BEGAN AND THE KNIFE SLOWED DOWN CONSIDERABLY. SOME OF THE STAPLES WERE FORMED, SOME WERE COMPLETELY OPEN, AND SOME WERE MANGLED. A HOLE WAS LEFT THAT HAD TO BE OVERSEWN ADDING TIME. DR STATES THE TISSUE WAS NOT THICK AND NOTHING WAS IN THE STOMACH OTHER THAN THE BOUGIE AND NOTHING WAS CLAMPED OR FIRED UPON OTHER THAN TISSUE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY COMPLETED WITH A DELAY OF FIVE MINUTES. ADDITIONAL INFORMATION REQUESTED.
18-12-2024 | Malfunction | ethicon endo-surgery | 03-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, A GOLD RELOAD WAS FIRED. DURING THE FIRING, THE SURGEON STATED THAT A GRINDING NOISE BEGAN AND THE KNIFE SLOWED DOWN CONSIDERABLY. SOME OF THE STAPLES WERE FORMED, SOME WERE COMPLETELY OPEN, AND SOME WERE MANGLED. A HOLE WAS LEFT THAT HAD TO BE OVERSEWN ADDING TIME. DR STATES THE TISSUE WAS NOT THICK AND NOTHING WAS IN THE STOMACH OTHER THAN THE BOUGIE AND NOTHING WAS CLAMPED OR FIRED UPON OTHER THAN TISSUE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY COMPLETED WITH A DELAY OF FIVE MINUTES. ADDITIONAL INFORMATION REQUESTED.
09-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PATIENT'S PROCEDURE INVOLVING THE EGIA60AMT EGIA 60 ARTIC MED THICK SULU RELOAD, THE STAPLER CHARGE FAILED TO FIRE. THE STANDARD TESTS WERE CONDUCTED TWICE, BUT THE ISSUE PERSISTED, NECESSITATING THE REPLACEMENT OF THE RELOAD WITH ANOTHER LOAD. THERE WAS NO PATIENT INJURY.
06-12-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE PURPLE RELOAD PRESENTED A DEFECT, WHICH WAS NOT BEING ABLE TO CLOSE THE STAPLER. EVEN CHANGING THE STAPLER, THE PROBLEM CONTINUED AND DETECTED THAT THE PROBLEM WAS IN THE RELOAD. THERE WAS NO PATIENT INJURY.
04-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL STOMACH-PRESERVING PANCREATICODUODENECTOMY (SSPPD) GASTRIC RESECTION, ON THE FIRST STAP LING, THE STAPLES DID NOT FORM A B-SHAPE, AND THE CUT END WAS IN A MESSED-UP STATE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY APPROXIMATELY TEN MINUTES DUE TO THE ISSUE.
04-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Failure to Cut | A050702 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A SUBTOTAL STOMACH-PRESERVING PANCREATICODUODENECTOMY (SSPPD) GASTRIC RESECTION, ON THE FIRST STAP LING, THE STAPLES DID NOT FORM A B-SHAPE, AND THE CUT END WAS IN A MESSED-UP STATE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY APPROXIMATELY TEN MINUTES DUE TO THE ISSUE.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN THE SECOND SHOT, THE LOAD DID NOT ADVANCE TO THE END OF THE STAPLE, AND THE KNIFE DID NOT RETRACT AS WELL, LOCKING THE JAWS OF THE DEVICE TO THE TISSUE. ALL POSSIBLE SOLUTIONS HAVE BEEN DONE, SUCH AS RESTARTING THE HANDLE, DISCONNECTING THE ADAPTER, AND USING THE MANUAL RETRACTION TOOL. SHELL REPLACEMENT WAS ALSO ATTEMPTED, BUT ALL OF THESE DID NOT RESOLVE THE ISSUE. TO COMPLETE THE CASE, A MANUAL STAPLER FROM ANOTHER MANUFACTURER WAS USED. IT WAS NOTED THAT THE PROCEDURE WAS EXTENDED FOR MORE THAN THREE HOURS.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN THE SECOND SHOT, THE LOAD DID NOT ADVANCE TO THE END OF THE STAPLE, AND THE KNIFE DID NOT RETRACT AS WELL, LOCKING THE JAWS OF THE DEVICE TO THE TISSUE. ALL POSSIBLE SOLUTIONS HAVE BEEN DONE, SUCH AS RESTARTING THE HANDLE, DISCONNECTING THE ADAPTER, AND USING THE MANUAL RETRACTION TOOL. SHELL REPLACEMENT WAS ALSO ATTEMPTED, BUT ALL OF THESE DID NOT RESOLVE THE ISSUE. TO COMPLETE THE CASE, A MANUAL STAPLER FROM ANOTHER MANUFACTURER WAS USED. IT WAS NOTED THAT THE PROCEDURE WAS EXTENDED FOR MORE THAN THREE HOURS.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN THE SECOND SHOT, THE LOAD DID NOT ADVANCE TO THE END OF THE STAPLE, AND THE KNIFE DID NOT RETRACT AS WELL, LOCKING THE JAWS OF THE DEVICE TO THE TISSUE. ALL POSSIBLE SOLUTIONS HAVE BEEN DONE, SUCH AS RESTARTING THE HANDLE, DISCONNECTING THE ADAPTER, AND USING THE MANUAL RETRACTION TOOL. SHELL REPLACEMENT WAS ALSO ATTEMPTED, BUT ALL OF THESE DID NOT RESOLVE THE ISSUE. TO COMPLETE THE CASE, A MANUAL STAPLER FROM ANOTHER MANUFACTURER WAS USED. IT WAS NOTED THAT THE PROCEDURE WAS EXTENDED FOR MORE THAN THREE HOURS.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN THE SECOND SHOT, THE LOAD DID NOT ADVANCE TO THE END OF THE STAPLE, AND THE KNIFE DID NOT RETRACT AS WELL, LOCKING THE JAWS OF THE DEVICE TO THE TISSUE. ALL POSSIBLE SOLUTIONS HAVE BEEN DONE, SUCH AS RESTARTING THE HANDLE, DISCONNECTING THE ADAPTER, AND USING THE MANUAL RETRACTION TOOL. SHELL REPLACEMENT WAS ALSO ATTEMPTED, BUT ALL OF THESE DID NOT RESOLVE THE ISSUE. TO COMPLETE THE CASE, A MANUAL STAPLER FROM ANOTHER MANUFACTURER WAS USED. IT WAS NOTED THAT THE PROCEDURE WAS EXTENDED FOR MORE THAN THREE HOURS.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN THE SECOND SHOT, THE LOAD DID NOT ADVANCE TO THE END OF THE STAPLE, AND THE KNIFE DID NOT RETRACT AS WELL, LOCKING THE JAWS OF THE DEVICE TO THE TISSUE. ALL POSSIBLE SOLUTIONS HAVE BEEN DONE, SUCH AS RESTARTING THE HANDLE, DISCONNECTING THE ADAPTER, AND USING THE MANUAL RETRACTION TOOL. SHELL REPLACEMENT WAS ALSO ATTEMPTED, BUT ALL OF THESE DID NOT RESOLVE THE ISSUE. TO COMPLETE THE CASE, A MANUAL STAPLER FROM ANOTHER MANUFACTURER WAS USED. IT WAS NOTED THAT THE PROCEDURE WAS EXTENDED FOR MORE THAN THREE HOURS.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IN THE SECOND SHOT, THE LOAD DID NOT ADVANCE TO THE END OF THE STAPLE, AND THE KNIFE DID NOT RETRACT AS WELL, LOCKING THE JAWS OF THE DEVICE TO THE TISSUE. ALL POSSIBLE SOLUTIONS HAVE BEEN DONE, SUCH AS RESTARTING THE HANDLE, DISCONNECTING THE ADAPTER, AND USING THE MANUAL RETRACTION TOOL. SHELL REPLACEMENT WAS ALSO ATTEMPTED, BUT ALL OF THESE DID NOT RESOLVE THE ISSUE. TO COMPLETE THE CASE, A MANUAL STAPLER FROM ANOTHER MANUFACTURER WAS USED. IT WAS NOTED THAT THE PROCEDURE WAS EXTENDED FOR MORE THAN THREE HOURS.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY LOBECTOMY, IN THE SECOND SHOT, THE LOAD DID NOT ADVANCE TO THE END OF THE STAPLE, AND THE KNIFE DID NOT RETRACT AS WELL, LOCKING THE JAWS OF THE DEVICE TO THE TISSUE. ALL POSSIBLE SOLUTIONS HAVE BEEN DONE, SUCH AS RESTARTING THE HANDLE, DISCONNECTING THE ADAPTER, AND USING THE MANUAL RETRACTION TOOL. SHELL REPLACEMENT WAS ALSO ATTEMPTED, BUT ALL OF THESE DID NOT RESOLVE THE ISSUE. TO COMPLETE THE CASE, A MANUAL STAPLER FROM ANOTHER MANUFACTURER WAS USED. THERE WAS NO ACTIVE BLEEDING DUE TO WHAT HAPPENED, BUT THE PROCEDURE WAS EXTENDED FOR MORE THAN THREE HOURS AND THE RETRACTION TOOL WORKED WITH OTHER DEVICES. IT WAS NOTED, AFTER MANUAL STAPLING, THE USER REPLACED THE ROD WITH A NEW ONE, AND THE PROCEDURE CONTINUED WITHOUT COMPLICATIONS.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY LOBECTOMY, IN THE SECOND SHOT, THE LOAD DID NOT ADVANCE TO THE END OF THE STAPLE, AND THE KNIFE DID NOT RETRACT AS WELL, LOCKING THE JAWS OF THE DEVICE TO THE TISSUE. ALL POSSIBLE SOLUTIONS HAVE BEEN DONE, SUCH AS RESTARTING THE HANDLE, DISCONNECTING THE ADAPTER, AND USING THE MANUAL RETRACTION TOOL. SHELL REPLACEMENT WAS ALSO ATTEMPTED, BUT ALL OF THESE DID NOT RESOLVE THE ISSUE. TO COMPLETE THE CASE, A MANUAL STAPLER FROM ANOTHER MANUFACTURER WAS USED. THERE WAS NO ACTIVE BLEEDING DUE TO WHAT HAPPENED, BUT THE PROCEDURE WAS EXTENDED FOR MORE THAN THREE HOURS AND THE RETRACTION TOOL WORKED WITH OTHER DEVICES. IT WAS NOTED, AFTER MANUAL STAPLING, THE USER REPLACED THE ROD WITH A NEW ONE, AND THE PROCEDURE CONTINUED WITHOUT COMPLICATIONS.
11-12-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY LOBECTOMY, IN THE SECOND SHOT, THE LOAD DID NOT ADVANCE TO THE END OF THE STAPLE, AND THE KNIFE DID NOT RETRACT AS WELL, LOCKING THE JAWS OF THE DEVICE TO THE TISSUE. ALL POSSIBLE SOLUTIONS HAVE BEEN DONE, SUCH AS RESTARTING THE HANDLE, DISCONNECTING THE ADAPTER, AND USING THE MANUAL RETRACTION TOOL. SHELL REPLACEMENT WAS ALSO ATTEMPTED, BUT ALL OF THESE DID NOT RESOLVE THE ISSUE. TO COMPLETE THE CASE, A MANUAL STAPLER FROM ANOTHER MANUFACTURER WAS USED. THERE WAS NO ACTIVE BLEEDING DUE TO WHAT HAPPENED, BUT THE PROCEDURE WAS EXTENDED FOR MORE THAN THREE HOURS AND THE RETRACTION TOOL WORKED WITH OTHER DEVICES. IT WAS NOTED, AFTER MANUAL STAPLING, THE USER REPLACED THE ROD WITH A NEW ONE, AND THE PROCEDURE CONTINUED WITHOUT COMPLICATIONS.
24-06-2024 | Malfunction | us surgical puerto rico | 03-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION PROCEDURE, AFTER THE RELOAD WAS INSTALLED, IT WAS UNABLE TO FIRE. AFTER REPLACING IT WITH A NEW CUTTING STAPLER, IT COULD BE USED NORMALLY.
24-06-2024 | Malfunction | covidien lp llc north haven | 03-01-2025 | UNKNOWN EGIA SU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG RESECTION PROCEDURE, AFTER THE RELOAD WAS INSTALLED, IT WAS UNABLE TO FIRE. AFTER REPLACING IT WITH A NEW CUTTING STAPLER, IT COULD BE USED NORMALLY.
12-12-2024 | Malfunction | us surgical puerto rico | 04-01-2025 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SIGMOID COLON CANCER RESECTION, DURING SIGMOID COLON RESECTION, THE JAWS OF THE DEVICE LOCKED ON THE TISSUE. TO REMOVE THE DEVICE, THE SURGEON USED THE SURGICAL HOOK TO HOOK THE I-BEAM BACK. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien lp llc north haven | 04-01-2025 | UNKNOWN EGIA TRI STAPLE | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC SIGMOID COLON CANCER RESECTION, DURING SIGMOID COLON RESECTION, THE JAWS OF THE DEVICE LOCKED ON THE TISSUE. TO REMOVE THE DEVICE, THE SURGEON USED THE SURGICAL HOOK TO HOOK THE I-BEAM BACK. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | us surgical puerto rico | 04-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, IN A SINGLE LOOP BYPASS, AFTER FIRING DURING STOMACH TRANSECTION, THE STAPLE LINE WAS BLEEDING. THE STAPLE LINE WAS MORE OOZY THAN NORMAL. TWO RELOADS HAD THE ISSUE. DIATHERMY AND SUTURING WERE DONE AS A STAPLE LINE REINFORCEMENT TO ACHIEVE HEMOSTASIS.
12-12-2024 | Malfunction | us surgical puerto rico | 04-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, IN A SINGLE LOOP BYPASS, AFTER FIRING DURING STOMACH TRANSECTION, THE STAPLE LINE WAS BLEEDING. THE STAPLE LINE WAS MORE OOZY THAN NORMAL. TWO RELOADS HAD THE ISSUE. DIATHERMY AND SUTURING WERE DONE AS A STAPLE LINE REINFORCEMENT TO ACHIEVE HEMOSTASIS.
10-10-2024 | Malfunction | covidien lp llc north haven | 04-01-2025 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, DURING INSERTION INTO THE SHELL, THERE WAS A GRAPHIC WITH A HANDLE AND A RED CIRCLE WITH A LINE GOING THROUGH IT. A NEW HANDLE WAS USED. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND: THE CAUSE OF THE OBSERVED ERROR WAS DUE TO A SOFTWARE RESTRICTIONS ON THE CAPACITY OF THE QUEUE. CAPA2023-019319 WAS INITIATED TO INVESTIGATE THIS FAILURE.
13-12-2024 | Malfunction | us surgical puerto rico | 05-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LIVER RESECTION, WHILE DETACHING INTRAHEPATIC DUCTS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE ISSUE HAPPENED WITH TWO HANDLES AND ONE RELOAD. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | us surgical puerto rico | 05-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LIVER RESECTION, WHILE DETACHING INTRAHEPATIC DUCTS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE ISSUE HAPPENED WITH TWO HANDLES AND ONE RELOAD. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | us surgical puerto rico | 05-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LIVER RESECTION, WHILE DETACHING INTRAHEPATIC DUCTS, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE ISSUE HAPPENED WITH TWO HANDLES AND ONE RELOAD. ANOTHER HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
17-10-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLORECTAL RESECTION AND LYMPHADENECTOMY PROCEDURE, THE STAPLES DID NOT COME OUT AFTER FIRING. THERE WAS NO PATIENT INJURY.
17-10-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLORECTAL RESECTION AND LYMPHADENECTOMY PROCEDURE, THE STAPLES DID NOT COME OUT AFTER FIRING. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE INSTRUMENT EXPERIENCED FIRING ISSUES, AND THE HANDLE REMAINED BLOCKED. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE. THERE WERE NO CONSEQUENCES TO THE PATIENT, AND THE OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
10-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | UNKNOWN EGIA SU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED, DURING A PROCEDURE, THAT THE INSTRUMENT EXPERIENCED FIRING ISSUES, AND THE HANDLE REMAINED BLOCKED. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE. THERE WERE NO CONSEQUENCES TO THE PATIENT, AND THE OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION PROCEDURE, WHEN THE RELOAD WAS ATTACHED TO THE ADAPTER AND WHEN AN ATTEMPT WAS MADE TO FIRE, THE FIRING COULD NOT BE DONE. UPON CHECKING THE POWERED HANDLE INDICATOR, THE RELOAD PART WAS NOT GREEN, BUT WAS CHANGED TO WHITE. IT WAS GREEN WHEN ATTACHED. WHEN AN ANOTHER RELOAD WAS ATTACHED, THERE WERE NO PROBLEMS WITH THE FIRING. THERE WAS NO PATIENT INJURY.
05-12-2024 | Injury | ethicon endo-surgery | 06-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A GYN PROCEDURE TO REMOVE AN OVARIAN MASS, THE MASS HAD GROWN INTO THE COLON. THEY HAD TO REMOVE PART OF THE COLON. THE DISTAL END. THE SURGEON FIRED THE POWERED CIRCULAR, THE LEAK TEST WAS DONE AND PASSED. SIX DAYS POST OP THE STAPLE LINE COMPLETELY BUSTED, IT BURST.
30-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | ILS 33MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-12-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR), THE ARTICULATION KNOB GOT BROKEN AFTER THE FIRST FIRING WHILE USING FOR THE SECOND FIRE. THE STAPLER DID NOT ROTATE WELL AND BROKE DOWN. THE DEVICE WAS NOT YET CLAMPED ONTO THE TISSUE AT THAT TIME. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR), THE ARTICULATION KNOB GOT BROKEN AFTER THE FIRST FIRING WHILE USING FOR THE SECOND FIRE. THE STAPLER DID NOT ROTATE WELL AND BROKE DOWN. THE DEVICE WAS NOT YET CLAMPED ONTO THE TISSUE AT THAT TIME. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-11-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE SHAFT OF THE RELOAD BROKE. NO BROKEN PART DISENGAGED. ANOTHER RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2024 | Injury | ethicon endo-surgery | 06-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC SIGMOID COLECTOMY PROCEDURE THAT WHEN THE LEAK TEST WAS DONE BUBBLES WERE SEEN. A SECOND DEVICE WAS USED AND AGAIN THERE WERE BUBBLES FOUND AFTER THE SECOND LEAK TEST. WITH BOTH FIRINGS THE AUDIBLE "CRUNCH" WAS HEARD AND THE DONUTS WERE INTACT. THERE WERE DIFFICULTIES REMOVING BOTH DEVICES THAT WERE NOT NORMAL. PATIENT SUFFERS FROM DIVERTICULITIS. SURGEON FINISHED THE CASE BY GIVING THE PATIENT A TEMPORARY COLOSTOMY.
30-12-2024 | Injury | ethicon endo-surgery | 06-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC SIGMOID COLECTOMY PROCEDURE THAT WHEN THE LEAK TEST WAS DONE BUBBLES WERE SEEN. A SECOND DEVICE WAS USED AND AGAIN THERE WERE BUBBLES FOUND AFTER THE SECOND LEAK TEST. WITH BOTH FIRINGS THE AUDIBLE "CRUNCH" WAS HEARD AND THE DONUTS WERE INTACT. THERE WERE DIFFICULTIES REMOVING BOTH DEVICES THAT WERE NOT NORMAL. PATIENT SUFFERS FROM DIVERTICULITIS. SURGEON FINISHED THE CASE BY GIVING THE PATIENT A TEMPORARY COLOSTOMY.
30-12-2024 | Injury | ethicon endo-surgery | 06-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC SIGMOID COLECTOMY PROCEDURE THAT WHEN THE LEAK TEST WAS DONE BUBBLES WERE SEEN. A SECOND DEVICE WAS USED AND AGAIN THERE WERE BUBBLES FOUND AFTER THE SECOND LEAK TEST. WITH BOTH FIRINGS THE AUDIBLE "CRUNCH" WAS HEARD AND THE DONUTS WERE INTACT. THERE WERE DIFFICULTIES REMOVING BOTH DEVICES THAT WERE NOT NORMAL. PATIENT SUFFERS FROM DIVERTICULITIS. SURGEON FINISHED THE CASE BY GIVING THE PATIENT A TEMPORARY COLOSTOMY.
30-12-2024 | Injury | ethicon endo-surgery | 06-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC SIGMOID COLECTOMY PROCEDURE THAT WHEN THE LEAK TEST WAS DONE BUBBLES WERE SEEN. A SECOND DEVICE WAS USED AND AGAIN THERE WERE BUBBLES FOUND AFTER THE SECOND LEAK TEST. WITH BOTH FIRINGS THE AUDIBLE "CRUNCH" WAS HEARD AND THE DONUTS WERE INTACT. THERE WERE DIFFICULTIES REMOVING BOTH DEVICES THAT WERE NOT NORMAL. PATIENT SUFFERS FROM DIVERTICULITIS. SURGEON FINISHED THE CASE BY GIVING THE PATIENT A TEMPORARY COLOSTOMY.
02-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN ENTERECTOMY PROCEDURE, THE COMPONENT DISENGAGED. IT WAS ALSO DESCRIBED AS PARTIALLY FIRE SINCE SOME PUSHERS. TO RESOLVE THE ISSUE, IT WAS REPLACED BY ANOTHER PRODUCT FROM THE SAME BATCH. THERE WAS NO PATIENT INJURY.
02-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN ENTERECTOMY PROCEDURE, THE COMPONENT DISENGAGED. IT WAS ALSO DESCRIBED AS PARTIALLY FIRE SINCE SOME PUSHERS. TO RESOLVE THE ISSUE, IT WAS REPLACED BY ANOTHER PRODUCT FROM THE SAME BATCH. THERE WAS NO PATIENT INJURY.
18-11-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN OPEN SURGERY, THE EXPERIENCED NURSE SET UP AND TRIED TO PERFORM FUNCTIONAL END TO END ANASTOMOSIS AS USUAL AT THE 1ST FIRING, BUT THE DEVICE LOCKED OUT, SO THE SURGERY WAS COMPLETED WITH A NEW DEVICE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
13-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR LUNG CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, HOWEVER THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY FOR LUNG CANCER, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, HOWEVER THE DEVICE DID NOT FIRE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2024 | Injury | ethicon endo-surgery | 06-01-2025 | ECHELON LINEAR CUTTER 80MM | Difficult to Insert | A150206 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A SMALL BOWEL RESECTION WITH GLC, A STAPLE LINE LEAK WAS REPORTED A FEW DAYS POST OP. THE DR WHO OPERATED SAID HE DID NOT BELIEVE IT WAS THE STAPLER'S FAULT. HOWEVER, THIS WAS HIS FIRST OPERATION IN 2 YEARS WHERE A LEAK OCCURRED AND THE ONLY CHANGE WAS USING THE NEW STAPLER. THE DEPARTMENT HEAD THINKS THE LEAK OCCURRED DUE TO THE GLC.
01-01-2024 | Injury | ethicon endo-surgery | 06-01-2025 | ECHELON LINEAR CUTTER 80MM | Failure to Fire | A050501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A SMALL BOWEL RESECTION WITH GLC, A STAPLE LINE LEAK WAS REPORTED A FEW DAYS POST OP. THE DR WHO OPERATED SAID HE DID NOT BELIEVE IT WAS THE STAPLER'S FAULT. HOWEVER, THIS WAS HIS FIRST OPERATION IN 2 YEARS WHERE A LEAK OCCURRED AND THE ONLY CHANGE WAS USING THE NEW STAPLER. THE DEPARTMENT HEAD THINKS THE LEAK OCCURRED DUE TO THE GLC.
01-01-2024 | Injury | ethicon endo-surgery | 06-01-2025 | ECHELON LINEAR CUTTER 80MM | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A SMALL BOWEL RESECTION WITH GLC, A STAPLE LINE LEAK WAS REPORTED A FEW DAYS POST OP. THE DR WHO OPERATED SAID HE DID NOT BELIEVE IT WAS THE STAPLER'S FAULT. HOWEVER, THIS WAS HIS FIRST OPERATION IN 2 YEARS WHERE A LEAK OCCURRED AND THE ONLY CHANGE WAS USING THE NEW STAPLER. THE DEPARTMENT HEAD THINKS THE LEAK OCCURRED DUE TO THE GLC.
17-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | PWRD 23MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY WHEN FIRING, THE STAPLES COULDN'T BE DEPLOYED. CHANGED ANOTHER DEVICE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
03-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY WHEN FIRING, THE STAPLES COULDN'T BE DEPLOYED. CHANGED ANOTHER DEVICE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-12-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE RELOAD, THERE WERE FIRING ISSUES, AN OPEN JAW ON THE DEVICE. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW RELOAD.
13-12-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC LOBECTOMY, WHEN BEING APPLIED TO DISCONNECT THE BLOOD VESSEL OF THE LUNG, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE DEFECTIVE HANDLE WITH ANOTHER HANDLE FROM THE COMPANY TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Injury | ethicon endo-surgery | 06-01-2025 | PWRD ECH FLEX 45MM EXACT CODE UNKNOWN | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST OP TO A GASTRIC BYPASS BLEEDING WAS IDENTIFIED POST-OPERATIVELY IN THE STAPLE LINE, EVIDENCED BY EPISODES OF BLOODY VOMITING. THE BLEEDING WAS LOCATED IN THE STAPLING OF THE GASTRIC POUCH, WHERE A BLUE RELOAD WAS USED. DUE TO THIS EVENT, THE PATIENT NEEDED AN ADDITIONAL DAY OF HOSPITALIZATION. HOWEVER, SHE WAS DISCHARGED WITHOUT ANY OTHER COMPLICATIONS, SINCE THE PROBLEM WAS PROPERLY RESOLVED. NO SURGICAL RE-APPROACH OR ADDITIONAL EXAMINATIONS WERE REQUIRED. THE PATIENT DID EXPERIENCE AN ADVERSE EVENT SUCH AS INFECTION, NON-UNION, ALLERGIC REACTION, OSTEOPOROSIS, OVERLOADING, PAIN, DEGENERATIVE DISEASES, BLEEDING OR OOZING.
10-12-2024 | Injury | ethicon endo-surgery | 06-01-2025 | PWRD ECH FLEX 45MM EXACT CODE UNKNOWN | Adverse Event Without Identified Device or Use Problem | A24 | Vomiting | Event Description: IT WAS REPORTED THAT POST OP TO A GASTRIC BYPASS BLEEDING WAS IDENTIFIED POST-OPERATIVELY IN THE STAPLE LINE, EVIDENCED BY EPISODES OF BLOODY VOMITING. THE BLEEDING WAS LOCATED IN THE STAPLING OF THE GASTRIC POUCH, WHERE A BLUE RELOAD WAS USED. DUE TO THIS EVENT, THE PATIENT NEEDED AN ADDITIONAL DAY OF HOSPITALIZATION. HOWEVER, SHE WAS DISCHARGED WITHOUT ANY OTHER COMPLICATIONS, SINCE THE PROBLEM WAS PROPERLY RESOLVED. NO SURGICAL RE-APPROACH OR ADDITIONAL EXAMINATIONS WERE REQUIRED. THE PATIENT DID EXPERIENCE AN ADVERSE EVENT SUCH AS INFECTION, NON-UNION, ALLERGIC REACTION, OSTEOPOROSIS, OVERLOADING, PAIN, DEGENERATIVE DISEASES, BLEEDING OR OOZING.
01-01-2024 | Injury | ethicon endo-surgery | 06-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT POST OP TO A GASTRECTOMY PROCEDURE THE DOCTOR INFORMED US THAT A PATIENT PRESENTED FISTULA 4 DAYS AFTER BARIATRIC PROCEDURE. ACCORDING TO THE DOCTOR, DURING THE PROCEDURE, IN THE LAST STAPLING OF THE STOMACH, HE NOTICED THAT AFTER THE SHOOTING THERE WAS A SMALL BLEEDING, BUT HE DID NOT TELL US WHICH METHOD OF HEMOSTASIS WAS USED TO CONTROL THE BLEEDING. WE DO NOT HAVE ANY ADDITIONAL INFORMATION, SINCE THE DOCTOR HAS COMMITTED TO SENDING US A MEDICAL REPORT ABOUT THE INCIDENT, PATIENT DATA AND ITS CURRENT STATUS.
19-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED AND THEN DURING THE SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED AND THEN DURING THE SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED AND THEN DURING THE SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED AND THEN DURING THE SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
20-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A RIGHT HEMICOLECTOMY SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT HEMICOLECTOMY SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A RIGHT HEMICOLECTOMY SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT HEMICOLECTOMY SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | GST GOLD RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DISTAL GASTRECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | POWERED PLUS STANDARD 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE SURGERY, THE STAPLES FORMED PARTIALLY AFTER THE FIRE. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CLAMP WAS USED FOR THE FIRST TIME WITHOUT ANY PROBLEMS. DURING THE ATTEMPT OF A SECOND USE THE ELECTRIC RELEASE CAUSED THE ADVANCE OF THE STAPLING ON ONLY NOT EVEN A QUARTER OF THE CLAMP THEN EVERYTHING IS BLOCKED COMPLETELY. THE UNBLOCKING WAS POSSIBLE ONLY MECHANICALLY AND WE HAD TO TAKE OUT A SECOND CLAMP TO FINISH THE STAPLING.
17-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CLAMP WAS USED FOR THE FIRST TIME WITHOUT ANY PROBLEMS. DURING THE ATTEMPT OF A SECOND USE THE ELECTRIC RELEASE CAUSED THE ADVANCE OF THE STAPLING ON ONLY NOT EVEN A QUARTER OF THE CLAMP THEN EVERYTHING IS BLOCKED COMPLETELY. THE UNBLOCKING WAS POSSIBLE ONLY MECHANICALLY AND WE HAD TO TAKE OUT A SECOND CLAMP TO FINISH THE STAPLING.
01-12-2024 | Injury | ethicon endo-surgery | 06-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Patient Device Interaction Problem | A01 | Wound Dehiscence | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE GASTRIC POUCH STAPLE LINE OPENED. THE PATIENT IS HOSPITALIZED IN THE ICU.
06-12-2024 | Injury | intuitive surgical | 06-01-2025 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED APPENDECTOMY AND HYSTERECTOMY PROCEDURE, A STAPLE LINE DEFECT WAS IDENTIFIED WHERE A BLUE SUREFORM 30 RELOAD WAS FIRED WITH A SUREFORM 30 STAPLER INSTRUMENT. THE SUREFORM 30 STAPLER INSTRUMENT WAS USED WITH A BLUE SUREFORM 30 RELOAD TO STAPLE THE APPENDICULAR ARTERY. THE FIRING SEQUENCE WAS COMPLETED WITHOUT ANY ERROR MESSAGES OR COMPLICATIONS. THE STAPLE LINE APPEARED COMPLETE BUT SOME BLEEDING OCCURRED FROM THE STAPLE LINE. THE ESTIMATED BLOOD LOSS WAS MINIMAL. COAGULATION WAS USED TO CAUTERIZE THE BLEEDING WHICH RESOLVED THE ISSUE. THE PROCEDURE WAS COMPLETED ROBOTICALLY. THE PATIENT IS RECOVERING WELL.
20-11-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, DURING DETACHMENT OF THE RECTUM, AFTER LOADING THE HANDLE AND THE RELOAD CORRECTLY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE THEN USED, BUT AFTER FIRING, THE STAPLES DID NOT FORM A PROPER B SHAPE, AND SEVERAL STAPLES WERE HANGING ON IN THE RELOAD THAT IT COULD NOT BE REMOVED. A NEW DEVICE FROM A DIFFERENT MANUFACTURER WAS USED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
20-11-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, DURING DETACHMENT OF THE RECTUM, AFTER LOADING THE HANDLE AND THE RELOAD CORRECTLY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, BUT THE HANDLE COULD NOT BE SQUEEZED AND THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE THEN USED, BUT AFTER FIRING, THE STAPLES DID NOT FORM A PROPER B SHAPE, AND SEVERAL STAPLES WERE HANGING ON IN THE RELOAD THAT IT COULD NOT BE REMOVED. A NEW DEVICE FROM A DIFFERENT MANUFACTURER WAS USED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
30-07-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING ROUTINE MAINTENANCE, AN END OF LIFE IMAGE APPEARED ON THE HANDLE'S SCREEN EVENT THOUGH THE DEVICE STILL HAS REMAINING USAGE/LIFE LEFT. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
30-07-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | SIGNIA | Premature End-of-Life Indicator | A160602 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING ROUTINE MAINTENANCE, AN END OF LIFE IMAGE APPEARED ON THE HANDLE'S SCREEN EVENT THOUGH THE DEVICE STILL HAS REMAINING USAGE/LIFE LEFT. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
01-11-2024 | Death | ethicon endo-surgery | 06-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A BARIATRIC SURGERY, THE PATIENT HAD DEHISCENCE OF THE VERTICAL STAPLE LINE OF THE GASTRIC POUCH IN THE SECOND POST-OPERATIVE DAY. HE DIED.
08-11-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE SET-UP ISSUES PRIOR TO FIRING WITH EGIAUSTND ENDOGIA ULTRA UNIV STD STAP. THE HANDLE THAT HAD AN UNKNOWN ISSUE WAS A MANUAL TYPE. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT A SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
08-11-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE SET-UP ISSUES PRIOR TO FIRING WITH EGIAUSTND ENDOGIA ULTRA UNIV STD STAP. THE HANDLE THAT HAD AN UNKNOWN ISSUE WAS A MANUAL TYPE. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT A SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
20-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | GST GREEN RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | GST WHITE RELOAD, 45MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-12-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | PREMIUM PLUS CEEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL RESECTION OF RIGHT HEMICOLON CANCER, WHEN THE DEVICE WAS BEING APPLIED TO THE END-TO-END ANASTOMOSIS THE STAPLE LINE WAS INCOMPLETE. THE ANASTOMOTIC STOMA WAS MANUALLY SUTURED TO RESOLVE THE ISSUE.
10-12-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | PREMIUM PLUS CEEA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL RESECTION OF RIGHT HEMICOLON CANCER, WHEN THE DEVICE WAS BEING APPLIED TO THE END-TO-END ANASTOMOSIS THE STAPLE LINE WAS INCOMPLETE. THE ANASTOMOTIC STOMA WAS MANUALLY SUTURED TO RESOLVE THE ISSUE.
20-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | ECHELON 3000 45MM COMPACT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING AN UNKNOWN RESPIRATORY SURGERY, THE JAWS BECAME NOT TO OPEN AFTER FIRING. THERE WAS NO RESPONSE EVEN THOUGH THE HOME BUTTON ETC. WERE TOUCHED. THE FIRING TRIGGER WAS GRASPED AND THE KNIFE MOVED ALL THE WAY TO THE TIP WITHOUT PROBLEM, BUT THE KNIFE RETURNED WITH STOPPING WHEN RETURNING. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
20-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | ECHELON 3000 45MM COMPACT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING AN UNKNOWN RESPIRATORY SURGERY, THE JAWS BECAME NOT TO OPEN AFTER FIRING. THERE WAS NO RESPONSE EVEN THOUGH THE HOME BUTTON ETC. WERE TOUCHED. THE FIRING TRIGGER WAS GRASPED AND THE KNIFE MOVED ALL THE WAY TO THE TIP WITHOUT PROBLEM, BUT THE KNIFE RETURNED WITH STOPPING WHEN RETURNING. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
11-12-2024 | Injury | intuitive surgical | 06-01-2025 | SUREFORM | Material Fragmentation | A040103 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SURGICAL PROCEDURE, A PART OF A BLUE SUREFORM 45 RELOAD HAD BROKEN OFF DURING THE FIRING PROCESS WITH A SUREFORM 45 STAPLER INSTRUMENT. ALL PIECES WERE RECOVERED DURING THE SAME PROCEDURE WHICH WAS COMPLETED ROBOTICALLY.
06-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AFTER FIRING STAPLER, THE JAW DIDN'T OPEN. THERE WAS NO SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
16-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP SLEEVE GASTRECTOMY, THE DISTAL RETAINING SUTURE ON THE ANVIL SIDE OF THE SIGTRSB60AXT 60MM XTR THK REINFORCED REL DID NOT RELEASE AT THE END OF THE FIRING, CAUSING TRS MATERIAL TO REMAIN ATTACHED TO THE RELOAD. THIS ISSUE OCCURRED DURING THE STAPLING OF THE STOMACH. AN ENDOSHEARS WAS USED TO MANUALLY CUT THE TRS STRIP OFF OF THE RELOAD TO REMOVE IT FROM THE TISSUE. IT WAS NOTED THAT IT OCCURRED ON THREE RELOADS. THE SIGNIA STAPLER WAS FULLY FIRED AND DOUBLE-CHECKED WITH EACH FIRING BY RELEASING AND THEN PRESSING DOWN AGAIN ON THE FIRING BUTTON WITH EACH RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | TRI-STAPLE 2.0 | Separation Failure | A150301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP SLEEVE GASTRECTOMY, THE DISTAL RETAINING SUTURE ON THE ANVIL SIDE OF THE SIGTRSB60AXT 60MM XTR THK REINFORCED REL DID NOT RELEASE AT THE END OF THE FIRING, CAUSING TRS MATERIAL TO REMAIN ATTACHED TO THE RELOAD. THIS ISSUE OCCURRED DURING THE STAPLING OF THE STOMACH. AN ENDOSHEARS WAS USED TO MANUALLY CUT THE TRS STRIP OFF OF THE RELOAD TO REMOVE IT FROM THE TISSUE. IT WAS NOTED THAT IT OCCURRED ON THREE RELOADS. THE SIGNIA STAPLER WAS FULLY FIRED AND DOUBLE-CHECKED WITH EACH FIRING BY RELEASING AND THEN PRESSING DOWN AGAIN ON THE FIRING BUTTON WITH EACH RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
12-11-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER TRIGGERING THE TRIGGER, THE DEVICE PARTIALLY FIRED THE LOAD AND LOCKED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP SLEEVE GASTRECTOMY, THE DISTAL RETAINING SUTURE ON THE ANVIL SIDE OF THE SIGTRSB60AXT 60MM XTR THK REINFORCED REL DID NOT RELEASE AT THE END OF THE FIRING, CAUSING TRS MATERIAL TO REMAIN ATTACHED TO THE RELOAD. THIS ISSUE OCCURRED DURING THE STAPLING OF THE STOMACH. AN ENDOSHEARS WAS USED TO MANUALLY CUT THE TRS STRIP OFF OF THE RELOAD TO REMOVE IT FROM THE TISSUE. IT WAS NOTED THAT IT OCCURRED ON THREE RELOADS. THE SIGNIA STAPLER WAS FULLY FIRED AND DOUBLE-CHECKED WITH EACH FIRING BY RELEASING AND THEN PRESSING DOWN AGAIN ON THE FIRING BUTTON WITH EACH RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | TRI-STAPLE 2.0 | Separation Failure | A150301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP SLEEVE GASTRECTOMY, THE DISTAL RETAINING SUTURE ON THE ANVIL SIDE OF THE SIGTRSB60AXT 60MM XTR THK REINFORCED REL DID NOT RELEASE AT THE END OF THE FIRING, CAUSING TRS MATERIAL TO REMAIN ATTACHED TO THE RELOAD. THIS ISSUE OCCURRED DURING THE STAPLING OF THE STOMACH. AN ENDOSHEARS WAS USED TO MANUALLY CUT THE TRS STRIP OFF OF THE RELOAD TO REMOVE IT FROM THE TISSUE. IT WAS NOTED THAT IT OCCURRED ON THREE RELOADS. THE SIGNIA STAPLER WAS FULLY FIRED AND DOUBLE-CHECKED WITH EACH FIRING BY RELEASING AND THEN PRESSING DOWN AGAIN ON THE FIRING BUTTON WITH EACH RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
20-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | GST WHITE RELOAD, 60MM, 6-ROW | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, BEFORE USED ON THE PATIENT, NOTED THE PACKING WAS DAMAGED. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
12-11-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER TRIGGERING THE TRIGGER, THE DEVICE PARTIALLY FIRED THE LOAD AND LOCKED. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP SLEEVE GASTRECTOMY, THE DISTAL RETAINING SUTURE ON THE ANVIL SIDE OF THE SIGTRSB60AXT 60MM XTR THK REINFORCED REL DID NOT RELEASE AT THE END OF THE FIRING, CAUSING TRS MATERIAL TO REMAIN ATTACHED TO THE RELOAD. THIS ISSUE OCCURRED DURING THE STAPLING OF THE STOMACH. AN ENDOSHEARS WAS USED TO MANUALLY CUT THE TRS STRIP OFF OF THE RELOAD TO REMOVE IT FROM THE TISSUE. IT WAS NOTED THAT IT OCCURRED ON THREE RELOADS. THE SIGNIA STAPLER WAS FULLY FIRED AND DOUBLE-CHECKED WITH EACH FIRING BY RELEASING AND THEN PRESSING DOWN AGAIN ON THE FIRING BUTTON WITH EACH RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
16-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | TRI-STAPLE 2.0 | Separation Failure | A150301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP SLEEVE GASTRECTOMY, THE DISTAL RETAINING SUTURE ON THE ANVIL SIDE OF THE SIGTRSB60AXT 60MM XTR THK REINFORCED REL DID NOT RELEASE AT THE END OF THE FIRING, CAUSING TRS MATERIAL TO REMAIN ATTACHED TO THE RELOAD. THIS ISSUE OCCURRED DURING THE STAPLING OF THE STOMACH. AN ENDOSHEARS WAS USED TO MANUALLY CUT THE TRS STRIP OFF OF THE RELOAD TO REMOVE IT FROM THE TISSUE. IT WAS NOTED THAT IT OCCURRED ON THREE RELOADS. THE SIGNIA STAPLER WAS FULLY FIRED AND DOUBLE-CHECKED WITH EACH FIRING BY RELEASING AND THEN PRESSING DOWN AGAIN ON THE FIRING BUTTON WITH EACH RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
19-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | ECHELON 3000 45MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, WHEN TRIED TO FIRE AT THE 1ST FIRING, THE KNIFE STOPPED ON THE WAY. THE DEVICE WAS CLAMPED THE LUNG, BUT THE KNIFE COULD BE RETURNED WITHOUT ANY PROBLEM. A NEW DEVICE WAS USED FROM THE 2ND FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
14-11-2024 | Injury | intuitive surgical | 06-01-2025 | SUREFORM | Firing Problem | A0505 | Pleural Empyema | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RIGHT LOWER PULMONARY LOBECTOMY PROCEDURE, THE PATIENT RETURNED TO THE HOSPITAL WITH PLEURAL EMPYEMA, NECESSITATING 2 ADDITIONAL PROCEDURES. THE INITIAL ROBOTIC PROCEDURE HAD NO COMPLICATIONS, AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. HOWEVER, ON POST-OPERATIVE DAY 22, THE PATIENT PRESENTED WITH A SIGNIFICANT FOUL ODOR FROM THE MOUTH AND A GENERAL DETERIORATION IN THEIR OVERALL CONDITION. A RIGHT THORACOTOMY REVEALED THAT THE BRONCHUS, WHICH HAD BEEN STAPLED USING A SUREFORM 45 CURVED-TIP STAPLE INSTRUMENT LOADED WITH A BLACK SUREFORM 45 RELOAD ACCESSORY, WAS WIDE OPEN. THE REMAINING BRONCHUS WAS INFLAMED AND ULCERATED, WITH ONLY A FEW CLOSED STAPLES OBSERVED IN THE THORAX. TO ADDRESS THESE ISSUES, AN EMPYEMECTOMY AND A REANASTOMOSIS OF THE MIDDLE LOBE TO THE INTERMEDIATE BRONCHUS WERE PERFORMED. ON POST-OPERATIVE DAY 3, A FLEXIBLE BRONCHOSCOPY WAS CONDUCTED, AND A 3CM LONG, 10MM COVERED STENT WAS IMPLANTED IN THE INTERMEDIATE BRONCHUS AND THE MIDDLE LOBE BRONCHUS FOR ANASTOMOSIS STENTING AND SEALING. SUBSEQUENTLY, A PERFORATION OF THE STENT OCCURRED DUE TO COVERED INSUFFICIENCY IN THE AREA OF THE INTERMEDIATE BRONCHUS, NECESSITATING AN ADDITIONAL PROCEDURE. ON POST-OPERATIVE DAY 5, A RIGHT RE-THORACOTOMY AND PLEUROLYSIS WERE PERFORMED, DURING WHICH THE RIGHT MIDDLE LOBE WAS RESECTED, AND THE EXISTING LATISSIMUS DORSI MUSCLE WAS USED TO COVER THE BRONCHIAL STUMP. SINCE THE COMPLETION OF THIS LAST PROCEDURE, THERE HAVE BEEN NO FURTHER COMPLICATIONS, AND THE PATIENT IS CURRENTLY STABLE.
14-11-2024 | Injury | intuitive surgical | 06-01-2025 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RIGHT LOWER PULMONARY LOBECTOMY PROCEDURE, THE PATIENT RETURNED TO THE HOSPITAL WITH PLEURAL EMPYEMA, NECESSITATING 2 ADDITIONAL PROCEDURES. THE INITIAL ROBOTIC PROCEDURE HAD NO COMPLICATIONS, AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. HOWEVER, ON POST-OPERATIVE DAY 22, THE PATIENT PRESENTED WITH A SIGNIFICANT FOUL ODOR FROM THE MOUTH AND A GENERAL DETERIORATION IN THEIR OVERALL CONDITION. A RIGHT THORACOTOMY REVEALED THAT THE BRONCHUS, WHICH HAD BEEN STAPLED USING A SUREFORM 45 CURVED-TIP STAPLE INSTRUMENT LOADED WITH A BLACK SUREFORM 45 RELOAD ACCESSORY, WAS WIDE OPEN. THE REMAINING BRONCHUS WAS INFLAMED AND ULCERATED, WITH ONLY A FEW CLOSED STAPLES OBSERVED IN THE THORAX. TO ADDRESS THESE ISSUES, AN EMPYEMECTOMY AND A REANASTOMOSIS OF THE MIDDLE LOBE TO THE INTERMEDIATE BRONCHUS WERE PERFORMED. ON POST-OPERATIVE DAY 3, A FLEXIBLE BRONCHOSCOPY WAS CONDUCTED, AND A 3CM LONG, 10MM COVERED STENT WAS IMPLANTED IN THE INTERMEDIATE BRONCHUS AND THE MIDDLE LOBE BRONCHUS FOR ANASTOMOSIS STENTING AND SEALING. SUBSEQUENTLY, A PERFORATION OF THE STENT OCCURRED DUE TO COVERED INSUFFICIENCY IN THE AREA OF THE INTERMEDIATE BRONCHUS, NECESSITATING AN ADDITIONAL PROCEDURE. ON POST-OPERATIVE DAY 5, A RIGHT RE-THORACOTOMY AND PLEUROLYSIS WERE PERFORMED, DURING WHICH THE RIGHT MIDDLE LOBE WAS RESECTED, AND THE EXISTING LATISSIMUS DORSI MUSCLE WAS USED TO COVER THE BRONCHIAL STUMP. SINCE THE COMPLETION OF THIS LAST PROCEDURE, THERE HAVE BEEN NO FURTHER COMPLICATIONS, AND THE PATIENT IS CURRENTLY STABLE.
16-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE RELOAD DURING SET-UP AND PRIOR TO FIRING. THE SPECIFIC PROBLEMS INCLUDED JAW ISSUES, THE JAWS WAS NOT CLOSING COMPLETELY. THE DEVICE WAS NOT IMPLANTABLE AND WAS ASSOCIATED WITH A POWERED STAPLER USING A STANDARD ADAPTER. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED. THERE WAS NO PATIENT INVOLVEMENT.
10-12-2024 | Malfunction | ethicon endo-surgery | 06-01-2025 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, CLOSE THE JAWS AND NOTED THE STAPLES CAME OUT AND THE KNIFE MOVED AUTOMATICALLY. . CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE RELOAD DURING SET-UP AND PRIOR TO FIRING. THE SPECIFIC PROBLEMS INCLUDED JAW ISSUES, THE JAWS WAS NOT CLOSING COMPLETELY. THE DEVICE WAS NOT IMPLANTABLE AND WAS ASSOCIATED WITH A POWERED STAPLER USING A STANDARD ADAPTER. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED. THERE WAS NO PATIENT INVOLVEMENT.
05-12-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE RELOAD DURING SET-UP AND PRIOR TO FIRING. THE SPECIFIC PROBLEMS INCLUDED JAW ISSUES, THE JAWS WAS NOT CLOSING COMPLETELY. THE DEVICE WAS NOT IMPLANTABLE AND WAS ASSOCIATED WITH A POWERED STAPLER USING A STANDARD ADAPTER. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED. THERE WAS NO PATIENT INVOLVEMENT.
03-01-2025 | Malfunction | covidien lp llc north haven | 06-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LEFT HEPATECTOMY, WHEN CUTTING OFF THE LEFT HEPATIC PEDICLE, THERE WAS POOR STAPLE FORMATION. THE STAPLES PARTLY FORMED. REINFORCEMENT WITH SUTURE WAS DONE TO RESOLVE THE ISSUE.
23-12-2021 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THE SURGEON WAS ABLE PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. ANOTHER HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS AN EXTENDED SURGICAL TIME. THERE WAS NO PATIENT INJURY.
06-12-2024 | Injury | intuitive surgical | 06-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Sepsis | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED SIGMOID COLECTOMY SURGICAL PROCEDURE, THE PATIENT DEVELOPED SEPSIS RESULTING IN ADDITIONAL PROCEDURES. THE INITIAL PROCEDURE WAS COMPLETED ROBOTICALLY. ON POST-OP DAY 3, A DIAGNOSTIC LAPAROSCOPY WITH ABDOMINAL WASHOUT RESULTED IN OVERSEWING OF THE STAPLE LINE AT THE POINT OF INITIAL BOWEL TRANSECTION DUE TO STAPLE LINE BREAKDOWN. THE PATIENT WAS LEFT WITH A LOOP ILEOSTOMY. THE PATIENT CONTINUED THE SEPSIS PATHWAY AFTER THE SECOND PROCEDURE. A THIRD PROCEDURE, WHERE THE ENTIRE ANASTOMOSIS WAS COMPLETELY RESECTED, RESULTED IN A LONG HARTMANN¿S POUCH, ILEOSTOMY, AND MULTIPLE DRAINS. THE PATIENT WAS SENT HOME WITH A PERIPHERALLY INSERTED CENTRAL CATHETER (PICC) LINE. THE PATIENT HAD A MEDICAL HISTORY OF CROHN¿S DISEASE, BUT PATHOLOGY FROM BOTH PROCEDURES SHOWED NO PATHOLOGY. THE PATIENT WAS ALSO ON IMMUNOSUPPRESSANTS AT THE TIME OF THE INITIAL PROCEDURE. THE SURGEON STATED THIS DISEASE PROCESS COULD HAVE RESULTED IN FRIABLE TISSUE.
12-11-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A NEPHRECTOMY PROCEDURE INVOLVING THE STAPLER, THERE WERE FIRING ISSUES WITH THE HANDLE. THE SURGEON WAS UNABLE TO SQUEEZE THE MANUAL HANDLE, ALTHOUGH THE GREEN SAFETY BUTTON WAS PRESSED SUCCESSFULLY. THE INSTRUMENT DID NOT FIRE. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW HANDLE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED. STAPLE PUSHERS WERE VISIBLE AT THE 4CM CUT LINE.
26-12-2024 | Malfunction | ethicon endo- surgery | 06-01-2025 | ETHICON ENDOPATH ETS -FLEX 45 ARTICULATING ENDOSCOPIC LINEAR CUTTER 45 MM 340 MM |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: KNIFE FUNCTION OF ENDOPATH ETS-FLEX45 ARTICULATING ENDOSCOPIC LINEAR CUTTER DID NOT FIRE. EQUIPMENT/DEVICE REMOVED FROM SERVICE EQUIPMENT/DEVICE STOPPED.
05-12-2024 | Injury | intuitive surgical | 06-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Sepsis | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED SIGMOID COLECTOMY SURGICAL PROCEDURE, THE PATIENT DEVELOPED SEPSIS RESULTING IN AN ADDITIONAL PROCEDURE. THE INITIAL PROCEDURE WAS COMPLETED ROBOTICALLY. ON POST-OP DAY 3, A DIAGNOSTIC LAPAROSCOPY WITH ABDOMINAL WASHOUT RESULTED IN OVERSEWING OF THE STAPLE LINE. A SMALL HOLE IN THE MIDDLE OF THE STAPLE LINE WAS DISCOVERED WITH MISSING STAPLES. THE PATIENT WAS LEFT WITH A LOOP ILEOSTOMY AND TRANSFERRED TO A REHABILITATION FACILITY. THE PATIENT WAS NOTED TO HAVE ¿SPECIAL NEEDS,¿ AND THE REHAB FACILITY WAS TO HELP WITH OSTOMY CARE.
23-11-2024 | Malfunction | covidien lp llc north haven | 06-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE OPEN STAPLER, THERE WERE FIRING ISSUES WHERE THE STAPLES DID NOT DEPLOY, RESULTING IN CUTTING WITHOUT STAPLING. THE PRESENCE OF STAPLES WAS CONFIRMED. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED WITH ANOTHER PRODUCT.
13-12-2024 | Malfunction | us surgical puerto rico | 06-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACIC SURGERY (VATS) LOBECTOMY, THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE, ALTHOUGH THEY WERE ABLE TO PRESS THE GREEN BUTTON. THE STAPLER WAS NOT POWERED, AND NO REINFORCEMENT MATERIAL WAS USED IN CONJUNCTION WITH THE DEVICE. THERE WERE NO ISSUES WITH LOADING OR UNLOADING THE DEVICE. THE HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Injury | intuitive surgical | 06-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED APPENDECTOMY PROCEDURE, AN INCOMPLETE STAPLE LINE WAS IDENTIFIED AFTER A BLUE SUREFORM 30 RELOAD WAS FIRED WITH A SUREFORM 30 STAPLER INSTRUMENT. THE MIDDLE SEGMENT OF THE STAPLE LINE WAS MISSING STAPLES. THE SURGEON OVERSEWED THE HOLE IN THE STAPLE LINE WITH V-LOC SUTURE. NO ERRORS WERE DISPLAYED BY THE SYSTEM DURING THE EVENT AND THERE WERE NO IRREGULARITIES WITH THE TARGET TISSUE NOTED. THIS EVENT CAUSED LESS THAN A 15 MINUTE DELAY IN THE CASE WHICH WAS COMPLETED ROBOTICALLY.
11-12-2024 | Malfunction | intuitive surgical | 07-01-2025 | SUREFORM | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE SUREFORM 60 STAPLER BLUE RELOAD FIRED HALFWAY WHEN A TISSUE TOO THICK MESSAGE APPEARED. THE CUSTOMER NOTICED SOMETHING IN THE JAWS AND USED ANOTHER STAPLER. IT WAS FOUND THAT THE RELOAD HAS A PLASTIC PIECE IN THE CHANNEL. THE USER COMPLETED THE PROCEDURE USING A BACKUP SUREFORM 60 STAPLER BLUE RELOADWITH NO FURTHER ISSUE REPORTED. NO FRAGMENT FELL INSIDE THE PATIENT. NO KNOWN IMPACT OR PATIENT CONSEQUENCE WAS REPORTED.
21-10-2024 | Malfunction | intuitive surgical | 07-01-2025 | SUREFORM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LOW ANTERIOR RESECTION SURGICAL PROCEDURE, THE SUREFORM 45 STAPLER DID NOT OPEN. THE USER COMPLETED THE PROCEDURE WITH NO FURTHER ISSUE REPORTED. NO FRAGMENT FELL INSIDE THE PATIENT. NO KNOWN IMPACT OR PATIENT CONSEQUENCE WAS REPORTED.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Failure to Cycle | A050703 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Failure to Cycle | A050703 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Misfire | A050502 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Failure to Cut | A050702 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END; WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED, THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Failure to Cycle | A050703 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END, WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED; THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END, WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED; THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END, WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED; THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END, WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED; THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END, WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED; THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Material Deformation | A0406 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END, WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED; THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Firing Problem | A0505 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS WEDGE AND SEGMENTECTOMY PROCEDURE ON THE LUNG, THERE WERE MULTIPLE PROCEDURAL AND DEVICE-RELATED ISSUES INVOLVING THE MANUAL HANDLE AND RELOAD. ON THE FIST HANDLE AND RELOAD, THE SURGEON COULD NOT COMPLETE FIRING NOR RETRACT BLACK KNOB TO OPEN THE JAWS. ADDITIONALLY, TISSUE AND TRS MATERIAL WADDED UP AT THE HALFWAY CUT LINE, AND THE PROXIMAL SUTURE OF THE TRS DID NOT RELEASE, REQUIRING MANUAL CUTTING. THE STAPLER HAD TO BE CUT OUT OF THE TISSUE. ON THE SECOND RELOAD AND HANDLE CARTRIDGE WAS USED AND WAS EXTREMELY DIFFICULT TO FIRE BUT DID MAKE IT TO THE END, WHEN SURGEON RETRACTED BLACK KNOB, THE PROXIMAL SUTURE OF THE TRS HAD NOT RELEASED AND HAD TO BE MANUAL CUT TO RELEASE. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. A SECOND AND THIRD HANDLE WERE OPENED; THE SECOND HANDLE WAS RETAINED AND RETURNED, WHILE THE FIRST AND THIRD HANDLES WERE NOT RETAINED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES. TO RESOLVE THESE ISSUES, THE DEVICE WAS REPLACED. THE THIRD HANDLE WAS USED SUCCESSFULLY WITH EIGHT SUBSEQUENT CARTRIDGES.
09-08-2024 | Death | us surgical puerto rico | 07-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS COMPLETED INVOLVING 100 CASES HAVING PREOPERATIVE ARTERIAL CALCIFICATIONS THAT UNDERWENT MINIMALLY INVASIVE IVOR LEWIS ESOPHAGECTOMIES FROM 2013 TO 2018. AN UNKNOWN LAPAROSCOPIC LINEAR STAPLER WAS USED TO CREATE A GASTRIC CONDUIT. THE 25 MM DST XL CIRCULAR STAPLER WAS USED TO CREATE END-SIDE ESOPHAGOGASTRIC ANASTOMOSIS. POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAK, GRAFT/CONDUIT NECROSIS, AND CONDUIT OBSTRUCTION. REOPERATION WAS REQUIRED FOR CONDUIT OBSTRUCTION/NECROSIS. NO REOPERATION WAS REPORTED FOR ANASTOMOTIC LEAK. READMISSIONS WERE ALSO REPORTED. OTHER COMPLICATIONS UNRELATED TO THE DEVICE INCLUDE ATRIAL FIBRILLATION, RESPIRATORY FAILURE, PLEURAL EFFUSIONS, COLONIC PSEUDO-OBSTRUCTION, CONDUIT VOLVULUS, CONDUIT REDUNDANCY, CECAL VOLVULUS, TOXIC MEGACOLON, CLOSED-LOOP OBSTRUCTION, WOUND INFECTION, AND LUNG DECORTICATION. DEATH OCCURRED IN FOUR CASES DUE TO UNSPECIFIED REASONS. TITLE: MAJOR CORONARY ARTERY CALCIFICATIONS AS PREDICTORS OF POSTOPERATIVE COMPLICATIONS IN IVOR LEWIS ESOPHAGECTOMIES: A FVE-YEAR RETROSPECTIVE ANALYSIS. KEOUNA PATHER1 · HAYTHAM ALABBAS1,3 · CARLOS GONZALEZ-BAERGA2 · MANUEL MENENDEZ2 · MAYUR K. VIRARKAR2 · IRAI SANTANA DE OLIVEIRA2 · ERIN M. MOBLEY1 · ZIAD T. AWAD1 SURGICAL ENDOSCOPY (2024) 38:6865¿6872 HTTPS://DOI.ORG/10.1007/S00464-024-11181-3.
09-08-2024 | Death | us surgical puerto rico | 07-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Necrosis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS COMPLETED INVOLVING 100 CASES HAVING PREOPERATIVE ARTERIAL CALCIFICATIONS THAT UNDERWENT MINIMALLY INVASIVE IVOR LEWIS ESOPHAGECTOMIES FROM 2013 TO 2018. AN UNKNOWN LAPAROSCOPIC LINEAR STAPLER WAS USED TO CREATE A GASTRIC CONDUIT. THE 25 MM DST XL CIRCULAR STAPLER WAS USED TO CREATE END-SIDE ESOPHAGOGASTRIC ANASTOMOSIS. POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAK, GRAFT/CONDUIT NECROSIS, AND CONDUIT OBSTRUCTION. REOPERATION WAS REQUIRED FOR CONDUIT OBSTRUCTION/NECROSIS. NO REOPERATION WAS REPORTED FOR ANASTOMOTIC LEAK. READMISSIONS WERE ALSO REPORTED. OTHER COMPLICATIONS UNRELATED TO THE DEVICE INCLUDE ATRIAL FIBRILLATION, RESPIRATORY FAILURE, PLEURAL EFFUSIONS, COLONIC PSEUDO-OBSTRUCTION, CONDUIT VOLVULUS, CONDUIT REDUNDANCY, CECAL VOLVULUS, TOXIC MEGACOLON, CLOSED-LOOP OBSTRUCTION, WOUND INFECTION, AND LUNG DECORTICATION. DEATH OCCURRED IN FOUR CASES DUE TO UNSPECIFIED REASONS. TITLE: MAJOR CORONARY ARTERY CALCIFICATIONS AS PREDICTORS OF POSTOPERATIVE COMPLICATIONS IN IVOR LEWIS ESOPHAGECTOMIES: A FVE-YEAR RETROSPECTIVE ANALYSIS. KEOUNA PATHER1 · HAYTHAM ALABBAS1,3 · CARLOS GONZALEZ-BAERGA2 · MANUEL MENENDEZ2 · MAYUR K. VIRARKAR2 · IRAI SANTANA DE OLIVEIRA2 · ERIN M. MOBLEY1 · ZIAD T. AWAD1 SURGICAL ENDOSCOPY (2024) 38:6865¿6872 HTTPS://DOI.ORG/10.1007/S00464-024-11181-3.
09-08-2024 | Death | us surgical puerto rico | 07-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Obstruction/Occlusion | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS COMPLETED INVOLVING 100 CASES HAVING PREOPERATIVE ARTERIAL CALCIFICATIONS THAT UNDERWENT MINIMALLY INVASIVE IVOR LEWIS ESOPHAGECTOMIES FROM 2013 TO 2018. AN UNKNOWN LAPAROSCOPIC LINEAR STAPLER WAS USED TO CREATE A GASTRIC CONDUIT. THE 25 MM DST XL CIRCULAR STAPLER WAS USED TO CREATE END-SIDE ESOPHAGOGASTRIC ANASTOMOSIS. POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAK, GRAFT/CONDUIT NECROSIS, AND CONDUIT OBSTRUCTION. REOPERATION WAS REQUIRED FOR CONDUIT OBSTRUCTION/NECROSIS. NO REOPERATION WAS REPORTED FOR ANASTOMOTIC LEAK. READMISSIONS WERE ALSO REPORTED. OTHER COMPLICATIONS UNRELATED TO THE DEVICE INCLUDE ATRIAL FIBRILLATION, RESPIRATORY FAILURE, PLEURAL EFFUSIONS, COLONIC PSEUDO-OBSTRUCTION, CONDUIT VOLVULUS, CONDUIT REDUNDANCY, CECAL VOLVULUS, TOXIC MEGACOLON, CLOSED-LOOP OBSTRUCTION, WOUND INFECTION, AND LUNG DECORTICATION. DEATH OCCURRED IN FOUR CASES DUE TO UNSPECIFIED REASONS. TITLE: MAJOR CORONARY ARTERY CALCIFICATIONS AS PREDICTORS OF POSTOPERATIVE COMPLICATIONS IN IVOR LEWIS ESOPHAGECTOMIES: A FVE-YEAR RETROSPECTIVE ANALYSIS. KEOUNA PATHER1 · HAYTHAM ALABBAS1,3 · CARLOS GONZALEZ-BAERGA2 · MANUEL MENENDEZ2 · MAYUR K. VIRARKAR2 · IRAI SANTANA DE OLIVEIRA2 · ERIN M. MOBLEY1 · ZIAD T. AWAD1 SURGICAL ENDOSCOPY (2024) 38:6865¿6872 HTTPS://DOI.ORG/10.1007/S00464-024-11181-3.
10-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE ROBOT ASSISTED LIVER RESECTION, THERE WAS A DAMAGED PART FROM THE RELOAD CONNECTION BEFORE FIRING. THE PRODUCT WAS REPLACED USING THE SAME PRODUCT FROM ANOTHER LOT. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING, THERE WAS A LEAK FROM THE STAPLE LINE; IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED, AND DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING, THERE WAS A LEAK FROM THE STAPLE LINE; IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED, AND DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING, THERE WAS A LEAK FROM THE STAPLE LINE; IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED, AND DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING, THERE WAS A LEAK FROM THE STAPLE LINE; IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED, AND DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
12-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC LOBECTOMY (UPPER LOBE, LOWER LOBE), AFTER THE FIRING THERE WAS LEAK FROM THE STAPLE LINE, IT WAS A MINOR ANASTOMOTIC LEAK. CONFIRMATION USING AIR BUBBLES/METHYLENE BLUE. THE LEAK SITE WAS SUTURED DURING THE SUTURING, THERE WERE MOMENTS WHEN THE LUNG WAS AVOIDED.
25-11-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GOLD RELOAD FOR ECHELON 60 | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION AND AFTER OPENING THE PACKING, THE SURGEON NOTED THAT IT WAS VERY DIFFICULT TO INSTALL THE CARTRIDGE. CHANGED ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE DEVICE WASN'T FIRING ALL THE WAY. UNKNOWN IF THE DEVICE DEPLOYED ANY STAPLES OR IF THE KNIFE CUT. ANOTHER LIKE DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
02-01-2025 | Malfunction | intuitive surgical | 07-01-2025 | SUREFORM |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: ROBOT STAPLER WOULD NOT WORK. KEPT SAYING RELOAD STAPLE. REF# (B)(4).
16-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC LOBECTOMY PROCEDURE, THE DEVICE WAS BEING APPLIED TO TRANSECT THE LOBAR BRONCHUS WHEN THE STAPLER COULD NOT BE FIRED, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. TO RESOLVE THE PROBLEM AND COMPLETE THE PROCEDURE, A NEW RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RADICAL RESECTION OF CARDIAC CANCER, WHILE ON DETACHMENT OF THE DUODENUM, THE HANDLE COMBINED WITH THE RELOAD. HOWEVER, COULD NOT BE ACTIVATED. AFTER REPLACING THE RELOAD, THE SAME ERROR OCCURRED. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS NOT ABLE TO SQUEEZE THE HANDLE. A NEW HANDLE WAS REPLACED TO COMPLETE THE ACTIVATION. DURING THE ANASTOMOSIS OF THE ESOPHAGUS AND JEJUNUM, THE STAPLE DID NOT FORM A GOOD B-SHAPE AFTER CIRCULAR STAPLER ACTIVATION AND BLEEDING OCCURRED. AT THE SAME TIME, THERE WAS NO COMPLETE CIRCULAR SPECIMEN TISSUE AFTER THE STAPLE ANVIL WAS REMOVED. IT WAS SUTURED MANUALLY USING CURB STITCHING TO RESOLVE THE ISSUE.
19-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | ABSORBATACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HERNIORRHAPHY, THE DEVICE BROKE WHILE FIXING THE MESH. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE, ALLOWING THE PROCEDURE TO CONTINUE WITHOUT ANY INJURY OR EXTENDED SURGICAL TIME.
19-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOBECTOMY, WHEN THE INTERLOBULAR FISSURE WAS DETACHED, THE STAPLE DID NOT FORM PROPERLY. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Wound Dehiscence | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Firing Problem | A0505 | Wound Dehiscence | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Firing Problem | A0505 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
10-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RESECTION ILEECTOMY IN A ROBOT-AIDED TOTAL CYSTECTOMY (ILEUM CONDUIT), FOR THE SECOND FIRING IN CREATING AN ILEAL CONDUIT, THE PRODUCT WAS CONNECTED TO THE STAPLING DEVICE, AND DURING THE OPENING AND CLOSING CHECK, IT WAS FOUND THAT THE JAWS COULD NOT OPEN OR CLOSE. AS A RESULT, ANOTHER PRODUCT WAS OPENED. THE THREE FOLLOWING FIRINGS WERE DONE WITHOUT ANY ISSUES AND PERFORMED AS USUAL. THERE WAS NO PATIENT INJURY.
30-09-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON 3000 60MM STANDARD | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS SURGERY AT THE INFANTA SOFÍA HOSPITAL, THE HEAD OF GENERAL SURGERY USED THE ECHELON 3000 DEVICE FOR THE FIRST TIME. THE DEVICE WAS INTRODUCED THROUGH THE TROCAR INTO THE PATIENT, ARTICULATED CORRECTLY, AND FIRED THE FIRST SHOT WITHOUT ISSUES. HOWEVER, WHEN ATTEMPTING TO RETURN THE DEVICE TO ITS ORIGINAL POSITION FOR REMOVAL FROM THE TROCAR, IT WAS NOT POSSIBLE. NONE OF THE USUAL SOLUTIONS WORKED: NEITHER THE ARTICULATION BUTTON, NOR THE RETURN-TO-HOME BUTTON, NOR REMOVING AND REINSERTING THE BATTERY, WHICH WAS ATTEMPTED SEVERAL TIMES. THE DEVICE COMPLETELY CEASED TO FUNCTION, AND IT HAD TO BE REMOVED FROM THE PATIENT¿S BODY WITH THE TROCAR STILL IN PLACE, AS IT COULD NOT BE EXTRACTED THROUGH THE TROCAR IN THE NORMAL MANNER. THE SURGERY WAS DELAYED 5 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
30-09-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON 3000 60MM STANDARD | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS SURGERY AT THE INFANTA SOFÍA HOSPITAL, THE HEAD OF GENERAL SURGERY USED THE ECHELON 3000 DEVICE FOR THE FIRST TIME. THE DEVICE WAS INTRODUCED THROUGH THE TROCAR INTO THE PATIENT, ARTICULATED CORRECTLY, AND FIRED THE FIRST SHOT WITHOUT ISSUES. HOWEVER, WHEN ATTEMPTING TO RETURN THE DEVICE TO ITS ORIGINAL POSITION FOR REMOVAL FROM THE TROCAR, IT WAS NOT POSSIBLE. NONE OF THE USUAL SOLUTIONS WORKED: NEITHER THE ARTICULATION BUTTON, NOR THE RETURN-TO-HOME BUTTON, NOR REMOVING AND REINSERTING THE BATTERY, WHICH WAS ATTEMPTED SEVERAL TIMES. THE DEVICE COMPLETELY CEASED TO FUNCTION, AND IT HAD TO BE REMOVED FROM THE PATIENT¿S BODY WITH THE TROCAR STILL IN PLACE, AS IT COULD NOT BE EXTRACTED THROUGH THE TROCAR IN THE NORMAL MANNER. THE SURGERY WAS DELAYED 5 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Wound Dehiscence | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
22-10-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, TRIED USING THE ECH60S AND IT DID NOT WORK, FAILED TO FIRE. OPENED ANOTHER DEVICE TO COMPLETE CASE. NO PATIENT HARM.
26-11-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP ANTERIOR RESECTION, NURSE REPORTED THAT A EC60A WITH GST60D WAS USED DURING THE OPERATION BY DR. SURGEON CLOSED THE JAW ON PATIENT TISSUE AND READY TO FIRE THE DEVICE. WHEN SURGEON WAS PRESSING THE FIRING TRIGGER, HE CAN FEEL THE BLADE WAS NOT ADVANCING. THE EC60A WAS JAMMED AND NEED TO REVERSE THE KNIFE TO OPEN JAWS. PHOTOS WERE TAKEN INTRA-OP AND ATTACHED FOR FURTHER INVESTIGATION. THE EC60A IS THEN REMOVED FROM THE PATIENT AND SURGEON USED ANOTHER ECHELON TO FINISH THE PROCEDURE.
26-11-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP ANTERIOR RESECTION, NURSE REPORTED THAT A EC60A WITH GST60D WAS USED DURING THE OPERATION BY DR. SURGEON CLOSED THE JAW ON PATIENT TISSUE AND READY TO FIRE THE DEVICE. WHEN SURGEON WAS PRESSING THE FIRING TRIGGER, HE CAN FEEL THE BLADE WAS NOT ADVANCING. THE EC60A WAS JAMMED AND NEED TO REVERSE THE KNIFE TO OPEN JAWS. PHOTOS WERE TAKEN INTRA-OP AND ATTACHED FOR FURTHER INVESTIGATION. THE EC60A IS THEN REMOVED FROM THE PATIENT AND SURGEON USED ANOTHER ECHELON TO FINISH THE PROCEDURE.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Wound Dehiscence | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
03-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A WIDE WEDGE RESECTION OF A LEFT LOWER LOBE NODULE, PERFORMED USING A COMBINATION OF EGIA 60 AND 45 TAN AND PURPLE TRI-STAPLE RELOADS, ISSUES WERE ENCOUNTERED WITH THE STAPLER. SPECIFICALLY, A TAN RELOAD WAS APPLIED NEAR THE LOWER LOBEPULMONARY ARTERY BRANCH, AND AT CLOSURE, THE ENTIRE STAPLE LINE DEHISCED. THIS REQUIRED REPAIR USING A 3.0 PROLENE SUTURE. THE PROCEDURE WAS EXTENDED BY 30 MINUTES DUE TO THESE ISSUES. THE REPORTED PRODUCT ISSUES INCLUDED FIRING DIFFICULTIES AND STAPLE FORMATION PROBLEMS. SUTURING WAS PERFORMED TO REPLACE THE STAPLE LINE.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, NO PROBLEM ON THE 1ST FIRING. WHEN IT WAS FIRED ON THE LEFT UPPER TRUNK OF PULMONARY ARTERY AT THE 2ND FIRING, THE KNIFE STOPPED AND THE DEVICE WAS LOCKED OUT. THE KNIFE AUTOMATICALLY RETURNED AFTER PUSHING THE KNIFE REVERSE SWITCH. AFTERWARDS, IT WAS REPRODUCED WITH THE SAME DEVICE AFTER THE SURGERY, BUT THE KNIFE STOPPED ON THE WAY AND THE DEVICE WAS LOCKED OUT. THE KNIFE RETURNED WITH THE KNIFE REVERSE SWITCH. THE CARTRIDGE COLOR WAS WHITE. ANOTHER COMPETITIVE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
19-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, NO PROBLEM ON THE 1ST FIRING. WHEN USED ON THE BRONCHIAL TUBE AT THE 2ND FIRING, THE KNIFE STOPPED ON THE WAY AND THE DEVICE WAS LOCKED OUT. ADDITIONAL SUTURE WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST GREEN RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST BLACK RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT IN THE COURSE OF A PROCEDURE USING THE PSEE60A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE WAS SUDDENLY STOPPED IN THE MIDDLE OF CARTRIDGE. IT OCCURRED TWO MORE TIMES. AS AN EMERGENCY MEASURE, THE SURGEON BACKED OFF THE BLADE AND DETACHED IT FROM THE TISSUE. IT WAS FURTHER REPORTED THAT THE TROCAR HAD BEEN BENT WHEN IT WAS INSERTED INTO THE BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST BLACK RELOAD 45MM 6 ROW | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT IN THE COURSE OF A PROCEDURE USING THE PSEE60A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE WAS SUDDENLY STOPPED IN THE MIDDLE OF CARTRIDGE. IT OCCURRED TWO MORE TIMES. AS AN EMERGENCY MEASURE, THE SURGEON BACKED OFF THE BLADE AND DETACHED IT FROM THE TISSUE. IT WAS FURTHER REPORTED THAT THE TROCAR HAD BEEN BENT WHEN IT WAS INSERTED INTO THE BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST GREEN RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT IN THE COURSE OF A PROCEDURE USING THE PSEE60A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE WAS SUDDENLY STOPPED IN THE MIDDLE OF CARTRIDGE. IT OCCURRED TWO MORE TIMES. AS AN EMERGENCY MEASURE, THE SURGEON BACKED OFF THE BLADE AND DETACHED IT FROM THE TISSUE. IT WAS FURTHER REPORTED THAT THE TROCAR HAD BEEN BENT WHEN IT WAS INSERTED INTO THE BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST GREEN RELOAD 45MM 6 ROW | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT IN THE COURSE OF A PROCEDURE USING THE PSEE60A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE WAS SUDDENLY STOPPED IN THE MIDDLE OF CARTRIDGE. IT OCCURRED TWO MORE TIMES. AS AN EMERGENCY MEASURE, THE SURGEON BACKED OFF THE BLADE AND DETACHED IT FROM THE TISSUE. IT WAS FURTHER REPORTED THAT THE TROCAR HAD BEEN BENT WHEN IT WAS INSERTED INTO THE BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 45MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT IN THE COURSE OF A PROCEDURE USING THE PSEE60A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE WAS SUDDENLY STOPPED IN THE MIDDLE OF CARTRIDGE. IT OCCURRED TWO MORE TIMES. AS AN EMERGENCY MEASURE, THE SURGEON BACKED OFF THE BLADE AND DETACHED IT FROM THE TISSUE. IT WAS FURTHER REPORTED THAT THE TROCAR HAD BEEN BENT WHEN IT WAS INSERTED INTO THE BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 45MM, 6 ROW | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT IN THE COURSE OF A PROCEDURE USING THE PSEE60A, AN ERROR OCCURRED IN THE PROCESS OF FIRING. WHEN SURGEON PULLED BACK THE SAFETY LOCK AND FIRED THE TRIGGER, THE BLADE WAS SUDDENLY STOPPED IN THE MIDDLE OF CARTRIDGE. IT OCCURRED TWO MORE TIMES. AS AN EMERGENCY MEASURE, THE SURGEON BACKED OFF THE BLADE AND DETACHED IT FROM THE TISSUE. IT WAS FURTHER REPORTED THAT THE TROCAR HAD BEEN BENT WHEN IT WAS INSERTED INTO THE BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST BLACK RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE RELOAD WAS PLACED IN THE STAPLER THEN THE DOCTOR INSERTED THE STAPLER AND TOOK A BITE ON THE STOMACH. THE DOCTOR WAITED 15 SEC AND FIRED, THE FIRST 2-3 ROWS CAME UP BUT THEN THE STAPLER STOPPED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
02-11-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GOLD RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX 45 STANDARD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT AFTER STOPPING IN THE MIDDLE OF OPERATION, THE BLADE DID NOT RETURN, SO THE SURGEON ENDED UP REMOVING THE BLADE WITH A MANUAL OVER-RIDE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT AFTER STOPPING IN THE MIDDLE OF OPERATION, THE BLADE DID NOT RETURN, SO THE SURGEON ENDED UP REMOVING THE BLADE WITH A MANUAL OVER-RIDE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Patient Device Interaction Problem | A01 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
04-12-2024 | Injury | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Patient Device Interaction Problem | A01 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS DR. PERFORMED THE PROCEDURE WITH A WHITE REFILL AND HAD TO PERFORM HEMOSTASIS WITH BIPOLAR BECAUSE THE TISSUE DID NOT STOP BLEEDING. THE SURGERY WAS DELAYED 10 MINUTES. THE SURGEON PERFORMED HEMOSTASIS WITH BIPOLAR AND WITH COMPRESSION AND GAUZE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED.
23-12-2024 | Malfunction | johnson & johnson medical (suzhou | 07-01-2025 | FINAL ASSY 60MM STD DEVICE | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | johnson & johnson medical (suzhou | 07-01-2025 | FINAL ASSY 60MM STD DEVICE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | johnson & johnson medical (suzhou | 07-01-2025 | FINAL ASSY 60MM STD DEVICE | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-12-2024 | Malfunction | johnson & johnson medical (suzhou | 07-01-2025 | FINAL ASSY 60MM STD DEVICE | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
27-11-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A LAPAROSCOPIC GASTRECTOMY, THE BATTERY WAS LOADED INTO THE DEVICE, BUT IT DID NOT ACTIVATED AT THE 1ST FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
17-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLER ABRUPTLY CUT THROUGH THE SEROSA AND MUSCULARIS AND APPEARED TO LEAVE A SMALL OPENING IN THE MUCOSA ON THE MEDIAL SIDE, THIS WAS THEN STAPLED WITH A NEW STAPLER AT THE BREECH. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO ADVERSE CONSEQUENCES FOR THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
25-11-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-11-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE SOUND OF THE MOTOR ONLY LASTED FOR LESS THAN ONE SECOND, AND THE KNIFE STOPPED MOVING FORWARD ONCE THE SURGEON FIRED THE DEVICE. THE SURGEON PUSHED THE REVERSE BUTTON AND REMOVED THE DEVICE FROM THE PATIENT. HOWEVER, ONCE IT SEEMED THAT THE BOARD HASN'T MOVED AND THE RELOADS HAVEN'T BEEN PUSHED OUT, BUT THE DEVICE STILL STUCK WHILE FIRING. THE HCP CHECKED AND CONFIRMED THAT THE KNIFE AND RELOADS HAVEN'T BEEN FIRED (NO STAPLE WAS FIRED) THROUGH THE BLOOD VESSEL, SO THE BLOOD VESSEL REMAINED UNHARMED. THERE WAS NO PATIENT CONSEQUENCE. ADDITIONAL INFORMATION REQUESTED.
29-10-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST WHITE RELOAD, 60MM, 6-ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A SEMI-OPEN PNEUMONECTOMY PROCEDURE, THE DEVICE STOPPED SUDDENLY ON THE PULMONARY ARTERY AT THE 1ST FIRING. THE DEVICE WAS LOCKED OUT. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
28-10-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE 1TH FIRING. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL COULD NOT FIRE AND WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE OF THE PSEE60A (C9E76Y) DID NOT RETURN. THE KNIFE DID NOT RETURN (THE JAWS DID NOT OPEN) EVEN IF THE MANUAL OVERRIDE LEVER WAS USED, AND AN ADDITIONAL RESECTION WAS PERFORMED. THE OTHER PSEE60A (C9E61V) WAS LOCKED OUT AT THE 4TH FIRING. (THE 1ST TO 3RD FIRING WERE NO PROBLEM.) THERE WAS NO BLEEDING. THE CARTRIDGE COLORS WERE BLUE AND GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
25-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A LAPAROSCOPIC COLECTOMY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE OF THE PSEE60A (C9E76Y) DID NOT RETURN. THE KNIFE DID NOT RETURN (THE JAWS DID NOT OPEN) EVEN IF THE MANUAL OVERRIDE LEVER WAS USED, AND AN ADDITIONAL RESECTION WAS PERFORMED. THE OTHER PSEE60A (C9E61V) WAS LOCKED OUT AT THE 4TH FIRING. (THE 1ST TO 3RD FIRING WERE NO PROBLEM.) THERE WAS NO BLEEDING. THE CARTRIDGE COLORS WERE BLUE AND GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
19-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE AND THE KNIFE MOVED TO THE DISTAL END BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE AND THE KNIFE MOVED TO THE DISTAL END BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE AND THE KNIFE MOVED TO THE DISTAL END BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
28-11-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY, THE FIRING TRIGGER DID NOT OPEN AFTER FIRING. THE FIRING TRIGGER WAS OPENED BY FORCIBLY RELEASING. THE DEVICE WAS LOCKED OUT. THERE WAS NO EFFECT ON THE PATIENT. THERE WAS NO BLEEDING. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN RESPIRATORY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS LOCKED OUT AT THE 3RD FIRING. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
28-10-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-10-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GREEN RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-11-2024 | Malfunction | johnson & johnson medical (suzhou | 07-01-2025 | FINAL ASSY 60MM STD DEVICE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE EFFECT OF NAILING WAS POOR DURING USE, AND HALF OF THE INTESTINAL TUBE WAS NOT CLOSED (BURSTING NAIL), AND THE STUMP WAS NOT CLOSED BY MANUAL SUTURE. ADDITIONAL INFORMATION REQUESTED.
14-11-2024 | Malfunction | johnson & johnson medical (suzhou | 07-01-2025 | FINAL ASSY 60MM STD DEVICE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE EFFECT OF NAILING WAS POOR DURING USE, AND HALF OF THE INTESTINAL TUBE WAS NOT CLOSED (BURSTING NAIL), AND THE STUMP WAS NOT CLOSED BY MANUAL SUTURE. ADDITIONAL INFORMATION REQUESTED.
28-11-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
25-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST GREEN RELOAD 45MM 6 ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TUMOR RESECTION PROCEDURE, OPENED THE PACKAGE AND NOTED THE METAL PART OF CARTRIDGE WAS DEFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
24-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | GST BLUE RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | PWRD ECH FLEX 60MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
16-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING APPENDECTOMY, INVOLVING THE MANUAL HANDLE AND RELOAD. AFTER FIRING, THE SURGEON IDENTIFIED SMALL COMPONENTS IN THE SITUS, WHICH WERE SUSPECTED TO HAVE ORIGINATED FROM EITHER THE HANDLE OR THE LOADING UNIT. ALL COMPONENTS WERE REMOVED FROM THE SITUS AND SUBSEQUENTLY DISCARDED. THERE WERE NO REPORTED CONSEQUENCES OR IMPACT TO THE PATIENT.
16-12-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING APPENDECTOMY, INVOLVING THE MANUAL HANDLE AND RELOAD. AFTER FIRING, THE SURGEON IDENTIFIED SMALL COMPONENTS IN THE SITUS, WHICH WERE SUSPECTED TO HAVE ORIGINATED FROM EITHER THE HANDLE OR THE LOADING UNIT. ALL COMPONENTS WERE REMOVED FROM THE SITUS AND SUBSEQUENTLY DISCARDED. THERE WERE NO REPORTED CONSEQUENCES OR IMPACT TO THE PATIENT.
14-11-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, WHILE CUTTING AND CLOSING THE STOMACH BODY, TWO CONSECUTIVE FIRINGS RESULTED IN POOR STAPLE FORMATION AND OOZING BLOOD IN THE TISSUE. TO RESOLVE THE ISSUE, AN ULTRASONIC SCALPEL WAS USED TO STOP BLEEDING. AFTER THE ENTIRE GASTRIC BODY WAS CUT AND CLOSED, SUTURES WERE USED FOR SUTURE EMBEDDING. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
06-09-2024 | Malfunction | us surgical puerto rico | 07-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, DURING STAPLING, THE HANDLE WARPED AND BROKE, MAKING IT IMPOSSIBLE TO USE. NO INTERVENTION WAS PERFORMED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT THE SITE OF THE INITIAL FIRING POST CLAMP-UP.
13-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE AND SHELL LIT UP GREEN WITHOUT ANY PROBLEMS. THE HANDLE AND ADAPTER ARE ALSO LIT IN GREEN, BUT WHEN THE CARTRIDGE WAS CONNECTED, IT RECOGNIZED THE CARTRIDGE, BUT THE USER COULD NOT OPEN OR CLOSE IT. THE SAME PHENOMENON OCCURRED REPEATEDLY WITH THE SAME HANDLE; THE ADAPTER WAS REMOVED FROM THE HANDLE AND WAS RE-INSERTED SEVERAL TIMES; IT RECOGNIZED THE CARTRIDGE. THE SAME ADAPTER AND ANOTHER HANDLE WERE USED WITHOUT ANY PROBLEMS. THE OPERATING ROOM POINTED OUT THE SAME DEFECT ABOUT ONE MONTH AGO, SO THE ACTUAL PRODUCT WAS CHECKED; HOWEVER, AT THAT TIME IT WAS ABLE TO RECOGNIZE THE CARTRIDGE WITHOUT ANY PROBLEMS, SO WE WAITED FOR A WHILE, BUT AFTER THAT, THE SAME THING HAPPENED SEVERAL TIMES. ANOTHER POWER HANDLE WAS USED TO RESOLVE THE ISSUE.
13-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE AND SHELL LIT UP GREEN WITHOUT ANY PROBLEMS. THE HANDLE AND ADAPTER ARE ALSO LIT IN GREEN, BUT WHEN THE CARTRIDGE WAS CONNECTED, IT RECOGNIZED THE CARTRIDGE, BUT THE USER COULD NOT OPEN OR CLOSE IT. THE SAME PHENOMENON OCCURRED REPEATEDLY WITH THE SAME HANDLE; THE ADAPTER WAS REMOVED FROM THE HANDLE AND WAS RE-INSERTED SEVERAL TIMES; IT RECOGNIZED THE CARTRIDGE. THE SAME ADAPTER AND ANOTHER HANDLE WERE USED WITHOUT ANY PROBLEMS. THE OPERATING ROOM POINTED OUT THE SAME DEFECT ABOUT ONE MONTH AGO, SO THE ACTUAL PRODUCT WAS CHECKED; HOWEVER, AT THAT TIME IT WAS ABLE TO RECOGNIZE THE CARTRIDGE WITHOUT ANY PROBLEMS, SO WE WAITED FOR A WHILE, BUT AFTER THAT, THE SAME THING HAPPENED SEVERAL TIMES. ANOTHER POWER HANDLE WAS USED TO RESOLVE THE ISSUE.
13-12-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, BEFORE PATIENT USE, THE HANDLE AND SHELL LIT UP GREEN WITHOUT ANY PROBLEMS. THE HANDLE AND ADAPTER ARE ALSO LIT IN GREEN, BUT WHEN THE CARTRIDGE WAS CONNECTED, IT RECOGNIZED THE CARTRIDGE, BUT THE USER COULD NOT OPEN OR CLOSE IT. THE SAME PHENOMENON OCCURRED REPEATEDLY WITH THE SAME HANDLE; THE ADAPTER WAS REMOVED FROM THE HANDLE AND WAS RE-INSERTED SEVERAL TIMES; IT RECOGNIZED THE CARTRIDGE. THE SAME ADAPTER AND ANOTHER HANDLE WERE USED WITHOUT ANY PROBLEMS. THE OPERATING ROOM POINTED OUT THE SAME DEFECT ABOUT 1 MONTH AGO, SO THE ACTUAL PRODUCT WAS CHECKED; HOWEVER, AT THAT TIME IT WAS ABLE TO RECOGNIZE THE CARTRIDGE WITHOUT ANY PROBLEMS, SO WE WAITED FOR A WHILE, BUT AFTER THAT, THE SAME THING HAPPENED SEVERAL TIMES. ANOTHER POWER HANDLE WAS USED TO RESOLVE THE ISSUE.
04-10-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL PANCREATECTOMY, WHILE FIRING THE STAPLER, THE ARTICULATING KNOB WAS BROKEN. THE BROKEN KNOB WAS DISENGAGED, BUT THERE WAS NO DEVICE COMPONENT THAT FELL INTO THE CAVITY OF THE PATIENT. SUTURING WAS DONE AS A STAPLE LINE REPLACEMENT. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT CLAMPING MECHANISM WAS DEFORMED AND THE RELOAD WAS PARTIALLY FIRED TO THE 2CM CUT LINE.
04-10-2024 | Malfunction | covidien lp llc north haven | 07-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DISTAL PANCREATECTOMY, WHILE FIRING THE STAPLER, THE ARTICULATING KNOB WAS BROKEN. THE BROKEN KNOB WAS DISENGAGED, BUT THERE WAS NO DEVICE COMPONENT THAT FELL INTO THE CAVITY OF THE PATIENT. SUTURING WAS DONE AS A STAPLE LINE REPLACEMENT. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT CLAMPING MECHANISM WAS DEFORMED AND THE RELOAD WAS PARTIALLY FIRED TO THE 2CM CUT LINE.
31-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ETHICONECHELON FLEX POWERED VASCULAR STAPLER |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLER WOULD NOT RELEASE. PHYSICIAN HAD TO IMPLEMENT EMERGENT RELEASE.
02-01-2025 | Malfunction | covidien | 07-01-2025 | GIA STAPLER |  |  | Fluid Discharge | Event Description: DURING A RIGHT COLON RESECTION WITH LAPAROSCOPIC ASSISTANCE CONVERTED TO OPEN PROCEDURE WITH ILEAL TRANSVERSE COLOSTOMY, THE SURGEON WRITES, 'A SEGMENT OF DISTAL ILEUM PROXIMAL TO THE SEROMUSCULAR TEAR WAS SELECTED AND DIVIDED WITH A GIA(GASTROINTESTINAL ANASTOMOSIS) STAPLER. UNFORTUNATELY, ONLY MINIMAL PARTIAL STAPLE LINE FIRED AND THERE WAS SOME EGRESS OF SMALL BOWEL FLUID. 2 LOADS OF THE TA (THORACIC ABDOMINA) WERE USED TO CLOSE THE PROXIMAL AND DISTAL ENDS.'
31-12-2024 | Malfunction | ethicon endo-surgery | 07-01-2025 | ETHICONECHELON FLEX POWERED VASCULAR STAPLER |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLER SELF-INJECTED LOAD.
28-11-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC LOBECTOMY, THERE WAS A PROCEDURAL/DEVICE-RELATED ISSUE INVOLVING THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP, WHERE POOR STAPLE FORMATION OCCURRED AS THE FIRING WAS NOT SMOOTH WHILE DETACHING A BLOOD VESSEL. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER WITH A NEW HANDLE TO COMPLETE THE PROCEDURE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
28-11-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC LOBECTOMY, THERE WAS A PROCEDURAL/DEVICE-RELATED ISSUE INVOLVING THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP, WHERE POOR STAPLE FORMATION OCCURRED AS THE FIRING WAS NOT SMOOTH WHILE DETACHING A BLOOD VESSEL. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER WITH A NEW HANDLE TO COMPLETE THE PROCEDURE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
01-01-2025 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE PEDIATRIC THORACIC RESECTION, WHEN ON THE PULMONARY VASCULAR CLOSURE, THERE WAS POOR STAPLE FORMATION AND AN INABILITY TO EFFECTIVELY ANASTOMOSE THE BLOOD VESSELS. ADDITIONALLY, THE STAPLE LINE WAS INCOMPLETE CENTRALLY. THERE WAS BLEEDING, BUT NOT MANY. THE INSTRUMENT WAS REPLACED, AND THE SUTURE WAS RE-FIRED TO RESOLVE THE ISSUE.
03-01-2025 | Malfunction | covidien | 08-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING AN OPEN WHIPPLE PROCEDURE, BEFORE FIRING THE JAWS WOULD NOT OPEN AND CLOSE COMPLETELY. A NEW HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING AN OPEN WHIPPLE PROCEDURE, BEFORE FIRING THE JAWS WOULD NOT OPEN AND CLOSE COMPLETELY. A NEW HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
31-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION, THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION, THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Retraction Problem | A0510 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION, THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION, THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION, THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION, THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Retraction Problem | A0510 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS REPAIR REVISION THE KNIFE BLADE WAS DIFFICULT TO RETRACT AND THE DEVICE LOCKED ON THE STOMACH TISSUE, PREVENTING REMOVAL. THE SURGEON HAD TO CUT THE TISSUE ON BOTH SIDES OF THE STAPLER WITH SCISSORS, RESULTING IN BLEEDING FROM THE STOMACH TISSUE. THE BLEEDING WAS DESCRIBED AS OOZING/MINIMAL/MODERATE. TO RESOLVE THE ISSUE, ADDITIONAL RESECTION AND RE-STAPLING WAS PERFORMED. THIS COMPLICATION EXTENDED THE SURGICAL TIME BY AN ADDITIONAL TEN MINUTES.
31-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC LOBECTOMY, THE RELOAD EXPERIENCED ISSUES WHILE DETACHING THE PULMONARY BLOOD VESSEL. THE STAPLER HANDLE COULD NOT BE ADVANCED AFTER PRESSING THE GREEN BUTTON, AND THE STAPLER WAS UNABLE TO FIRE. TO RESOLVE THE ISSUE, THE STAPLER WAS REPLACED WITH A RELOAD OF THE SAME MODEL, WHICH ALLOWED THE PROCEDURE TO BE COMPLETED. THERE WAS NO PATIENT INJURY.
01-12-2024 | Malfunction | covidien | 08-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE RELOAD DID NOT FIRE. ANOTHER RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE LAPAROSCOPIC COLON RESECTION, WHEN IN THE THIRD SHOT OF FUNCTIONAL END-TO-END ANASTOMOSIS (LATERAL ANASTOMOSIS), PRE FIRING EVENT OCCURRED. THERE WAS A STAPLING NEAR THE PROXIMAL SIDE OF THE CARTRIDGE¿S JAW. IT WAS NOTED ALSO THAT THE FIRING COULD NOT BE PERFORMED AT ALL AFTER PRESSING THE GREEN BUTTON. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY
24-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE USE OF THE EGIA60AVM EGIA 60 ARTIC VASC MED SULU STAPLER, THE STAPLER FIRED BUT DID NOT CUT OR FIRE STAPLES AND THE INSTRUMENT MADE A WEIRD NOISE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED, AND THE DEVICE WAS REPLACED. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT AND THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
24-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE USE OF THE EGIA60AVM EGIA 60 ARTIC VASC MED SULU STAPLER, THE STAPLER FIRED BUT DID NOT CUT OR FIRE STAPLES AND THE INSTRUMENT MADE A WEIRD NOISE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED, AND THE DEVICE WAS REPLACED. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT AND THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PROCEDURE, THE TACKER FAILED TO FIRE AND AFTER FIRING IT DID NOT ADHERE TO THE TISSUE. NO TACKS WERE DISENGAGED INTO THE PATIENT CAVITY. ANOTHER DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ABSORBATACK | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PROCEDURE, THE TACKER FAILED TO FIRE AND AFTER FIRING IT DID NOT ADHERE TO THE TISSUE. NO TACKS WERE DISENGAGED INTO THE PATIENT CAVITY. ANOTHER DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | covidien | 08-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER DURING A SLEEVE GASTRECTOMY, AFTER PERFORMING THE FIRST SHOT, THE STAPLE LINE WAS LEFT WITH SEGMENTS THAT WERE NOT STAPLED CORRECTLY AND LOOSE STAPLES. IT CAUSED SLIGHT BLEEDING AND THE SURGEON HAD TO PERFORM SUTURING AS A STAPLE LINE REINFORCEMENT. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES.
23-12-2024 | Malfunction | covidien | 08-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING A SLEEVE GASTRECTOMY, AFTER PERFORMING THE FIRST SHOT, THE STAPLE LINE WAS LEFT WITH SEGMENTS THAT WERE NOT STAPLED CORRECTLY AND LOOSE STAPLES. IT CAUSED SLIGHT BLEEDING AND THE SURGEON HAD TO PERFORM SUTURING AS A STAPLE LINE REINFORCEMENT. THE SURGICAL TIME WAS EXTENDED TO 30 MINUTES.
22-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER PRIOR TO USE OF THE DEVICE, THE HANDLE SHAFT OF THE STAPLER WAS OUT OF PLACE. WHEN THE OPERATING ROOM NURSE WAS PREPARING THE INSTRUMENT BEFORE THE SURGERY, IT WAS FOUND THAT THE HANDLE SHAFT WAS SEPARATED FROM THE HANDLE BODY. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED.
20-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE STAPLER WAS UNABLE TO FIRE WHEN DISCONNECTING THE ARTERY. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE HANDLE WITH A NEW ONE, AND THE PROCEDURE WAS COMPLETED WITHOUT FURTHER COMPLICATIONS. THERE WAS NO PATIENT INJURY.
24-07-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE DEMONSTRATION UNIT WAS SET ON THE CHARGER, THE RED LIGHT FLASHED AND A BATTERY ERROR MESSAGE APPEARED ON THE DISPLAY. THE DEVICE COULD NOT BE CHARGED. THE DEVICE WAS REPEATEDLY REBOOTED BUT ISSUE DID NOT IMPROVE. THE CHARGER COULD BE USED WITH OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ONE OF THE BATTERY CELLS TO BE VENTED.
24-07-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Battery Problem | A0705 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE DEMONSTRATION UNIT WAS SET ON THE CHARGER, THE RED LIGHT FLASHED AND A BATTERY ERROR MESSAGE APPEARED ON THE DISPLAY. THE DEVICE COULD NOT BE CHARGED. THE DEVICE WAS REPEATEDLY REBOOTED BUT ISSUE DID NOT IMPROVE. THE CHARGER COULD BE USED WITH OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ONE OF THE BATTERY CELLS TO BE VENTED.
24-07-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, WHEN THE DEMONSTRATION UNIT WAS SET ON THE CHARGER, THE RED LIGHT FLASHED AND A BATTERY ERROR MESSAGE APPEARED ON THE DISPLAY. THE DEVICE COULD NOT BE CHARGED. THE DEVICE WAS REPEATEDLY REBOOTED BUT ISSUE DID NOT IMPROVE. THE CHARGER COULD BE USED WITH OTHER HANDLES. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND ONE OF THE BATTERY CELLS TO BE VENTED.
17-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 08-01-2025 | CAPSURE | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, THE CAPSURE FIXATION DEVICE MISFIRES AND DEPLOYED TWO FASTENERS. IT WAS REPORTED THAT THE ADDITIONAL FASTENER WAS REMOVED FROM THE PATIENT. THE PROCEDURE WAS COMPLETED USING A NEW DEVICE FROM A DIFFERENT LOT NUMBER. IT WAS REPORTED THE SAMPLE WAS DISCARDED AND THERE WAS NO PATIENT HARM OR INJURY.
01-01-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST BLACK RELOAD 60MM 6 ROW | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE MISFIRED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST BLACK RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE MISFIRED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST BLACK RELOAD 60MM 6 ROW | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE MISFIRED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
13-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATICODUODENECTOMY, DURING INTESTINAL RESECTION, THE HANDLE AND THE CARTRIDGE COULD NOT BE CONNECTED. UPON VISUAL INSPECTION, THE SHAFT COVER WAS FOUND TO HAVE SHIFTED DOWNWARD. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATICODUODENECTOMY, DURING INTESTINAL RESECTION, THE HANDLE AND THE CARTRIDGE COULD NOT BE CONNECTED. UPON VISUAL INSPECTION, THE SHAFT COVER WAS FOUND TO HAVE SHIFTED DOWNWARD. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATICODUODENECTOMY, DURING INTESTINAL RESECTION, THE HANDLE AND THE CARTRIDGE COULD NOT BE CONNECTED. UPON VISUAL INSPECTION, THE SHAFT COVER WAS FOUND TO HAVE SHIFTED DOWNWARD. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
31-07-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SEGMENTAL SURGERY, THE RELOAD WAS DEFECTIVE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE CLAMPING MECHANISM WAS DEFORMED.
31-07-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SEGMENTAL SURGERY, THE RELOAD WAS DEFECTIVE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE CLAMPING MECHANISM WAS DEFORMED.
27-11-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RADICAL RECTAL SURGERY , DURING THE CLOSURE AND DETACHMENT OF THE RECTUM, THE STAPLER WAS UNABLE TO FIRE, AND THE SURGEON COULD NOT SQUEEZE THE HANDLE. THE HANDLE AND RELOAD WERE REPLACED WITH A NEW ONE TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
03-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PULMONARY, THE SINGLE USE LOADING UNIT AND STAPLING DEVICES LOCKED ON TISSUE AND COULD NOT BE REMOVED, AND THE HANDLE AND RELOAD COULD NOT BE RETURNED. THE ISSUE OCCURRED AFTER FIRING THE DEVICES, WITH THE RELOAD AND HANDLE COMPONENTS BEING AFFECTED. SUTURING WAS PERFORMED TO RESOLVE THE ISSUE WITH THE DEVICE.
03-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Retraction Problem | A0510 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PULMONARY, THE SINGLE USE LOADING UNIT AND STAPLING DEVICES LOCKED ON TISSUE AND COULD NOT BE REMOVED, AND THE HANDLE AND RELOAD COULD NOT BE RETURNED. THE ISSUE OCCURRED AFTER FIRING THE DEVICES, WITH THE RELOAD AND HANDLE COMPONENTS BEING AFFECTED. SUTURING WAS PERFORMED TO RESOLVE THE ISSUE WITH THE DEVICE.
03-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PULMONARY, THE SINGLE USE LOADING UNIT AND STAPLING DEVICES LOCKED ON TISSUE AND COULD NOT BE REMOVED, AND THE HANDLE AND RELOAD COULD NOT BE RETURNED. THE ISSUE OCCURRED AFTER FIRING THE DEVICES, WITH THE RELOAD AND HANDLE COMPONENTS BEING AFFECTED. SUTURING WAS PERFORMED TO RESOLVE THE ISSUE WITH THE DEVICE.
03-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Retraction Problem | A0510 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PULMONARY, THE SINGLE USE LOADING UNIT AND STAPLING DEVICES LOCKED ON TISSUE AND COULD NOT BE REMOVED, AND THE HANDLE AND RELOAD COULD NOT BE RETURNED. THE ISSUE OCCURRED AFTER FIRING THE DEVICES, WITH THE RELOAD AND HANDLE COMPONENTS BEING AFFECTED. SUTURING WAS PERFORMED TO RESOLVE THE ISSUE WITH THE DEVICE.
16-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, DURING CLOSURE AND DIVISION OF THE RECTUM, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE HANDLE COULD NOT BE SQUEEZED. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THEISSUE. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, DURING CLOSURE AND DIVISION OF THE RECTUM, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND THE HANDLE COULD NOT BE SQUEEZED. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THEISSUE. THERE WAS NO PATIENT INJURY.
07-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE 1TH FIRING. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL COULD NOT FIRE AND WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
21-11-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY IT WASN'T POSSIBLE TO FIRE THE STAPLER THE SURGERY WAS DELAYED 5 MINUTES. REPLACED BY ANOTHER INSTRUMENT. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
08-01-2025 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE POWERED HANDLE DID NOT CHARGE, THE POWER HANDLE WOULD NOT RECOGNIZED AND CHARGER BAY ERROR. IT WAS NOTED THAT CHARGER WAS WORKING WITH OTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THE VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE).
08-01-2025 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRE-OPERATIVELY, THE POWERED HANDLE DID NOT CHARGE, THE POWER HANDLE WOULD NOT RECOGNIZED AND CHARGER BAY ERROR. IT WAS NOTED THAT CHARGER WAS WORKING WITH OTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THE VENTED BATTERY WAS DETERMINED TO BE FROM BATTERY DEGRADATION (COMPONENT FAILURE).
16-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC PROCEDURE FOR LOW RECTAL CANCER RESECTION, THE STAPLER WAS UNABLE TO FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE BUT WAS ABLE TO PRESS THE GREEN BUTTON WHILE ATTEMPTING TO TRANSECT THE RECTUM. THE ISSUE WAS RESOLVED BY REPLACING THE STAPLER WITH A NEW CARTRIDGE, ALLOWING THE PROCEDURE TO BE COMPLETED WITHOUT EXTENDING SURGICAL TIME OR CAUSING INJURY TO THE PATIENT.
17-09-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Failure to Charge | A070603 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE POWER HANDLE DID NOT CHARGE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
17-09-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Complete Loss of Power | A070801 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, THE POWER HANDLE DID NOT CHARGE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
16-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | EEA | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS INVOLVING THE EEA STAPLER, AFTER FIRING THE STAPLER THE DEVICE WAS DIFFICULT TO REMOVE FROM THE CAVITY AND THE STAPLE LINE WAS FOUND TO BE INCOMPLETE WITH A HOLE PRESENT. AS A RESOLUTION, SUTURING WAS PERFORMED AS A REPLACEMENT FOR THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 10-15 MINUTES.
16-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | EEA | Difficult to Remove | A150207 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS INVOLVING THE EEA STAPLER, AFTER FIRING THE STAPLER THE DEVICE WAS DIFFICULT TO REMOVE FROM THE CAVITY AND THE STAPLE LINE WAS FOUND TO BE INCOMPLETE WITH A HOLE PRESENT. AS A RESOLUTION, SUTURING WAS PERFORMED AS A REPLACEMENT FOR THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 10-15 MINUTES.
16-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | EEA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS INVOLVING THE EEA STAPLER, AFTER FIRING THE STAPLER THE DEVICE WAS DIFFICULT TO REMOVE FROM THE CAVITY AND THE STAPLE LINE WAS FOUND TO BE INCOMPLETE WITH A HOLE PRESENT. AS A RESOLUTION, SUTURING WAS PERFORMED AS A REPLACEMENT FOR THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 10-15 MINUTES.
16-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | EEA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS INVOLVING THE EEA STAPLER, AFTER FIRING THE STAPLER THE DEVICE WAS DIFFICULT TO REMOVE FROM THE CAVITY AND THE STAPLE LINE WAS FOUND TO BE INCOMPLETE WITH A HOLE PRESENT. AS A RESOLUTION, SUTURING WAS PERFORMED AS A REPLACEMENT FOR THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 10-15 MINUTES.
03-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 08-01-2025 | CAPSURE | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A PROCEDURE, THE FASTENERS OF CAPSURE DEVICE CAME OUT CONNECTED ONE TO ANOTHER (TWO FASTENERS AT A TIME). THERE WAS NO PATIENT INJURY. AS REPORTED, ANOTHER CAPSURE WAS USED TO COMPLETE THE CASE. ADDENDUM: AS REPORTED, THE PROCEDURE PERFORMED WAS A POSTOPERATIVE VENTRAL HERNIA REPAIR (POVH) AND THE FIRST FOUR FASTENERS WERE SUCCESSFULLY FIXATED. IT WAS REPORTED THAT THE PROBLEM OCCURRED AFTER 5TH FASTENER DEPLOYED, THE FASTENERS CAME OUT CONNECT TO EACH OTHER WERE NOT FIXATED AND REMOVED FROM THE PATIENT'S BODY. ANOTHER CAPSURE WAS USED TO COMPLETE THE CASE. THERE WAS NO REPORTED PATIENT INJURY.
01-08-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE COULD BE PRESSED THE GREEN SAFETY, BUT COULD NOT BE FIRED, THE OPERATION WAS STOPPED. THE DEVICE WAS REPLACED WITH A NEW CUTTING STAPLER TO COMPLETE THE OPERATION. A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
01-08-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE COULD BE PRESSED THE GREEN SAFETY, BUT COULD NOT BE FIRED, THE OPERATION WAS STOPPED. THE DEVICE WAS REPLACED WITH A NEW CUTTING STAPLER TO COMPLETE THE OPERATION. A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
01-08-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE COULD BE PRESSED THE GREEN SAFETY, BUT COULD NOT BE FIRED. THE OPERATION WAS STOPPED. THE DEVICE WAS REPLACED WITH A NEW CUTTING STAPLER TO COMPLETE THE OPERATION. A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
01-08-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PROCEDURE, THE DEVICE COULD BE PRESSED THE GREEN SAFETY, BUT COULD NOT BE FIRED. THE OPERATION WAS STOPPED. THE DEVICE WAS REPLACED WITH A NEW CUTTING STAPLER TO COMPLETE THE OPERATION. A NEW HANDLE AND RELOAD WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY, THE DEVICE FIRED AT FIRST AND THEN STOPPED FIRING. THE ISSUE HAPPENED WITH TWO RELOADS. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien | 08-01-2025 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, IT FIRED AT FIRST AND THEN STOPPED FIRING. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-11-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE RADICAL RESECTION OF THE LEFT UPPER LOBE LUNG CANCER, THE DEVICE WAS USED TO CUT THE PULMONARY BLOOD VESSEL. AFTER PRESSING THE FIRING SAFETY BUTTON, THE RELOAD COULD NOT BE FIRED. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-11-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE RADICAL RESECTION OF THE LEFT UPPER LOBE LUNG CANCER, THE DEVICE WAS USED TO CUT THE PULMONARY BLOOD VESSEL. AFTER PRESSING THE FIRING SAFETY BUTTON, THE RELOAD COULD NOT BE FIRED. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-11-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LUNG RESECTION PROCEDURE, IT FAILED. IT WAS NOT POSSIBLE TO FIRE THE FIRST LOAD. A NEW HANDLE WAS USE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATICODUODENECTOMY, DURING INTESTINAL RESECTION, THE HANDLE AND THE CARTRIDGE COULD NOT BE CONNECTED. UPON VISUAL INSPECTION, THE SHAFT COVER WAS FOUND TO HAVE SHIFTED DOWNWARD. THE GRAY CYLINDER OF THE RELOAD HAD SLIPPED DOWNWARD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATICODUODENECTOMY, DURING INTESTINAL RESECTION, THE HANDLE AND THE CARTRIDGE COULD NOT BE CONNECTED. UPON VISUAL INSPECTION, THE SHAFT COVER WAS FOUND TO HAVE SHIFTED DOWNWARD. THE GRAY CYLINDER OF THE RELOAD HAD SLIPPED DOWNWARD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATICODUODENECTOMY, DURING INTESTINAL RESECTION, THE HANDLE AND THE CARTRIDGE COULD NOT BE CONNECTED. UPON VISUAL INSPECTION, THE SHAFT COVER WAS FOUND TO HAVE SHIFTED DOWNWARD. THE GRAY CYLINDER OF THE RELOAD HAD SLIPPED DOWNWARD. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING AN UNKNOWN RESPIRATORY PROCEDURE, IT WAS OBSERVED THAT THE FIRING ON THE DEVICE STOPPED ON THE WAY. THE DEVICE WAS RELEASED WITH THE KNIFE REVERSE SWITCH. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
25-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING AN UNKNOWN RESPIRATORY PROCEDURE, IT WAS OBSERVED THAT THE FIRING ON THE DEVICE STOPPED ON THE WAY. THE DEVICE WAS RELEASED WITH THE KNIFE REVERSE SWITCH. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
13-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INTESTINAL PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED HALF. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
26-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ECHELON 3000 60MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN RESPIRATORY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED OUT AND THE KNIFE STOPPED DURING USE. AFTER RE-ATTACHING ANOTHER CARTRIDGE, THE DEVICE COULD BE FIRED AS USUAL. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN RESPIRATORY PROCEDURE WITH UNIPORTAL VATS (U-VATS), IT WAS OBSERVED THAT WHEN THE DEVICE WAS FIRED AT THE 1ST FIRING, THE KNIFE MOVED FORWARD SLIGHTLY AND THEN STOPPED. THE KNIFE REVERSE SWITCH WAS USED TO RETURN THE KNIFE AND RELEASE THE CLAMP. AFTER REPLACING THE CARTRIDGE WITH A NEW ONE, THE DEVICE COULD BE USED. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-07-2024 | Injury | covidien lp llc north haven | 08-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Pneumothorax | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS CONDUCTED BETWEEN 2015 AND 2020 ON 125 CASES OF VIDEO-ASSISTED THORACOSCOPIC SURGERY PLEURODESIS FOR PRIMARY SPONTANEOUS PNEUMOTHORAX. BLEBS, BULLAE, AND EMPHYSEMATOUS-LIKE CHANGES WERE RESECTED USING EITHER ENDOGIA OR A COMPETITOR STAPLER. SUBSEQUENT MECHANICAL ABRASIVE PLEURODESIS WAS PERFORMED USING AN UNKNOWN MESH AND A CHEST DRAIN WAS PLACED. COMPLICATIONS INCLUDE POSTOPERATIVE AIR LEAK IN 13 PATIENTS IN DIGITAL CHEST DRAIN GROUP AND 20 PATIENTS IN CONVENTIONAL CHEST DRAIN GROUP. PROLONGED CHEST DRAIN WAS REQUIRED FOR PROLONGED AIR LEAK. THERE WERE 4 PATIENTS FROM THE DIGITAL CHEST DRAIN SYSTEM GROUP AND 12 PATIENTS FROM THE CONVENTIONAL CHEST DRAIN SYSTEM GROUP WHO EXPERIENCED RECURRENT PNEUMOTHORAX.
01-07-2024 | Injury | covidien lp llc north haven | 08-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS CONDUCTED BETWEEN 2015 AND 2020 ON 125 CASES OF VIDEO-ASSISTED THORACOSCOPIC SURGERY PLEURODESIS FOR PRIMARY SPONTANEOUS PNEUMOTHORAX. BLEBS, BULLAE, AND EMPHYSEMATOUS-LIKE CHANGES WERE RESECTED USING EITHER ENDOGIA OR A COMPETITOR STAPLER. SUBSEQUENT MECHANICAL ABRASIVE PLEURODESIS WAS PERFORMED USING AN UNKNOWN MESH AND A CHEST DRAIN WAS PLACED. COMPLICATIONS INCLUDE POSTOPERATIVE AIR LEAK IN 13 PATIENTS IN DIGITAL CHEST DRAIN GROUP AND 20 PATIENTS IN CONVENTIONAL CHEST DRAIN GROUP. PROLONGED CHEST DRAIN WAS REQUIRED FOR PROLONGED AIR LEAK. THERE WERE 4 PATIENTS FROM THE DIGITAL CHEST DRAIN SYSTEM GROUP AND 12 PATIENTS FROM THE CONVENTIONAL CHEST DRAIN SYSTEM GROUP WHO EXPERIENCED RECURRENT PNEUMOTHORAX.
01-07-2024 | Injury | covidien lp llc north haven | 08-01-2025 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Pneumothorax | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS CONDUCTED BETWEEN 2015 AND 2020 ON 125 CASES OF VIDEO-ASSISTED THORACOSCOPIC SURGERY PLEURODESIS FOR PRIMARY SPONTANEOUS PNEUMOTHORAX. BLEBS, BULLAE, AND EMPHYSEMATOUS-LIKE CHANGES WERE RESECTED USING EITHER ENDOGIA OR A COMPETITOR STAPLER. SUBSEQUENT MECHANICAL ABRASIVE PLEURODESIS WAS PERFORMED USING AN UNKNOWN MESH AND A CHEST DRAIN WAS PLACED. COMPLICATIONS INCLUDE POSTOPERATIVE AIR LEAK IN 13 PATIENTS IN DIGITAL CHEST DRAIN GROUP AND 20 PATIENTS IN CONVENTIONAL CHEST DRAIN GROUP. PROLONGED CHEST DRAIN WAS REQUIRED FOR PROLONGED AIR LEAK. THERE WERE 4 PATIENTS FROM THE DIGITAL CHEST DRAIN SYSTEM GROUP AND 12 PATIENTS FROM THE CONVENTIONAL CHEST DRAIN SYSTEM GROUP WHO EXPERIENCED RECURRENT PNEUMOTHORAX.
01-07-2024 | Injury | covidien lp llc north haven | 08-01-2025 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS CONDUCTED BETWEEN 2015 AND 2020 ON 125 CASES OF VIDEO-ASSISTED THORACOSCOPIC SURGERY PLEURODESIS FOR PRIMARY SPONTANEOUS PNEUMOTHORAX. BLEBS, BULLAE, AND EMPHYSEMATOUS-LIKE CHANGES WERE RESECTED USING EITHER ENDOGIA OR A COMPETITOR STAPLER. SUBSEQUENT MECHANICAL ABRASIVE PLEURODESIS WAS PERFORMED USING AN UNKNOWN MESH AND A CHEST DRAIN WAS PLACED. COMPLICATIONS INCLUDE POSTOPERATIVE AIR LEAK IN 13 PATIENTS IN DIGITAL CHEST DRAIN GROUP AND 20 PATIENTS IN CONVENTIONAL CHEST DRAIN GROUP. PROLONGED CHEST DRAIN WAS REQUIRED FOR PROLONGED AIR LEAK. THERE WERE 4 PATIENTS FROM THE DIGITAL CHEST DRAIN SYSTEM GROUP AND 12 PATIENTS FROM THE CONVENTIONAL CHEST DRAIN SYSTEM GROUP WHO EXPERIENCED RECURRENT PNEUMOTHORAX.
26-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ECHELON 3000 60MM STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED OUT AFTER FIRING. THEY HAD TO USE MANUAL OVERRIDE TO OPEN. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO ADVERSE CONSEQUENCES FOR THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
26-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ECHELON 3000 60MM STANDARD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED OUT AFTER FIRING. THEY HAD TO USE MANUAL OVERRIDE TO OPEN. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO ADVERSE CONSEQUENCES FOR THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
09-11-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
16-12-2024 | Malfunction | us surgical puerto rico | 08-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATICODUODENECTOMY, WHILE ON CLOSURE AND DETACHMENT OF DUODENUM, THE DEVICE DID NOT FIRE. A NEW HANDLE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTROPLASTY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE STOPPED DURING FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
09-11-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST WHITE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
16-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RADICAL RECTAL SURGERY , DURING THE CLOSURE AND DETACHMENT OF THE RECTUM, THE STAPLER WAS UNABLE TO FIRE AND THE SURGEON COULD NOT SQUEEZE THE HANDLE. THE DEVICE WAS REPLACED WITH A NEW ONE TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY PROCEDURE, A PVE35A WAS USED DURING VATS LOBECTOMY ON THE PULMONARY VEIN. THE BLOOD VESSEL WAS CLAMPED, AND WHEN TRYING TO FIRE AT THE 1ST FIRING, THE KNIFE MOVED FORWARD SLIGHTLY AND THEN STOPPED (LOCK-OUT). THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
27-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE STOPPED ON THE WAY OF THE 1ST FIRING. THE BATTERY WAS RELOADED, BUT THERE WAS NO STARTUP SOUND. THE KNIFE WAS RETURNED WITH THE MANUAL OVERRIDE LEVER, AND THE JAWS WERE OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
27-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | ECHELON 3000 60MM STANDARD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A LAPAROSCOPIC APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE STOPPED ON THE WAY OF THE 1ST FIRING. THE BATTERY WAS RELOADED, BUT THERE WAS NO STARTUP SOUND. THE KNIFE WAS RETURNED WITH THE MANUAL OVERRIDE LEVER, AND THE JAWS WERE OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
27-12-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN LAPAROSCOPIC PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED OUT ON THE WAY AT RECTUM DISSECTION. THE KNIFE WAS RETURNED WITH THE KNIFE REVERSE SWITCH. THE CARTRIDGE COLOR WAS GREEN. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
09-11-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED WITH IRREGULAR SHAPE (WITH STRAIGHT LEG) AFTER FIRING. THEY USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-11-2024 | Malfunction | ethicon endo-surgery | 08-01-2025 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
29-10-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS DIFFICULT TO FIRE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND FAILURE TO COMPLETELY FIRE THE RELOAD WILL RESULT IN AN INCOMPLETE CUT AND/OR INCOMPLETE STAPLE FORMATION, WHICH MAY RESULT IN POOR HEMOSTASIS AND/OR LEAKAGE.
29-10-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS DIFFICULT TO FIRE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND FAILURE TO COMPLETELY FIRE THE RELOAD WILL RESULT IN AN INCOMPLETE CUT AND/OR INCOMPLETE STAPLE FORMATION, WHICH MAY RESULT IN POOR HEMOSTASIS AND/OR LEAKAGE.
30-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, THE DEVICE WAS IMPOSSIBLE TO FIRE ON THE PULMONARY ARTERY, LEADING TO EXTENDED SURGICAL TIME OF LESS THAN 30 MINUTES. THERE WERE NO DIFFICULTIES DURING PREVIOUS FIRINGS. TO RESOLVE THE ISSUE, THE HANDLE AND RELOAD WERE CHANGED TO COMPLETE THE PROCEDURE. THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
30-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, THE DEVICE WAS IMPOSSIBLE TO FIRE ON THE PULMONARY ARTERY, LEADING TO EXTENDED SURGICAL TIME OF LESS THAN 30 MINUTES. THERE WERE NO DIFFICULTIES DURING PREVIOUS FIRINGS. TO RESOLVE THE ISSUE, THE HANDLE AND RELOAD WERE CHANGED TO COMPLETE THE PROCEDURE. THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Retraction Problem | A0510 | Insufficient Information | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Retraction Problem | A0510 | Appropriate Term / Code Not Available | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Misfire | A050502 | Appropriate Term / Code Not Available | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | Appropriate Term / Code Not Available | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Retraction Problem | A0510 | Insufficient Information | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Retraction Problem | A0510 | Appropriate Term / Code Not Available | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Misfire | A050502 | Insufficient Information | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Misfire | A050502 | Appropriate Term / Code Not Available | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | Appropriate Term / Code Not Available | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Retraction Problem | A0510 | Insufficient Information | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Retraction Problem | A0510 | Appropriate Term / Code Not Available | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | Misfire | A050502 | Appropriate Term / Code Not Available | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: DUPLICATE OF (B)(4).
11-12-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | Appropriate Term / Code Not Available | Event Description: DUPLICATE OF (B)(4).
30-10-2024 | Malfunction | covidien surgical | 08-01-2025 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING/MAINTENANCE, THE HANDLE DOESN'T CHARGE ONLY LIGHTS UP RED, THE OTHER HANDLE DISPLAY SCREEN SHUT OFF. IT WAS NOTED THAT THE CHARGER WAS WORKING WITH OTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT. PLI10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT USE WHEN THE DEVICE BATTERY VENTS ARE FOUND TO BE OPEN.
30-10-2024 | Malfunction | covidien surgical | 08-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING TESTING/MAINTENANCE, THE HANDLE DOESN'T CHARGE ONLY LIGHTS UP RED, THE OTHER HANDLE DISPLAY SCREEN SHUT OFF. IT WAS NOTED THAT THE CHARGER WAS WORKING WITH OTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT. PLI10 MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT USE WHEN THE DEVICE BATTERY VENTS ARE FOUND TO BE OPEN.
30-10-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TESTING/MAINTENANCE, THE HANDLE DOESN'T CHARGE ONLY LIGHTS UP RED. IT WAS NOTED THAT THE CHARGER WAS WORKING WITH OTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT. PRODUCT WAS RECEIVED WITHOUT ACCOMPANYING INFORMATION.
30-10-2024 | Malfunction | covidien lp llc north haven | 08-01-2025 | SIGNIA | Complete Loss of Power | A070801 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING TESTING/MAINTENANCE, THE HANDLE DOESN'T CHARGE ONLY LIGHTS UP RED. IT WAS NOTED THAT THE CHARGER WAS WORKING WITH OTHER HANDLE. THERE WAS NO PATIENT INVOLVEMENT. PRODUCT WAS RECEIVED WITHOUT ACCOMPANYING INFORMATION.
01-01-2024 | Injury | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX 45 STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE THE DOCTOR REPORTED BLEEDING. BLEEDING IDENTIFIED IN THE POSTOPERATIVE PERIOD, WITHOUT FURTHER DETAILS.
01-01-2024 | Injury | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX 45 STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE THE DOCTOR REPORTED BLEEDING. BLEEDING IDENTIFIED IN THE POSTOPERATIVE PERIOD, WITHOUT FURTHER DETAILS.
01-01-2024 | Injury | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX 45 STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE THE DOCTOR REPORTED BLEEDING. BLEEDING IDENTIFIED IN THE POSTOPERATIVE PERIOD, WITHOUT FURTHER DETAILS.
01-01-2024 | Injury | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX 45 STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE THE DOCTOR REPORTED BLEEDING. BLEEDING IDENTIFIED IN THE POSTOPERATIVE PERIOD, WITHOUT FURTHER DETAILS.
01-01-2024 | Injury | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX 45 STANDARD | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE THE DOCTOR REPORTED BLEEDING. BLEEDING IDENTIFIED IN THE POSTOPERATIVE PERIOD, WITHOUT FURTHER DETAILS.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE INVOLVING THE SIGNIA ADAPTER, THERE WERE MULTIPLE MALFUNCTIONS RELATED TO THE FIRING MECHANISM. SPECIFICALLY, THE RELOAD WAS PARTIALLY FIRED AND STOPPED 3/4 OF THE WAY THROUGH ON TWO OCCASIONS. THE SURGEON ENCOUNTERED FIRING ISSUES WITH THE ADAPTER, WHICH LED TO PARTIAL FIRING OF THE RELOAD. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE RELOADS AND WAS ABLE TO COMPLETE THE FIRINGS SUCCESSFULLY. THERE WERE NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Separation Failure | A150301 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING THE SIGMOID COLECTOMY WITH ROBOT SUPPORT, WHILE ON COLON ANASTOMOSIS, AFTER FIRING, THE REINFORCEMENT MATERIAL WAS NOT RELEASED. REANASTOMOSIS WAS DONE. ADDITIONAL RESECTION WAS PERFORMED, AND ADDITIONAL SUTURING WAS DONE WITH A STAPLER TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 20 MINUTES AS A RESULT.
16-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING THE SIGMOID COLECTOMY WITH ROBOT SUPPORT, WHILE ON COLON ANASTOMOSIS, AFTER FIRING, THE REINFORCEMENT MATERIAL WAS NOT RELEASED. REANASTOMOSIS WAS DONE. ADDITIONAL RESECTION WAS PERFORMED, AND ADDITIONAL SUTURING WAS DONE WITH A STAPLER TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 20 MINUTES AS A RESULT.
30-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THEN DURING THE SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE WERE BOTH COMPLETE, NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
30-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THEN DURING THE SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE WERE BOTH COMPLETE, NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
30-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | PWRD ECH FLEX 45MM 34CM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THEN DURING THE SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE WERE BOTH COMPLETE, NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
30-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | PWRD ECH FLEX 45MM 34CM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THEN DURING THE SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE WERE BOTH COMPLETE, NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
16-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 09-01-2025 | CAPSURE | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A TRANSABDOMINAL PREPERITONEAL (TAPP) PROCEDURE, WHEN THE SURGEON TRIED TO FIXATE THE 3DMAX LIGHT MESH TO THE ABDOMINAL WALL, THE CAPSURE DEVICE DEPLOYED TWO FASTENERS AT THE SAME TIME. IT WAS REPORTED THAT THE SAME ISSUE OCCURRED TWICE. IT WAS ALSO REPORTED THAT THE PROCEDURE WAS COMPLETED USING SAME DEVICE. THERE WAS NO PATIENT INJURY.
28-10-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS THE DEVICE GOT JAMMED. DURING THE FIRST STAPLING, THERE WAS A SUSPICIOUS NOISE AND THE STAPLING WAS SLOW. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE.
28-10-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS THE DEVICE GOT JAMMED. DURING THE FIRST STAPLING, THERE WAS A SUSPICIOUS NOISE AND THE STAPLING WAS SLOW. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE.
28-10-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS THE DEVICE GOT JAMMED. DURING THE FIRST STAPLING, THERE WAS A SUSPICIOUS NOISE AND THE STAPLING WAS SLOW. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE.
30-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. NOTED THE BATTERY WAS WITH ABNORMAL HIGH TEMPERATURE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
30-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Excessive Heating | A1002 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. NOTED THE BATTERY WAS WITH ABNORMAL HIGH TEMPERATURE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
27-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, THE FIRST STAPLES FIRED AND REMOVED WITH NO ISSUES. HOWEVER, IT WAS NOTED THAT THERE WAS A SMALL CORNER OF BOWEL STILL ATTACHED; AND THEREFORE, A RELOAD FROM THE SAME BATCH WAS INSERTED. FLEX INSERTED THROUGH THE PORT; JAWS REOPENED AND PLACED IN POSITION ON THE BOWEL. HOWEVER, THESE STAPLES DID NOT DEPLOY. OPENING OF THE JAWS OF THE ECHELON WAS ATTEMPTED TO NO AVAIL. THE MANUAL OVERRIDE WAS USED, THIS DID NOTHING. THEN AFTER A BIT OF FIDDLING, THE JAWS FINALLY OPENED. THE REMAINING BOWEL WAS REMOVED USING AN ENDO GIA WITH 45MM RELOAD. BOWEL CHECKED FULLY WITH NO DAMAGE NOTED. THE OPERATIVE PROCEDURE WAS COMPLETED WITH NO FURTHER ISSUES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
27-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION PROCEDURE, THE FIRST STAPLES FIRED AND REMOVED WITH NO ISSUES. HOWEVER, IT WAS NOTED THAT THERE WAS A SMALL CORNER OF BOWEL STILL ATTACHED; AND THEREFORE, A RELOAD FROM THE SAME BATCH WAS INSERTED. FLEX INSERTED THROUGH THE PORT; JAWS REOPENED AND PLACED IN POSITION ON THE BOWEL. HOWEVER, THESE STAPLES DID NOT DEPLOY. OPENING OF THE JAWS OF THE ECHELON WAS ATTEMPTED TO NO AVAIL. THE MANUAL OVERRIDE WAS USED, THIS DID NOTHING. THEN AFTER A BIT OF FIDDLING, THE JAWS FINALLY OPENED. THE REMAINING BOWEL WAS REMOVED USING AN ENDO GIA WITH 45MM RELOAD. BOWEL CHECKED FULLY WITH NO DAMAGE NOTED. THE OPERATIVE PROCEDURE WAS COMPLETED WITH NO FURTHER ISSUES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
27-08-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | GST BLUE RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED THE NEW ONE TO COMPLETE SURGERY. ADDITIONAL INFORMATION REQUESTED.
31-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | WHITE RELOAD FOR ECHELON 60 | Delivered as Unsterile Product | A020701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE PACKAGING WAS DAMAGED ON THE DEVICE BEFORE USE ON THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | ENDO GIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHEN THE FIRING BUTTON WAS PRESSED AGAIN TO CHECK AFTER THE FIRING WAS COMPLETED, THE ARTICULATION STARTED WHILE MAKING AN ABNORMAL SOUND AND IT WAS BENT TO 45. THE PARTS DROPPED IN THE PATIENT'S BODY. ALL PARTS WERE REMOVED AND WAS COLLECTED WITH FORCEPS. PLI 50: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. SOME STAPLE PUSHERS WERE PUSHED BACK TO THE CARTRIDGE. THREE STAPLES REMAINED ON THE CARTRIDGE AND ANVIL SURFACE. THE TWO STAPLES ON THE RELOAD WERE PROPERLY FORMED AND THE ONE WAS OVER CRIMPED.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | ENDO GIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHEN THE FIRING BUTTON WAS PRESSED AGAIN TO CHECK AFTER THE FIRING WAS COMPLETED, THE ARTICULATION STARTED WHILE MAKING AN ABNORMAL SOUND AND IT WAS BENT TO 45. THE PARTS DROPPED IN THE PATIENT'S BODY. ALL PARTS WERE REMOVED AND WAS COLLECTED WITH FORCEPS. PLI 50: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. SOME STAPLE PUSHERS WERE PUSHED BACK TO THE CARTRIDGE. THREE STAPLES REMAINED ON THE CARTRIDGE AND ANVIL SURFACE. THE TWO STAPLES ON THE RELOAD WERE PROPERLY FORMED AND THE ONE WAS OVER CRIMPED.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | ENDO GIA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHEN THE FIRING BUTTON WAS PRESSED AGAIN TO CHECK AFTER THE FIRING WAS COMPLETED, THE ARTICULATION STARTED WHILE MAKING AN ABNORMAL SOUND AND IT WAS BENT TO 45. THE PARTS DROPPED IN THE PATIENT'S BODY. ALL PARTS WERE REMOVED AND WAS COLLECTED WITH FORCEPS. PLI 50: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. SOME STAPLE PUSHERS WERE PUSHED BACK TO THE CARTRIDGE. THREE STAPLES REMAINED ON THE CARTRIDGE AND ANVIL SURFACE. THE TWO STAPLES ON THE RELOAD WERE PROPERLY FORMED AND THE ONE WAS OVER CRIMPED.
04-08-2024 | Injury | us surgical puerto rico | 09-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 31 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK AND ANASTOMOTIC STRICTURE IN 1 PATIENT. ANASTOMOTIC STRICTURES REQUIRED RADIATION EXPOSURE DURING A SWALLOW STUDY AND ONCE OR MORE THAN ONE ENDOSCOPIC BALLOON DILATATIONS.
04-08-2024 | Injury | us surgical puerto rico | 09-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Stenosis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 31 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK AND ANASTOMOTIC STRICTURE IN 1 PATIENT. ANASTOMOTIC STRICTURES REQUIRED RADIATION EXPOSURE DURING A SWALLOW STUDY AND ONCE OR MORE THAN ONE ENDOSCOPIC BALLOON DILATATIONS.
04-08-2024 | Injury | us surgical puerto rico | 09-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 31 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK AND ANASTOMOTIC STRICTURE IN 1 PATIENT. ANASTOMOTIC STRICTURES REQUIRED RADIATION EXPOSURE DURING A SWALLOW STUDY AND ONCE OR MORE THAN ONE ENDOSCOPIC BALLOON DILATATIONS.
04-08-2024 | Injury | us surgical puerto rico | 09-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 25 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK AND ANASTOMOTIC STRICTURE IN 12 PATIENTS. ANASTOMOTIC STRICTURES REQUIRED RADIATION EXPOSURE DURING A SWALLOW STUDY AND ONCE OR MORE THAN ONE ENDOSCOPIC BALLOON DILATATIONS.
04-08-2024 | Injury | us surgical puerto rico | 09-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Stenosis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 25 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK AND ANASTOMOTIC STRICTURE IN 12 PATIENTS. ANASTOMOTIC STRICTURES REQUIRED RADIATION EXPOSURE DURING A SWALLOW STUDY AND ONCE OR MORE THAN ONE ENDOSCOPIC BALLOON DILATATIONS.
04-08-2024 | Injury | us surgical puerto rico | 09-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 25 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK AND ANASTOMOTIC STRICTURE IN 12 PATIENTS. ANASTOMOTIC STRICTURES REQUIRED RADIATION EXPOSURE DURING A SWALLOW STUDY AND ONCE OR MORE THAN ONE ENDOSCOPIC BALLOON DILATATIONS.
04-08-2024 | Injury | us surgical puerto rico | 09-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 28 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK AND ANASTOMOTIC STRICTURE IN 23 PATIENTS. ANASTOMOTIC STRICTURES REQUIRED RADIATION EXPOSURE DURING A SWALLOW STUDY AND ONCE OR MORE THAN ONE ENDOSCOPIC BALLOON DILATATIONS.
04-08-2024 | Injury | us surgical puerto rico | 09-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Stenosis | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 28 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK AND ANASTOMOTIC STRICTURE IN 23 PATIENTS. ANASTOMOTIC STRICTURES REQUIRED RADIATION EXPOSURE DURING A SWALLOW STUDY AND ONCE OR MORE THAN ONE ENDOSCOPIC BALLOON DILATATIONS.
04-08-2024 | Injury | us surgical puerto rico | 09-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 28 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED ANASTOMOTIC LEAK AND ANASTOMOTIC STRICTURE IN 23 PATIENTS. ANASTOMOTIC STRICTURES REQUIRED RADIATION EXPOSURE DURING A SWALLOW STUDY AND ONCE OR MORE THAN ONE ENDOSCOPIC BALLOON DILATATIONS.
24-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BODY DISSECTION PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE WOULD NOT OPEN WHEN TRYING TO CHANGE THE POSITION OF THE CLAMPED POINT. THE OPENING AND CLOSING WERE RETRIED, AND THE KNIFE WAS RETURNED, BUT THE SITUATION COULD NOT BE HANDLED, SO THE JAWS WERE FORCIBLY OPENED WITH FORCEPS, AND THE CLAMPED TISSUE WAS REMOVED. IT WAS CHANGED TO SURE FORM AND HANDLED THE SITUATION AFTERWARDS. SINCE THE CUT LINE WAS ENLARGED BY ABOUT 1 CENTIMETER, THERE WAS AN EFFECT ON THE PATIENT SLIGHTLY. WHEN INTERVIEWED THE SURGEON SAID THAT THE HANDLE HAD BEEN OPENED FULLY, BUT IT COULD HAVE BEEN SOLVED IF IT HAD BEEN MANUALLY OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BODY DISSECTION PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE WOULD NOT OPEN WHEN TRYING TO CHANGE THE POSITION OF THE CLAMPED POINT. THE OPENING AND CLOSING WERE RETRIED, AND THE KNIFE WAS RETURNED, BUT THE SITUATION COULD NOT BE HANDLED, SO THE JAWS WERE FORCIBLY OPENED WITH FORCEPS, AND THE CLAMPED TISSUE WAS REMOVED. IT WAS CHANGED TO SURE FORM AND HANDLED THE SITUATION AFTERWARDS. SINCE THE CUT LINE WAS ENLARGED BY ABOUT 1 CENTIMETER, THERE WAS AN EFFECT ON THE PATIENT SLIGHTLY. WHEN INTERVIEWED THE SURGEON SAID THAT THE HANDLE HAD BEEN OPENED FULLY, BUT IT COULD HAVE BEEN SOLVED IF IT HAD BEEN MANUALLY OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON 3000 60MM LONG | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BODY DISSECTION PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE WOULD NOT OPEN WHEN TRYING TO CHANGE THE POSITION OF THE CLAMPED POINT. THE OPENING AND CLOSING WERE RETRIED, AND THE KNIFE WAS RETURNED, BUT THE SITUATION COULD NOT BE HANDLED, SO THE JAWS WERE FORCIBLY OPENED WITH FORCEPS, AND THE CLAMPED TISSUE WAS REMOVED. IT WAS CHANGED TO SURE FORM AND HANDLED THE SITUATION AFTERWARDS. SINCE THE CUT LINE WAS ENLARGED BY ABOUT 1 CENTIMETER, THERE WAS AN EFFECT ON THE PATIENT SLIGHTLY. WHEN INTERVIEWED THE SURGEON SAID THAT THE HANDLE HAD BEEN OPENED FULLY, BUT IT COULD HAVE BEEN SOLVED IF IT HAD BEEN MANUALLY OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON 3000 60MM LONG | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BODY DISSECTION PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE WOULD NOT OPEN WHEN TRYING TO CHANGE THE POSITION OF THE CLAMPED POINT. THE OPENING AND CLOSING WERE RETRIED, AND THE KNIFE WAS RETURNED, BUT THE SITUATION COULD NOT BE HANDLED, SO THE JAWS WERE FORCIBLY OPENED WITH FORCEPS, AND THE CLAMPED TISSUE WAS REMOVED. IT WAS CHANGED TO SURE FORM AND HANDLED THE SITUATION AFTERWARDS. SINCE THE CUT LINE WAS ENLARGED BY ABOUT 1 CENTIMETER, THERE WAS AN EFFECT ON THE PATIENT SLIGHTLY. WHEN INTERVIEWED THE SURGEON SAID THAT THE HANDLE HAD BEEN OPENED FULLY, BUT IT COULD HAVE BEEN SOLVED IF IT HAD BEEN MANUALLY OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON 3000 60MM LONG | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BODY DISSECTION PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE WOULD NOT OPEN WHEN TRYING TO CHANGE THE POSITION OF THE CLAMPED POINT. THE OPENING AND CLOSING WERE RETRIED, AND THE KNIFE WAS RETURNED, BUT THE SITUATION COULD NOT BE HANDLED, SO THE JAWS WERE FORCIBLY OPENED WITH FORCEPS, AND THE CLAMPED TISSUE WAS REMOVED. IT WAS CHANGED TO SURE FORM AND HANDLED THE SITUATION AFTERWARDS. SINCE THE CUT LINE WAS ENLARGED BY ABOUT 1 CENTIMETER, THERE WAS AN EFFECT ON THE PATIENT SLIGHTLY. WHEN INTERVIEWED THE SURGEON SAID THAT THE HANDLE HAD BEEN OPENED FULLY, BUT IT COULD HAVE BEEN SOLVED IF IT HAD BEEN MANUALLY OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON 3000 60MM LONG | Adverse Event Without Identified Device or Use Problem | A24 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BODY DISSECTION PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE WOULD NOT OPEN WHEN TRYING TO CHANGE THE POSITION OF THE CLAMPED POINT. THE OPENING AND CLOSING WERE RETRIED, AND THE KNIFE WAS RETURNED, BUT THE SITUATION COULD NOT BE HANDLED, SO THE JAWS WERE FORCIBLY OPENED WITH FORCEPS, AND THE CLAMPED TISSUE WAS REMOVED. IT WAS CHANGED TO SURE FORM AND HANDLED THE SITUATION AFTERWARDS. SINCE THE CUT LINE WAS ENLARGED BY ABOUT 1 CENTIMETER, THERE WAS AN EFFECT ON THE PATIENT SLIGHTLY. WHEN INTERVIEWED THE SURGEON SAID THAT THE HANDLE HAD BEEN OPENED FULLY, BUT IT COULD HAVE BEEN SOLVED IF IT HAD BEEN MANUALLY OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON 3000 60MM LONG | No Apparent Adverse Event | A25 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BODY DISSECTION PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE WOULD NOT OPEN WHEN TRYING TO CHANGE THE POSITION OF THE CLAMPED POINT. THE OPENING AND CLOSING WERE RETRIED, AND THE KNIFE WAS RETURNED, BUT THE SITUATION COULD NOT BE HANDLED, SO THE JAWS WERE FORCIBLY OPENED WITH FORCEPS, AND THE CLAMPED TISSUE WAS REMOVED. IT WAS CHANGED TO SURE FORM AND HANDLED THE SITUATION AFTERWARDS. SINCE THE CUT LINE WAS ENLARGED BY ABOUT 1 CENTIMETER, THERE WAS AN EFFECT ON THE PATIENT SLIGHTLY. WHEN INTERVIEWED THE SURGEON SAID THAT THE HANDLE HAD BEEN OPENED FULLY, BUT IT COULD HAVE BEEN SOLVED IF IT HAD BEEN MANUALLY OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON 3000 60MM LONG | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BODY DISSECTION PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE WOULD NOT OPEN WHEN TRYING TO CHANGE THE POSITION OF THE CLAMPED POINT. THE OPENING AND CLOSING WERE RETRIED, AND THE KNIFE WAS RETURNED, BUT THE SITUATION COULD NOT BE HANDLED, SO THE JAWS WERE FORCIBLY OPENED WITH FORCEPS, AND THE CLAMPED TISSUE WAS REMOVED. IT WAS CHANGED TO SURE FORM AND HANDLED THE SITUATION AFTERWARDS. SINCE THE CUT LINE WAS ENLARGED BY ABOUT 1 CENTIMETER, THERE WAS AN EFFECT ON THE PATIENT SLIGHTLY. WHEN INTERVIEWED THE SURGEON SAID THAT THE HANDLE HAD BEEN OPENED FULLY, BUT IT COULD HAVE BEEN SOLVED IF IT HAD BEEN MANUALLY OPENED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. ADDITIONAL INFORMATION REQUESTED.
26-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE MALFUNCTIONED AND DID NOT FIRE BLUE LOAD MID-PROCEDURE. THE SURGICAL TECH ATTEMPTED TO RESOLVE THE ISSUE BY MANUALLY ATTEMPTING TO OPEN THE JAWS OF THE STAPLER, BUT THEY WOULD NOT OPEN. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCES NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
26-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE MALFUNCTIONED AND DID NOT FIRE BLUE LOAD MID-PROCEDURE. THE SURGICAL TECH ATTEMPTED TO RESOLVE THE ISSUE BY MANUALLY ATTEMPTING TO OPEN THE JAWS OF THE STAPLER, BUT THEY WOULD NOT OPEN. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCES NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
16-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RUL (RIGHT UPPER LOBECTOMY) PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT FIRE WITH ITS REINFORCEMENT. THE FIRST FIRE WITH THE DEVICE STOPPED IN THE BEGINNING OF THE FIRE SEQUENCE; AND THEY HAD TO USE THE MANUAL OVERRIDE TO UNLOCK THE DEVICE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
16-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RUL (RIGHT UPPER LOBECTOMY) PROCEDURE, IT WAS OBSERVED THAT THE DEVICE DID NOT FIRE WITH ITS REINFORCEMENT. THE FIRST FIRE WITH THE DEVICE STOPPED IN THE BEGINNING OF THE FIRE SEQUENCE; AND THEY HAD TO USE THE MANUAL OVERRIDE TO UNLOCK THE DEVICE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
02-01-2025 | Malfunction | ethicon endo-surgery | 09-01-2025 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE AND NOTED THE STAPLES MALFORMED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
02-01-2025 | Malfunction | ethicon endo-surgery | 09-01-2025 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE AND NOTED THE STAPLES MALFORMED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
12-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INCISIONAL HERNIA REPAIR, THE TACKS DID NOT FIXATE THE MESH ON THE ABDOMINAL WALL AND FELL INTO THE PATIENT'S CAVITY. THE TACK WAS RETRIEVED AND REMOVED THROUGH THE MAIN PORT. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC INCISIONAL HERNIA REPAIR, THE TACKS DID NOT FIXATE THE MESH ON THE ABDOMINAL WALL AND FELL INTO THE PATIENT'S CAVITY. THE TACK WAS RETRIEVED AND REMOVED THROUGH THE MAIN PORT. THERE WAS NO PATIENT INJURY.
28-11-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE MATERIAL WAS DEFECTIVE. THERE WAS NO SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
24-12-2024 | Malfunction | johnson & johnson medical (suzhou | 09-01-2025 | FINAL ASSY 445MM STD DEVICE | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A PANCREATECTOMY PROCEDURE, OOZING WAS OBSERVED AFTER FIRING WHEN BOTH STAPLE LINE AND CUT LINE WERE BOTH COMPLETED. SUTURE WAS USED TO OVERSEW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | johnson & johnson medical (suzhou | 09-01-2025 | FINAL ASSY 445MM STD DEVICE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PANCREATECTOMY PROCEDURE, OOZING WAS OBSERVED AFTER FIRING WHEN BOTH STAPLE LINE AND CUT LINE WERE BOTH COMPLETED. SUTURE WAS USED TO OVERSEW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | johnson & johnson medical (suzhou | 09-01-2025 | FINAL ASSY 445MM STD DEVICE | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A PANCREATECTOMY PROCEDURE, OOZING WAS OBSERVED AFTER FIRING WHEN BOTH STAPLE LINE AND CUT LINE WERE BOTH COMPLETED. SUTURE WAS USED TO OVERSEW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
24-12-2024 | Malfunction | johnson & johnson medical (suzhou | 09-01-2025 | FINAL ASSY 445MM STD DEVICE | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PANCREATECTOMY PROCEDURE, OOZING WAS OBSERVED AFTER FIRING WHEN BOTH STAPLE LINE AND CUT LINE WERE BOTH COMPLETED. SUTURE WAS USED TO OVERSEW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
17-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | EEA | Difficult to Remove | A150207 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y GASTRIC BYPASS (RYGB), IT WAS VERY DIFFICULT TO REMOVE THE STAPLER FROM THE ANASTOMOSIS. THE SURGEON NEEDED TO USE GRASPERS AND AN ASSISTANT TO HELP HOLD THE TISSUE FORWARD, AS WELL AS PULL AND WIGGLE IT OUT, SO THAT THE USER COULD REMOVE THE DEVICE, WHICH WAS PUTTING A LOT OF STRESS ON THE ANASTOMOSIS. A POSITIVE AIR LEAK TEST WAS NOTED. TO RESOLVE THE ISSUE, SUTURING WAS DONE TO THE PATIENT WHICH CAUSES AN EXTRA 5-10 MINUTE SURGICAL TIME.
17-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | EEA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y GASTRIC BYPASS (RYGB), IT WAS VERY DIFFICULT TO REMOVE THE STAPLER FROM THE ANASTOMOSIS. THE SURGEON NEEDED TO USE GRASPERS AND AN ASSISTANT TO HELP HOLD THE TISSUE FORWARD, AS WELL AS PULL AND WIGGLE IT OUT, SO THAT THE USER COULD REMOVE THE DEVICE, WHICH WAS PUTTING A LOT OF STRESS ON THE ANASTOMOSIS. A POSITIVE AIR LEAK TEST WAS NOTED. TO RESOLVE THE ISSUE, SUTURING WAS DONE TO THE PATIENT WHICH CAUSES AN EXTRA 5-10 MINUTE SURGICAL TIME.
29-01-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE HANDLE, THE HANDLE SEIZING AFTER FOUR ATTEMPTS TO RELOAD IT. THIS ISSUE WAS IDENTIFIED AS A FIRING PROBLEM WITH THE MANUAL HANDLE. THE HANDLE WAS SUBSEQUENTLY REPLACED WITH ANOTHER OF THE SAME REFERENCE AND LOT NUMBER RESULTED IN LOSS OF TIME. THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
01-01-2024 | Injury | ethicon endo-surgery | 09-01-2025 | PROXIMATE*PPH PROCEDURE SET | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDOPEXY PROCEDURE THE DEVICE FAILED ON A PATIENT WHO WAS TAKEN BACK TO THEATRE. THE PROCEDURE WAS DONE AND PATIENT WAS DISCHARGED HOME. PATIENT CAME BACK WITH PER-RECTAL BLEEDING, RETURNED BACK TO THEATRE FOR AN EXAMINATION UNDER ANAESTHESIA, THE SURGEON STATED THAT HE COULD CLEARLY SEE THAT THE STAPLES HAD NOT COMPLETELY FORMED AND SOME HAD GIVEN WAY UPON REDO PROCTOSCOPY.
01-01-2024 | Injury | ethicon endo-surgery | 09-01-2025 | PROXIMATE*PPH PROCEDURE SET | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDOPEXY PROCEDURE THE DEVICE FAILED ON A PATIENT WHO WAS TAKEN BACK TO THEATRE. THE PROCEDURE WAS DONE AND PATIENT WAS DISCHARGED HOME. PATIENT CAME BACK WITH PER-RECTAL BLEEDING, RETURNED BACK TO THEATRE FOR AN EXAMINATION UNDER ANAESTHESIA, THE SURGEON STATED THAT HE COULD CLEARLY SEE THAT THE STAPLES HAD NOT COMPLETELY FORMED AND SOME HAD GIVEN WAY UPON REDO PROCTOSCOPY.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY RESECTION, THE FIRING COULD NOT BE PERFORMED AT ALL. ANOTHER CARTRIDGE WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien | 09-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLE DID NOT FORM. IT WAS NOTED THAT TWO RELOADS WERE USED AND HAD THE SAME ISSUE. SUTURING WAS DONE TO RESOLVE THE ISSUE.
16-12-2024 | Malfunction | covidien | 09-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLE DID NOT FORM. IT WAS NOTED THAT TWO RELOADS WERE USED AND HAD THE SAME ISSUE. SUTURING WAS DONE TO RESOLVE THE ISSUE.
13-12-2024 | Injury | intuitive surgical | 09-01-2025 | SUREFORM | Unintended Movement | A0512 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED TRANSTHORACIC ESOPHAGECTOMY WITH NECK ANASTOMOSIS, BLEEDING OCCURRED AND THE PROCEDURE WAS CONVERTED TO OPEN. A SUREFORM 45 STAPLER INSTRUMENT WAS INSERTED INTO THE PATIENT AND IT MOVED IN AN UNINTENDED DIRECTION, THEN DID NOT MOVE. WHILE INSERTING THE INSTRUMENT, A YELLOW LIGHT CAME ON AND THE INSTRUMENT DID NOT MOVE AS INTENDED. THE CUSTOMER HAD DIFFICULTY REMOVING THE INSTRUMENT AND ATTEMPTED TO FORCIBLY REMOVE IT FROM THE PATIENT AND TROCAR. THE THYROID ARTERY WAS DAMAGED WHILE ATTEMPTING TO REMOVE THE INSTRUMENT CAUSING A LARGE AMOUNT OF BLEEDING. THE STAPLER INSTRUMENT WAS REMOVED WITH THE TROCAR FROM THE PATIENT. THE STAPLER WRIST WAS TORN WHILE REMOVING THE INSTRUMENT. THE CUSTOMER ASSUMES THE INSTRUMENT WRIST DID NOT STRAIGHTEN DURING THIS EVENT. THE DA VINCI WAS UNDOCKED AND THE PROCEDURE WAS CONVERTED TO OPEN. HEMOSTASIS OF THE BLEEDING WAS ACHIEVED WITH COMPRESSION AND UNSPECIFIED REPAIR. THE PROCEDURE WAS CONTINUED OPEN AFTER THE BLEEDING WAS CONTROLLED. THE PATIENT IS RECOVERING WELL, AND THEIR CARE PLAN WAS NOT CHANGED DUE TO THIS EVENT. THE NURSE DID NOT KNOW IF THE SUREFORM 45 STAPLER WAS COLLIDING WITH ANOTHER INSTRUMENT DURING THIS EVENT. THE NURSE DID NOT KNOW IF THE DA VINCI SYSTEM ARMS WERE CHECKING FOR COLLISIONS AT THE TIME OF THIS EVENT. THE NURSE DID NOT KNOW ABOUT HOW FAR INTO THE PROCEDURE THIS EVENT OCCURRED.
13-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON HOWEVER THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE FIRING ISSUES WITH THE HANDLE AND RELOAD DURING A VATS LOBECTOMY PROCEDURE. THE INSTRUMENTS DID NOT FIRE, AND ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON, USER WERE UNABLE TO SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW HANDLE AND RELOAD. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT.
09-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE FIRING ISSUES WITH THE HANDLE AND RELOAD DURING A VATS LOBECTOMY PROCEDURE. THE INSTRUMENTS DID NOT FIRE, AND ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON, USER WERE UNABLE TO SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW HANDLE AND RELOAD. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT.
12-12-2024 | Injury | us surgical puerto rico | 09-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Abdominal Pain | Event Description: IT WAS REPORTED THAT AFTER A PROCEDURE INVOLVING A SINGLE USE LOADING UNIT, AN INCONSISTENT STAPLE LINE WAS DETECTED, NECESSITATING ADDITIONAL SURGERY. ANOTHER DEVICE FROM A DIFFERENT MANUFACTURER WAS USED TO COMPLETE THE CASE. IT WAS NOTED THAT THE PATIENT WAS CURRENTLY EXTUBATED, STILL IN THE ICU FOR REASONS OF HEMODYNAMIC SURVEILLANCE, ALREADY ACCEPTING AN ENTERAL DIET. ACCORDING TO THE SURGEON'S REPORT, THIS PATIENT WILL LATER NEED A NEW SURGERY TO RECONSTRUCT THE INTESTINAL TRACT, RESECTION OF ILEOCOLONIC ANASTOMOSIS AND CLOSURE OF A DOUBLE-MOUTH ILEOSTOMY. ON THE FOURTH DAY AFTER SURGERY, THE PATIENT BEGAN TO EXPERIENCE ABDOMINAL PAIN, ABDOMINAL DISTENSION AND WORSENING OF THE LEUKOGRAM. ON THE FIFTH DAY AFTER SURGERY, THE PATIENT UNDERWENT A TOMOGRAPHY, WHICH SHOWED A FREE PNEUMOPERITONEUM. IT WAS THEN DECIDED TO RE-APPROACH, WHEREAS SHOWN IN THE VIDEO SENT, THE CLAMP LINE WAS OPEN. BETWEEN THE ILEUM AND THE COLON IT WAS FRAYED.
12-12-2024 | Injury | us surgical puerto rico | 09-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT AFTER A PROCEDURE INVOLVING A SINGLE USE LOADING UNIT, AN INCONSISTENT STAPLE LINE WAS DETECTED, NECESSITATING ADDITIONAL SURGERY. ANOTHER DEVICE FROM A DIFFERENT MANUFACTURER WAS USED TO COMPLETE THE CASE. IT WAS NOTED THAT THE PATIENT WAS CURRENTLY EXTUBATED, STILL IN THE ICU FOR REASONS OF HEMODYNAMIC SURVEILLANCE, ALREADY ACCEPTING AN ENTERAL DIET. ACCORDING TO THE SURGEON'S REPORT, THIS PATIENT WILL LATER NEED A NEW SURGERY TO RECONSTRUCT THE INTESTINAL TRACT, RESECTION OF ILEOCOLONIC ANASTOMOSIS AND CLOSURE OF A DOUBLE-MOUTH ILEOSTOMY. ON THE FOURTH DAY AFTER SURGERY, THE PATIENT BEGAN TO EXPERIENCE ABDOMINAL PAIN, ABDOMINAL DISTENSION AND WORSENING OF THE LEUKOGRAM. ON THE FIFTH DAY AFTER SURGERY, THE PATIENT UNDERWENT A TOMOGRAPHY, WHICH SHOWED A FREE PNEUMOPERITONEUM. IT WAS THEN DECIDED TO RE-APPROACH, WHEREAS SHOWN IN THE VIDEO SENT, THE CLAMP LINE WAS OPEN. BETWEEN THE ILEUM AND THE COLON IT WAS FRAYED.
20-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 09-01-2025 | CAPSURE | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A TRANSABDOMINAL PREPERITONEAL (TAPP) PROCEDURE, THE CAPSURE DEVICE DEPLOYED TWO FASTENERS AT A TIME. IT WAS REPORTED THAT THOSE FASTENERS WERE REMOVED, AND A SECOND DEVICE WAS USED TO COMPLETE THE PROCEDURE WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 09-01-2025 | CAPSURE | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC INGUINAL HERNIA REPAIR, THE CAPSURE DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT WHEN THE SHAFT WAS INSERTED INTO THE PATIENT'S BODY THROUGH THE TROCAR AND ATTEMPTED TO DEPLOY, THE DEVICE DEPLOYED TWO FASTENERS STUCK TOGETHER AND OVERLAPPED. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien | 09-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING GASTRIC BYPASS, THE INSTRUMENT DID NOT FIRE DESPITE THE SURGEON BEING ABLE TO PRESS THE GREEN SAFETY BUTTON AND SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW RELOAD. THERE WAS NO PATIENT INJURY.
01-01-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE JAMMED. THERE WAS DIFFICULTY OPENING THE DEVICE. THE DEVICE WAS OPENED MANUALLY.
17-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A DISTAL GASTRECTOMY PROCEDURE, TWO STAPLERS, THE MANUAL HANDLE AND RELOAD, WERE UNABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE OF BOTH DEVICES, ALTHOUGH THE GREEN BUTTON COULD BE PRESSED. THE HANDLE AND RELOAD WERE REPLACED. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DISTAL GASTRECTOMY PROCEDURE, TWO STAPLERS, THE MANUAL HANDLE AND RELOAD, WERE UNABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE OF BOTH DEVICES, ALTHOUGH THE GREEN BUTTON COULD BE PRESSED. THE HANDLE AND RELOAD WERE REPLACED. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A DISTAL GASTRECTOMY PROCEDURE, TWO STAPLERS, THE MANUAL HANDLE AND RELOAD, WERE UNABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE OF BOTH DEVICES, ALTHOUGH THE GREEN BUTTON COULD BE PRESSED. THE HANDLE AND RELOAD WERE REPLACED. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DISTAL GASTRECTOMY PROCEDURE, TWO STAPLERS, THE MANUAL HANDLE AND RELOAD, WERE UNABLE TO FIRE. THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE OF BOTH DEVICES, ALTHOUGH THE GREEN BUTTON COULD BE PRESSED. THE HANDLE AND RELOAD WERE REPLACED. THERE WAS NO PATIENT INJURY.
04-12-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY BY VATS, THE DEVICE WAS LOCKED OUT AT THE 1ST STROKE OF THE 1ST FIRING DURING BLOOD VESSEL PROCESSING. THE SLEDS MAY HAVE MOVED WHEN LOADING THE CARTRIDGE. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE INVOLVING THE SIGNIA ADAPTER, THERE WERE MULTIPLE MALFUNCTIONS RELATED TO THE FIRING MECHANISM. SPECIFICALLY, THE RELOAD WAS PARTIALLY FIRED AND STOPPED 3/4 OF THE WAY THROUGH ON TWO OCCASIONS. THE SURGEON ENCOUNTERED FIRING ISSUES WITH THE ADAPTER, WHICH LED TO PARTIAL FIRING OF THE RELOAD. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE RELOADS AND WAS ABLE TO COMPLETE THE FIRINGS SUCCESSFULLY. THERE WERE NO PATIENT INJURY.
04-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE MANUAL HANDLE, THE INSTRUMENT DID NOT FIRE AFTER MULTIPLE ATTEMPTS. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED, AND A NEW RELOAD WAS OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
04-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE MANUAL HANDLE, THE INSTRUMENT DID NOT FIRE AFTER MULTIPLE ATTEMPTS. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. TO RESOLVE THE ISSUE, THE DEVICE WAS REPLACED, AND A NEW RELOAD WAS OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE INVOLVING THE SIGNIA ADAPTER, THERE WERE MULTIPLE MALFUNCTIONS RELATED TO THE FIRING MECHANISM. SPECIFICALLY, THE RELOAD WAS PARTIALLY FIRED AND STOPPED 3/4 OF THE WAY THROUGH ON TWO OCCASIONS. THE SURGEON ENCOUNTERED FIRING ISSUES WITH THE ADAPTER, WHICH LED TO PARTIAL FIRING OF THE RELOAD. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE RELOADS AND WAS ABLE TO COMPLETE THE FIRINGS SUCCESSFULLY. THERE WERE NO PATIENT INJURY.
01-01-2024 | Death | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN LIVER PROCEDURE THAT ON THE SECOND FIRING NEAR THE PORTAL VEIN THE DEVICE MISFIRED. WHEN THE DEVICE WAS REMOVED THE TISSUE STARTED BLEEDING PROFUSELY AND PATIENT DIED 2HR LATER. DURING THE PROCEDURE THE PATIENT WAS GIVEN 4 ROUNDS OF MTP.
01-01-2024 | Death | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN LIVER PROCEDURE THAT ON THE SECOND FIRING NEAR THE PORTAL VEIN THE DEVICE MISFIRED. WHEN THE DEVICE WAS REMOVED THE TISSUE STARTED BLEEDING PROFUSELY AND PATIENT DIED 2HR LATER. DURING THE PROCEDURE THE PATIENT WAS GIVEN 4 ROUNDS OF MTP.
01-01-2024 | Death | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Insufficient Information | E2401 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN LIVER PROCEDURE THAT ON THE SECOND FIRING NEAR THE PORTAL VEIN THE DEVICE MISFIRED. WHEN THE DEVICE WAS REMOVED THE TISSUE STARTED BLEEDING PROFUSELY AND PATIENT DIED 2HR LATER. DURING THE PROCEDURE THE PATIENT WAS GIVEN 4 ROUNDS OF MTP.
01-01-2024 | Death | ethicon endo-surgery | 09-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN LIVER PROCEDURE THAT ON THE SECOND FIRING NEAR THE PORTAL VEIN THE DEVICE MISFIRED. WHEN THE DEVICE WAS REMOVED THE TISSUE STARTED BLEEDING PROFUSELY AND PATIENT DIED 2HR LATER. DURING THE PROCEDURE THE PATIENT WAS GIVEN 4 ROUNDS OF MTP.
25-11-2024 | Malfunction | ethicon endo-surgery | 09-01-2025 | ECHELON 3000 60MM STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, THE BATTERY WAS NOT ACTIVATED DURING USE, AND COULD NOT BE USED. THE CARTRIDGE WAS REMOVED, AND A REPLACEMENT WAS SET AND USED. SINCE THEN, THE SURGERY WAS COMPLETED WITHOUT ANY PROBLEMS AND THE PATIENT WAS NOT AFFECTED. IT SEEMS THAT THE BATTERY WAS NOT RELOADED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
15-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION, THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED WITH THE SAME RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA | Break | A0401 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC RADICAL SURGERY FOR LUNG CANCER, THE SURGEON PREPARED TO USE THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD FOR INTRAOPERATIVE CUTTING, ANASTOMOSIS, AND HEMOSTASIS. AFTER THE HANDWASHING, THE NURSE ROUTINELY CHECKED THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD AND CONFIRMED THAT IT WAS INTACT AND HANDED IT TO THE SURGEON. WHEN THE SURGEON USED THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD TO CUT THE PATIENT'S LUNG TISSUE, IT WAS FOUND THAT THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD COULD NOT CUT THEN THE STAPLES DID NOT DEPLOY AT ALL. THE STAPLER WAS REMOVED FROM THE CHEST CAVITY, AND IT WAS FOUND THAT THE CONNECTION BETWEEN THE RELOAD AND THE STAPLER HANDLE WAS LOOSE AND THE RELOAD WAS DIFFICULT TO REMOVE. AFTER REPEATED ATTEMPTS, THE RELOAD WAS FINALLY REMOVED AND FOUND TO BE BROKEN WITH INTACT EDGES. THE OPERATION TIME WAS EXTENDED BY ABOUT 5 MINUTES. A NEW RELOAD WITH THE SAME HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS INCREASED TISSUE DAMAGE.
09-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA | Device Difficult to Setup or Prepare | A2201 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC RADICAL SURGERY FOR LUNG CANCER, THE SURGEON PREPARED TO USE THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD FOR INTRAOPERATIVE CUTTING, ANASTOMOSIS, AND HEMOSTASIS. AFTER THE HANDWASHING, THE NURSE ROUTINELY CHECKED THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD AND CONFIRMED THAT IT WAS INTACT AND HANDED IT TO THE SURGEON. WHEN THE SURGEON USED THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD TO CUT THE PATIENT'S LUNG TISSUE, IT WAS FOUND THAT THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD COULD NOT CUT THEN THE STAPLES DID NOT DEPLOY AT ALL. THE STAPLER WAS REMOVED FROM THE CHEST CAVITY, AND IT WAS FOUND THAT THE CONNECTION BETWEEN THE RELOAD AND THE STAPLER HANDLE WAS LOOSE AND THE RELOAD WAS DIFFICULT TO REMOVE. AFTER REPEATED ATTEMPTS, THE RELOAD WAS FINALLY REMOVED AND FOUND TO BE BROKEN WITH INTACT EDGES. THE OPERATION TIME WAS EXTENDED BY ABOUT 5 MINUTES. A NEW RELOAD WITH THE SAME HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS INCREASED TISSUE DAMAGE.
09-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC RADICAL SURGERY FOR LUNG CANCER, THE SURGEON PREPARED TO USE THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD FOR INTRAOPERATIVE CUTTING, ANASTOMOSIS, AND HEMOSTASIS. AFTER THE HANDWASHING, THE NURSE ROUTINELY CHECKED THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD AND CONFIRMED THAT IT WAS INTACT AND HANDED IT TO THE SURGEON. WHEN THE SURGEON USED THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD TO CUT THE PATIENT'S LUNG TISSUE, IT WAS FOUND THAT THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD COULD NOT CUT THEN THE STAPLES DID NOT DEPLOY AT ALL. THE STAPLER WAS REMOVED FROM THE CHEST CAVITY, AND IT WAS FOUND THAT THE CONNECTION BETWEEN THE RELOAD AND THE STAPLER HANDLE WAS LOOSE AND THE RELOAD WAS DIFFICULT TO REMOVE. AFTER REPEATED ATTEMPTS, THE RELOAD WAS FINALLY REMOVED AND FOUND TO BE BROKEN WITH INTACT EDGES. THE OPERATION TIME WAS EXTENDED BY ABOUT 5 MINUTES. A NEW RELOAD WITH THE SAME HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS INCREASED TISSUE DAMAGE.
09-12-2024 | Malfunction | us surgical puerto rico | 09-01-2025 | ENDO GIA | Human-Device Interface Problem | A22 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THORACOSCOPIC RADICAL SURGERY FOR LUNG CANCER, THE SURGEON PREPARED TO USE THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD FOR INTRAOPERATIVE CUTTING, ANASTOMOSIS, AND HEMOSTASIS. AFTER THE HANDWASHING, THE NURSE ROUTINELY CHECKED THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD AND CONFIRMED THAT IT WAS INTACT AND HANDED IT TO THE SURGEON. WHEN THE SURGEON USED THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD TO CUT THE PATIENT'S LUNG TISSUE, IT WAS FOUND THAT THE LAPAROSCOPIC CUTTING STAPLER AND DISPOSABLE RELOAD COULD NOT CUT THEN THE STAPLES DID NOT DEPLOY AT ALL. THE STAPLER WAS REMOVED FROM THE CHEST CAVITY, AND IT WAS FOUND THAT THE CONNECTION BETWEEN THE RELOAD AND THE STAPLER HANDLE WAS LOOSE AND THE RELOAD WAS DIFFICULT TO REMOVE. AFTER REPEATED ATTEMPTS, THE RELOAD WAS FINALLY REMOVED AND FOUND TO BE BROKEN WITH INTACT EDGES. THE OPERATION TIME WAS EXTENDED BY ABOUT 5 MINUTES. A NEW RELOAD WITH THE SAME HANDLE WERE USED TO RESOLVE THE ISSUE. THERE WAS INCREASED TISSUE DAMAGE.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE INVOLVING THE SIGNIA ADAPTER, THERE WERE MULTIPLE MALFUNCTIONS RELATED TO THE FIRING MECHANISM. SPECIFICALLY, THE RELOAD WAS PARTIALLY FIRED AND STOPPED 3/4 OF THE WAY THROUGH ON TWO OCCASIONS. THE SURGEON ENCOUNTERED FIRING ISSUES WITH THE ADAPTER, WHICH LED TO PARTIAL FIRING OF THE RELOAD. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE RELOADS AND WAS ABLE TO COMPLETE THE FIRINGS SUCCESSFULLY. THERE WERE NO PATIENT INJURY.
04-11-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, ARTICULATION JOINT OF THE RELOAD BROKE DURING RETRACTION AFTER STAPLING. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE JAWS WERE OPEN.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | ABSORBATACK | Positioning Failure | A150201 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC BILATERAL TRANSABDOMINAL PREPERITONEAL HERNIA CASE (TAPP), USING THE DEVICE, THE TACKER M ALFUNCTIONED, AND TACKS WERE NOT BEING RELEASED. THIS OCCURRED WHILE TACKING THE FLAP AFTER THE INSERTION OF THE MESH. AS A RESULT, ABSORBABLE TACKS FELL INTO THE CAVITY OF THE PATIENT, AND THE FLAP TISSUE WAS TORN. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT ISSUE. THE ISSUE WAS RESOLVED BY COMPLETING THE PROCEDURE WITH SUTURING, AND NO ADDITIONAL OPERATION WAS PLANNED.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | ABSORBATACK | Detachment of Device or Device Component | A0501 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC BILATERAL TRANSABDOMINAL PREPERITONEAL HERNIA CASE (TAPP), USING THE DEVICE, THE TACKER M ALFUNCTIONED, AND TACKS WERE NOT BEING RELEASED. THIS OCCURRED WHILE TACKING THE FLAP AFTER THE INSERTION OF THE MESH. AS A RESULT, ABSORBABLE TACKS FELL INTO THE CAVITY OF THE PATIENT, AND THE FLAP TISSUE WAS TORN. THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES DUE TO THE PRODUCT ISSUE. THE ISSUE WAS RESOLVED BY COMPLETING THE PROCEDURE WITH SUTURING, AND NO ADDITIONAL OPERATION WAS PLANNED.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHEN THE FIRING BUTTON WAS PRESSED AGAIN TO CHECK AFTER THE FIRING WAS COMPLETED, THE ARTICULATION STARTED WHILE MAKING AN ABNORMAL SOUND, AND IT WAS BENT TO 45. THE PARTS DROPPED IN THE PATIENT'S BODY. ALL PARTS WERE REMOVED AND WAS COLLECTED WITH FORCEPS. PLI 50: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. SOME STAPLE PUSHERS WERE PUSHED BACK TO THE CARTRIDGE. THREE STAPLES REMAINED ON THE CARTRIDGE AND ANVIL SURFACE. THE TWO STAPLES ON THE RELOAD WERE PROPERLY FORMED AND THE ONE WAS OVER CRIMPED.
25-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED PANCREATICODUODENECTOMY (WHIPPLE PROCEDURE), IN GASTRIC RESECTION, WHILE FIRING THE REINFORCED BLACK 60 MILLIMETER (MM) RELOAD, THE FIRING STOPPED DUE TO EXCESSIVE FORCE DETECTED. A NEW NON-REINFORCED BLACK RELOAD OF THE SAME LENGTH WAS USED, BUT IT ALSO STOPPED FIRING MIDWAY WITH THE SAME EXCESSIVE FORCE. TO RESOLVE THE ISSUE, A NEW HANDLE, A DAPTER, AND BLACK RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED PANCREATICODUODENECTOMY (WHIPPLE PROCEDURE), IN GASTRIC RESECTION, WHILE FIRING THE REINFORCED BLACK 60 MILLIMETER (MM) RELOAD, THE FIRING STOPPED DUE TO EXCESSIVE FORCE DETECTED. A NEW NON-REINFORCED BLACK RELOAD OF THE SAME LENGTH WAS USED, BUT IT ALSO STOPPED FIRING MIDWAY WITH THE SAME EXCESSIVE FORCE. TO RESOLVE THE ISSUE, A NEW HANDLE, A DAPTER, AND BLACK RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN GASTRECTOMY, WHEN THE FIRING BUTTON WAS PRESSED AGAIN TO CHECK AFTER THE FIRING WAS COMPLETED, THE ARTICULATION WAS STARTED WHILE MAKING AN ABNORMAL SOUND, AND IT WAS BENT TO 45. THE PARTS DROPPED IN THE PATIENT'S BODY. ALL PARTS WERE REMOVED AND WAS COLLECTED WITH FORCEPS. PLI 50: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. SOME STAPLE PUSHERS WERE PUSHED BACK TO THE CARTRIDGE. THREE STAPLES REMAINED ON THE CARTRIDGE AND ANVIL SURFACE. THE TWO STAPLES ON THE RELOAD WERE PROPERLY FORMED AND THE ONE WAS OVER CRIMPED.
25-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED PANCREATICODUODENECTOMY (WHIPPLE PROCEDURE), IN GASTRIC RESECTION, WHILE FIRING THE REINFORCED BLACK 60 MILLIMETER (MM) RELOAD, THE FIRING STOPPED DUE TO EXCESSIVE FORCE DETECTED. A NEW NON-REINFORCED BLACK RELOAD OF THE SAME LENGTH WAS USED, BUT IT ALSO STOPPED FIRING MIDWAY WITH THE SAME EXCESSIVE FORCE. TO RESOLVE THE ISSUE, A NEW HANDLE, A DAPTER, AND BLACK RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED PANCREATICODUODENECTOMY (WHIPPLE PROCEDURE), IN GASTRIC RESECTION, WHILE FIRING THE REINFORCED BLACK 60 MILLIMETER (MM) RELOAD, THE FIRING STOPPED DUE TO EXCESSIVE FORCE DETECTED. A NEW NON-REINFORCED BLACK RELOAD OF THE SAME LENGTH WAS USED, BUT IT ALSO STOPPED FIRING MIDWAY WITH THE SAME EXCESSIVE FORCE. TO RESOLVE THE ISSUE, A NEW HANDLE, A DAPTER, AND BLACK RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED PANCREATICODUODENECTOMY (WHIPPLE PROCEDURE), IN GASTRIC RESECTION, WHILE FIRING THE REINFORCED BLACK 60 MILLIMETER (MM) RELOAD, THE FIRING STOPPED DUE TO EXCESSIVE FORCE DETECTED. A NEW NON-REINFORCED BLACK RELOAD OF THE SAME LENGTH WAS USED, BUT IT ALSO STOPPED FIRING MIDWAY WITH THE SAME EXCESSIVE FORCE. TO RESOLVE THE ISSUE, A NEW HANDLE, A DAPTER, AND BLACK RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED PANCREATICODUODENECTOMY (WHIPPLE PROCEDURE), IN GASTRIC RESECTION, WHILE FIRING THE REINFORCED BLACK 60 MILLIMETER (MM) RELOAD, THE FIRING STOPPED DUE TO EXCESSIVE FORCE DETECTED. A NEW NON-REINFORCED BLACK RELOAD OF THE SAME LENGTH WAS USED, BUT IT ALSO STOPPED FIRING MIDWAY WITH THE SAME EXCESSIVE FORCE. TO RESOLVE THE ISSUE, A NEW HANDLE, A DAPTER, AND BLACK RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED PANCREATICODUODENECTOMY (WHIPPLE PROCEDURE), IN GASTRIC RESECTION, WHILE FIRING THE REINFORCED BLACK 60 MILLIMETER (MM) RELOAD, THE FIRING STOPPED DUE TO EXCESSIVE FORCE DETECTED. A NEW NON-REINFORCED BLACK RELOAD OF THE SAME LENGTH WAS USED, BUT IT ALSO STOPPED FIRING MIDWAY WITH THE SAME EXCESSIVE FORCE. TO RESOLVE THE ISSUE, A NEW HANDLE, A DAPTER, AND BLACK RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED PANCREATICODUODENECTOMY (WHIPPLE PROCEDURE), IN GASTRIC RESECTION, WHILE FIRING THE REINFORCED BLACK 60 MILLIMETER (MM) RELOAD, THE FIRING STOPPED DUE TO EXCESSIVE FORCE DETECTED. A NEW NON-REINFORCED BLACK RELOAD OF THE SAME LENGTH WAS USED, BUT IT ALSO STOPPED FIRING MIDWAY WITH THE SAME EXCESSIVE FORCE. TO RESOLVE THE ISSUE, A NEW HANDLE, A DAPTER, AND BLACK RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
29-01-2024 | Malfunction | covidien lp llc north haven | 09-01-2025 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE HANDLE, THE HANDLE SEIZING AFTER FOUR ATTEMPTS TO RELOAD IT. THIS ISSUE WAS IDENTIFIED AS A FIRING PROBLEM WITH THE MANUAL HANDLE. THE HANDLE WAS SUBSEQUENTLY REPLACED WITH ANOTHER OF THE SAME REFERENCE AND LOT NUMBER RESULTED IN LOSS OF TIME. THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
31-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING DEVICE TESTING PRIOR TO USE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
18-12-2024 | Injury | ethicon endo-surgery | 10-01-2025 | ECHELON 3000 60MM STANDARD | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ANTERIOR RESECTION OF RECTUM, RS OF RECTUM WAS DISSECTED. (NO PREOPERATIVE THERAPY, AND RECTAL IRRIGATION IS PERFORMED.) THE DEVICE COULD BE DISSECTED AT THE 1ST FIRING WITHOUT ANY PROBLEM. THE 2ND FIRING WAS ALSO PERFORMED IN THE SAME WAY THAT WAITED FOR 30 SECONDS, FIRED, AND WAITED FOR 30 SECONDS, THEN THE JAWS WERE OPENED, BUT THE STOOL LEAKED FROM THE CUT END. THE SEALING WAS DONE AT THE AFFECTED AREA BY HAR1136, AND WHEN IT STOPPED, THE 3RD FIRING WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITHOUT ANY PROBLEMS, AND THE DEVICE WAS CONVERTED TO CIRCULATOR. ALL CARTRIDGES WERE BLUE, THE DEVICE WAS NOT REPLACED, AND ALL THREE FIRING WERE USED. THE CARTRIDGES WERE NOT WASHED. THE DOCTOR SAID THAT ALTHOUGH THE HOME BUTTON PUSHED WITH THE JAWS ARTICULATED AFTER THE 1ST FIRING, BUT EVEN MORE ARTICULATED, THEN RETURNED TO THE HOME POSITION. IT IS POSSIBLE THAT THE EVENT OCCURRED COMPOSITELY BY THE PATIENT'S CONDITION AND NAIVE WASHING. THE STAPLES WERE MALFORMED AT THE DISTAL END OF THE CARTRIDGE. THE CARTRIDGE COLOR WAS BLUE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. NO FURTHER INFORMATION IS AVAILABLE.
18-12-2024 | Injury | ethicon endo-surgery | 10-01-2025 | ECHELON 3000 60MM STANDARD | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC ANTERIOR RESECTION OF RECTUM, RS OF RECTUM WAS DISSECTED. (NO PREOPERATIVE THERAPY, AND RECTAL IRRIGATION IS PERFORMED.) THE DEVICE COULD BE DISSECTED AT THE 1ST FIRING WITHOUT ANY PROBLEM. THE 2ND FIRING WAS ALSO PERFORMED IN THE SAME WAY THAT WAITED FOR 30 SECONDS, FIRED, AND WAITED FOR 30 SECONDS, THEN THE JAWS WERE OPENED, BUT THE STOOL LEAKED FROM THE CUT END. THE SEALING WAS DONE AT THE AFFECTED AREA BY HAR1136, AND WHEN IT STOPPED, THE 3RD FIRING WAS PERFORMED. THE PROCEDURE WAS COMPLETED WITHOUT ANY PROBLEMS, AND THE DEVICE WAS CONVERTED TO CIRCULATOR. ALL CARTRIDGES WERE BLUE, THE DEVICE WAS NOT REPLACED, AND ALL THREE FIRING WERE USED. THE CARTRIDGES WERE NOT WASHED. THE DOCTOR SAID THAT ALTHOUGH THE HOME BUTTON PUSHED WITH THE JAWS ARTICULATED AFTER THE 1ST FIRING, BUT EVEN MORE ARTICULATED, THEN RETURNED TO THE HOME POSITION. IT IS POSSIBLE THAT THE EVENT OCCURRED COMPOSITELY BY THE PATIENT'S CONDITION AND NAIVE WASHING. THE STAPLES WERE MALFORMED AT THE DISTAL END OF THE CARTRIDGE. THE CARTRIDGE COLOR WAS BLUE. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. NO FURTHER INFORMATION IS AVAILABLE.
25-11-2024 | Malfunction | ethicon endo-surgery | 10-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-12-2024 | Malfunction | ethicon endo-surgery | 10-01-2025 | PWRD ECH FLEX 60MM 28CM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
05-12-2024 | Malfunction | ethicon endo-surgery | 10-01-2025 | PWRD ECH FLEX 60MM 28CM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2024 | Malfunction | ethicon endo-surgery | 10-01-2025 | ECHELON 3000 60MM LONG | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE RELOAD WOULDN'T FIT AFTER FIRST FIRE. THEY GOT ANOTHER DEVICE TO COMPLETE THE CASE WITHOUT PATIENT HARM.
20-10-2024 | Malfunction | ethicon endo-surgery | 10-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE EGIAUSHORT ENDOGIA ULTRA UNIV SHT STAP AND THE EGIA45AMT EGIA 45 ARTIC MED THICK SULU, DID NOT FIRE. THE ISSUE WAS RESOLVED BY REPLACING THE HANDLE AND RELOAD. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE EGIAUSHORT ENDOGIA ULTRA UNIV SHT STAP AND THE EGIA45AMT EGIA 45 ARTIC MED THICK SULU, DID NOT FIRE. THE ISSUE WAS RESOLVED BY REPLACING THE HANDLE AND RELOAD. THERE WAS NO PATIENT INJURY.
04-08-2024 | Death | us surgical puerto rico | 10-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 28 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED FOUR (2.3%) 30-DAY MORTALITY RATE AND 8 (4.7%) 90-DAY MORTALITY FOR PATIENTS WITH NO ANASTOMOTIC STRICTURE. THE CAUSE OF DEATHS WAS NOT REPORTED.
04-08-2024 | Death | us surgical puerto rico | 10-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 31 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED FOUR (2.3%) 30-DAY MORTALITY RATE AND 8 (4.7%) 90-DAY MORTALITY FOR PATIENTS WITH NO ANASTOMOTIC STRICTURE. THE CAUSE OF DEATHS WAS NOT REPORTED.
15-10-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Battery Problem | A0705 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, WHILE BOTH HANDLE WAS ON THE CHARGER, THE SCREEN SHOWED A RED BATTERY SYMBOL WITH A LINE GOING THROUGH IT FOR BOTH HANDLE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED BATTERY CELLS WERE FOUND TO BE VENTED.
15-10-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, WHILE BOTH HANDLE WAS ON THE CHARGER, THE SCREEN SHOWED A RED BATTERY SYMBOL WITH A LINE GOING THROUGH IT FOR BOTH HANDLE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED BATTERY CELLS WERE FOUND TO BE VENTED.
04-08-2024 | Death | us surgical puerto rico | 10-01-2025 | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 25 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED FOUR (2.3%) 30-DAY MORTALITY RATE AND 8 (4.7%) 90-DAY MORTALITY FOR PATIENTS WITH NO ANASTOMOTIC STRICTURE. THE CAUSE OF DEATHS WAS NOT REPORTED.
19-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC COLON RESECTION, WHEN THE 60MM RELOAD WAS ATTACHED INTO THE MANUAL HANDLE, AND THE STAPLER WAS INSERTED INTO THE COLON, RESECTION WAS ATTEMPTED, HOWEVER, THE HANDLE COULD NOT BE SQUEEZED. THE DEVICE DID NOT FIRE. IT WAS NOTED THAT THE GREEN BUTTON WAS ABLE TO BE PUSHED. A NEW HANDLE AND THE SAME RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED DISTAL GASTRECTOMY PROCEDURE, IN DUODENOTOMY, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, AND EVEN THOUGH THE PREPARATION FOR FIRING WAS COMPLETED, THE DEVICE DID NOT FIRE WHEN THE FIRING BUTTON WAS PRESSED. THERE WAS NO ERROR MESSAGE DISPLAYED. TO RESOLVE THE ISSUE, A NEW DEVICE FROM ANOTHER MANUFACTURER WAS THEN USED. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CARDIATRIC GASTRECTOMY, IN ESOPHAGECTOMY, THE DEVICE STOPPED FIRING HALFWAY. TO RESOLVE THE ISSUE, A NEW RELOAD OF THE SAME TYPE WAS USED WITH THE SAME ADAPTER AND HANDLE. THERE WAS NO PATIENT INJURY.
13-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | UNKNOWN SIGNIA EGIA ADAPTER | Unintended System Motion | A051206 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, AT THE FIRING OF SIGNIA, THE RELOAD WAS TWITCHING WHEN THE KNIFE ADVANCED TO THE DISTAL END. IT WAS NOTED THAT THE MOVEMENT WAS BIGGER THAN BEFORE. AS THERE WAS NO PROBLEM WITH THE DEVICE OPERATION ITSELF, THE DEVICE WAS CONTINUED TO BE USED AS IT WAS.
13-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | UNKNOWN SIGNIA EGIA ADAPTER | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, AT THE FIRING OF SIGNIA, THE RELOAD WAS TWITCHING WHEN THE KNIFE ADVANCED TO THE DISTAL END. IT WAS NOTED THAT THE MOVEMENT WAS BIGGER THAN BEFORE. AS THERE WAS NO PROBLEM WITH THE DEVICE OPERATION ITSELF, THE DEVICE WAS CONTINUED TO BE USED AS IT WAS.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CARDIATRIC GASTRECTOMY, IN ESOPHAGECTOMY, THE DEVICE STOPPED FIRING HALFWAY. TO RESOLVE THE ISSUE, A NEW RELOAD OF THE SAME TYPE WAS USED WITH THE SAME ADAPTER AND HANDLE. THERE WAS NO PATIENT INJURY.
13-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Unintended System Motion | A051206 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, AT THE FIRING OF SIGNIA, THE RELOAD WAS TWITCHING WHEN THE KNIFE ADVANCED TO THE DISTAL END. IT WAS NOTED THAT THE MOVEMENT WAS BIGGER THAN BEFORE. AS THERE WAS NO PROBLEM WITH THE DEVICE OPERATION ITSELF, THE DEVICE WAS CONTINUED TO BE USED AS IT WAS.
13-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, AT THE FIRING OF SIGNIA, THE RELOAD WAS TWITCHING WHEN THE KNIFE ADVANCED TO THE DISTAL END. IT WAS NOTED THAT THE MOVEMENT WAS BIGGER THAN BEFORE. AS THERE WAS NO PROBLEM WITH THE DEVICE OPERATION ITSELF, THE DEVICE WAS CONTINUED TO BE USED AS IT WAS.
13-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Unintended System Motion | A051206 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, AT THE FIRING OF SIGNIA, THE RELOAD WAS TWITCHING WHEN THE KNIFE ADVANCED TO THE DISTAL END. IT WAS NOTED THAT THE MOVEMENT WAS BIGGER THAN BEFORE. AS THERE WAS NO PROBLEM WITH THE DEVICE OPERATION ITSELF, THE DEVICE WAS CONTINUED TO BE USED AS IT WAS.
13-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, AT THE FIRING OF SIGNIA, THE RELOAD WAS TWITCHING WHEN THE KNIFE ADVANCED TO THE DISTAL END. IT WAS NOTED THAT THE MOVEMENT WAS BIGGER THAN BEFORE. AS THERE WAS NO PROBLEM WITH THE DEVICE OPERATION ITSELF, THE DEVICE WAS CONTINUED TO BE USED AS IT WAS.
26-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER DURING A LEFT UPPER LOBECTOMY, AFTER DEPLOYING THE STAPLER, THERE WAS AN ISSUE WITH STAPLE FORMATION. IT WAS NOTED THAT THE STAPLE LINE HAD DEHISCED ON THE BRONCHUS. SUTURING WAS DONE AS A STAPLE LINE REPLACEMENT AND TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES.
26-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING A LEFT UPPER LOBECTOMY, AFTER DEPLOYING THE STAPLER, THERE WAS AN ISSUE WITH STAPLE FORMATION. IT WAS NOTED THAT THE STAPLE LINE HAD DEHISCED ON THE BRONCHUS. SUTURING WAS DONE AS A STAPLE LINE REPLACEMENT AND TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES.
26-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER DURING A LEFT UPPER LOBECTOMY, AFTER DEPLOYING THE STAPLER, THERE WAS AN ISSUE WITH STAPLE FORMATION. IT WAS NOTED THAT THE STAPLE LINE HAD DEHISCED ON THE BRONCHUS. SUTURING WAS DONE AS A STAPLE LINE REPLACEMENT AND TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES.
26-11-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING A LEFT UPPER LOBECTOMY, AFTER DEPLOYING THE STAPLER, THERE WAS AN ISSUE WITH STAPLE FORMATION. IT WAS NOTED THAT THE STAPLE LINE HAD DEHISCED ON THE BRONCHUS. SUTURING WAS DONE AS A STAPLE LINE REPLACEMENT AND TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR 30 MINUTES.
11-12-2024 | Injury | covidien lp llc north haven | 10-01-2025 | EEA ORVIL | Detachment of Device or Device Component | A0501 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE ORAL ANVIL, TUBING AND RETRIEVAL STRING DISENGAGED. THE BROKEN PIECES DID NOT FALL INTO THE PATIENT CAVITY. SUTURING WAS DONE FOR STAPLE LINE REPLACEMENT AND TO RESOLVE THE ISSUE. THERE WAS EXTENDED HOSPITALIZATION FOR AT LEAST ONE WEEK.
19-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Cycle | A050703 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC STOMACH WEDGE RESECTION, THERE WAS AN ISSUE WITH THE SIG POWER SIGADAPTSTND LINEAR ADAPTER DURING ITS USE. IT WAS NOTED THAT IT KEPT GO FIRING EVEN THOUGH TISSUE WAS SLIPPING FROM THE JAW. THE DEVICE WAS REPLACED WITH ANOTHER STAPLER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-10-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT TURN ON OR CHARGE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
16-10-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Failure to Power Up | A070803 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT TURN ON OR CHARGE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
16-10-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT TURN ON OR CHARGE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
17-12-2024 | Injury | covidien lp llc north haven | 10-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE INVOLVING THE 30 TAN STAPLER ON THE PULMONARY ARTERY BRANCH, THERE WAS SIGNIFICANT BLEEDING IMMEDIATELY PROXIMAL TO THE STAPLE LINE. THIS ISSUE NECESSITATED A CONVERSION FROM LAPAROSCOPIC TO OPEN SURGERY, REQUIRING A THORACOTOMY AND SUTURING TO CONTROL THE BLEEDING. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES DUE TO THESE COMPLICATIONS.
17-12-2024 | Injury | covidien lp llc north haven | 10-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE INVOLVING THE 30 TAN STAPLER ON THE PULMONARY ARTERY BRANCH, THERE WAS SIGNIFICANT BLEEDING IMMEDIATELY PROXIMAL TO THE STAPLE LINE. THIS ISSUE NECESSITATED A CONVERSION FROM LAPAROSCOPIC TO OPEN SURGERY, REQUIRING A THORACOTOMY AND SUTURING TO CONTROL THE BLEEDING. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES DUE TO THESE COMPLICATIONS.
18-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2024 | Injury | us surgical puerto rico | 10-01-2025 | ENDO GIA | Break | A0401 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER DURING THE PROCEDURE, THE DEVICE ARTICULATION WAS NOT SMOOTH. IT WAS ALSO NOTED THAT THE SHIPPING WEDGE OF THE RELOAD WERE BROKEN AND FELL INSIDE THE PATIENT'S CAVITY AND WAS NOT RETRIEVED. AS A RESOLUTION THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE.
19-12-2024 | Injury | us surgical puerto rico | 10-01-2025 | ENDO GIA | Positioning Problem | A1502 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER DURING THE PROCEDURE, THE DEVICE ARTICULATION WAS NOT SMOOTH. IT WAS ALSO NOTED THAT THE SHIPPING WEDGE OF THE RELOAD WERE BROKEN AND FELL INSIDE THE PATIENT'S CAVITY AND WAS NOT RETRIEVED. AS A RESOLUTION THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE.
18-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A BARIATRIC PROCEDURE, THE USE OF THE EGIA45AMT EGIA 45 ARTIC MED THICK SULU, THE STAPLE LINE WAS BLEEDING AFTER FIRING. TO RESOLVE THE ISSUE, MANUAL SUTURE WAS PERFORMED.
13-11-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID RESECTION PROCEDURE INVOLVING THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP, A DEVICE COMPONENT BROKE OFF, RESULTING IN AN EXTENDED SURGICAL TIME OF 5 MINUTES. THE STAPLER WAS NOT ABLE TO FIRE, ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE COMPONENT DID NOT FALL INTO THE PATIENT CAVITY. THE HANDLE WAS REPLACED BY ANOTHER PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-11-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA ULTRA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID RESECTION PROCEDURE INVOLVING THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP, A DEVICE COMPONENT BROKE OFF, RESULTING IN AN EXTENDED SURGICAL TIME OF 5 MINUTES. THE STAPLER WAS NOT ABLE TO FIRE, ALTHOUGH THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE COMPONENT DID NOT FALL INTO THE PATIENT CAVITY. THE HANDLE WAS REPLACED BY ANOTHER PRODUCT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2024 | Injury | davis & geck caribe | 10-01-2025 | UNKNOWN ENDO STITCH INSTRUMENT | Break | A0401 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL LAPAROSCOPIC HYSTERECTOMY, TWO BROKEN NEEDLES FELL INTO THE PATIENT CAVITY WHILE CLOSING THE VAGINAL CUFF. IT WAS NOTED THAT THE SURGICAL TIME WAS EXTENDED BY MORE THAN 30 MINUTES TO FIND THE NEEDLE BUT IT WAS UNABLE TO BE RETRIEVED. A NEW NEEDLE WAS USED TO RESOLVE THE ISSUE.
18-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMI HEPATECTOMY, IN CLOSING AND DETACHING THE RIGHT HEPATIC VEIN, AFTER CLOSING AND ENTERING THE TROCAR, THE JAWS OF THE DEVICE COULD NOT BE OPENED. AFTER REMOVING THE DEVICE FROM THE TROCAR, THE USER TRIED TO OPEN THE JAWS, BUT IT STILL COULD NOT BE OPENED. ADDITIONALLY, THE DEVICE COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMI HEPATECTOMY, IN CLOSING AND DETACHING THE RIGHT HEPATIC VEIN, AFTER CLOSING AND ENTERING THE TROCAR, THE JAWS OF THE DEVICE COULD NOT BE OPENED. AFTER REMOVING THE DEVICE FROM THE TROCAR, THE USER TRIED TO OPEN THE JAWS, BUT IT STILL COULD NOT BE OPENED. ADDITIONALLY, THE DEVICE COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMI HEPATECTOMY, IN CLOSING AND DETACHING THE RIGHT HEPATIC VEIN, AFTER CLOSING AND ENTERING THE TROCAR, THE JAWS OF THE DEVICE COULD NOT BE OPENED. AFTER REMOVING THE DEVICE FROM THE TROCAR, THE USER TRIED TO OPEN THE JAWS, BUT IT STILL COULD NOT BE OPENED. ADDITIONALLY, THE DEVICE COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT HEMI HEPATECTOMY, IN CLOSING AND DETACHING THE RIGHT HEPATIC VEIN, AFTER CLOSING AND ENTERING THE TROCAR, THE JAWS OF THE DEVICE COULD NOT BE OPENED. AFTER REMOVING THE DEVICE FROM THE TROCAR, THE USER TRIED TO OPEN THE JAWS, BUT IT STILL COULD NOT BE OPENED. ADDITIONALLY, THE DEVICE COULD NOT BE REMOVED. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
19-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE USE OF THE RELOAD, A TIGHT JAW WAS OBSERVED. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH A NEW RELOAD. THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED DISTAL GASTRECTOMY PROCEDURE, IN DUODENOTOMY, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, AND EVEN THOUGH THE PREPARATION FOR FIRING WAS COMPLETED, THE DEVICE DID NOT FIRE WHEN THE FIRING BUTTON WAS PRESSED. THERE WAS NO ERROR MESSAGE DISPLAYED. TO RESOLVE THE ISSUE, A NEW DEVICE FROM ANOTHER MANUFACTURER WAS THEN USED. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED DISTAL GASTRECTOMY PROCEDURE, IN DUODENOTOMY, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, AND EVEN THOUGH THE PREPARATION FOR FIRING WAS COMPLETED, THE DEVICE DID NOT FIRE WHEN THE FIRING BUTTON WAS PRESSED. THERE WAS NO ERROR MESSAGE DISPLAYED. TO RESOLVE THE ISSUE, A NEW DEVICE FROM ANOTHER MANUFACTURER WAS THEN USED. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED DISTAL GASTRECTOMY PROCEDURE, IN DUODENOTOMY, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, AND EVEN THOUGH THE PREPARATION FOR FIRING WAS COMPLETED, THE DEVICE DID NOT FIRE WHEN THE FIRING BUTTON WAS PRESSED. THERE WAS NO ERROR MESSAGE DISPLAYED. TO RESOLVE THE ISSUE, A NEW DEVICE FROM ANOTHER MANUFACTURER WAS THEN USED. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED DISTAL GASTRECTOMY PROCEDURE, IN DUODENOTOMY, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, AND EVEN THOUGH THE PREPARATION FOR FIRING WAS COMPLETED, THE DEVICE DID NOT FIRE WHEN THE FIRING BUTTON WAS PRESSED. THERE WAS NO ERROR MESSAGE DISPLAYED. TO RESOLVE THE ISSUE, A NEW DEVICE FROM ANOTHER MANUFACTURER WAS THEN USED. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED DISTAL GASTRECTOMY PROCEDURE, IN DUODENOTOMY, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, AND EVEN THOUGH THE PREPARATION FOR FIRING WAS COMPLETED, THE DEVICE DID NOT FIRE WHEN THE FIRING BUTTON WAS PRESSED. THERE WAS NO ERROR MESSAGE DISPLAYED. TO RESOLVE THE ISSUE, A NEW DEVICE FROM ANOTHER MANUFACTURER WAS THEN USED. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED DISTAL GASTRECTOMY PROCEDURE, IN DUODENOTOMY, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, AND EVEN THOUGH THE PREPARATION FOR FIRING WAS COMPLETED, THE DEVICE DID NOT FIRE WHEN THE FIRING BUTTON WAS PRESSED. THERE WAS NO ERROR MESSAGE DISPLAYED. TO RESOLVE THE ISSUE, A NEW DEVICE FROM ANOTHER MANUFACTURER WAS THEN USED. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED DISTAL GASTRECTOMY PROCEDURE, IN DUODENOTOMY, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, AND EVEN THOUGH THE PREPARATION FOR FIRING WAS COMPLETED, THE DEVICE DID NOT FIRE WHEN THE FIRING BUTTON WAS PRESSED. THERE WAS NO ERROR MESSAGE DISPLAYED. TO RESOLVE THE ISSUE, A NEW DEVICE FROM ANOTHER MANUFACTURER WAS THEN USED. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED DISTAL GASTRECTOMY PROCEDURE, IN DUODENOTOMY, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, AND EVEN THOUGH THE PREPARATION FOR FIRING WAS COMPLETED, THE DEVICE DID NOT FIRE WHEN THE FIRING BUTTON WAS PRESSED. THERE WAS NO ERROR MESSAGE DISPLAYED. TO RESOLVE THE ISSUE, A NEW DEVICE FROM ANOTHER MANUFACTURER WAS THEN USED. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE TISSUE WAS LIFT UP BY THE GRASPER FORCEPS, THE STAPLER WAS THEN HOLD FOR 10 SECONDS BUT THE GUN MECHANISM (TRIGGER) BROKE. HENCE, THE STAPLES DID NOT DEPLOY AT ALL. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE TISSUE WAS LIFT UP BY THE GRASPER FORCEPS, THE STAPLER WAS THEN HOLD FOR 10 SECONDS BUT THE GUN MECHANISM (TRIGGER) BROKE. HENCE, THE STAPLES DID NOT DEPLOY AT ALL. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE TISSUE WAS LIFT UP BY THE GRASPER FORCEPS, THE STAPLER WAS THEN HOLD FOR 10 SECONDS BUT THE GUN MECHANISM (TRIGGER) BROKE. HENCE, THE STAPLES DID NOT DEPLOY AT ALL. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | UNKNOWN EGIA TRI STAPLE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE TISSUE WAS LIFT UP BY THE GRASPER FORCEPS, THE STAPLER WAS THEN HOLD FOR 10 SECONDS BUT THE GUN MECHANISM (TRIGGER) BROKE. HENCE, THE STAPLES DID NOT DEPLOY AT ALL. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | UNKNOWN EGIA TRI STAPLE | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE TISSUE WAS LIFT UP BY THE GRASPER FORCEPS, THE STAPLER WAS THEN HOLD FOR 10 SECONDS BUT THE GUN MECHANISM (TRIGGER) BROKE. HENCE, THE STAPLES DID NOT DEPLOY AT ALL. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | UNKNOWN EGIA TRI STAPLE | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE TISSUE WAS LIFT UP BY THE GRASPER FORCEPS, THE STAPLER WAS THEN HOLD FOR 10 SECONDS BUT THE GUN MECHANISM (TRIGGER) BROKE. HENCE, THE STAPLES DID NOT DEPLOY AT ALL. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE FIRST FIRING AT THE PYLORUS, THE TWO RELOADS COULD NOT BE FIRED WHEN IT WAS FIRED TO THE MIDDLE POSITION. TO RESOLVE THE ISSUE, A NEW STAPLER WAS USED, AND A MANUAL SUTURING WAS DONE. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | covidien lp llc north haven | 10-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE FIRST FIRING AT THE PYLORUS, THE TWO RELOADS COULD NOT BE FIRED WHEN IT WAS FIRED TO THE MIDDLE POSITION. TO RESOLVE THE ISSUE, A NEW STAPLER WAS USED, AND A MANUAL SUTURING WAS DONE. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, ON THE FIRST FIRING AT THE PYLORUS, THE TWO RELOADS COULD NOT BE FIRED WHEN IT WAS FIRED TO THE MIDDLE POSITION. TO RESOLVE THE ISSUE, A NEW STAPLER WAS USED, AND A MANUAL SUTURING WAS DONE. THERE WAS NO PATIENT INJURY.
26-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC SEGMENTECTOMY, IN DETACHING THE INTERLOBAR FISSURE, THE STAPLES ON THE DISTAL END WERE POORLY FORMED (SHAPE OF A DOOR). ANOTHER TYPE OF RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. TO RESOLVE THE ISSUE, MANUAL SUTURING WAS DONE.
26-12-2024 | Malfunction | us surgical puerto rico | 10-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC SEGMENTECTOMY, IN DETACHING THE INTERLOBAR FISSURE, THE STAPLES ON THE DISTAL END WERE POORLY FORMED (SHAPE OF A DOOR). ANOTHER TYPE OF RELOAD WAS USED, BUT THE SAME ISSUE OCCURRED. TO RESOLVE THE ISSUE, MANUAL SUTURING WAS DONE.
25-11-2024 | Injury | standard bariatrics | 10-01-2025 | TITAN SGS RB STAPLER 23 CM | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: AWARENESS OF A POST-OPERATIVE LEAK WAS COMMUNICATED ON (B)(6) 2024. THE PATIENT UNDERWENT A LAPAROSCOPIC SLEEVE GASTRECTOMY ON (B)(6) 2024 AND APPROXIMATELY 1 WEEK LATER, THE PATIENT WAS BROUGHT BACK TO THE HOSPITAL WHERE A LEAK WAS DISCOVERED. THE PATIENT RECEIVED SURGERY TO FIX THE LEAK AND WAS TRANSFERRED TO ANOTHER FACILITY WHERE THE PATIENT WAS SEPTIC. THE PATIENT IS RECOVERING IN THE HOSPITAL.
16-10-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE HANDLE COULD NOT BE SQUEEZED AND FIRING COULD NOT BE PERFORMED AT ALL. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-08-2024 | Malfunction | covidien surgical | 13-01-2025 | GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS FIRING ISSUE. SUTURING WAS DONE AS STAPLE LINE REPLACEMENT AND TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-08-2024 | Malfunction | covidien surgical | 13-01-2025 | GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THERE WAS FIRING ISSUE. SUTURING WAS DONE AS STAPLE LINE REPLACEMENT AND TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC THORACOSCOPIC RIGHT LOWER LOBECTOMY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, HOWEVER THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE SIG POWER SIGADAPTSTND LINEAR ADAPTER WAS DAMAGED BECAUSE THE EGIA60AMT EGIA 60 ARTIC MED THICK SULU BROKE AT THE ARTICULATION JOINT, AND COULD NOT BE REMOVED FROM THE ADAPTER. TO RESOLVE THE ISSUE, ANOTHER STAPLER FORM A DIFFERENT MANUFACTURER WAS USED. THERE WAS NO PATIENT INJURY.
01-01-2024 | Injury | ethicon endo-surgery | 13-01-2025 | PROXIMATE*PPH PROCEDURE SET | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT A DR. LODGED A COMPLAINT TO THE SURGICAL DIRECTOR AT THE (B)(6) HOSPITAL REGARDING PPHO3 STAPLER. HE CLAIMED THAT IT FAILED ON ONE OF HIS PATIENTS WHO WAS TAKEN BACK TO THEATER AFTER A STAPLED HEMORRHOIDOPEXY PROCEDURE USING PPH03 STAPLER. THE PROCEDURE WAS DONE AND PATIENT WAS DISCHARGED HOME. PATIENT CAME BACK TO CASUALTY WITH PER-RECTAL BLEEDING, RETURNED BACK TO THEATER FOR AN EXAMINATION UNDER ANESTHESIA, HE STATED THAT HE COULD CLEARLY SEE THAT THE STAPLES HAD NOT COMPLETELY FORMED AND SOME HAD GIVEN WAY UPON REDO PROCTOSCOPY.
01-01-2024 | Injury | ethicon endo-surgery | 13-01-2025 | PROXIMATE*PPH PROCEDURE SET | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT A DR. LODGED A COMPLAINT TO THE SURGICAL DIRECTOR AT THE (B)(6) HOSPITAL REGARDING PPHO3 STAPLER. HE CLAIMED THAT IT FAILED ON ONE OF HIS PATIENTS WHO WAS TAKEN BACK TO THEATER AFTER A STAPLED HEMORRHOIDOPEXY PROCEDURE USING PPH03 STAPLER. THE PROCEDURE WAS DONE AND PATIENT WAS DISCHARGED HOME. PATIENT CAME BACK TO CASUALTY WITH PER-RECTAL BLEEDING, RETURNED BACK TO THEATER FOR AN EXAMINATION UNDER ANESTHESIA, HE STATED THAT HE COULD CLEARLY SEE THAT THE STAPLES HAD NOT COMPLETELY FORMED AND SOME HAD GIVEN WAY UPON REDO PROCTOSCOPY.
18-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LOWER ANTERIOR RESECTION, THE DEVICE HANDLE SHOWED A FULL BATTERY. BUT DEPLETED COMPLETELY AFTER TWO RELOADS. AS A RESOLUTION, A NEW HANDLE AND SHELL WERE USED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | SIGNIA | Unraveled Material | A040508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE POWER SHELL WAS UNABLE TO CLOSE. ANOTHER SHELL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
23-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | ABSORBATACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH (IPOM) PROCEDURE, THE SHAFT OF THE DEVICE BROKE WHILE ATTEMPTING TO FIRE THE TACKER. TO RESOLVE THE ISSUE, ANOTHER TACKER WAS USED. THERE WERE NO INJURIES.
18-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LOWER ANTERIOR RESECTION, THE DEVICE HANDLE SHOWED A FULL BATTERY. BUT DEPLETED COMPLETELY AFTER TWO RELOADS. AS A RESOLUTION, A NEW HANDLE AND SHELL WERE USED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INJURY.
22-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER RESECTION, THE DEVICE PARTIALLY FIRED. A NEW RELOAD WAS LOADED ONTO THE SAME HANDLE AND THE DEVICE FIRED WITH NO ISSUE. THERE WAS NO PATIENT INJURY.
14-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE DEVICE, THE STAPLER JAMMED. THE LOAD WAS CHANGED BUT THE ISSUE PERSISTED. THE STAPLER WAS SUBSEQUENTLY REPLACED, RESOLVING THE PROBLEM TEMPORARILY. HOWEVER, AFTER SOME USE, ANOTHER LOAD JAMMED. THE LOAD WAS CHANGED AGAIN, AND THE SURGERY WAS COMPLETED. THERE WAS NO PATIENT INJURY.
14-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE DEVICE, THE STAPLER JAMMED. THE LOAD WAS CHANGED BUT THE ISSUE PERSISTED. THE STAPLER WAS SUBSEQUENTLY REPLACED, RESOLVING THE PROBLEM TEMPORARILY. HOWEVER, AFTER SOME USE, ANOTHER LOAD JAMMED. THE LOAD WAS CHANGED AGAIN, AND THE SURGERY WAS COMPLETED. THERE WAS NO PATIENT INJURY.
18-12-2024 | Injury | intuitive surgical | 13-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.
18-12-2024 | Injury | intuitive surgical | 13-01-2025 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.
23-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH REPAIR, WHILE USING THE TACKER FOR MESH FIXATION, THE REMAINING TACKS WERE NOT COMING PROPERLY, AND THE TRIGGER BECAME HARD, RESULTING IN THE TACKS NOT PROVIDING TIGHT FIXATION TO THE MESH. ANOTHER FIXATION DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | ABSORBATACK | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC INTRAPERITONEAL ONLAY MESH REPAIR, WHILE USING THE TACKER FOR MESH FIXATION, THE REMAINING TACKS WERE NOT COMING PROPERLY, AND THE TRIGGER BECAME HARD, RESULTING IN THE TACKS NOT PROVIDING TIGHT FIXATION TO THE MESH. ANOTHER FIXATION DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
20-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY (PD), IN GASTRECTOMY, AFTER FIRING THE DEVICE, THE KNIFE WAS PULLED BACK ONLY HALFWAY, THE TISSUE BECAME ENTANGLED, AND THE JAWS OF THE DEVICE COULD NOT BE REMOVED. IT BECAME DIFFICULT TO BE PULLED BACK FROM THE MIDDLE, THE USER HAD TO FORCIBLY PULLED BACK THE KNIFE. TO COMPLETE THE CASE, A COLON RESECTION WAS PERFORMED.
20-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY (PD), IN GASTRECTOMY, AFTER FIRING THE DEVICE, THE KNIFE WAS PULLED BACK ONLY HALFWAY, THE TISSUE BECAME ENTANGLED, AND THE JAWS OF THE DEVICE COULD NOT BE REMOVED. IT BECAME DIFFICULT TO BE PULLED BACK FROM THE MIDDLE, THE USER HAD TO FORCIBLY PULLED BACK THE KNIFE. TO COMPLETE THE CASE, A COLON RESECTION WAS PERFORMED.
20-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY (PD), IN GASTRECTOMY, AFTER FIRING THE DEVICE, THE KNIFE WAS PULLED BACK ONLY HALFWAY, THE TISSUE BECAME ENTANGLED, AND THE JAWS OF THE DEVICE COULD NOT BE REMOVED. IT BECAME DIFFICULT TO BE PULLED BACK FROM THE MIDDLE, THE USER HAD TO FORCIBLY PULLED BACK THE KNIFE. TO COMPLETE THE CASE, A COLON RESECTION WAS PERFORMED.
20-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY (PD), IN GASTRECTOMY, AFTER FIRING THE DEVICE, THE KNIFE WAS PULLED BACK ONLY HALFWAY, THE TISSUE BECAME ENTANGLED, AND THE JAWS OF THE DEVICE COULD NOT BE REMOVED. IT BECAME DIFFICULT TO BE PULLED BACK FROM THE MIDDLE, THE USER HAD TO FORCIBLY PULLED BACK THE KNIFE. TO COMPLETE THE CASE, A COLON RESECTION WAS PERFORMED.
14-03-2025 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: IT WAS REPORTED THAT THERE WAS AN UNKNOWN ISSUE INVOLVING THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
14-03-2025 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT THERE WAS AN UNKNOWN ISSUE INVOLVING THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
14-03-2025 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Mechanics Altered | A0507 | Insufficient Information | Event Description: IT WAS REPORTED THAT THERE WAS AN UNKNOWN ISSUE INVOLVING THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
19-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE FIRST RELOAD FROM THE SLEEVE KIT DID NOT FIRE PROPERLY. THE FIRING PROCESS WAS DIFFICULT AND A NOISE COULD BE HEARD. THE SECOND RELOAD HAD THE SAME RESULT. THE TISSUE WAS PARTIALLY STAPLED AND NOT COMPLETELY SEVERED. A NEW HANDLE AND A NEW RELOAD WERE USED TO RESOLVE THE ISSUE.
01-01-2025 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT THORACOSCOPIC LOBECTOMY PROCEDURE, WHEN CUTTING LUNG TISSUE, THE STAPLER DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CLAMP TEST PRIOR TO USE, THE JAWS DID NOT CLOSE COMPLETELY. A NEW RELOAD FROM A DIFFERENT BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE CLAMPING MECHANISM WAS DEFORMED.
10-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING CLAMP TEST PRIOR TO USE, THE JAWS DID NOT CLOSE COMPLETELY. A NEW RELOAD FROM A DIFFERENT BATCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE CLAMPING MECHANISM WAS DEFORMED.
01-01-2025 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RIGHT THORACOSCOPIC LOBECTOMY PROCEDURE, WHEN CUTTING LUNG TISSUE, THE STAPLER DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
19-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE FIRST RELOAD FROM THE SLEEVE KIT DID NOT FIRE PROPERLY. THE FIRING PROCESS WAS DIFFICULT AND A NOISE COULD BE HEARD. THE SECOND RELOAD HAD THE SAME RESULT. THE TISSUE WAS PARTIALLY STAPLED AND NOT COMPLETELY SEVERED. A NEW HANDLE AND A NEW RELOAD WERE USED TO RESOLVE THE ISSUE.
19-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE FIRST RELOAD FROM THE SLEEVE KIT DID NOT FIRE PROPERLY. THE FIRING PROCESS WAS DIFFICULT AND A NOISE COULD BE HEARD. THE SECOND RELOAD HAD THE SAME RESULT. THE TISSUE WAS PARTIALLY STAPLED AND NOT COMPLETELY SEVERED. A NEW HANDLE AND A NEW RELOAD WERE USED TO RESOLVE THE ISSUE.
19-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE FIRST RELOAD FROM THE SLEEVE KIT DID NOT FIRE PROPERLY. THE FIRING PROCESS WAS DIFFICULT AND A NOISE COULD BE HEARD. THE SECOND RELOAD HAD THE SAME RESULT. THE TISSUE WAS PARTIALLY STAPLED AND NOT COMPLETELY SEVERED. A NEW HANDLE AND A NEW RELOAD WERE USED TO RESOLVE THE ISSUE.
16-12-2024 | Injury | covidien lp llc north haven | 13-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, A DAY AFTER AN OPEN RIGHT HEMICOLECTOMY, THERE WAS ANASTOMOTIC LEAK. THE ANASTOMOSIS STAPLE LINE WAS OPEN WITH NO VISIBLE STAPLES. IT WAS NOTED THAT THERE WAS SEVERE ANASTOMOSIS FAILURE. THE SURGEON HAD TO CREATE A STOMA AND PATIENT ON ANTIBIOTICS. PATIENT STILL IN HOSPITAL.
16-12-2024 | Injury | covidien lp llc north haven | 13-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, A DAY AFTER AN OPEN RIGHT HEMICOLECTOMY, THERE WAS ANASTOMOTIC LEAK. THE ANASTOMOSIS STAPLE LINE WAS OPEN WITH NO VISIBLE STAPLES. IT WAS NOTED THAT THERE WAS SEVERE ANASTOMOSIS FAILURE. THE SURGEON HAD TO CREATE A STOMA AND PATIENT ON ANTIBIOTICS. PATIENT STILL IN HOSPITAL.
16-12-2024 | Injury | covidien lp llc north haven | 13-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, A DAY AFTER AN OPEN RIGHT HEMICOLECTOMY, THERE WAS ANASTOMOTIC LEAK. THE ANASTOMOSIS STAPLE LINE WAS OPEN WITH NO VISIBLE STAPLES. IT WAS NOTED THAT THERE WAS SEVERE ANASTOMOSIS FAILURE. THE SURGEON HAD TO CREATE A STOMA AND PATIENT ON ANTIBIOTICS. PATIENT STILL IN HOSPITAL.
16-12-2024 | Injury | covidien lp llc north haven | 13-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, A DAY AFTER AN OPEN RIGHT HEMICOLECTOMY, THERE WAS ANASTOMOTIC LEAK. THE ANASTOMOSIS STAPLE LINE WAS OPEN WITH NO VISIBLE STAPLES. IT WAS NOTED THAT THERE WAS SEVERE ANASTOMOSIS FAILURE. THE SURGEON HAD TO CREATE A STOMA AND PATIENT ON ANTIBIOTICS. PATIENT STILL IN HOSPITAL.
26-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY INVOLVING THE HEM3335 33MM HEMORRHOID STAPLER, THE INSTRUMENT DID NOT FIRE. THIS RESULTED IN AN EXTENDED SURGICAL TIME OF 30 MINUTES. THE DEVICE WAS REPLACED BY ANOTHER MANUFACTURER'S PRODUCT. THERE WAS NO PATIENT INJURY.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | PROXIMATE*CUTTER RELOAD UNIT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | PROXIMATE*CUTTER RELOAD UNIT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | PROXIMATE*CUTTER RELOAD UNIT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | ENDO ILS 25MM, CURVED | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE SURGERY, OPEN THE PACKING AND NOTED ANVIL CANNOT BE ATTACHED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | WHITE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT HEMIHEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED HALFWAY. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
02-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SEGMENTECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS BROKEN WHILE CARRYING OUT AN OPENING OF THE LOAD. THEY MADE AN ATTEMPT TO FIT THE STAPLE BUT IT DID NOT FIT. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
13-12-2024 | Malfunction | ethicon endo-surgery | 13-01-2025 | ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2024 | Malfunction | ethicon endo-surgery | 13-01-2025 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTOCOLECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
13-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE JAWS ON THE DEVICE LOCKED-OUT AND WOULD NOT OPEN WITH PART OF THE SPECIMEN INSIDE THE JAWS. THERE WERE NO PATIENT CONSEQUENCES NOR SURGICAL DELAY REPORTED. THE EXACT DATE OF THE EVENT WAS UNKNOWN. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2024 | Malfunction | ethicon endo-surgery | 13-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE JAMMED DURING FIRING, MANUAL OVERRIDE ENGAGED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2024 | Malfunction | ethicon endo-surgery | 13-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE JAMMED DURING FIRING, MANUAL OVERRIDE ENGAGED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
02-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | ECHELON FLEX 60 LONG | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE JAW ON THE DEVICE WAS NOT OPENING IN FOURTH FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
03-01-2025 | Malfunction | johnson & johnson medical (suzhou | 13-01-2025 | CARTRIDGE FINAL ASSY 45MM GREEN | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAD A CARTRIDGE PAN SEPARATION AFTER OPENING THE PACKAGING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
16-12-2024 | Malfunction | ethicon endo-surgery | 13-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED (STUCK) AFTER FIRING; AND COULD NOT BE USED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED (STUCK) AFTER FIRING; AND COULD NOT BE USED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AND STUCK IN THE DEVICE. THE SURGEON CHANGED TO ANOTHER DEVICE TO CONTINUE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED (STUCK) AFTER FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED (STUCK) AFTER FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | GST BLACK RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED (STUCK) AFTER FIRING; AND COULD NOT BE USED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | GST BLACK RELOAD 45MM 6 ROW | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED (STUCK) AFTER FIRING; AND COULD NOT BE USED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
03-02-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN LAPAROSCOPIC PROCEDURE, IT WAS OBSERVED THAT THE KNIFE STOPPED WHEN PERFORMING THE INTESTINE TRANSECTION (THE 4TH TIMES). THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. THE EXACT DATE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION REQUESTED.
08-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | ENDO ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2024 | Injury | intuitive surgical | 13-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY AND PARAESOPHAGEAL HIATAL HERNIA REPAIR PROCEDURE, THE STAPLE LINE WHERE A WHITE SUREFORM 60 RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT HAD MISSING STAPLES. SOME OF THE MIDDLE STAPLES DID NOT DEPLOY IN THE FIRING SEQUENCE AND IT WAS OBSERVED THAT THE STAPLER RELOAD HAD A CRACK ON IT. IN ADDITION, THE TARGET TISSUE WAS NOT RELEASED IMMEDIATELY DUE TO TISSUE STICKING TO STAPLES THAT WERE NOT DEPLOYED. THE TISSUE WAS RELEASED WITHOUT ISSUE. THERE WERE NO ERROR MESSAGES THAT OCCURRED DURING THE EVENT. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. THERE WAS NO BLEEDING OBSERVED. THE SURGEON OVERSEWED THE STAPLE LINE DUE TO NOT ALL OF THE STAPLES DEPLOYING. THERE WAS NO UNEXPECTED TISSUE REMOVAL. THIS EVENT CAUSED LESS THAN A 15 MINUTE DELAY IN THE PROCEDURE. THERE WERE NO PHOTOGRAPHS TAKEN. THE CASE WAS COMPLETED ROBOTICALLY.
17-12-2024 | Injury | intuitive surgical | 13-01-2025 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY AND PARAESOPHAGEAL HIATAL HERNIA REPAIR PROCEDURE, THE STAPLE LINE WHERE A WHITE SUREFORM 60 RELOAD WAS FIRED WITH A SUREFORM 60 STAPLER INSTRUMENT HAD MISSING STAPLES. SOME OF THE MIDDLE STAPLES DID NOT DEPLOY IN THE FIRING SEQUENCE AND IT WAS OBSERVED THAT THE STAPLER RELOAD HAD A CRACK ON IT. IN ADDITION, THE TARGET TISSUE WAS NOT RELEASED IMMEDIATELY DUE TO TISSUE STICKING TO STAPLES THAT WERE NOT DEPLOYED. THE TISSUE WAS RELEASED WITHOUT ISSUE. THERE WERE NO ERROR MESSAGES THAT OCCURRED DURING THE EVENT. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. THERE WAS NO BLEEDING OBSERVED. THE SURGEON OVERSEWED THE STAPLE LINE DUE TO NOT ALL OF THE STAPLES DEPLOYING. THERE WAS NO UNEXPECTED TISSUE REMOVAL. THIS EVENT CAUSED LESS THAN A 15 MINUTE DELAY IN THE PROCEDURE. THERE WERE NO PHOTOGRAPHS TAKEN. THE CASE WAS COMPLETED ROBOTICALLY.
08-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-12-2024 | Malfunction | ethicon endo-surgery | 13-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC COLON RESECTION PROCEDURE THAT AFTER THE FIRST FIRING ONCE REMOVED FROM THE PATIENT THE DEVICE COULD NOT BE OPENED TO RELOAD. ANOTHER LIKE DEVICE WAS USED TO CONTINUE WITH THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES REPORTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 13-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE WHEN CONNECTING TROCAR TO ANVIL IT TOOK ADDITIONAL WORK TO GOT THEM CONNECTED. THERE HAVEN'T EARLIER BEEN THIS KIND OF DIFFICULTIES WITH CONNECTING. THE DEVICE FINALLY FIRED NORMALLY. THERE WERE NO PATIENT CONSEQUENCES.
20-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PNEUMONECTOMY PROCEDURE, IN SEVERING THE ARTERY, THE RELOAD WAS STUCK DURING FIRING AND COULD NOT CONTINUE. TO RESOLVE THE ISSUE, A NEW RELOAD WAS USED WITH THE SAME HANDLE. THERE WAS NO PATIENT INJURY.
19-12-2024 | Malfunction | davis & geck caribe | 13-01-2025 | UNKNOWN ENDO STITCH SULU | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE INSTRUMENT, A NEEDLE FROM A VLOC SUTURE FELL INTO THE CAVITY OF THE PATIENT. THERE WAS NO REPORTED INJURY TO THE PATIENT, AND THE PATIENT OUTCOME WAS NOTED AS ALIVE WITH NO INJURY.
19-12-2024 | Malfunction | davis & geck caribe | 13-01-2025 | UNKNOWN ENDO STITCH SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE INSTRUMENT, A NEEDLE FROM A VLOC SUTURE FELL INTO THE CAVITY OF THE PATIENT. THERE WAS NO REPORTED INJURY TO THE PATIENT, AND THE PATIENT OUTCOME WAS NOTED AS ALIVE WITH NO INJURY.
18-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE SETTING UP, GREEN LIGHT DID NOT LIGHT UP, SO IT COULD NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
18-12-2024 | Malfunction | us surgical puerto rico | 13-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE SETTING UP, GREEN LIGHT DID NOT LIGHT UP, SO IT COULD NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
15-12-2024 | Injury | intuitive surgical | 13-01-2025 | SUREFORM | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED CHOLECYSTECTOMY, THE BLADE OF A SUREFORM 30 RELOAD WAS FOUND TO HAVE FALLEN INSIDE THE PATIENT. WHEN FIRING THE WHITE SUREFORM 30 RELOAD, THE SURGEON RECEIVED A TISSUE TOO THICK TO CONTINUE ERROR MESSAGE AND CHANGED TO A BLUE SUREFORM 30 RELOAD. THE CUSTOMER WAS UNSURE IF THE BLADE CAME FROM A BLUE OR WHITE SUREFORM 30 RELOAD. THE BLADE WAS RETRIEVED DURING THE SAME PROCEDURE WHICH WAS COMPLETED ROBOTICALLY.
16-12-2024 | Malfunction | covidien lp llc north haven | 13-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC CARDIATRIC GASTRECTOMY, IN ESOPHAGECTOMY, THE DEVICE STOPPED FIRING HALFWAY. TO RESOLVE THE ISSUE, A NEW RELOAD OF THE SAME TYPE WAS USED WITH THE SAME ADAPTER AND HANDLE. THERE WAS NO PATIENT INJURY.
19-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING SET-UP PRIOR TO FIRING, THE RELOAD JAWS WOULD NOT OPEN AT ALL. THERE WAS NO TISSUE INVOLVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING SET-UP PRIOR TO FIRING, THE RELOAD JAWS WOULD NOT OPEN AT ALL. THERE WAS NO TISSUE INVOLVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
19-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, DURING SET-UP PRIOR TO FIRING, THE RELOAD JAWS WOULD NOT OPEN AT ALL. THERE WAS NO TISSUE INVOLVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
03-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
03-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
22-08-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | ENDO GIA | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC RECTAL RESECTION THROUGH THE ABDOMEN AND PERINEUM, THE STAPLES DID NOT COME OUT AFTER FIRING. THE PRODUCT WAS NOT USABLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE.
29-11-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SIGMOID RESECTION, THE TWO STAPLERS COULD NOT BE FIRED, AND A COMPONENT OF THE DEVICE BROKE OFF. ADDITIONALLY, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLES. THE PRODUCTS WERE REPLACED WITH ANOTHER HANDLE TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR SEVEN MINUTES. THERE WAS NO PATIENT INJURY.
29-11-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SIGMOID RESECTION, THE TWO STAPLERS COULD NOT BE FIRED, AND A COMPONENT OF THE DEVICES BROKE OFF. AD DITIONALLY, THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLES. THE PRODUCTS WERE REPLACED WITH ANOTHER HANDLE TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR SEVEN MINUTES. THERE WAS NO PATIENT INJURY.
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN RECTUM PROCEDURE THE INSTRUMENT HAS CUT BUT NOT CLAMPED BUT AFTERWARDS THE MAGAZINE COULD NOT BE CHANGED, OR ONLY WITH GREAT FORCE AND PLIERS. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED. OPENED A NEW INSTRUMENT AND PROCEDURE FINISHED WITHOUT PROBLEMS.
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN RECTUM PROCEDURE THE INSTRUMENT HAS CUT BUT NOT CLAMPED BUT AFTERWARDS THE MAGAZINE COULD NOT BE CHANGED, OR ONLY WITH GREAT FORCE AND PLIERS. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED. OPENED A NEW INSTRUMENT AND PROCEDURE FINISHED WITHOUT PROBLEMS.
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN RECTUM PROCEDURE THE INSTRUMENT HAS CUT BUT NOT CLAMPED BUT AFTERWARDS THE MAGAZINE COULD NOT BE CHANGED, OR ONLY WITH GREAT FORCE AND PLIERS. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED. OPENED A NEW INSTRUMENT AND PROCEDURE FINISHED WITHOUT PROBLEMS
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN RECTUM PROCEDURE THE INSTRUMENT HAS CUT BUT NOT CLAMPED BUT AFTERWARDS THE MAGAZINE COULD NOT BE CHANGED, OR ONLY WITH GREAT FORCE AND PLIERS. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED. OPENED A NEW INSTRUMENT AND PROCEDURE FINISHED WITHOUT PROBLEMS
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN RECTUM PROCEDURE THE INSTRUMENT HAS CUT BUT NOT CLAMPED BUT AFTERWARDS THE MAGAZINE COULD NOT BE CHANGED, OR ONLY WITH GREAT FORCE AND PLIERS. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED. OPENED A NEW INSTRUMENT AND PROCEDURE FINISHED WITHOUT PROBLEMS.
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN RECTUM PROCEDURE THE INSTRUMENT HAS CUT BUT NOT CLAMPED BUT AFTERWARDS THE MAGAZINE COULD NOT BE CHANGED, OR ONLY WITH GREAT FORCE AND PLIERS. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED. OPENED A NEW INSTRUMENT AND PROCEDURE FINISHED WITHOUT PROBLEMS.
27-12-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ROBOTIC SIGMO RESECTION, PROCEDURE WHEN THE GUT WAS TO BE STAPLED TOGETHER, IT WAS NOT POSSIBLE TO PRESS THE FIRING BUTTON OF THE CIRCULAR STAPLER. THEY CHANGED TO A NEW CIRCULAR STAPLER WHICH WORKED FINE. THE SURGERY WAS DELAYED 15 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
08-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED DURING AN EMERGENCY ILEUS OPERATION, WHERE THEY DECIDED TO USE ECHELON CONTOUR TO RESECT THE SMALL INTESTINE. THE INSTRUMENT WAS NEW. ONCE THE INSTRUMENT WAS ACTIVATED, THEY COULD NOT OPEN THE INSTRUMENT JAWS. THEY DECIDED TO FORCE THE RELOAD OUT OF THE JAW WHICH WAS LEFT ON THE SMALL INTESTINE. AFTER THAT, THEY DECIDED TO LOWER THE RESECTION CUT UNDER THE RELOAD AND WITH ANOTHER STAPLER (LINEAR, THEY DID NOT GIVE THE INFORMATION WHICH ONE) CUT THE INTESTINE. THE RELOAD FROM GCS40G WAS SENT TO THE PATHOLOGY DUE TO THE IMPOSSIBILITY OF REMOVING IT. THE PATIENT WAS NOT AT RISK BECAUSE THEY WERE USING IT ON THE SMALL INTESTINE AND THEY HAD THE ABILITY TO MOVE THE RESECTION INCISION.
08-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Insufficient Information | E2401 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED DURING AN EMERGENCY ILEUS OPERATION, WHERE THEY DECIDED TO USE ECHELON CONTOUR TO RESECT THE SMALL INTESTINE. THE INSTRUMENT WAS NEW. ONCE THE INSTRUMENT WAS ACTIVATED, THEY COULD NOT OPEN THE INSTRUMENT JAWS. THEY DECIDED TO FORCE THE RELOAD OUT OF THE JAW WHICH WAS LEFT ON THE SMALL INTESTINE. AFTER THAT, THEY DECIDED TO LOWER THE RESECTION CUT UNDER THE RELOAD AND WITH ANOTHER STAPLER (LINEAR, THEY DID NOT GIVE THE INFORMATION WHICH ONE) CUT THE INTESTINE. THE RELOAD FROM GCS40G WAS SENT TO THE PATHOLOGY DUE TO THE IMPOSSIBILITY OF REMOVING IT. THE PATIENT WAS NOT AT RISK BECAUSE THEY WERE USING IT ON THE SMALL INTESTINE AND THEY HAD THE ABILITY TO MOVE THE RESECTION INCISION.
15-10-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
27-11-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, WHEN FIRING, THE CARTRIDGE WAS HIGHER THAN EXPECTED, AND THE KNIFE DID NOT MOVE FORWARD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
19-11-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | ILS 29MM, CURVED | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
19-11-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | ILS 29MM, CURVED | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-12-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | 55MM SELECTABLE NEW TLC | Structural Problem | A0511 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, BEFORE FIRING AND NOTED THAT SOME STAPLES FELL OFF. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
03-01-2025 | Malfunction | covidien lp | 14-01-2025 | EEA |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: EEA CIRCULAR STAPLER WITH TRI-STAPLE MISFIRE.
17-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, WHEN STAPLING THE LUNG, BOTH RELOADS ONLY FIRED HALFWAY AND WOULD NOT FIRE ANY FURTHER. A THIRD RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | ENDOPATH |  |  | Laceration(s) | Event Description: THE STAPLER ABRUPTLY CUT THROUGH THE SEROSA AND MUSCULARIS, AND APPEAR TO LEAVE A SMALL OPENING IN THE MUCOSA ON THE MEDIAL SIDE. THIS SMALL OPENING WAS THEN STAPLED WITH A NEW STAPLER AT THE BREECH.
17-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, WHEN STAPLING THE LUNG, BOTH RELOADS ONLY FIRED HALFWAY AND WOULD NOT FIRE ANY FURTHER. A THIRD RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, WHEN STAPLING THE LUNG, BOTH RELOADS ONLY FIRED HALFWAY AND WOULD NOT FIRE ANY FURTHER. A THIRD RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LUNG WEDGE RESECTION, WHEN STAPLING THE LUNG, BOTH RELOADS ONLY FIRED HALFWAY AND WOULD NOT FIRE ANY FURTHER. A THIRD RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION SURGERY, WHEN RESECTING THE VESSEL, DURING THE FIRING, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE JAWS CLOSED COMPLETELY, HOWEVER, THE FIRING COULD NOT BE PERFORMED AND THE STAPLE COME OUT AROUND THE PROXIMAL SIDE OF THE CARTRIDGE¿S JAWS. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION SURGERY, WHEN RESECTING THE VESSEL, DURING THE FIRING, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND THE JAWS CLOSED COMPLETELY, HOWEVER, THE FIRING COULD NOT BE PERFORMED AND THE STAPLE COME OUT AROUND THE PROXIMAL SIDE OF THE CARTRIDGE¿S JAWS. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THEY CHANGED TO ANOTHER TWO TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
07-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THEY CHANGED TO ANOTHER TWO TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
07-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. THEY CHANGED TO ANOTHER TWO TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
27-12-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 45MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A SEMI-OPEN RESPIRATORY PROCEDURE, IT WAS OBSERVED THAT THE KNIFE STOPPED IN THE MIDDLE OF THE LUNG DISSECTION. THE KNIFE WAS MANAGED TO BE RETURNED, AND THE PATIENT WAS NOT AFFECTED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION REQUESTED.
06-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, IT WAS OBSERVED THAT DEFORMATION OF THE STAPLES AND AN INCOMPLETE STAPLE LINE OCCURRED DURING FIRING WHICH OCCURRED REPEATEDLY IN TWO CARTRIDGES. THE INNER ROW OF CLAMPS WAS INCOMPLETE BUT ONLY IN SOME PLACES. THERE WAS A DELAY OF 10 MINUTES TO COMPLETE THE PROCEDURE BY TREATING WITH SUTURE AND CLIPS. THERE WAS NO PATIENT CONSEQUENCE. NO ADDITIONAL INFORMATION PROVIDED.
06-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, IT WAS OBSERVED THAT DEFORMATION OF THE STAPLES AND AN INCOMPLETE STAPLE LINE OCCURRED DURING FIRING WHICH OCCURRED REPEATEDLY IN TWO CARTRIDGES. THE INNER ROW OF CLAMPS WAS INCOMPLETE BUT ONLY IN SOME PLACES. THERE WAS A DELAY OF 10 MINUTES TO COMPLETE THE PROCEDURE BY TREATING WITH SUTURE AND CLIPS. THERE WAS NO PATIENT CONSEQUENCE. NO ADDITIONAL INFORMATION PROVIDED.
20-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | EEA ORVIL | Adverse Event Without Identified Device or Use Problem | A24 | Internal Organ Perforation | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY THE RESIDENT ANESTHESIOLOGIST PASSED THE ORAL ANVIL AND PERFORATED THE LARYNX. THE PROCEDURE WAS EXTENDED BY FORTY-FIVE MINUTES. AND ULTIMATELY, THE SURGEON DECIDED NOT TO USE THE ORAL ANVIL AND PERFORMED A HAND SEWN ANASTOMOSIS TO COMPLETE THE PROCEDURE.
30-12-2024 | Injury | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 60MM LONG | Difficult to Open or Close | A051104 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THE FORCEPS NOT WORKING, DIFFICULTY CLOSING FORCEPS DURING STAPLING. THE STAPLE LINE RIPPED THE STOMACH TISSUE AND DAMAGED THE GASTRIC TISSUE. BECAUSE OF THE TISSUE DAMAGE CAUSED BY THE STAPLER, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM THE DAY AFTER HER SLEEVE AND A LARGER-THAN-EXPECTED CUT WAS MADE, PROLONGING HER HOSPITAL STAY.
30-12-2024 | Injury | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 60MM LONG | Difficult to Open or Close | A051104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THE FORCEPS NOT WORKING, DIFFICULTY CLOSING FORCEPS DURING STAPLING. THE STAPLE LINE RIPPED THE STOMACH TISSUE AND DAMAGED THE GASTRIC TISSUE. BECAUSE OF THE TISSUE DAMAGE CAUSED BY THE STAPLER, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM THE DAY AFTER HER SLEEVE AND A LARGER-THAN-EXPECTED CUT WAS MADE, PROLONGING HER HOSPITAL STAY.
30-12-2024 | Injury | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 60MM LONG | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THE FORCEPS NOT WORKING, DIFFICULTY CLOSING FORCEPS DURING STAPLING. THE STAPLE LINE RIPPED THE STOMACH TISSUE AND DAMAGED THE GASTRIC TISSUE. BECAUSE OF THE TISSUE DAMAGE CAUSED BY THE STAPLER, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM THE DAY AFTER HER SLEEVE AND A LARGER-THAN-EXPECTED CUT WAS MADE, PROLONGING HER HOSPITAL STAY.
30-12-2024 | Injury | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 60MM LONG | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THE FORCEPS NOT WORKING, DIFFICULTY CLOSING FORCEPS DURING STAPLING. THE STAPLE LINE RIPPED THE STOMACH TISSUE AND DAMAGED THE GASTRIC TISSUE. BECAUSE OF THE TISSUE DAMAGE CAUSED BY THE STAPLER, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM THE DAY AFTER HER SLEEVE AND A LARGER-THAN-EXPECTED CUT WAS MADE, PROLONGING HER HOSPITAL STAY.
30-12-2024 | Injury | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 60MM LONG | Patient Device Interaction Problem | A01 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THE FORCEPS NOT WORKING, DIFFICULTY CLOSING FORCEPS DURING STAPLING. THE STAPLE LINE RIPPED THE STOMACH TISSUE AND DAMAGED THE GASTRIC TISSUE. BECAUSE OF THE TISSUE DAMAGE CAUSED BY THE STAPLER, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM THE DAY AFTER HER SLEEVE AND A LARGER-THAN-EXPECTED CUT WAS MADE, PROLONGING HER HOSPITAL STAY.
30-12-2024 | Injury | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 60MM LONG | Patient Device Interaction Problem | A01 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THE FORCEPS NOT WORKING, DIFFICULTY CLOSING FORCEPS DURING STAPLING. THE STAPLE LINE RIPPED THE STOMACH TISSUE AND DAMAGED THE GASTRIC TISSUE. BECAUSE OF THE TISSUE DAMAGE CAUSED BY THE STAPLER, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM THE DAY AFTER HER SLEEVE AND A LARGER-THAN-EXPECTED CUT WAS MADE, PROLONGING HER HOSPITAL STAY.
30-12-2024 | Injury | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 60MM LONG | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THE FORCEPS NOT WORKING, DIFFICULTY CLOSING FORCEPS DURING STAPLING. THE STAPLE LINE RIPPED THE STOMACH TISSUE AND DAMAGED THE GASTRIC TISSUE. BECAUSE OF THE TISSUE DAMAGE CAUSED BY THE STAPLER, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM THE DAY AFTER HER SLEEVE AND A LARGER-THAN-EXPECTED CUT WAS MADE, PROLONGING HER HOSPITAL STAY.
30-12-2024 | Injury | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 60MM LONG | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY THE FORCEPS NOT WORKING, DIFFICULTY CLOSING FORCEPS DURING STAPLING. THE STAPLE LINE RIPPED THE STOMACH TISSUE AND DAMAGED THE GASTRIC TISSUE. BECAUSE OF THE TISSUE DAMAGE CAUSED BY THE STAPLER, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM THE DAY AFTER HER SLEEVE AND A LARGER-THAN-EXPECTED CUT WAS MADE, PROLONGING HER HOSPITAL STAY.
02-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
18-12-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A LITTLE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
13-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG RESECTION, THE DEVICE COULD NOT BE FIRED AS THERE WAS NO RESISTANCE ON THE HANDLE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK DUE TO THICK TISSUE.
13-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LUNG RESECTION, THE DEVICE COULD NOT BE FIRED AS THERE WAS NO RESISTANCE ON THE HANDLE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK DUE TO THICK TISSUE.
07-01-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | GST GOLD RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE STOPPED WHILE FIRING. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
27-11-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A RIGHT UPPER LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLER STOPPED VERY QUICKLY MID-FIRING ON THE ARTERY. THEY GOT A NEW STAPLER AND RELOAD READY FOR WHEN THEY REMOVED THE INITIAL STAPLER. THERE WAS SOME BLEEDING WHEN THEY OPENED THE JAWS AND THEY WERE ABLE TO USE THE NEW STAPLER AND RELOAD TO FINISH THE TRANSECTION OF THE ARTERY TO COMPLETE THE PROCEDURE WITH A DELAY OF THREE MINUTES. THERE WAS NO PATIENT CONSEQUENCE. ADDITIONAL INFORMATION REQUESTED.
27-11-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT UPPER LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLER STOPPED VERY QUICKLY MID-FIRING ON THE ARTERY. THEY GOT A NEW STAPLER AND RELOAD READY FOR WHEN THEY REMOVED THE INITIAL STAPLER. THERE WAS SOME BLEEDING WHEN THEY OPENED THE JAWS AND THEY WERE ABLE TO USE THE NEW STAPLER AND RELOAD TO FINISH THE TRANSECTION OF THE ARTERY TO COMPLETE THE PROCEDURE WITH A DELAY OF THREE MINUTES. THERE WAS NO PATIENT CONSEQUENCE. ADDITIONAL INFORMATION REQUESTED.
27-11-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A RIGHT UPPER LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLER STOPPED VERY QUICKLY MID-FIRING ON THE ARTERY. THEY GOT A NEW STAPLER AND RELOAD READY FOR WHEN THEY REMOVED THE INITIAL STAPLER. THERE WAS SOME BLEEDING WHEN THEY OPENED THE JAWS AND THEY WERE ABLE TO USE THE NEW STAPLER AND RELOAD TO FINISH THE TRANSECTION OF THE ARTERY TO COMPLETE THE PROCEDURE WITH A DELAY OF THREE MINUTES. THERE WAS NO PATIENT CONSEQUENCE. ADDITIONAL INFORMATION REQUESTED.
27-11-2024 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RIGHT UPPER LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLER STOPPED VERY QUICKLY MID-FIRING ON THE ARTERY. THEY GOT A NEW STAPLER AND RELOAD READY FOR WHEN THEY REMOVED THE INITIAL STAPLER. THERE WAS SOME BLEEDING WHEN THEY OPENED THE JAWS AND THEY WERE ABLE TO USE THE NEW STAPLER AND RELOAD TO FINISH THE TRANSECTION OF THE ARTERY TO COMPLETE THE PROCEDURE WITH A DELAY OF THREE MINUTES. THERE WAS NO PATIENT CONSEQUENCE. ADDITIONAL INFORMATION REQUESTED.
07-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | FLEX 60 ARTICULATING | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING THE DEVICE. CHANGED TO ANOTHER ONE TO CONTINUE THE SURGERY BUT THE DEVICE COULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING THE DEVICE. CHANGED TO ANOTHER ONE TO CONTINUE THE SURGERY BUT THE DEVICE COULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, A HEPATIC STAPLER WAS USED FOUR TIMES. WHEN THEY FIRED ON THE HEPATIC VEIN ON THE 4TH FIRING, IT WAS OBSERVED THAT BOTH STAPLES AND VESSEL OPENED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
29-11-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A NEPHRECTOMY, THE DEVICE WAS DIFFICULT TO FIRE. TO RESOLVE THE ISSUE A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE DUE TO THICK TISSUE.
29-11-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A NEPHRECTOMY, THE DEVICE WAS DIFFICULT TO FIRE. TO RESOLVE THE ISSUE A NEW HANDLE WAS USED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE DUE TO THICK TISSUE.
06-01-2025 | Malfunction | ethicon endo-surgery | 14-01-2025 | GST BLUE RELOAD 60MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY PROCEDURE, IT WAS OBSERVED THAT RELOAD OF THE ENDOSTAPLER PRESENTED PROBLEMS AT THE TIME OF STAPLING. THE DEVICE WAS ASSEMBLED ACCORDING TO INSTRUCTION BUT AT THE TIME OF SHOOTING THERE WAS NO CORRECT FORMATION OF THE STAPLE LINE. IT WAS NECESSARY TO OPEN ANOTHER EXTRA RELOAD. THE INCIDENT DID NOT AFFECT THE PATIENT OR SURGEON IN A PREJUDICIAL WAY, THERE WERE NO FRAGMENTS FORMED IN THE PATIENT. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A THORACOSCOPIC RIGHT UPPER LOBECTOMY, WHEN IN THE LUNG PARENCHYMA, THE DEVICE WAS STIFF TO SQUEEZE. THERE WAS NO ISSUE WITH THE STAPLE LINE BUT THERE WAS FAILURE TO RETRACT KNIFE BLADE AFTER FIRING. THE JAWS COULD NOT BE REMOVED FROM THE TISSUE. THE RELOAD HAD TISSUE AND A STAPLE REMAINING AT THE ROOT OF THE JAWS. THE RETURN KNOB WAS PULLED BACK FORCEFULLY WITH BOTH HANDS TO REMOVE THE DEVICE FROM THE TISSUE. THE SURGICAL TEAM STRUGGLED TO PULL BACK THE DEVICE FOR 15 MINUTES. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A THORACOSCOPIC RIGHT UPPER LOBECTOMY, WHEN IN THE LUNG PARENCHYMA, THE DEVICE WAS STIFF TO SQUEEZE. THERE WAS NO ISSUE WITH THE STAPLE LINE BUT THERE WAS FAILURE TO RETRACT KNIFE BLADE AFTER FIRING. THE JAWS COULD NOT BE REMOVED FROM THE TISSUE. THE RELOAD HAD TISSUE AND A STAPLE REMAINING AT THE ROOT OF THE JAWS. THE RETURN KNOB WAS PULLED BACK FORCEFULLY WITH BOTH HANDS TO REMOVE THE DEVICE FROM THE TISSUE. THE SURGICAL TEAM STRUGGLED TO PULL BACK THE DEVICE FOR 15 MINUTES. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A THORACOSCOPIC RIGHT UPPER LOBECTOMY, WHEN IN THE LUNG PARENCHYMA, THE DEVICE WAS STIFF TO SQUEEZE. THERE WAS NO ISSUE WITH THE STAPLE LINE BUT THERE WAS FAILURE TO RETRACT KNIFE BLADE AFTER FIRING. THE JAWS COULD NOT BE REMOVED FROM THE TISSUE. THE RELOAD HAD TISSUE AND A STAPLE REMAINING AT THE ROOT OF THE JAWS. THE RETURN KNOB WAS PULLED BACK FORCEFULLY WITH BOTH HANDS TO REMOVE THE DEVICE FROM THE TISSUE. THE SURGICAL TEAM STRUGGLED TO PULL BACK THE DEVICE FOR 15 MINUTES. THERE WAS NO PATIENT INJURY.
17-07-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE HANDLE WAS CHARGED, AND BEFORE THE FIRST FIRING, IT WENT TO ZERO (0) BATTERY. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT AN ANALYSIS OF THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
17-07-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THE HANDLE WAS CHARGED, AND BEFORE THE FIRST FIRING, IT WENT TO ZERO (0) BATTERY. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT AN ANALYSIS OF THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
25-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A THORACOSCOPIC RIGHT UPPER LOBECTOMY, WHEN IN THE LUNG PARENCHYMA, THE DEVICE WAS STIFF TO SQUEEZE. THERE WAS NO ISSUE WITH THE STAPLE LINE BUT THERE WAS FAILURE TO RETRACT KNIFE BLADE AFTER FIRING. THE JAWS COULD NOT BE REMOVED FROM THE TISSUE. THE RELOAD HAD TISSUE AND A STAPLE REMAINING AT THE ROOT OF THE JAWS. THE RETURN KNOB WAS PULLED BACK FORCEFULLY WITH BOTH HANDS TO REMOVE THE DEVICE FROM THE TISSUE. THE SURGICAL TEAM STRUGGLED TO PULL BACK THE DEVICE FOR 15 MINUTES. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | ENDO GIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A THORACOSCOPIC RIGHT UPPER LOBECTOMY, WHEN IN THE LUNG PARENCHYMA, THE DEVICE WAS STIFF TO SQUEEZE. THERE WAS NO ISSUE WITH THE STAPLE LINE BUT THERE WAS FAILURE TO RETRACT KNIFE BLADE AFTER FIRING. THE JAWS COULD NOT BE REMOVED FROM THE TISSUE. THE RELOAD HAD TISSUE AND A STAPLE REMAINING AT THE ROOT OF THE JAWS. THE RETURN KNOB WAS PULLED BACK FORCEFULLY WITH BOTH HANDS TO REMOVE THE DEVICE FROM THE TISSUE. THE SURGICAL TEAM STRUGGLED TO PULL BACK THE DEVICE FOR 15 MINUTES. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A THORACOSCOPIC RIGHT UPPER LOBECTOMY, WHEN IN THE LUNG PARENCHYMA, THE DEVICE WAS STIFF TO SQUEEZE. THERE WAS NO ISSUE WITH THE STAPLE LINE BUT THERE WAS FAILURE TO RETRACT KNIFE BLADE AFTER FIRING. THE JAWS COULD NOT BE REMOVED FROM THE TISSUE. THE RELOAD HAD TISSUE AND A STAPLE REMAINING AT THE ROOT OF THE JAWS. THE RETURN KNOB WAS PULLED BACK FORCEFULLY WITH BOTH HANDS TO REMOVE THE DEVICE FROM THE TISSUE. THE SURGICAL TEAM STRUGGLED TO PULL BACK THE DEVICE FOR 15 MINUTES. THERE WAS NO PATIENT INJURY.
22-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE COMPUTED TOMOGRAPHY, THERE WERE FIRING ISSUES WITH THE RELOAD COMPONENT OF THE STAPLER ON THE ARTERY. THE STAPLES DEPLOYED ON THE PATIENT TISSUE, BUT A KNIFE DID NOT PASS BETWEEN THEM. AS A RESULT, THE SURGEON REMOVED THE STAPLER AND RELOAD, OPTING TO USE TWO SUTURES AS A STAPLE LINE REPLACEMENT. THE PROCEDURE CONTINUED WITH THE USE OF THE STAPLER AND SHORT ADAPTER, AND ADDITIONAL RELOADS WERE USED WITHOUT FURTHER ISSUES.
20-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY (PD), IN GASTRECTOMY, AFTER FIRING THE DEVICE, THE USER WAS UNABLE TO PULL BACK THE KNIFE BLADE. THE TISSUE BECAME ENTANGLED, AND THE JAWS OF THE DEVICE COULD NOT BE REMOVED. THERE WAS NO PATIENT INJURY.
20-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY (PD), IN GASTRECTOMY, AFTER FIRING THE DEVICE, THE USER WAS UNABLE TO PULL BACK THE KNIFE BLADE. THE TISSUE BECAME ENTANGLED, AND THE JAWS OF THE DEVICE COULD NOT BE REMOVED. THERE WAS NO PATIENT INJURY.
20-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY (PD), IN GASTRECTOMY, AFTER FIRING THE DEVICE, THE USER WAS UNABLE TO PULL BACK THE KNIFE BLADE. THE TISSUE BECAME ENTANGLED, AND THE JAWS OF THE DEVICE COULD NOT BE REMOVED. THERE WAS NO PATIENT INJURY.
20-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROTOMY PANCREATICODUODENECTOMY (PD), IN GASTRECTOMY, AFTER FIRING THE DEVICE, THE USER WAS UNABLE TO PULL BACK THE KNIFE BLADE. THE TISSUE BECAME ENTANGLED, AND THE JAWS OF THE DEVICE COULD NOT BE REMOVED. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | covidien | 14-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZED. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | covidien | 14-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZED. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | covidien | 14-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZED. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | covidien | 14-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZED. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZE. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZE. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZE. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZE. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZE. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | us surgical puerto rico | 14-01-2025 | ENDO GIA ULTRA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY OF A VIDEO-ASSISTED THORACIC SURGERY (VATS) PARTIAL LUNG RESECTION, FIRING COULD NOT BE PERFORMED AT ALL, AS THE HANDLE COULD NOT BE SQUEEZE. IT WAS ALSO NOTED THAT IT WAS NOT POSSIBLE TO OPEN THE JAWS IN THE ABDOMINAL CAVITY. A NEW ADDITIONAL RELOAD WAS BROUGHT IN AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-10-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE HANDLE TURNED ON, BUT THE DISPLAY DID NOT OPERATE. THE LCD OF THE POWER HANDLE WAS NOT SHOWING. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AFTER TAKING A PART THE HANDLE, MOTOR TWO WAS FOUND TO BE LOOSE.
03-10-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Application Program Freezes, Becomes Nonfunctional |  | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE HANDLE TURNED ON, BUT THE DISPLAY DID NOT OPERATE. THE LCD OF THE POWER HANDLE WAS NOT SHOWING. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AFTER TAKING A PART THE HANDLE, MOTOR TWO WAS FOUND TO BE LOOSE.
06-01-2025 | Injury | ethicon endo-surgery | 14-01-2025 | ECHELON 3000 60MM LONG | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A ROBOT-ASSISTED TOTAL GASTRECTOMY, THE SALES-REP HEARD FROM ANOTHER CUSTOMER THAT THE PATIENT WHO HAD USED ECHELON AT THE END OF THE YEAR OCCURRED A LEAKAGE FROM THE DUODENUM STUMP AND TOOK THE RE-OPERATION, THE PRIMARY SURGEON, COULD NOT MEET, SO AN INTERVIEW FOR DETAILS IS SCHEDULED FOR A LATER DATE. THE CARTRIDGE COLOR WAS BLUE. NO FURTHER INFORMATION IS AVAILABLE.
25-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Unraveled Material | A040508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED PANCREATICODUODENECTOMY, WHEN THE POWER HANDLE WAS INSERTED INTO THE SHELL, AN ATTEMPT WAS MADE TO CLOSE THE SHELL, BUT IT COULD NOT BE PROPERLY CLOSED. A NEW SHELL WAS OPENED AND USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, AFTER FIRING AT THE DUODENAL RESECTION, THE RELOAD MOVED IN AN UNEXPECTED DIRECTION. IT WAS NOTED THAT THE TREATMENT INTERVENTION WAS NOT PERFORMED/NOT NECESSARY. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | SIGNIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, AFTER FIRING AT THE DUODENAL RESECTION, THE RELOAD MOVED IN AN UNEXPECTED DIRECTION. IT WAS NOTED THAT THE TREATMENT INTERVENTION WAS NOT PERFORMED/NOT NECESSARY. THERE WAS NO PATIENT INJURY.
17-12-2024 | Malfunction | covidien lp llc north haven | 14-01-2025 | ENDO GIA | Unintended System Motion | A051206 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, AFTER FIRING AT THE DUODENAL RESECTION, THE RELOAD MOVED IN AN UNEXPECTED DIRECTION. IT WAS NOTED THAT THE TREATMENT INTERVENTION WAS NOT PERFORMED/NOT NECESSARY. THERE WAS NO PATIENT INJURY.
17-12-2024 | Injury | intuitive surgical | 14-01-2025 | SUREFORM | Firing Problem | A0505 | Hemorrhage/Blood Loss/Bleeding | Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.
17-12-2024 | Injury | intuitive surgical | 14-01-2025 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.
24-12-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHEN DISSOCIATING BLOOD VESSELS, THE TWO RELOADS DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. IT WAS NOTED THAT THE TWO RELOADS WERE USED FOR THE SAME PATIENT AND PROCEDURE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC THORACOSCOPIC LOBECTOMY, WHEN DISSOCIATING BLOOD VESSELS, THE TWO RELOADS DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. IT WAS NOTED THAT THE TWO RELOADS WERE USED FOR THE SAME PATIENT AND PROCEDURE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER INTRAOPERATIVELY TO THE ULTRA-LOW ANTERIOR RESECTION, THE SURGEON FIRED THE DEVICE BUT IT SUDDENLY STOP. LATER THE SURGEON RETRIEVED THE BLADE AND REMOVE THE CARTRIDGE A SECOND RELOAD WAS USED BUT IT HAD THE SAME ISSUE. A THIRD RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER INTRAOPERATIVELY TO THE ULTRA-LOW ANTERIOR RESECTION, THE SURGEON FIRED THE DEVICE BUT IT SUDDENLY STOP. LATER THE SURGEON RETRIEVED THE BLADE AND REMOVE THE CARTRIDGE A SECOND RELOAD WAS USED BUT IT HAD THE SAME ISSUE. A THIRD RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER INTRAOPERATIVELY TO THE ULTRA-LOW ANTERIOR RESECTION, THE SURGEON FIRED THE DEVICE BUT IT SUDDENLY STOP. LATER THE SURGEON RETRIEVED THE BLADE AND REMOVE THE CARTRIDGE A SECOND RELOAD WAS USED BUT IT HAD THE SAME ISSUE. A THIRD RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
24-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER INTRAOPERATIVELY TO THE ULTRA-LOW ANTERIOR RESECTION, THE SURGEON FIRED THE DEVICE BUT IT SUDDENLY STOP. LATER THE SURGEON RETRIEVED THE BLADE AND REMOVE THE CARTRIDGE A SECOND RELOAD WAS USED BUT IT HAD THE SAME ISSUE. A THIRD RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
18-11-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED STOMACH RESECTION, THE JAWS DID NOT CLOSE PROPERLY. IT WAS TO THE EXTENT THAT IT COULD NOT FIT THROUGH THE PORT. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-11-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED STOMACH RESECTION, THE JAWS DID NOT CLOSE PROPERLY. IT WAS TO THE EXTENT THAT IT COULD NOT FIT THROUGH THE PORT. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-11-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED STOMACH RESECTION, THE JAWS DID NOT CLOSE PROPERLY. IT WAS TO THE EXTENT THAT IT COULD NOT FIT THROUGH THE PORT. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-11-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOT-ASSISTED STOMACH RESECTION, THE JAWS DID NOT CLOSE PROPERLY. IT WAS TO THE EXTENT THAT IT COULD NOT FIT THROUGH THE PORT. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-11-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE JAWS DID NOT CLOSE PROPERLY. IT WAS TO THE EXTENT THAT IT COULDN'T FIT THROUGH THE PORT. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE.
18-11-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A PROCEDURE, THE JAWS DID NOT CLOSE PROPERLY. IT WAS TO THE EXTENT THAT IT COULDN'T FIT THROUGH THE PORT. A NEW CARTRIDGE WAS USED TO RESOLVE THE ISSUE.
23-01-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHEN BEFORE THE LUNG WAS RESECTED, THE FIRING STOPPED PARTIALLY. IT WAS NOTED THAT THE KNIFE BLADE WAS RETRACTED. A NEW RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-01-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHEN BEFORE THE LUNG WAS RESECTED, THE FIRING STOPPED PARTIALLY. IT WAS NOTED THAT THE KNIFE BLADE WAS RETRACTED. A NEW RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
23-01-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHEN BEFORE THE LUNG WAS RESECTED, THE FIRING STOPPED PARTIALLY. IT WAS NOTED THAT THE KNIFE BLADE WAS RETRACTED. A NEW RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC PARTIAL LUNG/PULMONARY RESECTION, WHEN CUTTING THE LUNG PARENCHYMA THE TISSUE WAS CLAMPED. THE GREEN BUTTON WAS PRESSED AND THE FIRING WAS TRIED. BUT THE HANDLE COULD NOT BE SQUEEZED. UPON CHECKING THE PROXIMAL SIDE OF THE CARTRIDGE, IT WAS FOUND TO BE PRE-FIRED. AFTER TAKING OUT A NEW CARTRIDGE AND CONNECTING IT TO THE SAME HANDLE, IT WAS ABLE TO BE USED WITHOUT ANY ISSUES. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A THORACOSCOPIC PARTIAL LUNG/PULMONARY RESECTION, WHEN CUTTING THE LUNG PARENCHYMA THE TISSUE WAS CLAMPED. THE GREEN BUTTON WAS PRESSED AND THE FIRING WAS TRIED. BUT THE HANDLE COULD NOT BE SQUEEZED. UPON CHECKING THE PROXIMAL SIDE OF THE CARTRIDGE, IT WAS FOUND TO BE PRE-FIRED. AFTER TAKING OUT A NEW CARTRIDGE AND CONNECTING IT TO THE SAME HANDLE, IT WAS ABLE TO BE USED WITHOUT ANY ISSUES. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | ethicon endo-surgery | 15-01-2025 | ECHELON FLEX |  |  | Foreign Body In Patient | Event Description: WE WERE DOING A LIVER RESECTION CASE AND THE ETHICON STAPLER MISFIRED. WE HAD TO GO IN AND DO THE MANUAL OVERRIDE TO RELEASE THE STAPLER AFTER IT HAD MISFIRED. ITEM: ETHICON POWERED VASCULAR STAPLER. LOT #: C9E623, REFERENCE #: PVE35A. MANUFACTURER RESPONSE FOR STAPLE, IMPLANTABLE, ECHELON FLEX (PER SITE REPORTER). THEY HAVE PROVIDED A RETURN KIT FOR THE SAMPLE AND WE ARE AWAITING EVALUATION.
09-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATIC TAIL RESECTION, ON THE PANCREAS, IT WAS NOTED THAT THE KNIFE DID NOT RUN WHEN THIS PRODUCT WAS USED FOR PANCREATIC TAIL RESECTION. THE CARTRIDGE WAS CHECKED, BUT THERE WAS A POSSIBILITY THAT THE FIRING HAD NOT BEEN COMPLETED IN THE FIRST PLACE. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PANCREATIC TAIL RESECTION, ON THE PANCREAS, IT WAS NOTED THAT THE KNIFE DID NOT RUN WHEN THIS PRODUCT WAS USED FOR PANCREATIC TAIL RESECTION. THE CARTRIDGE WAS CHECKED, BUT THERE WAS A POSSIBILITY THAT THE FIRING HAD NOT BEEN COMPLETED IN THE FIRST PLACE. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOID PROCEDURE, THE HEMORRHOID STAPLER FAILED TO FIRE. THE HANDLE WAS FULLY SQUEEZED UNTIL IT CAME IN CONTACT WITH THE INSTRUMENT BODY, AND SIZERS WERE USED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY OF RECTAL CANCER, WHILE DETACHING THE TISSUE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. A NEW RELOAD WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
25-12-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
02-12-2024 | Injury | covidien lp llc north haven | 15-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2021 AND 2023, A PROSPECTIVE STUDY REVIEWED 88 CASES OF FAILED PRIMARY RESTRICTIVE PROCEDURES THAT SUBSEQUENTLY WERE CANDIDATES FOR BARIATRIC SURGERY REVISION BETWEEN 2021 AND 2023. PATIENTS UNDERWENT EITHER ROUX-EN-Y GASTRIC BYPASS OR ONE-ANASTOMOSIS GASTRIC BYPASS. THE ENDO GIA WAS USED TO STAPLE THE STOMACH. ENDO GIA WAS USED FOR ANASTOMOSES. POTENTIAL DEVICE RELATED POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAKAGE, BLEEDING, AND SUBPHRENIC COLLECTION. NON-DEVICE RELATED COMPLICATIONS INCLUDED DEEP VEIN THROMBOSIS AND PORT SITE INFECTION. NO INTERVENTIONS WERE MENTIONED IN THE ARTICLE. THE ADVERSE EFFECT WAS PROBABLY RELATED TO THE PROCEDURE BUT NOT DEVICE-RELATED. LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS VERSUS ONE ANASTOMOSIS GASTRIC BYPASS FOR REVISIONAL BARIATRIC SURGERY: A PROPENSITY SCORE MATCHED STUDY. DR. MOHAMED ABDALLA SALMAN, 2024, BARIATRIC SURGICAL PRACTICE AND PATIENT CARE, 10.1089/BARI.2023.0060.
02-12-2024 | Injury | covidien lp llc north haven | 15-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2021 AND 2023, A PROSPECTIVE STUDY REVIEWED 88 CASES OF FAILED PRIMARY RESTRICTIVE PROCEDURES THAT SUBSEQUENTLY WERE CANDIDATES FOR BARIATRIC SURGERY REVISION BETWEEN 2021 AND 2023. PATIENTS UNDERWENT EITHER ROUX-EN-Y GASTRIC BYPASS OR ONE-ANASTOMOSIS GASTRIC BYPASS. THE ENDO GIA WAS USED TO STAPLE THE STOMACH. ENDO GIA WAS USED FOR ANASTOMOSES. POTENTIAL DEVICE RELATED POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAKAGE, BLEEDING, AND SUBPHRENIC COLLECTION. NON-DEVICE RELATED COMPLICATIONS INCLUDED DEEP VEIN THROMBOSIS AND PORT SITE INFECTION. NO INTERVENTIONS WERE MENTIONED IN THE ARTICLE. THE ADVERSE EFFECT WAS PROBABLY RELATED TO THE PROCEDURE BUT NOT DEVICE-RELATED. LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS VERSUS ONE ANASTOMOSIS GASTRIC BYPASS FOR REVISIONAL BARIATRIC SURGERY: A PROPENSITY SCORE MATCHED STUDY. DR. MOHAMED ABDALLA SALMAN, 2024, BARIATRIC SURGICAL PRACTICE AND PATIENT CARE, 10.1089/BARI.2023.0060.
02-12-2024 | Injury | covidien lp llc north haven | 15-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2021 AND 2023, A PROSPECTIVE STUDY REVIEWED 88 CASES OF FAILED PRIMARY RESTRICTIVE PROCEDURES THAT SUBSEQUENTLY WERE CANDIDATES FOR BARIATRIC SURGERY REVISION BETWEEN 2021 AND 2023. PATIENTS UNDERWENT EITHER ROUX-EN-Y GASTRIC BYPASS OR ONE-ANASTOMOSIS GASTRIC BYPASS. THE ENDO GIA WAS USED TO STAPLE THE STOMACH. ENDO GIA WAS USED FOR ANASTOMOSES. POTENTIAL DEVICE RELATED POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAKAGE, BLEEDING, AND SUBPHRENIC COLLECTION. NON-DEVICE RELATED COMPLICATIONS INCLUDED DEEP VEIN THROMBOSIS AND PORT SITE INFECTION. NO INTERVENTIONS WERE MENTIONED IN THE ARTICLE. THE ADVERSE EFFECT WAS PROBABLY RELATED TO THE PROCEDURE BUT NOT DEVICE-RELATED. LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS VERSUS ONE ANASTOMOSIS GASTRIC BYPASS FOR REVISIONAL BARIATRIC SURGERY: A PROPENSITY SCORE MATCHED STUDY. DR. MOHAMED ABDALLA SALMAN, 2024, BARIATRIC SURGICAL PRACTICE AND PATIENT CARE, 10.1089/BARI.2023.0060.
02-12-2024 | Injury | covidien lp llc north haven | 15-01-2025 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Failure to Anastomose | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2021 AND 2023, A PROSPECTIVE STUDY REVIEWED 88 CASES OF FAILED PRIMARY RESTRICTIVE PROCEDURES THAT SUBSEQUENTLY WERE CANDIDATES FOR BARIATRIC SURGERY REVISION BETWEEN 2021 AND 2023. PATIENTS UNDERWENT EITHER ROUX-EN-Y GASTRIC BYPASS OR ONE-ANASTOMOSIS GASTRIC BYPASS. THE ENDO GIA WAS USED TO STAPLE THE STOMACH. ENDO GIA WAS USED FOR ANASTOMOSES. POTENTIAL DEVICE RELATED POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAKAGE, BLEEDING, AND SUBPHRENIC COLLECTION. NON-DEVICE RELATED COMPLICATIONS INCLUDED DEEP VEIN THROMBOSIS AND PORT SITE INFECTION. NO INTERVENTIONS WERE MENTIONED IN THE ARTICLE. THE ADVERSE EFFECT WAS PROBABLY RELATED TO THE PROCEDURE BUT NOT DEVICE-RELATED. LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS VERSUS ONE ANASTOMOSIS GASTRIC BYPASS FOR REVISIONAL BARIATRIC SURGERY: A PROPENSITY SCORE MATCHED STUDY. DR. MOHAMED ABDALLA SALMAN, 2024, BARIATRIC SURGICAL PRACTICE AND PATIENT CARE, 10.1089/BARI.2023.0060.
02-12-2024 | Injury | covidien lp llc north haven | 15-01-2025 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Abscess | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2021 AND 2023, A PROSPECTIVE STUDY REVIEWED 88 CASES OF FAILED PRIMARY RESTRICTIVE PROCEDURES THAT SUBSEQUENTLY WERE CANDIDATES FOR BARIATRIC SURGERY REVISION BETWEEN 2021 AND 2023. PATIENTS UNDERWENT EITHER ROUX-EN-Y GASTRIC BYPASS OR ONE-ANASTOMOSIS GASTRIC BYPASS. THE ENDO GIA WAS USED TO STAPLE THE STOMACH. ENDO GIA WAS USED FOR ANASTOMOSES. POTENTIAL DEVICE RELATED POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAKAGE, BLEEDING, AND SUBPHRENIC COLLECTION. NON-DEVICE RELATED COMPLICATIONS INCLUDED DEEP VEIN THROMBOSIS AND PORT SITE INFECTION. NO INTERVENTIONS WERE MENTIONED IN THE ARTICLE. THE ADVERSE EFFECT WAS PROBABLY RELATED TO THE PROCEDURE BUT NOT DEVICE-RELATED. LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS VERSUS ONE ANASTOMOSIS GASTRIC BYPASS FOR REVISIONAL BARIATRIC SURGERY: A PROPENSITY SCORE MATCHED STUDY. DR. MOHAMED ABDALLA SALMAN, 2024, BARIATRIC SURGICAL PRACTICE AND PATIENT CARE, 10.1089/BARI.2023.0060.
02-12-2024 | Injury | covidien lp llc north haven | 15-01-2025 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY PERFORMED BETWEEN 2021 AND 2023, A PROSPECTIVE STUDY REVIEWED 88 CASES OF FAILED PRIMARY RESTRICTIVE PROCEDURES THAT SUBSEQUENTLY WERE CANDIDATES FOR BARIATRIC SURGERY REVISION BETWEEN 2021 AND 2023. PATIENTS UNDERWENT EITHER ROUX-EN-Y GASTRIC BYPASS OR ONE-ANASTOMOSIS GASTRIC BYPASS. THE ENDO GIA WAS USED TO STAPLE THE STOMACH. ENDO GIA WAS USED FOR ANASTOMOSES. POTENTIAL DEVICE RELATED POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAKAGE, BLEEDING, AND SUBPHRENIC COLLECTION. NON-DEVICE RELATED COMPLICATIONS INCLUDED DEEP VEIN THROMBOSIS AND PORT SITE INFECTION. NO INTERVENTIONS WERE MENTIONED IN THE ARTICLE. THE ADVERSE EFFECT WAS PROBABLY RELATED TO THE PROCEDURE BUT NOT DEVICE-RELATED. LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS VERSUS ONE ANASTOMOSIS GASTRIC BYPASS FOR REVISIONAL BARIATRIC SURGERY: A PROPENSITY SCORE MATCHED STUDY. DR. MOHAMED ABDALLA SALMAN, 2024, BARIATRIC SURGICAL PRACTICE AND PATIENT CARE, 10.1089/BARI.2023.0060.
20-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE COLECTOMY, WHEN APPLIED ON THE COLON, AND FOR SIDE-TO-SIDE ANASTOMOSIS, THE TISSUE WAS CAUGHT ON THE TIP OF THE CARTRIDGE AND WAS IN A STATE, WHERE THE JAWS COULD NOT BE REMOVED, ADDITIONALLY, THE RETRACTION OF KNIFE BLADE WAS DIFFICULT OR NOT POSSIBLE FROM THE TISSUE. SO IT WAS RESECTED AND REMOVED. THE RESECTED PART WAS MANUALLY SUTURED. THE SURGICAL TIME WAS EXTENDED FOR ABOUT 20 MINUTES.
20-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE COLECTOMY, WHEN APPLIED ON THE COLON, AND FOR SIDE-TO-SIDE ANASTOMOSIS, THE TISSUE WAS CAUGHT ON THE TIP OF THE CARTRIDGE AND WAS IN A STATE, WHERE THE JAWS COULD NOT BE REMOVED, ADDITIONALLY, THE RETRACTION OF KNIFE BLADE WAS DIFFICULT OR NOT POSSIBLE FROM THE TISSUE. SO IT WAS RESECTED AND REMOVED. THE RESECTED PART WAS MANUALLY SUTURED. THE SURGICAL TIME WAS EXTENDED FOR ABOUT 20 MINUTES.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Human-Device Interface Problem | A22 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Break | A0401 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH WEDGE RESECTION, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM DID NOT RETRACT, AND THE CARTRIDGE JAWS DID NOT OPEN. THE SURGEON USED AN ENERGY DEVICE TO CUT THE STAPLED AREA AND THEN SUTURED IT WITH SUTURE. THE SURGEON TRIED TO OPEN THE CARTRIDGE USING A MANUAL TOOL BUT FAILED.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH WEDGE RESECTION, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM DID NOT RETRACT, AND THE CARTRIDGE JAWS DID NOT OPEN. THE SURGEON USED AN ENERGY DEVICE TO CUT THE STAPLED AREA AND THEN SUTURED IT WITH SUTURE. THE SURGEON TRIED TO OPEN THE CARTRIDGE USING A MANUAL TOOL BUT FAILED.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH WEDGE RESECTION, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM DID NOT RETRACT, AND THE CARTRIDGE JAWS DID NOT OPEN. THE SURGEON USED AN ENERGY DEVICE TO CUT THE STAPLED AREA AND THEN SUTURED IT WITH SUTURE. THE SURGEON TRIED TO OPEN THE CARTRIDGE USING A MANUAL TOOL BUT FAILED.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH WEDGE RESECTION, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM DID NOT RETRACT, AND THE CARTRIDGE JAWS DID NOT OPEN. THE SURGEON USED AN ENERGY DEVICE TO CUT THE STAPLED AREA AND THEN SUTURED IT WITH SUTURE. THE SURGEON TRIED TO OPEN THE CARTRIDGE USING A MANUAL TOOL BUT FAILED.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH WEDGE RESECTION, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM DID NOT RETRACT, AND THE CARTRIDGE JAWS DID NOT OPEN. THE SURGEON USED AN ENERGY DEVICE TO CUT THE STAPLED AREA AND THEN SUTURED IT WITH SUTURE. THE SURGEON TRIED TO OPEN THE CARTRIDGE USING A MANUAL TOOL BUT FAILED.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH WEDGE RESECTION, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM DID NOT RETRACT, AND THE CARTRIDGE JAWS DID NOT OPEN. THE SURGEON USED AN ENERGY DEVICE TO CUT THE STAPLED AREA AND THEN SUTURED IT WITH SUTURE. THE SURGEON TRIED TO OPEN THE CARTRIDGE USING A MANUAL TOOL BUT FAILED.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A STOMACH WEDGE RESECTION, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM DID NOT RETRACT, AND THE CARTRIDGE JAWS DID NOT OPEN. THE SURGEON USED AN ENERGY DEVICE TO CUT THE STAPLED AREA AND THEN SUTURED IT WITH SUTURE. THE SURGEON TRIED TO OPEN THE CARTRIDGE USING A MANUAL TOOL BUT FAILED.
20-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING THE COLECTOMY, WHEN APPLIED ON THE COLON AND FOR SIDE-TO-SIDE ANASTOMOSIS, THE TISSUE WAS CAUGHT ON THE TIP OF THE CARTRIDGE AND WAS IN A STATE WHERE THE JAWS COULD NOT BE REMOVED. ADDITIONALLY, THE RETRACTION OF KNIFE BLADE WAS DIFFICULT OR NOT POSSIBLE FROM THE TISSUE. SO IT WAS RESECTED AND REMOVED. THE RESECTED PART WAS MANUALLY SUTURED. THE SURGICAL TIME WAS EXTENDED FOR ABOUT 20 MINUTES.
20-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING THE COLECTOMY, WHEN APPLIED ON THE COLON AND FOR SIDE-TO-SIDE ANASTOMOSIS, THE TISSUE WAS CAUGHT ON THE TIP OF THE CARTRIDGE AND WAS IN A STATE WHERE THE JAWS COULD NOT BE REMOVED. ADDITIONALLY, THE RETRACTION OF KNIFE BLADE WAS DIFFICULT OR NOT POSSIBLE FROM THE TISSUE. SO IT WAS RESECTED AND REMOVED. THE RESECTED PART WAS MANUALLY SUTURED. THE SURGICAL TIME WAS EXTENDED FOR ABOUT 20 MINUTES.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Human-Device Interface Problem | A22 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
25-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Insufficient Information | E2401 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RECTAL RESECTION, AFTER THE PHYSICIAN APPROACHED THE RECTUM, AND CLOSED THE CARTRIDGE, HOWEVER THE GREEN LIGHT DID NOT LIGHT UP AND THEREFORE THE DEVICE WAS NOT FIRED. THE STAPLER WAS REMOVED FROM THE PATIENT AND WAS REPLACED BY ANOTHER RELOAD. UPON CHECKING THE DEVICE, IT WAS FOUND THAT RELOAD WAS BENT TO THE RIGHT AND THE CONNECTION WITH THE ADAPTER WAS LOOSE. THE REPLACEMENT RELOAD WAS THEN USED AND SUCCESSFULLY FIRE, HOWEVER THE KNIFE OF THE RELOAD DID NOT RETURN AFTER FIRING. ALL POSSIBLE MEASURES (FORCED RESTART, RECONNECTING THE HANDLE AND ADAPTER, MANUAL ADAPTER TOOL) WERE ATTEMPTED, BUT THE KNIFE COULD NOT BE PULLED BACK. THE MOUTH AND ANAL BOWELS OF THE RELOAD THAT WAS LOCKED ON THE TISSUE WAS SEPARATED USING ANOTHER COMPANY'S STAPLER. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE TISSUE.
11-12-2024 | Malfunction | ethicon endo-surgery | 15-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A NEPHRECTOMY THE STAPLER DIDN'T FIRE, THEY OPENED A NEW ONE AND IT WORKED. THE SURGERY WAS DELAYED 3 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
17-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING OPEN GASTRECTOMY WHEN THE FIRING BUTTON WAS PRESSED AGAIN TO CHECK AFTER THE FIRING WAS COMPLETED, THE ARTICULATION WAS STARTED WHILE MAKING AN ABNORMAL SOUND AND IT WAS BENT TO 45. THE PARTS DROPPED IN THE PATIENT'S BODY. ALL PARTS WERE REMOVED AND WAS COLLECTED WITH FORCEPS. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. SOME STAPLE PUSHERS WERE PUSHED BACK TO THE CARTRIDGE. THREE STAPLES REMAINED ON THE CARTRIDGE AND ANVIL SURFACE. THE TWO STAPLES ON THE RELOAD WERE PROPERLY FORMED AND THE ONE WAS OVER CRIMPED.
17-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING OPEN GASTRECTOMY WHEN THE FIRING BUTTON WAS PRESSED AGAIN TO CHECK AFTER THE FIRING WAS COMPLETED, THE ARTICULATION WAS STARTED WHILE MAKING AN ABNORMAL SOUND AND IT WAS BENT TO 45. THE PARTS DROPPED IN THE PATIENT'S BODY. ALL PARTS WERE REMOVED AND WAS COLLECTED WITH FORCEPS. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. SOME STAPLE PUSHERS WERE PUSHED BACK TO THE CARTRIDGE. THREE STAPLES REMAINED ON THE CARTRIDGE AND ANVIL SURFACE. THE TWO STAPLES ON THE RELOAD WERE PROPERLY FORMED AND THE ONE WAS OVER CRIMPED.
07-01-2025 | Malfunction | ethicon endo-surgery | 15-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, IT WAS OBSERVED THAT DURING THE FIRING OF THE FIRST AND THIRD BLUE CARTRIDGE, A COMPLETE STAPLE LINE OF THREE ROWS WAS NOT FORMED. THE INNER ROW WAS INCOMPLETE AND THE STAPLES DEFORMED AND MISALIGNED. THE PROCEDURE WAS COMPLETED WITH A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2025 | Malfunction | ethicon endo-surgery | 15-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS PROCEDURE, IT WAS OBSERVED THAT DURING THE FIRING OF THE FIRST AND THIRD BLUE CARTRIDGE, A COMPLETE STAPLE LINE OF THREE ROWS WAS NOT FORMED. THE INNER ROW WAS INCOMPLETE AND THE STAPLES DEFORMED AND MISALIGNED. THE PROCEDURE WAS COMPLETED WITH A DELAY OF 10 MINUTES. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, BEFORE A RESPIRATORY LOBECTOMY SURGERY, THE JAWS DID NOT CLOSE COMPLETELY. ANOTHER RELOAD WAS USED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
29-11-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE DEVICE, THERE WERE ISSUES WITH THE HANDLE AND RELOAD. SPECIFICALLY, THERE WERE SET-UP AND PRIOR TO FIRING ISSUES, THE JAWS WAS DIFFICULT TO OPEN, IT WAS NOT OPENING SMOOTHLY. THE SURGERY WAS COMPLETED USING A NEW HANDLE OF THE SAME TYPE WITH THE SAME RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S EVALUATION FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE.
29-11-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE DEVICE, THERE WERE ISSUES WITH THE HANDLE AND RELOAD. SPECIFICALLY, THERE WERE SET-UP AND PRIOR TO FIRING ISSUES, THE JAWS WAS DIFFICULT TO OPEN, IT WAS NOT OPENING SMOOTHLY. THE SURGERY WAS COMPLETED USING A NEW HANDLE OF THE SAME TYPE WITH THE SAME RELOAD. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S EVALUATION FOUND THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE.
09-12-2024 | Malfunction | us surgical puerto rico | 15-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT FIRE AT ALL. THE HANDLE WAS REPLACED WITH ANOTHER ONE TO RESOLVE THE ISSUE.
07-10-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE ANASTOMOSIS DUODENAL ILEAL (SADI) PROCEDURE, WHILE THE DUODENO-ILEAL ANASTOMOSIS WAS BEING PERFORMED, AFTER STAPLING THE ENTIRE STOMACH AND DURING THE LAST SHOT, THE STAPLER STOPPED, AND A MESSAGE APPEARED ON THE SCREEN: "PRESS THE 2 GREEN BUTTONS UNTIL THE SYSTEM RESTARTS." THE SURGEON SIMULTANEOUSLY PRESSED THE SAFETY BUTTONS FOR THE SHOT, AND THE STAPLER RESTARTED. THE BLADE THEN RETRACTED, AND THE LOAD OPENED. IT WAS NOTED THAT THE STAPLING OF THE PIECE WAS COMPLETE ALL THE WAY TO THE END. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT AN ANALYSIS OF THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
07-10-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SINGLE ANASTOMOSIS DUODENAL ILEAL (SADI) PROCEDURE, WHILE THE DUODENO-ILEAL ANASTOMOSIS WAS BEING PERFORMED, AFTER STAPLING THE ENTIRE STOMACH AND DURING THE LAST SHOT, THE STAPLER STOPPED, AND A MESSAGE APPEARED ON THE SCREEN: "PRESS THE 2 GREEN BUTTONS UNTIL THE SYSTEM RESTARTS." THE SURGEON SIMULTANEOUSLY PRESSED THE SAFETY BUTTONS FOR THE SHOT, AND THE STAPLER RESTARTED. THE BLADE THEN RETRACTED, AND THE LOAD OPENED. IT WAS NOTED THAT THE STAPLING OF THE PIECE WAS COMPLETE ALL THE WAY TO THE END. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND THAT AN ANALYSIS OF THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
03-01-2025 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY, THE POWERED STAPLER SHOWED IT FIRED, BUT THERE WERE NO STAPLES OR BLADE DEPLOYED. THE INSIDE OF THE NECK (ARTICULATION JOINT) OF THE RELOAD WAS COMPLETELY UNRAVELED OUTSIDE OF THE RELOAD AND INTO THE PATIENT CAVITY. THE GUTS REMAINED ATTACHED BUT WERE VISIBLE ON THE SCREEN. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND TAKE IT OFF FROM THE TISSUE, BUT THE DEVICE COULD NOT BE ARTICULATED BACK TO NEUTRAL AND THEREFORE COULD NOT GET THE RELOAD OUT FROM THE ABDOMINAL CAVITY WITHOUT TAKING OUT THE TROCAR. THE RELOAD WAS STUCK INTO THE ADAPTER. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY, THE POWERED STAPLER SHOWED IT FIRED, BUT THERE WERE NO STAPLES OR BLADE DEPLOYED. THE INSIDE OF THE NECK (ARTICULATION JOINT) OF THE RELOAD WAS COMPLETELY UNRAVELED OUTSIDE OF THE RELOAD AND INTO THE PATIENT CAVITY. THE GUTS REMAINED ATTACHED BUT WERE VISIBLE ON THE SCREEN. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND TAKE IT OFF FROM THE TISSUE, BUT THE DEVICE COULD NOT BE ARTICULATED BACK TO NEUTRAL AND THEREFORE COULD NOT GET THE RELOAD OUT FROM THE ABDOMINAL CAVITY WITHOUT TAKING OUT THE TROCAR. THE RELOAD WAS STUCK INTO THE ADAPTER. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY, THE POWERED STAPLER SHOWED IT FIRED, BUT THERE WERE NO STAPLES OR BLADE DEPLOYED. THE INSIDE OF THE NECK (ARTICULATION JOINT) OF THE RELOAD WAS COMPLETELY UNRAVELED OUTSIDE OF THE RELOAD AND INTO THE PATIENT CAVITY. THE GUTS REMAINED ATTACHED BUT WERE VISIBLE ON THE SCREEN. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND TAKE IT OFF FROM THE TISSUE, BUT THE DEVICE COULD NOT BE ARTICULATED BACK TO NEUTRAL AND THEREFORE COULD NOT GET THE RELOAD OUT FROM THE ABDOMINAL CAVITY WITHOUT TAKING OUT THE TROCAR. THE RELOAD WAS STUCK INTO THE ADAPTER. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY, THE POWERED STAPLER SHOWED IT FIRED, BUT THERE WERE NO STAPLES OR BLADE DEPLOYED. THE INSIDE OF THE NECK (ARTICULATION JOINT) OF THE RELOAD WAS COMPLETELY UNRAVELED OUTSIDE OF THE RELOAD AND INTO THE PATIENT CAVITY. THE GUTS REMAINED ATTACHED BUT WERE VISIBLE ON THE SCREEN. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND TAKE IT OFF FROM THE TISSUE, BUT THE DEVICE COULD NOT BE ARTICULATED BACK TO NEUTRAL AND THEREFORE COULD NOT GET THE RELOAD OUT FROM THE ABDOMINAL CAVITY WITHOUT TAKING OUT THE TROCAR. THE RELOAD WAS STUCK INTO THE ADAPTER. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY, THE POWERED STAPLER SHOWED IT FIRED, BUT THERE WAS NO STAPLES OR BLADE DEPLOYED. THE INSIDE OF THE NECK (ARTICULATION JOINT) OF THE RELOAD WAS COMPLETELY UNRAVELED OUTSIDE OF THE RELOAD AND INTO THE PATIENT CAVITY. THE GUTS REMAINED ATTACHED BUT WERE VISIBLE ON THE SCREEN. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND TAKE IT OFF FROM THE TISSUE, BUT THE DEVICE COULD NOT BE ARTICULATED BACK TO NEUTRAL AND THEREFORE COULD NOT GET THE RELOAD OUT FROM THE ABDOMINAL CAVITY WITHOUT TAKING OUT THE TROCAR. THE RELOAD WAS STUCK INTO THE ADAPTER. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY, THE POWERED STAPLER SHOWED IT FIRED, BUT THERE WAS NO STAPLES OR BLADE DEPLOYED. THE INSIDE OF THE NECK (ARTICULATION JOINT) OF THE RELOAD WAS COMPLETELY UNRAVELED OUTSIDE OF THE RELOAD AND INTO THE PATIENT CAVITY. THE GUTS REMAINED ATTACHED BUT WERE VISIBLE ON THE SCREEN. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND TAKE IT OFF FROM THE TISSUE, BUT THE DEVICE COULD NOT BE ARTICULATED BACK TO NEUTRAL AND THEREFORE COULD NOT GET THE RELOAD OUT FROM THE ABDOMINAL CAVITY WITHOUT TAKING OUT THE TROCAR. THE RELOAD WAS STUCK INTO THE ADAPTER. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | covidien lp llc north haven | 15-01-2025 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY, THE POWERED STAPLER SHOWED IT FIRED, BUT THERE WAS NO STAPLES OR BLADE DEPLOYED. THE INSIDE OF THE NECK (ARTICULATION JOINT) OF THE RELOAD WAS COMPLETELY UNRAVELED OUTSIDE OF THE RELOAD AND INTO THE PATIENT CAVITY. THE GUTS REMAINED ATTACHED BUT WERE VISIBLE ON THE SCREEN. THE SURGEON WAS ABLE TO OPEN THE RELOAD AND TAKE IT OFF FROM THE TISSUE, BUT THE DEVICE COULD NOT BE ARTICULATED BACK TO NEUTRAL AND THEREFORE COULD NOT GET THE RELOAD OUT FROM THE ABDOMINAL CAVITY WITHOUT TAKING OUT THE TROCAR. THE RELOAD WAS STUCK INTO THE ADAPTER. A NEW ADAPTER AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-11-2024 | Malfunction | covidien | 15-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS THICK TISSUE. THE CLAMPING MECHANISMS WERE DEFORMED. STAPLE PUSHERS WERE PARTIALLY FLUSH WITH THE CARTRIDGE.
16-11-2024 | Malfunction | covidien | 15-01-2025 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS THICK TISSUE. THE CLAMPING MECHANISMS WERE DEFORMED. STAPLE PUSHERS WERE PARTIALLY FLUSH WITH THE CARTRIDGE.
23-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, BOTH RELOAD ONLY FIRED FOR ABOUT 15MM THEN IT GOT STUCKED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES DUE TO THE ISSUE. THERE WAS NO PATIENT INJURY.
23-12-2024 | Malfunction | covidien lp llc north haven | 15-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, BOTH RELOAD ONLY FIRED FOR ABOUT 15MM THEN IT GOT STUCKED. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED BY 30 MINUTES DUE TO THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2025 | Malfunction | ethicon endo-surgery | 16-01-2025 | ECHELON FLEX |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: 60MM STAPLER FIRED ALL STAPLES, BUT KNIFE BLADE DID NOT RETRACT. SURGEON DID NOT TAKE OUT BATTERY OR USE REVERSE BUTTON. SURGEON WENT STRAIGHT TO MANUAL OVERRIDE. SURGEON SAYS STAPLE LINE LOOKED GOOD.
09-05-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, WHEN THE DEVICE WAS PLACED INTO THE CHARGER, THE LED INITIALLY LIT UP GREEN, BUT AFTER A FEW MINUTES, THE LIGHT TURNED RED WITHOUT CHARGING. THE HANDLE DID NOT INDICATE ANYTHING ON THE SCREEN. IT WAS COMPLETELY TURNED OFF. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. ONE OF THE BATTERY CELLS WAS FOUND TO BE VENTED.
09-05-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, WHEN THE DEVICE WAS PLACED INTO THE CHARGER, THE LED INITIALLY LIT UP GREEN, BUT AFTER A FEW MINUTES, THE LIGHT TURNED RED WITHOUT CHARGING. THE HANDLE DID NOT INDICATE ANYTHING ON THE SCREEN. IT WAS COMPLETELY TURNED OFF. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. ONE OF THE BATTERY CELLS WAS FOUND TO BE VENTED.
20-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE COLECTOMY, WHEN APPLIED ON THE COLON, AND FOR SIDE-TO-SIDE ANASTOMOSIS, THE TISSUE WAS CAUGHT ON THE TIP OF THE CARTRIDGE AND WAS IN A STATE, WHERE THE JAWS COULD NOT BE REMOVED, ADDITIONALLY, THE RETRACTION OF KNIFE BLADE WAS DIFFICULT OR NOT POSSIBLE FROM THE TISSUE. SO IT WAS RESECTED AND REMOVED. THE RESECTED PART WAS MANUALLY SUTURED. THE SURGICAL TIME WAS EXTENDED FOR ABOUT 20 MINUTES.
20-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE COLECTOMY, WHEN APPLIED ON THE COLON, AND FOR SIDE-TO-SIDE ANASTOMOSIS, THE TISSUE WAS CAUGHT ON THE TIP OF THE CARTRIDGE AND WAS IN A STATE, WHERE THE JAWS COULD NOT BE REMOVED, ADDITIONALLY, THE RETRACTION OF KNIFE BLADE WAS DIFFICULT OR NOT POSSIBLE FROM THE TISSUE. SO IT WAS RESECTED AND REMOVED. THE RESECTED PART WAS MANUALLY SUTURED. THE SURGICAL TIME WAS EXTENDED FOR ABOUT 20 MINUTES.
16-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SURGERY, THE STAPLES BURST OUT, CAUSING THE CUTTING TO FAIL AND THE RUPTURE OF THE PATIENT'S GASTRIC WALL IN THE SURGICAL AREA. IT OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, MANUAL SUTURING AND SEROMUSCULAR LAYER EMBEDDING REINFORCEMENT WERE DONE.
17-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SURGERY, THE STAPLES BURST OUT, CAUSING THE CUTTING TO FAIL AND THE RUPTURE OF THE PATIENT'S GASTRIC WALL IN THE SURGICAL AREA. IT OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, MANUAL SUTURING AND SEROMUSCULAR LAYER EMBEDDING REINFORCEMENT WERE DONE.
01-01-2025 | Injury | ethicon endo-surgery | 16-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Self-Activation or Keying | A150104 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE POWERED CIRCULAR STAPLER MALFUNCTIONED DURING THE PROCEDURE. ACCORDING TO THE CUSTOMER, THE STAPLER SPONTANEOUSLY FIRED BEFORE THEY WERE ABLE TO CLOSE THE ANVIL. STAPLES PIERCED THE TISSUE AT THE RECTAL STUMP, BUT THEY DID NOT CLOSE. THE KNIFE WAS ALSO NOT INITIATED.
01-01-2025 | Injury | ethicon endo-surgery | 16-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE POWERED CIRCULAR STAPLER MALFUNCTIONED DURING THE PROCEDURE. ACCORDING TO THE CUSTOMER, THE STAPLER SPONTANEOUSLY FIRED BEFORE THEY WERE ABLE TO CLOSE THE ANVIL. STAPLES PIERCED THE TISSUE AT THE RECTAL STUMP, BUT THEY DID NOT CLOSE. THE KNIFE WAS ALSO NOT INITIATED.
01-01-2025 | Injury | ethicon endo-surgery | 16-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Cut | A050702 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE POWERED CIRCULAR STAPLER MALFUNCTIONED DURING THE PROCEDURE. ACCORDING TO THE CUSTOMER, THE STAPLER SPONTANEOUSLY FIRED BEFORE THEY WERE ABLE TO CLOSE THE ANVIL. STAPLES PIERCED THE TISSUE AT THE RECTAL STUMP, BUT THEY DID NOT CLOSE. THE KNIFE WAS ALSO NOT INITIATED.
01-01-2025 | Injury | ethicon endo-surgery | 16-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A LOW ANTERIOR RESECTION THE POWERED CIRCULAR STAPLER MALFUNCTIONED DURING THE PROCEDURE. ACCORDING TO THE CUSTOMER, THE STAPLER SPONTANEOUSLY FIRED BEFORE THEY WERE ABLE TO CLOSE THE ANVIL. STAPLES PIERCED THE TISSUE AT THE RECTAL STUMP, BUT THEY DID NOT CLOSE. THE KNIFE WAS ALSO NOT INITIATED.
18-12-2024 | Injury | davis & geck caribe | 16-01-2025 | UNKNOWN ENDO STITCH SULU | Component Missing | A020602 | Foreign Body In Patient | Event Description: ACCORDING TO THE REPORTER, THE DEVICE JAMMED AND WHEN THE SURGEON REMOVED THE DEVICE FROM THE PATIENT, THE STREAM PART HAD NO NEEDLE, HALF OF THE NEEDLE WAS IN THE PRODUCT. X-RAYS WERE PERFORMED ON THE PATIENT TO SEE IF THE NEEDLE WAS INSIDE THE PATIENT BUT COULD NOT FIND IT.
13-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, OPEN PYLORIC GASTRECTOMY, ON THE DUODENUM, THE CARTRIDGE COULD NOT BE OPENED AFTER THE FIRING. TO RESOLVE THE ISSUE, THE CARTRIDGE JAW WAS OPENED USING THE MANUAL ADAPTER TOOL AND AFTERWARDS, USED ANOTHER DEVICE FROM A DIFFERENT MANUFACTURER. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, OPEN PYLORIC GASTRECTOMY, ON THE DUODENUM, THE CARTRIDGE COULD NOT BE OPENED AFTER THE FIRING. TO RESOLVE THE ISSUE, THE CARTRIDGE JAW WAS OPENED USING THE MANUAL ADAPTER TOOL AND AFTERWARDS, USED ANOTHER DEVICE FROM A DIFFERENT MANUFACTURER. THERE WAS NO PATIENT INJURY.
01-12-2024 | Malfunction | us surgical puerto rico | 16-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY TO THE LAPAROSCOPIC APPENDECTOMY, THE RELOAD OR THE STAPLER WAS DEFECTIVE RESULTED TO THE RELOAD GETTING STUCK ON THE STAPLER. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT HARM.
01-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY TO THE LAPAROSCOPIC APPENDECTOMY, THE RELOAD OR THE STAPLER WAS DEFECTIVE RESULTED TO THE RELOAD GETTING STUCK ON THE STAPLER. A NEW DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT HARM.
24-12-2024 | Malfunction | covidien | 16-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, WHEN RESECTING THE ISSUE, THE JAWS COULD NOT BE CLOSED AND IT WAS DIFFICULT TO INSERT IT INTO THE TROCAR. TO RESOLVE THE ISSUE, THE RELOAD WAS REPLACED WITH THE SAME HANDLE AND ADAPTER. THERE WAS NO PATIENT INJURY.
24-12-2024 | Malfunction | us surgical puerto rico | 16-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE STAPLING THE APPENDIX, THE INSTRUMENT DID NOT FIRE AND THE PROCEDURE WAS EXTENDED. THE FIRST ASSIST INDICATED THAT THE APPENDIX RIPPED OFF DUE TO MULTIPLE ATTEMPTS TO FIRE THE STAPLER. TO COMPLETE THE PROCEDURE, ANOTHER STAPLER WAS USED FROM A DIFFERENT MANUFACTURER.
24-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, WHEN RESECTING THE ISSUE, THE JAWS COULD NOT BE CLOSED AND IT WAS DIFFICULT TO INSERT IT INTO THE TROCAR. TO RESOLVE THE ISSUE, THE RELOAD WAS REPLACED WITH THE SAME HANDLE AND ADAPTER. THERE WAS NO PATIENT INJURY.
24-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRAOPERATIVELY, WHEN RESECTING THE ISSUE, THE JAW COULD NOT BE CLOSED AND IT WAS DIFFICULT TO INSERT IT INTO THE TROCAR. TO RESOLVE THE ISSUE, THE RELOAD WAS REPLACED WITH THE SAME HANDLE AND ADAPTER. THERE WAS NO PATIENT INJURY.
13-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | TRI-STAPLE 2.0 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, OPEN PYLORIC GASTRECTOMY, ON THE DUODENUM, THE CARTRIDGE COULD NOT BE OPENED AFTER THE FIRING. TO RESOLVE THE ISSUE, THE CARTRIDGE JAW WAS OPENED USING THE MANUAL ADAPTER TOOL AND AFTERWARDS, USED ANOTHER DEVICE FROM A DIFFERENT MANUFACTURER. THERE WAS NO PATIENT INJURY.
01-01-2025 | Injury | ethicon endo-surgery | 16-01-2025 | PROXIMATE*PPH PROCEDURE SET | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY, THE PATIENT'S BLED AFTER SURGERY. THE FIRST PATIENT BLED AFTER TWELVE HOURS. THE RETURNED FOR THE STAPLE LINES TO BE REINFORCED.
06-01-2025 | Injury | ethicon endo-surgery | 16-01-2025 | PROXIMATE*PPH PROCEDURE SET | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A HEMORRHOIDECTOMY THE PATIENTS BLED AFTER SURGERY. THE PATIENT BLED AFTER TWENTY FOUR HOURS. PATIENT RETURNED FOR THE STAPLE LINES TO BE REINFORCED.
01-01-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | GST GREEN RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT GLOBAL BRAND PROTECTION LATAM TEAM HAS PERFORMED A TEST PURCHASE IN MEXICO, IN TWO NON-AUTHORIZED AND SUSPICIOUS SELLERS.
01-01-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT GLOBAL BRAND PROTECTION LATAM TEAM HAS PERFORMED A TEST PURCHASE IN MEXICO, IN TWO NON-AUTHORIZED AND SUSPICIOUS SELLERS.
01-01-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT GLOBAL BRAND PROTECTION LATAM TEAM HAS PERFORMED A TEST PURCHASE IN MEXICO, IN TWO NON-AUTHORIZED AND SUSPICIOUS SELLERS.
01-01-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT GLOBAL BRAND PROTECTION LATAM TEAM HAS PERFORMED A TEST PURCHASE IN MEXICO, IN TWO NON-AUTHORIZED AND SUSPICIOUS SELLERS.
01-01-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT GLOBAL BRAND PROTECTION LATAM TEAM HAS PERFORMED A TEST PURCHASE IN MEXICO, IN TWO NON-AUTHORIZED AND SUSPICIOUS SELLERS.
16-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SURGERY, THE STAPLES BURST OUT, CAUSING THE CUTTING TO FAIL AND THE RUPTURE OF THE PATIENT'S GASTRIC WALL IN THE SURGICAL AREA. IT OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, MANUAL SUTURING AND SEROMUSCULAR LAYER EMBEDDING REINFORCEMENT WERE DONE.
15-12-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY SURGERY, AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END, BUT COULD NOT RETURN KNIFE AUTOMATICALLY. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2025 | Malfunction | us surgical puerto rico | 16-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACIC LOBECTOMY AND VASCULAR DETACHMENT PROCEDURE, THE STAPLER DID NOT FIRE. A NEW RELOAD WAS USED WITH THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
20-10-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT GLOBAL BRAND PROTECTION LATAM TEAM HAS PERFORMED A TEST PURCHASE IN MEXICO, IN TWO NON-AUTHORIZED AND SUSPICIOUS SELLERS.
07-01-2025 | Malfunction | covidien lp llc north haven | 16-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE STAPLE LINE WAS INCOMPLETE.
07-01-2025 | Malfunction | covidien lp llc north haven | 16-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE, THE STAPLE LINE WAS INCOMPLETE.
25-12-2024 | Malfunction | us surgical puerto rico | 16-01-2025 | ENDO GIA | Break | A0401 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING SMALL INTESTINE RESECTION, INSERTION PORT CLOSURE, IT BROKE BEFORE CUTTING. IT WAS REPORTED THAT THERE WAS A PART WHERE ONLY ONE ROW WAS STAPLED. AN ADDITIONAL RESECTION WAS PERFORMED, AND ADDITIONAL SUTURES WERE PERFORMED WITH A STAPLER.
25-12-2024 | Malfunction | us surgical puerto rico | 16-01-2025 | ENDO GIA | Mechanics Altered | A0507 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING SMALL INTESTINE RESECTION, INSERTION PORT CLOSURE, IT BROKE BEFORE CUTTING. IT WAS REPORTED THAT THERE WAS A PART WHERE ONLY ONE ROW WAS STAPLED. AN ADDITIONAL RESECTION WAS PERFORMED, AND ADDITIONAL SUTURES WERE PERFORMED WITH A STAPLER.
20-12-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
24-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, WHILE STAPLING THE APPENDIX, THE INSTRUMENT DID NOT FIRE AND THE PROCEDURE WAS EXTENDED. THE FIRST ASSIST INDICATED THAT THE APPENDIX RIPPED OFF DUE TO MULTIPLE ATTEMPTS TO FIRE THE STAPLER. TO COMPLETE THE PROCEDURE, ANOTHER STAPLER WAS USED FROM A DIFFERENT MANUFACTURER.
19-11-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A WEDGE RESECTION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-11-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-11-2024 | Malfunction | ethicon endo-surgery | 16-01-2025 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE SIGNIA POWER HANDLE, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM(KNIFE BLADE) DID NOT RETRACT, AND THE CARTRIDGE JAW DID NOT OPEN. THE DEVICE LOCKED ON TISSUE AND COULD NOT BE REMOVED. THE DEVICE WAS REMOVED FROM THE CAVITY WITH THE JAWS STILL CLOSED ON THE TISSUE. THE DEVICE WAS REPLACED WITH A STAPLER FROM ANOTHER MANUFACTURER TO RESOLVE THE ISSUE.
26-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE SIGNIA POWER HANDLE, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM(KNIFE BLADE) DID NOT RETRACT, AND THE CARTRIDGE JAW DID NOT OPEN. THE DEVICE LOCKED ON TISSUE AND COULD NOT BE REMOVED. THE DEVICE WAS REMOVED FROM THE CAVITY WITH THE JAWS STILL CLOSED ON THE TISSUE. THE DEVICE WAS REPLACED WITH A STAPLER FROM ANOTHER MANUFACTURER TO RESOLVE THE ISSUE.
26-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE SIGNIA POWER HANDLE, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM(KNIFE BLADE) DID NOT RETRACT, AND THE CARTRIDGE JAW DID NOT OPEN. THE DEVICE LOCKED ON TISSUE AND COULD NOT BE REMOVED. THE DEVICE WAS REMOVED FROM THE CAVITY WITH THE JAWS STILL CLOSED ON THE TISSUE. THE DEVICE WAS REPLACED WITH A STAPLER FROM ANOTHER MANUFACTURER TO RESOLVE THE ISSUE.
26-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE SIGNIA POWER HANDLE, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM(KNIFE BLADE) DID NOT RETRACT, AND THE CARTRIDGE JAW DID NOT OPEN. THE DEVICE LOCKED ON TISSUE AND COULD NOT BE REMOVED. THE DEVICE WAS REMOVED FROM THE CAVITY WITH THE JAWS STILL CLOSED ON THE TISSUE. THE DEVICE WAS REPLACED WITH A STAPLER FROM ANOTHER MANUFACTURER TO RESOLVE THE ISSUE.
26-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE SIGNIA POWER HANDLE, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM (KNIFE BLADE) DID NOT RETRACT, AND THE CARTRIDGE JAW DID NOT OPEN. THE DEVICE LOCKED ON TISSUE AND COULD NOT BE REMOVED. THE DEVICE WAS REMOVED FROM THE CAVITY WITH THE JAWS STILL CLOSED ON THE TISSUE. THE DEVICE WAS REPLACED WITH A STAPLER FROM ANOTHER MANUFACTURER TO RESOLVE THE ISSUE.
26-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE SIGNIA POWER HANDLE, AFTER FIRING THE STAPLER INTO THE TISSUE, THE BEAM (KNIFE BLADE) DID NOT RETRACT, AND THE CARTRIDGE JAW DID NOT OPEN. THE DEVICE LOCKED ON TISSUE AND COULD NOT BE REMOVED. THE DEVICE WAS REMOVED FROM THE CAVITY WITH THE JAWS STILL CLOSED ON THE TISSUE. THE DEVICE WAS REPLACED WITH A STAPLER FROM ANOTHER MANUFACTURER TO RESOLVE THE ISSUE.
22-12-2024 | Malfunction | johnson & johnson medical (suzhou | 16-01-2025 | CARTRIDGE FINAL ASSY 60MM BLUE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC GASTRIC SURGERY, THE STAPLES BURST OUT CAUSING THE CUTTING TO FAIL AND THE RUPTURE OF THE PATIENT'S GASTRIC WALL IN THE SURGICAL AREA. IT OCCURRED ON TWO RELOADS. TO RESOLVE THE ISSUE, MANUAL SUTURING AND SEROMUSCULAR LAYER EMBEDDING REINFORCEMENT WERE DONE.
06-01-2025 | Malfunction | covidien lp llc north haven | 16-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A DUODENAL SWITCH LAPAROSCOPIC PROCEDURE, THE STAPLE LINE WAS INCOMPLETE DISTALLY AFTER THE FIRST FIRING OF THE SIGNIA STAPLING SYSTEM WITH THE EGIA60AMT RELOAD ON THE STOMACH. THE ISSUE REQUIRED RESECTION AND RE-STAPLING TO ADDRESS THE INCOMPLETE STAPLE LINE.
26-12-2024 | Malfunction | us surgical puerto rico | 16-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLON RESECTION, DURING OPENING AND CLOSING OUTSIDE THE BODY AT THE TIME OF INTESTINAL RESECTION , AFTER RE-CLOSING THE DEVICE, THE DEVICE DID NOT OPEN ANYMORE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLON RESECTION, DURING OPENING AND CLOSING OUTSIDE THE BODY AT THE TIME OF INTESTINAL RESECTION , AFTER RE-CLOSING THE DEVICE, THE DEVICE DID NOT OPEN ANYMORE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
26-12-2024 | Malfunction | covidien lp llc north haven | 16-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC COLON RESECTION, DURING OPENING AND CLOSING OUTSIDE THE BODY AT THE TIME OF INTESTINAL RESECTION, AFTER RE-CLOSING THE DEVICE, THE DEVICE DID NOT OPEN ANYMORE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-12-2024 | Malfunction | us surgical puerto rico | 16-01-2025 | ENDO GIA ULTRA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC DISTAL GASTRECTOMY, WHILE DETACHING THE GASTRIC TISSUE, THE DEVICE WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION. THE FIRING WAS NOT SMOOTH. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
03-12-2024 | Malfunction | us surgical puerto rico | 16-01-2025 | ENDO GIA ULTRA | Firing Problem | A0505 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC DISTAL GASTRECTOMY, WHILE DETACHING THE GASTRIC TISSUE, THE DEVICE WAS ABLE TO FIRE BUT THERE WAS POOR STAPLE FORMATION. THE FIRING WAS NOT SMOOTH. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK.
01-01-2025 | Injury | ethicon endo-surgery | 17-01-2025 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Patient Device Interaction Problem | A01 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT POST OP TO A LAP HEMICOLECTOMY TWO DAYS PATIENT WAS REOPERATED BECAUSE DE-ANASTOMOSIS. COLON WAS IN GOOD CONDITION, BUT THE ANASTOMOSIS HAD LOOSENED. DURING THE SURGERY HE DIDN'T FEEL ANY DIFFERENCE COMPARED TO PREVIOUS TIME HE USED IT. HE IS A FREQUENT USER OF ECHELON POWERED CIRCULAR STAPLER.
11-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 17-01-2025 | CAPSURE | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A LAPAROSCOPIC VENTRAL HERNIA REPAIR, CAPSURE WAS USED TO FIXATE THE MESH TO THE ABDOMINAL WALL. IT WAS REPORTED THAT THE DEVICE DEPLOYED TWO FASTENERS OVERLAPPING WHEN THE TRIGGER WAS PULLED. IT WAS REPORTED THAT THE DEPLOYED MULTIPLE FASTENERS WERE REMOVED FROM THE PATIENT'S BODY. THE PROCEDURE WAS COMPLETED AS IT WAS AND THERE WAS NO PATIENT INJURY.
15-10-2024 | Malfunction | covidien lp llc north haven | 17-01-2025 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE SERVICING OF THE DEVICE, THE HANDLE WOULD NOT CHARGE. THE CHARGER WAS ABLE TO CHARGE OTHER HANDLE WITHOUT ISSUE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED BATTERY CELLS WERE FOUND TO BE VENTED.
15-10-2024 | Malfunction | covidien lp llc north haven | 17-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE SERVICING OF THE DEVICE, THE HANDLE WOULD NOT CHARGE. THE CHARGER WAS ABLE TO CHARGE OTHER HANDLE WITHOUT ISSUE. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED BATTERY CELLS WERE FOUND TO BE VENTED.
27-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 17-01-2025 | OPTIFIX STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING ONE INGUINAL HERNIA REPAIR PROCEDURE, THE FASTENERS OF OPTIFIX FIXATION DEVICE WERE NOT FIXATING (DEVICE #1, #2 - 15 COUNT, DEVICE #3 - 30 COUNT). IT WAS REPORTED THAT SURGEON WAS AN EXPERIENCED USER OF THE DEVICE BUT HAD ISSUES THIS TIME DURING USE. AS REPORTED, THE AREA OF FIXATION WAS NOT ANY BONE OR LIGAMENT AND THE FASTENERS THAT WERE NOT FIXATING WERE REMOVED FROM THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY.
27-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 17-01-2025 | OPTIFIX STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING ONE INGUINAL HERNIA REPAIR PROCEDURE, THE FASTENERS OF OPTIFIX FIXATION DEVICE WERE NOT FIXATING (DEVICE #1, #2 - (B)(4) COUNT, DEVICE #3 - (B)(4) COUNT). IT WAS REPORTED THAT SURGEON WAS AN EXPERIENCED USER OF THE DEVICE BUT HAD ISSUES THIS TIME DURING USE. AS REPORTED, THE AREA OF FIXATION WAS NOT ANY BONE OR LIGAMENT AND THE FASTENERS THAT WERE NOT FIXATING WERE REMOVED FROM THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY.
27-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 17-01-2025 | OPTIFIX STRAIGHT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING ONE INGUINAL HERNIA REPAIR PROCEDURE, THE FASTENERS OF OPTIFIX FIXATION DEVICE WERE NOT FIXATING (DEVICE #1, #2 - 15 COUNT, DEVICE #3 - 30 COUNT). IT WAS REPORTED THAT SURGEON WAS AN EXPERIENCED USER OF THE DEVICE BUT HAD ISSUES THIS TIME DURING USE. AS REPORTED, THE AREA OF FIXATION WAS NOT ANY BONE OR LIGAMENT AND THE FASTENERS THAT WERE NOT FIXATING WERE REMOVED FROM THE PATIENT. THERE WAS NO REPORTED PATIENT INJURY.
06-01-2025 | Malfunction | ethicon endo-surgery | 17-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN RESPIRATORY PROCEDURE, WHEN THE DEVICE WAS USED FOR PULMONARY ARTERY DISSECTION, THE KNIFE OF STAPLE STOPPED ON THE WAY. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
04-10-2024 | Malfunction | covidien lp llc north haven | 17-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY (VATS) WEDGE RESECTION PROCEDURE, THERE WAS LO ADING DIFFICULTY WITH ONE OF THE PURPLE RELOADS. WHILE STAPLING THE LUNG, USING ANOTHER PURPLE RELOAD, IT FIRED PAST THE STAPLE LINE, WHICH CAUSED SLIGHT BLEEDING IN THE TISSUE. THIS RELOAD DID NOT FULLY OPEN EITHER. ADDITIONALLY, A REINFORCED BLACK RELOAD HAD LOADING DIFFICULTY AND COULD NOT BE REMOVED FROM THE ADAPTER. TO RESOLVE THE ISSUE, A MANUAL HANDLE AND RELOAD WERE USED. MEDTRONIC'S EVALUATION FOUND THAT THE RELOAD WAS PARTIALLY APPLIED.
04-10-2024 | Malfunction | covidien lp llc north haven | 17-01-2025 | TRI-STAPLE 2.0 | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY (VATS) WEDGE RESECTION PROCEDURE, THERE WAS LO ADING DIFFICULTY WITH ONE OF THE PURPLE RELOADS. WHILE STAPLING THE LUNG, USING ANOTHER PURPLE RELOAD, IT FIRED PAST THE STAPLE LINE, WHICH CAUSED SLIGHT BLEEDING IN THE TISSUE. THIS RELOAD DID NOT FULLY OPEN EITHER. ADDITIONALLY, A REINFORCED BLACK RELOAD HAD LOADING DIFFICULTY AND COULD NOT BE REMOVED FROM THE ADAPTER. TO RESOLVE THE ISSUE, A MANUAL HANDLE AND RELOAD WERE USED. MEDTRONIC'S EVALUATION FOUND THAT THE RELOAD WAS PARTIALLY APPLIED.
31-07-2024 | Malfunction | covidien lp llc north haven | 17-01-2025 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPY PROCEDURE, THE LOAD WAS DEFECTIVE. THE DEVICE WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE INTERLOCK WAS OVERRIDDEN AND THE RELOAD WAS APPLIED TO TEST MEDIA.
01-01-2025 | Malfunction | ethicon endo-surgery | 17-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED AND BROKE INSIDE A PATIENT. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 17-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE LOCKED AND BROKE INSIDE A PATIENT. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
03-12-2024 | Malfunction | ethicon | 17-01-2025 | ETHICON SECURESTRAP | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT INGUINAL HERNIORRAPHY ON (B)(6) 2024 AND ABSORBABLE STRAPS WERE USED. IT WAS OBSERVED THAT THE STAPLES DID NOT PENETRATE THE TISSUES. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.
29-12-2024 | Malfunction | us surgical puerto rico | 17-01-2025 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC RADICAL SURGERY OF RECTAL CANCER, WHEN REMOVING THE MESORECTUM. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE, THEREFORE THE DEVICE WAS NOT ABLE TO BE FIRED. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-11-2024 | Malfunction | us surgical puerto rico | 17-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN SAFETY BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE. TO COMPLETE THE PROCEDURE A NEW HANDLE WAS USED. THERE IS NO PATIENT INJURY.
15-11-2024 | Malfunction | us surgical puerto rico | 17-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN SAFETY BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE. TO COMPLETE THE PROCEDURE A NEW HANDLE WAS USED. THERE IS NO PATIENT INJURY.
15-11-2024 | Malfunction | us surgical puerto rico | 17-01-2025 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN SAFETY BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE. TO COMPLETE THE PROCEDURE A NEW HANDLE WAS USED. THERE IS NO PATIENT INJURY.
15-11-2024 | Malfunction | us surgical puerto rico | 17-01-2025 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC APPENDECTOMY, THE DEVICE DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN SAFETY BUTTON AND WAS ABLE TO SQUEEZE THE HANDLE. TO COMPLETE THE PROCEDURE A NEW HANDLE WAS USED. THERE IS NO PATIENT INJURY.
15-11-2024 | Malfunction | us surgical puerto rico | 17-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN SAFETY BUTTON AND AND WAS ABLE TO SQUEEZE THE HANDLE. TO COMPLETE THE PROCEDURE A NEW DEVICE WAS USED.
15-11-2024 | Malfunction | us surgical puerto rico | 17-01-2025 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN SAFETY BUTTON AND AND WAS ABLE TO SQUEEZE THE HANDLE. TO COMPLETE THE PROCEDURE A NEW DEVICE WAS USED.
16-09-2024 | Malfunction | us surgical puerto rico | 19-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, BOTH THE HANDLE AND RELOAD WERE DEFECTIVE. IT WAS NECESSARY TO CHANGE THE HANDLE AND RELOAD TO RESOLVE THE ISSUE AND COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS HANDLE THICK TISSUE. THE INSTRUMENT WAS LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
16-09-2024 | Malfunction | us surgical puerto rico | 19-01-2025 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PROCEDURE, BOTH THE HANDLE AND RELOAD WERE DEFECTIVE. IT WAS NECESSARY TO CHANGE THE HANDLE AND RELOAD TO RESOLVE THE ISSUE AND COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS HANDLE THICK TISSUE. THE INSTRUMENT WAS LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
22-11-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 20-01-2025 | CAPSURE | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING LAPAROSCOPIC INGUINAL HERNIA REPAIR, CAPSURE FIXATION DEVICE WAS USED AND WHEN IT WAS TRIGGERED, TWO OVERLAPPED FASTENERS WERE DEPLOYED. IT WAS REPORTED THAT FIRST SHOT WAS FIXED TO THE MESH, THE SECOND SHOT OVERLAPPED THE FIRST AND DID NOT REACH THE MESH. THE SECOND TRIGGER SHOT WAS A BLANK, AND THE THIRD SHOT AFTER THAT WAS FIRED NORMALLY. IT WAS REPORTED THAT ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE SINCE THIS DEVICE RAN OUT OF FASTENERS. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY, THE DEVICE FIRED AT FIRST AND THEN STOPPED FIRING. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY, THE DEVICE FIRED AT FIRST AND THEN STOPPED FIRING. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | ENDO GIA | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY, THE DEVICE FIRED AT FIRST AND THEN STOPPED FIRING. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A TOTAL GASTRECTOMY, THE DEVICE FIRED AT FIRST AND THEN STOPPED FIRING. A NEW MANUAL HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-01-2025 | Malfunction | covidien lp llc north haven | 20-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACIC LOBECTOMY, THE INSTRUMENT DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-01-2025 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THORACIC LOBECTOMY, THE INSTRUMENT DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE. THE HANDLE AND RELOAD WERE REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE SEPARATING THE TARGET TISSUE FROM ANOTHER ADJACENT TISSUE, THE HANDLE WAS STUCK DURING THE FIRING PROCESS AND WAS NOT USED WELL. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, WHILE SEPARATING THE TARGET TISSUE FROM ANOTHER ADJACENT TISSUE, THE HANDLE WAS STUCK DURING THE FIRING PROCESS AND WAS NOT USED WELL. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
02-01-2025 | Malfunction | covidien lp llc north haven | 20-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT UPPER LOBECTOMY PROCEDURE, THE RELOAD WAS USED TO DETACH THE INTERLOBAR FISSURE, AND THE STAPLES BURST OUT AND DID NOT FORM A GOOD B SHAPE, AND THERE WAS BLEEDING. AFTER PRESSING WITH GAUZE, THE LUNG PARENCHYMA WAS RE-DETACHED AFTER REPLACEMENT OF THE NEW HANDLE AND STAPLE AT THE ADJACENT POSITION TO STOP THE BLEEDING IMMEDIATELY AND FINALLY THE OPERATION WAS COMPLETED. THE INCISION WAS EXTENDED DUE TO THE PRODUCT PROBLEM.
02-01-2025 | Malfunction | covidien lp llc north haven | 20-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LEFT UPPER LOBECTOMY PROCEDURE, THE RELOAD WAS USED TO DETACH THE INTERLOBAR FISSURE, AND THE STAPLES BURST OUT AND DID NOT FORM A GOOD B SHAPE, AND THERE WAS BLEEDING. AFTER PRESSING WITH GAUZE, THE LUNG PARENCHYMA WAS RE-DETACHED AFTER REPLACEMENT OF THE NEW HANDLE AND STAPLE AT THE ADJACENT POSITION TO STOP THE BLEEDING IMMEDIATELY AND FINALLY THE OPERATION WAS COMPLETED. THE INCISION WAS EXTENDED DUE TO THE PRODUCT PROBLEM.
30-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A LOWER ANTERIOR RESECTION, THE DISTAL PART OF STAPLE LINE MISFIRED. TO RESOLVE THE ISSUE, SUTURING AS A STAPLE LINE REPLACEMENT WAS DONE.
03-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LOBECTOMY, WHEN DETACHING THE BLOOD VESSELS, THERE WAS DIFFICULTY UNLOADING THE RELOAD FROM THE HANDLE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
03-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPIC LOBECTOMY, WHEN DETACHING THE BLOOD VESSELS, THERE WAS DIFFICULTY UNLOADING THE RELOAD FROM THE HANDLE. ANOTHER HANDLE WAS USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
30-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTAL RESECTION, WHILE ON DISSOCIATING SEGMENTAL BLOOD VESSELS, THE SURGICAL HOOK HOOKED THE I-SHAPED BEAM BACK. IT WAS NOTED THAT THERE WAS UNSMOOTH FIRING. THE SURGICAL HOOK RETURNED THE I-SHAPED BEAM TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTAL RESECTION, WHILE ON DISSOCIATING SEGMENTAL BLOOD VESSELS, THE SURGICAL HOOK HOOKED THE I-SHAPED BEAM BACK. IT WAS NOTED THAT THERE WAS UNSMOOTH FIRING. THE SURGICAL HOOK RETURNED THE I-SHAPED BEAM TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | covidien lp llc north haven | 20-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL LUNG CANCER SURGERY, THE FIRST RELOAD HAD POOR STAPLE FORMATION AND STAPLE LINE INCOMPLETE ISSUE. THE SECOND RELOAD HAD AN UNKNOWN ISSUE. THE DEVICE WAS REPLACED TO COMPLETE THE PROCEDURE.
22-10-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THERE WAS AN UNDEFINED ADAPTER ERROR AND THE RELOAD WAS NOT RECOGNIZED. ANOTHER POWER HANDLE AND ADAPTER WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AN ANALYSIS OF THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
22-10-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | SIGNIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THERE WAS AN UNDEFINED ADAPTER ERROR AND THE RELOAD WAS NOT RECOGNIZED. ANOTHER POWER HANDLE AND ADAPTER WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AN ANALYSIS OF THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
22-10-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PREOPERATIVELY, THERE WAS AN UNDEFINED ADAPTER ERROR AND THE RELOAD WAS NOT RECOGNIZED. ANOTHER POWER HANDLE AND ADAPTER WERE USED TO COMPLETE THE CASE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT AN ANALYSIS OF THE SYSTEM DATA INDICATED THAT THERE WAS AN ERROR FOR THE SYSTEM QUEUE BEING FULL.
31-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | UNKNOWN GIA PRODUCT | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE BLADE OF THE CARTRIDGE POPPED UP THE MOMENT THE HANDLE WAS INSTALLED AND OPENED AFTER THE STAPLE WAS MOUNTED. THE CARTRIDGE WAS LOADED SUCCESSFULLY BUT THEN POPPED OFF WHEN THE HANDLE WAS ATTEMPTED TO BE PUT TOGETHER. ANOTHER CARTRIDGE WAS USED BUT THE SAME ISSUE OCCURRED. THERE WAS NO PATIENT INJURY.
31-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | UNKNOWN GIA PRODUCT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE BLADE OF THE CARTRIDGE POPPED UP THE MOMENT THE HANDLE WAS INSTALLED AND OPENED AFTER THE STAPLE WAS MOUNTED. THE CARTRIDGE WAS LOADED SUCCESSFULLY BUT THEN POPPED OFF WHEN THE HANDLE WAS ATTEMPTED TO BE PUT TOGETHER. ANOTHER CARTRIDGE WAS USED BUT THE SAME ISSUE OCCURRED. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | covidien lp llc north haven | 20-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A RADICAL LUNG CANCER SURGERY, WHEN DISSOCIATED THE LUNG PARENCHYMA, SOME STAPLES DID NOT HAVE A B-SHAPED STAPLE. ANOTHER HANDLE WAS USED AND MANUAL SUTURING WAS DONE TO RESOLVE THE ISSUE.
30-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING USE, A FIRE OCCURRED POTENTIALLY INVOLVING THE DEVICE AS THE SOURCE, CAUSING DAMAGE BOTH TO THE POWER HANDLE AND CHARGER. TO RESOLVE THE ISSUE, THE PRODUCTS WERE REPLACED. A HANDPIECE AND CHARGER FROM CONSIGNMENT WAS PROVIDED. THERE WAS NO PATIENT INVOLVEMENT.
30-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | SIGNIA | Overheating of Device | A1005 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING USE, A FIRE OCCURRED POTENTIALLY INVOLVING THE DEVICE AS THE SOURCE, CAUSING DAMAGE BOTH TO THE POWER HANDLE AND CHARGER. TO RESOLVE THE ISSUE, THE PRODUCTS WERE REPLACED. A HANDPIECE AND CHARGER FROM CONSIGNMENT WAS PROVIDED. THERE WAS NO PATIENT INVOLVEMENT.
30-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING USE, A FIRE OCCURRED POTENTIALLY INVOLVING THE DEVICE AS THE SOURCE, CAUSING DAMAGE BOTH TO THE POWER HANDLE AND CHARGER. TO RESOLVE THE ISSUE, THE PRODUCTS WERE REPLACED. A HANDPIECE AND CHARGER FROM CONSIGNMENT WAS PROVIDED. THERE WAS NO PATIENT INVOLVEMENT.
30-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | SIGNIA | Overheating of Device | A1005 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING USE, A FIRE OCCURRED POTENTIALLY INVOLVING THE DEVICE AS THE SOURCE, CAUSING DAMAGE BOTH TO THE POWER HANDLE AND CHARGER. TO RESOLVE THE ISSUE, THE PRODUCTS WERE REPLACED. A HANDPIECE AND CHARGER FROM CONSIGNMENT WAS PROVIDED. THERE WAS NO PATIENT INVOLVEMENT.
16-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC RADICAL RESECTION OF RIGHT UPPER LUNG CANCER, WHEN THE DOCTOR OPENED THE STAPLER AND INSTALLED THE FIRST RELOAD, THE CUTTING WAS COMPLETED SMOOTHLY. WHEN THE SECOND RELOAD WAS INSTALLED FOR THE SECOND TIME, THE CUTTING WAS ALSO COMPLETED SMOOTHLY. WHEN THE THIRD RELOAD WAS INSTALLED FOR THE THIRD TIME, THE STAPLER COULD NOT CUT. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS UNABLE TO FIRE. THE DOCTOR IMMEDIATELY REPLACED THE RELOAD WITH A NEW ONE, BUT THE STAPLER STILL COULD NOT CUT. THE DOCTOR ASKED THE NURSE TO GET A NEW STAPLER AND THE THIRD UNSUCCESSFUL RELOAD WAS REINSTALLED. AFTER THE RELOAD WAS PLACED ON THE NEW STAPLER, THE CUTTING OPERATION WAS SUCCESSFULLY COMPLETED. THE PATIENT'S OPERATION TIME WAS PROLONGED BY ABOUT 7 MINUTES AND INCREASED THE BLEEDING VOLUME BY ABOUT 5ML AND WAS RESOLVED WITH THE USE OF HEMOSTATIC PRODUCT. ADDITIONALLY THERE WAS UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT.
16-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE THORACOSCOPIC RADICAL RESECTION OF RIGHT UPPER LUNG CANCER, WHEN THE DOCTOR OPENED THE STAPLER AND INSTALLED THE FIRST RELOAD, THE CUTTING WAS COMPLETED SMOOTHLY. WHEN THE SECOND RELOAD WAS INSTALLED FOR THE SECOND TIME, THE CUTTING WAS ALSO COMPLETED SMOOTHLY. WHEN THE THIRD RELOAD WAS INSTALLED FOR THE THIRD TIME, THE STAPLER COULD NOT CUT. IT WAS NOTED THAT THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON AND SQUEEZE THE HANDLE BUT THE STAPLER WAS UNABLE TO FIRE. THE DOCTOR IMMEDIATELY REPLACED THE RELOAD WITH A NEW ONE, BUT THE STAPLER STILL COULD NOT CUT. THE DOCTOR ASKED THE NURSE TO GET A NEW STAPLER AND THE THIRD UNSUCCESSFUL RELOAD WAS REINSTALLED. AFTER THE RELOAD WAS PLACED ON THE NEW STAPLER, THE CUTTING OPERATION WAS SUCCESSFULLY COMPLETED. THE PATIENT'S OPERATION TIME WAS PROLONGED BY ABOUT 7 MINUTES AND INCREASED THE BLEEDING VOLUME BY ABOUT 5ML AND WAS RESOLVED WITH THE USE OF HEMOSTATIC PRODUCT. ADDITIONALLY THERE WAS UNANTICIPATED TISSUE LOSS AND TISSUE DAMAGE AS A RESULT.
11-12-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A SIGMOID COLECTOMY (TYPICALLY LAP FOR DISSECTION WITH EEA INSERTION TRANSANALLY), DURING ANASTOMOSIS, AFTER FIRING, A LEAK TEST WAS PERFORMED WITH NEGATIVE RESULTS. TWO DAYS LATER, THE PATIENT SHOWED SYMPTOMS; THE PATIENT BECAME TACHYCARDIC. ON THE THIRD DAY, THERE WAS FLUID IN THE PATIENT¿S STOMACH, KIDNEY FAILURE - TANKED THROUGH, A PNEUMONIA SYMPTOM, INCREASING RESPIRATORY DISTRESS, AND NO BOWEL MOVEMENT. A WEEK AFTER THE INITIAL PROCEDURE, THE COMPUTED TOMOGRAPHY SCAN (CT) SHOWED AIR AROUND THE ANASTOMOSIS. THE PATIENT WAS READMITTED TO THE HOSPITAL AND WAS RE-OPERATED ON DUE TO THE ISSUE. IT WAS ALSO FOUND THAT THE PATIENT HAD A BOWEL INJURY THAT WAS LEADING TO FURTHER COMPLICATIONS. UPON SECOND SURGERY, IT WAS DISCOVERED THAT THE ANASTOMOSIS WAS HEALING EXTREMELY WELL AND THERE WAS NOT A LEAK IN THE ANASTOMOSIS.
20-01-2025 | Malfunction | covidien lp llc north haven | 20-01-2025 | SIGNIA | Retraction Problem | A0510 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE USE OF AN ARTICULATING MEDIUM THICK RELOAD WITH A POWERED STAPLER AND STANDARD ADAPTER, THERE WAS DIFFICULTY RETRACTING THE KNIFE BLADE AFTER FIRING. THE RELOAD WAS REPLACED WITH ANOTHER RELOAD TO RESOLVE THE ISSUE.
20-01-2025 | Malfunction | covidien lp llc north haven | 20-01-2025 | SIGNIA | Retraction Problem | A0510 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE USE OF AN ARTICULATING MEDIUM THICK RELOAD WITH A POWERED STAPLER AND STANDARD ADAPTER, THERE WAS DIFFICULTY RETRACTING THE KNIFE BLADE AFTER FIRING. THE RELOAD WAS REPLACED WITH ANOTHER RELOAD TO RESOLVE THE ISSUE.
26-03-2025 | Malfunction | covidien lp llc north haven | 20-01-2025 | ENDO GIA | Retraction Problem | A0510 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE USE OF AN ARTICULATING MEDIUM THICK RELOAD WITH A POWERED STAPLER AND STANDARD ADAPTER, THERE WAS DIFFICULTY RETRACTING THE KNIFE BLADE AFTER FIRING. THE RELOAD WAS REPLACED WITH ANOTHER RELOAD TO RESOLVE THE ISSUE.
05-09-2024 | Malfunction | covidien lp llc north haven | 20-01-2025 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IT WAS FOUND THAT IT COULD NOT BE FIRED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
05-09-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IT WAS FOUND THAT IT COULD NOT BE FIRED. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
30-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF AN ANTERIOR RESECTION, WHILE IN THE COLON, THE DEVICE FIRED COMPLETELY, BUT THERE WAS POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE. THE STAPLE LINE WAS SUTURED TO RESOLVE THE ISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE.
30-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE INVOLVING THE HANDLE AND RELOAD, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, AND THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, ARTICULATION ISSUES WERE NOTED WITH THE HANDLE, AS THE DEVICE DID NOT ARTICULATE. THE DEVICE WAS REPLACED, AND THE SURGEON USED ANOTHER HANDLE AND RELOAD TO CONTINUE THE CASE.
30-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA | Physical Resistance/Sticking | A0509 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE INVOLVING THE HANDLE AND RELOAD, THE STAPLER WAS NOT ABLE TO FIRE. THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, AND THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, ARTICULATION ISSUES WERE NOTED WITH THE HANDLE, AS THE DEVICE DID NOT ARTICULATE. THE DEVICE WAS REPLACED, AND THE SURGEON USED ANOTHER HANDLE AND RELOAD TO CONTINUE THE CASE.
30-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE INVOLVING THE HANDLE AND RELOAD, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, AND THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, ARTICULATION ISSUES WERE NOTED WITH THE HANDLE, AS THE DEVICE DID NOT ARTICULATE. THE DEVICE WAS REPLACED, AND THE SURGEON USED ANOTHER HANDLE AND RELOAD TO CONTINUE THE CASE.
30-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE INVOLVING THE HANDLE AND RELOAD, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, AND THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, ARTICULATION ISSUES WERE NOTED WITH THE HANDLE, AS THE DEVICE DID NOT ARTICULATE. THE DEVICE WAS REPLACED, AND THE SURGEON USED ANOTHER HANDLE AND RELOAD TO CONTINUE THE CASE.
30-12-2024 | Malfunction | us surgical puerto rico | 20-01-2025 | ENDO GIA ULTRA | Physical Resistance/Sticking | A0509 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION PROCEDURE INVOLVING THE HANDLE AND RELOAD, THE STAPLER WAS NOT ABLE TO FIRE, THE SURGEON WAS UNABLE TO PRESS THE GREEN BUTTON, AND THE HANDLE COULD NOT BE SQUEEZED. ADDITIONALLY, ARTICULATION ISSUES WERE NOTED WITH THE HANDLE, AS THE DEVICE DID NOT ARTICULATE. THE DEVICE WAS REPLACED, AND THE SURGEON USED ANOTHER HANDLE AND RELOAD TO CONTINUE THE CASE.
01-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGUS PROCEDURE SURGERY, COULD NOT FIRE WITH ABNORMAL WEAK SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGUS PROCEDURE SURGERY, COULD NOT FIRE WITH ABNORMAL WEAK SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
30-12-2024 | Malfunction | covidien lp llc north haven | 21-01-2025 | UNKNOWN EGIA TRI STAPLE | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SEGMENTAL RESECTION, WHILE ON DISSOCIATING SEGMENTAL BLOOD VESSELS, THE SURGICAL HOOK HOOKED THE I-SHAPED BEAM BACK. IT WAS NOTED THAT THERE WAS UNSMOOTH FIRING. THE SURGICAL HOOK RETURNED THE I-SHAPED BEAM TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON 3000 60MM STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A MEDIASTINOSCOPIC SUBTOTAL ESOPHAGECTOMY, THE KNIFE DID NOT RETURN AFTER THE 2ND FIRING DURING A STOMACH RESECTION, AND THE JAWS WOULD NOT OPEN. THE JAWS DID NOT OPEN EVEN AFTER PUSHING THE KNIFE REVERSE SWITCH, SO THE KNIFE WAS MANUALLY RETURNED WITH THE MANUAL OVERRIDE LEVER AND THE JAWS WERE OPENED. THERE WAS NO EFFECT ON THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
24-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON 3000 60MM STANDARD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A MEDIASTINOSCOPIC SUBTOTAL ESOPHAGECTOMY, THE KNIFE DID NOT RETURN AFTER THE 2ND FIRING DURING A STOMACH RESECTION, AND THE JAWS WOULD NOT OPEN. THE JAWS DID NOT OPEN EVEN AFTER PUSHING THE KNIFE REVERSE SWITCH, SO THE KNIFE WAS MANUALLY RETURNED WITH THE MANUAL OVERRIDE LEVER AND THE JAWS WERE OPENED. THERE WAS NO EFFECT ON THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
25-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY, THE PRIMARY SURGEON PULLED THE FIRING TRIGGER AND RELEASED IT AFTER THE SOUND STOPPED, BUT THE KNIFE DID NOT RETURN AND THE JAWS DID NOT OPEN. THE OPERATION CHANGED TO THE ASSISTANT DOCTOR AND PULLED THE FIRING TRIGGER, AND IT WAS FELT THAT THE KNIFE MOVED SLIGHTLY. WHEN THE FIRING TRIGGER WAS RELEASED, THE KNIFE RETURNED AND THE JAWS OPENED. (THIS EVENT WAS THE 1ST FIRING.) FROM THE NEXT FIRING, A NEW DEVICE WAS PUT OUT TO HANDLE THE SITUATION. THE DEVICE WAS LOCKED OUT. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
25-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY, THE PRIMARY SURGEON PULLED THE FIRING TRIGGER AND RELEASED IT AFTER THE SOUND STOPPED, BUT THE KNIFE DID NOT RETURN AND THE JAWS DID NOT OPEN. THE OPERATION CHANGED TO THE ASSISTANT DOCTOR AND PULLED THE FIRING TRIGGER, AND IT WAS FELT THAT THE KNIFE MOVED SLIGHTLY. WHEN THE FIRING TRIGGER WAS RELEASED, THE KNIFE RETURNED AND THE JAWS OPENED. (THIS EVENT WAS THE 1ST FIRING.) FROM THE NEXT FIRING, A NEW DEVICE WAS PUT OUT TO HANDLE THE SITUATION. THE DEVICE WAS LOCKED OUT. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
25-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN PNEUMONECTOMY, THE PRIMARY SURGEON PULLED THE FIRING TRIGGER AND RELEASED IT AFTER THE SOUND STOPPED, BUT THE KNIFE DID NOT RETURN AND THE JAWS DID NOT OPEN. THE OPERATION CHANGED TO THE ASSISTANT DOCTOR AND PULLED THE FIRING TRIGGER, AND IT WAS FELT THAT THE KNIFE MOVED SLIGHTLY. WHEN THE FIRING TRIGGER WAS RELEASED, THE KNIFE RETURNED AND THE JAWS OPENED. (THIS EVENT WAS THE 1ST FIRING.) FROM THE NEXT FIRING, A NEW DEVICE WAS PUT OUT TO HANDLE THE SITUATION. THE DEVICE WAS LOCKED OUT. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
25-12-2024 | Malfunction | covidien lp llc north haven | 21-01-2025 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC TRANSABDOMINAL PREPERITONEAL (TAPP), WHEN TACKING THE MESH, THE FIRST TWO FIRINGS THAT WERE TACKED DID NOT GOT DETACHED FROM THE TACKER'S HANDLE AND WAS STUCK HALFWAY OUT INTO THE TISSUE. FOR THE TIME BEING, IT WAS FORCIBLY REMOVED FROM THE HANDLE, AND FIRING WAS PERFORMED WITH THE SAME HANDLE, AND THE PROCEDURE WAS COMPLETED. THERE WAS NO PATIENT INJURY.
02-01-2025 | Malfunction | covidien lp llc north haven | 21-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A MINI GASTRIC BYPASS (GASTROJEJUNEM ANASTOMOSIS), DURING THE USE OF THE RELOAD, THERE WERE ISSUES AFTER FIRING SPECIFICALLY WITH STAPLE FORMATION. THE STAPLES ON THE DISTAL TIP OF THE STAPLE LINE WERE NOT WELL-FORMED, CAUSING A LEAK, AND THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END.
02-01-2025 | Malfunction | covidien lp llc north haven | 21-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A MINI GASTRIC BYPASS (GASTROJEJUNEM ANASTOMOSIS), DURING THE USE OF THE RELOAD, THERE WERE ISSUES AFTER FIRING SPECIFICALLY WITH STAPLE FORMATION. THE STAPLES ON THE DISTAL TIP OF THE STAPLE LINE WERE NOT WELL-FORMED, CAUSING A LEAK, AND THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END.
31-12-2024 | Malfunction | covidien lp llc north haven | 21-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE AND WAS REPLACED WITH A NEW ONE. THERE WAS NO PATIENT INJURY.
02-01-2025 | Malfunction | covidien lp llc north haven | 21-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A MINI GASTRIC BYPASS (GASTROJEJUNEM ANASTOMOSIS), DURING THE USE OF THE RELOAD, THERE WERE ISSUES AFTER FIRING, SPECIFICALLY WITH STAPLE FORMATION. THE STAPLES ON THE DISTAL TIP OF THE STAPLE LINE WERE NOT WELL-FORMED, CAUSING A LEAK, AND THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END.
02-01-2025 | Malfunction | covidien lp llc north haven | 21-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A MINI GASTRIC BYPASS (GASTROJEJUNEM ANASTOMOSIS), DURING THE USE OF THE RELOAD, THERE WERE ISSUES AFTER FIRING, SPECIFICALLY WITH STAPLE FORMATION. THE STAPLES ON THE DISTAL TIP OF THE STAPLE LINE WERE NOT WELL-FORMED, CAUSING A LEAK, AND THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END.
28-11-2024 | Malfunction | covidien lp llc north haven | 21-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, BOTH STAPLE RELOADS WOULD NOT WOULD NOT CLOSE CORRECTLY WHICH CAUSED BLEEDING OF 150ML WHEN ATTEMPTING TO CUT THE BLOOD VESSELS WHICH WAS CONTROLLED QUICKLY. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
28-11-2024 | Malfunction | covidien lp llc north haven | 21-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) LOBECTOMY, BOTH STAPLE RELOADS WOULD NOT WOULD NOT CLOSE CORRECTLY WHICH CAUSED BLEEDING OF 150ML WHEN ATTEMPTING TO CUT THE BLOOD VESSELS WHICH WAS CONTROLLED QUICKLY. ANOTHER STAPLER WAS USED TO RESOLVE THE ISSUE.
13-01-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ILS 29MM, CURVED | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE AN ESOPHAGEAL PROCEDURE, OPEN THE PACKING AND NOTED ANVIL CANNOT BE ATTACHED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2025 | Malfunction | us surgical puerto rico | 21-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PULMONARY LOBECTOMY, WHEN SEVERING THE LUNG LOBE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT WAS UNABLE TO SQUEEZE THE HANDLE, AND THE DEVICE DID NOT FIRE. ANOTHER RELOAD WITH THE SAME HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
15-10-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RADICAL RESECTION OF RECTAL CANCER SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2025 | Malfunction | covidien | 21-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHEN DETACHING THE LUNG LOBE, THE STAPLER DID NOT FIRE. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2025 | Malfunction | ethicon endo-surgery | 21-01-2025 | ENDO ILS 25MM, CURVED | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE THE LAP HEMICOLECTOMY SURGERY, OPEN THE PACKING AND NOTED ANVIL CANNOT BE ATTACHED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
06-01-2025 | Malfunction | us surgical puerto rico | 21-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LOBECTOMY, WHEN DETACHING THE LUNG LOBE, THE STAPLER DID NOT FIRE. ANOTHER HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2025 | Malfunction | us surgical puerto rico | 21-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IN SEVERING THE GASTRIC BODY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
01-01-2024 | Injury | ethicon endo-surgery | 21-01-2025 | ECHELON LINEAR CUTTER 80MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY THERE WAS MORE OOZING/BLEEDING ON THE STAPLE LINE WITH THE ECHELON GLC THAN WITH THE OLDER TYPES OF NTLC AND TLC. IT IS UNCLEAR WHETHER THE ISSUE IS DEVICE-RELATED OR PATIENT-RELATED.
01-01-2024 | Injury | ethicon endo-surgery | 21-01-2025 | ECHELON LINEAR CUTTER 100MM | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY THERE WAS MORE OOZING/BLEEDING ON THE STAPLE LINE WITH THE ECHELON GLC THAN WITH THE OLDER TYPES OF NTLC AND TLC. IT IS UNCLEAR WHETHER THE ISSUE IS DEVICE-RELATED OR PATIENT-RELATED.
01-01-2025 | Malfunction | ethicon endo-surgery | 21-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE A 29 ECHELON CIRCULAR POWERED WAS OPENED AND DIDN¿T WORK, THEY CHOSE TO USE A 31 AND WHEN THEY CHANGED THE BATTERY FROM 31 TO 29 IT STILL DIDN¿T WORK. THERE WERE NO PATIENT CONSEQUENCES.
02-01-2025 | Malfunction | ethicon endo-surgery | 21-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTECTOMY PROCEDURE AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-01-2025 | Malfunction | ethicon endo-surgery | 21-01-2025 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-12-2024 | Malfunction | covidien | 21-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PARTIAL LUNG RESECTION FOR PNEUMOTHORAX, THERE WERE NO PROBLEMS WITH THE FIRST SHOT, HOWEVER, SECOND SHOT COULD NOT BE FIRED AT ALL. ADDITIONALLY, THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE RELOAD JAW. BOTH THE HANDLE RELOAD WERE REPLACED WITH NEW ONES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2024 | Malfunction | covidien | 21-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PARTIAL LUNG RESECTION FOR PNEUMOTHORAX, THERE WERE NO PROBLEMS WITH THE FIRST SHOT, HOWEVER, SECOND SHOT COULD NOT BE FIRED AT ALL. ADDITIONALLY, THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE RELOAD JAW. BOTH THE HANDLE RELOAD WERE REPLACED WITH NEW ONES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2024 | Malfunction | covidien | 21-01-2025 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PARTIAL LUNG RESECTION FOR PNEUMOTHORAX, THERE WERE NO PROBLEMS WITH THE FIRST SHOT, HOWEVER, SECOND SHOT COULD NOT BE FIRED AT ALL. ADDITIONALLY, THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE RELOAD JAW. BOTH THE HANDLE RELOAD WERE REPLACED WITH NEW ONES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2024 | Malfunction | us surgical puerto rico | 21-01-2025 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PARTIAL LUNG RESECTION FOR PNEUMOTHORAX, THERE WERE NO PROBLEMS WITH THE FIRST SHOT, HOWEVER, SECOND SHOT COULD NOT BE FIRED AT ALL. ADDITIONALLY, THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE RELOAD JAW. BOTH THE HANDLE RELOAD WERE REPLACED WITH NEW ONES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-12-2024 | Malfunction | us surgical puerto rico | 21-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE PARTIAL LUNG RESECTION FOR PNEUMOTHORAX, THERE WERE NO PROBLEMS WITH THE FIRST SHOT, HOWEVER, SECOND SHOT COULD NOT BE FIRED AT ALL. ADDITIONALLY, THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE RELOAD JAW. BOTH THE HANDLE RELOAD WERE REPLACED WITH NEW ONES TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2025 | Injury | ethicon endo-surgery | 21-01-2025 | ETHICON ENDO-SURGERY PROXIMATE PPH PROCEDURE FOR PROLAPSE AND HEMORRHOIDS SET |  |  | Hemorrhage/Blood Loss/Bleeding | Event Description: PATIENT HAD HEMORRHOIDOPEXY WITH ETHICON PPH STAPLER, CPT 46947. WHILE IN PACU(POST-ANESTHESIA CARE UNIT) HAD RECTAL BLEEDING AND RETURNED TO THE OPERATING ROOM FOR EVACUATION OF CLOT, ANORECTAL EXAM AND REINFORCEMENT OF STAPLE LINE WITH SUTURES. THE ONLY ITEM IMPLANTED WAS THE STAPLES. PRE-OP HGB 13.7. HOURS AFTER SURGERY ON 1-6 WAS 12.9 AND DAY AFTER RETURN TO SURGERY WAS 11.7.
06-01-2025 | Injury | ethicon endo-surgery | 21-01-2025 | PROXIMATE PPH PROCEDURE FOR PROLAPSE AND HEMORRHOIDS SET |  |  | Hemorrhage/Blood Loss/Bleeding | Event Description: SHE HAD STAPLED HEMORROIDOPEXY WITH ETHICON HEMORRHOID PPH STAPLER CPT 46947 AND RETURNED TO THE ED THE FOLLOWING DAY WITH RECTAL BLEEDING. THIS WAS REPORTED FROM HOME AND OBSERVED IN THE ED. PATIENT TO THE OR FOR ANORECTAL EXAM AND, ALTHOUGH NO ACTIVE BLEEDING WAS NOTED AT THAT TIME, THE STAPLE LINE WAS REINFORCED WITH SUTURES AND SHE WAS DISCHARGED. PATIENT PRE-OP HEMOGLOBIN WAS 16, ON ARRIVAL TO EMERGENCY ROOM ON 1-7 WAS 14 AND THEN TO 12.4.
06-01-2025 | Injury | ethicon endo-surgery | 21-01-2025 | PROXIMATE PPH PROCEDURE FOR PROLAPSE AND HEMORRHOIDS SET |  |  | Rectal Anastomotic Leakage | Event Description: SHE HAD STAPLED HEMORROIDOPEXY WITH ETHICON HEMORRHOID PPH STAPLER CPT 46947 AND RETURNED TO THE ED THE FOLLOWING DAY WITH RECTAL BLEEDING. THIS WAS REPORTED FROM HOME AND OBSERVED IN THE ED. PATIENT TO THE OR FOR ANORECTAL EXAM AND, ALTHOUGH NO ACTIVE BLEEDING WAS NOTED AT THAT TIME, THE STAPLE LINE WAS REINFORCED WITH SUTURES AND SHE WAS DISCHARGED. PATIENT PRE-OP HEMOGLOBIN WAS 16, ON ARRIVAL TO EMERGENCY ROOM ON 1-7 WAS 14 AND THEN TO 12.4.
08-01-2025 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
08-01-2025 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
08-01-2025 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
08-01-2025 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
03-12-2024 | Malfunction | johnson & johnson medical (suzhou | 21-01-2025 | CARTRIDGE FINAL ASSY 45MM WHITE | Structural Problem | A0511 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, BEFORE FIRING AND NOTED THAT SOME STAPLES PROTRUDING. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
04-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | GREEN RELOAD FOR ECHELON 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PANCREAS PROCEDURE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON FLEX 60 LONG | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, AFTER FIRED, COULD NOT OPEN THE JAW. REPEATED MANY TIMES TO OPEN THE JAW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
09-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | GREEN RELOAD FOR ECHELON 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | BLUE RELOAD FOR ECHELON 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LIVER RESECTION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | BLUE RELOAD FOR ECHELON 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LIVER RESECTION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
05-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | BLUE RELOAD FOR ECHELON 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
02-12-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 21-01-2025 | CAPSURE | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING INTRAPERITONEAL-ONLAY-MESH REPAIR (IPOM PLUS), THE FASTENERS OF CAPSURE DEVICE DID NOT COME OUT AFTER 3 TO 4 FASTENERS GOT FIRED SUCCESSFULLY. IT WAS REPORTED THAT TRIGGER FUNCTIONED NORMAL AND SURGEON DRY FIRED THE DEVICE OUTSIDE TO CHECK BUT PROBLEM REMAIN SAME. THE SURGEON USED ANOTHER FIXATION TO COMPLETE SURGERY, WHICH INCREASED SURGERY TIME BY 35-40 MINUTES. IT WAS REPORTED THAT NO OTHER MEDICAL INTERVENTION WAS PERFORMED.
04-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON 3000 45 STANDARD | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER THE RECTUM WAS DISSECTED WITH 2 FIRINGS OF ECHELON 45, THE ANGLE DID NOT RETURN TO STRAIGHT (IT DID NOT COME OUT FROM THE TROCAR). THE ARTICULATION DID NOT FUNCTION PROPERLY. A REPLACEMENT WAS PUT OUT AND THE SURGERY WAS COMPLETED WITHOUT ANY PROBLEMS. THE CARTRIDGE COLOR WAS BLUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
04-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON 3000 45 STANDARD | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER THE RECTUM WAS DISSECTED WITH 2 FIRINGS OF ECHELON 45, THE ANGLE DID NOT RETURN TO STRAIGHT (IT DID NOT COME OUT FROM THE TROCAR). THE ARTICULATION DID NOT FUNCTION PROPERLY. A REPLACEMENT WAS PUT OUT AND THE SURGERY WAS COMPLETED WITHOUT ANY PROBLEMS. THE CARTRIDGE COLOR WAS BLUE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
22-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON 3000 45MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN RESPIRATORY SURGERY, THE DEVICE COULD BE BENT, BUT THE KNIFE DID NOT MOVE FORWARD AND THE DEVICE COULD NOT BE FIRED. THE CARTRIDGE COLOR WAS GOLD. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
13-01-2025 | Malfunction | ethicon endo-surgery | 21-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, THE KNIFE DID NOT MOVE FORWARD AND THE DEVICE COULD NOT BE FIRED. THE DEVICE STILL WOULD NOT WORK EVEN CHANGING CARTRIDGE. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-02-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE RELOAD DID NOT FIT, IT BECAME LOOSE. THE PRODUCT WAS REPLACED TO FINISH THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES.
26-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON FLEX 45 STANDARD | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AT THE TIME OF STAPLING, THE STAPLER TRIGGER WAS LOOSE. THERE WAS NO SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
26-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
27-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | GST GREEN RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A WEDGE RESECTION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | GST GREEN RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A WEDGE RESECTION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTROPLASTY THE STAPLER LOCKED AND DID NOT UNLOCK ANY MORE, EVEN MAKING THE RED NOTION MECHANISM DOWN TO UNLOCK.
12-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTROPLASTY THE STAPLER LOCKED AND DID NOT UNLOCK ANY MORE, EVEN MAKING THE RED NOTION MECHANISM DOWN TO UNLOCK.
29-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | GST WHITE RELOAD, 45MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ECHELON FLEX 45 STANDARD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2020 | Malfunction | us surgical puerto rico | 21-01-2025 | UNKNOWN EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
12-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CAVE IS BROKEN OFF.
12-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CAVE IS BROKEN OFF.
12-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CAVE IS BROKEN OFF.
12-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CAVE IS BROKEN OFF.
12-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CAVE IS BROKEN OFF.
12-11-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CAVE IS BROKEN OFF.
01-01-2024 | Malfunction | ethicon endo-surgery | 21-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Mechanical Problem | A05 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER BLADE DID NOT RETRACT AFTER 2 RELOADS USED. THERE WERE NO PATIENT CONSEQUENCES.
05-01-2025 | Malfunction | us surgical puerto rico | 21-01-2025 | UNKNOWN GIA DST PRODUCT | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WAS AN EVENT INVOLVING THE RELOAD, THERE WERE MISFIRE PRODUCT ISSUES. THE PATIENT WAS REPORTED TO BE ALIVE WITH NO INJURY.
05-01-2025 | Malfunction | us surgical puerto rico | 21-01-2025 | UNKNOWN GIA DST | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR) PROCEDURE, THE STAPLER RELOADS WERE NOT ABLE TO FIRE PROPERLY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-01-2025 | Malfunction | us surgical puerto rico | 21-01-2025 | UNKNOWN GIA DST | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION (LAR) PROCEDURE, THE STAPLER RELOADS WERE NOT ABLE TO FIRE PROPERLY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | ethicon | 21-01-2025 | ETHICON SECURESTRAP 5MM ABSORBABLE STRAP | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HERNIORRAPHY PROCEDURE ON (B)(6) 2024 AND ABSORBABLE STRAPS WERE USED. PROBLEM DURING STAPLING WAS NOTED. THE STAPLES WERE NOT FIXING THE SCREEN, SEVERAL ATTEMPTS WERE MADE BUT WITHOUT SUCCESS, THE DOCTOR TRIGGERED THE SHOT AND THE CLIP FELL INTO THE CAVITY. HE USED ANOTHER STAPLER AND THE SURGERY FINISHED WELL AND WITHOUT FURTHER COMPLICATIONS. PER MEDICAL REPORT: "WHEN USING THE STAPLER, WE WERE SURPRISED BY THE FAILURE OF THE STAPLER DURING ITS USE, ALTERNATING A CLIP WITH THE SHOT FAILURE OF THE NEXT." THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION WAS PROVIDED.
06-01-2025 | Malfunction | us surgical puerto rico | 21-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, IN SEVERING THE GASTRIC BODY, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE AND RELOAD WERE USED. THERE WAS NO PATIENT INJURY.
16-12-2024 | Malfunction | covidien lp llc north haven | 21-01-2025 | ENDO GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING ESOPHAGECTOMY THE STAPLER FIRED BUT THE BLADE DID NOT CUT BETWEEN THE STAPLE LINES. THE BLADE ONLY CUT ABOUT 20MM OF THE 60MM RELOAD, IT WAS ALSO NOTED THAT THE STAPLE LINE WAS HIGHLY DISTORTED. AS A RESOLUTION A NEW STAPLER AND RELOAD WERE OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THERE WERE UNKNOWN ISSUES PRIOR TO FIRING WITH THE EGIA HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH ANOTHER ONE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
12-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA ULTRA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THERE WERE UNKNOWN ISSUES PRIOR TO FIRING WITH THE EGIA HANDLE. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE WITH ANOTHER ONE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
17-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE INSTRUMENT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON AND SQUEEZE THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-12-2024 | Injury | intuitive surgical | 22-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY SURGICAL PROCEDURE, POSTOPERATIVELY THE PATIENT SUSTAINED AN ILEOCOLIC ANASTOMOSIS DEHISCENCE AND REQUIRED A SUBSEQUENT PROCEDURE. A SUREFORM 60 STAPLER INSTRUMENT WITH A SUREFORM 60 BLUE RELOAD WAS UTILIZED DURING THE INITIAL PROCEDURE. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
03-12-2024 | Injury | intuitive surgical | 22-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY SURGICAL PROCEDURE, POSTOPERATIVELY THE PATIENT SUSTAINED AN ILEOCOLIC ANASTOMOSIS DEHISCENCE AND REQUIRED A SUBSEQUENT PROCEDURE. A SUREFORM 60 STAPLER INSTRUMENT WITH A SUREFORM 60 BLUE RELOAD WAS UTILIZED DURING THE INITIAL PROCEDURE. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Material Fragmentation | A040103 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WAS SHAVING OF METAL THAT CAME OUT OF THE TOP OF THE STAPLER RELOAD I-BEAM CAVITY. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Material Fragmentation | A040103 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WAS SHAVING OF METAL THAT CAME OUT OF THE TOP OF THE STAPLER RELOAD I-BEAM CAVITY. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Material Fragmentation | A040103 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WAS SHAVING OF METAL THAT CAME OUT OF THE TOP OF THE STAPLER RELOAD I-BEAM CAVITY. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WAS SHAVING OF METAL THAT CAME OUT OF THE TOP OF THE STAPLER RELOAD I-BEAM CAVITY. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WAS SHAVING OF METAL THAT CAME OUT OF THE TOP OF THE STAPLER RELOAD I-BEAM CAVITY. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WAS SHAVING OF METAL THAT CAME OUT OF THE TOP OF THE STAPLER RELOAD I-BEAM CAVITY. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Material Fragmentation | A040103 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WERE ISSUES WITH THE DEVICE NOT STAPLING PROPERLY. SPECIFICALLY, THE STAPLES "CRUMPLED" AND DID NOT FORM PROPER FUNCTIONAL STAPLES WHEN FIRED. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Material Fragmentation | A040103 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WERE ISSUES WITH THE DEVICE NOT STAPLING PROPERLY. SPECIFICALLY, THE STAPLES "CRUMPLED" AND DID NOT FORM PROPER FUNCTIONAL STAPLES WHEN FIRED. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Material Fragmentation | A040103 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WERE ISSUES WITH THE DEVICE NOT STAPLING PROPERLY. SPECIFICALLY, THE STAPLES "CRUMPLED" AND DID NOT FORM PROPER FUNCTIONAL STAPLES WHEN FIRED. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WERE ISSUES WITH THE DEVICE NOT STAPLING PROPERLY. SPECIFICALLY, THE STAPLES "CRUMPLED" AND DID NOT FORM PROPER FUNCTIONAL STAPLES WHEN FIRED. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WERE ISSUES WITH THE DEVICE NOT STAPLING PROPERLY. SPECIFICALLY, THE STAPLES "CRUMPLED" AND DID NOT FORM PROPER FUNCTIONAL STAPLES WHEN FIRED. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
06-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THORACIC SURGERY INVOLVING THE RELOADS, THERE WERE ISSUES WITH THE DEVICE NOT STAPLING PROPERLY. SPECIFICALLY, THE STAPLES "CRUMPLED" AND DID NOT FORM PROPER FUNCTIONAL STAPLES WHEN FIRED. FOUR LOADS WERE OPENED ON THE TABLE, AND TWO OF THE RELOADS MALFUNCTIONED.
26-11-2024 | Injury | intuitive surgical | 22-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY SURGICAL PROCEDURE, THE PATIENT SUSTAINED AN ILEOCOLIC ANASTOMOSIS DEHISCENCE AND REQUIRED A SUBSEQUENT PROCEDURE. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
26-11-2024 | Injury | intuitive surgical | 22-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Unspecified Tissue Injury | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY SURGICAL PROCEDURE, THE PATIENT SUSTAINED AN ILEOCOLIC ANASTOMOSIS DEHISCENCE AND REQUIRED A SUBSEQUENT PROCEDURE. INTUITIVE SURGICAL, INC. (ISI) HAS ATTEMPTED TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, AS OF THE DATE OF THIS REPORT, NO FURTHER DETAILS HAVE BEEN RECEIVED.
30-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING A MINI GASTRIC BYPASS, THERE WERE ISSUES AFTER FIRING SPECIFICALLY WITH STAPLE FORMATION. THE STAPLES ON THE DISTAL TIP OF THE STAPLE LINE WERE NOT WELL-FORMED CAUSING A LEAK AND THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END.
30-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | TRI-STAPLE 2.0 | Failure to Form Staple | A050704 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A MINI GASTRIC BYPASS, THERE WERE ISSUES AFTER FIRING SPECIFICALLY WITH STAPLE FORMATION. THE STAPLES ON THE DISTAL TIP OF THE STAPLE LINE WERE NOT WELL-FORMED CAUSING A LEAK AND THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END.
30-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER DURING A MINI GASTRIC BYPASS, THERE WERE ISSUES AFTER FIRING SPECIFICALLY WITH STAPLE FORMATION. THE STAPLES ON THE DISTAL TIP OF THE STAPLE LINE WERE NOT WELL-FORMED CAUSING A LEAK AND THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END.
30-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | TRI-STAPLE 2.0 | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A MINI GASTRIC BYPASS, THERE WERE ISSUES AFTER FIRING SPECIFICALLY WITH STAPLE FORMATION. THE STAPLES ON THE DISTAL TIP OF THE STAPLE LINE WERE NOT WELL-FORMED CAUSING A LEAK AND THERE WAS AN INCOMPLETE STAPLE LINE AT THE DISTAL END.
01-01-2020 | Malfunction | us surgical puerto rico | 22-01-2025 | UNKNOWN EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | us surgical puerto rico | 22-01-2025 | UNKNOWN EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | us surgical puerto rico | 22-01-2025 | UNKNOWN EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: DUPLICATE PE OF 0707447694 ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | us surgical puerto rico | 22-01-2025 | UNKNOWN EEA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: DUPLICATE PE OF 0707447694 ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC TUBE ANGIOPLASTY FOR ESOPHAGEAL CANCER, THE RELEASE BUTTON COULD NOT PRESS AND THE JAWS COULD OPEN AFTER FIRING, SO THE GRASPED TISSUE WAS PULLED OUT THROUGH THE GAP. THE STAPLES WERE MALFORMED. IT WAS USED ON GASTRIC BODY. ADDITIONAL HAND SUTURE WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
10-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC TUBE ANGIOPLASTY FOR ESOPHAGEAL CANCER, THE RELEASE BUTTON COULD NOT PRESS AND THE JAWS COULD OPEN AFTER FIRING, SO THE GRASPED TISSUE WAS PULLED OUT THROUGH THE GAP. THE STAPLES WERE MALFORMED. IT WAS USED ON GASTRIC BODY. ADDITIONAL HAND SUTURE WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
10-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Difficult to Open or Close | A051104 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC TUBE ANGIOPLASTY FOR ESOPHAGEAL CANCER, THE RELEASE BUTTON COULD NOT PRESS AND THE JAWS COULD OPEN AFTER FIRING, SO THE GRASPED TISSUE WAS PULLED OUT THROUGH THE GAP. THE STAPLES WERE MALFORMED. IT WAS USED ON GASTRIC BODY. ADDITIONAL HAND SUTURE WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
10-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC TUBE ANGIOPLASTY FOR ESOPHAGEAL CANCER, THE RELEASE BUTTON COULD NOT PRESS AND THE JAWS COULD OPEN AFTER FIRING, SO THE GRASPED TISSUE WAS PULLED OUT THROUGH THE GAP. THE STAPLES WERE MALFORMED. IT WAS USED ON GASTRIC BODY. ADDITIONAL HAND SUTURE WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
10-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC TUBE ANGIOPLASTY FOR ESOPHAGEAL CANCER, THE RELEASE BUTTON COULD NOT PRESS AND THE JAWS COULD OPEN AFTER FIRING, SO THE GRASPED TISSUE WAS PULLED OUT THROUGH THE GAP. THE STAPLES WERE MALFORMED. IT WAS USED ON GASTRIC BODY. ADDITIONAL HAND SUTURE WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
10-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC GASTRIC TUBE ANGIOPLASTY FOR ESOPHAGEAL CANCER, THE RELEASE BUTTON COULD NOT PRESS AND THE JAWS COULD OPEN AFTER FIRING, SO THE GRASPED TISSUE WAS PULLED OUT THROUGH THE GAP. THE STAPLES WERE MALFORMED. IT WAS USED ON GASTRIC BODY. ADDITIONAL HAND SUTURE WAS PERFORMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
31-12-2024 | Malfunction | covidien lp llc north haven | 22-01-2025 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
31-12-2024 | Malfunction | covidien lp llc north haven | 22-01-2025 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE AND ADAPTER TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-12-2024 | Injury | covidien lp llc north haven | 22-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A HEMICOLECTOMY WITH LATERO-LATERAL ANASTOMOSIS, THE STAPLE LINE WAS SEPARATED AND THERE WAS A SUTURE DEHISCENCE AT THE OPENING OF THE STITCHES OF THE SURGICAL SCAR. THE STAPLE LINE WAS RE-STAPLED AND RESECTED AND THE PATIENT HAD THREE OSTOMIES APPROACHED. HOSPITALIZATION WAS EXTENDED FOR 20 DAYS, AND THE PATIENT WAS IN THE INTENSIVE CARE UNIT.
13-12-2024 | Injury | covidien lp llc north haven | 22-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Laceration(s) | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A HEMICOLECTOMY WITH LATERO-LATERAL ANASTOMOSIS, THE STAPLE LINE WAS SEPARATED AND THERE WAS A SUTURE DEHISCENCE AT THE OPENING OF THE STITCHES OF THE SURGICAL SCAR. THE STAPLE LINE WAS RE-STAPLED AND RESECTED AND THE PATIENT HAD THREE OSTOMIES APPROACHED. HOSPITALIZATION WAS EXTENDED FOR 20 DAYS, AND THE PATIENT WAS IN THE INTENSIVE CARE UNIT.
13-12-2024 | Injury | covidien lp llc north haven | 22-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Organ Dehiscence | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A HEMICOLECTOMY WITH LATERO-LATERAL ANASTOMOSIS, THE STAPLE LINE WAS SEPARATED AND THERE WAS A SUTURE DEHISCENCE AT THE OPENING OF THE STITCHES OF THE SURGICAL SCAR. THE STAPLE LINE WAS RE-STAPLED AND RESECTED AND THE PATIENT HAD THREE OSTOMIES APPROACHED. HOSPITALIZATION WAS EXTENDED FOR 20 DAYS, AND THE PATIENT WAS IN THE INTENSIVE CARE UNIT.
13-12-2024 | Injury | covidien lp llc north haven | 22-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, FOLLOWING A HEMICOLECTOMY WITH LATERO-LATERAL ANASTOMOSIS, THE STAPLE LINE WAS SEPARATED AND THERE WAS A SUTURE DEHISCENCE AT THE OPENING OF THE STITCHES OF THE SURGICAL SCAR. THE STAPLE LINE WAS RE-STAPLED AND RESECTED AND THE PATIENT HAD THREE OSTOMIES APPROACHED. HOSPITALIZATION WAS EXTENDED FOR 20 DAYS, AND THE PATIENT WAS IN THE INTENSIVE CARE UNIT.
31-12-2024 | Malfunction | covidien lp llc north haven | 22-01-2025 | SIGNIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE AND WAS REPLACED WITH A NEW ONE. THERE WAS NO PATIENT INJURY.
31-12-2024 | Malfunction | covidien lp llc north haven | 22-01-2025 | SIGNIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE DEVICE DID NOT FIRE AND WAS REPLACED WITH A NEW ONE. THERE WAS NO PATIENT INJURY.
20-12-2024 | Malfunction | davis & geck caribe | 22-01-2025 | UNKNOWN ENDO STITCH SULU | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC HYSTERECTOMY, THE NEEDLE DID NOT STAY ON THE DEVICE. A NEW ENDOSTITCH WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | us surgical puerto rico | 22-01-2025 | UNKNOWN EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
26-12-2024 | Malfunction | covidien lp llc north haven | 22-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, WHEN THE STOMACH WAS RESECTED AND THE FIRING WAS ABOUT HALFWAY COMPLETE, THE FIRING STOPPED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | us surgical puerto rico | 22-01-2025 | UNKNOWN EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
26-12-2024 | Malfunction | covidien lp llc north haven | 22-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, WHEN THE STOMACH WAS RESECTED AND THE FIRING WAS ABOUT HALFWAY COMPLETE, THE FIRING STOPPED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-11-2024 | Malfunction | covidien | 22-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE RELOAD COMPONENT DURING SET-UP AND PRIOR TO FIRING, SPECIFICALLY THAT THE RELOAD WAS NOT RECOGNIZED. TO RESOLVE THE ISSUE, THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS UNREPORTED CONDITION OF THICK TISSUE. THE CLAMPING MECHANISM WAS DEFORMED. THE RELOAD HAD DAMAGE TO THE CUTTING EDGE OF THE KNIFE BLADE.
01-11-2024 | Malfunction | covidien | 22-01-2025 | ENDO GIA | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WERE ISSUES WITH THE RELOAD COMPONENT DURING SET-UP AND PRIOR TO FIRING, SPECIFICALLY THAT THE RELOAD WAS NOT RECOGNIZED. TO RESOLVE THE ISSUE, THE RELOAD WAS REPLACED. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS UNREPORTED CONDITION OF THICK TISSUE. THE CLAMPING MECHANISM WAS DEFORMED. THE RELOAD HAD DAMAGE TO THE CUTTING EDGE OF THE KNIFE BLADE.
01-01-2020 | Malfunction | us surgical puerto rico | 22-01-2025 | UNKNOWN EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
01-01-2020 | Malfunction | us surgical puerto rico | 22-01-2025 | UNKNOWN EEA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE DID NOT FIRE. THERE WAS NO PATIENT INJURY.
14-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TOTAL COLECTOMY FOR THE REMOVAL OF THE ENTIRE COLON IN A PATIENT WITH FAP (FAMILIAL ADENOMATOUS POLYPOSIS), PERFORMED BY PROFESSOR AND DR., A LOWER ANASTOMOSIS BETWEEN THE ILEUM AND THE ANAL CANAL (ILEOANAL ANASTOMOSIS) WAS CARRIED OUT THROUGH A PFANNENSTIEL INCISION. THE INITIAL ANASTOMOSIS WAS ATTEMPTED USING THE CDH29P DEVICE. AFTER SUTURING THE ANVIL TO THE ILEUM, THE DEVICE BODY WAS INSERTED THROUGH THE ANAL CANAL, AND THEN THE TROCAR WAS REMOVED IN A ROTATIONAL MOTION TO CONNECT IT TO THE ANVIL. ONCE THE TROCAR WAS ATTACHED TO THE DEVICE (A "CLICK" WAS HEARD INDICATING PROPER ATTACHMENT), THE SURGEONS BEGAN ROTATING THE HANDLE TO PERFORM THE ANASTOMOSIS. HOWEVER, AFTER A QUARTER TURN, THE ANVIL AND TROCAR IN THE DEVICE DETACHED. THE SURGEONS TRIED TO PERFORM THE PROCEDURE SEVERAL TIMES, BUT EACH TIME THE TROCAR DETACHED WHEN THE HANDLE WAS ROTATED COUNTERCLOCKWISE. PROFESSOR VERIFIED THAT THE ANVIL WAS NOT CONTAMINATED WITH OIL AND EVEN REQUESTED GAUZE TO CLEAN THE ANVIL AND ENSURE IT WAS CLEAN. AT THIS POINT, A COMPANY REPRESENTATIVE WAS CALLED TO THE OPERATING ROOM TO CONFIRM THAT THE DEVICE WAS FUNCTIONING CORRECTLY. AFTER SEEING NO ISSUE, THEY TRIED TO PERFORM THE ANASTOMOSIS AGAIN, BUT ONCE AGAIN THE TROCAR AND ANVIL DETACHED. THE DEVICE WAS THEN REPLACED WITH A NEW CDH29P, AND THE ANASTOMOSIS WAS ATTEMPTED AGAIN, BUT IT FAILED FOR THE SAME REASON. AT THIS POINT, THE CIRCULAR STAPLER WAS REPLACED WITH AN ECS DEVICE, AND THE ANASTOMOSIS WAS SUCCESSFULLY COMPLETED. THE SURGERY PROCEEDED AS PLANNED WITH A SLIGHT DELAY OF ABOUT HALF AN HOUR DUE TO THE DEVICE REPLACEMENT. THERE WAS NO HARM TO THE PATIENT. THERE WAS A 30 MINUTE SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
14-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A TOTAL COLECTOMY FOR THE REMOVAL OF THE ENTIRE COLON IN A PATIENT WITH FAP (FAMILIAL ADENOMATOUS POLYPOSIS), PERFORMED BY PROFESSOR AND DR., A LOWER ANASTOMOSIS BETWEEN THE ILEUM AND THE ANAL CANAL (ILEOANAL ANASTOMOSIS) WAS CARRIED OUT THROUGH A PFANNENSTIEL INCISION. THE INITIAL ANASTOMOSIS WAS ATTEMPTED USING THE CDH29P DEVICE. AFTER SUTURING THE ANVIL TO THE ILEUM, THE DEVICE BODY WAS INSERTED THROUGH THE ANAL CANAL, AND THEN THE TROCAR WAS REMOVED IN A ROTATIONAL MOTION TO CONNECT IT TO THE ANVIL. ONCE THE TROCAR WAS ATTACHED TO THE DEVICE (A "CLICK" WAS HEARD INDICATING PROPER ATTACHMENT), THE SURGEONS BEGAN ROTATING THE HANDLE TO PERFORM THE ANASTOMOSIS. HOWEVER, AFTER A QUARTER TURN, THE ANVIL AND TROCAR IN THE DEVICE DETACHED. THE SURGEONS TRIED TO PERFORM THE PROCEDURE SEVERAL TIMES, BUT EACH TIME THE TROCAR DETACHED WHEN THE HANDLE WAS ROTATED COUNTERCLOCKWISE. PROFESSOR VERIFIED THAT THE ANVIL WAS NOT CONTAMINATED WITH OIL AND EVEN REQUESTED GAUZE TO CLEAN THE ANVIL AND ENSURE IT WAS CLEAN. AT THIS POINT, A COMPANY REPRESENTATIVE WAS CALLED TO THE OPERATING ROOM TO CONFIRM THAT THE DEVICE WAS FUNCTIONING CORRECTLY. AFTER SEEING NO ISSUE, THEY TRIED TO PERFORM THE ANASTOMOSIS AGAIN, BUT ONCE AGAIN THE TROCAR AND ANVIL DETACHED. THE DEVICE WAS THEN REPLACED WITH A NEW CDH29P, AND THE ANASTOMOSIS WAS ATTEMPTED AGAIN, BUT IT FAILED FOR THE SAME REASON. AT THIS POINT, THE CIRCULAR STAPLER WAS REPLACED WITH AN ECS DEVICE, AND THE ANASTOMOSIS WAS SUCCESSFULLY COMPLETED. THE SURGERY PROCEEDED AS PLANNED WITH A SLIGHT DELAY OF ABOUT HALF AN HOUR DUE TO THE DEVICE REPLACEMENT. THERE WAS NO HARM TO THE PATIENT. THERE WAS A 30 MINUTE SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
13-01-2025 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT AN ANTERIOR LOWER RESECTION, WHEN SHOOTING THE CONTOUR, IT WAS ALREADY ACTIVATED, WITH THE STAPLES OUTGOING, NOT FORMING STAPLING LINE CORRECTLY AND MUST CHANGE THE RELOAD. THERE WAS A 10 MINUTE SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
15-10-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-11-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE CLOSURE OF AN ILEOSTOMY CASE, THE SURGEON WAS UNABLE TO FIRE THE FIRST RELOAD. A SCRUBBING NURSE NOTED THAT THE STAPLER ASSEMBLY FELT SLIGHTLY DIFFERENT THAN USUAL. A NEW STAPLER WAS OPENED, AND THE FIRING WAS SUCCESSFULLY COMPLETED WITH IT. THERE WERE NO PATIENT CONSEQUENCES.
07-01-2025 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WEDGE RESECTION PROCEDURE, THE PHYSICIAN WAS ABLE TO OPEN AND CLOSE THE RELOAD (BLACK 60MM), ARTICULATE AND ROTATE, BUT UPON CLOSING ON THE TISSUE AND PRESSING THE GREEN BUTTON THE STAPLER WOULD NOT ENGAGE THE FIRING SEQUENCE. THE ISSUE PERSISTED EVEN WHEN THE DEVICE WAS TESTED OUTSIDE OF THE PATIENT CAVITY. THE DOCTOR HAD A NEW HANDLE OPEN AND ATTACHED THE SAME BLACK 60MM RELOAD FROM THE PREVIOUS HANDLE AND WAS ABLE TO FIRE THE RELOAD, WHICH ALLOWED THE PROCEDURE TO BE COMPLETED SUCCESSFULLY WITH MULTIPLE FIRINGS WITHOUT ISSUE. THERE WAS NO PATIENT INJURY.
10-12-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | 75MM SELECTABLE NEW TLC | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
17-12-2024 | Malfunction | covidien lp llc north haven | 22-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE INSTRUMENT DID NOT FIRE, THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON AND SQUEEZE THE HANDLE. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-12-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, WHEN THE THE JEJUNAL STUMP WAS TO BE RESECTED WITH THE DEVICE. UPON PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, AN ABNORMAL "GRINDING" SOUND OCCURRED FROM THE STAPLER, AND THE HANDLE COULD BE SQUEEZED WITHOUT RESISTANCE. AFTERWARDS THE HANDLE AND CARTRIDGE WERE REPLACED WITH NEW ONES AND THE RESECTION WAS SUCCESSFULLY COMPLETED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK DUE TO THICK TISSUE.
16-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, WHEN THE THE JEJUNAL STUMP WAS TO BE RESECTED WITH THE DEVICE. UPON PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, AN ABNORMAL "GRINDING" SOUND OCCURRED FROM THE STAPLER, AND THE HANDLE COULD BE SQUEEZED WITHOUT RESISTANCE. AFTERWARDS THE HANDLE AND CARTRIDGE WERE REPLACED WITH NEW ONES AND THE RESECTION WAS SUCCESSFULLY COMPLETED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK DUE TO THICK TISSUE.
16-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC TOTAL GASTRECTOMY, WHEN THE THE JEJUNAL STUMP WAS TO BE RESECTED WITH THE DEVICE. UPON PRESSING THE GREEN BUTTON AND SQUEEZING THE HANDLE, AN ABNORMAL "GRINDING" SOUND OCCURRED FROM THE STAPLER, AND THE HANDLE COULD BE SQUEEZED WITHOUT RESISTANCE. AFTERWARDS THE HANDLE AND CARTRIDGE WERE REPLACED WITH NEW ONES AND THE RESECTION WAS SUCCESSFULLY COMPLETED. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK DUE TO THICK TISSUE.
08-11-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | GREEN RELOAD FOR ECHELON 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A WEDGE RESECTION SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
31-10-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON 3000 45MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC PNEUMONECTOMY, THE BATTERY WAS INSERTED BUT THERE WAS NO OPERATING SOUND. THE BATTERY WAS INSERTED AGAIN BUT THERE WAS NO SOUND. THE BUTTON ON THE SIDE DID NOT WORK EITHER. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
06-01-2025 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RIGHT THORACOSCOPIC PARTIAL LUNG RESECTION, THE FIRING WAS TRIED BUT COULD NOT BE DONE. WHEN ANOTHER CARTRIDGE WAS USED WITH THE SAID HANDLE, IT COULD BE USED WITHOUT ANY PROBLEMS, SO IT WAS DETERMINED THAT THERE WAS NO PROBLEM WITH THE HANDLE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | TRI-STAPLE 2.0 | Failure to Cut | A050702 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE SIGTRSB45AXT 45MM XTR THK REINFORCED RELOAD, THERE WERE FIRING ISSUES WITH THE RELOAD COMPONENT, AND IT DID NOT CUT AT ALL BUT STAPLED. THE DEVICE WAS REPLACED WITH A NEW ONE.
26-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRECTOMY PROCEDURE, WHEN THE STOMACH WAS RESECTED AND THE FIRING WAS ABOUT HALFWAY COMPLETE, THE FIRING STOPPED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
27-10-2024 | Injury | us surgical puerto rico | 22-01-2025 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Hematoma | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY OF RIGHT SUBCOSTAL INCISIONAL HERNIAS FROM 2014-2024 COMPARED OUTCOMES OF 46 TOTALLY ENDOSCOPIC RETRO MUSCULAR HERNIA REPAIRS WITH TRANSVERSUS ABDOMINIS RELEASE (ETEP GROUP) AND 51 CONVENTIONAL LAPAROSCOPIC INCISIONAL HERNIA REPAIRS WITH DEFECT CLOSURE (IPOM+ GROUP). IN THE IPOM+ GROUP, THE HERNIA DEFECT WAS CLOSED WITH A NONABSORBABLE VLOC SUTURE. SYMBOTEX COMPOSITE OR A COMPETITOR MESH WAS USED TO COVER THE CLOSED DEFECT AND NONABSORBABLE PROTACK WAS USED FOR MESH FIXATION. IN THE ETEP GROUP, A SPACEMAKER WAS USED TO RELEASE THE RETRO MUSCULAR SPACE AND SUBSEQUENTLY REMOVED. THE ANTERIOR AND PERITONEAL DEFECTS WERE CLOSED USING A VLOC SUTURE AND SYMBOTEX OR A COMPETITOR MESH WAS GLUED IN TO COVER THE RETRO MUSCULAR SPACE. POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING, ILEUS, SEROMAS, HEMATOMAS, HERNIA RECURRENCE, AND POSTOPERATIVE CHRONIC PAIN. INTRAOPERATIVE BLEEDING DUE TACK FIXATION WAS RESOLVED INTRAOPERATIVELY. READMISSIONS WERE REQUIRED FOR ILEUS AND POSTOPERATIVE PAIN. NO ADDITIONAL INTERVENTIONS WERE REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED PERITONEAL AND SEROSAL TEARS AND INTRAOPERATIVE VESSEL INJURY. THERE WERE NO COMPLICATIONS RELATED TO THE SPACEMAKER DEVICE.
27-10-2024 | Injury | us surgical puerto rico | 22-01-2025 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY OF RIGHT SUBCOSTAL INCISIONAL HERNIAS FROM 2014-2024 COMPARED OUTCOMES OF 46 TOTALLY ENDOSCOPIC RETRO MUSCULAR HERNIA REPAIRS WITH TRANSVERSUS ABDOMINIS RELEASE (ETEP GROUP) AND 51 CONVENTIONAL LAPAROSCOPIC INCISIONAL HERNIA REPAIRS WITH DEFECT CLOSURE (IPOM+ GROUP). IN THE IPOM+ GROUP, THE HERNIA DEFECT WAS CLOSED WITH A NONABSORBABLE VLOC SUTURE. SYMBOTEX COMPOSITE OR A COMPETITOR MESH WAS USED TO COVER THE CLOSED DEFECT AND NONABSORBABLE PROTACK WAS USED FOR MESH FIXATION. IN THE ETEP GROUP, A SPACEMAKER WAS USED TO RELEASE THE RETRO MUSCULAR SPACE AND SUBSEQUENTLY REMOVED. THE ANTERIOR AND PERITONEAL DEFECTS WERE CLOSED USING A VLOC SUTURE AND SYMBOTEX OR A COMPETITOR MESH WAS GLUED IN TO COVER THE RETRO MUSCULAR SPACE. POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING, ILEUS, SEROMAS, HEMATOMAS, HERNIA RECURRENCE, AND POSTOPERATIVE CHRONIC PAIN. INTRAOPERATIVE BLEEDING DUE TACK FIXATION WAS RESOLVED INTRAOPERATIVELY. READMISSIONS WERE REQUIRED FOR ILEUS AND POSTOPERATIVE PAIN. NO ADDITIONAL INTERVENTIONS WERE REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED PERITONEAL AND SEROSAL TEARS AND INTRAOPERATIVE VESSEL INJURY. THERE WERE NO COMPLICATIONS RELATED TO THE SPACEMAKER DEVICE.
27-10-2024 | Injury | us surgical puerto rico | 22-01-2025 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Pain | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY OF RIGHT SUBCOSTAL INCISIONAL HERNIAS FROM 2014-2024 COMPARED OUTCOMES OF 46 TOTALLY ENDOSCOPIC RETRO MUSCULAR HERNIA REPAIRS WITH TRANSVERSUS ABDOMINIS RELEASE (ETEP GROUP) AND 51 CONVENTIONAL LAPAROSCOPIC INCISIONAL HERNIA REPAIRS WITH DEFECT CLOSURE (IPOM+ GROUP). IN THE IPOM+ GROUP, THE HERNIA DEFECT WAS CLOSED WITH A NONABSORBABLE VLOC SUTURE. SYMBOTEX COMPOSITE OR A COMPETITOR MESH WAS USED TO COVER THE CLOSED DEFECT AND NONABSORBABLE PROTACK WAS USED FOR MESH FIXATION. IN THE ETEP GROUP, A SPACEMAKER WAS USED TO RELEASE THE RETRO MUSCULAR SPACE AND SUBSEQUENTLY REMOVED. THE ANTERIOR AND PERITONEAL DEFECTS WERE CLOSED USING A VLOC SUTURE AND SYMBOTEX OR A COMPETITOR MESH WAS GLUED IN TO COVER THE RETRO MUSCULAR SPACE. POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING, ILEUS, SEROMAS, HEMATOMAS, HERNIA RECURRENCE, AND POSTOPERATIVE CHRONIC PAIN. INTRAOPERATIVE BLEEDING DUE TACK FIXATION WAS RESOLVED INTRAOPERATIVELY. READMISSIONS WERE REQUIRED FOR ILEUS AND POSTOPERATIVE PAIN. NO ADDITIONAL INTERVENTIONS WERE REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED PERITONEAL AND SEROSAL TEARS AND INTRAOPERATIVE VESSEL INJURY. THERE WERE NO COMPLICATIONS RELATED TO THE SPACEMAKER DEVICE.
27-10-2024 | Injury | us surgical puerto rico | 22-01-2025 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Seroma | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY OF RIGHT SUBCOSTAL INCISIONAL HERNIAS FROM 2014-2024 COMPARED OUTCOMES OF 46 TOTALLY ENDOSCOPIC RETRO MUSCULAR HERNIA REPAIRS WITH TRANSVERSUS ABDOMINIS RELEASE (ETEP GROUP) AND 51 CONVENTIONAL LAPAROSCOPIC INCISIONAL HERNIA REPAIRS WITH DEFECT CLOSURE (IPOM+ GROUP). IN THE IPOM+ GROUP, THE HERNIA DEFECT WAS CLOSED WITH A NONABSORBABLE VLOC SUTURE. SYMBOTEX COMPOSITE OR A COMPETITOR MESH WAS USED TO COVER THE CLOSED DEFECT AND NONABSORBABLE PROTACK WAS USED FOR MESH FIXATION. IN THE ETEP GROUP, A SPACEMAKER WAS USED TO RELEASE THE RETRO MUSCULAR SPACE AND SUBSEQUENTLY REMOVED. THE ANTERIOR AND PERITONEAL DEFECTS WERE CLOSED USING A VLOC SUTURE AND SYMBOTEX OR A COMPETITOR MESH WAS GLUED IN TO COVER THE RETRO MUSCULAR SPACE. POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING, ILEUS, SEROMAS, HEMATOMAS, HERNIA RECURRENCE, AND POSTOPERATIVE CHRONIC PAIN. INTRAOPERATIVE BLEEDING DUE TACK FIXATION WAS RESOLVED INTRAOPERATIVELY. READMISSIONS WERE REQUIRED FOR ILEUS AND POSTOPERATIVE PAIN. NO ADDITIONAL INTERVENTIONS WERE REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED PERITONEAL AND SEROSAL TEARS AND INTRAOPERATIVE VESSEL INJURY. THERE WERE NO COMPLICATIONS RELATED TO THE SPACEMAKER DEVICE.
27-10-2024 | Injury | us surgical puerto rico | 22-01-2025 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Hernia | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY OF RIGHT SUBCOSTAL INCISIONAL HERNIAS FROM 2014-2024 COMPARED OUTCOMES OF 46 TOTALLY ENDOSCOPIC RETRO MUSCULAR HERNIA REPAIRS WITH TRANSVERSUS ABDOMINIS RELEASE (ETEP GROUP) AND 51 CONVENTIONAL LAPAROSCOPIC INCISIONAL HERNIA REPAIRS WITH DEFECT CLOSURE (IPOM+ GROUP). IN THE IPOM+ GROUP, THE HERNIA DEFECT WAS CLOSED WITH A NONABSORBABLE VLOC SUTURE. SYMBOTEX COMPOSITE OR A COMPETITOR MESH WAS USED TO COVER THE CLOSED DEFECT AND NONABSORBABLE PROTACK WAS USED FOR MESH FIXATION. IN THE ETEP GROUP, A SPACEMAKER WAS USED TO RELEASE THE RETRO MUSCULAR SPACE AND SUBSEQUENTLY REMOVED. THE ANTERIOR AND PERITONEAL DEFECTS WERE CLOSED USING A VLOC SUTURE AND SYMBOTEX OR A COMPETITOR MESH WAS GLUED IN TO COVER THE RETRO MUSCULAR SPACE. POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING, ILEUS, SEROMAS, HEMATOMAS, HERNIA RECURRENCE, AND POSTOPERATIVE CHRONIC PAIN. INTRAOPERATIVE BLEEDING DUE TACK FIXATION WAS RESOLVED INTRAOPERATIVELY. READMISSIONS WERE REQUIRED FOR ILEUS AND POSTOPERATIVE PAIN. NO ADDITIONAL INTERVENTIONS WERE REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED PERITONEAL AND SEROSAL TEARS AND INTRAOPERATIVE VESSEL INJURY. THERE WERE NO COMPLICATIONS RELATED TO THE SPACEMAKER DEVICE.
27-10-2024 | Injury | us surgical puerto rico | 22-01-2025 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY OF RIGHT SUBCOSTAL INCISIONAL HERNIAS FROM 2014-2024 COMPARED OUTCOMES OF 46 TOTALLY ENDOSCOPIC RETRO MUSCULAR HERNIA REPAIRS WITH TRANSVERSUS ABDOMINIS RELEASE (ETEP GROUP) AND 51 CONVENTIONAL LAPAROSCOPIC INCISIONAL HERNIA REPAIRS WITH DEFECT CLOSURE (IPOM+ GROUP). IN THE IPOM+ GROUP, THE HERNIA DEFECT WAS CLOSED WITH A NONABSORBABLE VLOC SUTURE. SYMBOTEX COMPOSITE OR A COMPETITOR MESH WAS USED TO COVER THE CLOSED DEFECT AND NONABSORBABLE PROTACK WAS USED FOR MESH FIXATION. IN THE ETEP GROUP, A SPACEMAKER WAS USED TO RELEASE THE RETRO MUSCULAR SPACE AND SUBSEQUENTLY REMOVED. THE ANTERIOR AND PERITONEAL DEFECTS WERE CLOSED USING A VLOC SUTURE AND SYMBOTEX OR A COMPETITOR MESH WAS GLUED IN TO COVER THE RETRO MUSCULAR SPACE. POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING, ILEUS, SEROMAS, HEMATOMAS, HERNIA RECURRENCE, AND POSTOPERATIVE CHRONIC PAIN. INTRAOPERATIVE BLEEDING DUE TACK FIXATION WAS RESOLVED INTRAOPERATIVELY. READMISSIONS WERE REQUIRED FOR ILEUS AND POSTOPERATIVE PAIN. NO ADDITIONAL INTERVENTIONS WERE REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED PERITONEAL AND SEROSAL TEARS AND INTRAOPERATIVE VESSEL INJURY. THERE WERE NO COMPLICATIONS RELATED TO THE SPACEMAKER DEVICE.
27-10-2024 | Injury | us surgical puerto rico | 22-01-2025 | UNKNOWN PROTACK | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY OF RIGHT SUBCOSTAL INCISIONAL HERNIAS FROM 2014-2024 COMPARED OUTCOMES OF 46 TOTALLY ENDOSCOPIC RETRO MUSCULAR HERNIA REPAIRS WITH TRANSVERSUS ABDOMINIS RELEASE (ETEP GROUP) AND 51 CONVENTIONAL LAPAROSCOPIC INCISIONAL HERNIA REPAIRS WITH DEFECT CLOSURE (IPOM+ GROUP). IN THE IPOM+ GROUP, THE HERNIA DEFECT WAS CLOSED WITH A NONABSORBABLE VLOC SUTURE. SYMBOTEX COMPOSITE OR A COMPETITOR MESH WAS USED TO COVER THE CLOSED DEFECT AND NONABSORBABLE PROTACK WAS USED FOR MESH FIXATION. IN THE ETEP GROUP, A SPACEMAKER WAS USED TO RELEASE THE RETRO MUSCULAR SPACE AND SUBSEQUENTLY REMOVED. THE ANTERIOR AND PERITONEAL DEFECTS WERE CLOSED USING A VLOC SUTURE AND SYMBOTEX OR A COMPETITOR MESH WAS GLUED IN TO COVER THE RETRO MUSCULAR SPACE. POSTOPERATIVE COMPLICATIONS INCLUDED BLEEDING, ILEUS, SEROMAS, HEMATOMAS, HERNIA RECURRENCE, AND POSTOPERATIVE CHRONIC PAIN. INTRAOPERATIVE BLEEDING DUE TACK FIXATION WAS RESOLVED INTRAOPERATIVELY. READMISSIONS WERE REQUIRED FOR ILEUS AND POSTOPERATIVE PAIN. NO ADDITIONAL INTERVENTIONS WERE REPORTED. OTHER COMPLICATIONS NOT RELATED TO THE DEVICE INCLUDED PERITONEAL AND SEROSAL TEARS AND INTRAOPERATIVE VESSEL INJURY. THERE WERE NO COMPLICATIONS RELATED TO THE SPACEMAKER DEVICE.
31-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING TWO EGIA60AMT EGIA 60 ARTIC MED THICK SULU, THE STAPLE LINE DID NOT HOLD OR OPENED UP AFTER FIRING.
31-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING TWO EGIA60AMT EGIA 60 ARTIC MED THICK SULU, THE STAPLE LINE DID NOT HOLD OR OPENED UP AFTER FIRING.
07-01-2025 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, THERE WAS A CLICK SOUND WHEN THE HANDLE WAS HALFWAY FIRED. WHEN THE LEVER WAS PRESSED, AND THERE WAS NO RESPONSE, THE STAPLES WERE STUCK AND FORMED INCOMPLETELY. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | us surgical puerto rico | 22-01-2025 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, THERE WAS A CLICK SOUND WHEN THE HANDLE WAS HALFWAY FIRED. WHEN THE LEVER WAS PRESSED, AND THERE WAS NO RESPONSE, THE STAPLES WERE STUCK AND FORMED INCOMPLETELY. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, THERE WAS A CLICK SOUND WHEN THE HANDLE WAS HALFWAY FIRED. WHEN THE LEVER WAS PRESSED, AND THERE WAS NO RESPONSE, THE STAPLES WERE STUCK AND FORMED INCOMPLETELY. A NEW HANDLE AND RELOAD WERE USED TO RESOLVE THE ISSUE AND COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
21-08-2024 | Malfunction | covidien | 22-01-2025 | ENDO GIA |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: STAPLER RELOAD HAS A YELLOW REMOVABLE TAB USED FOR RELOADING. AFTER FIRING, A YELLOW PIECE OF PLASTIC FELL OFF THE STAPLER ONTO THE STERILE FIELD. THE PIECE HAD BROKEN OFF THE REMOVABLE TAB FROM THE RELOAD.
01-01-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE DID NOT WORK. THERE WAS NO PATIENT CONSEQUENCES REPORTED. SURGICAL DELAY UNKNOWN.
12-08-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-08-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-08-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
25-12-2024 | Injury | covidien lp llc north haven | 22-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Organ Dehiscence | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE SINGLE USE RELOADABLE STAPLER AND THE SINGLE USE LOADING UNIT, THE STAPLE LINE DID NOT HOLD OR OPENED UP. THIS NECESSITATED RESECTION AND RE-STAPLING TWICE. THE PATIENT REQUIRED EXTENDED HOSPITALIZATION OF AN UNSPECIFIED DURATION AND WAS ADMITTED TO THE INTENSIVE CARE UNIT.
25-12-2024 | Injury | covidien lp llc north haven | 22-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE SINGLE USE RELOADABLE STAPLER AND THE SINGLE USE LOADING UNIT, THE STAPLE LINE DID NOT HOLD OR OPENED UP. THIS NECESSITATED RESECTION AND RE-STAPLING TWICE. THE PATIENT REQUIRED EXTENDED HOSPITALIZATION OF AN UNSPECIFIED DURATION AND WAS ADMITTED TO THE INTENSIVE CARE UNIT.
25-12-2024 | Injury | covidien lp llc north haven | 22-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Organ Dehiscence | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE SINGLE USE RELOADABLE STAPLER AND THE SINGLE USE LOADING UNIT, THE STAPLE LINE DID NOT HOLD OR OPENED UP. THIS NECESSITATED RESECTION AND RE-STAPLING TWICE. THE PATIENT REQUIRED EXTENDED HOSPITALIZATION OF AN UNSPECIFIED DURATION AND WAS ADMITTED TO THE INTENSIVE CARE UNIT.
25-12-2024 | Injury | covidien lp llc north haven | 22-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE SINGLE USE RELOADABLE STAPLER AND THE SINGLE USE LOADING UNIT, THE STAPLE LINE DID NOT HOLD OR OPENED UP. THIS NECESSITATED RESECTION AND RE-STAPLING TWICE. THE PATIENT REQUIRED EXTENDED HOSPITALIZATION OF AN UNSPECIFIED DURATION AND WAS ADMITTED TO THE INTENSIVE CARE UNIT.
19-12-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
11-12-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A THORACOSCOPY PNEUMONECTOMY, THE DEVICE WAS LOCKED OUT AT THE 1ST FIRING WHEN PULMONARY VEIN DISSECTION FOR VATS. THE DEVICE COULD NOT BE FIRED AND THE KNIFE DID NOT MOVE FORWARD. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE, AND A NEW DEVICE WAS REPLACED TO HANDLE. THERE WAS NO BLEEDING. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
11-12-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING A THORACOSCOPY PNEUMONECTOMY, THE DEVICE WAS LOCKED OUT AT THE 1ST FIRING WHEN PULMONARY VEIN DISSECTION FOR VATS. THE DEVICE COULD NOT BE FIRED AND THE KNIFE DID NOT MOVE FORWARD. THE MANUAL OVERRIDE LEVER WAS USED TO RELEASE THE DEVICE, AND A NEW DEVICE WAS REPLACED TO HANDLE. THERE WAS NO BLEEDING. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
10-12-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | WHITE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-12-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE FIRING. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL COULD NOT FIRE AND WITHOUT MOTOR SOUND. CHANGED ANOTHER ONE TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
01-01-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | ECHELON 3000 60MM LONG | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING AN UNKNOWN PROCEDURE THAT THE DEVICE WAS NOT ARTICULATING. NO PATIENT CONSEQUENCE WAS REPORTED.
23-11-2024 | Malfunction | ethicon endo-surgery | 22-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE 1ST FIRING. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL COULD NOT FIRE AND WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
08-01-2025 | Malfunction | davol inc sub c.r bard inc -1213643 | 22-01-2025 | OPTIFIX AT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A PROCEDURE, THE OPTIFIX AT FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE FASTENER DID NOT FIXATE THE MESH AND THERE WAS TOO MUCH CONCERN REGARDING A POTENTIAL PATIENT SAFETY ISSUE. THERE WAS NO PATIENT INJURY. THIS FILE REPRESENTS DEVICE #1.
17-12-2024 | Malfunction | us surgical puerto rico | 22-01-2025 | UNKNOWN ENDO GI | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHILE SETTING UP, THE GREEN SAFETY BUTTON ON THE HANDLE COULD NOT BE PRESSED OR DID NOT WORK. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE. THERE WAS NO PATIENT INVOLVEMENT.
17-12-2024 | Malfunction | covidien lp llc north haven | 22-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHILE SETTING UP, THE GREEN SAFETY BUTTON ON THE HANDLE COULD NOT BE PRESSED OR DID NOT WORK. THE ISSUE WAS RESOLVED BY REPLACING THE DEVICE. THERE WAS NO PATIENT INVOLVEMENT.
07-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE HAD PARTIAL FIRING. IT WAS NOTED THAT THE SURGEON FIRED THE RELOAD BUT IT ONLY STAPLED HALF OF THE STAPLE LINE. THE SURGEON CONTINUED STAPLING THE SAME STAPLE LINE WITH ANOTHER RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien lp llc north haven | 22-01-2025 | SIGNIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE HAD PARTIAL FIRING. IT WAS NOTED THAT THE SURGEON FIRED THE RELOAD BUT IT ONLY STAPLED HALF OF THE STAPLE LINE. THE SURGEON CONTINUED STAPLING THE SAME STAPLE LINE WITH ANOTHER RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien | 22-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE HAD PARTIAL FIRING. IT WAS NOTED THAT THE SURGEON FIRED THE RELOAD BUT IT ONLY STAPLED HALF OF THE STAPLE LINE. THE SURGEON CONTINUED STAPLING THE SAME STAPLE LINE WITH ANOTHER RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-01-2025 | Malfunction | davol inc sub c.r bard inc -1213643 | 22-01-2025 | OPTIFIX AT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A PROCEDURE, THE OPTIFIX AT FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE FASTENER DID NOT FIXATE THE MESH AND THERE WAS TOO MUCH CONCERN REGARDING A POTENTIAL PATIENT SAFETY ISSUE. THERE WAS NO PATIENT INJURY. THIS FILE REPRESENTS DEVICE # 3.
08-01-2025 | Malfunction | davol inc sub c.r bard inc -1213643 | 22-01-2025 | OPTIFIX AT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A PROCEDURE, THE OPTIFIX AT FIXATION DEVICE WAS USED TO FIXATE THE MESH. IT WAS REPORTED THAT THE FASTENER DID NOT FIXATE THE MESH AND THERE WAS TOO MUCH CONCERN REGARDING A POTENTIAL PATIENT SAFETY ISSUE. THERE WAS NO PATIENT INJURY. THIS FILE REPRESENTS DEVICE #2.
09-01-2025 | Malfunction | us surgical puerto rico | 23-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A PARTIAL PULMONARY RESECTION, DURING THE FIRING OF THE SECOND SHOT, THE CARTRIDGE HAD A LARGE TIP OPENING WHEN CHECKING THE OPENING AND CLOSING AT THE TIME OF CONNECTION. EXCESSIVE FORCE INDICATION APPEARED WHEN THE CARTRIDGE WAS SET IN THE TISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED.
09-01-2025 | Malfunction | us surgical puerto rico | 23-01-2025 | ENDO GIA | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A PARTIAL PULMONARY RESECTION, DURING THE FIRING OF THE SECOND SHOT, THE CARTRIDGE HAD A LARGE TIP OPENING WHEN CHECKING THE OPENING AND CLOSING AT THE TIME OF CONNECTION. EXCESSIVE FORCE INDICATION APPEARED WHEN THE CARTRIDGE WAS SET IN THE TISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED.
09-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A PARTIAL PULMONARY RESECTION, DURING THE FIRING OF THE SECOND SHOT, THE CARTRIDGE HAD A LARGE TIP OPENING WHEN CHECKING THE OPENING AND CLOSING AT THE TIME OF CONNECTION. EXCESSIVE FORCE INDICATION APPEARED WHEN THE CARTRIDGE WAS SET IN THE TISSUE. THERE WAS NO PATIENT INJURY.
09-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A PARTIAL PULMONARY RESECTION, DURING THE FIRING OF THE SECOND SHOT, THE CARTRIDGE HAD A LARGE TIP OPENING WHEN CHECKING THE OPENING AND CLOSING AT THE TIME OF CONNECTION. EXCESSIVE FORCE INDICATION APPEARED WHEN THE CARTRIDGE WAS SET IN THE TISSUE. THERE WAS NO PATIENT INJURY.
09-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A PARTIAL PULMONARY RESECTION, DURING THE FIRING OF THE SECOND SHOT, THE CARTRIDGE HAD A LARGE TIP OPENING WHEN CHECKING THE OPENING AND CLOSING AT THE TIME OF CONNECTION. EXCESSIVE FORCE INDICATION APPEARED WHEN THE CARTRIDGE WAS SET IN THE TISSUE. THERE WAS NO PATIENT INJURY.
09-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A PARTIAL PULMONARY RESECTION, DURING THE FIRING OF THE SECOND SHOT, THE CARTRIDGE HAD A LARGE TIP OPENING WHEN CHECKING THE OPENING AND CLOSING AT THE TIME OF CONNECTION. EXCESSIVE FORCE INDICATION APPEARED WHEN THE CARTRIDGE WAS SET IN THE TISSUE. THERE WAS NO PATIENT INJURY.
09-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | ENDO GIA | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A PARTIAL PULMONARY RESECTION, DURING THE FIRING OF THE SECOND SHOT, THE CARTRIDGE HAD A LARGE TIP OPENING WHEN CHECKING THE OPENING AND CLOSING AT THE TIME OF CONNECTION. EXCESSIVE FORCE INDICATION APPEARED WHEN THE CARTRIDGE WAS SET IN THE TISSUE. THERE WAS NO PATIENT INJURY.
09-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY OF A PARTIAL PULMONARY RESECTION, DURING THE FIRING OF THE SECOND SHOT, THE CARTRIDGE HAD A LARGE TIP OPENING WHEN CHECKING THE OPENING AND CLOSING AT THE TIME OF CONNECTION. EXCESSIVE FORCE INDICATION APPEARED WHEN THE CARTRIDGE WAS SET IN THE TISSUE. THERE WAS NO PATIENT INJURY.
01-12-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE DEVICE WOULD ONLY STAPLE HALF THE LENGTH OF THE STAPLER. THEY FIRED TWICE AND IT DID NOT EVEN GET A DECENT LENGTH ANASTOMOSIS. THIS WAS WITH A BLUE RELOAD. CASE WAS COMPLETED WITH A LIKE DEVICE. NO PATIENT HARM WAS REPORTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 23-01-2025 | CIRCULAR MECH STD 29MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE COULDN'T FIRE.
16-08-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | CONTOUR CURVED CUTTER STAPLER | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION AN ATTEMPT IS MADE TO SECTION THE COLORECTAL STUMP WITH A CONTOUR STAPLER, TWO SHOTS, BUT IT DOES NOT WORK WELL DUE TO STAPLE DYSFUNCTION. THE SURGERY WAS DELAYED 20 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
16-08-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | CONTOUR CURVED CUTTER STAPLER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION AN ATTEMPT IS MADE TO SECTION THE COLORECTAL STUMP WITH A CONTOUR STAPLER, TWO SHOTS, BUT IT DOES NOT WORK WELL DUE TO STAPLE DYSFUNCTION. THE SURGERY WAS DELAYED 20 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
04-10-2024 | Malfunction | covidien lp llc north haven | 23-01-2025 | ENDO GIA | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY (VATS) WEDGE RESECTION PROCEDURE, THERE WAS LOADING DIFFICULTY WITH ONE OF THE PURPLE RELOADS. WHILE STAPLING THE LUNG, USING ANOTHER PURPLE RELOAD, IT FIRED PAST THE STAPLE LINE, WHICH CAUSED SLIGHT BLEEDING IN THE TISSUE. THIS RELOAD DID NOT FULLY OPEN EITHER. ADDITIONALLY, A REINFORCED BLACK RELOAD HAD LOADING DIFFICULTY AND COULD NOT BE REMOVED FROM THE ADAPTER. TO RESOLVE THE ISSUE, A MANUAL HANDLE AND RELOAD WERE USED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA RELOAD THICK TISSUE. THE LOADING UNIT HAD DAMAGE TO THE CUTTING EDGE OF THE KNIFE BLADE. THE CLAMPING MECHANISM WAS DEFORMED. THE ANVIL WAS BOWED.
04-10-2024 | Malfunction | covidien lp llc north haven | 23-01-2025 | ENDO GIA | Component or Accessory Incompatibility | A1701 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC VIDEO ASSISTED THORACOSCOPIC SURGERY (VATS) WEDGE RESECTION PROCEDURE, THERE WAS LOADING DIFFICULTY WITH ONE OF THE PURPLE RELOADS. WHILE STAPLING THE LUNG, USING ANOTHER PURPLE RELOAD, IT FIRED PAST THE STAPLE LINE, WHICH CAUSED SLIGHT BLEEDING IN THE TISSUE. THIS RELOAD DID NOT FULLY OPEN EITHER. ADDITIONALLY, A REINFORCED BLACK RELOAD HAD LOADING DIFFICULTY AND COULD NOT BE REMOVED FROM THE ADAPTER. TO RESOLVE THE ISSUE, A MANUAL HANDLE AND RELOAD WERE USED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA RELOAD THICK TISSUE. THE LOADING UNIT HAD DAMAGE TO THE CUTTING EDGE OF THE KNIFE BLADE. THE CLAMPING MECHANISM WAS DEFORMED. THE ANVIL WAS BOWED.
07-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING THE THORACOSCOPIC LUNG RESECTION, THE USER WAS ABLE TO PRESS THE GREEN BUTTON BUT FIRING COULD NOT BE PERFORMED AT ALL. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-10-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY SURGERY, COULD NOT FIRE FARTHER AFTER FIRED A HALF. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | GST GREEN RELOAD, 45MM, 6-ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
16-12-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRO-INTESTINAL SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END BUT COULD NOT RETURN KNIFE AUTOMATICALLY AND COULD NOT OPEN THE JAW. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE, OPEN THE JAW AND REMOVED THE DEVICE. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END BUT COULD NOT RETURN KNIFE AUTOMATICALLY AND COULD NOT OPEN THE JAW. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE, OPEN THE JAW AND REMOVED THE DEVICE. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | POWERED PLUS STANDARD 45 | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END BUT COULD NOT RETURN KNIFE AUTOMATICALLY AND COULD NOT OPEN THE JAW. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE, OPEN THE JAW AND REMOVED THE DEVICE. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
26-12-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | POWERED PLUS STANDARD 45 | Activation Failure | A150101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION AFTER FIRED, THE KNIFE MOVED TO THE DISTAL END BUT COULD NOT RETURN KNIFE AUTOMATICALLY AND COULD NOT OPEN THE JAW. KNIFE REVERSE BUTTON COULD NOT WORK. USED MANUAL OVERRIDE LEVER TO RETURN KNIFE, OPEN THE JAW AND REMOVED THE DEVICE. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2024 | Malfunction | ethicon endo-surgery | 23-01-2025 | WHITE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP HEMICOLECTOMY SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-01-2025 | Injury | ethicon endo-surgery | 23-01-2025 | ECHELON 3000 45MM COMPACT | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN OPEN RADICAL NEPHRECTOMY PROCEDURE THAT THE STAPLE LINE WAS MALFORMED RESULTING IN BLOOD LOSS. PATIENT WAS GIVEN AN UNKNOWN NUMBER OF UNITS OF BLOOD. A VASCULAR SURGEON WAS CALLED IN FOR VESSEL REPAIR. UNKNOWN IF THIS REPAIR WAS ON THE STAPLE LINE. PATIENT STATUS IS CURRENTLY UNKNOWN.
08-01-2025 | Injury | ethicon endo-surgery | 23-01-2025 | ECHELON 3000 45MM COMPACT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN OPEN RADICAL NEPHRECTOMY PROCEDURE THAT THE STAPLE LINE WAS MALFORMED RESULTING IN BLOOD LOSS. PATIENT WAS GIVEN AN UNKNOWN NUMBER OF UNITS OF BLOOD. A VASCULAR SURGEON WAS CALLED IN FOR VESSEL REPAIR. UNKNOWN IF THIS REPAIR WAS ON THE STAPLE LINE. PATIENT STATUS IS CURRENTLY UNKNOWN.
13-01-2025 | Injury | ethicon endo-surgery | 23-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A HYBRID ROBOTIC PROCEDURE NEEDED TO BE CONVERTED TO OPEN. THE DEVICE WAS FIRED ON THICK TISSUE ACCORDING TO SURGEON THEN JAMMED AND THE JAWS WOULD NOT OPEN. WHILST TRYING TO OPEN THE JAWS THE SURGEON BROKE THE FIRING TRIGGER. SURGEON HAD TO CONVERT TO OPEN PROCEDURE TO REMOVE THE DEVICE FROM THE PATIENT. PATIENTS PROCEDURE WAS CONVERTED FROM ROBOTIC TO OPEN IN ORDER TO REMOVE THE DEVICE.THE SURGERY WAS DELAYED AN UNKNOWN NUMBER OF MINUTES.
13-01-2025 | Injury | ethicon endo-surgery | 23-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Firing Problem | A0505 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A HYBRID ROBOTIC PROCEDURE NEEDED TO BE CONVERTED TO OPEN. THE DEVICE WAS FIRED ON THICK TISSUE ACCORDING TO SURGEON THEN JAMMED AND THE JAWS WOULD NOT OPEN. WHILST TRYING TO OPEN THE JAWS THE SURGEON BROKE THE FIRING TRIGGER. SURGEON HAD TO CONVERT TO OPEN PROCEDURE TO REMOVE THE DEVICE FROM THE PATIENT. PATIENTS PROCEDURE WAS CONVERTED FROM ROBOTIC TO OPEN IN ORDER TO REMOVE THE DEVICE.THE SURGERY WAS DELAYED AN UNKNOWN NUMBER OF MINUTES.
21-11-2024 | Injury | ethicon endo-surgery | 23-01-2025 | PWRD ECH FLEX 45MM 34CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS THE DOCTOR REPORTED THAT BLEEDING WAS IDENTIFIED IN THE POSTOPERATIVE PERIOD AT THE STAPLING LINE. PROCEDURE EVIDENCED BY TOMOGRAPHY EXAM, AND BY A LAPAROSCOPY, IN WHICH INFECTED BLOOD CLOTS WERE IDENTIFIED. BLEEDING WAS LOCATED IN THE STAPLING. DUE TO THIS EVENT, THE PATIENT NEEDED TO STAY ADDITIONAL DAYS OF HOSPITALIZATION, TOTALING 30 DAYS MORE THAN ICU. HOWEVER, HE WAS DISCHARGED WITHOUT OTHER COMPLICATIONS, AND THE PATIENT IS CURRENTLY DOING WELL.
21-11-2024 | Injury | ethicon endo-surgery | 23-01-2025 | PWRD ECH FLEX 45MM 34CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS THE DOCTOR REPORTED THAT BLEEDING WAS IDENTIFIED IN THE POSTOPERATIVE PERIOD AT THE STAPLING LINE. PROCEDURE EVIDENCED BY TOMOGRAPHY EXAM, AND BY A LAPAROSCOPY, IN WHICH INFECTED BLOOD CLOTS WERE IDENTIFIED. BLEEDING WAS LOCATED IN THE STAPLING. DUE TO THIS EVENT, THE PATIENT NEEDED TO STAY ADDITIONAL DAYS OF HOSPITALIZATION, TOTALING 30 DAYS MORE THAN ICU. HOWEVER, HE WAS DISCHARGED WITHOUT OTHER COMPLICATIONS, AND THE PATIENT IS CURRENTLY DOING WELL.
16-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | TRI-STAPLE 2.0 | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE RELOAD USED WITH VENTRICULUS USING THE RELOAD WITH A POWERED STAPLER AND SHORT ADAPTER, THERE WERE ISSUES AFTER FIRING. SPECIFICALLY, IT WAS DIFFICULT TO RETRACT THE KNIFE BLADE, AND THE FAULTY RELOAD HAD TO BE CUT FREE AS THE KNIFE WOULD NOT RETRACT. THE DEVICE WAS REPLACED WITH A NEW RELOAD. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT, AND THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
16-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | TRI-STAPLE 2.0 | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE RELOAD USED WITH VENTRICULUS USING THE RELOAD WITH A POWERED STAPLER AND SHORT ADAPTER, THERE WERE ISSUES AFTER FIRING. SPECIFICALLY, IT WAS DIFFICULT TO RETRACT THE KNIFE BLADE, AND THE FAULTY RELOAD HAD TO BE CUT FREE AS THE KNIFE WOULD NOT RETRACT. THE DEVICE WAS REPLACED WITH A NEW RELOAD. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT, AND THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
16-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE RELOAD USED WITH VENTRICULUS USING THE RELOAD WITH A POWERED STAPLER AND SHORT ADAPTER, THERE WERE ISSUES AFTER FIRING. SPECIFICALLY, IT WAS DIFFICULT TO RETRACT THE KNIFE BLADE, AND THE FAULTY RELOAD HAD TO BE CUT FREE AS THE KNIFE WOULD NOT RETRACT. THE DEVICE WAS REPLACED WITH A NEW RELOAD. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT, AND THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
16-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE RELOAD USED WITH VENTRICULUS USING THE RELOAD WITH A POWERED STAPLER AND SHORT ADAPTER, THERE WERE ISSUES AFTER FIRING. SPECIFICALLY, IT WAS DIFFICULT TO RETRACT THE KNIFE BLADE, AND THE FAULTY RELOAD HAD TO BE CUT FREE AS THE KNIFE WOULD NOT RETRACT. THE DEVICE WAS REPLACED WITH A NEW RELOAD. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT, AND THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
01-01-2025 | Injury | ethicon endo-surgery | 23-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UPPER LEFT VATS LOBECTOMY ON THE THIRD FIRING OF THE PVE35A NO STAPLES WERE VISIBLE ON THE PULMONARY FIRING. THE SUBSEQUENT BLEEDING REQUIRED CONVERSION TO PNEUMONECTOMY. PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A DELAY GREATER THAN 20 MINUTES.
01-01-2025 | Injury | ethicon endo-surgery | 23-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING AN UPPER LEFT VATS LOBECTOMY ON THE THIRD FIRING OF THE PVE35A NO STAPLES WERE VISIBLE ON THE PULMONARY FIRING. THE SUBSEQUENT BLEEDING REQUIRED CONVERSION TO PNEUMONECTOMY. PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A DELAY GREATER THAN 20 MINUTES.
16-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Entrapment of Device | A150208 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE RELOAD USED WITH VENTRICULUS USING THE RELOAD WITH A POWERED STAPLER AND SHORT ADAPTER, THERE WERE ISSUES AFTER FIRING. SPECIFICALLY, IT WAS DIFFICULT TO RETRACT THE KNIFE BLADE, AND THE FAULTY RELOAD HAD TO BE CUT FREE AS THE KNIFE WOULD NOT RETRACT. THE DEVICE WAS REPLACED WITH A NEW RELOAD. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT, AND THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
16-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Retraction Problem | A0510 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE RELOAD USED WITH VENTRICULUS USING THE RELOAD WITH A POWERED STAPLER AND SHORT ADAPTER, THERE WERE ISSUES AFTER FIRING. SPECIFICALLY, IT WAS DIFFICULT TO RETRACT THE KNIFE BLADE, AND THE FAULTY RELOAD HAD TO BE CUT FREE AS THE KNIFE WOULD NOT RETRACT. THE DEVICE WAS REPLACED WITH A NEW RELOAD. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT, AND THE PATIENT OUTCOME WAS REPORTED AS ALIVE WITH NO INJURY.
01-01-2025 | Injury | ethicon endo-surgery | 23-01-2025 | ECHELON 45 (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: A CASE OF GASTROINTESTINAL STROMAL TUMOR (GIST) SUBCARDINAL TREATED BY SINGLE-PORT LAPAROSCOPIC ENDOGASTRIC SURGERY. AUTHORS: HIDENORI OSHIO, TAIJI TOKUYAMA (NAOKI OKUMURA, NITTO TAKANO, MEGUMI SUGIYAMA, KAZUHIRO ISHIHARA. CITATION: THE JOURNAL OF PRACTICAL MEDICINE, VOL. 72 5, PAGE 408 ~ 414 2024. 1. THIS STUDY REPORTS A CASE OF SINGLE-INCISION LAPAROSCOPIC ENDOGASTRIC SURGERY FOR GASTROINTESTINAL STROMAL TUMOR (HEREINAFTER REFERRED TO AS GIST) BENEATH THE CARDIA. A FEMALE IN HER 70S WHO HAD BEEN VISITING THE DEPARTMENT OF CARDIOVASCULAR INTERNAL MEDICINE FOR HYPERTENSION AND ATRIAL FIBRILLATION. THE PATIENT BECAME AWARE OF ABDOMINAL PAIN AND UNDERWENT CT, WHICH REVEALED A 3 CM SIZED MASS IN THE STOMACH. PATIENT UNDERWENT SINGLE-PORT LAPAROSCOPIC ENDOGASTRIC SURGERY. AN AUTOMATIC SUTURE DEVICE (ECHELON R BLUE45MM) WAS USED TO ENSURE TO PRESERVE THE CARDIA. FOLLOW-UP IS UNKNOWN. REPORTED COMPLICATIONS INCLUDED A 70 YEAR-OLD FEMALE WITH HEMATEMESIS AND WORSENING FROM THE STAPLE LINE AT THE TUMOR RESECTION SITE IN CONCLUSION, FOR GIST JUST BELOW THE GASTRIC CARDIA, GIST, AND LAPAROSCOPIC INTRAGASTRIC SURGERY WAS PERFORMED. LAPAROSCOPIC ENDOGASTRIC SURGERY WAS CONSIDERED TO BE A USEFUL SURGICAL PROCEDURE BECAUSE IT CAN PRESERVE THE CARDIAC FUNCTION WITHOUT FAIL, AND FROM THE VIEWPOINT OF AESTHETICS AND LOW INVASIVENESS.
01-01-2025 | Injury | ethicon endo-surgery | 23-01-2025 | ECHELON 45 (EXACT CODE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | A24 | Insufficient Information | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: A CASE OF GASTROINTESTINAL STROMAL TUMOR (GIST) SUBCARDINAL TREATED BY SINGLE-PORT LAPAROSCOPIC ENDOGASTRIC SURGERY. AUTHORS: HIDENORI OSHIO, TAIJI TOKUYAMA (NAOKI OKUMURA, NITTO TAKANO, MEGUMI SUGIYAMA, KAZUHIRO ISHIHARA. CITATION: THE JOURNAL OF PRACTICAL MEDICINE, VOL. 72 5, PAGE 408 ~ 414 2024. 1. THIS STUDY REPORTS A CASE OF SINGLE-INCISION LAPAROSCOPIC ENDOGASTRIC SURGERY FOR GASTROINTESTINAL STROMAL TUMOR (HEREINAFTER REFERRED TO AS GIST) BENEATH THE CARDIA. A FEMALE IN HER 70S WHO HAD BEEN VISITING THE DEPARTMENT OF CARDIOVASCULAR INTERNAL MEDICINE FOR HYPERTENSION AND ATRIAL FIBRILLATION. THE PATIENT BECAME AWARE OF ABDOMINAL PAIN AND UNDERWENT CT, WHICH REVEALED A 3 CM SIZED MASS IN THE STOMACH. PATIENT UNDERWENT SINGLE-PORT LAPAROSCOPIC ENDOGASTRIC SURGERY. AN AUTOMATIC SUTURE DEVICE (ECHELON R BLUE45MM) WAS USED TO ENSURE TO PRESERVE THE CARDIA. FOLLOW-UP IS UNKNOWN. REPORTED COMPLICATIONS INCLUDED A 70 YEAR-OLD FEMALE WITH HEMATEMESIS AND WORSENING FROM THE STAPLE LINE AT THE TUMOR RESECTION SITE IN CONCLUSION, FOR GIST JUST BELOW THE GASTRIC CARDIA, GIST, AND LAPAROSCOPIC INTRAGASTRIC SURGERY WAS PERFORMED. LAPAROSCOPIC ENDOGASTRIC SURGERY WAS CONSIDERED TO BE A USEFUL SURGICAL PROCEDURE BECAUSE IT CAN PRESERVE THE CARDIAC FUNCTION WITHOUT FAIL, AND FROM THE VIEWPOINT OF AESTHETICS AND LOW INVASIVENESS.
21-11-2024 | Injury | ethicon endo-surgery | 23-01-2025 | PWRD ECH FLEX 45MM 34CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS THE DOCTOR REPORTED THAT BLEEDING WAS IDENTIFIED IN THE POSTOPERATIVE PERIOD AT THE STAPLING LINE. PROCEDURE EVIDENCED BY TOMOGRAPHY EXAM, AND BY A LAPAROSCOPY, IN WHICH INFECTED BLOOD CLOTS WERE IDENTIFIED. BLEEDING WAS LOCATED IN THE STAPLING. DUE TO THIS EVENT, THE PATIENT NEEDED TO STAY ADDITIONAL DAYS OF HOSPITALIZATION, TOTALING 30 DAYS MORE THAN ICU. HOWEVER, HE WAS DISCHARGED WITHOUT OTHER COMPLICATIONS, AND THE PATIENT IS CURRENTLY DOING WELL.
21-11-2024 | Injury | ethicon endo-surgery | 23-01-2025 | PWRD ECH FLEX 45MM 34CM SHAFT | Adverse Event Without Identified Device or Use Problem | A24 | Unspecified Infection | Event Description: IT WAS REPORTED THAT DURING A GASTRIC BYPASS THE DOCTOR REPORTED THAT BLEEDING WAS IDENTIFIED IN THE POSTOPERATIVE PERIOD AT THE STAPLING LINE. PROCEDURE EVIDENCED BY TOMOGRAPHY EXAM, AND BY A LAPAROSCOPY, IN WHICH INFECTED BLOOD CLOTS WERE IDENTIFIED. BLEEDING WAS LOCATED IN THE STAPLING. DUE TO THIS EVENT, THE PATIENT NEEDED TO STAY ADDITIONAL DAYS OF HOSPITALIZATION, TOTALING 30 DAYS MORE THAN ICU. HOWEVER, HE WAS DISCHARGED WITHOUT OTHER COMPLICATIONS, AND THE PATIENT IS CURRENTLY DOING WELL.
07-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER RESECTION, THE INSTRUMENT PARTIALLY FIRED AND STAPLES DID NOT DEPLOY FULLY. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | us surgical puerto rico | 23-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION (LAP LAR) PROCEDURE INVOLVING THE UPPER COLON, WHEN THE JAW GOT THE TISSUE, THE JAWS WOULD NOT CLOSE COMPLETELY. TO RESOLVE THE PROBLEM, THE SURGEON REPLACED THE DEVICE WITH ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
07-10-2024 | Malfunction | us surgical puerto rico | 23-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING THE SEGMENTECTOMY, THE HANDLE/TRIGGER OF THE STAPLER WAS BROKEN. THE HANDLE WAS REPLACED WITH ANOTHER PRODUCT FROM THE SAME COMPANY. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS HANDLE THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
07-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION (LAP LAR) PROCEDURE INVOLVING THE UPPER COLON, WHEN THE JAW GOT THE TISSUE, THE JAWS WOULD NOT CLOSE COMPLETELY. TO RESOLVE THE PROBLEM, THE SURGEON REPLACED THE DEVICE WITH ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien lp llc north haven | 23-01-2025 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION (LAP LAR) PROCEDURE INVOLVING THE UPPER COLON, WHEN THE JAW GOT THE TISSUE, THE JAWS WOULD NOT CLOSE COMPLETELY. TO RESOLVE THE PROBLEM, THE SURGEON REPLACED THE DEVICE WITH ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
26-12-2024 | Death | covidien lp llc north haven | 23-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY WITH LATEROLATERAL ANASTOMOSIS PROCEDURE, THE DEVICE PERFORMED AS USUAL. POSTOPERATIVELY, THE STAPLE LINE DID NOT HOLD OR OPENED UP (DEHISCENCE). THE PATIENT EXPERIENCED EXTENDED HOSPITALIZATION FOR 8 DAYS AND WAS ADMITTED TO THE INTENSIVE CARE UNIT. PLANS TO RE-OPERATE ON THE PATIENT WERE MADE, BUT IT WAS NOT POSSIBLE. THE PATIENT OUTCOME WAS DEATH, WITH THE CAUSE OF DEATH CURRENTLY UNKNOWN.
26-12-2024 | Death | covidien lp llc north haven | 23-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Fistula | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY WITH LATEROLATERAL ANASTOMOSIS PROCEDURE, THE DEVICE PERFORMED AS USUAL. POSTOPERATIVELY, THE STAPLE LINE DID NOT HOLD OR OPENED UP (DEHISCENCE). THE PATIENT EXPERIENCED EXTENDED HOSPITALIZATION FOR 8 DAYS AND WAS ADMITTED TO THE INTENSIVE CARE UNIT. PLANS TO RE-OPERATE ON THE PATIENT WERE MADE, BUT IT WAS NOT POSSIBLE. THE PATIENT OUTCOME WAS DEATH, WITH THE CAUSE OF DEATH CURRENTLY UNKNOWN.
26-12-2024 | Death | covidien lp llc north haven | 23-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Sepsis | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY WITH LATEROLATERAL ANASTOMOSIS PROCEDURE, THE DEVICE PERFORMED AS USUAL. POSTOPERATIVELY, THE STAPLE LINE DID NOT HOLD OR OPENED UP (DEHISCENCE). THE PATIENT EXPERIENCED EXTENDED HOSPITALIZATION FOR 8 DAYS AND WAS ADMITTED TO THE INTENSIVE CARE UNIT. PLANS TO RE-OPERATE ON THE PATIENT WERE MADE, BUT IT WAS NOT POSSIBLE. THE PATIENT OUTCOME WAS DEATH, WITH THE CAUSE OF DEATH CURRENTLY UNKNOWN.
26-12-2024 | Death | covidien lp llc north haven | 23-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Peritonitis | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY WITH LATEROLATERAL ANASTOMOSIS PROCEDURE, THE DEVICE PERFORMED AS USUAL. POSTOPERATIVELY, THE STAPLE LINE DID NOT HOLD OR OPENED UP (DEHISCENCE). THE PATIENT EXPERIENCED EXTENDED HOSPITALIZATION FOR 8 DAYS AND WAS ADMITTED TO THE INTENSIVE CARE UNIT. PLANS TO RE-OPERATE ON THE PATIENT WERE MADE, BUT IT WAS NOT POSSIBLE. THE PATIENT OUTCOME WAS DEATH, WITH THE CAUSE OF DEATH CURRENTLY UNKNOWN.
26-12-2024 | Death | covidien lp llc north haven | 23-01-2025 | GIA | Adverse Event Without Identified Device or Use Problem | A24 | Organ Dehiscence | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY WITH LATEROLATERAL ANASTOMOSIS PROCEDURE, THE DEVICE PERFORMED AS USUAL. POSTOPERATIVELY, THE STAPLE LINE DID NOT HOLD OR OPENED UP (DEHISCENCE). THE PATIENT EXPERIENCED EXTENDED HOSPITALIZATION FOR 8 DAYS AND WAS ADMITTED TO THE INTENSIVE CARE UNIT. PLANS TO RE-OPERATE ON THE PATIENT WERE MADE, BUT IT WAS NOT POSSIBLE. THE PATIENT OUTCOME WAS DEATH, WITH THE CAUSE OF DEATH CURRENTLY UNKNOWN.
31-12-2024 | Malfunction | covidien lp llc north haven | 23-01-2025 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC SLEEVE GASTRECTOMY, DURING SET-UP, THERE WAS A LOOSE CONNECTION FROM SHELL TO ADAPTER. A NEW SHELL WAS OPENED, AND THE SAME HANDLE AND ADAPTER WERE USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
24-10-2024 | Malfunction | covidien lp llc north haven | 23-01-2025 | ENDO GIA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A HEMICOLECTOMY PROCEDURE, THE DEVICE PARTIALLY CUT, BUT THE STAPLE LINE WAS COMPLETE. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
30-12-2024 | Injury | intuitive surgical | 23-01-2025 | SUREFORM | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LEFT LOWER PULMONARY LOBECTOMY, THE DISTAL EDGE OF THE STAPLE LINE FROM THE SUREFORM 30 INSTRUMENT FAILED AND BLED UNTIL ALL PULMONARY ARTERY BRANCHES TO THE LOWER LOBE WERE TAKEN/STAPLED. ACCORDING TO THE CUSTOMER, THE SUREFORM 30 STAPLER INSTRUMENT WITH A WHITE SUREFORM 30 RELOAD "DIDN'T FEEL AS EXPECTED." THE PROCEDURE WAS FINISHED ROBOTICALLY.
24-12-2024 | Injury | intuitive surgical | 23-01-2025 | SUREFORM | Misfire | A050502 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS (ROUX-EN-Y) PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT MISFIRED AND AS A RESULT, THE SURGEON HAD TO PERFORM A PARTIAL GASTRECTOMY. THE PROCEDURE WAS DELAYED ABOUT 30 MINUTES TO AN HOUR BUT WAS COMPLETED ROBOTICALLY. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN MADE; HOWEVER, NO FURTHER DETAILS HAVE BEEN RECEIVED.
17-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | ENDO ILS 25MM, CURVED | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT BEFORE A HEMICOLECTOMY SURGERY, OPEN THE PACKING AND NOTED ANVIL CANNOT BE ATTACHED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
30-12-2024 | Malfunction | intuitive surgical | 24-01-2025 | ENDOWRIST |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: THE STAPLE LOAD MALFUNCTIONED AND WOULD NOT FIRE OR RELEASE FROM THE TISSUE.
15-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER DID NOT TRIGGER, ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. NO PATIENT HARM OR SURGICAL DELAY HAVE BEEN REPORTED. THE NON-TRIGGERING OF THE STACKER WAS DUE TO A TECHNICALLY FAULTY CONNECTION OF THE HEAD TO THE STAPLER. THE CUSTOMER HAD USED ACOUSTIC FEEDBACK TO CHECK THAT THE HEAD AND STAPLER WERE SECURELY CONNECTED, BUT THE HEAD STILL SLIPPED OFF THE MANDREL WHEN THEY WERE SCREWED TOGETHER, SO THAT THE STAPLER COULD NOT BE TRIGGERED.
15-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER DID NOT TRIGGER, ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. NO PATIENT HARM OR SURGICAL DELAY HAVE BEEN REPORTED. THE NON-TRIGGERING OF THE STACKER WAS DUE TO A TECHNICALLY FAULTY CONNECTION OF THE HEAD TO THE STAPLER. THE CUSTOMER HAD USED ACOUSTIC FEEDBACK TO CHECK THAT THE HEAD AND STAPLER WERE SECURELY CONNECTED, BUT THE HEAD STILL SLIPPED OFF THE MANDREL WHEN THEY WERE SCREWED TOGETHER, SO THAT THE STAPLER COULD NOT BE TRIGGERED.
15-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER DID NOT TRIGGER, ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. NO PATIENT HARM OR SURGICAL DELAY HAVE BEEN REPORTED. THE NON-TRIGGERING OF THE STACKER WAS DUE TO A TECHNICALLY FAULTY CONNECTION OF THE HEAD TO THE STAPLER. THE CUSTOMER HAD USED ACOUSTIC FEEDBACK TO CHECK THAT THE HEAD AND STAPLER WERE SECURELY CONNECTED, BUT THE HEAD STILL SLIPPED OFF THE MANDREL WHEN THEY WERE SCREWED TOGETHER, SO THAT THE STAPLER COULD NOT BE TRIGGERED.
15-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER DID NOT TRIGGER, ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. NO PATIENT HARM OR SURGICAL DELAY HAVE BEEN REPORTED. THE NON-TRIGGERING OF THE STACKER WAS DUE TO A TECHNICALLY FAULTY CONNECTION OF THE HEAD TO THE STAPLER. THE CUSTOMER HAD USED ACOUSTIC FEEDBACK TO CHECK THAT THE HEAD AND STAPLER WERE SECURELY CONNECTED, BUT THE HEAD STILL SLIPPED OFF THE MANDREL WHEN THEY WERE SCREWED TOGETHER, SO THAT THE STAPLER COULD NOT BE TRIGGERED.
13-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION THE DEVICE WOULD NOT FIRE OUT THE RELOADS. THE PROCEDURE WAS DELAYED BY FIVE MINUTES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY.
13-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLON RESECTION THE DEVICE WOULD NOT FIRE OUT THE RELOADS. THE PROCEDURE WAS DELAYED BY FIVE MINUTES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY.
16-12-2024 | Malfunction | ethicon endo-surgery | 24-01-2025 | ILS 33MM, CURVED |  |  | Failure to Anastomose | Event Description: THE ANVIL AND SPIKE PUSHED THROUGH THE PROXIMAL TRANSECTION SITE POSTERIOR TO THE STAPLE LINE. AFTER SEVERAL ATTEMPTS THE ANVIL SPIKE WOULD NOT DISENGAGE FROM THE ANVIL UTILIZING VARIOUS TECHNIQUES. A NEW STAPLER WAS OBTAINED AND PLACED. THE ANVIL PLACED WITHOUT THE SPIKE USING A PURSE STRING TO CLOSE AROUND THE ENTEROTOMY. THE ENDOSCOPIC ENDONASAL APPROACH (EEA) WAS PLACED VIA THE RECTUM AND THE DISTAL SPIKE DEPLOYED. THE TWO ENDS WERE CONNECTED, THE STAPLER CLOSED, AND THE STAPLER FIRED. INSPECTION OF THE ANASTOMOTIC DONUTS SHOWED AN INTACT DISTAL DONUT, BUT THE PROXIMAL DONUT WAS NOT INTACT. LEAK TEST WAS POSITIVE, AND THE DEFECT FOUND ANTERIORLY.
16-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING AN UNKNOWN RESPIRATORY SURGERY; THE DEVICE COULD BE FIRED AT 1ST FIRING WITHOUT PROBLEM. THE DEVICE COULD NOT BE FIRED AND LOCKED OUT AT 2ND FIRING, THE KNIFE DID NOT MOVE AND STOPPED IN THE MIDDLE, AND THE REVERSE BUTTON WAS USED TO RETURN THE KNIFE TO THE ORIGINAL POSITION. A NEW DEVICE AND CARTRIDGE WERE USED TO COMPLETE THE CASE. NO FURTHER INFORMATION IS AVAILABLE.
16-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING AN UNKNOWN RESPIRATORY SURGERY; THE DEVICE COULD BE FIRED AT 1ST FIRING WITHOUT PROBLEM. THE DEVICE COULD NOT BE FIRED AND LOCKED OUT AT 2ND FIRING, THE KNIFE DID NOT MOVE AND STOPPED IN THE MIDDLE, AND THE REVERSE BUTTON WAS USED TO RETURN THE KNIFE TO THE ORIGINAL POSITION. A NEW DEVICE AND CARTRIDGE WERE USED TO COMPLETE THE CASE. NO FURTHER INFORMATION IS AVAILABLE.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER APPLYING ALL THREE STAPLES IN THE UPPER LOBE OF THE LUNG, AT CHEST CLOSURE, THERE WAS NO AIR LEAK. HOWEVER, SUBSEQUENTLY WHILE THE PATIENT WAS STILL ON THE OPERATION TABLE, THERE WAS A MASSIVE AIR LEAK NOTED FOR ALL DEVICES. THE CHEST OF THE PATIENT HAD TO BE RE-OPENED TO REPAIR THE STAPLE LINE WHICH HAD DEHISCED. SUTURING WAS PERFORMED AS REINFORCEMENT OF THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE ISSUE.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER APPLYING ALL THREE STAPLES IN THE UPPER LOBE OF THE LUNG, AT CHEST CLOSURE, THERE WAS NO AIR LEAK. HOWEVER, SUBSEQUENTLY WHILE THE PATIENT WAS STILL ON THE OPERATION TABLE, THERE WAS A MASSIVE AIR LEAK NOTED FOR ALL DEVICES. THE CHEST OF THE PATIENT HAD TO BE RE-OPENED TO REPAIR THE STAPLE LINE WHICH HAD DEHISCED. SUTURING WAS PERFORMED AS REINFORCEMENT OF THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE ISSUE.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, THE JAWS WAS CLOSED, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE FIRING COULD NOT BE DONE. BOTH HANDLE AND RELOAD WERE REPLACED WITH THE SAME MODEL HOWEVER THE SAME ISSUE OCCURRED. ANOTHER STAPLER FROM A COMPETITOR WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT FOR BOTH PAIR OF THE DEVICE, THAT THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE CARTRIDGE JAWS AFTER THE ATTEMPT TO FIRE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, THE JAWS WAS CLOSED, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE FIRING COULD NOT BE DONE. BOTH HANDLE AND RELOAD WERE REPLACED WITH THE SAME MODEL HOWEVER THE SAME ISSUE OCCURRED. ANOTHER STAPLER FROM A COMPETITOR WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT FOR BOTH PAIR OF THE DEVICE, THAT THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE CARTRIDGE JAWS AFTER THE ATTEMPT TO FIRE. THERE WAS NO PATIENT INJURY.
03-12-2024 | Malfunction | ethicon endo-surgery | 24-01-2025 | POWERED PLUS COMPACT 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN THE BATTERY IS ADJUSTED AND DOES NOT TRIGGER THE SHOT, THE BATTERY IS REMOVED AND ASSEMBLED AGAIN, DESPITE SEVERAL ATTEMPTS IT DOES NOT WORK; ANOTHER ECHELON IS PASSED OUT AND WORKED PERFECTLY. THERE WERE NO PATIENT CONSEQUENCES.
09-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | SIGNIA | Display or Visual Feedback Problem | A0902 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A FILE ERROR ISSUE OCCURRED WITH THE HANDLE DURING PRE-OPERATIVE PREPARATION. THE UNIT CONSISTENTLY DISPLAYED A FILE ERROR AND SUBSEQUENTLY LOST ALL CHARGE. THERE WAS NO PATIENT INVOLVEMENT. THE HANDLE WAS TESTED ON BOTH THE USER'S AND A CUSTOMER'S FOUR-PORT CHARGER, BUT NO POSITIVE RESULTS WERE OBTAINED.
09-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT A FILE ERROR ISSUE OCCURRED WITH THE HANDLE DURING PRE-OPERATIVE PREPARATION. THE UNIT CONSISTENTLY DISPLAYED A FILE ERROR AND SUBSEQUENTLY LOST ALL CHARGE. THERE WAS NO PATIENT INVOLVEMENT. THE HANDLE WAS TESTED ON BOTH THE USER'S AND A CUSTOMER'S FOUR-PORT CHARGER, BUT NO POSITIVE RESULTS WERE OBTAINED.
23-10-2024 | Malfunction | ethicon endo-surgery | 24-01-2025 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE STAPLER IS PASSED ON, THE STAPLES ARE MOUNTED, IT IS ASSEMBLED ACCORDING TO THE MANUFACTURER (POWER SUPPLY), THE SAFETY IS REMOVED AND IT DOES NOT FIRE, IT IS CHECKED AGAIN AND IT DOES NOT FIRE. AT THE END THE SURGEON DESISTS FROM ITS USE AND USES INSTEAD A TITANIUM LIGATURE CLIP WITH CLAMP. THE SURGICAL PROCEDURE WAS DELAYED 20 MINUTES. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
23-10-2024 | Malfunction | ethicon endo-surgery | 24-01-2025 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DOCTOR STATES THAT THE SHEARS WERE STOPPED AND DID NOT WORK. THERE WAS NO SURGICAL DELAY. THERE WERE NO PATIENT CONSEQUENCES.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, THE JAWS WAS CLOSED, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE FIRING COULD NOT BE DONE. BOTH HANDLE AND RELOAD WERE REPLACED WITH THE SAME MODEL HOWEVER THE SAME ISSUE OCCURRED. ANOTHER STAPLER FROM A COMPETITOR WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT FOR BOTH PAIR OF THE DEVICE, THAT THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE CARTRIDGE JAWS AFTER THE ATTEMPT TO FIRE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | TRI-STAPLE 2.0 | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, THE JAWS WAS CLOSED, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE FIRING COULD NOT BE DONE. BOTH HANDLE AND RELOAD WERE REPLACED WITH THE SAME MODEL HOWEVER THE SAME ISSUE OCCURRED. ANOTHER STAPLER FROM A COMPETITOR WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT FOR BOTH PAIR OF THE DEVICE, THAT THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE CARTRIDGE JAWS AFTER THE ATTEMPT TO FIRE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, THE JAWS WAS CLOSED, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE FIRING COULD NOT BE DONE. BOTH HANDLE AND RELOAD WERE REPLACED WITH THE SAME MODEL HOWEVER THE SAME ISSUE OCCURRED. ANOTHER STAPLER FROM A COMPETITOR WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT FOR BOTH PAIR OF THE DEVICE, THAT THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE CARTRIDGE JAWS AFTER THE ATTEMPT TO FIRE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC LUNG RESECTION, THE JAWS WAS CLOSED, THE USER WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE FIRING COULD NOT BE DONE. BOTH HANDLE AND RELOAD WERE REPLACED WITH THE SAME MODEL HOWEVER THE SAME ISSUE OCCURRED. ANOTHER STAPLER FROM A COMPETITOR WAS USED TO RESOLVE THE ISSUE. IT WAS NOTED THAT FOR BOTH PAIR OF THE DEVICE, THAT THE STAPLES WERE PROTRUDING NEAR THE BASE OF THE CARTRIDGE JAWS AFTER THE ATTEMPT TO FIRE. THERE WAS NO PATIENT INJURY.
09-01-2025 | Injury | us surgical puerto rico | 24-01-2025 | ENDO GIA | Break | A0401 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER DURING A ROBOTIC GASTRECTOMY, WHEN CLOSING THE INSERTION PORT OF DELTA RECONSTRUCTION, THE RELOAD WAS ARTICULATED TO THE LEFT AND WAS FIRED. THE DEVICE WAS ARTICULATED SIGNIFICANTLY TO THE LEFT, MORE THAN THE USUAL TWITCHING. AFTER THAT WHEN TRYING TO RETURN TO THE NEUTRAL POSITION, IT DID NOT BECOME STRAIGHT SO THE CENTER CONTROL WAS PRESSED TO THE RIGHT, STRAIGHTENED IT AND REMOVED FROM THE TROCAR. THE SURGEON FELT UNCOMFORTABLE AND UPON OBSERVING THE ABDOMINAL CAVITY, DISCOVERED A METAL FRAGMENT DIRECTLY UNDER THE TROCAR FROM WHICH THE STAPLER WAS REMOVED. THE METAL FRAGMENTS WERE REMOVED FROM THE PATIENT. THE PATIENT WAS CHECKED WITH AN X-RAY AND NO OTHER METAL REACTIONS WERE OBSERVED AND SO THE PATIENT WAS CLOSED AS IS. WHEN CHECKING THE CARTRIDGE, IT APPEARED THAT THE INTERNAL PART OF THE ARTICULATED PART WAS DAMAGED.
09-01-2025 | Injury | us surgical puerto rico | 24-01-2025 | ENDO GIA | Detachment of Device or Device Component | A0501 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER DURING A ROBOTIC GASTRECTOMY, WHEN CLOSING THE INSERTION PORT OF DELTA RECONSTRUCTION, THE RELOAD WAS ARTICULATED TO THE LEFT AND WAS FIRED. THE DEVICE WAS ARTICULATED SIGNIFICANTLY TO THE LEFT, MORE THAN THE USUAL TWITCHING. AFTER THAT WHEN TRYING TO RETURN TO THE NEUTRAL POSITION, IT DID NOT BECOME STRAIGHT SO THE CENTER CONTROL WAS PRESSED TO THE RIGHT, STRAIGHTENED IT AND REMOVED FROM THE TROCAR. THE SURGEON FELT UNCOMFORTABLE AND UPON OBSERVING THE ABDOMINAL CAVITY, DISCOVERED A METAL FRAGMENT DIRECTLY UNDER THE TROCAR FROM WHICH THE STAPLER WAS REMOVED. THE METAL FRAGMENTS WERE REMOVED FROM THE PATIENT. THE PATIENT WAS CHECKED WITH AN X-RAY AND NO OTHER METAL REACTIONS WERE OBSERVED AND SO THE PATIENT WAS CLOSED AS IS. WHEN CHECKING THE CARTRIDGE, IT APPEARED THAT THE INTERNAL PART OF THE ARTICULATED PART WAS DAMAGED.
09-01-2025 | Injury | us surgical puerto rico | 24-01-2025 | ENDO GIA | Human-Device Interface Problem | A22 | Unintended Radiation Exposure | Event Description: ACCORDING TO THE REPORTER DURING A ROBOTIC GASTRECTOMY, WHEN CLOSING THE INSERTION PORT OF DELTA RECONSTRUCTION, THE RELOAD WAS ARTICULATED TO THE LEFT AND WAS FIRED. THE DEVICE WAS ARTICULATED SIGNIFICANTLY TO THE LEFT, MORE THAN THE USUAL TWITCHING. AFTER THAT WHEN TRYING TO RETURN TO THE NEUTRAL POSITION, IT DID NOT BECOME STRAIGHT SO THE CENTER CONTROL WAS PRESSED TO THE RIGHT, STRAIGHTENED IT AND REMOVED FROM THE TROCAR. THE SURGEON FELT UNCOMFORTABLE AND UPON OBSERVING THE ABDOMINAL CAVITY, DISCOVERED A METAL FRAGMENT DIRECTLY UNDER THE TROCAR FROM WHICH THE STAPLER WAS REMOVED. THE METAL FRAGMENTS WERE REMOVED FROM THE PATIENT. THE PATIENT WAS CHECKED WITH AN X-RAY AND NO OTHER METAL REACTIONS WERE OBSERVED AND SO THE PATIENT WAS CLOSED AS IS. WHEN CHECKING THE CARTRIDGE, IT APPEARED THAT THE INTERNAL PART OF THE ARTICULATED PART WAS DAMAGED.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GASTRIC SEPTATION, THE INSTRUMENT DID NOT FIRE, THE WAS SURGEON UNABLE TO PRESS THE GREEN SAFETY BUTTON BUT COULD NOT SQUEEZE THE HANDLE. A NEW RELOAD WAS USED TO RESOLVED THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER AFTER THE EJ (ESOPHAGOJEJUNOSTOMY) SIDE-TO-SIDE ANASTOMOSIS, LEAKAGE OCCURRED AT THE END TIP OF THE EJ JUN CTION STAPLE LINE. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE PATIENT DEVELOPED ABSCESSES IN THE DIAPHRAGM AND LUNG AREA. NO INTERVENTION WAS DONE/REQUIRED.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Abscess | Event Description: ACCORDING TO THE REPORTER AFTER THE EJ (ESOPHAGOJEJUNOSTOMY) SIDE-TO-SIDE ANASTOMOSIS, LEAKAGE OCCURRED AT THE END TIP OF THE EJ JUN CTION STAPLE LINE. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE PATIENT DEVELOPED ABSCESSES IN THE DIAPHRAGM AND LUNG AREA. NO INTERVENTION WAS DONE/REQUIRED.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER AFTER THE EJ (ESOPHAGOJEJUNOSTOMY) SIDE-TO-SIDE ANASTOMOSIS, LEAKAGE OCCURRED AT THE END TIP OF THE EJ JUN CTION STAPLE LINE. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE PATIENT DEVELOPED ABSCESSES IN THE DIAPHRAGM AND LUNG AREA. NO INTERVENTION WAS DONE/REQUIRED.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE REPORTER AFTER THE EJ (ESOPHAGOJEJUNOSTOMY) SIDE-TO-SIDE ANASTOMOSIS, LEAKAGE OCCURRED AT THE END TIP OF THE EJ JUN CTION STAPLE LINE. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE PATIENT DEVELOPED ABSCESSES IN THE DIAPHRAGM AND LUNG AREA. NO INTERVENTION WAS DONE/REQUIRED.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER APPLYING ALL THREE STAPLES IN THE UPPER LOBE OF THE LUNG, AT CHEST CLOSURE, THERE WAS NO AIR LEAK. HOWEVER, SUBSEQUENTLY WHILE THE PATIENT WAS STILL ON THE OPERATION TABLE, THERE WAS A MASSIVE AIR LEAK NOTED FOR ALL DEVICES. THE CHEST OF THE PATIENT HAD TO BE RE-OPENED TO REPAIR THE STAPLE LINE WHICH HAD DEHISCED. SUTURING WAS PERFORMED AS REINFORCEMENT OF THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE ISSUE.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER APPLYING ALL THREE STAPLES IN THE UPPER LOBE OF THE LUNG, AT CHEST CLOSURE, THERE WAS NO AIR LEAK. HOWEVER, SUBSEQUENTLY WHILE THE PATIENT WAS STILL ON THE OPERATION TABLE, THERE WAS A MASSIVE AIR LEAK NOTED FOR ALL DEVICES. THE CHEST OF THE PATIENT HAD TO BE RE-OPENED TO REPAIR THE STAPLE LINE WHICH HAD DEHISCED. SUTURING WAS PERFORMED AS REINFORCEMENT OF THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE ISSUE.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Wound Dehiscence | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER APPLYING ALL THREE STAPLES IN THE UPPER LOBE OF THE LUNG, AT CHEST CLOSURE, THERE WAS NO AIR LEAK. HOWEVER, SUBSEQUENTLY WHILE THE PATIENT WAS STILL ON THE OPERATION TABLE, THERE WAS A MASSIVE AIR LEAK NOTED FOR ALL DEVICES. THE CHEST OF THE PATIENT HAD TO BE RE-OPENED TO REPAIR THE STAPLE LINE WHICH HAD DEHISCED. SUTURING WAS PERFORMED AS REINFORCEMENT OF THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE ISSUE.
07-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER APPLYING ALL THREE STAPLES IN THE UPPER LOBE OF THE LUNG, AT CHEST CLOSURE, THERE WAS NO AIR LEAK. HOWEVER, SUBSEQUENTLY WHILE THE PATIENT WAS STILL ON THE OPERATION TABLE, THERE WAS A MASSIVE AIR LEAK NOTED FOR ALL DEVICES. THE CHEST OF THE PATIENT HAD TO BE RE-OPENED TO REPAIR THE STAPLE LINE WHICH HAD DEHISCED. SUTURING WAS PERFORMED AS REINFORCEMENT OF THE STAPLE LINE. THE SURGICAL TIME WAS EXTENDED BY 45 MINUTES DUE TO THE ISSUE.
31-12-2024 | Malfunction | us surgical puerto rico | 24-01-2025 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PARTIAL GASTRECTOMY, THE STAPLER DID NOT FIRE. IT WAS ALSO NOTED THAT THE BLACK THUMB PIECE WAS UNABLE TO MOVE AT ALL. AS A RESOLUTION, A NEW IDENTICAL DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC WEDGE RESECTION OF THE RIGHT UPPER LUNG LOBE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING IT WITH A NEW DEVICE. THERE WAS NO PATIENT INJURY.
22-08-2024 | Malfunction | covidien lp llc north haven | 24-01-2025 | SIGNIA | Computer Software Problem | A11 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY SURGERY, THE POWERED STAPLER MALFUNCTIONED. WHEN PLACING THE RELOAD, WHEN IT WAS ABOUT TO BE CLIPPED, INFORMATION APPEARED ON THE DISPLAY TO FIT THE RELOAD ON THE ADAPTER, BUT THE RELOAD WAS ALREADY FITTED ONTO THE ADAPTER. DISASSEMBLY AND ASSEMBLY WERE CARRIED OUT AGAIN, AND IT WORKED. BUT WHEN ANOTHER RELOAD WAS USED, THE DEFECT APPEARED AGAIN. DISASSEMBLY AND ASSEMBLY WERE CARRIED OUT AGAIN, BUT IT DID NOT WORK. IT WAS NOTED THAT IT WAS INSTRUCTED TO DO THE PROCESS AGAIN OR CHANGE THE ADAPTER. THE PROCESS WAS REDONE, THE ADAPTER WAS CHANGED, AND THE STAPLER DID NOT WORK. TO RESOLVE THE ISSUE, A MANUAL HANDLE WAS USED TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CAUSE OF THE ARTICULATION CALIBRATION AND CALIBRATION TURNS ERRORS AS WELL AS THE DIFFICULTY DURING THE CLAMP TEST CAN BE TRACED TO FPGA RESET/REPROGRAM FAILURES. THE ROOT CAUSE OF THE FPGA REPROGRAM/RESET FAILURES COULD NOT BE ESTABLISHED.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER AFTER THE EJ (ESOPHAGOJEJUNOSTOMY) SIDE-TO-SIDE ANASTOMOSIS, LEAKAGE OCCURRED AT THE END TIP OF THE EJ JUN CTION STAPLE LINE. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE PATIENT DEVELOPED ABSCESSES IN THE DIAPHRAGM AND LUNG AREA. NO INTERVENTION WAS DONE/REQUIRED.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Abscess | Event Description: ACCORDING TO THE REPORTER AFTER THE EJ (ESOPHAGOJEJUNOSTOMY) SIDE-TO-SIDE ANASTOMOSIS, LEAKAGE OCCURRED AT THE END TIP OF THE EJ JUN CTION STAPLE LINE. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE PATIENT DEVELOPED ABSCESSES IN THE DIAPHRAGM AND LUNG AREA. NO INTERVENTION WAS DONE/REQUIRED.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Failure to Anastomose | Event Description: ACCORDING TO THE REPORTER AFTER THE EJ (ESOPHAGOJEJUNOSTOMY) SIDE-TO-SIDE ANASTOMOSIS, LEAKAGE OCCURRED AT THE END TIP OF THE EJ JUN CTION STAPLE LINE. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE PATIENT DEVELOPED ABSCESSES IN THE DIAPHRAGM AND LUNG AREA. NO INTERVENTION WAS DONE/REQUIRED.
09-01-2025 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA | Adverse Event Without Identified Device or Use Problem | A24 | Abscess | Event Description: ACCORDING TO THE REPORTER AFTER THE EJ (ESOPHAGOJEJUNOSTOMY) SIDE-TO-SIDE ANASTOMOSIS, LEAKAGE OCCURRED AT THE END TIP OF THE EJ JUN CTION STAPLE LINE. THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE PATIENT DEVELOPED ABSCESSES IN THE DIAPHRAGM AND LUNG AREA. NO INTERVENTION WAS DONE/REQUIRED.
09-01-2025 | Malfunction | covidien | 24-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC WEDGE RESECTION OF THE RIGHT UPPER LUNG LOBE, THE STAPLER WAS UNABLE TO FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. THE ISSUE WAS RESOLVED BY REPLACING IT WITH A NEW DEVICE. THERE WAS NO PATIENT INJURY.
17-10-2024 | Malfunction | covidien lp llc north haven | 24-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PARTIAL GASTRECTOMY, THE STAPLER WOULD NOT FIRE THE FIRST STAPLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-10-2024 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PARTIAL GASTRECTOMY, THE STAPLER WOULD NOT FIRE THE FIRST STAPLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-10-2024 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA ULTRA | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC PARTIAL GASTRECTOMY, THE STAPLER WOULD NOT FIRE THE FIRST STAPLE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | SIGNIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | ENDO GIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
08-01-2025 | Malfunction | covidien | 24-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THERE WERE FIRING ISSUES WITH THE RELOAD COMPONENT OF THE POWERED STAPLER, SPECIFICALLY, THE INSTRUMENT DID NOT FIRE. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
05-12-2024 | Malfunction | covidien lp llc north haven | 24-01-2025 | ABSORBATACK | Failure to Advance | A150204 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE PASSING THE DEVICE, ONE OF THE TACKS FELL WITHOUT BEING FIRED. AFTER THAT, WHEN THE TACKER WAS FIRED, THE TACKS WERE STUCK AND IT DID NOT ALLOW OTHER TACKS TO EXIT. THE TACKS WERE NOT PLACED ON THE PATIENT. A NEW TACKER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2024 | Malfunction | covidien lp llc north haven | 24-01-2025 | ABSORBATACK | Failure to Advance | A150204 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE PASSING THE DEVICE, ONE OF THE TACKS FELL WITHOUT BEING FIRED. AFTER THAT, WHEN THE TACKER WAS FIRED, THE TACKS WERE STUCK AND IT DID NOT ALLOW OTHER TACKS TO EXIT. THE TACKS WERE NOT PLACED ON THE PATIENT. A NEW TACKER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2024 | Malfunction | covidien lp llc north haven | 24-01-2025 | ABSORBATACK | Detachment of Device or Device Component | A0501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE PASSING THE DEVICE, ONE OF THE TACKS FELL WITHOUT BEING FIRED. AFTER THAT, WHEN THE TACKER WAS FIRED, THE TACKS WERE STUCK AND IT DID NOT ALLOW OTHER TACKS TO EXIT. THE TACKS WERE NOT PLACED ON THE PATIENT. A NEW TACKER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-12-2024 | Malfunction | covidien lp llc north haven | 24-01-2025 | ABSORBATACK | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHILE PASSING THE DEVICE, ONE OF THE TACKS FELL WITHOUT BEING FIRED. AFTER THAT, WHEN THE TACKER WAS FIRED, THE TACKS WERE STUCK AND IT DID NOT ALLOW OTHER TACKS TO EXIT. THE TACKS WERE NOT PLACED ON THE PATIENT. A NEW TACKER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOBECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC LOBECTOMY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO ANOTHER DEVICE TO CONTINUE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC LOBECTOMY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO ANOTHER DEVICE TO CONTINUE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
07-01-2025 | Malfunction | ethicon endo-surgery | 24-01-2025 | GST BLUE RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC LOBECTOMY, AFTER FIRED, NOTED THE STAPLES MALFORMED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
28-11-2024 | Malfunction | us surgical puerto rico | 24-01-2025 | ENDO GIA ULTRA | Insufficient Information | E2401 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE RELOAD AND MANUAL HANDLE, THERE WERE ISSUES WITH THE PRODUCT HANDLES. THE STEERING WHEEL DID NOT MOVE SMOOTHLY. THE SURGERY WAS COMPLETED USING A NEW RELOAD WITH THE SAME HANDLE, AND THE DEVICE WAS REPLACED. THERE WERE NO REPORTED CONSEQUENCES OR IMPACT TO THE PATIENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS HANDLE THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
19-12-2024 | Malfunction | ethicon endo-surgery | 24-01-2025 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE PNEUMONECTOMY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 24-01-2025 | GREEN RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PNEUMONECTOMY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 24-01-2025 | GREEN RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE PNEUMONECTOMY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
19-12-2024 | Malfunction | ethicon endo-surgery | 24-01-2025 | GREEN RELOAD FOR ECHELON 60 | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PNEUMONECTOMY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
29-12-2024 | Malfunction | ethicon endo-surgery | 24-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE LOBECTOMY, AFTER FIRED, NOTED THE STAPLES MALFORMED. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-01-2025 | Malfunction | covidien lp llc north haven | 24-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, INTRA-OPERATIVELY, THE DEVICE DID NOT FIRE. ANOTHER RELOAD WAS USED TO COMPLETE THE CASE.
10-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | SIGNIA | Unintended Movement | A0512 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A RECTAL RESECTION WITH ROBOT SUPPORT, DURING FIRING AT THE RECTUM, THERE WAS A BREAKAGE OR DAMAGE AT THE PIN. IT OCCURRED ON THE TWO ADAPTERS. MOREOVER, THE 60MM PURPLE RELOAD DID NOT OPEN AT ALL INSIDE THE ABDOMINAL CAVITY, AND THE 45MM PURPLE RELOAD HAD AN UNINTENDED ARTICULATION OR UNCONTROLLABLE PROBLEM. ANOTHER COMPANY'S PRODUCT WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT LOBE RESECTION, DURING MAIN PULMONARY VEIN RESECTION, AFTER APPROACHING THE RIGHT PULMONARY VEIN, WHEN AN ATTEMPT WAS MADE TO RESECTION, TRIED TO SQUEEZE THE HANDLE AFTER SETTING IT TO FIRING MODE ACCORDING TO THE PROCEDURE, BUT THE RESISTANCE WAS STRONG AND THE FIRING COULD NOT BE PERFORMED. THE EVENT OCCURRED DURING THE SECOND AND THIRD SHOTS OF THE FIRING OPERATION DURING SURGERY. WHEN THE HANDLE AND THE CARTRIDGE WERE REPLACED TO A NEW ONES, WAS ABLE TO PERFORM FIRING WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT LOBE RESECTION, DURING MAIN PULMONARY VEIN RESECTION, AFTER APPROACHING THE RIGHT PULMONARY VEIN, WHEN AN ATTEMPT WAS MADE TO RESECTION, TRIED TO SQUEEZE THE HANDLE AFTER SETTING IT TO FIRING MODE ACCORDING TO THE PROCEDURE, BUT THE RESISTANCE WAS STRONG AND THE FIRING COULD NOT BE PERFORMED. THE EVENT OCCURRED DURING THE SECOND AND THIRD SHOTS OF THE FIRING OPERATION DURING SURGERY. WHEN THE HANDLE AND THE CARTRIDGE WERE REPLACED TO A NEW ONES, WAS ABLE TO PERFORM FIRING WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
02-12-2024 | Malfunction | intuitive surgical | 27-01-2025 | SUREFORM |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: FROM STAFF: DURING A ROBOTICALLY ASSISTED RIGHT SIDE VAT WEDGE RESECTION, THE SUREFORM 45 CURVED TIP STAPLER WAS BEING USED WITH A BLACK LOAD. THE STAPLER WOULD NOT FIRE DUE TO TISSUE BEING TOO THICK. THE NOTICE APPEARED ON THE SCREEN THAT TISSUE WAS TOO THICK HALFWAY THROUGH THE FIRE. THE STAPLER WAS UNCLAMPED AND REMOVED. THE STAPLE LOAD HAD PARTIALLY FIRED. THERE WAS NO BLEEDING FROM THE STAPLE LINE. WHEN THE STAPLER WAS REMOVED, THE BLADE FOR THE STAPLER WAS STILL EJECTED AND VISIBLE WITHIN THE STAPLE LOAD THAT WAS INSIDE THE STAPLER. THE STAPLER WAS REMOVED FROM THE SURGICAL FIELD AND REPLACED IMMEDIATELY. FROM OPERATIVE REPORT: OF NOTE THE BLACK STAPLER LOADING USUALLY SUFFICIENT FOR EVEN DENSE PARENCHYMA COULD NOT COMPLETE THE FIRING AND REQUIRED A SEPARATE STAPLE LINE WHICH EVENTUALLY SEALED SEEMINGLY SATISFACTORY WITH GOOD HEMOSTASIS AND HEMOSTASIS ACHIEVED.
09-01-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DAVINCI THORACOSCOPIC RIGHT MIDDLE LOBECTOMY, THE FIRING COULD NOT BE DONE AT ALL. ANOTHER RELOAD HAD THE GREEN BUTTON PRESSED AND COULD NOT BE FIRED (2ND FIRING). IT WAS NOTED THAT HANDLE AND ADAPTER WAS WORKING. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
09-01-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING DAVINCI THORACOSCOPIC RIGHT MIDDLE LOBECTOMY, THE FIRING COULD NOT BE DONE AT ALL. ANOTHER RELOAD HAD THE GREEN BUTTON PRESSED, AND COULD NOT BE FIRED (2ND FIRING). IT WAS NOTED THAT HANDLE AND ADAPTER WAS WORKING. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SPLENECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLER WAS CLOSED AND FIRED, BUT WOULD NOT OPEN. REQUIRED EMERGENCY MANUAL OVERRIDE. ON CLOSER INSPECTION AFTER PASSING IT OFF THE FIELD, IT APPEARED THAT THE STAPLE LOAD BROKE. THERE WERE NO PATIENT CONSEQUENCES NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
12-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | GST WHITE RELOAD, 60MM, 6 ROW | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SPLENECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLER WAS CLOSED AND FIRED, BUT WOULD NOT OPEN. REQUIRED EMERGENCY MANUAL OVERRIDE. ON CLOSER INSPECTION AFTER PASSING IT OFF THE FIELD, IT APPEARED THAT THE STAPLE LOAD BROKE. THERE WERE NO PATIENT CONSEQUENCES NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
22-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | WHITE RELOAD FOR ECHELON 60 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE MALFUNCTIONED AND COULD NOT CONTINUE TO BE USED NORMALLY WHEN THE ANASTOMOSIS TISSUE WAS HALFWAY THROUGH PROLONGING THE OPERATION TIME AND INCREASING THE RISK OF SURGERY. CHANGED TO ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
09-01-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC-ASSISTED DISTAL GASTRECTOMY (LADG) PROCEDURE, THE RELOAD JAWS WOULD NOT CLOSE COMPLETELY. TO RESOLVE THE ISSUE, THE SURGEON REPLACED THE DEVICE WITH ANOTHER RELOAD. NO ADDITIONAL OPERATION WAS REQUIRED, AND THERE WERE NO ADVERSE OUTCOMES OR INJURIES REPORTED.
10-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGUS RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS BLOCKED DURING THE SECOND FIRING. THE MOTOR WAS ON AND THE KNIFE CUT BUT STOPPED AFTER 1CM. THEY HAD TO USE MANUAL OVERRIDE TO OPEN THE DEVICE AGAIN. A NEW STAPLER AND NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
10-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN ESOPHAGUS RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS BLOCKED DURING THE SECOND FIRING. THE MOTOR WAS ON AND THE KNIFE CUT BUT STOPPED AFTER 1CM. THEY HAD TO USE MANUAL OVERRIDE TO OPEN THE DEVICE AGAIN. A NEW STAPLER AND NEW RELOAD WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
07-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE SLEEVE GASTRECTOMY, THE INSTRUMENT DID NOT FIRE. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON BUT UNABLE TO SQUEEZE THE HANDLE. A NEW HANDLE AND RELOADS WERE OPENED. IT WAS NOTED THAT THE HANDLE WAS TESTED WITH OTHER RELOADS AND WAS FOUND TO BE FIRING, WHILE THE RELOAD WAS TESTED WITH OTHER HANDLES AND WAS FOUND NOT TO BE FIRING. THERE WAS NO PATIENT INJURY.
13-08-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE DEVICE WAS DIFFICULT TO FIRE. THE ISSUE WAS RESOLVED BY USING A NEW RELOAD WITH THE SAME HANDLE AND ADAPTER. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT. MEDTRONIC'S EVALUATION FOUND THAT THE RELOAD WAS PARTIALLY FIRED TO THE 2CM CUT LINE. THE CLAMPING MECHANISM WAS DEFORMED.
13-08-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE DEVICE WAS DIFFICULT TO FIRE. THE ISSUE WAS RESOLVED BY USING A NEW RELOAD WITH THE SAME HANDLE AND ADAPTER. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT. MEDTRONIC'S EVALUATION FOUND THAT THE RELOAD WAS PARTIALLY FIRED TO THE 2CM CUT LINE. THE CLAMPING MECHANISM WAS DEFORMED.
13-08-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE DEVICE WAS DIFFICULT TO FIRE. THE ISSUE WAS RESOLVED BY USING A NEW RELOAD WITH THE SAME HANDLE AND ADAPTER. THERE WERE NO CONSEQUENCES OR IMPACT TO THE PATIENT. MEDTRONIC'S EVALUATION FOUND THAT THE RELOAD WAS PARTIALLY FIRED TO THE 2CM CUT LINE. THE CLAMPING MECHANISM WAS DEFORMED.
06-01-2025 | Malfunction | ethicon endo-surgery | 27-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS FOR LOBECTOMY, WHEN THE PULMONARY ARTERY PROCESS WAS PERFORMED, THE KNIFE SOUNDED A SLIGHT MOVED FORWARD, AND THEN STOPPED (AT THE 3RD FIRING). THE SURGEON WAS NOT SURE HOW FAR THE KNIFE HAD MOVED FORWARD AND USED THE MANUAL OVERRIDE LEVER. THEN, THE ANVIL WAS OPENED, AND IT WAS CONFIRMED THAT THE KNIFE HAD NOT MOVED FORWARD. THERE WAS NO BLEEDING FROM THE BLOOD VESSEL. AFTERWARDS, A NEW ONE WAS PUT OUT AND THE SURGERY WAS COMPLETED WITHOUT ANY PROBLEMS. AFTER RECEIVING THE DEVICE INVOLVED AFTER THE SURGERY, THE SALES REP CHECKED THE CARTRIDGE AND THE SLEDS HAD MOVED FORWARD SLIGHTLY, SO IT IS CONSIDERED THE DEVICE MIGHT BE LOCKED OUT. HOWEVER, THE REASON OF THE SLEDS HAD MOVED FORWARD WAS POSSIBLE THAT IT WAS TOUCHED BY THE NURSE WHEN LOADED, BUT IT DID NOT SEEM TO TOUCH THE CARTRIDGE, AND THE DEVICE ITSELF WAS REMOVING THE REMAINS OF THE STAPLES AT EACH FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
06-01-2025 | Malfunction | ethicon endo-surgery | 27-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC VATS FOR LOBECTOMY, WHEN THE PULMONARY ARTERY PROCESS WAS PERFORMED, THE KNIFE SOUNDED A SLIGHT MOVED FORWARD, AND THEN STOPPED (AT THE 3RD FIRING). THE SURGEON WAS NOT SURE HOW FAR THE KNIFE HAD MOVED FORWARD AND USED THE MANUAL OVERRIDE LEVER. THEN, THE ANVIL WAS OPENED, AND IT WAS CONFIRMED THAT THE KNIFE HAD NOT MOVED FORWARD. THERE WAS NO BLEEDING FROM THE BLOOD VESSEL. AFTERWARDS, A NEW ONE WAS PUT OUT AND THE SURGERY WAS COMPLETED WITHOUT ANY PROBLEMS. AFTER RECEIVING THE DEVICE INVOLVED AFTER THE SURGERY, THE SALES REP CHECKED THE CARTRIDGE AND THE SLEDS HAD MOVED FORWARD SLIGHTLY, SO IT IS CONSIDERED THE DEVICE MIGHT BE LOCKED OUT. HOWEVER, THE REASON OF THE SLEDS HAD MOVED FORWARD WAS POSSIBLE THAT IT WAS TOUCHED BY THE NURSE WHEN LOADED, BUT IT DID NOT SEEM TO TOUCH THE CARTRIDGE, AND THE DEVICE ITSELF WAS REMOVING THE REMAINS OF THE STAPLES AT EACH FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
27-01-2025 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE STOPPED TO FIRE ON THE WAY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT NOR SURGICAL DELAY REPORTED. THE EXACT DATE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION REQUESTED.
10-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | BLUE RELOAD FOR ECHELON 60 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRED. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
07-01-2025 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED 60 ECHELON +, 440MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PANCREAS PROCEDURE, IT WAS OBSERVED THAT STAPLER WAS PERFORMING THEN STOPPED ONCE IT HIT THE STAPLER LINE AND STOPPED WORKING. CIRCULATED IN A NEW DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCES REPORTED NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
10-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | GREEN RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED AFTER FIRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
10-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | GREEN RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED AFTER FIRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
10-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | GREEN RELOAD FOR ECHELON 45 | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED AFTER FIRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
10-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | GREEN RELOAD FOR ECHELON 45 | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, IT WAS OBSERVED THAT SOME STAPLES MALFORMED AND OOZING OCCURRED AFTER FIRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
08-01-2025 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED PLUS STANDARD 45 | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS DAMAGED (MIDDLE OF THE DECK UNWRAP) WHEN FIRING . ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
16-10-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | GOLD RELOAD FOR ECHELON 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PNEUMONECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE COULD NOT FIRE FARTHER AFTER FIRED A HALF. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-11-2024 | Malfunction | davol inc sub c.r bard inc -1213643 | 27-01-2025 | CAPSURE | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: AS REPORTED, DURING A SACROCOLPOPEXY SURGERY, THE CAPSURE FIXATION DEVICE DID NOT FIRE THE FASTENER AND TRIGGER WAS STUCK. IT WAS REPORTED THAT A COMPETITOR PRODUCT WAS USED TO COMPLETE THE SURGERY. THERE WAS NO PATIENT INJURY. DURING SAMPLE EVALUATION IT WAS FOUND THAT THERE WAS A DETACHED AND MISSING INTERNAL COMPONENT THAT YIELDED A NONFUNCTIONAL DEVICE.
14-01-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Expulsion | A010401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC-ASSISTED LAPAROSCOPIC ILEAL MALIGNANT TUMOR RESECTION, AT THE TIME OF ILEOCECAL RESECTION, AFTER CONNECTING THE RELOAD INTO THE POWERED STAPLER, THE DEVICE WAS HANDED OVER TO THE SURGEON. IT WAS FOUND THAT THE DEVICE ALREADY ISSUED A STAPLE AND THE KNIFE ADVANCED. THE NURSE REPORTED THAT THE GREEN BUTTON WAS NOT TOUCHED. THE RELOAD WAS REPLACED WITH A NEW AND CONNECTED IT TO THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
14-01-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Protective Measures Problem | A16 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC-ASSISTED LAPAROSCOPIC ILEAL MALIGNANT TUMOR RESECTION, AT THE TIME OF ILEOCECAL RESECTION, AFTER CONNECTING THE RELOAD INTO THE POWERED STAPLER, THE DEVICE WAS HANDED OVER TO THE SURGEON. IT WAS FOUND THAT THE DEVICE ALREADY ISSUED A STAPLE AND THE KNIFE ADVANCED. THE NURSE REPORTED THAT THE GREEN BUTTON WAS NOT TOUCHED. THE RELOAD WAS REPLACED WITH A NEW AND CONNECTED IT TO THE SAME HANDLE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
12-01-2025 | Malfunction | johnson and johnson / ethicon | 27-01-2025 | PROXIMATE 75 MM LINEAR CUTTIER |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: FIVE ETHICON STAPLERS OF THIS LOT NUMBER FAILED TO FIRE OR FIRED BENT STAPLES THAT WERE NOT SUITABLE FOR USE. REFERENCE REPORTS: MW5165269, MW5165270, MW5165271, MW5165272 .
12-01-2025 | Malfunction | johnson and johnson / ethicon | 27-01-2025 | PROXIMATE 75 MM LINEAR CUTTIER |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: FIVE ETHICON STAPLERS OF THIS LOT NUMBER FAILED TO FIRE OR FIRED BENT STAPLES THAT WERE NOT SUITABLE FOR USE. REFERENCE REPORTS: MW5165268, MW5165270, MW5165271, MW5165272.
10-01-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC-ASSISTED DISTAL GASTRECTOMY (LADG) PROCEDURE INVOLVING THE STOMACH, THERE WERE ISSUES WITH THE RELOAD DEVICE. SPECIFICALLY, THERE WERE SET-UP ISSUES PRIOR TO FIRING, THE JAWS NOT CLOSING COMPLETELY. THE SURGEON RESOLVED THE ISSUE BY USING ANOTHER RELOAD DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO NEED FOR ADDITIONAL OPERATIONS, AND THE PATIENT DID NOT EXPERIENCE ANY INJURY OR EXTENDED HOSPITALIZATION.
12-01-2025 | Malfunction | johnson and johnson / ethicon | 27-01-2025 | PROXIMATE 75 MM LINEAR CUTTIER |  |  | Insufficient Information | Event Description: FIVE ETHICON STAPLERS OF THIS LOT NUMBER FAILED TO FIRE OR FIRED BENT STAPLES THAT WERE NOT SUITABLE FOR USE. REFERENCE REPORTS: MW5165268, MW5165269, MW5165271, MW5165272.
12-01-2025 | Malfunction | johnson and johnson / ethicon | 27-01-2025 | PROXIMATE 75 MM LINEAR CUTTIER |  |  | Insufficient Information | Event Description: FIVE ETHICON STAPLERS OF THIS LOT NUMBER FAILED TO FIRE OR FIRED BENT STAPLES THAT WERE NOT SUITABLE FOR USE. REFERENCE REPORTS: MW5165268, MW5165269, MW5165270, MW5165272.
12-01-2025 | Malfunction | johnson and johnson / ethicon | 27-01-2025 | PROXIMATE 75 MM LINEAR CUTTIER |  |  | Insufficient Information | Event Description: FIVE ETHICON STAPLERS OF THIS LOT NUMBER FAILED TO FIRE OR FIRED BENT STAPLES THAT WERE NOT SUITABLE FOR USE. REFERENCE REPORTS: MW5165268, MW5165269, MW5165270, MW5165271.
10-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC-ASSISTED DISTAL GASTRECTOMY (LADG) PROCEDURE INVOLVING THE STOMACH, THERE WERE ISSUES WITH THE RELOAD DEVICE. SPECIFICALLY, THERE WERE SET-UP ISSUES PRIOR TO FIRING, THE JAWS NOT CLOSING COMPLETELY. THE SURGEON RESOLVED THE ISSUE BY USING ANOTHER RELOAD DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO NEED FOR ADDITIONAL OPERATIONS, AND THE PATIENT DID NOT EXPERIENCE ANY INJURY OR EXTENDED HOSPITALIZATION.
22-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER REMOVING THE SHIPPING WEDGE BEFORE USE, IT APPEARS TO BE FULLY REMOVED. HOWEVER WHEN FIRING THE LOAD, A PIECE FALLS INTO THE ABDOMEN. THE PIECE WAS SUCCESSFULLY REMOVED FROM THE PATIENT'S ABDOMEN AND RELOAD WAS SUCCESSFUL FIRED. THERE WAS NO PATIENT INJURY.
22-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER REMOVING THE SHIPPING WEDGE BEFORE USE, IT APPEARS TO BE FULLY REMOVED. HOWEVER WHEN FIRING THE LOAD, A PIECE FALLS INTO THE ABDOMEN. THE PIECE WAS SUCCESSFULLY REMOVED FROM THE PATIENT'S ABDOMEN AND RELOAD WAS SUCCESSFUL FIRED. THERE WAS NO PATIENT INJURY.
27-10-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO ENTERECTOMY, WHILE STAPLING THE STOMACH, THE RELOAD DID NOT PROGRESS ALL THE WAY. IT STARTED FIRING THEN STOPPED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
10-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC-ASSISTED DISTAL GASTRECTOMY (LADG) PROCEDURE INVOLVING THE STOMACH, THERE WERE ISSUES WITH THE RELOAD DEVICE. SPECIFICALLY, THERE WERE SET-UP ISSUES PRIOR TO FIRING, THE JAWS NOT CLOSING COMPLETELY. THE SURGEON RESOLVED THE ISSUE BY USING ANOTHER RELOAD DEVICE TO COMPLETE THE PROCEDURE. THERE WAS NO NEED FOR ADDITIONAL OPERATIONS, AND THE PATIENT DID NOT EXPERIENCE ANY INJURY OR EXTENDED HOSPITALIZATION.
03-12-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GENERAL SURGERY, DURING CUTTING AND CLOSING THE DIGESTIVE TRACT TISSUE, THE FIRING WAS NOT SMOOTH. A NEW HANDLE OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
03-12-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC GENERAL SURGERY, DURING CUTTING AND CLOSING THE DIGESTIVE TRACT TISSUE, THE FIRING WAS NOT SMOOTH. A NEW HANDLE OF THE SAME MODEL WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT IS THAT THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
23-12-2024 | Malfunction | johnson & johnson medical (suzhou | 27-01-2025 | CARTRIDGE FINAL ASSY 60MM BLUE | Structural Problem | A0511 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, OPEN THE PACKING AND NOTED THAT SOME SLED FELL OFF. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER REMOVING THE SHIPPING WEDGE BEFORE USE, IT APPEARS TO BE FULLY REMOVED. HOWEVER WHEN FIRING THE LOAD, A PIECE FALLS INTO THE ABDOMEN. THE PIECE WAS SUCCESSFULLY REMOVED FROM THE PATIENT'S ABDOMEN AND RELOAD WAS SUCCESSFUL FIRED. THERE WAS NO PATIENT INJURY.
21-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER REMOVING THE SHIPPING WEDGE BEFORE USE, IT APPEARS TO BE FULLY REMOVED. HOWEVER WHEN FIRING THE LOAD, A PIECE FALLS INTO THE ABDOMEN. THE PIECE WAS SUCCESSFULLY REMOVED FROM THE PATIENT'S ABDOMEN AND RELOAD WAS SUCCESSFUL FIRED. THERE WAS NO PATIENT INJURY.
12-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC HEPATECTOMY SURGERY, COULD NOT FIRE WITHOUT MOTOR SOUND ON THE FIRING. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL COULD NOT FIRE AND WITHOUT MOTOR SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
23-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER REMOVING THE SHIPPING WEDGE BEFORE USE, IT APPEARS TO BE FULLY REMOVED. HOWEVER WHEN FIRING THE LOAD, A PIECE FALLS INTO THE ABDOMEN. THE PIECE WAS SUCCESSFULLY REMOVED FROM THE PATIENT'S ABDOMEN AND RELOAD WAS SUCCESSFUL FIRED. THERE WAS NO PATIENT INJURY.
23-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | TRI-STAPLE 2.0 | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER REMOVING THE SHIPPING WEDGE BEFORE USE, IT APPEARS TO BE FULLY REMOVED. HOWEVER WHEN FIRING THE LOAD, A PIECE FALLS INTO THE ABDOMEN. THE PIECE WAS SUCCESSFULLY REMOVED FROM THE PATIENT'S ABDOMEN AND RELOAD WAS SUCCESSFUL FIRED. THERE WAS NO PATIENT INJURY.
05-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SIGMOIDECTOMY THAT THE DEVICE APPARENTLY LOCKED OUT BEFORE USING IT DURING LAP SIGMOID. THEY GOT A NEW DEVICE TO COMPLETE THE CASE WITHOUT PATIENT HARM.
16-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | ECHELON 3000 45MM COMPACT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A SEMI-OPEN PNEUMONECTOMY, AFTER THE CARTRIDGE WAS USED AT THE 1ST FIRING, THE CARTRIDGE COULD NOT BE LOADED FOR THE 2ND FIRING. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
09-01-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE EGIA60AMT EGIA 60 ARTIC MED THICK SULU EXPERIENCED ISSUES WHERE THE JAW DID NOT CLOSE. THE DEVICE WAS REPLACED WITH A NEW RELOAD TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
28-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG WEDGE RESECTION PROCEDURE INVOLVING FOUR RELOADS AND THREE MANUAL HANDLES, THE RELOAD WAS PLACED ON THE HANDLE, BUT WHEN THE RELEASE BUTTON WAS PRESSED, A CLICKING NOISE WAS HEARD FROM THE HANDLE. AFTER PRESSING THE HANDLE AGAIN, THERE WAS NO RESISTANCE, AND NOTHING HAPPENED. THE TISSUE WAS NOT CUT OR STAPLED. A NEW HANDLE WAS USED WITH THE OLD RELOAD, BUT THE ISSUE PERSISTED. THE RELOAD WAS NOT THE PROBLEM, AS IT OPENED AND CLOSED WITHOUT ANY ISSUES. THE PROCEDURE WAS COMPLETED USING A COMPETITOR'S PRODUCT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
28-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG WEDGE RESECTION PROCEDURE INVOLVING FOUR RELOADS AND THREE MANUAL HANDLES, THE RELOAD WAS PLACED ON THE HANDLE, BUT WHEN THE RELEASE BUTTON WAS PRESSED, A CLICKING NOISE WAS HEARD FROM THE HANDLE. AFTER PRESSING THE HANDLE AGAIN, THERE WAS NO RESISTANCE, AND NOTHING HAPPENED. THE TISSUE WAS NOT CUT OR STAPLED. A NEW HANDLE WAS USED WITH THE OLD RELOAD, BUT THE ISSUE PERSISTED. THE RELOAD WAS NOT THE PROBLEM, AS IT OPENED AND CLOSED WITHOUT ANY ISSUES. THE PROCEDURE WAS COMPLETED USING A COMPETITOR'S PRODUCT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
01-01-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | GST GREEN RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WAS A DEFECTIVE COUNTERFEIT PRODUCT THAT WAS OBTAINED BY HOMELAND SECURITY AND BORDER PATROL. NO PATIENT INVOLVEMENT.
01-01-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | FLEX 60 ARTICULATING | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THERE WAS A DEFECTIVE COUNTERFEIT PRODUCT THAT WAS OBTAINED BY HOMELAND SECURITY AND BORDER PATROL. NO PATIENT INVOLVEMENT.
28-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG WEDGE RESECTION PROCEDURE INVOLVING FOUR RELOADS AND THREE MANUAL HANDLES, THE RELOAD WAS PLACED ON THE HANDLE, BUT WHEN THE RELEASE BUTTON WAS PRESSED, A CLICKING NOISE WAS HEARD FROM THE HANDLE. AFTER PRESSING THE HANDLE AGAIN, THERE WAS NO RESISTANCE, AND NOTHING HAPPENED. THE TISSUE WAS NOT CUT OR STAPLED. A NEW HANDLE WAS USED WITH THE OLD RELOAD, BUT THE ISSUE PERSISTED. THE RELOAD WAS NOT THE PROBLEM, AS IT OPENED AND CLOSED WITHOUT ANY ISSUES. THE PROCEDURE WAS COMPLETED USING A COMPETITOR'S PRODUCT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
28-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG WEDGE RESECTION PROCEDURE INVOLVING FOUR RELOADS AND THREE MANUAL HANDLES, THE RELOAD WAS PLACED ON THE HANDLE, BUT WHEN THE RELEASE BUTTON WAS PRESSED, A CLICKING NOISE WAS HEARD FROM THE HANDLE. AFTER PRESSING THE HANDLE AGAIN, THERE WAS NO RESISTANCE, AND NOTHING HAPPENED. THE TISSUE WAS NOT CUT OR STAPLED. A NEW HANDLE WAS USED WITH THE OLD RELOAD, BUT THE ISSUE PERSISTED. THE RELOAD WAS NOT THE PROBLEM, AS IT OPENED AND CLOSED WITHOUT ANY ISSUES. THE PROCEDURE WAS COMPLETED USING A COMPETITOR'S PRODUCT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
28-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG WEDGE RESECTION PROCEDURE INVOLVING FOUR RELOADS AND THREE MANUAL HANDLES, THE RELOAD WAS PLACED ON THE HANDLE, BUT WHEN THE RELEASE BUTTON WAS PRESSED, A CLICKING NOISE WAS HEARD FROM THE HANDLE. AFTER PRESSING THE HANDLE AGAIN, THERE WAS NO RESISTANCE, AND NOTHING HAPPENED. THE TISSUE WAS NOT CUT OR STAPLED. A NEW HANDLE WAS USED WITH THE OLD RELOAD, BUT THE ISSUE PERSISTED. THE RELOAD WAS NOT THE PROBLEM, AS IT OPENED AND CLOSED WITHOUT ANY ISSUES. THE PROCEDURE WAS COMPLETED USING A COMPETITOR'S PRODUCT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
28-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Device Slipped | A051204 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG WEDGE RESECTION PROCEDURE INVOLVING FOUR RELOADS AND THREE MANUAL HANDLES, THE RELOAD WAS PLACED ON THE HANDLE, BUT WHEN THE RELEASE BUTTON WAS PRESSED, A CLICKING NOISE WAS HEARD FROM THE HANDLE. AFTER PRESSING THE HANDLE AGAIN, THERE WAS NO RESISTANCE, AND NOTHING HAPPENED. THE TISSUE WAS NOT CUT OR STAPLED. A NEW HANDLE WAS USED WITH THE OLD RELOAD, BUT THE ISSUE PERSISTED. THE RELOAD WAS NOT THE PROBLEM, AS IT OPENED AND CLOSED WITHOUT ANY ISSUES. THE PROCEDURE WAS COMPLETED USING A COMPETITOR'S PRODUCT. MEDTRONIC'S INITIAL EVALUATION OF THE RETURNED PRODUCT FOUND SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
06-12-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN RESPIRATORY SURGERY, THE CARTRIDGE COULD NOT BE LOADED INTO THE JAW AND REMOVED AT THE 2ND FIRING. THE SALES REP THINKS THAT IT MAY NOT BE POSSIBLE TO LOAD PHYSICALLY. THE DEVICE WAS NOT USED FOR THE PATIENT. THE CARTRIDGE COLOR WAS WHITE. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED PLUS COMPACT 45 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE KNIFE STUCK. THERE WERE NO PATIENT CONSEQUENCES.
10-01-2025 | Malfunction | covidien lp llc north haven | 27-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE USE OF AN ARTICULATING MEDIUM THICK RELOAD, THE JAWS OF THE RELOAD BROKE. THE DEVICE WAS REPLACED WITH ANOTHER RELOAD TO RESOLVE THE ISSUE. THERE WERE NO REPORTED CONSEQUENCES OR IMPACT TO THE PATIENT.
01-01-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED PLUS STANDARD 45 | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE KNIFE STUCK. THERE WERE NO PATIENT CONSEQUENCES.
24-01-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, IN A LAPAROSCOPICALLY ASSISTED DISTAL GASTRECTOMY, DURING GASTRECTOMY, THE FIRING COULD NOT BE PERFORMED AT ALL. ANOTHER RELOAD WAS USED WITH THE SAME HANDLE TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY.
13-11-2024 | Malfunction | ethicon endo-surgery | 27-01-2025 | ETS FLEX ARTICNG LNR CUTR 45MM | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE MISFIRED.
08-01-2025 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY SURGERY, THE JAWS COULD NOT BE OPENED. THERE WAS NOT STAPLED, AND THE JAWS RELEASED WITH THE KNIFE REVERSE SWITCH. THERE WAS NO EFFECT ON THE PATIENT. THE CARTRIDGE COLOR WAS GREEN. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
08-01-2025 | Malfunction | ethicon endo-surgery | 27-01-2025 | POWERED 60 ECHELON +, 280MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY SURGERY, THE JAWS COULD NOT BE OPENED. THERE WAS NOT STAPLED, AND THE JAWS RELEASED WITH THE KNIFE REVERSE SWITCH. THERE WAS NO EFFECT ON THE PATIENT. THE CARTRIDGE COLOR WAS GREEN. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
13-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS DIFFICULT TO FIRE. THE SURGERY WAS COMPLETED WITHOUT ANY PROBLEMS AFTER USING A NEW PRODUCT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
13-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Firing Problem | A0505 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE WAS DIFFICULT TO FIRE. THE SURGERY WAS COMPLETED WITHOUT ANY PROBLEMS AFTER USING A NEW PRODUCT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED.
29-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP, THERE WERE ISSUES WITH THE HANDLE. SPECIFICALLY, THERE WAS DIFFICULTY OPENING THE JAWS. THE JAWS WOULD NOT OPEN SMOOTHLY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
29-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE INVOLVING THE EGIAUSTND ENDOGIA ULTRA UNIV STD STAP, THERE WERE ISSUES WITH THE HANDLE. SPECIFICALLY, THERE WAS DIFFICULTY OPENING THE JAWS. THE JAWS WOULD NOT OPEN SMOOTHLY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE SHEARED TEETH WERE VISIBLE ON THE FIRING RACK.
03-01-2025 | Malfunction | intuitive surgical | 27-01-2025 | SUREFORM | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.
03-12-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLER WAS DIFFICULT TO FIRE. THE HANDLE WAS REPLACED WITH A NEW ONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS HANDLE THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
03-12-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE STAPLER WAS DIFFICULT TO FIRE. THE HANDLE WAS REPLACED WITH A NEW ONE TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS HANDLE THICK TISSUE. THE INSTRUMENT LOADED, CLAMPED, YET WOULD NOT CYCLE DUE TO THE SHEARED FIRING RACK TEETH DAMAGE. AN ACCESS HOLE WAS CUT INTO THE INSTRUMENT BODY FOR VISUALIZATION OF THE FIRING RACK. SHEARED TEETH WERE VISIBLE ON THE FIRING RACK AT SITE OF INITIAL FIRING POST CLAMP UP.
16-11-2024 | Malfunction | covidien | 27-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
16-11-2024 | Malfunction | covidien | 27-01-2025 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
16-11-2024 | Malfunction | covidien | 27-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
16-11-2024 | Malfunction | covidien | 27-01-2025 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT.
16-11-2024 | Malfunction | covidien | 27-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. THE CLAMPING MECHANISMS WERE DEFORMED. STAPLE PUSHERS WERE PARTIALLY FLUSH WITH THE CARTRIDGE.
16-11-2024 | Malfunction | covidien | 27-01-2025 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. THE CLAMPING MECHANISMS WERE DEFORMED. STAPLE PUSHERS WERE PARTIALLY FLUSH WITH THE CARTRIDGE.
18-03-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION, THE JAWS OF THE DEVICE WOULD NOT OPEN WHEN MANIPULATED FOR TESTING. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK DUE TO THICK TISSUE.
18-03-2025 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA ULTRA | Mechanics Altered | A0507 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LOW ANTERIOR RESECTION, THE JAWS OF THE DEVICE WOULD NOT OPEN WHEN MANIPULATED FOR TESTING. ANOTHER DEVICE WAS USED TO RESOLVE THE ISSUE. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND INTERNAL COMPONENTS REVEALED SHEARED TEETH ON THE FIRING RACK DUE TO THICK TISSUE.
16-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. THE CLAMPING MECHANISMS WERE DEFORMED. STAPLE PUSHERS WERE PARTIALLY FLUSH WITH THE CARTRIDGE.
16-11-2024 | Malfunction | us surgical puerto rico | 27-01-2025 | ENDO GIA | Mechanics Altered | A0507 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, WHEN SETTING UP, THE RELOADS HAD AN UNKNOWN ISSUE. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THERE WAS EGIA THICK TISSUE. THE CLAMPING MECHANISMS WERE DEFORMED. STAPLE PUSHERS WERE PARTIALLY FLUSH WITH THE CARTRIDGE.
03-01-2025 | Malfunction | intuitive surgical | 27-01-2025 | SUREFORM | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.
26-12-2024 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EGIAUSHORT ENDOGIA ULTRA UNIV SHT STAP, THE DEVICE DID NOT FIRE. THE HANDLE WAS REPLACED WITH A NEW ONE. THERE WAS NO PATIENT INJURY.
26-12-2024 | Malfunction | covidien lp llc north haven | 28-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EGIAUSHORT ENDOGIA ULTRA UNIV SHT STAP, THE DEVICE DID NOT FIRE. THE DEVICE WAS REPLACED WITH A NEW ONE. THERE WAS NO PATIENT INJURY.
26-12-2024 | Malfunction | covidien lp llc north haven | 28-01-2025 | UNKNOWN ENDO GIA SULU | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EGIAUSHORT ENDOGIA ULTRA UNIV SHT STAP, THE DEVICE DID NOT FIRE. THE DEVICE WAS REPLACED WITH A NEW ONE. THERE WAS NO PATIENT INJURY.
19-12-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE CONTOUR STAPLER WAS NOT ABLE TO BE PROPERLY FIRED, THE DEVICE WOULD NOT CLOSE DOWN PROPERLY ON THE TISSUE NOR ALLOW THE SURGEON TO FULLY FIRE THE STAPLES. NO PATIENT CONSEQUENCE WAS REPORTED.
19-12-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THAT THE CONTOUR STAPLER WAS NOT ABLE TO BE PROPERLY FIRED, THE DEVICE WOULD NOT CLOSE DOWN PROPERLY ON THE TISSUE NOR ALLOW THE SURGEON TO FULLY FIRE THE STAPLES. NO PATIENT CONSEQUENCE WAS REPORTED.
14-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | CONTOUR CURVED CUTTER STAPLER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | CONTOUR CURVED CUTTER STAPLER | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | CONTOUR CURVED CUTTER STAPLER | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ILS 25MM, CURVED | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
14-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ILS 25MM, CURVED | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTUM PROCEDURE SURGERY. AFTER FIRED, NOTED THAT A FEW STAPLES FORMED WITH IRREGULAR SHAPES AND ANASTOMOTIC SITE WAS OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WAS NO PATIENT CONSEQUENCE REPORTED, NO ADDITIONAL INFORMATION COULD BE PROVIDED.
17-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | CIRCULAR MECH STD 25MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, CDH25B WAS USED. USUALLY THE DOCTOR USES PCDH, BUT THE NURSE OPENED CDH25B BY MISTAKE AND PCDH WAS USED. IT WAS USED ON RECTUM. THE DOCTOR DID NOT GRIP THE FIRE RING UNTIL THE END BECAUSE THE DOCTOR WHO FIRED THE ANASTOMOTIC WAS NOT USED TO USING IT. THE STAPLE WAS MALFORMED AND FAILED TO DISSECT. SINCE THE RECTUM COULD NOT BE DISSECTED, THE RECTUM WAS DISSECTED BY TRANSANAL ANASTOMOSIS. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
17-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | CIRCULAR MECH STD 25MM | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, CDH25B WAS USED. USUALLY THE DOCTOR USES PCDH, BUT THE NURSE OPENED CDH25B BY MISTAKE AND PCDH WAS USED. IT WAS USED ON RECTUM. THE DOCTOR DID NOT GRIP THE FIRE RING UNTIL THE END BECAUSE THE DOCTOR WHO FIRED THE ANASTOMOTIC WAS NOT USED TO USING IT. THE STAPLE WAS MALFORMED AND FAILED TO DISSECT. SINCE THE RECTUM COULD NOT BE DISSECTED, THE RECTUM WAS DISSECTED BY TRANSANAL ANASTOMOSIS. ADDITIONAL HAND SUTURE WAS PERFORMED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
03-12-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Failure to Read Input Signal | A1301 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN THEY WERE UNABLE TO NAIL IN CASE OF POWER FAILURE DURING OPERATION. THERE WERE NO PATIENT CONSEQUENCES.
07-01-2025 | Malfunction | covidien | 28-01-2025 | ENDO GIA ULTRA |  |  | No Clinical Signs, Symptoms or Conditions | Event Description: PER SURGEON'S NOTES: AFTER APPROPRIATE PREOPERATIVE WORKUP WAS PERFORMED AND AFTER INFORMED CONSENT WAS OBTAINED FROM THE PATIENT, HE WAS TAKEN TO THE OPERATING ROOM AND PLACED IN THE SUPINE POSITION. A TIMEOUT WAS COMPLETED VERIFYING CORRECT PATIENT, SITE AND OPERATIVE PROCEDURE. GENERAL ENDOTRACHEAL TUBE ANESTHESIA WAS ADMINISTERED. AN ARTERIAL LINE AND A FOLEY CATHETER WERE PLACED. BRONCHOSCOPY WAS INITIALLY PERFORMED. ALL PRESSURE POINTS WERE ADEQUATELY PROTECTED AND PADDED. THE ENTIRE LEFT CHEST WAS THEN PREPPED AND DRAPED IN THE USUAL STERILE FASHION. WEDGE RESECTION WAS THEN INITIATED. WHEN IT WAS TIME TO USE THE ENDOGIA STAPLER, THE HANDLE DID NOT CLOSE/ CLIP THE STAPLE. THE TEAM HAD TO OPEN ANOTHER STAPLE HANDLE IN ORDER TO USE THE STAPLE CLIP (RELOAD). THE PROCEDURE WAS COMPLETED WITHOUT ANY FURTHER INCIDENT. NO PATIENT HARM. MANUFACTURER RESPONSE FOR STAPLE, IMPLANTABLE, ENDO GIA ULTRA (PER SITE REPORTER) FIELD MED REPRESENTATIVE WAS INFORMED BY OR SUPERVISOR AND HE WILL COME AND PICK UP THE DEVICE IN THE NEAR FUTURE.
02-12-2024 | Injury | covidien mfg dc boulder | 28-01-2025 | UNKNOWN LIGASURE INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Hemorrhage/Blood Loss/Bleeding | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY REVIEWED 88 CASES OF FAILED PRIMARY RESTRICTIVE PROCEDURES THAT SUBSEQUENTLY WERE CANDIDATES FOR BARIATRIC SURGERY REVISION BETWEEN 2021 AND 2023. PATIENTS UNDERWENT EITHER ROUX-EN-Y GASTRIC BYPASS OR ONE-ANASTOMOSIS GASTRIC BYPASS. PROCEDURES WERE SIMILAR BETWEEN THE TWO GROUPS: A LIGASURE WAS USED FOR DISSECTION AND THE ENDO GIA WAS USED TO STAPLE THE STOMACH. ENDO GIA OR VLOC WERE USED FOR ANASTOMOSES. POTENTIAL DEVICE RELATED POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAKAGE, BLEEDING, AND SUBPHRENIC COLLECTION. LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS VERSUS ONE ANASTOMOSIS GASTRIC BYPASS FOR REVISIONAL BARIATRIC SURGERY: A PROPENSITY SCORE MATCHED STUDY. DR. (B)(6) SALMAN, 2024, BARIATRIC SURGICAL PRACTICE AND PATIENT CARE, 10.1089/BARI.2023.0060.
02-12-2024 | Injury | covidien mfg dc boulder | 28-01-2025 | UNKNOWN LIGASURE INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Fluid Discharge | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY REVIEWED 88 CASES OF FAILED PRIMARY RESTRICTIVE PROCEDURES THAT SUBSEQUENTLY WERE CANDIDATES FOR BARIATRIC SURGERY REVISION BETWEEN 2021 AND 2023. PATIENTS UNDERWENT EITHER ROUX-EN-Y GASTRIC BYPASS OR ONE-ANASTOMOSIS GASTRIC BYPASS. PROCEDURES WERE SIMILAR BETWEEN THE TWO GROUPS: A LIGASURE WAS USED FOR DISSECTION AND THE ENDO GIA WAS USED TO STAPLE THE STOMACH. ENDO GIA OR VLOC WERE USED FOR ANASTOMOSES. POTENTIAL DEVICE RELATED POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAKAGE, BLEEDING, AND SUBPHRENIC COLLECTION. LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS VERSUS ONE ANASTOMOSIS GASTRIC BYPASS FOR REVISIONAL BARIATRIC SURGERY: A PROPENSITY SCORE MATCHED STUDY. DR. (B)(6) SALMAN, 2024, BARIATRIC SURGICAL PRACTICE AND PATIENT CARE, 10.1089/BARI.2023.0060.
02-12-2024 | Injury | covidien mfg dc boulder | 28-01-2025 | UNKNOWN LIGASURE INSTRUMENT | Adverse Event Without Identified Device or Use Problem | A24 | Appropriate Term / Code Not Available | Event Description: ACCORDING TO THE LITERATURE, A PROSPECTIVE STUDY REVIEWED 88 CASES OF FAILED PRIMARY RESTRICTIVE PROCEDURES THAT SUBSEQUENTLY WERE CANDIDATES FOR BARIATRIC SURGERY REVISION BETWEEN 2021 AND 2023. PATIENTS UNDERWENT EITHER ROUX-EN-Y GASTRIC BYPASS OR ONE-ANASTOMOSIS GASTRIC BYPASS. PROCEDURES WERE SIMILAR BETWEEN THE TWO GROUPS: A LIGASURE WAS USED FOR DISSECTION AND THE ENDO GIA WAS USED TO STAPLE THE STOMACH. ENDO GIA OR VLOC WERE USED FOR ANASTOMOSES. POTENTIAL DEVICE RELATED POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAKAGE, BLEEDING, AND SUBPHRENIC COLLECTION. LAPAROSCOPIC ROUX-EN-Y GASTRIC BYPASS VERSUS ONE ANASTOMOSIS GASTRIC BYPASS FOR REVISIONAL BARIATRIC SURGERY: A PROPENSITY SCORE MATCHED STUDY. DR. (B)(6) SALMAN, 2024, BARIATRIC SURGICAL PRACTICE AND PATIENT CARE, 10.1089/BARI.2023.0060.
06-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION (LAP LAR) PROCEDURE INVOLVING THE UPPER COLON, WHEN THE JAW GOT THE TISSUE, THE JAWS WOULD NOT CLOSE COMPLETELY. TO RESOLVE THE PROBLEM, THE SURGEON REPLACED THE DEVICE WITH ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | UNKNOWN ENDO GIA SULU | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT LOBE RESECTION, DURING MAIN PULMONARY VEIN RESECTION, AFTER APPROACHING THE RIGHT PULMONARY VEIN, WHEN AN ATTEMPT WAS MADE TO RESECTION, TRIED TO SQUEEZE THE HANDLE AFTER SETTING IT TO FIRING MODE ACCORDING TO THE PROCEDURE, BUT THE RESISTANCE WAS STRONG AND THE FIRING COULD NOT BE PERFORMED. THE EVENT OCCURRED DURING THE SECOND AND THIRD SHOTS OF THE FIRING OPERATION DURING SURGERY. WHEN THE HANDLE AND THE CARTRIDGE WERE REPLACED TO A NEW ONES, WAS ABLE TO PERFORM FIRING WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
06-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION (LAP LAR) PROCEDURE INVOLVING THE UPPER COLON, WHEN THE JAW GOT THE TISSUE, THE JAWS WOULD NOT CLOSE COMPLETELY. TO RESOLVE THE PROBLEM, THE SURGEON REPLACED THE DEVICE WITH ANOTHER RELOAD. THEREWAS NO PATIENT INJURY.
07-01-2025 | Malfunction | us surgical puerto rico | 28-01-2025 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC RIGHT LOBE RESECTION, DURING MAIN PULMONARY VEIN RESECTION, AFTER APPROACHING THE RIGHT PULMONARY VEIN, WHEN AN ATTEMPT WAS MADE TO RESECTION, TRIED TO SQUEEZE THE HANDLE AFTER SETTING IT TO FIRING MODE ACCORDING TO THE PROCEDURE, BUT THE RESISTANCE WAS STRONG AND THE FIRING COULD NOT BE PERFORMED. THE EVENT OCCURRED DURING THE SECOND AND THIRD SHOTS OF THE FIRING OPERATION DURING SURGERY. WHEN THE HANDLE AND THE CARTRIDGE WERE REPLACED TO A NEW ONES, WAS ABLE TO PERFORM FIRING WITHOUT ANY PROBLEMS. THERE WAS NO PATIENT INJURY.
06-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | UNK-SIGADAPT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOWER ANTERIOR RESECTION (LAP LAR) PROCEDURE INVOLVING THE UPPER COLON, WHEN THE JAW GOT THE TISSUE, THE JAWS WOULD NOT CLOSE COMPLETELY. TO RESOLVE THE PROBLEM, THE SURGEON REPLACED THE DEVICE WITH ANOTHER RELOAD. THERE WAS NO PATIENT INJURY.
10-10-2024 | Malfunction | us surgical puerto rico | 28-01-2025 | UNKNOWN ENDO GIA INSTRUMENT | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THERE WERE FIRING ISSUES WITH THE RELOAD COMPONENT OF THE STAPLER, THE RELOAD WAS NOT TRIGGER WHEN THE STAPLER WAS ACTIVATED.
10-10-2024 | Malfunction | us surgical puerto rico | 28-01-2025 | UNKNOWN ENDO GIA INSTRUMENT | No Apparent Adverse Event | A25 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, THERE WERE FIRING ISSUES WITH THE RELOAD COMPONENT OF THE STAPLER, THE RELOAD WAS NOT TRIGGER WHEN THE STAPLER WAS ACTIVATED.
08-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | GST GOLD RELOAD 60MM 6 ROW | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, AFTER LOADING THE CARTRIDGE, THE CARTRIDGE FELL OUT BEFORE THE TISSUE WAS CLAMPED, SO THE CARTRIDGE WAS REPLACED WITH A NEW ONE AND THE SURGERY WAS SUCCESSFULLY COMPLETED. THERE WAS NO EFFECT ON THE PATIENT. THE CARTRIDGE COLOR WAS GOLD. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED.
03-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | POWERED PLUS STANDARD 45 | Unintended Ejection | A050302 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, THE RELOAD FELL COMPLETELY APART AFTER FIRING THE FIRST TIME. NOTHING FELL INTO THE PATIENT. CASE WAS COMPLETED WITH A LIKE DEVICE. NO INFORMATION WAS SAVED ABOUT THE RELOAD. NO ADDITIONAL INFORMATION PROVIDED.
09-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | GST BLACK RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE NOTED THE STAPLES MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED.
04-11-2024 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER PRIOR TO USE, THE STAPLER WAS DIFFICULT TO FIRE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED.
04-11-2024 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER PRIOR TO USE, THE STAPLER WAS DIFFICULT TO FIRE. ANOTHER RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT THE CLAMPING MECHANISM WAS DEFORMED.
07-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON FLEX 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE, THE SURGEON TRIED TO CLAMP DOWN AND THEY MAYBE CLOSED ON A CLIP AND THEN ATTEMPTED TO FIRE THE DEVICE. THE DEVICE LOCKED OUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. DURING THE CASE THAT SECOND DEVICE LOCKED OUT. CASE WAS COMPLETED WITH A THIRD DEVICE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
07-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON FLEX 45 STANDARD | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE, THE SURGEON TRIED TO CLAMP DOWN AND THEY MAYBE CLOSED ON A CLIP AND THEN ATTEMPTED TO FIRE THE DEVICE. THE DEVICE LOCKED OUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. DURING THE CASE THAT SECOND DEVICE LOCKED OUT. CASE WAS COMPLETED WITH A THIRD DEVICE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
07-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON FLEX 45 STANDARD | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN WHIPPLE PROCEDURE, THE SURGEON TRIED TO CLAMP DOWN AND THEY MAYBE CLOSED ON A CLIP AND THEN ATTEMPTED TO FIRE THE DEVICE. THE DEVICE LOCKED OUT. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. DURING THE CASE THAT SECOND DEVICE LOCKED OUT. CASE WAS COMPLETED WITH A THIRD DEVICE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
08-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SEGMENTAL RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS LOCKED OUT DURING THE PULMONARY ARTERY PROCESS. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO DEVICE WILL BE RETURNING. ADDITIONAL INFORMATION PROVIDED.
08-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SEGMENTAL RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WAS LOCKED OUT DURING THE PULMONARY ARTERY PROCESS. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO DEVICE WILL BE RETURNING. ADDITIONAL INFORMATION PROVIDED.
13-01-2025 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC DUODENECTOMY, FOR BOTH PAIR OF DEVICE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE HANDLE COULD NOT BE SQUEEZED. BOTH PAIR OF DEVICE WAS NOT ABLE TO FIRE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2025 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC DUODENECTOMY, FOR BOTH PAIR OF DEVICE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE HANDLE COULD NOT BE SQUEEZED. BOTH PAIR OF DEVICE WAS NOT ABLE TO FIRE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2025 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC DUODENECTOMY, FOR BOTH PAIR OF DEVICES, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE HANDLE COULD NOT BE SQUEEZED. BOTH PAIR OF DEVICES WAS NOT ABLE TO FIRE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2025 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATIC DUODENECTOMY, FOR BOTH PAIR OF DEVICES, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON, HOWEVER THE HANDLE COULD NOT BE SQUEEZED. BOTH PAIR OF DEVICES WAS NOT ABLE TO FIRE. ANOTHER DEVICE FROM ANOTHER MANUFACTURER WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE STARTED THE RELEASE PROCESS AND THEN STOPPED DURING THE THIRD LOADING UNIT - THE STAPLER COULD NOT BE RELEASED ANY FURTHER. WITH THE HELP OF THE RESERVE BUTTON, THE BLADE COULD BE MOVED BACK AND THE DEVICE OPENED. A NEW STAPLER WAS OPENED AND THE OPERATION CONTINUED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
03-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE STARTED THE RELEASE PROCESS AND THEN STOPPED DURING THE THIRD LOADING UNIT - THE STAPLER COULD NOT BE RELEASED ANY FURTHER. WITH THE HELP OF THE RESERVE BUTTON, THE BLADE COULD BE MOVED BACK AND THE DEVICE OPENED. A NEW STAPLER WAS OPENED AND THE OPERATION CONTINUED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED.
01-01-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON CONTOUR W/ GREEN RELOAD | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOID RESECTION THE GUIDE ROD BENT. DEVICE COULD NOT BE CLOSED. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED.
08-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | FLEX 60 ARTICULATING | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY (VATS) PROCEDURE, IT WAS OBSERVED THAT THE JAWS ON THE DEVICE WERE NOT ABLE TO OPEN PROPERLY SO THE DEVICE COULD NOT BE ABLE TO USE. THERE WAS NO PATIENT CONSEQUENCE. ADDITIONAL INFORMATION REQUESTED.
11-12-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SEMI-OPEN SURGERY, THE DEVICE COULD NOT BE FIRED AND THE FIRING KNOB DID NOT MOVE FORWARD AT THE 3RD FIRING FOR FEEA. SO, A NEW CARTRIDGE WAS TAKEN OUT TO COMPLETE THE RECONSTRUCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
11-12-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SEMI-OPEN SURGERY, THE DEVICE COULD NOT BE FIRED AND THE FIRING KNOB DID NOT MOVE FORWARD AT THE 3RD FIRING FOR FEEA. SO, A NEW CARTRIDGE WAS TAKEN OUT TO COMPLETE THE RECONSTRUCTION. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
21-10-2024 | Malfunction | covidien lp llc north haven | 28-01-2025 | ABSORBATACK | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC REPAIR OF AN INCISIONAL HERNIA WITH COMPOSITE MESH, SOME TACKS FELL INTO THE PERITONEUM OF THE PATIENT AND WERE RETRIEVED USING AN INSTRUMENT. IT WAS ALSO NOTED THAT THERE WERE PERSISTENT FAILURE TO DEPLOY THE TACKS DESPITE MULTIPLE ATTEMPTS. ADDITIONALLY, THE TACKER SEEMS TO HAD A CRACKED AT THE MIDDLE PART. THE TACKER WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
21-10-2024 | Malfunction | covidien lp llc north haven | 28-01-2025 | ABSORBATACK | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC REPAIR OF AN INCISIONAL HERNIA WITH COMPOSITE MESH, SOME TACKS FELL INTO THE PERITONEUM OF THE PATIENT AND WERE RETRIEVED USING AN INSTRUMENT. IT WAS ALSO NOTED THAT THERE WERE PERSISTENT FAILURE TO DEPLOY THE TACKS DESPITE MULTIPLE ATTEMPTS. ADDITIONALLY, THE TACKER SEEMS TO HAD A CRACKED AT THE MIDDLE PART. THE TACKER WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
21-10-2024 | Malfunction | covidien lp llc north haven | 28-01-2025 | ABSORBATACK | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC REPAIR OF AN INCISIONAL HERNIA WITH COMPOSITE MESH, SOME TACKS FELL INTO THE PERITONEUM OF THE PATIENT AND WERE RETRIEVED USING AN INSTRUMENT. IT WAS ALSO NOTED THAT THERE WERE PERSISTENT FAILURE TO DEPLOY THE TACKS DESPITE MULTIPLE ATTEMPTS. ADDITIONALLY, THE TACKER SEEMS TO HAD A CRACKED AT THE MIDDLE PART. THE TACKER WAS REPLACED TO RESOLVE THE ISSUE AND COMPLETE THE PROCEDURE. THERE WAS NO PATIENT INJURY.
17-10-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | 75 MM NTLC SELECTABLE RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE DID NOT FIRE. THERE WERE NO PATIENT CONSEQUENCES.
10-10-2024 | Malfunction | covidien | 28-01-2025 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: IT WAS REPORTED THAT DURING THE LAPAROSCOPIC RETROPERITONEAL LYMPHADENECTOMY, THERE WERE FIRING ISSUES WITH THE RELOAD COMPONENT OF THE STAPLER. THE RELOAD WAS NOT TRIGGERED WHEN THE STAPLER WAS ACTIVATED. THE RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
07-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE CLAMP DID NOT OPEN WHEN INTRODUCED THROUGH THE COELIOSCOPY TROCAR, NECESSITATING THE REMOVAL AND REPLACEMENT OF THE RELOAD. THE SURGEON WAS ABLE TO PRESS THE GREEN SAFETY BUTTON AND SQUEEZE THE HANDLE. THE SURGEON REMOVED THE RELOAD, PERFORMED A STAPLE/CUT "IN THE VOID" TO REMOVE THE RELOAD, AND THEN USED A NEW RELOAD, WHICH FUNCTIONED CORRECTLY WITHOUT FURTHER ISSUES. THERE WAS NO PATIENT INJURY.
13-01-2025 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, THE TWO STAPLERS DID NOT FIRE WHEN RESECTING RECTAL TISSUE. A NEW DEVICE WITH ANOTHER MANUFACTURER'S WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2025 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, THE TWO STAPLERS DID NOT FIRE WHEN RESECTING RECTAL TISSUE. A NEW DEVICE WITH ANOTHER MANUFACTURER'S WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2025 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, THE TWO STAPLERS DID NOT FIRE WHEN RESECTING RECTAL TISSUE. A NEW DEVICE WITH ANOTHER MANUFACTURER'S WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
13-01-2025 | Malfunction | us surgical puerto rico | 28-01-2025 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER DURING A LAPAROSCOPIC RADICAL RESECTION OF RECTAL CANCER, THE TWO STAPLERS DID NOT FIRE WHEN RESECTING RECTAL TISSUE. A NEW DEVICE WITH ANOTHER MANUFACTURER'S WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
06-01-2025 | Malfunction | intuitive surgical | 28-01-2025 | ENDOWRIST | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY SURGICAL PROCEDURE, THE CUSTOMER WAS ATTEMPTING TO STAPLER THE BRONCHUS USING THE STAPLER 30 INSTRUMENT AND GREEN RELOAD. DURING FIRING, THE INSTRUMENT STOPPED WITH A MESSAGE ¿REMOVE STAPLER WARNING BLADE MAY BE EXPOSED." SUBSEQUENTLY, A MESSAGE APPEARED ¿UNABLE TO COMPLETE FIRE. TAP BLUE PEDAL TO UNCLAMP THEN REMOVE STAPLER. WARNING BLADE MAY BE EXPOSED." THEREFORE, THE CUSTOMER REMOVED THE STAPLER AND REPLACED IT WITH A SUREFORM 60 GREEN RELOAD AND STARTED STAPLING. IT WAS OBSERVED THAT SOME OF THE TITANIUM STAPLES WERE STUCK IN THE STAPLER/RELOAD WITHOUT BEING ENGAGED. NO FRAGMENT FELL INTO THE PATIENT. THE CUSTOMER COMPLETED THE PROCEDURE, AND NO KNOWN IMPACT OR PATIENT CONSEQUENCE WAS REPORTED. INTUITIVE SURGICAL, INC. (ISI) FOLLOWED UP WITH THE INITIAL REPORTER AND OBTAINED THE FOLLOWING ADDITIONAL INFORMATION: THE ISSUE OCCURRED ON THE FIRST STAPLER FIRE. BRONCHUS STAPLING WAS BEING PERFORMED AT THE TIME OF THE EVENT. THE BRONCHUS WAS THE TARGET TISSUE. THE TISSUE WAS NOT CALCIFIED. THE TISSUE WAS NOT EXPOSED TO ANY RADIATION OR CHEMOTHERAPY PRIOR TO THE PROCEDURE. THERE WAS NOT ANY TISSUE TENSION. THERE WAS NOT ANY TISSUE BUNCHING. THE ISSUE INVOLVED A PARTIAL FIRE. THE ISSUE RESULTED IN UNFORMED STAPLES THAT WERE IN THE ¿U¿ SHAPE/FORM. THE SURGEON DID NOT ENCOUNTER ANY OBSTRUCTIONS SUCH AS CLIPS, STAPLES, OR OTHER HARD MATERIAL BETWEEN THE INSTRUMENT JAWS. THE SURGEON DID NOT FIRE OVER AN EXISTING STAPLE LINE. THE BUTTRESS MATERIAL WAS NOT USED. THE SURGEON DID NOT EXPERIENCE ANY CLAMPING ISSUES PRIOR TO FIRING THE STAPLER RELOAD. THERE WAS NOT ANY BLEEDING OBSERVED ON THE STAPLE LINE DUE TO THE STAPLER ISSUE. THERE WAS NO UNPLANNED TISSUE REMOVAL DUE TO THE STAPLER FIRING FAILURE.
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY BYPASS, POWERED STAPLER FIRED THE LOAD BUT THE BLADE DID NOT RETRACT NOR OPEN THE LOAD. THE LOAD OPENING BUTTON ALSO DID NOT WORK. THE SURGEON TRIED THE RETRACTION KEY TO OPEN THE LOAD, BUT THAT DID NOT WORK EITHER. THE LOAD REMAINED STUCK ON THE PATIENT'S TISSUE, THIS RESULTED TO AN HOUR PROLONGED SURGERY TIME. AS A RESOLUTION, THE SHELL WAS REPLACED AND SEVERAL STAPLES WERE USED AROUND THE STUCK LOAD TO RELEASE THE TISSUE.
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY BYPASS, POWERED STAPLER FIRED THE LOAD BUT THE BLADE DID NOT RETRACT NOR OPEN THE LOAD. THE LOAD OPENING BUTTON ALSO DID NOT WORK. THE SURGEON TRIED THE RETRACTION KEY TO OPEN THE LOAD, BUT THAT DID NOT WORK EITHER. THE LOAD REMAINED STUCK ON THE PATIENT'S TISSUE, THIS RESULTED TO AN HOUR PROLONGED SURGERY TIME. AS A RESOLUTION, THE SHELL WAS REPLACED AND SEVERAL STAPLES WERE USED AROUND THE STUCK LOAD TO RELEASE THE TISSUE.
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY BYPASS, POWERED STAPLER FIRED THE LOAD BUT THE BLADE DID NOT RETRACT NOR OPEN THE LOAD. THE LOAD OPENING BUTTON ALSO DID NOT WORK. THE SURGEON TRIED THE RETRACTION KEY TO OPEN THE LOAD, BUT THAT DID NOT WORK EITHER. THE LOAD REMAINED STUCK ON THE PATIENT'S TISSUE, THIS RESULTED TO AN HOUR PROLONGED SURGERY TIME. AS A RESOLUTION, THE SHELL WAS REPLACED AND SEVERAL STAPLES WERE USED AROUND THE STUCK LOAD TO RELEASE THE TISSUE.
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: DUPLICATE OF PE (B)(4).
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: DUPLICATE OF PE (B)(4).
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | SIGNIA | No Apparent Adverse Event | A25 | Tissue Breakdown | Event Description: DUPLICATE OF PE (B)(4).
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY BYPASS, POWERED STAPLER FIRED THE LOAD BUT THE BLADE DID NOT RETRACT NOR OPEN THE LOAD. THE LOAD OPENING BUTTON ALSO DID NOT WORK. THE SURGEON TRIED THE RETRACTION KEY TO OPEN THE LOAD, BUT THAT DID NOT WORK EITHER. THE LOAD REMAINED STUCK ON THE PATIENT'S TISSUE, THIS RESULTED TO AN HOUR PROLONGED SURGERY TIME. AS A RESOLUTION, THE SHELL WAS REPLACED AND SEVERAL STAPLES WERE USED AROUND THE STUCK LOAD TO RELEASE THE TISSUE.
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY BYPASS, POWERED STAPLER FIRED THE LOAD BUT THE BLADE DID NOT RETRACT NOR OPEN THE LOAD. THE LOAD OPENING BUTTON ALSO DID NOT WORK. THE SURGEON TRIED THE RETRACTION KEY TO OPEN THE LOAD, BUT THAT DID NOT WORK EITHER. THE LOAD REMAINED STUCK ON THE PATIENT'S TISSUE, THIS RESULTED TO AN HOUR PROLONGED SURGERY TIME. AS A RESOLUTION, THE SHELL WAS REPLACED AND SEVERAL STAPLES WERE USED AROUND THE STUCK LOAD TO RELEASE THE TISSUE.
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | ENDO GIA | No Apparent Adverse Event | A25 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING GASTROPLASTY BYPASS, POWERED STAPLER FIRED THE LOAD BUT THE BLADE DID NOT RETRACT NOR OPEN THE LOAD. THE LOAD OPENING BUTTON ALSO DID NOT WORK. THE SURGEON TRIED THE RETRACTION KEY TO OPEN THE LOAD, BUT THAT DID NOT WORK EITHER. THE LOAD REMAINED STUCK ON THE PATIENT'S TISSUE, THIS RESULTED TO AN HOUR PROLONGED SURGERY TIME. AS A RESOLUTION, THE SHELL WAS REPLACED AND SEVERAL STAPLES WERE USED AROUND THE STUCK LOAD TO RELEASE THE TISSUE.
17-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED DURING A RECTUM PROCEDURE THE STAPLER COULD NOT BE TRIGGERED. NO FUNCTION CONTACT ERROR OF THE BATTERY. NO PATIENT HARM NOR SIGNIFICANT DELAY REPORTED. PROCEDURE FINISHED WITH SAME LIKE DEVICE
16-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | TRI-STAPLE 2.0 | Device Displays Incorrect Message | A090201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE, AFTER PRESSING THE GREEN BUTTON, THE RELOAD DID NOT FIRE AS EXPECTED. THE SURGEON FOUND THAT THE PANEL FOR THE RELOAD WAS WHITE INDICATING THAT THE RELOAD HAD BEEN FIRED. A NEW RELOAD WAS USES TO RESOLVE THE ISSUE. SURGICAL TIME WAS EXTENDED BY 10 MINUTES. THERE WAS NO PATIENT INJURY.
16-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A GASTRIC SLEEVE, AFTER PRESSING THE GREEN BUTTON, THE RELOAD DID NOT FIRE AS EXPECTED. THE SURGEON FOUND THAT THE PANEL FOR THE RELOAD WAS WHITE INDICATING THAT THE RELOAD HAD BEEN FIRED. A NEW RELOAD WAS USES TO RESOLVE THE ISSUE. SURGICAL TIME WAS EXTENDED BY 10 MINUTES. THERE WAS NO PATIENT INJURY.
17-10-2024 | Malfunction | covidien lp llc north haven | 28-01-2025 | TRI-STAPLE 2.0 | Misfire | A050502 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED WITH THE SAME RELOAD. IT WAS NOTED THAT THE PROCEDURE TIME WAS EXTENDED FOR TEN MINUTES. THERE WAS NO PATIENT INJURY. PLI-20: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED. THE JAWS WERE OPEN. STAPLE PUSHERS WERE VISIBLE AT THE 2.2CM CUT LINE.
17-10-2024 | Malfunction | covidien lp llc north haven | 28-01-2025 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROUX-EN-Y PROCEDURE, THE SURGEON WAS ABLE TO PRESS THE GREEN BUTTON BUT COULD NOT SQUEEZE THE HANDLE. THE DEVICE DID NOT FIRE. TO RESOLVE THE ISSUE, A NEW HANDLE WAS USED WITH THE SAME RELOAD. IT WAS NOTED THAT THE PROCEDURE TIME WAS EXTENDED FOR TEN MINUTES. THERE WAS NO PATIENT INJURY. PLI-20: MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE RELOAD WAS PARTIALLY FIRED WITH THE INTERLOCK ENGAGED. THE JAWS WERE OPEN. STAPLE PUSHERS WERE VISIBLE AT THE 2.2CM CUT LINE.
03-12-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | CIRCULAR MECH STD 29MM | Failure to Fire | A050501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AFTER USING THE STAPLER, A LARGE ANASTOMOTIC INSUFFICIENCY OCCURRED BECAUSE THE STAPLER DID NOT CUT OR CLOSE THE ENTIRE BOWEL IN A CIRCULAR FASHION. NO PATIENT HARM AS THE INSUFFICIENCY COULD BE SUTURED OVER.
03-12-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | CIRCULAR MECH STD 29MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE AFTER USING THE STAPLER, A LARGE ANASTOMOTIC INSUFFICIENCY OCCURRED BECAUSE THE STAPLER DID NOT CUT OR CLOSE THE ENTIRE BOWEL IN A CIRCULAR FASHION. NO PATIENT HARM AS THE INSUFFICIENCY COULD BE SUTURED OVER.
16-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Difficult to Insert | A150206 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMI-COLECTOMY BY LAPAROTOMY AFTER THE FIRST STAPLING, IT WAS DIFFICULT TO REMOVE THE FIRST STAPLE ROW (ALREADY RELOADED ON THE DEVICE INSIDE THE BLISTER). RELOADING OF A NEW STAPLE RELOAD: IT WAS IMPOSSIBLE TO CLOSE THE DEVICE AND CUT. THE COLIC CUT WAS PERFORMED MANUALLY WITH A COLD BLADE. NO PATIENT CONSEQUENCE.
16-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMI-COLECTOMY BY LAPAROTOMY AFTER THE FIRST STAPLING, IT WAS DIFFICULT TO REMOVE THE FIRST STAPLE ROW (ALREADY RELOADED ON THE DEVICE INSIDE THE BLISTER). RELOADING OF A NEW STAPLE RELOAD: IT WAS IMPOSSIBLE TO CLOSE THE DEVICE AND CUT. THE COLIC CUT WAS PERFORMED MANUALLY WITH A COLD BLADE. NO PATIENT CONSEQUENCE.
16-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMI-COLECTOMY BY LAPAROTOMY AFTER THE FIRST STAPLING, IT WAS DIFFICULT TO REMOVE THE FIRST STAPLE ROW (ALREADY RELOADED ON THE DEVICE INSIDE THE BLISTER). RELOADING OF A NEW STAPLE RELOAD: IT WAS IMPOSSIBLE TO CLOSE THE DEVICE AND CUT. THE COLIC CUT WAS PERFORMED MANUALLY WITH A COLD BLADE. NO PATIENT CONSEQUENCE.
16-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LEFT HEMI-COLECTOMY BY LAPAROTOMY AFTER THE FIRST STAPLING, IT WAS DIFFICULT TO REMOVE THE FIRST STAPLE ROW (ALREADY RELOADED ON THE DEVICE INSIDE THE BLISTER). RELOADING OF A NEW STAPLE RELOAD: IT WAS IMPOSSIBLE TO CLOSE THE DEVICE AND CUT. THE COLIC CUT WAS PERFORMED MANUALLY WITH A COLD BLADE. NO PATIENT CONSEQUENCE.
06-12-2024 | Malfunction | covidien lp llc north haven | 28-01-2025 | GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE STAPLES DID NOT CLOSE, OR CLOSED PARTIALLY IN THE TISSUE, IN WHICH THE SURGEON HAD TO USE MORE STAPLES THAN NECESSARY. A NEW DEVICE WAS USED TO RESOLVE THE ISSUE.
22-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 23MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A GASTRIC PROCEDURE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
20-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCTOCOLECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, THE ISSUE WAS REPORTED BUT THE DETAILS ARE UNKNOWN AT THIS TIME, SO IT IS BEING CONFIRMED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | ILS 29MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
21-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | CIRCULAR MECH STD 29MM | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL ANASTOMOSIS AFTER STAPLING, THE COLLARS WERE COMPLETE. THE SURGEON DISCOVERED STAPLES UNCLOSED AT THE ANASTOMOSIS LEVEL SHOWING A POTENTIAL DEFAULT FROM THE DEVICE. INEFFICIENT ANASTOMOSIS. MANUAL REFECTION.
21-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | CIRCULAR MECH STD 29MM | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLORECTAL ANASTOMOSIS AFTER STAPLING, THE COLLARS WERE COMPLETE. THE SURGEON DISCOVERED STAPLES UNCLOSED AT THE ANASTOMOSIS LEVEL SHOWING A POTENTIAL DEFAULT FROM THE DEVICE. INEFFICIENT ANASTOMOSIS. MANUAL REFECTION.
23-01-2025 | Malfunction | covidien | 28-01-2025 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE FIRING COULD NOT BE DONE AT ALL. ANOTHER RELOAD WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
18-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PROXIMATE*75MM LINEAR CUTTER | Failure to Form Staple | A050704 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE FAILED TO DEPLOY THE STAPLER PROPERLY, THIS MADE THE ANASTOMOSIS FAIL. THIS MADE THE SURGERY LONGER THAN WAS ANTICIPATED, THEREFORE A PROLONGED ANESTHETIC AS WELL.
18-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PROXIMATE*75MM LINEAR CUTTER | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE DEVICE FAILED TO DEPLOY THE STAPLER PROPERLY, THIS MADE THE ANASTOMOSIS FAIL. THIS MADE THE SURGERY LONGER THAN WAS ANTICIPATED, THEREFORE A PROLONGED ANESTHETIC AS WELL.
16-09-2024 | Malfunction | ethicon endo-surgery | 28-01-2025 | 75MM SELECTABLE NEW TLC | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A WHIPPLE PROCEDURE THE DEVICE MISFIRED.
01-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Detachment of Device or Device Component | A0501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE DURING AN INTERVENTION (LEFT COLON), THE DEVICE WAS DEFECTIVE. RETRACTION AFTER STAPLING WAS NOT DONE CORRECTLY. AFTER A MANIPULATION BY THE SURGEON, AND THE EXTRACTION OF THE DEVICE, THE OPERATOR FOUND THAT A WHITE TIP HAS BECOME DETACHED (JUNCTION BETWEEN THE PROXIMAL AND DISTAL PART OF THE COLLARS) . FUNCTIONAL ELEMENT BEFORE ANAL INTRODUCTION. THIS RESULTED IN A NON-FUNCTIONAL ANASTOMOSIS, A MANUAL RESUMPTION AND THEREFORE A SURGICAL DELAY.
01-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Detachment of Device or Device Component | A0501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE DURING AN INTERVENTION (LEFT COLON), THE DEVICE WAS DEFECTIVE. RETRACTION AFTER STAPLING WAS NOT DONE CORRECTLY. AFTER A MANIPULATION BY THE SURGEON, AND THE EXTRACTION OF THE DEVICE, THE OPERATOR FOUND THAT A WHITE TIP HAS BECOME DETACHED (JUNCTION BETWEEN THE PROXIMAL AND DISTAL PART OF THE COLLARS) . FUNCTIONAL ELEMENT BEFORE ANAL INTRODUCTION. THIS RESULTED IN A NON-FUNCTIONAL ANASTOMOSIS, A MANUAL RESUMPTION AND THEREFORE A SURGICAL DELAY.
01-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Insufficient Information | E2401 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE DURING AN INTERVENTION (LEFT COLON), THE DEVICE WAS DEFECTIVE. RETRACTION AFTER STAPLING WAS NOT DONE CORRECTLY. AFTER A MANIPULATION BY THE SURGEON, AND THE EXTRACTION OF THE DEVICE, THE OPERATOR FOUND THAT A WHITE TIP HAS BECOME DETACHED (JUNCTION BETWEEN THE PROXIMAL AND DISTAL PART OF THE COLLARS) . FUNCTIONAL ELEMENT BEFORE ANAL INTRODUCTION. THIS RESULTED IN A NON-FUNCTIONAL ANASTOMOSIS, A MANUAL RESUMPTION AND THEREFORE A SURGICAL DELAY.
01-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Insufficient Information | E2401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE DURING AN INTERVENTION (LEFT COLON), THE DEVICE WAS DEFECTIVE. RETRACTION AFTER STAPLING WAS NOT DONE CORRECTLY. AFTER A MANIPULATION BY THE SURGEON, AND THE EXTRACTION OF THE DEVICE, THE OPERATOR FOUND THAT A WHITE TIP HAS BECOME DETACHED (JUNCTION BETWEEN THE PROXIMAL AND DISTAL PART OF THE COLLARS) . FUNCTIONAL ELEMENT BEFORE ANAL INTRODUCTION. THIS RESULTED IN A NON-FUNCTIONAL ANASTOMOSIS, A MANUAL RESUMPTION AND THEREFORE A SURGICAL DELAY.
01-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE DURING AN INTERVENTION (LEFT COLON), THE DEVICE WAS DEFECTIVE. RETRACTION AFTER STAPLING WAS NOT DONE CORRECTLY. AFTER A MANIPULATION BY THE SURGEON, AND THE EXTRACTION OF THE DEVICE, THE OPERATOR FOUND THAT A WHITE TIP HAS BECOME DETACHED (JUNCTION BETWEEN THE PROXIMAL AND DISTAL PART OF THE COLLARS) . FUNCTIONAL ELEMENT BEFORE ANAL INTRODUCTION. THIS RESULTED IN A NON-FUNCTIONAL ANASTOMOSIS, A MANUAL RESUMPTION AND THEREFORE A SURGICAL DELAY.
01-01-2025 | Malfunction | ethicon endo-surgery | 28-01-2025 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE DURING AN INTERVENTION (LEFT COLON), THE DEVICE WAS DEFECTIVE. RETRACTION AFTER STAPLING WAS NOT DONE CORRECTLY. AFTER A MANIPULATION BY THE SURGEON, AND THE EXTRACTION OF THE DEVICE, THE OPERATOR FOUND THAT A WHITE TIP HAS BECOME DETACHED (JUNCTION BETWEEN THE PROXIMAL AND DISTAL PART OF THE COLLARS) . FUNCTIONAL ELEMENT BEFORE ANAL INTRODUCTION. THIS RESULTED IN A NON-FUNCTIONAL ANASTOMOSIS, A MANUAL RESUMPTION AND THEREFORE A SURGICAL DELAY.
06-01-2025 | Injury | intuitive surgical | 28-01-2025 | SUREFORM | Firing Problem | A0505 | Unspecified Tissue Injury | Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY BYPASS, THE RELOAD WAS FIRED, BUT THE RELOAD'S CUTTING BLADE DID NOT RETRACT AND THE RELOAD WAS LOCKED ON TISSUE. THE RELOAD REMAINED CLOSED IN THE PATIENT'S TISSUE AND COULD NOT BE REMOVED. RESTARTED THE ENTIRE PROCESS USING THE MANUAL RETRACTION KEY, BUT IT DID NOT WORK. SEVERAL STAPLES WERE USED AROUND THE ENTIRE RELOAD TO BE ABLE TO RELEASE FROM THE TISSUE, AND RESECTION WAS DONE TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE DEVICE ISSUE.
13-01-2025 | Malfunction | covidien lp llc north haven | 28-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY BYPASS, THE RELOAD WAS FIRED, BUT THE RELOAD'S CUTTING BLADE DID NOT RETRACT AND THE RELOAD WAS LOCKED ON TISSUE. THE RELOAD REMAINED CLOSED IN THE PATIENT'S TISSUE AND COULD NOT BE REMOVED. RESTARTED THE ENTIRE PROCESS USING THE MANUAL RETRACTION KEY, BUT IT DID NOT WORK. SEVERAL STAPLES WERE USED AROUND THE ENTIRE RELOAD TO BE ABLE TO RELEASE FROM THE TISSUE, AND RESECTION WAS DONE TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE DEVICE ISSUE.
13-01-2025 | Malfunction | covidien surgical | 28-01-2025 | SIGNIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY BYPASS, THE RELOAD WAS FIRED, BUT THE RELOAD'S CUTTING BLADE DID NOT RETRACT AND THE RELOAD WAS LOCKED ON TISSUE. THE RELOAD REMAINED CLOSED IN THE PATIENT'S TISSUE AND COULD NOT BE REMOVED. RESTARTED THE ENTIRE PROCESS USING THE MANUAL RETRACTION KEY, BUT IT DID NOT WORK. SEVERAL STAPLES WERE USED AROUND THE ENTIRE RELOAD TO BE ABLE TO RELEASE FROM THE TISSUE, AND RESECTION WAS DONE TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE DEVICE ISSUE.
13-01-2025 | Malfunction | covidien surgical | 28-01-2025 | SIGNIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY BYPASS, THE RELOAD WAS FIRED, BUT THE RELOAD'S CUTTING BLADE DID NOT RETRACT AND THE RELOAD WAS LOCKED ON TISSUE. THE RELOAD REMAINED CLOSED IN THE PATIENT'S TISSUE AND COULD NOT BE REMOVED. RESTARTED THE ENTIRE PROCESS USING THE MANUAL RETRACTION KEY, BUT IT DID NOT WORK. SEVERAL STAPLES WERE USED AROUND THE ENTIRE RELOAD TO BE ABLE TO RELEASE FROM THE TISSUE, AND RESECTION WAS DONE TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE DEVICE ISSUE.
13-01-2025 | Malfunction | covidien | 28-01-2025 | ENDO GIA | Entrapment of Device | A150208 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY BYPASS, THE RELOAD WAS FIRED, BUT THE RELOAD'S CUTTING BLADE DID NOT RETRACT AND THE RELOAD WAS LOCKED ON TISSUE. THE RELOAD REMAINED CLOSED IN THE PATIENT'S TISSUE AND COULD NOT BE REMOVED. RESTARTED THE ENTIRE PROCESS USING THE MANUAL RETRACTION KEY, BUT IT DID NOT WORK. SEVERAL STAPLES WERE USED AROUND THE ENTIRE RELOAD TO BE ABLE TO RELEASE FROM THE TISSUE, AND RESECTION WAS DONE TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE DEVICE ISSUE.
13-01-2025 | Malfunction | covidien | 28-01-2025 | ENDO GIA | Retraction Problem | A0510 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A GASTROPLASTY BYPASS, THE RELOAD WAS FIRED, BUT THE RELOAD'S CUTTING BLADE DID NOT RETRACT AND THE RELOAD WAS LOCKED ON TISSUE. THE RELOAD REMAINED CLOSED IN THE PATIENT'S TISSUE AND COULD NOT BE REMOVED. RESTARTED THE ENTIRE PROCESS USING THE MANUAL RETRACTION KEY, BUT IT DID NOT WORK. SEVERAL STAPLES WERE USED AROUND THE ENTIRE RELOAD TO BE ABLE TO RELEASE FROM THE TISSUE, AND RESECTION WAS DONE TO RESOLVE THE ISSUE. THE SURGICAL TIME WAS EXTENDED FOR ONE HOUR DUE TO THE DEVICE ISSUE.
16-01-2025 | Malfunction | ethicon endo-surgery | 29-01-2025 | CIRCULAR MECH STD 33MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A COLECTOMY DURING THE STAPLING OF THE RECTUM, THERE WAS A FAILURE IN THE CIRCULAR STAPLER NUMBER 33. IT PERFORMED THE SECTION, BUT DID NOT STAPLE THE TISSUE, LEAVING THE INTESTINAL LOOP OPEN, DOCUMENTED IN THE PHOTO. FOR THIS REASON, ANOTHER STAPLER WAS REQUESTED TO PERFORM THE COLORECTAL ANASTOMOSIS. IMMEDIATE ITEM EXCHANGE, THERE WAS NO SURGICAL DELAY. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES.
01-01-2024 | Malfunction | ethicon endo-surgery | 29-01-2025 | 55MM SELECTABLE NEW TLC | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE ANVIL WAS DAMAGED.
01-01-2025 | Malfunction | ethicon endo-surgery | 29-01-2025 | ECHELON CONTOUR W/ BLUE RELOAD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE THE CONTOUR WAS USED TO TRANSECT THE RECTUM, PROTOCOL WAS FOLLOWED AND THE STAPLER SEEMED TO HAVE FIRED PROPERLY. AFTER FIRING, THE STAPLE LINE LOOKED OKAY BUT ONCE THE END TO END ANASTOMOSIS HAD BEEN CREATED, THE CONTOUR STAPLE LINE RELEASED. THE SURGEON NOTED THAT THE MUCOSA WAS VISIBLE AND THAT THE STAPLES WERE NOT COMPLETELY FORMED. THE ANASTOMOSES HAD TO BE REDONE BY HAND SEWING. THE SURGERY WAS DELAYED 60 MINUTES.
29-01-2025 | Malfunction | covidien lp llc north haven | 29-01-2025 | UNKNOWN ENDO GIA SULU | Difficult to Open or Close | A051104 | Insufficient Information | Event Description: IT WAS REPORTED THAT THE JAW DID NOT OPEN WHEN MANIPULATED FOR TESTING DURING SURGICAL PREPARATION.
09-01-2025 | Malfunction | covidien lp llc north haven | 29-01-2025 | SIGNIA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC GASTRECTOMY, WHEN CLOSING THE INSERTION PORT OF DELTA RECONSTRUCTION, THE RELOAD WAS ARTICULATED TO THE LEFT AND WAS FIRED. THE DEVICE WAS ARTICULATED SIGNIFICANTLY TO THE LEFT, MORE THAN THE USUAL TWITCHING. AFTER THAT, WHEN TRYING TO RETURN TO THE NEUTRAL POSITION, IT DID NOT BECOME STRAIGHT, SO THE CENTER CONTROL WAS PRESSED TO THE RIGHT, STRAIGHTENED IT AND REMOVED FROM THE TROCAR. THE SURGEON, FELT UNCOMFORTABLE, AND UPON OBSERVING THE ABDOMINAL CAVITY, DISCOVERED A METAL FRAGMENT DIRECTLY UNDER THE TROCAR FROM WHICH THE STAPLER WAS REMOVED. THE METAL FRAGMENTS WERE REMOVED FROM THE PATIENT. THE PATIENT WAS CHECKED WITH AN X-RAY AND NO OTHER METAL REACTIONS WERE OBSERVED AND SO THE PATIENT WAS CLOSED AS IS. WHEN CHECKING THE CARTRIDGE, IT APPEARED THAT THE INTERNAL PART OF THE ARTICULATED PART WAS DAMAGED.
06-01-2025 | Malfunction | ethicon endo-surgery | 29-01-2025 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-01-2025 | Malfunction | covidien lp llc north haven | 29-01-2025 | SIGNIA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A ROBOTIC GASTRECTOMY, WHEN CLOSING THE INSERTION PORT OF DELTA RECONSTRUCTION, THE RELOAD WAS ARTICULATED TO THE LEFT AND WAS FIRED. THE DEVICE WAS ARTICULATED SIGNIFICANTLY TO THE LEFT, MORE THAN THE USUAL TWITCHING. AFTER THAT, WHEN TRYING TO RETURN TO THE NEUTRAL POSITION, IT DID NOT BECOME STRAIGHT, SO THE CENTER CONTROL WAS PRESSED TO THE RIGHT, STRAIGHTENED IT AND REMOVED FROM THE TROCAR. THE SURGEON, FELT UNCOMFORTABLE, AND UPON OBSERVING THE ABDOMINAL CAVITY, DISCOVERED A METAL FRAGMENT DIRECTLY UNDER THE TROCAR FROM WHICH THE STAPLER WAS REMOVED. THE METAL FRAGMENTS WERE REMOVED FROM THE PATIENT. THE PATIENT WAS CHECKED WITH AN X-RAY AND NO OTHER METAL REACTIONS WERE OBSERVED AND SO THE PATIENT WAS CLOSED AS IS. WHEN CHECKING THE CARTRIDGE, IT APPEARED THAT THE INTERNAL PART OF THE ARTICULATED PART WAS DAMAGED.
17-01-2025 | Malfunction | us surgical puerto rico | 29-01-2025 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC LOBECTOMY, WHEN THE PULMONARY LOBE TISSUE WAS DETACHED AND CLOSED, THE STAPLES HAD POOR FORMATION AND NO GOOD ANASTOMOSIS COULD BE OBTAINED. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE.
11-12-2025 | Malfunction | covidien | 02-01-2026 | UNKNOWN ENDO GIA SULU | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | UNKNOWN ENDO GIA SULU | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | UNKNOWN ENDO GIA SULU | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Difficult to Remove | A150207 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Misfire | A050502 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
11-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A 6 CM TUMOR, THREE HANDLES AND ONE RELOAD WERE DIFFICULT TO FIRE, RESULTING IN AN EXTENDED SURGICAL TIME. THE SURGEON USED MULTIPLE CARTRIDGES BECAUSE THE DEVICE COULD NOT BE FULLY ACTIVATED. WHEN THE CARTRIDGE WAS CLOSED ON THE LUNG, THE ANVIL BOUNCED BACK. IT WAS NOTED THAT THE HANDLE WAS CRACKING. THE DEVICE WAS REPLACED TWICE TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.
05-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | CIRCULAR MECH STD 25MM | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN PROCEDURE, THE ANVIL COULD NOT BE REMOVED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
01-01-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE OPERATOR TRIED TO CONNECT THE ANVIL TO THE TROCAR WITH A GRASPER WHILE AVOIDING THE LOCKING-SPRING AREA, BUT THEY DID NOT ENGAGE PROPERLY. NO PATIENT CONSEQUENCES WERE REPORTED.
21-09-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON LINEAR CUTTER 80MM | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNSPECIFIED PROCEDURE, THE STAPLER BROKE WHEN THE SURGEON WENT TO FIRE IT. THEY USED ANOTHER STAPLER TO COMPLETE THE SURGERY. NO PATIENT CONSEQUENCES WERE REPORTED.
11-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
22-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT UPPER LOBECTOMY, WHILE DETACHING THE APICAL ANTERIOR ARTERY, THE JAWS OF THE RELOAD DID NOT CLOSE AT ALL DURING THE HANDLE TEST. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF THE PANCREAS, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
22-12-2025 | Malfunction | covidien surgical | 02-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC RIGHT UPPER LOBECTOMY, THE THE JAWS OF THE RELOAD DID NOT CLOSE AT ALL DURING THE HANDLE TEST. A NEW HANDLE AND RELOAD WERE USED TO COMPLETE THE CASE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
01-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: MALFORMED STAPLES. IT WAS REPORTED THAT DURING THE SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED ANOTHER TWO TO CONTINUE THE SURGERY, THE SAME PROBLEM HAPPENED AGAIN. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
12-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT OPEN WHEN THEY FIRED THE DEVICE ON TO THE TISSUE. IT WAS REPORTED THAT IT HAPPENED TO THE THREE DIFFERENT DEVICES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
12-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN APPENDECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT OPEN WHEN THEY FIRED THE DEVICE ON TO THE TISSUE. IT WAS REPORTED THAT IT HAPPENED TO THE THREE DIFFERENT DEVICES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM STANDARD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP GASTRECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED. CHANGED TO ANOTHER FOUR TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VATS PROCEDURE, IT WAS OBSERVED THAT THE KNIFE DID NOT RETURN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
10-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Mechanical Jam | A0506 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A VATS PROCEDURE, IT WAS OBSERVED THAT THE KNIFE DID NOT RETURN. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
13-10-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | FLEX 60 ARTICULATING | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAS AN UNSPECIFIED MALFUNCTION. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-09-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM STANDARD | Separation Problem | C070603 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE OF LARGE INTESTINE, IT WAS OBSERVED THAT THE CARTRIDGE ADHERED TO THE TISSUE AND COULD NOT BE REMOVED WHILE CUTTING THE TISSUE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
25-10-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
18-10-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | GST GOLD RELOAD 60MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
03-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
12-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | GST BLACK RELOAD 45MM 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
18-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM STANDARD | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SIGMOIDECTOMY PROCEDURE, THE STAPLER FIRED AS EXPECTED WHEN THE FIRING SEQUENCE WAS COMPLETED. HOWEVER, THE DEVICE BEGAN MAKING UNUSUAL SOUNDS WHEN THE STAPLE KNIFE WAS DEPLOYED. THERE WERE NO REPORTED PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAD AN UNSPECIFIED MALFUNCTION. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | POWERED 60 ECHELON +, 280MM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE HAD AN UNSPECIFIED MALFUNCTION. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
19-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM LONG | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAP SLEEVE GASTRECTOMY PROCEDURE, THE ECHELON 3000 WAS USED WITH ITS FIRST FIRING ON THE ANTRUM OF THE STOMACH. THE DEVICE WAS POSITIONED AND THE JAW CLOSED WITHOUT ISSUE. UPON FIRING THE MOTOR SEEMED VERY LABORED AND WAS 3 - 4 TIMES AS SLOW TO COMPLETE THE FIRING WHICH THE SURGEON USES THE PULSE TECHNIQUE WHEN FIRING. THE ECHELON COMPLETED IT'S FIRING CYCLE AND UPON RELEASE OF THE FIRING TRIGGER THE DEVICE WAS AGAIN VERY SLOW IN RETURNING THE KNIFE BLADE ASSEMBLY TO THE START/FINISH POSITION (APPROXIMATELY A MINUTE TO DO SO). THE DEVICE WAS THEN OPENED AND ARTICULATED INTO THE NEUTRAL POSITION WITHOUT ANY FURTHER ISSUES. THE THOUGHT IN THEATRE WAS THAT THIS IS A BATTERY ISSUE AND THEREFORE A NEW DEVICE WAS OPENED TO AVOID ANY FURTHER POTENTIAL ISSUES WITH A STAPLER AND BATTERY THAT HAD CAUSED PROBLEMS. THE FAULTY DEVICE WAS REMOVED AND REPLACED WITH NEW DEVICE TO COMPLETE THE PROCEDURE WITH A DELAY OF FIVE MINUTES. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
08-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ENDO ILS 25MM, CURVED | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE COLON RESECTION SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ENDO ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE COLON RESECTION SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ENDO ILS 25MM, CURVED | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE COLON RESECTION SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ENDO ILS 25MM, CURVED | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE COLON RESECTION SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
18-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | GST WHITE RELOAD, 60MM, 6 ROW | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC ASSISTED NEPHRECTOMY PROCEDURE, THE PVS LOCKED OUT ON RENAL VEIN. TRIED REVERSE 3 TIMES, MANUAL OVERRIDE AND OPENED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A ROBOTIC ASSISTED NEPHRECTOMY PROCEDURE, THE PVS LOCKED OUT ON RENAL VEIN. TRIED REVERSE 3 TIMES, MANUAL OVERRIDE AND OPENED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
05-12-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PREOPERATIVELY, THE CUSTOMER STATED THAT THE HANDLE HAD BEEN PLACED IN SEVERAL WORKING CHARGERS BUT WOULD NOT TAKE CHARGE OR TURN ON. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT FOUND THAT THE HANDLE WAS RECEIVED WITH A FULLY DEPLETED BATTERY. THE HANDLE WAS OPENED UP AND BOTH BATTERY CELLS WERE FOUND TO BE VENTED.
26-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | GST BLUE RELOAD 60MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE ARRIVED NOT FROM THE JOHNSON AND JOHNSON WAREHOUSE. THERE WAS NO PATIENT INVOLVEMENT. NO ADDITIONAL INFORMATION PROVIDED.
15-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | GST BLUE RELOAD 45MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE ARRIVED NOT FROM THE JOHNSON AND JOHNSON WAREHOUSE. THERE WAS NO PATIENT INVOLVEMENT. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | WHITE RELOAD FOR ECHELON 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE CLOSED STAPLES UNRAVELED WHEN VESSEL WAS CLOSED INTRAOPERATIVELY. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-12-2025 | Malfunction | johnson & johnson medical (suzhou | 02-01-2026 | CARTRIDGE FINAL ASSY 45MM WHITE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
15-12-2025 | Malfunction | johnson & johnson medical (suzhou | 02-01-2026 | CARTRIDGE FINAL ASSY 45MM BLUE | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES MALFORMED AFTER FIRING. CHANGED TO ANOTHER ONE TO CONTINUE THE PROCEDURE BUT THE SAME PROBLEM HAPPENED AGAIN. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
14-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 45MM COMPACT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | nan
14-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 45MM COMPACT | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
15-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | POWERED PLUS STANDARD 45 | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | nan
15-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-12-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Material Integrity Problem | A04 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE MANUAL OVERRIDE COVER WAS OPEN. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-12-2025 | Malfunction | gdw | 02-01-2026 | ETHICON 4000 STAPLER STRAIGHT 60 STANDARD | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A SLEEVE GASTRECTOMY PROCEDURE, AFTER THE FIRST FIRING IT SEEMED THAT THE SURGEON KIND OF OVERLAPPED THE PREVIOUS STAPLE LINES WITH EACH SEQUENTIAL FIRING AND THE STAPLES DID NOT PROPERLY GO INTO THE TISSUE AND WERE KIND OF HANGING OFF AT THE BEGINNING OF EACH STAPLE LINE. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
17-11-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | POWERED 60 ECHELON +, 340MM SHAFT | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE AND HAD NO MOTOR SOUND ON THE SECOND FIRING. THE SURGEON REMOVED THE BATTERY AND ROTATED FOR 180 DEGREE AND REINSERTED THE BATTERY. THE DEVICE STILL WOULD NOT FIRE AND HAD NO MOTOR SOUND. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED.
21-11-2025 | Malfunction | johnson & johnson medical (suzhou | 02-01-2026 | CARTRIDGE FINAL ASSY 45MM GREEN | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
14-08-2025 | Malfunction | johnson & johnson medical (suzhou | 02-01-2026 | CARTRIDGE FINAL ASSY 60MM BLUE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
23-08-2025 | Malfunction | johnson & johnson medical (suzhou | 02-01-2026 | CARTRIDGE FINAL ASSY 60MM WHITE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LUNG RESECTION PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
01-12-2025 | Malfunction | johnson & johnson medical (suzhou | 02-01-2026 | CARTRIDGE FINAL ASSY 45MM BLUE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
10-12-2025 | Malfunction | johnson & johnson medical (suzhou | 02-01-2026 | FINAL ASSY 445MM STD DEVICE | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE FARTHER AFTER IT FIRED A LITTLE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
22-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREATICODUODENECTOMY (WHIPPLE), IN TRANSECTION OF THE PANCREAS TAIL, THE DEVICE FAILED TO FIRE AND MADE AN ABNORMAL SOUND. TO RESOLVE THE ISSUE, A NEW HANDLE OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
22-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Noise, Audible | A0508 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN PANCREATICODUODENECTOMY (WHIPPLE), IN TRANSECTION OF THE PANCREAS TAIL, THE DEVICE FAILED TO FIRE AND MADE AN ABNORMAL SOUND. TO RESOLVE THE ISSUE, A NEW HANDLE OF THE SAME TYPE WAS USED. THERE WAS NO PATIENT INJURY.
20-10-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ECHELON STAPLER FAILED TO FIRE THROUGH BOTH SIDES OF TISSUE. THIS LEFT A HOLE IN THE ANASTOMOSIS THAT NEEDED TO BE REPAIRED. STAPLER AND RELOAD WERE REMOVED FROM FIELD. REPAIR WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-10-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | Perforation | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ECHELON STAPLER FAILED TO FIRE THROUGH BOTH SIDES OF TISSUE. THIS LEFT A HOLE IN THE ANASTOMOSIS THAT NEEDED TO BE REPAIRED. STAPLER AND RELOAD WERE REMOVED FROM FIELD. REPAIR WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-10-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM LONG | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ECHELON STAPLER FAILED TO FIRE THROUGH BOTH SIDES OF TISSUE. THIS LEFT A HOLE IN THE ANASTOMOSIS THAT NEEDED TO BE REPAIRED. STAPLER AND RELOAD WERE REMOVED FROM FIELD. REPAIR WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-10-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM LONG | No Apparent Adverse Event | A25 | Failure to Anastomose | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ECHELON STAPLER FAILED TO FIRE THROUGH BOTH SIDES OF TISSUE. THIS LEFT A HOLE IN THE ANASTOMOSIS THAT NEEDED TO BE REPAIRED. STAPLER AND RELOAD WERE REMOVED FROM FIELD. REPAIR WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-10-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM LONG | No Apparent Adverse Event | A25 | Perforation | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ECHELON STAPLER FAILED TO FIRE THROUGH BOTH SIDES OF TISSUE. THIS LEFT A HOLE IN THE ANASTOMOSIS THAT NEEDED TO BE REPAIRED. STAPLER AND RELOAD WERE REMOVED FROM FIELD. REPAIR WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
20-10-2025 | Malfunction | ethicon endo-surgery | 02-01-2026 | ECHELON 3000 60MM LONG | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE ECHELON STAPLER FAILED TO FIRE THROUGH BOTH SIDES OF TISSUE. THIS LEFT A HOLE IN THE ANASTOMOSIS THAT NEEDED TO BE REPAIRED. STAPLER AND RELOAD WERE REMOVED FROM FIELD. REPAIR WAS COMPLETED. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
17-12-2025 | Malfunction | covidien surgical | 02-01-2026 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN OPEN DISTAL GASTRECTOMY, IN DUODENAL RESECTION, AFTER STARTING THE FIRING, WHILE THE CLAMP COVER ON THE CARTRIDGE SIDE WAS ADVANCING, THE KNIFE CAME LOOSE. THE DEVICE FIRING WAS STILL COMPLETED AND NO INTERVENTION WAS PERFORMED. THERE WAS NO PATIENT INJURY.
22-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE SINGLE-PORT THORACOSCOPIC LOBECTOMY, AFTER FIRING, BOTH RELOADS HAVE POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE SURGEON REINFORCED MANUAL STITCHING TO RESOLVE THE ISSUE.
22-12-2025 | Malfunction | covidien surgical | 02-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: ACCORDING TO THE REPORTER, DURING THE SINGLE-PORT THORACOSCOPIC LOBECTOMY, AFTER FIRING, BOTH RELOADS HAVE POOR STAPLE FORMATION AND THE STAPLE LINE WAS INCOMPLETE DISTALLY. THE SURGEON REINFORCED MANUAL STITCHING TO RESOLVE THE ISSUE.
12-12-2025 | Malfunction | covidien surgical | 02-01-2026 | TRI-STAPLE 2.0 | Break | A0401 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATECTOMY, WHILE ON THE PREPARATION STAGE, THE SHIPPING WEDGE WAS FOUND TO BE DAMAGED. THE YELLOW TAB FRAGMENTS WERE REMOVED, AND THE PRODUCT WAS USED. THERE WAS NO PATIENT INJURY.
12-12-2025 | Malfunction | covidien surgical | 02-01-2026 | TRI-STAPLE 2.0 | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PANCREATECTOMY, WHILE ON THE PREPARATION STAGE, THE SHIPPING WEDGE WAS FOUND TO BE DAMAGED. THE YELLOW TAB FRAGMENTS WERE REMOVED, AND THE PRODUCT WAS USED. THERE WAS NO PATIENT INJURY.
18-12-2025 | Malfunction | covidien surgical | 02-01-2026 | TRI-STAPLE 2.0 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A THORACOSCOPIC PARTIAL PULMONARY RESECTION, PULMONARY PARENCHYMA, THE FIRING COULD NOT BE DONE AT ALL. A NEW RELOAD WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
11-12-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY VESSEL DISSECTION, WHEN ATTEMPTING TO OPEN THE JAWS AFTER DISSECTING A BLOOD VESSEL WITH THE RELOAD, THE ADAPTER SUDDENLY DETACHED. SINCE THE RELOAD COULD NOT BE REMOVED FROM THE TISSUE BEFORE THE JAWS OPENED, IT WAS RECONNECTED, AND AFTER PRESSING A BUTTON, THE JAWS OPENED AND IT COULD BE REMOVED FROM THE TISSUE. THE SCREEN SHOWED A YELLOW ADAPTER SWIMLANE. THERE WAS NO PATIENT INJURY.
11-12-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Device Difficult to Setup or Prepare | A2201 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING LOBECTOMY VESSEL DISSECTION, WHEN ATTEMPTING TO OPEN THE JAWS AFTER DISSECTING A BLOOD VESSEL WITH THE RELOAD, THE ADAPTER SUDDENLY DETACHED. SINCE THE RELOAD COULD NOT BE REMOVED FROM THE TISSUE BEFORE THE JAWS OPENED, IT WAS RECONNECTED, AND AFTER PRESSING A BUTTON, THE JAWS OPENED AND IT COULD BE REMOVED FROM THE TISSUE. THE SCREEN SHOWED A YELLOW ADAPTER SWIMLANE. THERE WAS NO PATIENT INJURY.
10-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Failure to Cut | A050702 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER CANCER RESECTION SURGERY, THE STAPLER WAS UNABLE TO CUT AND COULD NOT BE USED PROPERLY. THE STAPLER WAS REPLACED, AND THE SURGERY CONTINUED. AS A RESULT, THE OPERATION TIME WAS PROLONGED BUT NOT MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
10-12-2025 | Malfunction | covidien | 02-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LIVER CANCER RESECTION SURGERY, THE STAPLER WAS UNABLE TO CUT AND COULD NOT BE USED PROPERLY. THE STAPLER WAS REPLACED, AND THE SURGERY CONTINUED. AS A RESULT, THE OPERATION TIME WAS PROLONGED BUT NOT MORE THAN 30 MINUTES. THERE WAS NO PATIENT INJURY.
27-10-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Display or Visual Feedback Problem | A0902 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, THE POWERED STAPLER EXPERIENCED AN ISSUE INVOLVING FOUR HANDLES THAT WERE NOT WORKING. THE SPECIFIC MALFUNCTIONS INCLUDED FILE ERROR, BATTERY ERROR MESSAGES, AND THE DEVICES NOT BEING RECOGNIZED OR UNABLE TO BE CHARGED WITH EITHER THE FOUR-BAY OR SINGLE-BAY CHARGERS. THE AFFECTED HANDLES WERE REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHILE THE HANDLE WAS INSERTED INTO A CHARGER, THE BATTERY VOLTAGE DRAINED BELOW OPERABLE LEVELS OVER TIME.
27-10-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Battery Problem | A0705 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, THE POWERED STAPLER EXPERIENCED AN ISSUE INVOLVING FOUR HANDLES THAT WERE NOT WORKING. THE SPECIFIC MALFUNCTIONS INCLUDED FILE ERROR, BATTERY ERROR MESSAGES, AND THE DEVICES NOT BEING RECOGNIZED OR UNABLE TO BE CHARGED WITH EITHER THE FOUR-BAY OR SINGLE-BAY CHARGERS. THE AFFECTED HANDLES WERE REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHILE THE HANDLE WAS INSERTED INTO A CHARGER, THE BATTERY VOLTAGE DRAINED BELOW OPERABLE LEVELS OVER TIME.
27-10-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, THE POWERED STAPLER EXPERIENCED AN ISSUE INVOLVING FOUR HANDLES THAT WERE NOT WORKING. THE SPECIFIC MALFUNCTIONS INCLUDED FILE ERROR, BATTERY ERROR MESSAGES, AND THE DEVICES NOT BEING RECOGNIZED OR UNABLE TO BE CHARGED WITH EITHER THE FOUR-BAY OR SINGLE-BAY CHARGERS. THE AFFECTED HANDLES WERE REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHILE THE HANDLE WAS INSERTED INTO A CHARGER, THE BATTERY VOLTAGE DRAINED BELOW OPERABLE LEVELS OVER TIME.
27-10-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Display or Visual Feedback Problem | A0902 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, THE POWERED STAPLER EXPERIENCED AN ISSUE INVOLVING FOUR HANDLES THAT WERE NOT WORKING. THE SPECIFIC MALFUNCTIONS INCLUDED FILE ERROR, BATTERY ERROR MESSAGES, AND THE DEVICES NOT BEING RECOGNIZED OR UNABLE TO BE CHARGED WITH EITHER THE FOUR-BAY OR SINGLE-BAY CHARGERS. THE AFFECTED HANDLES WERE REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHILE THE HANDLE WAS INSERTED INTO A CHARGER, THE BATTERY VOLTAGE DRAINED BELOW OPERABLE LEVELS OVER TIME.
27-10-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Battery Problem | A0705 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, THE POWERED STAPLER EXPERIENCED AN ISSUE INVOLVING FOUR HANDLES THAT WERE NOT WORKING. THE SPECIFIC MALFUNCTIONS INCLUDED FILE ERROR, BATTERY ERROR MESSAGES, AND THE DEVICES NOT BEING RECOGNIZED OR UNABLE TO BE CHARGED WITH EITHER THE FOUR-BAY OR SINGLE-BAY CHARGERS. THE AFFECTED HANDLES WERE REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHILE THE HANDLE WAS INSERTED INTO A CHARGER, THE BATTERY VOLTAGE DRAINED BELOW OPERABLE LEVELS OVER TIME.
27-10-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Complete Loss of Power | A070801 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT PRIOR TO USE, THE POWERED STAPLER EXPERIENCED AN ISSUE INVOLVING FOUR HANDLES THAT WERE NOT WORKING. THE SPECIFIC MALFUNCTIONS INCLUDED FILE ERROR, BATTERY ERROR MESSAGES, AND THE DEVICES NOT BEING RECOGNIZED OR UNABLE TO BE CHARGED WITH EITHER THE FOUR-BAY OR SINGLE-BAY CHARGERS. THE AFFECTED HANDLES WERE REPLACED. THERE WAS NO PATIENT INVOLVED. MEDTRONIC'S INITIAL EVALUATION OF THE INCIDENT DEVICE FOUND THAT WHILE THE HANDLE WAS INSERTED INTO A CHARGER, THE BATTERY VOLTAGE DRAINED BELOW OPERABLE LEVELS OVER TIME.
09-12-2025 | Malfunction | covidien surgical | 02-01-2026 | GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, THE DEVICE DID NOT FIRE. ANOTHER STAPLER WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
17-12-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Failure to Charge | A070603 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER FIRING, EVEN WHEN THE CENTER CONTROL BUTTON IS CONTINUOUSLY PRESSED UPWARD TO RETURN THE ARTICULATED CARTRIDGE TO THE NEUTRAL POSITION, IT DID NOT FUNCTION. THE EVENT OCCURRED AFTER FIRING, AND SINCE THE PROCEDURE HAD ALREADY BEEN COMPLETED, NO MEDICAL INTERVENTION WAS PERFORMED OR REQUIRED TO RESOLVE THE EVENT. ADDITIONALLY, THE CHARGE OF THE POWERED HANDLE RAN OUT IMMEDIATELY; THE CHARGE COULD NOT BE MAINTAINED. THERE WAS NO PATIENT INJURY.
17-12-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGNIA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER FIRING, EVEN WHEN THE CENTER CONTROL BUTTON IS CONTINUOUSLY PRESSED UPWARD TO RETURN THE ARTICULATED CARTRIDGE TO THE NEUTRAL POSITION, IT DID NOT FUNCTION. THE EVENT OCCURRED AFTER FIRING, AND SINCE THE PROCEDURE HAD ALREADY BEEN COMPLETED, NO MEDICAL INTERVENTION WAS PERFORMED OR REQUIRED TO RESOLVE THE EVENT. ADDITIONALLY, THE CHARGE OF THE POWERED HANDLE RAN OUT IMMEDIATELY; THE CHARGE COULD NOT BE MAINTAINED. THERE WAS NO PATIENT INJURY.
17-12-2025 | Malfunction | covidien surgical | 02-01-2026 | SIGADAPTSTND | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER FIRING, EVEN WHEN THE CENTER CONTROL BUTTON IS CONTINUOUSLY PRESSED UPWARD TO RETURN THE ARTICULATED CARTRIDGE TO THE NEUTRAL POSITION, IT DID NOT FUNCTION. THE EVENT OCCURRED AFTER FIRING, AND SINCE THE PROCEDURE HAD ALREADY BEEN COMPLETED, NO MEDICAL INTERVENTION WAS PERFORMED OR REQUIRED TO RESOLVE THE EVENT. ADDITIONALLY, THE CHARGE OF THE POWERED HANDLE RAN OUT IMMEDIATELY; THE CHARGE COULD NOT BE MAINTAINED. THERE WAS NO PATIENT INJURY.
17-12-2025 | Malfunction | covidien surgical | 02-01-2026 | EGIA60AMT | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC LOW ANTERIOR RESECTION, AFTER FIRING, EVEN WHEN THE CENTER CONTROL BUTTON IS CONTINUOUSLY PRESSED UPWARD TO RETURN THE ARTICULATED CARTRIDGE TO THE NEUTRAL POSITION, IT DID NOT FUNCTION. THE EVENT OCCURRED AFTER FIRING, AND SINCE THE PROCEDURE HAD ALREADY BEEN COMPLETED, NO MEDICAL INTERVENTION WAS PERFORMED OR REQUIRED TO RESOLVE THE EVENT. ADDITIONALLY, THE CHARGE OF THE POWERED HANDLE RAN OUT IMMEDIATELY; THE CHARGE COULD NOT BE MAINTAINED. THERE WAS NO PATIENT INJURY.
12-12-2025 | Malfunction | covidien surgical | 02-01-2026 | UNKNOWN SIGNIA EGIA ADAPTER | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING LUNG RESECTION, THE FIRST RELOAD WAS DIFFICULT TO FIRE. THE OTHER RELOAD WAS NOTED TO HAVE A JAW THAT DID NOT COMPLETELY CLOSE. THERE WAS NO PATIENT INJURY.
12-12-2025 | Malfunction | covidien surgical | 02-01-2026 | UNKNOWN SIGNIA EGIA ADAPTER | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING LUNG RESECTION, THE FIRST RELOAD WAS DIFFICULT TO FIRE. THE OTHER RELOAD WAS NOTED TO HAVE A JAW THAT DID NOT COMPLETELY CLOSE. THERE WAS NO PATIENT INJURY.
12-12-2025 | Malfunction | covidien surgical | 03-01-2026 | SIGNIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING LUNG RESECTION, THE FIRST RELOAD WAS DIFFICULT TO FIRE. THE OTHER RELOAD WAS NOTED TO HAVE A JAW THAT DID NOT COMPLETELY CLOSE. THERE WAS NO PATIENT INJURY.
12-12-2025 | Malfunction | covidien surgical | 03-01-2026 | SIGNIA | Firing Problem | A0505 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING LUNG RESECTION, THE FIRST RELOAD WAS DIFFICULT TO FIRE. THE OTHER RELOAD WAS NOTED TO HAVE A JAW THAT DID NOT COMPLETELY CLOSE. THERE WAS NO PATIENT INJURY.
12-12-2025 | Malfunction | covidien surgical | 03-01-2026 | ENDO GIA | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING LUNG RESECTION, THE FIRST RELOAD WAS DIFFICULT TO FIRE. THE OTHER RELOAD WAS NOTED TO HAVE A JAW THAT DID NOT COMPLETELY CLOSE. THERE WAS NO PATIENT INJURY.
03-12-2025 | Malfunction | covidien | 03-01-2026 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IT WAS UNABLE TO ARTICULATE BACK TO NEUTRAL POSITION AFTER USING THE DEVICE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
03-12-2025 | Malfunction | covidien | 03-01-2026 | ENDO GIA ULTRA | Human-Device Interface Problem | A22 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, IT WAS UNABLE TO ARTICULATE BACK TO NEUTRAL POSITION AFTER USING THE DEVICE. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
05-11-2025 | Malfunction | covidien surgical | 03-01-2026 | TRI-STAPLE 2.0 | Misfire | A050502 | Insufficient Information | nan
05-11-2025 | Malfunction | covidien surgical | 03-01-2026 | TRI-STAPLE 2.0 | Insufficient Information | E2401 | Insufficient Information | nan
30-12-2025 | Malfunction | covidien | 03-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC SEGMENTECTOMY FOR LUNG SEGMENT DETACHMENT, THE HANDLE AND RELOAD WERE UNABLE TO FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE DEFECTIVE STAPLERS WERE REPLACED WITH NEW HANDLE AND RELOAD TO COMPLETE THE PROCEDURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
30-12-2025 | Malfunction | covidien surgical | 03-01-2026 | ENDO GIA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A THORACOSCOPIC SEGMENTECTOMY FOR LUNG SEGMENT DETACHMENT, THE HANDLE AND RELOAD WERE UNABLE TO FIRE, AND THE SURGEON WAS UNABLE TO SQUEEZE THE HANDLE. THE DEFECTIVE STAPLERS WERE REPLACED WITH NEW HANDLE AND RELOAD TO COMPLETE THE PROCEDURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.
21-12-2025 | Malfunction | covidien surgical | 05-01-2026 | ENDO GIA | Failure to Form Staple | A050704 | Tissue Breakdown | Event Description: IT WAS REPORTED THAT DURING A SINGLE-PORT THORACOSCOPIC LOBECTOMY, THE DEVICES EXHIBITED POOR STAPLE FORMATION AT THE DISTAL END AND AN INCOMPLETE STAPLE LINE AT THE DISTAL END. THE ISSUE WAS RESOLVED BY REINFORCED STITCHING WITH SUTURE.
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ILS 33MM CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ENDO ILS 25MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE PROCTRECTOMY PROCEDURE, COULD NOT FIRE. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ENDO ILS 29MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE RADICAL ESOPHAGECTOMY SURGERY, AFTER FIRED, NOTED THE STAPLES MALFORMED. USED SUTURE TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ENDO ILS 25MM, CURVED | Failure to Form Staple | A050704 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ENDO ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ENDO ILS 25MM, CURVED | No Apparent Adverse Event | A25 | Hemorrhage/Blood Loss/Bleeding | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ENDO ILS 25MM, CURVED | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. SUTURE WAS USED TO OVERSEW. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
08-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ENDO ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE UNK SURGERY, AFTER FIRED, NOTED SOME STAPLES MALFORMED AND OOZING OCCURRED. CHANGED THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
01-01-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE AFTER INSERTING THE BATTERY, AFTER THE ANVIL REMOVED AND DURING INSERT CDH29P FOR TO ANASTOMOSIS, DEVICE FIRED AUTOMATICALLY WITHOUT PRESSING THE TRIGGER (EVEN THOUGH THE ANVIL HAD NOT BEEN INSTALLED.) NO PATIENT CONSEQUENCES WERE REPORTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 29MM CURVED CIRCULAR, 18CM SHAFT | Self-Activation or Keying | A150104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE AFTER INSERTING THE BATTERY, AFTER THE ANVIL REMOVED AND DURING INSERT CDH29P FOR TO ANASTOMOSIS, DEVICE FIRED AUTOMATICALLY WITHOUT PRESSING THE TRIGGER (EVEN THOUGH THE ANVIL HAD NOT BEEN INSTALLED.) NO PATIENT CONSEQUENCES WERE REPORTED.
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Break | A0401 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN PROCEDURE, THE TROCAR WAS LOOSE. THE TROCAR GOT IN AND IT WAS NOT POSSIBLE TO COMBINE THEM A FEW TIMES WHEN COMBINING THE ANVIL. THE REPORTED DEVICE WAS USED AS IT WAS TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD 25MM CURVED CIRCULAR, 18CM SHAFT | Ejection Problem | A0503 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT, DURING AN UNKNOWN PROCEDURE, THE TROCAR WAS LOOSE. THE TROCAR GOT IN AND IT WAS NOT POSSIBLE TO COMBINE THEM A FEW TIMES WHEN COMBINING THE ANVIL. THE REPORTED DEVICE WAS USED AS IT WAS TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
09-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON CONTOUR W/ GREEN RELOAD | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A RECTOSIGMOIDECTOMY ONCOLOGY, USING ECHELON CONTOUR CURVED CUTTER AND THE STAPLER DID NOT STAPLE THE COLON. NO PATIENT CONSEQUENCES WERE REPORTED.
01-01-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON LINEAR CUTTER 80MM | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED, DURING AN UNKNOWN PROCEDURE, THE DEVICE COULD NOT BE FIRED. ANOTHER DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE.
23-10-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | PWRD ECH FLEX 45MM 34CM SHAFT | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
11-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON CONTOUR W/ GREEN RELOAD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNK PROCEDURE, THE DEVICE COULDN'T CLOSE. IT WAS BLOCKED. THEY DIDN'T USE IT ON A PATIENT. AFTER SEVERAL ATTEMPTS AND USE OF GREAT FORCE, THEY WERE ABLE TO CLOSE IT AND FIRE. NO PATIENT CONSEQUENCES WERE REPORTED.
15-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST BLUE RELOAD 45MM 6 ROW | No Apparent Adverse Event | A25 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT THE DEVICE ARRIVED NOT FROM THE JOHNSON AND JOHNSON WAREHOUSE. THERE WAS NO PATIENT INVOLVEMENT. NO ADDITIONAL INFORMATION PROVIDED.
21-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ILS 21MM, CURVED | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE COMBINED PANCREATICODUODENECTOMY SURGERY, COULD NOT FIRE. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
12-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ENDO ILS 25MM, CURVED | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING THE TOTAL GASTRECTOMY SURGERY, WHEN ANASTOMOSIS OF INTESTINE AND ESOPHAGUS, AFTER FIRED, NOTED THE STAPLES MALFORMED. CHANGED TO THE NEW ONE TO COMPLETE SURGERY. THERE WAS NO PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION CAN BE PROVIDED.
10-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST GOLD RELOAD 45MM 6 ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC LOBECTOMY PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
07-01-2026 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 60MM STANDARD | Device Markings/Labelling Problem | A2101 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT AN ECH45S WAS INSIDE THE PACKAGE OF AN ECH60S. THIS EVENT WAS FOUND WHEN THE PACKAGE WAS OPENED FOR USE DURING SURGERY. A REPLACEMENT PRODUCT WAS USED THE SURGERY WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION IS AVAILABLE.
17-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST WHITE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
25-11-2025 | Malfunction | covidien surgical | 05-01-2026 | ENDO GIA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SEGMENTECTOMY, THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED, AND THE SURGERY WAS COMPLETED.
25-11-2025 | Malfunction | covidien | 05-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | Insufficient Information | Event Description: ACCORDING TO THE REPORTER, DURING A SEGMENTECTOMY, THE DEVICE DID NOT FIRE. A NEW HANDLE AND RELOAD WERE USED, AND THE SURGERY WAS COMPLETED.
27-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 45 STANDARD | Positioning Failure | A150201 | No Clinical Signs, Symptoms or Conditions | nan
27-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 45 STANDARD | Component or Accessory Incompatibility | A1701 | No Clinical Signs, Symptoms or Conditions | nan
27-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 45 STANDARD | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
27-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON 3000 45 STANDARD | Activation Problem | A1501 | No Clinical Signs, Symptoms or Conditions | nan
17-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan
19-11-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | POWERED PLUS STANDARD 45 | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
15-12-2025 | Malfunction | covidien | 05-01-2026 | ENDO GIA ULTRA | Failure to Fire | A050501 | No Clinical Signs, Symptoms or Conditions | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, WHEN THE SURGEON WAS USING THE MANUAL STAPLER, THE DEVICE JAMMED, THE TRIGGER COULD NOT BE SQUEEZED, AND IT GOT LOCKED. A NEW HANDLE WAS USED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY.
01-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST BLUE RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE DEVICE WOULD NOT FIRE. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
09-12-2025 | Injury | intuitive surgical | 05-01-2026 | SUREFORM | Material Fragmentation | A040103 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED GASTRIC BYPASS ROUY EN-Y SURGICAL PROCEDURE, WHILE STAPLING, THE SURGEON SAW A BLUE PLASTIC PIECE ON THE STAPLING LINE. HE TOOK IT OUT AND SAW A BROKEN PIECE FROM THE STAPLER RELOAD. THE PROCEDURE WAS COMPLETED.
01-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | GST GOLD RELOAD, 60MM, 6-ROW | Failure to Form Staple | A050704 | No Clinical Signs, Symptoms or Conditions | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, IT WAS OBSERVED THAT THE STAPLES WERE MALFORMED AFTER FIRING. CHANGED TO A NEW ONE TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. NO ADDITIONAL INFORMATION PROVIDED.
10-12-2025 | Malfunction | ethicon endo-surgery | 05-01-2026 | POWERED PLUS STANDARD 45 | Difficult to Open or Close | A051104 | No Clinical Signs, Symptoms or Conditions | nan